## in
Flow Status | Submission Date | DATE | CUSTOMER | ADDRESS | REGION | PROVINCE | DEPARTMENT | BILLABLE | TYPE OF SERVICE | PRINCIPAL | FSE NAME | CONTACT NO. | FSE EMAIL | ADDITIONAL CUSTOMER EMAIL | IF, OTHERS: | SR NUMBER | ACCOUNT TYPE | CUSTOMER EMAIL | AVERAGE MODULE / ROOM TEMPERATURE | ASSAY STATISTICS | PARTS REFERENCE DOCUMENT DATE | MODEL | SERIAL NO | REPORTED PROBLEM/S | CORRECTIVE ACTION/S | SERVICE/S DONE | UNIT'S CONDITION AND STATUS | ADDITIONAL REMARKS | PARTS USED | PARTS REFERENCE DOCUMENT | SERVICE ENGINEER NAME | SERVICE ENGINEER SIGN | CUSTOMER NAME | CUSTOMER SIGN
| 2026-04-17 | 2026-04-17 | Tanay Rural Health Unit | Tanay, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor | 9067346538 | jmdoctor@macare-medicals.com |  |  | 32546 | PBSP | maryjunnette@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1489 ASSAY: Xpert HIV-1 Viral Load, COUNT: 52 | 2026-04-17 | GX-IV-D-6C | 807593 | Module A4 for replacement | Replaced module A4 | Performed maintenance Replaced module A4 Performed self test Generated reports | Unit is operational. Module A2/862098 not operational,for reporting. |  | Old: A4/642666 New: A4/612866 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6522213259617317353/6522213259617317353_signature_28.png | Mary Junnette Fulgueras | https://www.jotform.com/uploads/macarewecare/222927978372067/6522213259617317353/6522213259617317353_signature_31.png
| 2026-04-17 | 2026-04-17 | Region II Trauma and Medical Center | Magsaysay, Bayombong, Nueva Vizcaya | 2 | Nueva Vizcaya | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas | 9062144116 | roxasmatthew659@yahoo.com.ph | r2tmcmolec@gmail.com |  | 32545 | PBSP | r2tmcmolec@gmail.com | 25 |  | 2026-04-14 | GXIV-6C | 834621 | Due for Calibration | Performed Calibration (3 Modules Passed and Module B4/784302 Failed) | Peformed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module B4/784302 for Replacement | Xpertcheck kit(Genexpert Calibrator) | 13880 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6522207734712337142/6522207734712337142_signature_28.png | Karen O. Ramiro, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6522207734712337142/6522207734712337142_signature_31.png
| 2026-04-17 | 2026-04-17 | Binan CHO II | Binan, Laguna | 4A | Laguna | Laboratory | NO | REPAIR | CEPHEID | Harold Toco | 9362871122 | htoco@macare-medicals.com |  |  | 32544 | PBSP | montanezaira016@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 581 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1628 | 2026-04-17 | GXIV | 821278 | CPU booting issue. | Replaced Bios battery. | Replaced Bios Battery. Performed system restart. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6522144523915188602/6522144523915188602_signature_28.png | Jasmine Aira Montañez | https://www.jotform.com/uploads/macarewecare/222927978372067/6522144523915188602/6522144523915188602_signature_31.png
| 2026-04-17 | 2026-04-17 | The Medical City South Luzon | Santa Rosa City, Laguna | 4A | Laguna | Molecular Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Harold Toco | 9174795232 | htoco@macare-medicals.com |  |  | 32543 | COMMERCIAL | tmcsl-molecularlab@themedicalcity.com | 25 | ASSAY: Xpert HBV VL, COUNT: 28 ASSAY: Xpert HIV-1 Viral Load, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 249 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 2 | 2026-04-17 | GXIV | 839233 | For unit xpertcheck calibration and maintenance. | Performed xpertcheck calibration. | Performed maintennce. Performed xpertcheck calibration. | Modules B1 and B2--failed in xpertcheck calibration, recommended for replacement. Module B3-skipped and retest required, needs addional xpertcheck cartridge. Module B4- not performed, needs addional xpertcheck cartridge. |  | 1kit Xprtcheck | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6522062403918624052/6522062403918624052_signature_28.png | Ruskin Relucio | https://www.jotform.com/uploads/macarewecare/222927978372067/6522062403918624052/6522062403918624052_signature_31.png
| 2026-04-17 | 2026-04-17 | Cainta Rural Health Unit | Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor | 0992 594 5027 | jmdoctor@macare-medicals.com |  |  | 32542 | PBSP | ballesterrichard18@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 9o4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 106 |  | GX-IV-D-6C | 838182 | For swapping of modules | Swapped modules | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module 701436 with 807131 | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6522045199618207104/6522045199618207104_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/6522045199618207104/6522045199618207104_signature_31.png
| 2026-04-17 | 2026-04-17 | Cainta Rural Health Unit | Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor | 0992 594 5027 | jmdoctor@macare-medicals.com |  |  | 32541 | DOH | ballesterrichard18@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 109 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1146 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 |  | GX-IV-D-6C | 831624 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Monitor is shutting down frequently. Monitor SN:CNK832243R |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6522044439613332287/6522044439613332287_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/6522044439613332287/6522044439613332287_signature_31.png
| 2026-04-17 | 2026-04-17 | Cainta Rural Health Unit | Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor | 0992 594 5027 | jmdoctor@macare-medicals.com |  |  | 32540 | PBSP | ballesterrichard18@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 953 |  | GX-IV-D-6C | 804687 | Module B2 has error 1002 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B2/888288 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6522043939611817209/6522043939611817209_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/6522043939611817209/6522043939611817209_signature_31.png
| 2026-04-17 | 2026-04-17 | Apalit Rural Health Unit | Municipal Grounds Apalit Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon | 9452482577 | mark10venzon@gmail.com |  |  | 32539 | DOH | duenasblesilda@gmail.com | 25 |  |  | GXIV-D 6C | 831900 | B2/758106 and B4/674002 for module replacement | Preventive maintenance Module replacement | Performed preventive maintenance Performed module Gathered reports Performed functionality | Unit is operational B2/734293 B4/730935 Newly replaced modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6522012076415815557/6522012076415815557_signature_28.png | Blesilda Duenñas | https://www.jotform.com/uploads/macarewecare/222927978372067/6522012076415815557/6522012076415815557_signature_31.png
| 2026-04-17 | 2026-04-14 | Cardinal Santos Medical Center | Wilson St, Greenhills, San Juan | NCR | NCR 2nd - San Juan | Microbiology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor | 639566106402 | jmdoctor@macare-medicals.com |  |  | 32538 | COMMERCIAL | renalyn.cabaltea@medilinx.com.ph | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1241 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules B3 and B4 not operational. | Recommended to replace modules B1-B4 due to errors and failed in XpertCheck last October 2025. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6522010709618569374/6522010709618569374_signature_28.png | Renalyn Cabaltea | https://www.jotform.com/uploads/macarewecare/222927978372067/6522010709618569374/6522010709618569374_signature_31.png
| 2026-04-16 | 2026-04-16 | Calamba Medical Center | Calamba, Laguna | 4A | Laguna | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Harold Toco | 9096195573 | htoco@macare-medicals.com |  |  | 32537 | COMMERCIAL | aivee.alcantara@medilinx.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 729 | 2026-04-16 | GXIV-2 Modules | 812916 | For unit maintenance and Xpertcheck Calibration | Performed preventive maintenance. Performed xpertcheck calibration. | Performed unit preventive maintenance and xpertcheck calibration. | All 2 modules Passed. |  | 1 kit Xpertcheck Calibrator | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521357553911242378/6521357553911242378_signature_28.png | Aivee Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/6521357553911242378/6521357553911242378_signature_31.png
| 2026-04-16 | 2026-04-16 | San Rafael Rural Health Unit | San Rafael | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar | 9778118287 | rcb.macare@gmail.com |  |  | 32536 | DOH | tatsdayrit25@gmail.com | 27 |  | 2026-04-13 | GXIV | 831692 | Module B1(793141) for replacement | Replaced module B1 | - replaced module B1 by 677963 - generated reports - performed self test | Unit is operational |  | GeneXpert Module | G 8402 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521301006991590594/6521301006991590594_signature_28.png | Ma. Suzie Villaraga | https://www.jotform.com/uploads/macarewecare/222927978372067/6521301006991590594/6521301006991590594_signature_31.png
| 2026-04-16 | 2026-04-16 | Batangas City Health Office | Batangas City | 4A | Batangas | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Harold Toco | 9084819818 | htoco@macare-medicals.com |  |  | 32535 | PBSP | lumanglasmariejane@gmsil.com | 25 |  | 2026-04-16 | GXIV | 821317 | For check up due to Error 5007. Probe check error. | Performed maintenance. Clean ICORE lense. | Performed maitenance. Clean ICORE lense. | Module A2 with HW error. For reporting. Midules A1,A3,A4 for actual sample result monitoring. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521267073916206442/6521267073916206442_signature_28.png | Marie Jsne Lumanglas | https://www.jotform.com/uploads/macarewecare/222927978372067/6521267073916206442/6521267073916206442_signature_31.png
| 2026-04-16 | 2026-04-16 | Lupon Rural Health Unit | Poblacion, Lupon Davao Oriental | 11 | Davao Oriental | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE Pull out Modules (Defective and Buffer) | CEPHEID | Roi Kenneth Fagutao | 9381973374 | rkfagutao.macare@gmail.com |  |  | 32534 | PBSP | karentemplado16@gmail.com | 29.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 303 |  | GX-IV 6C 4D | 815496 | For Preventive Maintenance For pullout GeneXpert Module | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Pull out the defective Module SN:694691 And Buffer Working Module SN:210060725 | Perform Preventive Maintenance Perform pull out Modules | All Modules are operational Machine is Operational High Room Temperature, enduser was advised to achieve ideal room temperature before operating the machine | Pull out modules SN 694691 / Defective and SN 210060725 / working buffer |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521245093313758800/6521245093313758800_signature_28.png | Josephine C. Millan,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6521245093313758800/6521245093313758800_signature_31.png
| 2026-04-16 | 2026-04-16 | Olongapo City Health Office | West Bajac-bajac Olongapo City | 3 | Zambales | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon | 639177196172 | mark10venzon@gmail.com |  |  | 32533 | PBSP | olongapochotmlandrtdl@gmail.com | 25 |  |  | GXIV-D 6C | 804700 | D2/617387 for replacement | Preventive maintenance Module replacement | Performed preventive maintenance Performed module replacement Gathered Performed functionality test | Unit is operational D2/695675 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521199518614426717/6521199518614426717_signature_28.png | Ann Malinay | https://www.jotform.com/uploads/macarewecare/222927978372067/6521199518614426717/6521199518614426717_signature_31.png
| 2026-04-16 | 2026-04-16 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar | 9205020505 | rcb.macare@gmail.com |  |  | 32532 | PBSP | mandaluyongcityhealthlab@gmail.com | 26 |  | 2026-04-13 | GXIV | 805062 | Module A1 (210101730) for replacement | Replaced module A1 | - replaced module A1 by 642313 - checked and cleaned cartridge bay, plunger rod, filter - generated reports | Unit is operational |  | GeneXpert Module | G 8400 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521188646993755842/6521188646993755842_signature_28.png | Rosemarie Guerrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6521188646993755842/6521188646993755842_signature_31.png
| 2026-04-16 | 2026-04-16 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar | 9205020505 | rcb.macare@gmail.com |  |  | 32531 | PBSP | mandaluyongcityhealthlab@gmail.com | 27 |  |  | GXIV | 831596 | Modules C1(619782) and C3(765956) for replacement | Replaced modules | - replaced modules by 210034563 and 210033908 - checked and cleaned cartridge bay, plunger rod, filter - generated reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521188206997568889/6521188206997568889_signature_28.png | Rosemarie Guerrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6521188206997568889/6521188206997568889_signature_31.png
| 2026-04-16 | 2026-04-16 | Taysan Super Health Center | Taysan, Batangas | 4A | Batangas | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Harold Toco | 9493994407 | htoco@macare-medicals.com |  |  | 32530 | PBSP | kiajaikaatienza0928@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 | 2026-04-16 | GXIV | 838890 | For Xpertcheck calibration and maintenance. | Performed xpertcheck calibration. | Performed maintenance and cleaning. Performed xprtcheck calibration | 2 modules Failed. 1 module skipped and retest required. 1 module not performed due to insufficient cartridge. |  | Xpettcheck kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521165550509128712/6521165550509128712_signature_28.png | KIA JAIKA A. PEREZ | https://www.jotform.com/uploads/macarewecare/222927978372067/6521165550509128712/6521165550509128712_signature_31.png
| 2026-04-16 | 2026-04-16 | Tinambacan Main Health Center | Tinambacan, Calbayog city | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap | 9175128474 | johnpaullabutap@gmail.com |  |  | 32529 | PBSP | porlaressteph@gmail.com | 27 |  |  | GXIV 6c 2SITES | 834824 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521136716122418436/6521136716122418436_signature_28.png | stephanie jis rosalez | https://www.jotform.com/uploads/macarewecare/222927978372067/6521136716122418436/6521136716122418436_signature_31.png
| 2026-04-16 | 2026-04-16 | Bindoy District Hospital | Bindoy, Negros Oriental | 7 | Negros Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga | 9338620615 | gselga@macare-medicals.com |  |  | 32528 | DOH | kristinemorales032118@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 678 |  | GX4 6C R2 | 838475 | CMOS battery for replacement, causing date-time result issues. | Preventive Maintenance - cleaned optics, module bays, filter, plunger rods; archive tests | Replaced CMOS battery. Configure date and time settings. Generate reports. | All 4 modules in working condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521135691017813351/6521135691017813351_signature_28.png | Kristine Katrina M. Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/6521135691017813351/6521135691017813351_signature_31.png
| 2026-04-16 | 2026-04-16 | Abucay Rural Health Unit( Roving) | Dacena Ave. Abucay Bataan | 3 | Bataan | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon | 639687252456 | mark10venzon@gmail.com |  |  | 32527 | PBSP | parreraledelina22176@gmail.com | 25 |  |  | GXIV-D 6C | 807466 | A3/701939 for module replacement | Preventive maintenance Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational A3/729668 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521125078619394686/6521125078619394686_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/6521125078619394686/6521125078619394686_signature_31.png
| 2026-04-16 | 2026-04-16 | St. Louis Molecular Diagnostic Laboratoy | Tacurong City, Sultan Kudarat | 12 | Sultan Kudarat | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas | 9217166490 | jessontupas.macaremed@gmail.com |  |  | 32526 | COMMERCIAL | stlouismoleculardiagnosticslab@gmail.com | 20.1 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 38 |  | GxlV-10 color | 110024270 | None | Perform genexpert machine maintenance and cleaning Module cleaning Plunger rod and optics cleaning Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational modulea |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6521123522907422720/6521123522907422720_signature_28.png | Johana Shane De Pedro | https://www.jotform.com/uploads/macarewecare/222927978372067/6521123522907422720/6521123522907422720_signature_31.png
| 2026-04-15 | 2026-04-15 | Velasquez Health Center | Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar | 9169671831 | rcb.macare@gmail.com |  |  | 32525 | DOH | wcpatron11x4@gmail.com | 38 |  |  | GXIV | 831534 | Module A2 (616338) for replacement | Replaced module A2 Relocated unit | - replaced module by 908646 - checked and cleaned cartrige bay, plunger rod, filter - generated reports - performed self test - relocated unit | Unit is operational Temporarily relocated unit for appropriate air condition (Current module temperatures exceed 30 degrees) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6520431905816436764/6520431905816436764_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/6520431905816436764/6520431905816436764_signature_31.png
| 2026-04-15 | 2026-04-15 | San Pablo Social Hygiene Clinic | San Pablo City, Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION | CEPHEID | Harold Toco | 9462488973 | htoco@macare-medicals.com |  |  | 32524 | MACARE-HBDC | chrisnoeaninias@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5 | 2026-04-15 | GXXVI | 120007829 | For unit installation. | Performed unit installatiion. | Performed unit installation. Performed functionality test. | Unit is operational | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6520386296511262287/6520386296511262287_signature_28.png | Chris Noe C. Aninias | https://www.jotform.com/uploads/macarewecare/222927978372067/6520386296511262287/6520386296511262287_signature_31.png
| 2026-04-15 | 2026-04-15 | Canossa Health Center | Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar | 9292140735 | rcb.macare@gmail.com |  |  | 32523 | DOH | acapuli.pbsp@gmail.com | 30 |  |  | GXIV | 812561 | Module A3(210068328) for replacement | Replaced module A3 | - replaced module A3 by (625517) - checked and cleaned cartridge bay, plunger rod, filter - generated reports - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6520364255013447140/6520364255013447140_signature_28.png | Alexis Rica Claire Apuli - Ang | https://www.jotform.com/uploads/macarewecare/222927978372067/6520364255013447140/6520364255013447140_signature_31.png
| 2026-04-15 | 2026-04-15 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | Wynnona Reoma | 9489831726 | wbreoma.2020@gmail.com |  |  | 32522 | DOH | sajecopooh20@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1390 |  | GXIV-4D 6C | 838555 | None | None | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6520236489816310440/6520236489816310440_signature_28.png | Sarah Jean G. Pasana | https://www.jotform.com/uploads/macarewecare/222927978372067/6520236489816310440/6520236489816310440_signature_31.png
| 2026-04-14 | 2026-04-14 | SLH-NRL SACCL | Quiricada St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | National Reference Laboratory | YES | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Julius Marion Biteng | 9213604482 | techsupport@macare-medicals.com |  |  | 32521 | DOH | sacclkitevaluation@slh.doh.gov.ph | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 | 2026-04-14 | GX4 | 821257 | Loss Communication Error | Reconnected Lan Cable GX and Laptop. Informed End user to check and reconnect the Lan cable if loss communication error occurs. Informed end user to follow the sequence when turning on the machine. | Checked Lan Cable Connection Checked IP Address (IP address is already set) | Machine is Operational | Observe Module D3 Error 4014 encountered and UPS Fuse for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6519631186223170401/6519631186223170401_signature_28.png | Janine Navarro | https://www.jotform.com/uploads/macarewecare/222927978372067/6519631186223170401/6519631186223170401_signature_31.png
| 2026-04-14 | 2026-04-14 | Dr. Jose N. Rodriguez Memorial Hospital | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar | 9560792440 | rcb.macare@gmail.com |  |  | 32520 | PBSP | molecular.lab@djnrmhs.doh.gov.ph | 26 |  |  | GXIV-10C | 110020277 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Unit is due for calibration |  |  |  |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/6519589767111467157/6519589767111467157_signature_28.png | Noreen Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/6519591177114300050/6519591177114300050_signature_31.png
| 2026-04-14 | 2026-04-14 | Dr. Jose N. Rodriguez Memorial Hospital | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar | 9560792440 | rcb.macare@gmail.com |  |  | 32519 | PBSP | molecular.lab@djnrmhs.doh.gov.ph | 26 |  |  | GXIV | 834431 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Unit is due for calibration |  |  |  |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/6519589767111467157/6519589767111467157_signature_28.png | Noreen Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/6519590677119292426/6519590677119292426_signature_31.png
| 2026-04-14 | 2026-04-14 | Dr. Jose N. Rodriguez Memorial Hospital | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar | 9560792440 | rcb.macare@gmail.com |  |  | 32518 | PBSP | molecular.lab@djnrmhs.doh.gov.ph | 26 |  |  | GXIV | 814159 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Unit is due for calibration |  |  |  |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/6519589767111467157/6519589767111467157_signature_28.png | Noreen Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/6519590227114558362/6519590227114558362_signature_31.png
| 2026-04-14 | 2026-04-14 | Dr. Jose N. Rodriguez Memorial Hospital | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar | 9560792440 | rcb.macare@gmail.com |  |  | 32517 | PBSP | molecular.lab@djnrmhs.doh.gov.ph | 26 |  |  | GXIV | 834432 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Unit is due for calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6519589767111467157/6519589767111467157_signature_28.png | Noreen Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/6519589767111467157/6519589767111467157_signature_31.png
| 2026-04-14 | 2026-04-14 | Cardinal Santos Medical Center | Wilson St., Greenhills, San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor | 9088963079 | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 32516 | COMMERCIAL | csmclab@csmc.ph | 22 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Run simulation Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6519547498984191556/6519547498984191556_signature_28.png | Angel Lynn Grace Pagulayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6519547498984191556/6519547498984191556_signature_31.png
| 2026-04-14 | 2026-04-14 | Olongapo City Health Office | West Bajac--Bajac Olongapo City | 3 | Zambales | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon | 639177196172 | mark10venzon@gmail.com |  |  | 32515 | PBSP | olongapochotmlandrtdl@gmail.com | 25 |  |  | GXIV-D 10c | 120000606 | B1/220013803 for module eplacement | Preventive maintenance Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational B1/668577 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6519531854017288523/6519531854017288523_signature_28.png | Roselyn Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/6519531854017288523/6519531854017288523_signature_31.png
| 2026-04-14 | 2026-04-14 | DOST - Industrial Technology Development Institute | DOST Compound, General Santos Avenue, Brgy. Central Bicutan, Taguig City | NCR | NCR 4th - Taguig | Industrial Technology Development Institute | NO | INSTALLATION | OTHERS | Carlos Leniel Origenes | 9760543099 | carlos.leniel.origenes@gmail.com |  | Miltenyi Biotec | 32514 | COMMERCIAL | jbpadaon.dost@gmail.com | 24 |  |  | gentleMACS Octo Dissociator with Heaters | 472 | N/A | None | Installed instrument based on gentleMACS Octo Dissociator with Heaters User Manual Original Instructions, Issued February 2024, 140-006-564-004.04 Performed functionality test | Instrument is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6519465106958831844/6519465106958831844_signature_28.png | DOST-ITDI | https://www.jotform.com/uploads/macarewecare/222927978372067/6519465106958831844/6519465106958831844_signature_31.png
| 2026-04-14 | 2026-04-14 | Maysan 3S Health Station | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar | 9186064241 | rcb.macare@gmail.com |  |  | 32513 | PBSP | neldimatera@gmail.com | 27 |  |  | GXIV 10C | 120000727 | Module B3 for replacement | Replaced module B3 | - replaced B3 (220013986) by (635735) - generated reports - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6519446751113454881/6519446751113454881_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/6519446751113454881/6519446751113454881_signature_31.png
| 2026-04-13 | 2026-04-13 | Chong Hua Hospital-Mandaue | Mandaue, Cebu City | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito | 9688590770 | daymon899@gmail.com |  |  | 32512 | COMMERCIAL | irl.maderazo@gmail.com | 26 |  |  | Mcs+ | 11J217 | Due Date for Apheresis Calibration | Correcting the out of range calibration datas | Preventive Maintenance Calibration of Apheresis Machine | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6518710870746668853/6518710870746668853_signature_28.png | Inna Rose Maderazo | https://www.jotform.com/uploads/macarewecare/222927978372067/6518710870746668853/6518710870746668853_signature_31.png
| 2026-04-13 | 2026-04-10 | QC Jail | Payatas rd QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca | . | nfloresca@macare-medicals.com |  |  | 32511 | PBSP | aubreydurian@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2159 |  | GX4D 6C | 803089 | For module replacement A3 661626 - 220043288 | Performed module replacement ‎Cleaned unit ‎Cleaned plunger rod,filter,icore,cartridge bay ‎Gathered reports ‎Test run unit | Performed module replacement | A1 disabled reported already A2, A3, A4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6518542483611375231/6518542483611375231_signature_28.png | Aubrey D Singh | https://www.jotform.com/uploads/macarewecare/222927978372067/6518542483611375231/6518542483611375231_signature_31.png
| 2026-04-13 | 2026-04-08 | Masambong Health Center | Masambong QC | NCR | NCR 2nd - Quezon City | Labortory | NO | REPAIR | CEPHEID | Noel Floresca | . | nfloresca@macare-medicals.com |  |  | 32510 | PBSP | michellebonrostro@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1422 ASSAY: Xpert MTB-XDR, COUNT: 134 |  | GX4 10 C | 120000173 | For module replacement B1 220011408-644800, B3 220011590-652195, B4 220011633-667004 | Performed module replacement ‎Cleaned unit ‎Cleaned plunger rod,filter,icore,cartridge bay ‎Gathered reports ‎Test run unit | Performed module replacement | B2 dusabled due to temp error 2014 already reported thru jotform B1, B3, B4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6518541513618740021/6518541513618740021_signature_28.png | Michelle Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/6518541513618740021/6518541513618740021_signature_31.png
| 2026-04-13 | 2026-04-08 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Blood donor room | NO | MAINTENANCE | HAEMONETICS | Noel Floresca | . | nfloresca@macare-medicals.com |  |  | 32509 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 26 |  |  | MCS + | 15H029G | Due for maintenance Urethane rollers yellowish need to be replace | Cleaned unit ‎Checked program parameters ‎Checked and adjust line sensor value ‎Checked rollers,dpm,spm,centrifuge,weigher ‎Performed functionality test ‎Performed simulation running | Performed maintenance | Oerational Urethane rollers for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6518540663619407340/6518540663619407340_signature_28.png | PBC | https://www.jotform.com/uploads/macarewecare/222927978372067/6518540663619407340/6518540663619407340_signature_31.png
| 2026-04-13 | 2026-04-08 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Blood donor room | NO | MAINTENANCE | HAEMONETICS | Noel Floresca | . | nfloresca@macare-medicals.com |  |  | 32508 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 26 |  |  | MCS+ | 13C269 | Due for maintenance | Cleaned unit ‎Checked program parameters ‎Checked and adjust line sensor value ‎Checked rollers,dpm,spm,centrifuge,weigher ‎Performed functionality test ‎Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6518540083615412873/6518540083615412873_signature_28.png | PBC | https://www.jotform.com/uploads/macarewecare/222927978372067/6518540083615412873/6518540083615412873_signature_31.png
| 2026-04-13 | 2026-04-08 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Blood donor room | NO | MAINTENANCE | HAEMONETICS | Noel Floresca | . | nfloresca@macare-medicals.com |  |  | 32507 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 26 |  |  | MCS + | 16L505SPG | Due for maintenance | Cleaned unit ‎Checked program parameters ‎Checked and adjust line sensor value ‎Checked rollers,dpm,spm,centrifuge,weigher ‎Performed functionality test ‎Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6518539423614648715/6518539423614648715_signature_28.png | PBC | https://www.jotform.com/uploads/macarewecare/222927978372067/6518539423614648715/6518539423614648715_signature_31.png
| 2026-04-13 | 2026-04-08 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Blood donor room | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Noel Floresca | . | nfloresca@macare-medicals.com |  |  | 32506 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 27 |  |  | MCS+ | 15H030G | Due for maintenance Missing screw lock of weigher Urethane rollers yellowish color need to be replace | Cleaned unit ‎Checked program parameters ‎Checked and adjust line sensor value ‎Checked rollers,dpm,spm,centrifuge,weigher Replaced screw lock ‎Performed functionality test ‎Performed simulation running | Performed maintenance | Operational Urethane rollers need to replace |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6518531893619144368/6518531893619144368_signature_28.png | PBC | https://www.jotform.com/uploads/macarewecare/222927978372067/6518531893619144368/6518531893619144368_signature_31.png
| 2026-04-13 | 2026-04-08 | Mandaue City Health Office | Mandaue City, Cebu | 7 | Cebu | Laboratoy | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32505 | PBSP | mchoclinical.laboratory@gmail.com | 25 |  |  | Gx4 | 815275 | Error code 2124 and 2120: Module 4 communication lost | Reset Lan of module 4 Reset gatewayboard | Reset Lan of module 4 Reset gatewayboard | Under observation module 4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6518379537618859977/6518379537618859977_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/6518379537618859977/6518379537618859977_signature_31.png
| 2026-04-10 | 2026-04-10 | Notre Dame De Chartres Hospital | Baguio City | CAR | Benguet | Blood Bank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 32504 | COMMERCIAL | jarrenmadiam1@gmail.com | 25 |  | 2026-04-07 | MCS + | 10M103 | For unit calibration and maintenance. | Unit calibration and naintenance. | Performed unit calibration. Replaced pump rotor roller. Peformed maintenance. | Unit is operational. |  | Pump rotor roller | SI 102474 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6516130791526520993/6516130791526520993_signature_28.png | Jarren Madiam | https://www.jotform.com/uploads/macarewecare/222927978372067/6516130791526520993/6516130791526520993_signature_31.png
| 2026-04-10 | 2026-04-10 | Calaanan Health Center | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32503 | DOH | rosaleslorina0917@gmail.com | 28 |  |  | GXIV | 831892 | GX Workstation for troubleshooting | Replaced BIOS battery Changed boot configuration | #NAME? | Unit is not operational Workstation unable to boot For reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6516044948652580191/6516044948652580191_signature_28.png | Jusmie Yomari Mar | https://www.jotform.com/uploads/macarewecare/222927978372067/6516044948652580191/6516044948652580191_signature_31.png
| 2026-04-10 | 2026-04-10 | San Jose Del Monte City Health Center XII | San Jose Del Monte | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32502 | PBSP | krizetteamorio@gmail.com | 25 |  |  | GX XVI | 844111 | Unit for maintenance | Cleaned filters | #NAME? | Unit is operational A3, B1, C3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6515980338653019394/6515980338653019394_signature_28.png | Krizette Marie Amorio | https://www.jotform.com/uploads/macarewecare/222927978372067/6515980338653019394/6515980338653019394_signature_31.png
| 2026-04-10 | 2026-04-08 | Marikina City Health Office | Shoe Ave Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32501 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 ASSAY: Xpert MTB-XDR, COUNT: 25 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892501634367055/6515892501634367055_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892501634367055/6515892501634367055_signature_31.png
| 2026-04-10 | 2026-04-08 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32497 | DOH | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1363 |  | GX-IV-D-6C | 831588 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451631975372/6515892451631975372_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451631975372/6515892451631975372_signature_31.png
| 2026-04-10 | 2026-04-08 | Taguig Pateros District Hospital | C5 Service Road, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32496 | DOH | ierlanegail@gmail.com | 24 |  | 2026-04-08 | GX-IV-D-6C | 831638 | Module A3 for replacement | Replaced module A3 | Performed maintenance Replaced module A3 Performed Self test Generated reports | Unit is operational. All modules operational. |  | Old: A3/694973 New: A3/908766 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451632065955/6515892451632065955_signature_28.png | Erlanegail Isla | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451632065955/6515892451632065955_signature_31.png
| 2026-04-10 | 2026-04-08 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32500 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 748 ASSAY: Xpert MTB-XDR, COUNT: 3 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451636816182/6515892451636816182_signature_28.png | Estrelita 6 | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451636816182/6515892451636816182_signature_31.png
| 2026-04-10 | 2026-04-08 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32498 | PBSP | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1363 | 2026-04-08 | GX-IV-D-6C | 805211 | Module A2 for replacement | Replaced module A2 | Performed maintenance Replaced module A2 Performed Self test Generated reports | Unit is operational. All modules operational. |  | Old: A2/646295 New:A2/210089348 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451637759431/6515892451637759431_signature_28.png | Esty Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451637759431/6515892451637759431_signature_31.png
| 2026-04-10 | 2026-04-08 | Marikina City Health Office | Shoe Ave, Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32499 | PBSP | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 723 |  | GX-IV-D-10C | 110020361 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451638645202/6515892451638645202_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6515892451638645202/6515892451638645202_signature_31.png
| 2026-04-08 | 2026-04-08 | Old Balara RTDL | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.nacare@gmail.com |  |  | 32495 | PBSP | obooogs@gmail.com | 27 |  |  | GXIV | 831719 | UPS for installation | Installed UPS | #NAME? | Unit is operational UPS under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514444121115677537/6514444121115677537_signature_28.png | Rushelle Obogne | https://www.jotform.com/uploads/macarewecare/222927978372067/6514444121115677537/6514444121115677537_signature_31.png
| 2026-04-08 | 2026-04-08 | Silliman University Medical Center | Aldecoa str. Dumaguete City Neg.Oriental | 7 | Negros Oriental | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32494 | COMMERCIAL | sumcmoleclab@gmail.com | 26 |  | 2026-04-08 | GXIV 6C | 821471 | Due for Calibration. | . Save File for XpertCal. . Save Installation Qualifications Report. . Save Assay Statistics Report | . Perform Xpertcheck for module A1,A3,A4 . A2 deffective | . Module A1,A3 passed . Module A4 calibration fail. | Hot room area for GeneXpert need for additional aircon due to temperature error. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514393253709015707/6514393253709015707_signature_28.png | Maryjune bangalando | https://www.jotform.com/uploads/macarewecare/222927978372067/6514393253709015707/6514393253709015707_signature_31.png
| 2026-04-08 | 2026-04-08 | Cebu South Medical Center | San Isidro, Talisay City Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | csmcmicro@gmail.com |  | 32493 | DOH | csmcgenexpert@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1623 |  | GX4 | 838559 | - For assessment of the GeneXpert machine as requirements for purchase of warranty over the machine and/ or modules. - Module 3 is available during the assessment and noted in the System log encountered error 4009 during interim period from November 2025 until April 3, 2026. - This error occurrence started when it was transfer from other room due additional machine in the laboratory. | - Inform end-user to transfer the machine as soon as possible to avoid overshooting of temperature in the Module 3 during the test procedure. - Isolate the GeneXpert machine from any heating device or heat source area. - Report the data to immediate supervisor that NCR was already inform the error 4009 several times but no personnel was immediately dispatched. - Module 3 required to be replaced after several error occurrence. | #NAME? | Unit modules are available during the self-test except Module 3 due to Error 4009. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514385350128697784/6514385350128697784_signature_28.png | Cebu South Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6514385350128697784/6514385350128697784_signature_31.png
| 2026-04-08 | 2026-04-08 | Wenceslao Dela Paz Health Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32492 | PBSP | ramosfely38@gmail.com | 34 |  |  | GXIV | 804688 | GX unit for switch | Installed GX | #NAME? | Unit is operational Room ambient temperature exceeds 30 deg C Room air conditioner is recommended for repair/replacement to prevent module damage |  |  |  |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/6508182296418282170/6508182296418282170_signature_28.png | Felicidad De Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/6514366695316677158/6514366695316677158_signature_31.png
| 2026-04-08 | 2026-04-08 | Salcedo Rural Health Unit | Salcedo Eastern Samar | 8 | Eastern Samar | TB laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | Johnpaullabutap@gmail.com |  |  | 32491 | PBSP | angelinegarcia_23@yahoo.com | 28 |  | 2026-04-08 | GXIV - 2 Sites 6c | 835061 | n/a | n/a | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514350051715314452/6514350051715314452_signature_28.png | Angeline G. Bandoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6514350051715314452/6514350051715314452_signature_31.png
| 2026-04-08 | 2026-04-08 | Mariveles Rural Health Unit | Mariveles Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32490 | DOH | angelicaroseconos@gmail.com | 25 |  |  | GXIV-D 6C | 831894 | B2/761029 for replacement | Preventive maintenance Module replacement | Performed preventive maintenance Replaced module B2 Gathered reports Performed functionality test | Unit is operational B2:210089345 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514334659207133795/6514334659207133795_signature_28.png | Angelica Rose Conos | https://www.jotform.com/uploads/macarewecare/222927978372067/6514334659207133795/6514334659207133795_signature_31.png
| 2026-04-08 | 2026-04-06 | Biopath Garden Heights | E. RodriguezQC | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | THERMO FISHER SCIENTIFIC | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32489 | COMMERCIAL | debuyanterahgrace@gmail.com | 35 |  |  | 88400D | 160414501161220 | Encounter high temp unable ro reach desire temp -80 High ambient temp | Checked temp setting Adjust to -50 deg. Checked amperage low side and high side of compressor Checked leakage | Perforedm troubleshooting | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514269436815270249/6514269436815270249_signature_28.png | Terah Debuyan | https://www.jotform.com/uploads/macarewecare/222927978372067/6514269436815270249/6514269436815270249_signature_31.png
| 2026-04-08 | 2026-04-07 | National Children's Hospital | E. Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed@nch.doh.gov.ph |  | 32488 | PBSP | ajunio16@yahoo.com | 26 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1749 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 15 | 2026-04-07 | GX4 6C | 803109 | For module replacement B3 607894 (def) - 650075 (NEW) | Performed module replacement Cleaned unit ‎Cleaned plunger rod,filter,icore,cartridge bay ‎Gathered reports ‎Test run unit | Performed module replacement | B1, B2, B3 operational B4 for bservation enable module |  | 900-0491Rs | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514255228714666183/6514255228714666183_signature_28.png | Armando V. Junio Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6514255228714666183/6514255228714666183_signature_31.png
| 2026-04-08 | 2026-04-07 | Philippine Tuberculosis Society Inc. | E. Rodriguez QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32487 | MACARE-HBDC | sannymisorena@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 360 ASSAY: Xpert MTB-XDR, COUNT: 42 |  | GX 4 10C | 110027248 | As per wnd user they encpunter 2005 error in module b3 | Cleaned unit ‎Cleaned plunger rod,filter,icore,cartridge bay ‎Gathered reports ‎Test run unit | Performed maintenance and troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514254878713476139/6514254878713476139_signature_28.png | Sanny Isorena | https://www.jotform.com/uploads/macarewecare/222927978372067/6514254878713476139/6514254878713476139_signature_31.png
| 2026-04-08 | 2026-04-08 | Limay Rural Health Unit | Pagcor Bldg. Limay Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32486 | PBSP | nelsonchavez40771@gmail.com | 25 |  |  | GXIV-D 6C | 805124 | C1/687261 C2/210021632 C3/210100613 C4/609681 For replacement | Preventive maintenance Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational C1/220003502 C2/210089363 C3/210101852 C4/210112370 Newly replaced module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514233552533198876/6514233552533198876_signature_28.png | Nelson Chavez | https://www.jotform.com/uploads/macarewecare/222927978372067/6514233552533198876/6514233552533198876_signature_31.png
| 2026-04-08 | 2026-04-08 | Samal Super Health Center | Brgy. Tabing Ilog, Samal Bataan | 3 | Bataan | Labor | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32485 | DOH | samalrhulab@gmail.com | 28 |  |  | GXIV-D 6C | 831729 | A4: 784697 for replacement | Preventive maintenance Module replacement | Performed preventive maintenance Replaced module A4 Gathered reports Performed functionality test | Unit is operational A4: 220001481 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6514168502532672338/6514168502532672338_signature_28.png | Maria Del Mar Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/6514168502532672338/6514168502532672338_signature_31.png
| 2026-04-07 | 2026-04-07 | Kalinga Health Hub | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32484 | CONSORTIUM | kalingahealth.manda@innovationsch.org | 27 |  | 2026-04-07 | GXIV | 110008420 | Module B1 (220000585) for replacement | Replaced module B1 | - replaced module by 628967 - checked and cleaned cartridge bay, plunger rod, filter - performed self test - generated reports | Unit is operational |  | GeneXpert Module SN: 628967 | 102479 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6513590423117505920/6513590423117505920_signature_28.png | Chin Lyn Corpin | https://www.jotform.com/uploads/macarewecare/222927978372067/6513590423117505920/6513590423117505920_signature_31.png
| 2026-04-07 | 2026-04-07 | Kalinga Health Hub | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32483 | CONSORTIUM | kalingahealth.manda@innovationsch.org | 27 |  | 2026-04-07 | GXIV | 110008415 | Module B2(220000649) for replacement | Replaced module B2 | - replaced module by 618089 - checked and cleaned cartridge bay, plunger rod, filter - performed self test - generated reports | Unit is operational |  | GeneXpert Module SN: 618089 | 102479 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6513587053112341412/6513587053112341412_signature_28.png | Chin Lyn Corpin | https://www.jotform.com/uploads/macarewecare/222927978372067/6513587053112341412/6513587053112341412_signature_31.png
| 2026-04-07 | 2026-04-07 | San Remigio Rural Health Unit | Poblacion | 7 | Cebu | TB Laboratory | NO | TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32482 | WHO | rodacanetecamay@gmail.com | 22.7 |  |  | GXIV 6C | 847333 | Lost communication all modules die to earth quake. | . Save System Log file . Save Installation Qualifications Report . Save Assay Statistics Report. | . Resit all wire in module's and check voltage supply of gateway board. . Run sample test for all module's for B1,B2,B3,B4. . B1 for observation lost communication error 2123. | All module are functional ok to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6513541458422005868/6513541458422005868_signature_28.png | Roda Camay | https://www.jotform.com/uploads/macarewecare/222927978372067/6513541458422005868/6513541458422005868_signature_31.png
| 2026-04-07 | 2026-04-07 | Ramos Rural Health Unit | Ramos Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32481 | DOH | jadeeh8143@yahoo.com | 26 |  |  | GXIV-D 6C | 831721 | B2: 757555 for module replacement | Preventive maintenance Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational New module B2:220000611 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6513524355911020766/6513524355911020766_signature_28.png | Emily Jade Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/6513524355911020766/6513524355911020766_signature_31.png
| 2026-04-07 | 2026-04-07 | Giporlos Rural Health Unit | Giporlos Eastern Samar | 8 | Eastern Samar | TB Laboratory | NO | CALIBRATION TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32480 | PBSP | chenniegabornes25@gmail.com | 26 |  |  | GXIV - 6c | 835060 | CPU - not working properly | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6513474668793150478/6513474668793150478_signature_28.png | Ma. Chennie I. Gabornes | https://www.jotform.com/uploads/macarewecare/222927978372067/6513474668793150478/6513474668793150478_signature_31.png
| 2026-04-07 | 2026-04-07 | Dasmariñas Social Hygiene Clinic | Green Breeze Village, 4 Sitio Piela, Barangay Sampaloc 3, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32479 | MACARE-HBDC | manzanomarkdave060790@gmail.com | 24 |  |  | GX-IV-D-10C | 120007702 | For installation | Installed the GeneXpert system | Installed the GeneXpert system -GeneXpert unit -Workstation -UPS -Printer -Barcode scanner Functionality test Import assay Generated IQR | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6513463800529201867/6513463800529201867_signature_28.png | Mark Dave Manzano | https://www.jotform.com/uploads/macarewecare/222927978372067/6513463800529201867/6513463800529201867_signature_31.png
| 2026-04-06 | 2026-03-26 | Malilipot Rural Health Unit | Malilipot, Albay | 5 | Albay | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32478 | PBSP | faithbuban2124@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3157 |  | GX4-6C | 808280 | Date and time reset Stuck cartridge | Replaced CRC BIOS battery Stuck cartridge readjust(resolved) Data collection | Performed troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512898543103634223/6512898543103634223_signature_28.png | Faith Buban | https://www.jotform.com/uploads/macarewecare/222927978372067/6512898543103634223/6512898543103634223_signature_31.png
| 2026-04-06 | 2026-04-01 | Alabang Health Center | Alabang, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32476 | DOH | cathycarasig@yahoo.com | 22 |  | 2026-04-01 | GX-IV-D-6C | 831913 | For CPU replacement | Replaced CPU | Replaced defective CPU Performed maintenance Functionality test Import assay Generated reports | Unit is operational. All modules operational. |  | Old CPU: 2UA8471W1J New CPU: CPH0625F59729 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512702023612682600/6512702023612682600_signature_28.png | Catherine Carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/6512702023612682600/6512702023612682600_signature_31.png
| 2026-04-06 | 2026-03-31 | Pasig Treatment Hub - Sumilang | Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | pathmedtech@gmail.com |  | 32474 | MACARE-HBDC | dansumulong.pasigcho@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: |  | GX-IV-D-10C | 110025128 | Error 1001/2014 on module A2 | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -icore -Filter Generated reports | Unit is operational. Module A2/210129252 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512702023614783598/6512702023614783598_signature_28.png | Dan Nathaniel Sumulong | https://www.jotform.com/uploads/macarewecare/222927978372067/6512702023614783598/6512702023614783598_signature_31.png
| 2026-04-06 | 2026-04-01 | Department of Health Treatment and Rehabilitation Center - Bicutan | Bicutan, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32477 | PBSP | dohtrcbicutan.lab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1754 |  | GX-IV-D-6C | 807452 | For storing to other room | Transferred GeneXpert system to other room | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -icore -Filter Generated reports Transferred GeneXpert system to other rooms | Unit temporarily stored in other room. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512702023617526244/6512702023617526244_signature_28.png | Ulric Lane D Angosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6512702023617526244/6512702023617526244_signature_31.png
| 2026-04-06 | 2026-03-31 | Pasig Treatment Hub - Champ | Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32475 | MACARE-HBDC | pathmedtech@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 135 ASSAY: Xpert HIV-1 Viral Load, COUNT: 645 |  | GX-IV-D-10C | 110025127 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512702023619906229/6512702023619906229_signature_28.png | Jhozza Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/6512702023619906229/6512702023619906229_signature_31.png
| 2026-04-06 | 2026-03-31 | Metro Health Specialist Hospital Inc. | Poblacion, Sorsogon City | 5 | Sorsogon | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32473 | COMMERCIAL | cmanalansan45@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 24 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 51 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 3 |  | GX4-10C | 120002716 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Self test module Generate data | Performed Preventive Maintenance | Unit is operational All 4 modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512559843107928125/6512559843107928125_signature_28.png | Shawn manalansan | https://www.jotform.com/uploads/macarewecare/222927978372067/6512559843107928125/6512559843107928125_signature_31.png
| 2026-04-06 | 2026-03-19 | Balud Rural Health Unit | Balud, Masbate | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32472 | PBSP | rhu_balud1@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 631 |  | GX4-6C | 808633 | 687101 for replacement Date and time reset | Module B1: 687101 >> 210114896 Replaced CRC BIOS battery Cleaned cartridge bays, plunger rod, icore, fan filter Generate Data | Performed module replacement Performed Preventive Maintenance Performed CRC BIOS battery replacement | Unit is operational All Gx modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512558973101004704/6512558973101004704_signature_28.png | Sharon Oliva | https://www.jotform.com/uploads/macarewecare/222927978372067/6512558973101004704/6512558973101004704_signature_31.png
| 2026-04-06 | 2026-03-17 | Bicol Medical Center | Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32471 | MACARE-HBDC | bamborjaph@gmail.com | 25 |  |  | MCS+ | 15H027G | Unit leaked fluid in centrifuge | Cleaned Unit surface Cleaned Centrifuge bowl Checked parameters | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512510583101256620/6512510583101256620_signature_28.png | Reynan Borja | https://www.jotform.com/uploads/macarewecare/222927978372067/6512510583101256620/6512510583101256620_signature_31.png
| 2026-04-06 | 2026-03-17 | Bicol Medical Center | Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32470 | MACARE-HBDC | bamborjaph@gmail.com | 25 |  |  | MCS+ | 15H031G | Dirty waste bag | Cleaned unit surface Cleanes centrifuge bowl Checked parameters Replaced waste bag | Replaced waste bag Performed preventive maintenance | Unit is operarional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512506433102445350/6512506433102445350_signature_28.png | REYNAN BORJA | https://www.jotform.com/uploads/macarewecare/222927978372067/6512506433102445350/6512506433102445350_signature_31.png
| 2026-04-06 | 2026-02-23 | St Lukes Extension Clinic | J. Bocobo St, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 32469 | MACARE-HBDC | laboratory@slec.ph | 22 |  |  | Infinity 48s | 809643 | Conveyor belt not working during operation | Adjusted damaged conveyor belt to avoid stuck up | Performed basic troubleshooting | Unit is operational, conveyor belt temporarily fixed but shows signs of wear and tear damage and needs replacing |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512470093618282457/6512470093618282457_signature_28.png | Ma. Sarah Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6512470093618282457/6512470093618282457_signature_31.png
| 2026-04-06 | 2026-04-01 | PGH MRL | Taft Avenue | NCR | NCR 1st - Manila | MRL | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32468 | PBSP | mcreyes@up.edu.ph | 30 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 189 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1475 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1272 |  | GX4 6C | 838145 | Due for xpertcheck and maintenance | Cleaned unit Cleaned plunger rod, filter, icore Gathered reports Performed xpertcheck Back up gxx files | Performed xpertcheck | B3 806677 failed in xpertcheck B1 ,B2 passed B4 unable to perform xpertcheck due to lack of cartridge |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512431363618273122/6512431363618273122_signature_28.png | Marjorie C. Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6512431363618273122/6512431363618273122_signature_31.png
| 2026-04-06 | 2026-04-01 | PGH MRL | Taft Avenue | NCR | NCR 1st - Manila | MRL | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32467 | PBSP | mcreyes@up.edu.ph | 30 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 189 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1475 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1272 |  | GX4 6C | 802511 | Due for xpertcheck and maintenance | Cleaned unit Cleaned filter, plunger rod, cartridge bay Gathered reports Back up gxx files | Performed xpertcheck and maintenance | C4 passed C1, C2, C3 unable to performed xpertcheck lack of kits |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6512428491227198919/6512428491227198919_signature_28.png | Marjorie C. Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6512428491227198919/6512428491227198919_signature_31.png
| 2026-04-01 | 2026-04-01 | Biopath Garden Heights | E. Rodriguez | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | THERMO FISHER SCIENTIFIC | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32466 | COMMERCIAL | debuyanterahgrace@gmail.com | 35 |  |  | 88400D | 160414501161220 | High temp due to high ambient temp Formation of ice in tubings | Check unit Checked connwction Checked amperage Shut down unit | Performed troubleshooting | Advise end user turn on the machine thus coming monday 4/6/2026 For further observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6508515433344252231/6508515433344252231_signature_28.png | Terah debuyan | https://www.jotform.com/uploads/macarewecare/222927978372067/6508515433344252231/6508515433344252231_signature_31.png
| 2026-04-01 | 2026-04-01 | PGH | Taft Avenue | NCR | NCR 1st - Manila | Central Lab | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32465 | MACARE-HBDC | Pghmolecularlab@gmail.com | 26 |  |  | GX16 | 844975 | B1, B3, C2 HW failed A3, A4, B2, B4, C1, C3, C4, D2, D3, D4 disabled due to error | Cleaned unit Cleaned filter, plunger rod, icore, cartridge bay Back up data Gathered reports | Performed maintenance | A1, A2, D1 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6508515023348710274/6508515023348710274_signature_28.png | Kleante Bautro | https://www.jotform.com/uploads/macarewecare/222927978372067/6508515023348710274/6508515023348710274_signature_31.png
| 2026-04-01 | 2026-04-01 | PGH | Taft Avenue | NCR | NCR 1st - Manila | Central lab | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32464 | MACARE-HBDC | Pghmolecularlab@gmail.com | 31 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 46 ASSAY: Xpert SARS-CoV-2, COUNT: 360 |  | GX4 6C | 834624 | B2, B3, B4 disabled Defective keyboard | Cleaned unit Cleaned filter, Icore, Plunger rod, Cartridge bay Gathered reports | Performed maintenance | B1 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6508514583342954148/6508514583342954148_signature_28.png | Kleante Bautro | https://www.jotform.com/uploads/macarewecare/222927978372067/6508514583342954148/6508514583342954148_signature_31.png
| 2026-04-01 | 2026-03-31 | Valeriano Fugoso Health Center RTDL | Lacson cor Arago st., Sta Cruz Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32463 | PBSP | julie19_esguerra@yahoo.com | 28 |  | 2026-03-31 | GX4 6C | 831618 | For module replacement 702259 - 210079099 650710 - 210087033 For monitor replacement s/n cnk832243y - cn-01hrfc-fcc00-4cj-eoux-A00 For CPU replacement 2ua8471v20 - CPH0625F59728 | Replaced module, CPU and monitor Upload Adf and user account Performed self test Cleaned plunger Gathered reports | Performed repair | Operational |  | Genexpert 6 color module Monitor Dell workstation | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6508513783342856922/6508513783342856922_signature_28.png | Julie o. Esguerra | https://www.jotform.com/uploads/macarewecare/222927978372067/6508513783342856922/6508513783342856922_signature_31.png
| 2026-04-01 | 2026-04-01 | UST Hospital | Espana, Mania | NCR | NCR 1st - Manila | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 32462 | COMMERCIAL | bloodbank@usthospital.com.ph | 25 |  | 2026-04-01 | MCS + | 02F069 | Unit for regular maintenance | Performed maintwnance. | Performed unit maintenance and cleaning. Checked pump rotor roller. Performed functionality test. | Unit is operational. |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6508376051118272218/6508376051118272218_signature_28.png | Cherrissa Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/6508376051118272218/6508376051118272218_signature_31.png
| 2026-04-01 | 2026-03-25 | Surigao Doctors Hospital | Brgy. Rizal, Surigao City | 13 | Surigao del Norte | Microbiology Laboratory | NO | INSTALLATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 32461 | COMMERCIAL | surigaodoctors@gmail.com | 27 |  |  | GX4 6 Colors 4 sites | 830967 | None | None | Installation of GeneXpert Machine and its peripherals. Discussing Preventive Maintenance and functionality of GX Software. | All modules and units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6508351821229157787/6508351821229157787_signature_28.png | Russ Wayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6508351821229157787/6508351821229157787_signature_31.png
| 2026-04-01 | 2026-04-01 | UST Hospital | Espana, Mania | NCR | NCR 1st - Manila | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 32460 | COMMERCIAL | ustimp@yahoo.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 21 ASSAY: Xpert HIV-1 Viral Load, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5 ASSAY: , COUNT: 744 | 2026-04-01 | GXIV | 11000711 | For unit xpertcheck calibration. | Performed unit mintenance and xpertcheck calibration. | Perforned unit maintanance and cleaning. Performed xpertcheck calibration. | Unit is operational. Module B2: 210004296 - service required. |  | Xpertcheck kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6508332190422444220/6508332190422444220_signature_28.png | SHERLEEN CALAUNAN | https://www.jotform.com/uploads/macarewecare/222927978372067/6508332190422444220/6508332190422444220_signature_31.png
| 2026-04-01 | 2026-04-01 | Aloguinsan Rural Health Unit | Poblacion Aloguinsan Cebu | 7 | Cebu | TB Laboratory | NO | CPU Pullout | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32459 | DOH | aloguinsanrhulaboratory@gmail.com | 24 |  | 2026-04-01 | GXIV 6C | 834812 | Cannot power ON cpu Serial number CPH119F04158 | Pull out CPU | Pull out CPU | For repair. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6508208711224614617/6508208711224614617_signature_28.png | Jennifer Sayson | https://www.jotform.com/uploads/macarewecare/222927978372067/6508208711224614617/6508208711224614617_signature_31.png
| 2026-04-01 | 2026-04-01 | Wenceslao Dela Paz Health Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32458 | PBSP | ramosfely38@gmail.com | 34 |  |  | GXIV | 831876 | CPU for replacement | Replaced CPU | #NAME? | Unit is operational Room ambient temperature exceeds 30 deg C Room air conditioner is recommended for repair/replacement to prevent module damage |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6508182296418282170/6508182296418282170_signature_28.png | Felicidad De Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/6508182296418282170/6508182296418282170_signature_31.png
| 2026-03-31 | 2026-03-27 | Isulan Rural Health Unit | Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32456 | UNOPS | isulanrhuclin.lab@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1224 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 |  | GxIV-10C | 210023520 | None | Perform Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507740157323381313/6507740157323381313_signature_28.png | FEBIE JEN APIADO | https://www.jotform.com/uploads/macarewecare/222927978372067/6507740157323381313/6507740157323381313_signature_31.png
| 2026-03-31 | 2026-03-26 | Norala District Hospital | Norala, South Cotabato | 12 | South Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32457 | DOH | noralaDHlaboratory@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1266 |  | GxIV-6C | 838543 | None | Perform Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507740157329196244/6507740157329196244_signature_28.png | Geraldine Magan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6507740157329196244/6507740157329196244_signature_31.png
| 2026-03-31 | 2026-03-31 | Pasig City General Hospital | Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32455 | PBSP | vilmarumadhay@gmail.com | 24 |  | 2026-03-31 | GX-IV-D-6C | 805010 | For CPU and Gateway board replacement | Replaced CPU and gateway board | Replaced CPU Replaced gateway board Configured gateway board SN Installed assay Performed maintenance Generated reports | Unit is operational. All modules operational |  | Old CPU SN: 2UA3520W1P New CPU SN: CPH0625F59563 Old gateway board SN: 720025735 New Gateway board SN: 720053450 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507471301323061400/6507471301323061400_signature_28.png | Vilma Umadhay | https://www.jotform.com/uploads/macarewecare/222927978372067/6507471301323061400/6507471301323061400_signature_31.png
| 2026-03-31 | 2026-03-31 | Klinika Project 7 Health Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32454 | MACARE-HBDC | freyaoriana@yahoo.com | 27 |  |  | GXIV-10C | 845446 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: Failed B2: Passed B3: Failed B4: Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507459929515502034/6507459929515502034_signature_28.png | Alyanna Faye Zerudo | https://www.jotform.com/uploads/macarewecare/222927978372067/6507459929515502034/6507459929515502034_signature_31.png
| 2026-03-31 | 2026-03-31 | Bogo City Health Office | Bogo City, Cebu | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32453 | PBSP | bogocholab@gmail.com | 26 |  |  | Gx 4 | 120000494 | retry modules 2 and 4 | Re-run for retry modules 2 and 4 Gather data IQR and calibration data Transfer location of genexpert machine from Bogo Provincial hospital to super health center | Re-run for retry modules 2 and 4 Gather data IQR and calibration data Transfer location of genexpert machine from Bogo Provincial hospital to super health center | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507396746214679673/6507396746214679673_signature_28.png | Reggie V. Dahili | https://www.jotform.com/uploads/macarewecare/222927978372067/6507396746214679673/6507396746214679673_signature_31.png
| 2026-03-31 | 2026-03-31 | Tarlac City Health Center X | Brgy. Tibag, Tarlac City | 3 | Tarlac | Laboratories | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32452 | DOH | jayteecornelius@gmail.com | 25 |  |  | GXIV-D 6C | 831587 | CPU:2UA8471VSN For replacement B2:784728 for replacement | Preventive maintenance Installation and setup Module replacement | Performed preventive maintenance Performed CPU installation and setup Performed module replacement Gathered reports Performed functionality test | Unit is operation New CPU:CPH0625F59558 New B2:736799 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507365547986760962/6507365547986760962_signature_28.png | Jaytee Cornelius Bachiller | https://www.jotform.com/uploads/macarewecare/222927978372067/6507365547986760962/6507365547986760962_signature_31.png
| 2026-03-31 | 2026-03-24 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32451 | COMMERCIAL | molecularpathology@asianhospital.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 10 ASSAY: Xpert CT/NG, COUNT: 86 ASSAY: Xpert HBV VL, COUNT: 177 ASSAY: Xpert HIV-1 Viral Load, COUNT: 56 ASSAY: Xpert HPV, COUNT: 201 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1294 ASSAY: Others, COUNT: 94 |  | GXIV-D-10C | 110014996 | Module 2 with Temperature Error(Error Code:2014) | Module Replacement. | Replaced Defective Module with New Module. Performed Self-test. Generate Reports. Performed Functionality Test. | B2: Def/210076074 -> New/210077663 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507330110911140939/6507330110911140939_signature_28.png | Trisha Faye Pastor | https://www.jotform.com/uploads/macarewecare/222927978372067/6507330110911140939/6507330110911140939_signature_31.png
| 2026-03-31 | 2026-03-31 | Quezon City Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32450 | MACARE-HBDC | molecularlab.cityhealth@quezoncity.gov.ph | 27 |  |  | GXIV-10C | 110020696 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507318739514968204/6507318739514968204_signature_28.png | Quezon City Molecular Diagnostics Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/6507318739514968204/6507318739514968204_signature_31.png
| 2026-03-31 | 2026-03-27 | Bayambang District Hospital | Bayambang, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32449 | PBSP | bevzneri@gmail.com | 28 |  |  | GXIV-4-D-6C | 812830 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Modules A1 and A4 currently disabled after failed XpertCheck on 08.22.2025 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507284893644654559/6507284893644654559_signature_28.png | Beverly A. Neri | https://www.jotform.com/uploads/macarewecare/222927978372067/6507284893644654559/6507284893644654559_signature_31.png
| 2026-03-31 | 2026-03-26 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32448 | COMMERCIAL | annabellelcatbagan@gmail.com | 26 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is currently Disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507284093645737837/6507284093645737837_signature_28.png | ANGELA KATE DELA PEÑA | https://www.jotform.com/uploads/macarewecare/222927978372067/6507284093645737837/6507284093645737837_signature_31.png
| 2026-03-31 | 2026-03-25 | Mangatarem Rural Health Unit | Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32447 | DOH | slasheiji1988new@gmail.com | 26 |  |  | GXIV-4-D-6C | 838284 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6507282843645785995/6507282843645785995_signature_28.png | Ken Lee Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/6507282843645785995/6507282843645785995_signature_31.png
| 2026-03-30 | 2026-03-28 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 32446 | COMMERCIAL | vsmmcbloodservices@gmail.com | 24 |  |  | MCS+ | 02A090 | Purple valve problem during startup of the procedure of apheresis. | Check pneumatic assembly in the MCS+. Restart machine and test in the diagnostic mode. Check voltage supply and UPS side. | Check pneumatic assembly in the MCS+. Restart machine and test in the diagnostic mode. Check voltage supply and UPS side. | MCS+ unit is operational. Pneumatic assembly sometimes noisy during startup. Require end-user to observe during the week. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506783513423956221/6506783513423956221_signature_28.png | Charlotte Mae Ygoña, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6506783513423956221/6506783513423956221_signature_31.png
| 2026-03-30 | 2026-03-27 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32444 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 22 |  |  | MCS+ | 04H070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506679153611717097/6506679153611717097_signature_28.png | Gracia Tina A. Angara | https://www.jotform.com/uploads/macarewecare/222927978372067/6506679153611717097/6506679153611717097_signature_31.png
| 2026-03-30 | 2026-03-29 | St. Luke's Medical Center -BGC | Rizal Dr., Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 32445 | COMMERCIAL | advancedmoleculardx@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 50 ASSAY: Xpert CT/NG, COUNT: 192 ASSAY: Xpert HBV VL, COUNT: 151 ASSAY: Xpert HPV, COUNT: 349 ASSAY: Xpert MTB/RIF Ultra, COUNT: 852 ASSAY: Xpert HCV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 146 |  | GX-IV-D-6C | 839232 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module A2/866696 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506679153612002138/6506679153612002138_signature_28.png | Jordan Purpura | https://www.jotform.com/uploads/macarewecare/222927978372067/6506679153612002138/6506679153612002138_signature_31.png
| 2026-03-30 | 2026-03-29 | St. Luke's Medical Center - BGC | Rizal Dr., Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 32442 | COMMERCIAL | advancedmoleculardx@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 158 ASSAY: Xpert CT/NG, COUNT: 549 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 190 ASSAY: Xpert HIV-1 Viral Load, COUNT: 153 ASSAY: Xpert HPV, COUNT: 788 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2075 ASSAY: Xpert MTB-XDR, COUNT: 34 ASSAY: Xpert HCV VL, COUNT: 47 ASSAY: Xpert HIV-1 Viral Load, COUNT: 237 |  | GX-IV-D-10C | 120002383 | Error 2120/2124 in mdoule B1, B2 and B3 | Reseat cables | Performed maintenance Checked connections Reseat cables Performed Self test Generated reports | Unit is operational. All modules operational but under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506679153613239209/6506679153613239209_signature_28.png | Jordan Purpura | https://www.jotform.com/uploads/macarewecare/222927978372067/6506679153613239209/6506679153613239209_signature_31.png
| 2026-03-30 | 2026-03-27 | Angono Rural Health Unit | Angono, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32443 | DOH | aimielyvonnedla@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1237 |  | GX-IV-D-6C | 819864 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -filter Run XpertCheck Generated reports | Unit is operational. Module A3/624401 failed in XpertCheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506679153617432621/6506679153617432621_signature_28.png | Aimiel Yvonne De Las Alas | https://www.jotform.com/uploads/macarewecare/222927978372067/6506679153617432621/6506679153617432621_signature_31.png
| 2026-03-30 | 2026-03-30 | Don Emilio Del Valle Memorial Hospital | Ubay Bohol | 7 | Cebu | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32441 | PBSP | drmpbrmtmd28@gmail.com | 25 |  | 2026-03-30 | GXIV 6C | 807603 | Module A2,A3,A4 continuous error code # 2120,2124,4014 | . Save Installation Qualifications Report . Save System log Report . Save Assay Statistics Report. | . Cleaning filter and plunger rod cleaning. | Under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506592247616894749/6506592247616894749_signature_28.png | Michelle P. Bacaling | https://www.jotform.com/uploads/macarewecare/222927978372067/6506592247616894749/6506592247616894749_signature_31.png
| 2026-03-30 | 2026-03-30 | Eastern Visayas Medical Center | Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32440 | PBSP | earosales80@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2991 ASSAY: Xpert MTB-XDR, COUNT: 40 |  | GXIV 10c | 120000177 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506568606715247604/6506568606715247604_signature_28.png | Emmylou Rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/6506568606715247604/6506568606715247604_signature_31.png
| 2026-03-30 | 2026-03-27 | San Mateo Super Health Center | GSIS rd, San Mateo | 4A | Rizal | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32439 | PBSP | superhealthlaboratory@gmail.com | 26 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2108 ASSAY: Xpert HIV-1 Viral Load, COUNT: 119 |  | GX4D | 808616 | As per end user they booting issue in workstation | Checked workstation Cleaned memory card Reset all cable Checked bios set up Test run unit | Performed troubleshooting | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506503494716002737/6506503494716002737_signature_28.png | San Mateo Super Health | https://www.jotform.com/uploads/macarewecare/222927978372067/6506503494716002737/6506503494716002737_signature_31.png
| 2026-03-30 | 2026-03-27 | Lung Center of the Philippines | Quezon Avenue | NCR | NCR 2nd - Quezon City | Microbiology | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32438 | MACARE-HBDC | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 24 ASSAY: Xpert MTB-XDR, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX4D | 110027386 | For printer installtion | Performed printer installation Test unit | Printer installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506503024719277193/6506503024719277193_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6506503024719277193/6506503024719277193_signature_31.png
| 2026-03-30 | 2026-03-26 | To Love Yourself | Anglo Shaw blvd | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | LW Scientific C3 select | 32437 | COMMERCIAL | laboratory@loveyourself.ph | 26 |  |  | EZ Swing 3k | 1101271 | Defective gas spring of lid assy Timer switch malfunction | Checked unit Cleaned unit Checked and test timer switch Checked Gas spring Performed functionality test Performed calibration Test run unit | Performed troubleshooting Performed calibration and maintenance | Checked availabilty of the parts For replacement of timer switch and gas spring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6506502454711989943/6506502454711989943_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6506502454711989943/6506502454711989943_signature_31.png
| 2026-03-27 | 2026-03-27 | St. Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 32436 | MACARE-HBDC | laboratory@slec.ph | 23 |  | 2026-03-24 | Infinity 48s | 809643 | Conveyor belt damaged | Replaced conveyor belt, checked conveyor belt functionality | Performed conveyor belt replacement | Unit is operational |  | 200-5571 golden yellow v-guided type 658 belt | DR 13833 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6504108386137497361/6504108386137497361_signature_28.png | Jasper Ian Sonieda | https://www.jotform.com/uploads/macarewecare/222927978372067/6504108386137497361/6504108386137497361_signature_31.png
| 2026-03-27 | 2026-03-27 | Northern Palawan Provincial Hospital | Taytay, Palawan | 4B | Palawan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 32435 | PBSP | markjakepaduga@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 278 | 2026-03-27 | GXIV | 807671 | Module B : 778977 for module replacement. | Replaced modile B2 by module sn: 887889 Performed functionality test. Generate reports. | Perfromed module replacement. Performed maintenance. | All 4 modules operational. |  | 1 module | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6504094435407970967/6504094435407970967_signature_28.png | Mark Jake Paduga | https://www.jotform.com/uploads/macarewecare/222927978372067/6504094435407970967/6504094435407970967_signature_31.png
| 2026-03-27 | 2026-03-27 | Roxas RHU | Roxas, Palawan | 4B | Palawan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 32434 | MACARE-HBDC | benjpaul.1994@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 678 | 2026-03-24 | GXIV | 110026464 | Module B2: 210139025 for module replacement | Performwd module replacement. Performed maintenance Generate reports. | Replaced module 2100139025 by moduke sn: 801928 | All 4 modules operational. |  | 1 module | DR 13811 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503959449013456705/6503959449013456705_signature_28.png | Paul John B. Benjamin | https://www.jotform.com/uploads/macarewecare/222927978372067/6503959449013456705/6503959449013456705_signature_31.png
| 2026-03-27 | 2026-03-27 | Southern Philippines Medical Center | JP Laurel Bajada Avenue | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32433 | PBSP | spmcrhivda@gmail.com | 209 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 329 |  | GX-IV 6C 4D | 80298 | Onsite Assistance Module B1/ Lost communication | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Reset network cable connection and gatewayboard Generate IQR, Assay Stat, and System Log | Perform Preventive Maintenance, and Onsite Troubleshooting | All Modules are operational | Pullout the Defective UPS for SN 110023506 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503934780328477243/6503934780328477243_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6503934780328477243/6503934780328477243_signature_31.png
| 2026-03-27 | 2026-03-27 | Bustos Rural Health Unit | Bustos | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32432 | DOH | ronald_galvez2002@yahoo.com | 28 |  |  | GXIV | 831874 | CPU for replacement | Replaced CPU | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503889567517206432/6503889567517206432_signature_28.png | Ronald Galvez | https://www.jotform.com/uploads/macarewecare/222927978372067/6503889567517206432/6503889567517206432_signature_31.png
| 2026-03-27 | 2026-03-26 | Morong Rural Health Unit | Morong, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32431 | DOH | moronglab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 618 |  | GX-IV-D-6C | 821213 | Boot device not found | Replaced CMOS battery | Replaced CMOS battery Configured boot options Performed maintenance Functional test Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503852147212142863/6503852147212142863_signature_28.png | Necitas V Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/6503852147212142863/6503852147212142863_signature_31.png
| 2026-03-27 | 2026-03-27 | Pulupandan PCF | Pulupandan, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR | CEPHEID | Gerald Fritz Selga |  | gselga@macare-medicals.com |  |  | 32430 | PBSP | pulupandantbdots@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 492 |  | GX4 6C | 847255 | UPS problem - No power | Reinstall chaeged batteries, functional test for UPS. | . Pullout and recharge 2 units 12 V 7AH batreries. | UPS operational, Genexpert system operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503848627164832078/6503848627164832078_signature_28.png | arriane gaye escanes | https://www.jotform.com/uploads/macarewecare/222927978372067/6503848627164832078/6503848627164832078_signature_31.png
| 2026-03-27 | 2026-03-27 | Apalit Rural Health Unit | San Juan Apalit Pampanga | 3 | Pampanga | Laboratory | NO | INSTALLATION REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32429 | DOH | duenasblesilda@gmail.com | 25 |  |  | GXIV-D 6C | 831900 | Due for calibration CPU for replacement Defective CPU SN:2UA8471VWH | CPU replacement Preventive maintenance Xpertcheck | Performed preventive maintenance Performed CPU replacement Performed xpertcheck Gathered reports Performed functionality test | Unit is operational New CPU: CPH0625F59565 Xpertcheck result B1: 210019116 PASSED B2: 758106 FAILED / 1 Retry B3: 663076 PASSED B4: 674002 FAILED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503837093726092848/6503837093726092848_signature_28.png | Blesilda Duenas | https://www.jotform.com/uploads/macarewecare/222927978372067/6503837093726092848/6503837093726092848_signature_31.png
| 2026-03-26 | 2026-03-26 | Sulu provincial hospital | Asturias ,scott rd, Jolo, sulu | BARMM | Sulu | Laboratory | YES | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32428 | PBSP | sphcl2016@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1980 | 2026-03-26 | GX 4 - 6c | 815497 | None | Site inspection, preventive Maintenence, checked machine status and data collection. | Site inspection, preventive Maintenence, checked machine status and data collection. | All modules are operational | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503470083411318648/6503470083411318648_signature_28.png | Anam H. Nuh | https://www.jotform.com/uploads/macarewecare/222927978372067/6503470083411318648/6503470083411318648_signature_31.png
| 2026-03-26 | 2026-03-26 | Sulu provincial hospital | Asturias ,scott rd, jolo, sulu | BARMM | Sulu | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32427 | IOM | sphcl2016@gmail.com | 29 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert MTB-XDR, COUNT: 12 | 2026-03-26 | GX 4 - 10c | 110008651 | Module 2, 3 and 4 are not in use due to communication lost error. | Site inspection, preventive Maintenence, checked machine status and data collection. | Site inspection, preventive Maintenence, checked machine status and data collection. | Module 1 is operational while module 2, 3 and 4 is not in use due to communication errors. | Only module 1 is operational while module 2, 3 and 4 gives communication errors. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503452543414965240/6503452543414965240_signature_28.png | Anam H. Nuh | https://www.jotform.com/uploads/macarewecare/222927978372067/6503452543414965240/6503452543414965240_signature_31.png
| 2026-03-26 | 2026-03-26 | Sulu Provincial Hospital | Asturias ,scott rd, Jolo, sulu | BARMM | Sulu | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32426 | PBSP | sphcl2016@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1150 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 | 2026-03-26 | GX 4 - 10c | 120000488 | module 3 gives an error of 2008, 2037, 2034 | Site inspection, Preventive Maintenence, checked machine status and data collection. | Site inspection, Preventive Maintenence, checked machine status and data collection. | Module 1, 2 and 4 are operational while module 3 gives an error of 2008, 2037, 2034. | module 3 gives an error of 2008, 2037, 2034. This genexpert machine must have a new Uninterrupted power supply as soon as possible. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503451513414818285/6503451513414818285_signature_28.png | Anam H. Nuh | https://www.jotform.com/uploads/macarewecare/222927978372067/6503451513414818285/6503451513414818285_signature_31.png
| 2026-03-26 | 2026-03-26 | Western Visayas Medical Center | Iloilo City | 6 | Iloilo | TB Culture Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 32425 | PBSP | meobordo.pbsp@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 346 ASSAY: Xpert HPV, COUNT: 69 |  | GX-IV | 706910 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503366525469565988/6503366525469565988_signature_28.png | Melnia Andrea E. Obordo,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6503366525469565988/6503366525469565988_signature_31.png
| 2026-03-26 | 2026-03-26 | St Pauls Hospital of Iloilo | ILOILO CITY | 6 | Iloilo | Blood Bank | NO | REPAIR | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 32424 | COMMERCIAL | sphi.bloodbank@yahoo.com | 25 |  |  | MCS+ | 98B020 | Blood Pump Ruber Rotor noise | #NAME? | -Inspect Pump rotors, confirm rubber wear for blood pump -Replace rubber roller(2 pcs used) -Perform functionality test, passed | Equipment is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503362615463488886/6503362615463488886_signature_28.png | Mariejol G. Dureza | https://www.jotform.com/uploads/macarewecare/222927978372067/6503362615463488886/6503362615463488886_signature_31.png
| 2026-03-26 | 2026-03-26 | San Rafael RHU | San Rafael, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32423 | DOH | tatsdayrit25@gmail.com | 27 |  |  | GXIV | 831692 | Desktop for replacement Unit for calibration | Replaced CPU Performed XpertCheck | #NAME? | Unit is operational All modules passed the calibration except B1 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503189721512063548/6503189721512063548_signature_28.png | Joana Dayrit | https://www.jotform.com/uploads/macarewecare/222927978372067/6503189721512063548/6503189721512063548_signature_31.png
| 2026-03-26 | 2026-03-26 | Davao Reproductive Health and Wellness Center | 838262 | 11 | Davao del Sur | TB-HIV | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32422 | DOH | rhwclab@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 408 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3228 |  | GX-IV 6C 4D | 838262 | Modules C3/ SN 807728 and C4/ SN: 210043526 Encounter Errors 5007/5017 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Onsite Assistance, Troubleshooting and Data Gathering | All Modules are operational | Sent datas and documentation to higher ups for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503103604789753647/6503103604789753647_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6503103604789753647/6503103604789753647_signature_31.png
| 2026-03-26 | 2026-03-26 | Davao Reproductive Health and Wellness Center | Jacinto St, Poblacion Davao City | 11 | Davao del Sur | TB-HIV | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32421 | MACARE-HBDC | rhwclab@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 408 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3228 |  | GX-IV 6C 4D | 839236 | Multiple Error on Modules B1/ SN:210003270 and B4/ SN:866835 | Errors codes are Error 5007, 5017. Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Onsite Assistance and Troubleshoot | Only 3 Modules are available, Module B3/ SN:866834 Hardware Failure | Data gathered sent to higherups for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503094124784542391/6503094124784542391_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6503094124784542391/6503094124784542391_signature_31.png
| 2026-03-26 | 2026-03-26 | Klinika Bernardo Health Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32420 | PBSP | klinikabernardolaboratory@gmail.com | 27 |  |  | GXIV | 808618 | Module D4: 667961 for replacement Gateway board for replacement | Replaced module Replaced gateway board | - replaced module by 912901 - replaced gateway board - generated reports - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503053711513138658/6503053711513138658_signature_28.png | Mylen Manongsong-Sibal | https://www.jotform.com/uploads/macarewecare/222927978372067/6503053711513138658/6503053711513138658_signature_31.png
| 2026-03-26 | 2026-03-26 | Leyte Provincial Hospital | Palo Leyte | 8 | Leyte | TB laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32419 | PBSP | lphlabdeptcovid19@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1613 |  | GXIV - 10c | 110018056 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503027838413010466/6503027838413010466_signature_28.png | Elnora Ortiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6503027838413010466/6503027838413010466_signature_31.png
| 2026-03-26 | 2026-03-26 | Abucay Rural Health Unit | Dacena Ave. Abucay, Bataan | 3 | Bataan | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32418 | PBSP | parreraledelina22176@gmail.com | 26 |  |  | GXIV-D 6C | 831878 | CPU for replacement CPU:2UA8471W2N | Repair and installation Preventive maintenance | Performed preventive maintenance Replaced defective CPU Gathered reports Performed functionality test | Unit is operational New CPU:CPH0625F59732 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503019603727736386/6503019603727736386_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/6503019603727736386/6503019603727736386_signature_31.png
| 2026-03-26 | 2026-03-26 | Sulu Sanitarium General Hospital | Upper san raymundo, jolo, sulu | BARMM | Sulu | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32417 | COMMERCIAL | ssghlaboratory@gmail.com | 28 |  | 2026-03-10 | GX 4 - 10c | 110015481 | None | Delivered Xpert MTB/RIF Ultra (50 test) 2kits with Lot no. 1001507217. Checked if all modules are working and in condition to operate. Preventive Maintenence | Delivered Xpert MTB/RIF Ultra (50 test) 2kits with Lot no. 1001507217. Checked if all modules are working and in condition to operate. Preventive Maintenance. | All modules are available and operational. | No test have been conducted in the machine since installation. | GXMTB/RIF-ULTRA-50 | 13809 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503012019729579106/6503012019729579106_signature_28.png | Shirley y. Julkanain | https://www.jotform.com/uploads/macarewecare/222927978372067/6503012019729579106/6503012019729579106_signature_31.png
| 2026-03-26 | 2026-03-26 | Abucay Rural Health Unit( Roving) | Dacena Ave. Abucay, Bataan | 3 | Bataan | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32416 | PBSP | parreraledelina22176@gmail.com | 26 |  |  | GXIV-D 6C | 897466 | CPU For replacements Defective CPU SN:2UA5201PHZ | CPU Replacement | Performed preventive maintenance Replaced defective CPU Gathered reports Performed functionality test | Unit is operational New CPU:CPH0625F59733 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6503011133722514463/6503011133722514463_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/6503011133722514463/6503011133722514463_signature_31.png
| 2026-03-25 | 2026-03-25 | CEBU CITY JAIL MALE DORM | Kalunasan Guadalupe Cebu City | 7 | Cebu | Tb laboratory | YES | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | jasperperanco@gmail.com |  | 32415 | PBSP | jasperperanco@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1710 |  | GX4 | 805214 | Module 4 SN# 865164 encountered Error 2126 communication error for several times. | #NAME? | #NAME? | All modules are available after checking. Required end-user to observe within the week. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6502381223211220088/6502381223211220088_signature_28.png | Jasper Peranco | https://www.jotform.com/uploads/macarewecare/222927978372067/6502381223211220088/6502381223211220088_signature_31.png
| 2026-03-25 | 2026-03-25 | Kananga Rural Health Unit | Kananga, Leyte | 8 | Leyte | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32414 | PBSP | retulla_den@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 690 |  | 6C GXIV-4D 2 Sites | 835085 | None | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6502320059911363662/6502320059911363662_signature_28.png | Dennis O. Retulla | https://www.jotform.com/uploads/macarewecare/222927978372067/6502320059911363662/6502320059911363662_signature_31.png
| 2026-03-24 | 2026-03-24 | Lugait RHU | Masilakon 1, poblacion Lugait, Misamis Oriental | 10 | Misamis Oriental | Laboratory | NO | ON-SITE ASSISTANCE Remove carthridge | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32413 | PBSP | sacabinlancelot@gmail.com | 31 |  |  | Gx 4 | 838562 | Stuck-up carthridge module 1 &4 | Removal of Stuck-up carthridge module 1 &4 Gather data IQR, System log | Removal of Stuck-up carthridge module 1 &4 Gather data IQR, System log | Systems Operational | Last Xpertcheck Calibration Oct. 10, 2024 by Engr. Mark Efren Ramas |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501304855428877932/6501304855428877932_signature_28.png | Lancelot Sacabin | https://www.jotform.com/uploads/macarewecare/222927978372067/6501304855428877932/6501304855428877932_signature_31.png
| 2026-03-24 | 2026-03-24 | Mayorga Rural Health Unit | Mayorga Leyte | 8 | Leyte | TB laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32412 | PBSP | irenepsia@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 772 | 2026-03-24 | GXIV - 6C 2module | 834819 | #NAME? | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501282094618129691/6501282094618129691_signature_28.png | IRENE P. SIA | https://www.jotform.com/uploads/macarewecare/222927978372067/6501282094618129691/6501282094618129691_signature_31.png
| 2026-03-24 | 2026-03-12 | Bulan Rural Health Unit | Poblacion, Bulan | 5 | Sorsogon | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32411 | PBSP | bing.gonzales33@gmail.com | 25 |  |  | GX4-6C | 807983 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational All 4 modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501263314318748709/6501263314318748709_signature_28.png | Claribel G Nacion | https://www.jotform.com/uploads/macarewecare/222927978372067/6501263314318748709/6501263314318748709_signature_31.png
| 2026-03-24 | 2026-03-19 | Benguet Provincial Health Office | La Trinidad Benguet | CAR | Benguet | TB DOTS | NO | Relocation | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32410 | UNOPS | benguetphortdl@gmail.com | 24 |  |  | GXIV-4-D-10C | 110023498 | For Relocation | Temporary Relocated GX to Molecular Laboratory for Training | Performed Relocation | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501261036616420725/6501261036616420725_signature_28.png | Tillber Baniaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6501261036616420725/6501261036616420725_signature_31.png
| 2026-03-24 | 2026-03-11 | San Jacinto Rural Health Unit | Poblacion, San Jacinto | 5 | Masbate | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32409 | PBSP | janet.bragais@yahoo.com | 25 |  |  | GX4-6C | 812793 | Module B2/712979 H/W failed | Replaced module SN: 712979 >> 781465 Cleaned cartridge bays, plunger rod Generate Data | Replaced defective module Performed Preventive Maintenance | Unit is Operational All 4-4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501260904312747182/6501260904312747182_signature_28.png | Janet A. Bragais | https://www.jotform.com/uploads/macarewecare/222927978372067/6501260904312747182/6501260904312747182_signature_31.png
| 2026-03-24 | 2026-03-17 | Cordova Primary Health Care | Poblacion, Cordova Cebu | 7 | Cebu | Tb laboratory | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | melodyjoynoelle0428@gmail.com |  | 32408 | PBSP | melodyjoynoelle0428@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 62 |  | GX4 | 808638 | Module 3 SN# 807134 sample overflow in the cartridge bay during the process. | - Test run Module 3 with sample, observed during the test. - Pull out the Module 1, Module 3 and Module 4 and found out the door hinges is defective - Recommend the replacement of the module 1, Module 3 and Module 4 due door problem. | - Troubleshoot the problem and conduct calibration of the Module 3 SN# 807134 - Pull out the Module 1, Module 3 and Module 4 and found out the door hinges is defective | All modules are operational but the door of Module 1 , 3 and 4 recommended for replacement. | Poblacion, Cordova Cebu |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501260733219416602/6501260733219416602_signature_28.png | Melody Arasan | https://www.jotform.com/uploads/macarewecare/222927978372067/6501260733219416602/6501260733219416602_signature_31.png
| 2026-03-24 | 2026-03-19 | Benguet General Hospital | La Trinidad, Benguet | CAR | Benguet | Molecular Laboratory | NO | Relocation | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32407 | MACARE-HBDC | bevlapisac@gmail.com | 23 |  |  | GXIV-4-D-10C | 120007537 | For Relocation | Transferred GX room to new designated room | Performed Relocation | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501258626614497282/6501258626614497282_signature_28.png | Beverlee Lapisac | https://www.jotform.com/uploads/macarewecare/222927978372067/6501258626614497282/6501258626614497282_signature_31.png
| 2026-03-24 | 2026-03-24 | San Isidro Rural Health Unit | San Isidro, Leyte | 8 | Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32406 | PBSP | rosel.alesna@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 547 |  | 6C GXIV-4D 2 Sites | 834816 | Module B1 error | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501243272357638336/6501243272357638336_signature_28.png | Rosel B. Alesna | https://www.jotform.com/uploads/macarewecare/222927978372067/6501243272357638336/6501243272357638336_signature_31.png
| 2026-03-24 | 2026-03-16 | Koronadal City Health Office | Koronadal City | 12 | South Cotabato | Tb-Lab | NO | INSTALLATION MAINTENANCE Software re installation | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32404 | PBSP | calvokenneth12@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2667 |  | GxIV-6C | 707529 | Laptop problem Module 4 problem | Set up computer for Genexpert use Gx software Installation | Genexpert re-installation Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501191527321023032/6501191527321023032_signature_28.png | Kenneth Paolo T. Calvo | https://www.jotform.com/uploads/macarewecare/222927978372067/6501191527321023032/6501191527321023032_signature_31.png
| 2026-03-24 | 2026-03-17 | Soccsksargen General Hospital | Surallah, South Cotabato | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32405 | PBSP | johnellamarisbautista@gmail.com | 24.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1934 ASSAY: Xpert MTB-XDR, COUNT: 22 | 2026-03-17 | GxIV-10C | 120000172 | Genexpert module B3/220011574 Defective | Replaced module B3/220011574 with new module SN:210034700 Perform Genexpert Cleaning and maintenance Checking of module status and problem Fan and filter cleaning Data collection | Genexpert machine module repair Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 10 color | Work Order: 336538 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6501191527326018273/6501191527326018273_signature_28.png | Johnella Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/6501191527326018273/6501191527326018273_signature_31.png
| 2026-03-23 | 2026-03-23 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32403 | DOH | luzvmindanavidad2@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 932 |  | 6C GXIV-4D | 838297 | None | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6500567739329037583/6500567739329037583_signature_28.png | Melinda Laroa | https://www.jotform.com/uploads/macarewecare/222927978372067/6500567739329037583/6500567739329037583_signature_31.png
| 2026-03-23 | 2026-03-23 | DUMALAG SUPER HEALTH CENTER | Dumalag, Capiz | 6 | Capiz | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 32402 | DOH | dumalag.rhucl@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-IV | 837615 | -UPS Defective -System Unit SN 2UA3520VY6 HP Z220 is defective | #NAME? | #NAME? | Gx operational, for system unit servicing. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6500472213113369102/6500472213113369102_signature_28.png | Rica Enate | https://www.jotform.com/uploads/macarewecare/222927978372067/6500472213113369102/6500472213113369102_signature_31.png
| 2026-03-23 | 2026-03-19 | Cardona Rural Health Unit | Cardona, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32401 | DOH | renalyn_villaran@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 479 |  | GX-IV-D-6C | 821215 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6500443453615412913/6500443453615412913_signature_28.png | Renalyn E. Villaran | https://www.jotform.com/uploads/macarewecare/222927978372067/6500443453615412913/6500443453615412913_signature_31.png
| 2026-03-23 | 2026-03-23 | SAN RAFAEL RURAL HEALTH UNIT | SAN RAFAEL, ILOILO | 6 | Iloilo | Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 32400 | UNOPS | apilmayfaith@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 283 |  | GX-IV | 110023536 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6500418726179090975/6500418726179090975_signature_28.png | May Faith Apil | https://www.jotform.com/uploads/macarewecare/222927978372067/6500418726179090975/6500418726179090975_signature_31.png
| 2026-03-23 | 2026-03-23 | Eastern Samar Provincial Hospital | Borongan City | 8 | Eastern Samar | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32399 | PBSP | elvalynlom@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1252 ASSAY: Xpert MTB-XDR, COUNT: 33 ASSAY: Xpert HIV-1 Viral Load, COUNT: 41 | 2026-03-23 | GXIV - 10C | 120000604 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6500410618616151950/6500410618616151950_signature_28.png | ELVALYN GRACIA G. LOMUNTAD | https://www.jotform.com/uploads/macarewecare/222927978372067/6500410618616151950/6500410618616151950_signature_31.png
| 2026-03-23 | 2026-03-18 | Earnshaw Health Center | Earnshaw St, Sampaloc, Manila | NCR | NCR 1st - Manila | Social Hygiene Clinic | NO | CALIBRATION MAINTENANCE XpertCheck | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32398 | PBSP | johndarrellturato@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 796 ASSAY: Xpert MTB/RIF Ultra, COUNT: 262 |  | GXIV-D-10C | 848262 | Unit is due for XpertCheck. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed XpertCheck. Performed Self-test. Generate Reports. | B1 and B2 Pass. B3 Service Required. B4 for Retry. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6500350783614193244/6500350783614193244_signature_28.png | John Darrell Turato | https://www.jotform.com/uploads/macarewecare/222927978372067/6500350783614193244/6500350783614193244_signature_31.png
| 2026-03-22 | 2026-03-18 | Perpetual Soccour Hospital | Cebu City | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 32397 | COMMERCIAL | pshclinlab@gmail.com | 26 |  |  | Mcs+ | 11H307 | Due date for Calibration/PM | Little adjustment for out of range sensors and weigher | Calibration and Preventive Maintenance | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6499917347893666522/6499917347893666522_signature_28.png | Debbie Joyce A. Baga | https://www.jotform.com/uploads/macarewecare/222927978372067/6499917347893666522/6499917347893666522_signature_31.png
| 2026-03-21 | 2026-03-19 | Albuera Rural Health Unit | Albuera, Leyte | 8 | Leyte | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32396 | PBSP | mariaana.barte@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 267 |  | 6C GXIV-2D | 834818 | None | None | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6498848289116365681/6498848289116365681_signature_28.png | Maria Ana S. Barte | https://www.jotform.com/uploads/macarewecare/222927978372067/6498848289116365681/6498848289116365681_signature_31.png
| 2026-03-20 | 2026-03-18 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32395 | MACARE-HBDC | lcpmicrolab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 19 ASSAY: Xpert MTB-XDR, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX4 D 10C | 110027386 | As per end user they encounter undetected module | Checked I.P address and and input ip address Plug and unplug all cables Performed functionality test Cleaned plunger Test run unit | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6498166296619476738/6498166296619476738_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6498166296619476738/6498166296619476738_signature_31.png
| 2026-03-20 | 2026-03-17 | Casimiro A. Ynares Memorial Hospital | Becco St.,Rodriguez Rizal | 4A | Rizal | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32394 | PBSP | caysmhlab@gmail.com | 25 |  | 2026-03-16 | HX4 D | 831642 | Due for xpertcheck For Workstation Replacement 2UA8471W3L - CPH0625F59726 Module 4 loss communication | Delivered and replaced workstation Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck and workstation replacement | Workstationoperational Module 4 for reporting Module B1,B2,B3 passed xpertcheck |  | Workstation dellcomputer | 55434 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6498165786618269035/6498165786618269035_signature_28.png | Aileen A. Docto,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6498165786618269035/6498165786618269035_signature_31.png
| 2026-03-19 | 2026-03-19 | Sulat Rural Health Unit | Sulat Eastern Samar | 8 | Eastern Samar | TB laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32393 | PBSP | katherinelupos@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 757 | 2026-03-19 | GXIV - 2 Module 6C | 835100 | #NAME? | #NAME? | - Troubleshooting - Preventive maintenance - | #NAME? | Unit out of Warranty |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6497125006713051206/6497125006713051206_signature_28.png | Katherine Ann L. Alday | https://www.jotform.com/uploads/macarewecare/222927978372067/6497125006713051206/6497125006713051206_signature_31.png
| 2026-03-19 | 2026-03-19 | Aloguinsan Rural Health Unit | Poblacion | 7 | Cebu | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32392 | PBSP | aloguinsanrhulaboratory@gmail.com | 23 |  |  | GXIV 6C | 808620 | Module B2,B4 same error's code error's 2126,2014 | . Save System Log Report . Save Installation Qualifications Report. . Save assay Statistics Report | . Perform preventive maintenance. . Clean filter . Check cable wire for gateway board . Check voltage supply . Run sample test for module B2, B4. | . Module B1,B3 okay to use . Module B2 hardware fail during run sample . Module B4 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6497093673528722147/6497093673528722147_signature_28.png | Jennifer Sayson | https://www.jotform.com/uploads/macarewecare/222927978372067/6497093673528722147/6497093673528722147_signature_31.png
| 2026-03-19 | 2026-03-19 | Camarin Health Center | Camarin | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32391 | PBSP | jasminfalala@gmail.com | 25 |  |  | GXIV | 815395 | Unit for maintenance | Replaced filter | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6497073505617403318/6497073505617403318_signature_28.png | Arianne Jasmin Roma | https://www.jotform.com/uploads/macarewecare/222927978372067/6497073505617403318/6497073505617403318_signature_31.png
| 2026-03-19 | 2026-03-19 | Hagonoy Rural Health Unit | Poblacion, Hagonoy | 11 | Davao del Sur | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32390 | WHO | acresofrosemt@gmail.com | 29.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 |  | GX-IV 6C 4D | 847256 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck on Four Modules | Perform Preventive Maintenance and XpertCheck Calibration | Module B1/ SN: 876411 For Service Required Module B4/ SN: 910190 skipped and retest required all Cartidges avaliable was used Module B2: Passed Module B3: Retry and Passed | High Room Temperature, enduser was advised to achieve the ideal operating temperature before using the machine to avoid Errors |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6497062617119438891/6497062617119438891_signature_28.png | Arnel Cresar Nerosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6497062617119438891/6497062617119438891_signature_31.png
| 2026-03-19 | 2026-03-19 | Ipil Rural Health Unit | Ipil, Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32389 | PBSP | rheasolana@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 122 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1434 ASSAY: Others, COUNT: 21 | 2026-03-19 | GX 4 - 6c | 808615 | Module 4 gives temperature error of 1001, 1002, 4017. | Preventive Maintenence as PM checklist Document #D20843. Data collection. Site inspection | Preventive Maintenence as PM checklist Document #D20843. Data collection. Site inspection. | Module 1, 2 and 3 are working properly while module 4 is under observation because of frequent temperature error. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496996078216127625/6496996078216127625_signature_28.png | Rheamay Solana | https://www.jotform.com/uploads/macarewecare/222927978372067/6496996078216127625/6496996078216127625_signature_31.png
| 2026-03-19 | 2026-03-19 | Casiguran RHU | Casiguran, Aurora | 3 | Aurora | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  | harolduno0905@gmail.com |  |  | 32388 | PBSP | cezalen@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 771 | 2026-03-19 | GXIV | 831844 | Module A2: 757662 for replacement. For unit maintenance. | Performed module replacement and maintenance. | Replaced module A2 by module sn: 903921 Performed unit maintenance. | All 4 modules opwrational. | NA | Module SN: 903921 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496985910615327005/6496985910615327005_signature_28.png | Princess Alen A. Cabunoc, RMT,MPA | https://www.jotform.com/uploads/macarewecare/222927978372067/6496985910615327005/6496985910615327005_signature_31.png
| 2026-03-19 | 2026-03-18 | Ormoc City Health Office | Ormoc City, Leyte | 8 | Leyte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32387 | PBSP | theresesaladaga@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3021 ASSAY: Xpert HIV-1 Viral Load, COUNT: 151 |  | 6C GXIV-4D | 804253 | Module B2 failed XpertCheck | -Replace module B2/210089520 with module SN: 897494 | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496855567757538894/6496855567757538894_signature_28.png | Ma. Therese A. Saladaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6496855567757538894/6496855567757538894_signature_31.png
| 2026-03-18 | 2026-03-18 | Northern Samar Provincial Hospital | Catarman Northern Samar | 8 | Northern Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32386 | PBSP | maryanthonettedellosa@gmail.com | 25 |  | 2026-03-18 | GXIV 10C | 120000480 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496283886416464126/6496283886416464126_signature_28.png | Mary Anthonette Dellosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6496283886416464126/6496283886416464126_signature_31.png
| 2026-03-18 | 2026-03-18 | Northern Samar Provincial Hospital | Catarman Northern Samar | 8 | Northern Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32385 | PBSP | maryanthonettedellosa@gmail.com | 25 |  | 2026-03-18 | GXIV 10C | 120000480 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496283866415686404/6496283866415686404_signature_28.png | Mary Anthonette Dellosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6496283866415686404/6496283866415686404_signature_31.png
| 2026-03-18 | 2026-03-18 | Pedro Gil Health Center | Pedro Gil | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32384 | PBSP | acg_medtech@yahoo.com | 24 |  |  | GXIV | 812829 | Module B1 868124 for replacement | Replaced module B1 | - checked and cleaned cartridge bay, plunger rod, filter - replaced module B1 by 210057345 - generated reports - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496267990216324499/6496267990216324499_signature_28.png | Agnes Ranches | https://www.jotform.com/uploads/macarewecare/222927978372067/6496267990216324499/6496267990216324499_signature_31.png
| 2026-03-18 | 2026-03-17 | Rosario Super Health Center | Rosario, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32383 | PBSP | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 893 |  | GX-IV-D-6C | 843924 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221973616267005/6496221973616267005_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221973616267005/6496221973616267005_signature_31.png
| 2026-03-18 | 2026-03-17 | Santolan Super Health Center | Santolan, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32379 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1683 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221963614047218/6496221963614047218_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221963614047218/6496221963614047218_signature_31.png
| 2026-03-18 | 2026-03-17 | Santolan Super Health Center | Santolan, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32382 | PBSP | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1349 |  | GX-IV-D-6C | 838898 | For maintenance | performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221973614389332/6496221973614389332_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221973614389332/6496221973614389332_signature_31.png
| 2026-03-18 | 2026-03-18 | Taytay Rural Health Unit 5-B | Taytay, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32381 | PBSP | zenresurreccion@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 70 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1452 ASSAY: Xpert HIV-1 Viral Load, COUNT: 83 |  | GX-IV-D-6C | 838136 | Error 1001 on module B4/684235 Monitor auto adjustment and buttons not functioning | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -filter Generated reports Checked monitor | Unit is operational. Module B4 not operational, for reporting. Monitor buttons not functioning and auto adjustment always shows, recommend to replace. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221973617144302/6496221973617144302_signature_28.png | Zenilynn R. Bobiles | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221973617144302/6496221973617144302_signature_31.png
| 2026-03-18 | 2026-03-17 | Rosario Super Health Center | Rosario, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32380 | MACARE-HBDC | consalaneljaine@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1258 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221973619012882/6496221973619012882_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6496221973619012882/6496221973619012882_signature_31.png
| 2026-03-18 | 2026-03-18 | Calbayog City Health Office | Calbayog Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32378 | PBSP | Bogart.sd.1970@gmail.com | 27 |  | 2026-03-18 | GXIV | 804265 | N/a | N/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6496095549715765181/6496095549715765181_signature_28.png | ma.leonora dejapin | https://www.jotform.com/uploads/macarewecare/222927978372067/6496095549715765181/6496095549715765181_signature_31.png
| 2026-03-17 | 2026-03-17 | San Jose City General Hospital | Malasin, San Jose City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32377 | PBSP | katsorzadp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2004 |  | GXIV-6C | 804173 | Module B2 failed Xpertcheck | Performed Module Replacement B2(643773 - 745192) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495425229113539216/6495425229113539216_signature_28.png | Katherine Dela Pena | https://www.jotform.com/uploads/macarewecare/222927978372067/6495425229113539216/6495425229113539216_signature_31.png
| 2026-03-17 | 2026-03-17 | Amparo Health Center | Amparo | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32376 | PBSP | josephgescuetarmt@gmail.com | 24 |  |  | GXIV | 838873 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Valve related errors (1018, 2006, 4004) found on Module B2 (706384) B2 for reporting/follow up |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495379139911339865/6495379139911339865_signature_28.png | Joseph Gherico Navarro Escueta | https://www.jotform.com/uploads/macarewecare/222927978372067/6495379139911339865/6495379139911339865_signature_31.png
| 2026-03-17 | 2026-03-17 | Mariveles District Hospital | Mariveles Bataan | 3 | Batanes | Molecular Laboratory | NO | Pull-out | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32375 | MACARE-HBDC | laboratoryomdh2020@gmail.com | 25 |  |  | GXIV-Laptop 6C | 806377 | For pull-out | Preventive maintenance Pull-out | Collect Data Performaned preventive maintenance Performed functionality test Pull-out machine | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495372620187088106/6495372620187088106_signature_28.png | Marigae Monilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6495372620187088106/6495372620187088106_signature_31.png
| 2026-03-17 | 2026-03-17 | Abucay Rural Health Unit ( Roving GX) | Decena Ave Abucay Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32374 | PBSP | parreraledelina22176@gmail.com | 25 |  |  | GXIV-D 6C | 807466 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational D1: 685708 PASSED D2: 210068621 PASSED D3: 701939 FAILED D4: 706514 PASSED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495358590188128419/6495358590188128419_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/6495358590188128419/6495358590188128419_signature_31.png
| 2026-03-17 | 2026-03-17 | Davao de Oro Provincial Hospital - Montevista | Poblacion, Davao de Oro | 11 | Davao de Oro | TB-DOTS | YES | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32373 | PBSP | cecillejagurilles1122@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1448 |  | GX-IV 6C 4D | 808624 | Defective Module SN 730686/ A4 due to Multiple temperature errors Defective Module SN 210068367/ A1 due to XpertCheck Calibration Failure | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replaced the Defective Module SN: 730686/A4 with new Module SN: 892960 Replaced the Defective Module SN: 210068367 with new Module SN: 802376 | Perform Module Replacement Perform Preventive Maintenance | All Modules are operational, machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495319736218883593/6495319736218883593_signature_28.png | Cecille Jagurilles,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6495319736218883593/6495319736218883593_signature_31.png
| 2026-03-17 | 2026-03-17 | Norzagaray Rural Health Unit | Norzagaray | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32372 | DOH | chelseajempolicarpio@gmail.com | 27 |  |  | GXIV | 838166 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495319209913039620/6495319209913039620_signature_28.png | Chelsea Policarpio | https://www.jotform.com/uploads/macarewecare/222927978372067/6495319209913039620/6495319209913039620_signature_31.png
| 2026-03-17 | 2026-03-17 | Abucay Rural Health Unit | Decena Ave. Abucay Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32371 | PBSP | parreraledelina22176@gmail.com | 25 |  |  | GXIV-D 6C | 831878 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational D1: 667624 PASSED D2: 210029873 PASSED D3: 757993 PASSED D4: 784680 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495297108213727686/6495297108213727686_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/6495297108213727686/6495297108213727686_signature_31.png
| 2026-03-17 | 2026-03-17 | Angat Rural Health Unit | Angat, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32370 | PBSP | eileenrowenaicba@yaoo.com | 25 |  |  | GXIV | 814161 | Module door for troubleshooting | Adjusted module door | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495268379917154521/6495268379917154521_signature_28.png | Eileen Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/6495268379917154521/6495268379917154521_signature_31.png
| 2026-03-17 | 2026-03-17 | Biliran Provincial Hospital | Naval, Biliran | 8 | Biliran | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32369 | PBSP | ghelmission13@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1231 ASSAY: Xpert MTB-XDR, COUNT: 14 ASSAY: Xpert HIV-1 Viral Load, COUNT: 53 |  | 10C GXIV-4 | 110020264 | Module B1 and B3 defective | -Replace defective module B1/210103041 with module SN: 210041632 -Replace defective module B3/210103139 with module SN: 210040553 | #NAME? | Machine is functional. Module B2 is defective for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495222171126520699/6495222171126520699_signature_28.png | Gheland Corto | https://www.jotform.com/uploads/macarewecare/222927978372067/6495222171126520699/6495222171126520699_signature_31.png
| 2026-03-17 | 2026-03-17 | Hermosa Super Health Center | Palihan Hermosa, Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32368 | PBSP | kryztynangela@gmail.com | 26 |  |  | GXIV-D 6C | 831645 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational B1: 756769 PASSED B2: 756768 PASSED B3: 651999 PASSED B4: 756766 PASSED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6495209748214331041/6495209748214331041_signature_28.png | Mary Ann Patambag | https://www.jotform.com/uploads/macarewecare/222927978372067/6495209748214331041/6495209748214331041_signature_31.png
| 2026-03-16 | 2026-03-16 | Talisay City Health Office | Poblacion | 6 | Negros Occidental | TB Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32367 | PBSP | hadmeoffguard@gmail.com | 24 |  | 2026-03-16 | GXIV 6C | 847222 | Defective module B1 | . Save Installation Qualifications Report. . Save System Log Report | Replace module B1 with new serial number 210074046 | All modules are functional okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494561920395236966/6494561920395236966_signature_28.png | Hadney Besa | https://www.jotform.com/uploads/macarewecare/222927978372067/6494561920395236966/6494561920395236966_signature_31.png
| 2026-03-16 | 2026-03-16 | Cebu TB Reference Lab-SRP | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32366 | PBSP | asmantiquilla.pbsp@gmail.com | 26 |  |  | Gx4-10c | 120000133 | Defective Module B4 | New Module Installed B4 Clean Fan Filter | New Module Installed B4 Clean Fan Filter | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494505685899829033/6494505685899829033_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6494505685899829033/6494505685899829033_signature_31.png
| 2026-03-16 | 2026-03-16 | Cebu TB Reference Lab | Cebu City-SRP | 7 | Cebu | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32365 | PBSP | asmantiquilla.pbsp@gmail.com | 26 |  |  | Gx4-10C | 120000180 | Defective Module B1,B3 & B4 | New Module Installed B1,B3 & B4 Clean Fan Filter | New Module Installed B1,B3 & B4 Clean Fan Filter | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494501185898600143/6494501185898600143_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6494501185898600143/6494501185898600143_signature_31.png
| 2026-03-16 | 2026-03-16 | Manapla Rural Health Unit | Poblacion | 6 | Negros Occidental | TB Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32364 | PBSP | magalonaruderic@gmail.com | 24 |  |  | GXIV 6C | 805093 | Module A2 and A3 Defective. | . Save Installation Qualifications Report. . Save System Log Report. . GXX file | Replace module A2 with new serial number 801491 A3 with new serial number 679135 | All modules are functional okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494493718299776092/6494493718299776092_signature_28.png | Ruderic Magalona | https://www.jotform.com/uploads/macarewecare/222927978372067/6494493718299776092/6494493718299776092_signature_31.png
| 2026-03-16 | 2026-03-09 | Cordova Primary Health Care | Poblacion, Cordova Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | melodyjoynoelle0428@gmail.com |  | 32363 | PBSP | melodyjoynoelle0428@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 500 |  | GX4 | 808638 | Printer cannot proceed to print operation | Installation of printer driver Generate IQR, System Log, and Gxx files. Test print after installation. | Installation of printer driver Generate IQR, System Log, and Gxx files. Test print after installation. | Unit machine and its printer is operational. | Poblacion, Cordova Cebu |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494486235895063042/6494486235895063042_signature_28.png | Melody Arasan | https://www.jotform.com/uploads/macarewecare/222927978372067/6494486235895063042/6494486235895063042_signature_31.png
| 2026-03-16 | 2026-03-13 | Vicente Mendiola Center for Health Infirmary | Poblacion Naga City, Cebu | 7 | Cebu | Tb laboratory | YES | INSTALLATION TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 32362 | PBSP | geccaprocianos26@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 500 |  | GX4 | 838564 | Printer cannot proceed to print due driver corrupted. | Installation of printer driver Generate IQR, System Log, and Gxx files. Test print after installation. Load test assay such HIV, HCV and MTB Ultra | Installation of printer driver Generate IQR, System Log, and Gxx files. Test print after installation. | Unit machine and its printer is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494483965899798517/6494483965899798517_signature_28.png | Gene Carmel T..Procianos | https://www.jotform.com/uploads/macarewecare/222927978372067/6494483965899798517/6494483965899798517_signature_31.png
| 2026-03-16 | 2026-03-16 | Macare in house | 1005 Aurora Grand,Aurora blvd QC | NCR | NCR 2nd - Quezon City | Macare 8n house | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com | APC UPS | 32361 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  | 2026-03-16 | UPS | 4B1011PS3727 | Defective battery | Replaced and installed battery with s/n 548682 | Performed repair | Operational |  | Battery | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494426683613845889/6494426683613845889_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6494426683613845889/6494426683613845889_signature_31.png
| 2026-03-16 | 2026-03-16 | Macare in house | 1005 Aurora grand Aurora blvd QC | NCR | NCR 2nd - Quezon City | Macare in house | NO | REPAIR TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com | Accupower UPS | 32360 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  | 2026-03-16 | 1KVA ups | AT1000190123ST0255 | No power defective battery | Replaced and installed battery witn s/n 573259/573260 | Performed repair and troubleshooting | Operational |  | 12v 7ah battery | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494421583611252951/6494421583611252951_signature_28.png | Macare on house | https://www.jotform.com/uploads/macarewecare/222927978372067/6494421583611252951/6494421583611252951_signature_31.png
| 2026-03-16 | 2026-03-16 | Macare in house | 1005 Aurora grand ,Aurora blvd QC | NCR | NCR 2nd - Quezon City | Macare in house | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com | Accupower UPS | 32359 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  | 2026-03-16 | 1kva | At1000190123ST0160 | No power Defective battery | Replaced defective battery witn s/n 573257/573258 | Performed troubleshooting and repair | Operational |  | 12v 7ah battery | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494418543619503011/6494418543619503011_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6494418543619503011/6494418543619503011_signature_31.png
| 2026-03-16 | 2026-03-13 | Cotabato Regional and Medical Center | Sinsuat Avenue, Cotabato City | BARMM | Maguindanao | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32358 | PBSP | crmcmolecularlab@gmail.com | 25.6 | ASSAY: Xpert HBV VL, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3673 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 |  | GX-XVI 6C | 839307 | Module B3/ 717133 encounter multiple temperature related errors | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Module Troubleshooting and Preventive Maintenance | Module B3/ SN 717133 still encounter multiple temperature related errors | Errors sent to Cepheid For further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494409621312784126/6494409621312784126_signature_28.png | Joseph Romy Pascasio | https://www.jotform.com/uploads/macarewecare/222927978372067/6494409621312784126/6494409621312784126_signature_31.png
| 2026-03-16 | 2026-03-13 | Cotabato Regional and Medical Center | Sinsuat Avenue, Cotabato City | BARMM | Maguindanao | TB-DOTS | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32357 | PBSP | crmcmolecularlab@gmail.com | 25.7 | ASSAY: Xpert HBV VL, COUNT: 595 ASSAY: Xpert HPV, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 40 ASSAY: Xpert HIV-1 Viral Load, COUNT: 513 |  | GX-IV 6C 4D | 807607 | Stuck Cartidges at Modules 2 and 4 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Remove the Stuck Cartidges at Modules 2 and 4 | Perform Troubleshooting and extraction of stuck cartridges | Successfully removed the Stucked Cartidges All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494408991317307065/6494408991317307065_signature_28.png | Joseph Romy Pascasio | https://www.jotform.com/uploads/macarewecare/222927978372067/6494408991317307065/6494408991317307065_signature_31.png
| 2026-03-16 | 2026-03-13 | Ilocos Training and Regional Medical Center | San Fernando, La Union | 1 | La Union | Molecular Laboratory | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com | raalili.pbsp@gmail.com |  | 32356 | PBSP | itrmc.tblaboratory@gmail.com | 24 |  | 2026-03-13 | GXIV-4-D-6C R1 | 707524 | For Gateway Board Replacement | Performed Gatewayboard Replacement | Performed Gatewayboard Replacement | Unit is Partially Operational Modules B3 and B4 are disabled |  | Gatewayboard | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494386126615953033/6494386126615953033_signature_28.png | Russel Alili | https://www.jotform.com/uploads/macarewecare/222927978372067/6494386126615953033/6494386126615953033_signature_31.png
| 2026-03-16 | 2026-03-12 | Ilocos Sur Provincial Hospital Gabriela Silang | Vigan, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32355 | PBSP | mary.alexis.tahil@gmail.com | 25 |  | 2026-03-12 | GXIV-4-D-10C | 120000590 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B2 220013578 >>> 795939 - Generated Reports | Performed Module Replacement | Unit is Fully Operational |  | GeneXpert Module 795939 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494383966611182881/6494383966611182881_signature_28.png | Mary Alexis F. TAHIL | https://www.jotform.com/uploads/macarewecare/222927978372067/6494383966611182881/6494383966611182881_signature_31.png
| 2026-03-16 | 2026-03-12 | Bantay Municipal Health Office | Bantay, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32354 | PBSP | lornarabara_23@yahoo.com | 27 |  |  | GXIV-4-D-6C | 807585 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Partially Operational | Module 4 for Reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494382736613822805/6494382736613822805_signature_28.png | Lorna Rabara | https://www.jotform.com/uploads/macarewecare/222927978372067/6494382736613822805/6494382736613822805_signature_31.png
| 2026-03-16 | 2026-03-12 | Juan Posadas Health Center | Rodriguez St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32353 | PBSP | lizjunio67@gmail.com | 25 |  |  | GXIV | 805060 | Unit for preventive maintenance. For module replacement of module A4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. | Def Module (632590), New Module (672062) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494382563619159304/6494382563619159304_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/6494382563619159304/6494382563619159304_signature_31.png
| 2026-03-16 | 2026-03-11 | Ilocos Sur District Hospital- Narvacan | Narvacan, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | CALIBRATION MAINTENANCE XpertCheck | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32352 | PBSP | lfm11141977@gmail.com | 26 |  | 2026-03-11 | GXIV-4-D-6C | 807578 | For XpertCheck | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Reprogrammed Gateway board to orginal SN (910765241 >>> 807578) - Performed XpertCheck - Generated Reports | Performed XpertCheck | Unit is Partially Operational | Module D2 failed XpertCheck | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494380096617837906/6494380096617837906_signature_28.png | Lorely Miranda | https://www.jotform.com/uploads/macarewecare/222927978372067/6494380096617837906/6494380096617837906_signature_31.png
| 2026-03-16 | 2026-03-11 | Moonwalk Health Center | St.Francis, Moonwalk, Parañaque | NCR | NCR 4th - Parañaque | Laboratory | NO | Site Assessment | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32350 | PBSP | mayeemariel17@gmail.com | 25 |  |  | GXIV-D-6C | 805113 | GeneXpert Unit for Assessment. | Data Gathering. | Generate Reports. Performed Self-test. | B1,B2 and B3 are Operational. B4/ 724769 for Reporting. Check for Warranty Status. | Last XpertCheck: 10/27/2025 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494318323617227054/6494318323617227054_signature_28.png | Mariel Murillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6494318323617227054/6494318323617227054_signature_31.png
| 2026-03-16 | 2026-03-12 | UST Hospital | Dapitan St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32351 | COMMERCIAL | ustimp@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 766 ASSAY: Xpert BCR-ABL Ultra, COUNT: 22 ASSAY: Xpert HBV VL, COUNT: 16 |  | GXIV-D-6C | 110000711 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | B1 is Disabled. B2,B3 and B4 are Good. | Data Assessment for Purchasing of Warranty. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494318323618135687/6494318323618135687_signature_28.png | Maria Liova Sherleen Calaunan | https://www.jotform.com/uploads/macarewecare/222927978372067/6494318323618135687/6494318323618135687_signature_31.png
| 2026-03-16 | 2026-03-11 | Parañaque City Clinical and Drug Testing Laboratory | San Antonio Ave, Parañaque | NCR | NCR 4th - Parañaque | Laboratory | NO | Site Assessment | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32348 | PBSP | mariannelibot@gmail.com | 29 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 224 ASSAY: Xpert HPV, COUNT: 765 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1547 |  | GXIV-D-6C | 838192 | For Site Assessment. | Data Gathering. | Generate Reports. Performed Self-test. | B1, B2, B3 and B4 are Operational. Unit needs to be XpertCheck. | Last XpertCheck: 11/06/2024 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494318313614167160/6494318313614167160_signature_28.png | MARIANNE TIMTIM | https://www.jotform.com/uploads/macarewecare/222927978372067/6494318313614167160/6494318313614167160_signature_31.png
| 2026-03-16 | 2026-03-11 | San Isidro Health Center | UPS5, San Isidro, Parañaque | NCR | NCR 4th - Parañaque | Laboratory | NO | Site Assessment | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32349 | PBSP | herminadoa@yahoo.com | 28 | ASSAY: Xpert HPV, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1858 ASSAY: Xpert MTB-XDR, COUNT: 62 |  | GXIV-D-10C | 110020387 | For Site Assessment. | Data Gathering. | Generate Data for Reporting. Performed Self-test. | B1,B2 and B3 are in Good Condition. B4 for Observation with 2014 Error Code occurred last 09/30/2025. | Last XpertCheck: 11/20/2025 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494318313617157810/6494318313617157810_signature_28.png | Amelia G.Herminado | https://www.jotform.com/uploads/macarewecare/222927978372067/6494318313617157810/6494318313617157810_signature_31.png
| 2026-03-16 | 2026-03-12 | MMC Medilinx | Amorsolo st , Makati | NCR | NCR 4th - Makati | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32347 | COMMERCIAL | reinaluz.bactol@medilinx.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 314 ASSAY: Others, COUNT: 2 ASSAY: Others, COUNT: 2 ASSAY: Xpert CT/NG, COUNT: 803 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 118 ASSAY: Xpert HIV-1 Viral Load, COUNT: 151 ASSAY: Xpert HPV, COUNT: 489 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4375 ASSAY: Others, COUNT: 21 ASSAY: Xpert HCV VL, COUNT: 213 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1874 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 12 |  | GX16D 10C | 110026448 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494302763618746960/6494302763618746960_signature_28.png | Maria Reinaluz Bactol | https://www.jotform.com/uploads/macarewecare/222927978372067/6494302763618746960/6494302763618746960_signature_31.png
| 2026-03-16 | 2026-03-13 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 32346 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 59 ASSAY: Xpert CT/NG, COUNT: 83 ASSAY: Xpert HBV VL, COUNT: 632 ASSAY: Xpert HIV-1 Viral Load, COUNT: 363 ASSAY: Xpert HPV, COUNT: 376 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6484 ASSAY: Xpert MTB-XDR, COUNT: 5 ASSAY: Others, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 186 ASSAY: Xpert HIV-1 Viral Load, COUNT: 903 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 920 |  | GX-XVI-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -Icore -Filter Generated reports | Unit is operational. Module D1/210082300 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6494280143612721980/6494280143612721980_signature_28.png | Xchosa Kay C. Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/6494280143612721980/6494280143612721980_signature_31.png
| 2026-03-14 | 2026-03-14 | Felipe Abrigo Memorial Hospital | Guiuan Eastern Samar | 8 | Eastern Samar | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32345 | PBSP | famhlaboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 676 | 2026-03-14 | GXIV 6C | 815271 | Default CPU | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6492675833788314949/6492675833788314949_signature_28.png | Myla M. Aparejado | https://www.jotform.com/uploads/macarewecare/222927978372067/6492675833788314949/6492675833788314949_signature_31.png
| 2026-03-13 | 2026-03-13 | Bagong Silang Ph. 9 Health Center | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32344 | DOH | bsphase9hclaboratory@gmail.com | 27 |  |  | GXIV | 831677 | Module B2 for replacement | Replaced module B2 | - replaced module by 903458 - generated reports - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6492008388225640914/6492008388225640914_signature_28.png | Sheryl Dela Cerna | https://www.jotform.com/uploads/macarewecare/222927978372067/6492008388225640914/6492008388225640914_signature_31.png
| 2026-03-13 | 2026-03-13 | St. Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 32343 | MACARE-HBDC | laboratory@slec.ph | 23 |  | 2026-03-12 | Infinity 48s | 809643 | Module A3 for replacement due to hardware failure due to temperature error | Replaced module A3 (698176 > 903325) | Performed module replacement Performed maintenance cleaning | Unit is operational |  | Certified pre-owned GeneXpert Module (S/N: 903325) Modified ejector Assy door GeneXpert module | DR 13812 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491966320424809840/6491966320424809840_signature_28.png | Ma. Sarah Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6491966320424809840/6491966320424809840_signature_31.png
| 2026-03-13 | 2026-03-13 | Balabac RTDL | Poblacion 1, Balabac | 4B | Palawan | TB Lab | NO | MAINTENANCE | CEPHEID | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 32342 | MACARE-HBDC | kalim.jaycylmae@gmail.com | 20 |  |  | GXIV 6C | 832028 | Due fpr Xpertcheck | None | Performed PM Cleaned cartridge bays, icores and plunger rods Performed xpertcheck | Units is operational. A1 passed. A2, A3, and A4 not available |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491912392013060675/6491912392013060675_signature_28.png | Jaycyl Mae A. Kalim | https://www.jotform.com/uploads/macarewecare/222927978372067/6491912392013060675/6491912392013060675_signature_31.png
| 2026-03-13 | 2026-03-12 | Ospital ng Palawan | Malvar, Puerto Princesa City, Palawan | 4B | Palawan | TB Culture Lab | NO | MAINTENANCE | CEPHEID | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 32341 | PBSP | onptbclab@gmail.com | 24 |  |  | GXIV 10C | 120000182 | Due for Xpertcheck | None | Performed preventive maintenance Cleaned cartridge bays, icores, plunger rods, and replaced filter Performed xpertcheck Performed functionality test | Unit is operational. All modules passed Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491911692017828423/6491911692017828423_signature_28.png | Karizz Mae Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6491911692017828423/6491911692017828423_signature_31.png
| 2026-03-13 | 2026-03-13 | Hagonoy Rural Health Unit | Hagonoy | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32340 | PBSP | lanielyncastro@yahoo.com | 27 |  |  | GXIV | 807620 | Module C1(662453) and C2(668985) for replacement | Replaced modules | - replaced modules C1, C2 by 753916 and 756989 - generated reports - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491906148221823076/6491906148221823076_signature_28.png | Lanielyn Abadam | https://www.jotform.com/uploads/macarewecare/222927978372067/6491906148221823076/6491906148221823076_signature_31.png
| 2026-03-13 | 2026-03-13 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 32339 | COMMERCIAL | bloodbank@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. | Recommend to calibrate the unit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491905557519837062/6491905557519837062_signature_28.png | Roxanne mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/6491905557519837062/6491905557519837062_signature_31.png
| 2026-03-13 | 2026-03-13 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32336 | PBSP | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 759 |  | GXIV-D-10C | 110020361 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Run XpertCheck Generated reports | Unit is operational. All modules passed. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491897007512348714/6491897007512348714_signature_28.png | Girlie De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/6491897007512348714/6491897007512348714_signature_31.png
| 2026-03-13 | 2026-03-13 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32338 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 557 ASSAY: Xpert MTB-XDR, COUNT: 26 ASSAY: Xpert HIV-1 Viral Load, COUNT: 405 |  | GX-IV-D-10C | 110024073 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Wll modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491897007516839740/6491897007516839740_signature_28.png | Girlie De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/6491897007516839740/6491897007516839740_signature_31.png
| 2026-03-13 | 2026-03-13 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32337 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 515 ASSAY: Xpert MTB-XDR, COUNT: 36 ASSAY: Xpert HIV-1 Viral Load, COUNT: 193 |  | GX-IV-D-10C | 110024072 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491897007518954192/6491897007518954192_signature_28.png | Girlie De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/6491897007518954192/6491897007518954192_signature_31.png
| 2026-03-13 | 2026-03-13 | Mary Mediatrix Medical Center | Lipa, Batangas | 4A | Batangas | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomedmgrmi.mmmc2@gmail.com |  | 32335 | COMMERCIAL | mmmcpcrlab@yahoo.com | 25 |  |  | GXIV-10C | 120002889 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational and ready to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491886875612009902/6491886875612009902_signature_28.png | Irish Kristine Culla | https://www.jotform.com/uploads/macarewecare/222927978372067/6491886875612009902/6491886875612009902_signature_31.png
| 2026-03-13 | 2026-03-13 | San Jose Rural Health Unit | San Jose, Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32334 | DOH | sanjoserhulab@yahoo.com | 25 |  |  | GXIV | 821420 | Unit for preventive maintenance. For module replacement. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. | Def Module 220024256, New Module 620116 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491882995619399220/6491882995619399220_signature_28.png | Paul T. Tajanlamgit | https://www.jotform.com/uploads/macarewecare/222927978372067/6491882995619399220/6491882995619399220_signature_31.png
| 2026-03-13 | 2026-03-13 | San Remegio Rural Health Unit | Poblacion San Remegio Cebu | 7 | Cebu | TB Laboratory | NO | LCD Monitor Installation | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32333 | WHO | rodacanetecamay@gmail.com | 23 |  |  | GXIV 6C | 847333 | Installation of New LCD Monitor | . Save setup pecture . Save Installation Qualifications Report. | Replace the deffective LCD Install new LCD monitor woth serial number CN-01HRFC-FCC00-4CJ-DYPX-A00 | LCD functional okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491792814318216467/6491792814318216467_signature_28.png | Roda Camay | https://www.jotform.com/uploads/macarewecare/222927978372067/6491792814318216467/6491792814318216467_signature_31.png
| 2026-03-12 | 2026-03-12 | Gingoog City Health Office | Gingoog City, Misamis Oriental | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 32332 | PBSP | mabellagbas@gmail.com | 28 |  |  | GX-4 6 Colors | 807572 | Module B3 found defective. | Module Replacement | Replacing Module B3 with SN: 782563 to SN: 220001429 Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491120546223213730/6491120546223213730_signature_28.png | Mabel Lagbas | https://www.jotform.com/uploads/macarewecare/222927978372067/6491120546223213730/6491120546223213730_signature_31.png
| 2026-03-12 | 2026-03-12 | Cotabato Regional Medical Center-Blood Center | Cotabato City, Maguindanao | BARMM | Cotabato | Blood Center | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32331 | COMMERCIAL | bloodcenter@ro12.doh.gov.ph | 26.4 |  |  | MCS+ Heamonetics | 15H042G | None | Check and adjust optimal value of: Line Sensor SPM DPM Perfom cycle test for Actual Functionality | Calibration Preventive maintenance | Heamonetics machine is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491112928026304751/6491112928026304751_signature_28.png | Rhea Lim-Mancion, RMT, MSMLS | https://www.jotform.com/uploads/macarewecare/222927978372067/6491112928026304751/6491112928026304751_signature_31.png
| 2026-03-12 | 2026-03-12 | Teresa Rural Health Unit | Teresa, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32329 | DOH | teresa_rhu@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1174 |  | GX-IV-D-6C | 820710 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Performed Self test Generated reports | Unit is operational. All modules operational. All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491075018413158907/6491075018413158907_signature_28.png | Jia Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/6491075018413158907/6491075018413158907_signature_31.png
| 2026-03-12 | 2026-03-12 | Baras Rural Health Unit | Baras, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32330 | PBSP | mitchz0323@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 454 |  | GX-IV-D-6C | 815389 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports Checked old CPU | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491075018416528989/6491075018416528989_signature_28.png | Michelle R. Mendoza-Avela | https://www.jotform.com/uploads/macarewecare/222927978372067/6491075018416528989/6491075018416528989_signature_31.png
| 2026-03-12 | 2026-03-12 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32328 | PBSP | laboratory@loveyourself.ph | 27 |  |  | GXIV | 808649 | Module B4 for replacement | Replaced module B4 | - replaced module B4(635671) by 210008410 - generated reports - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491068218225101522/6491068218225101522_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6491068218225101522/6491068218225101522_signature_31.png
| 2026-03-12 | 2026-03-12 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32327 | DOH | laboratory@loveyourself.ph | 27 |  |  | GXIV | 831895 | Unit for maintenance | - checked and cleaned cartridge bay, plunger rod, filter, icore slot - performed XpertCheck - replaced module A2 (756790) by 210017852 - generated reports | Module A2 for replacement Unit for calibration | Unit is operational A1: Passed A2: Newly replaced A3: Passed A4: Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491066278224230908/6491066278224230908_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6491066278224230908/6491066278224230908_signature_31.png
| 2026-03-12 | 2026-03-09 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank donor room | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32326 | MACARE-HBDC | blooddonorsection@gmail.com | 26 |  |  | MCS + | 15H044G | As per user they encounter lid lock problem unble to lock properly | Checked unit Checked and adjust screw Performed functionality test Test run unit ok | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491058187547467198/6491058187547467198_signature_28.png | Daniel Gabriel Querijero | https://www.jotform.com/uploads/macarewecare/222927978372067/6491058187547467198/6491058187547467198_signature_31.png
| 2026-03-12 | 2026-03-12 | Liloan RHu | Liloan | 7 | Cebu | Laboratory | NO | MAINTENANCE Gather data Gxx file | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32325 | PBSP | vickcabahugmt@gmail.com | 25 |  |  | Gx4-2 module | 835189 | N/A | Gather data Nxx file IQR and System log | Gather data Nxx file IQR and System log | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491035304617008995/6491035304617008995_signature_28.png | Vick Csbahug | https://www.jotform.com/uploads/macarewecare/222927978372067/6491035304617008995/6491035304617008995_signature_31.png
| 2026-03-12 | 2026-03-12 | Tukuran rural health unit | Barangay curved, tukuran, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 32324 | PBSP | donnamarielanzaderas@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 307 | 2026-03-12 | GX 4 - 6c | 805121 | Module 1 code 2120 communication error | Preventive Maintenence, data collection and repaired module 1. | Preventive Maintenence, data collection and repaired module 1. | All modules are available while module 1 is under observation. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6491018180756694906/6491018180756694906_signature_28.png | Donna Marie D. Lanzaderas, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6491018180756694906/6491018180756694906_signature_31.png
| 2026-03-12 | 2026-03-12 | Puerto Princesa City Health Office | Santa Monica, Puerto Princesa City, Palawan | 4B | Palawan | Lab | NO | MAINTENANCE | CEPHEID | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 32323 | MACARE-HBDC | jocal0324@gmail.com | 23 |  |  | GXIV 6C | 832026 | Due for Xpertcheck | None | Performed preventive maintenace Cleaned cartridge bays, icores, and replaced filter Performed Xpertcheck | All modules are available but modules b1, b2, and b3 are service required while module b4 is skipped and retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490971780818346262/6490971780818346262_signature_28.png | Jocelyn Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/6490971780818346262/6490971780818346262_signature_31.png
| 2026-03-12 | 2026-03-12 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32322 | PBSP | rose.guerrero597@gmail.com | 27 |  |  | GXIV | 805062 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational A1: Failed A2: Passed A3: Passed A4: Retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490954188228853092/6490954188228853092_signature_28.png | Rosemarie Guerrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6490954188228853092/6490954188228853092_signature_31.png
| 2026-03-12 | 2026-03-12 | Zamboanga del sur medical center | Zamboanga del sur medical center, pagadian city | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32321 | DOH | limboemillie@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3797 ASSAY: Xpert HIV-1 Viral Load, COUNT: 98 | 2026-03-12 | Gx 4 - 6c | 838204 | None | Checked genexpert machine gatewayboard revision number. Preventive Maintenance | Preventive Maintenence and data collection requested by ceiphied. | All modules are working properly. All modules are operational. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490902580753158187/6490902580753158187_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/6490902580753158187/6490902580753158187_signature_31.png
| 2026-03-12 | 2026-03-12 | Balamban Rural Health Unit | Poblacion Balamban Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32320 | DOH | climaco.mildred@yahoo.com | 23 |  |  | GXIV 6C | 838231 | Module A2 serial number 806987 hardware fail. | . Save System Log Report. . Save Installation Qualifications Report. . Assay statistics Report. . Save GXX file. | Perform preventive maintenance for GeneXpert machine. | Module A1,A3,A4 okay to use. Module A2 Hardware fail not okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490898570011996520/6490898570011996520_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6490898570011996520/6490898570011996520_signature_31.png
| 2026-03-11 | 2026-03-06 | Santa Rosa City Health Office 2 | Santa Rosa Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32319 | DOH | annie.soliman@yahoo.com | 25 |  |  | GXIV | 821156 | Unit for preventive maintenance. For module replacement. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. | Def Module 662327, New Module 210103817 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490483257424848977/6490483257424848977_signature_28.png | Annie Soliman | https://www.jotform.com/uploads/macarewecare/222927978372067/6490483257424848977/6490483257424848977_signature_31.png
| 2026-03-11 | 2026-03-05 | Talisay Rural Health Unit | Talisay Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32318 | DOH | rodriguezadett@gmail.com | 25 |  |  | GXIV | 820201 | Unit for preventive maintenance. For module replacement. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. | Def Module 721276, New Module 210021426 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490481657422402003/6490481657422402003_signature_28.png | Adettlou A. Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/6490481657422402003/6490481657422402003_signature_31.png
| 2026-03-11 | 2026-03-05 | Rosario Rural Health Unit | Rosario, Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32317 | DOH | liberty_yutuc@yahoo.com | 25 |  |  | GXIV | 821421 | Unit for preventive maintenance. For module replacement. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. | Def Module 784137, New Module 687705 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490480807426883501/6490480807426883501_signature_28.png | Liberty I. Yutuc | https://www.jotform.com/uploads/macarewecare/222927978372067/6490480807426883501/6490480807426883501_signature_31.png
| 2026-03-11 | 2026-03-11 | Ospital ng Makati | Sampaguita St Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 32316 | MACARE-HBDC | jinx1gs@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 739 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490271835731024937/6490271835731024937_signature_28.png | Cristina Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6490271835731024937/6490271835731024937_signature_31.png
| 2026-03-11 | 2026-03-11 | Chinese General Hospital | Blumentritt St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 32315 | MACARE-HBDC | cghmcbloodbank2016@gmail.com | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490264784811783700/6490264784811783700_signature_28.png | Katrina mae pineda | https://www.jotform.com/uploads/macarewecare/222927978372067/6490264784811783700/6490264784811783700_signature_31.png
| 2026-03-11 | 2026-03-11 | Chinese General Hospital | Blumentritt, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 32314 | COMMERCIAL | sabclaridad@gmail.com | 25 |  |  | GXIV-6C | 835604 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490225544813573041/6490225544813573041_signature_28.png | Sandra Abigail Claridad | https://www.jotform.com/uploads/macarewecare/222927978372067/6490225544813573041/6490225544813573041_signature_31.png
| 2026-03-11 | 2026-03-11 | Chinese General Hospital | Blumentritt St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 32313 | COMMERCIAL | sabclaridad@gmail.com | 25 |  |  | GXIV-10C | 120002026 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490225544815036845/6490225544815036845_signature_28.png | Sandra Abigail Claridad | https://www.jotform.com/uploads/macarewecare/222927978372067/6490225544815036845/6490225544815036845_signature_31.png
| 2026-03-11 | 2026-03-11 | Vicente Mendiola Center for Health Infirmary | Poblacion Naga Cebu | 7 | Cebu | Tb Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 32312 | PBSP | geccaprocianos26@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 500 |  | Gx4 | 838564 | Cpu Problem | Install the cpu after recovery of Operating System | Format and Operating System recovery | Unit cpu is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490223018024065425/6490223018024065425_signature_28.png | Gene Carmel T..Procianos | https://www.jotform.com/uploads/macarewecare/222927978372067/6490223018024065425/6490223018024065425_signature_31.png
| 2026-03-11 | 2026-03-11 | San Remigio Rural Health Unit | Poblacion San Remigio Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32311 | WHO | rodacanetecamay@gmail.com | 25 |  |  | GXIV 6C | 847333 | Due for preventive Maintenance. | . Save Installation Qualifications Report . Save Assay Statistics Report . Save system log report. | . Replace new filter of GX machine . Clean all modules inside the machine. . Clean all wire's of GX machine. | All modules are functional okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490113153814950556/6490113153814950556_signature_28.png | Roda Camay | https://www.jotform.com/uploads/macarewecare/222927978372067/6490113153814950556/6490113153814950556_signature_31.png
| 2026-03-11 | 2026-03-11 | Bogo City Health Office | Bogo City, Cebu | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32310 | PBSP | bogocholab@gmail.com | 25 |  |  | Gx4-10c | 120000494 | Due date for xpertcheck calibration | Xpertcheck calibration Gather data IQR,system log,assay stats And calibration data | Xpertcheck calibration Gather data IQR,system log,assay stats And calibration data | Systems operational Module 1&3 passed Module 2&4 retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490108064615999243/6490108064615999243_signature_28.png | Reggie V Dahili | https://www.jotform.com/uploads/macarewecare/222927978372067/6490108064615999243/6490108064615999243_signature_31.png
| 2026-03-11 | 2026-03-11 | Makati City NTP Office (From West Rembo) | Makati City | NCR | NCR 4th - Makati | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | ntpmakati@gmail.com |  | 32309 | DOH | kapid5sepe@yahoo.com | 24 |  |  | GX-IV-D-6C | 831712 | Boot device not found | Replaced BIOS battery | Replaced BIOS battery Change boot options Configure secure boot Functionality test | Unit is operational. UPS not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490102673426460101/6490102673426460101_signature_28.png | Marco Anonas | https://www.jotform.com/uploads/macarewecare/222927978372067/6490102673426460101/6490102673426460101_signature_31.png
| 2026-03-11 | 2026-03-11 | Makati City NTP Office (From East Rembo) | Makati City Hall | NCR | NCR 4th - Makati | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | ntpmakati@gmail.com |  | 32308 | PBSP | kapid5sepe@yahoo.com | 22 |  |  | GX-IV-D-6C | 807591 | Bios battery replacement and CPU not powering on sometimes | Troubleshoot CPU | Replaced BIOS battery Reseat internal connections Checked CPU | Unit is operational but under observation due to CPU not powering on sometimes. UPS not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490101943422185164/6490101943422185164_signature_28.png | Marco Anonas | https://www.jotform.com/uploads/macarewecare/222927978372067/6490101943422185164/6490101943422185164_signature_31.png
| 2026-03-11 | 2026-03-02 | Medellin Rural Health Unit | Poblacion Medellin Cebu | 7 | Cebu | Tb laboratory | NO | TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | rhumedellin@gmail.com |  | 32307 | DOH | rhumedellin@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 600 |  | GX4 | 838506 | CPU defective no power. | Need to replace the CPU Unit of the GeneXpert machine | Troubleshoot CPU power supply and motherboard. Power supply no output voltage. | Unit CPU is out of warranty need for quotation of the unit. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490074345895449733/6490074345895449733_signature_28.png | Raul Galano | https://www.jotform.com/uploads/macarewecare/222927978372067/6490074345895449733/6490074345895449733_signature_31.png
| 2026-03-11 | 2026-03-11 | TANJAY CITY HEALTH OFFICE | Tanjay City, Negros Oriental | 6 | Negros Oriental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 32306 | PBSP | skyskynemueco@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 780 | 2026-03-11 | GX-IV | 808621 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM XPERTCHECK CALIBRATION -MODULE 2 RETRY THEN FOR SERVICE -MODULE 1,3,4 PASSED | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK-CE-5 LOT 1001495845 2027-08-15 | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6490060245778800080/6490060245778800080_signature_28.png | Resurrecion F.Mueco | https://www.jotform.com/uploads/macarewecare/222927978372067/6490060245778800080/6490060245778800080_signature_31.png
| 2026-03-11 | 2026-03-10 | Taguig Pateros District Hospital | Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32304 | DOH | ierlanegail@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 132 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3335 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1210 | 2026-03-10 | GX-IV-D-6C | 831638 | Moduel A1/691267 and A2/681512 for replacement | Replaced module A1 and A2 | Replaced module A1 and A2 Performed maintenance Performed self test Generated reports | Unit is operational. All modules operational. A3 and A4 has error 2120 and 2037, for reporting. |  | Old:A1/291267, A2/681512 New: A1/210035807, A2/210052069 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6489981904644655755/6489981904644655755_signature_28.png | Erlanegail Isla | https://www.jotform.com/uploads/macarewecare/222927978372067/6489981904644655755/6489981904644655755_signature_31.png
| 2026-03-11 | 2026-03-10 | Department of Health - Treatment and Rehabilitation Center | Camp Bagong Diwa, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32305 | PBSP | dohtrcbicutan.lab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1748 |  | GXIV-D-6C | 807452 | For maintenance Error 2005 and 2008 in all modules Broken stand | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -icore -Filter Generated reports | Unit is operational. All modules operational but under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6489981904646231500/6489981904646231500_signature_28.png | April Chakee A. Javier | https://www.jotform.com/uploads/macarewecare/222927978372067/6489981904646231500/6489981904646231500_signature_31.png
| 2026-03-10 | 2026-03-09 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City tb | 12 | South Cotabato | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32303 | PBSP | renlizros11@gmail.com | 22.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GxIV-6C | 815428 | Due for Xpertcheck | Perform Genexpert annual Xpertcheck Perform Genexpert machine Cleaning and maintenance Module cleaning Fan and filter cleaning Plunger rod and Optics cleaning Checking of module status and errors Data collection | Genexpert machine Xpertcheck Calibration Genexpert machine Preventive maintenance | 3 operational modules Xpertcheck result: Module C1 SN:696269 -Requiring for Service Module C2,C3 and C4 -Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6489545057326007528/6489545057326007528_signature_28.png | Renee Liz Gomez | https://www.jotform.com/uploads/macarewecare/222927978372067/6489545057326007528/6489545057326007528_signature_31.png
| 2026-03-10 | 2026-03-10 | Soccsksargen General Hospital | Surallah, South Cotabato | 12 | South Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32302 | PBSP | sghlab21@gmail.com | 24.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1925 ASSAY: Xpert MTB-XDR, COUNT: 22 |  | GxIV-10C | 120000172 | Module 3 for Service Required | Perform Genexpert machine Cleaning and maintenance Module cleaning Fan and filter cleaning Plunger rod and Optics cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6489522887326689149/6489522887326689149_signature_28.png | SARAH MAE AGUILAR | https://www.jotform.com/uploads/macarewecare/222927978372067/6489522887326689149/6489522887326689149_signature_31.png
| 2026-03-10 | 2026-03-09 | Cordova RHU Primary Health Care | Poblacion Naga City Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 32301 | PBSP | melodyjoynoelle0428@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 200 |  | GX 4 | 808638 | Module 2 SN 882343 due for replacement due to error 2025 home position error | Replaced new module with new module SN # 750060 | Replaced new module with new module SN # 750060 | Unit modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6489512435895450575/6489512435895450575_signature_28.png | Melody Arasan | https://www.jotform.com/uploads/macarewecare/222927978372067/6489512435895450575/6489512435895450575_signature_31.png
| 2026-03-10 | 2026-03-10 | Kolambugan Rural Health Unit | Kolambugan, Lanao del Norte | 13 | Lanao del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 32300 | PBSP | yojecinaj@gmail.com | 22 |  |  | GX 4 6 Colors | 815272 | Module B1,B2 and B4 found Service Required. | Nine | XpertCheck Calibration Preventive Maintenance | 1 module working 3 modules for replacement due to xpc result |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6489237937702887099/6489237937702887099_signature_28.png | Janice Joy R. Siso | https://www.jotform.com/uploads/macarewecare/222927978372067/6489237937702887099/6489237937702887099_signature_31.png
| 2026-03-10 | 2026-03-10 | Perpetual Soccour Hospital | Cebu City | 7 | Cebu | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 32299 | COMMERCIAL | pshclinlab@gmail.com | 26 |  |  | Mcs+ | 02D167 | Possoble mixing og plasma and red cell | Calibrate and Preventive Maintenance | Calibrate and Preventive Maintenance | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6489170933423077552/6489170933423077552_signature_28.png | Debbie Joyce Baga | https://www.jotform.com/uploads/macarewecare/222927978372067/6489170933423077552/6489170933423077552_signature_31.png
| 2026-03-10 | 2026-03-10 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32298 | PBSP | cdh.laboratorydept@gmail.com | 25 |  |  | GXIV-D 6C | 807562 | C3:210068087 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performaned Module replacement Gathered reports Performaned functionality test | Unit is operational C3: 745839 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6489115437225930349/6489115437225930349_signature_28.png | Innah Coleen David | https://www.jotform.com/uploads/macarewecare/222927978372067/6489115437225930349/6489115437225930349_signature_31.png
| 2026-03-09 | 2026-02-27 | Tunasan Health Center | E.Rodriguez St, Tunasan, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32297 | PBSP | gtrazothc@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1417 |  | GXIV-D-6C | 807613 | Modules 2 and 4 Service Required in XpertCheck. | Module Replacement. | Replaced Modules 2 and 4. Performed Self-test. Generate Reports. | B2: 657230 replaced with 210071466. B4: 662484 replaced with 210071561. Modules 1 and 3 are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488553883616212910/6488553883616212910_signature_28.png | Geraldine T. Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/6488553883616212910/6488553883616212910_signature_31.png
| 2026-03-09 | 2026-03-03 | Capitol Medical Center | Quezon Ave., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32295 | COMMERCIAL | laboratory@capitolmedical.com.ph | 25 |  |  | MCS+ | 07C002 | Bowl Breakage during Procedure. | Preventive Maintenance. | Cleaned Centrifuge Chamber, Pumps, and Sensors. Checked Sensors Value. Performed Simulation Test. | Unit is Good. All Values in Range of Normal Operation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488548263612254090/6488548263612254090_signature_28.png | Eryll Joy Agojo | https://www.jotform.com/uploads/macarewecare/222927978372067/6488548263612254090/6488548263612254090_signature_31.png
| 2026-03-09 | 2026-03-06 | General Emilio Aguinaldo Memorial Hospital | Trece-Indang Road, Trece Martires, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32296 | PBSP | djaemiel@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 554 |  | GXIV-L-6C | 802860 | Unit for Gateway Board Replacement. | Repair. | Replaced Gateway Board. Run Software to Change Random SN. Performed Self-test. Generate Reports. | All Modules are Good. New Gateway Board. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488548263612376437/6488548263612376437_signature_28.png | Jaemiel Juvirens A. Dimaranan | https://www.jotform.com/uploads/macarewecare/222927978372067/6488548263612376437/6488548263612376437_signature_31.png
| 2026-03-09 | 2026-03-04 | San Dionisio Health Center | Pelaez St, San Dionisio, Parañaque | NCR | NCR 4th - Parañaque | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32294 | DOH | marubi1019@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 39 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1593 |  | GXIV-D-6C | 831898 | Module 1 is Defective and needs to be Replacement . | Module Replacement. Preventive Maintenance. | Replaced Module 1. Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Performed Self-test. Generate Reports. | Module 1: 210008304 -> 210030476 Modules 2, 3 and 4 are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488548263615440927/6488548263615440927_signature_28.png | Fatima Joyce Tresvalles | https://www.jotform.com/uploads/macarewecare/222927978372067/6488548263615440927/6488548263615440927_signature_31.png
| 2026-03-09 | 2026-03-05 | Taal Rural Health Unit | Carsuche, Taal, Batangas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32293 | PBSP | taalrhulaboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2178 |  | GXIV-D-6C | 805259 | Module 2 is Defective for Replacement. | Module Replacement. Preventive Maintenance. | Replaced Defective Module with New Module. Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Replaced Filter. Performed Self-test. Generate Reports. | Modules 1,3 and 4 are Good. Module 2: 647538(Defective) -> 657307(New). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488548263611092193/6488548263611092193_signature_28.png | Juanita Balbacal | https://www.jotform.com/uploads/macarewecare/222927978372067/6488548263611092193/6488548263611092193_signature_31.png
| 2026-03-09 | 2026-03-09 | Philippine Red Cross Tower | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | REPAIR | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32292 | MACARE-HBDC | bloodcollection.tower@gmail.com | 25 |  | 2026-03-05 | MCS+ | 01J038 | CPU Board for replacement | Replaced CPU board Checked calibration parameters | #NAME? | Unit is operational |  | MCS+ CPU Board Assembly SN: C111252810 | 13791 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488505413512152256/6488505413512152256_signature_28.png | ardel brecinio | https://www.jotform.com/uploads/macarewecare/222927978372067/6488505413512152256/6488505413512152256_signature_31.png
| 2026-03-09 | 2026-03-09 | Dauin Primary Care Facility | Dauin, Negros Oriental | 7 | Negros Oriental | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 32291 | PBSP | dkannahmae@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 326 |  | GX4 6c | 835059 | Module B1 Hardware Fail. Module B1 SN 0, Temp 320.0 System Date/time problem. | . | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488504342686638729/6488504342686638729_signature_28.png | kannah mae dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/6488504342686638729/6488504342686638729_signature_31.png
| 2026-03-09 | 2026-03-06 | Hilongos Rural Health Unit | Hilongos, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE Data Gather | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32290 | PBSP | leovillaflor@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 724 |  | 6C GXIV-4D | 838143 | None | None | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488496445894805749/6488496445894805749_signature_28.png | Leonides Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/6488496445894805749/6488496445894805749_signature_31.png
| 2026-03-09 | 2026-03-06 | Lucena City Health Office | Lucena City Government Center | 4A | Quezon | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32288 | PBSP | czarinapalillo@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1766 | 2026-03-06 | GX-IV-D-6C | 807462 | For xpertCheck Module D2 for replacement | Replaced module D2 Run xpertCheck | Performed maintenance Replaced module D2 Run xpertCheck Performed self test Generated reports | Unit is operational. All modules operational. Module D2 newly replaced not XpertChecked. |  | Old: D2/776975 New: D2/210106501 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323613587657/6488464323613587657_signature_28.png | Czarina Palillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323613587657/6488464323613587657_signature_31.png
| 2026-03-09 | 2026-03-06 | Sariaya Rural Health Unit | Sariaya, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32289 | DOH | smcl_2012@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1869 | 2026-03-06 | GX-IV-D-6C | 831890 | Module A1 for replacement | Replaced module A1 | Performed maintenance Replaced module A1 Performed self test Generated reports | Unit is operational. All modules operational. |  | Old: A1/758224 New: A1/65799 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323613774019/6488464323613774019_signature_28.png | Marie Linnet Abril | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323613774019/6488464323613774019_signature_31.png
| 2026-03-09 | 2026-03-05 | Calauag Municipal Hospital | Calauag, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32287 | DOH | jamila17carmel@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 513 | 2026-03-05 | GX-IV-D-6C | 831910 | Module A4 for replacement | Replaced module A4 | Performed maintenance Replaced module A4 Performed self test Generated reports | Unit is operational. All modules operational. Monitor is lagging and has a line on upper part. |  | Old: A4/758384 New: A4/633391 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323617302148/6488464323617302148_signature_28.png | Jamila Pedrezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323617302148/6488464323617302148_signature_31.png
| 2026-03-09 | 2026-03-05 | Quezon Medical Center | Lucena, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32285 | PBSP | mariannenaynes081774@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 ASSAY: Xpert HPV, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2224 | 2026-03-05 | GX-IV-D-6C | 804692 | Module A4 for replacement | Replaced module A4 | Performed maintenance Replaced module A4 Performed self test Generated reports | Unit is operational. All modules operational. |  | Old: A4/210070786 New: A4/220031976 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323617926369/6488464323617926369_signature_28.png | Marianne Naynes | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323617926369/6488464323617926369_signature_31.png
| 2026-03-09 | 2026-03-05 | Quezon Medical Center | Lucena, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32286 | PBSP | mariannenaynes081774@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1929 ASSAY: Xpert MTB-XDR, COUNT: 119 |  | GX-IV-D-10C | 10019627 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323618596032/6488464323618596032_signature_28.png | Marianne Naynes | https://www.jotform.com/uploads/macarewecare/222927978372067/6488464323618596032/6488464323618596032_signature_31.png
| 2026-03-09 | 2026-03-09 | West Dadiangas Rural Health Unit | Brgy.West Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32284 | DOH | renlizros11@gmail.com | 23.4 |  |  | GxIV-6C | 838570 | Due for Xpertcheck | Perform Genexpert annual Xpertcheck Perform Genexpert machine Cleaning and maintenance Module cleaning Fan and filter cleaning Plunger rod and Optics cleaning Checking of module status and errors Data collection | Genexpert machine Xpertcheck Genexpert machine Preventive maintenance | 4 operational modules Xpertcheck Results: All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488437554658638568/6488437554658638568_signature_28.png | Renee Liz Gomez | https://www.jotform.com/uploads/macarewecare/222927978372067/6488437554658638568/6488437554658638568_signature_31.png
| 2026-03-09 | 2026-03-05 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32283 | UNOPS | spmcrhivda@gmail.com | 27.8 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Xpert MTB-XDR, COUNT: 17 ASSAY: Xpert SARS-CoV-2, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1352 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2011 |  | GX-IV 10C 4D | 110023505 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488381660125565619/6488381660125565619_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6488381660125565619/6488381660125565619_signature_31.png
| 2026-03-09 | 2026-03-05 | Southern Philippines Medical Center | Bajada Davao City | 11 | Davao del Sur | RHIVDA | NO | REPAIR MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32282 | UNOPS | spmcrhivda@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1461 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1368 ASSAY: Xpert SARS-CoV-2, COUNT: 19 | 2026-02-11 | GX-IV 10C 4D | 110023506 | Module B3 SN:210122278 / For Service Required Due for Preventive Maintenance | Replaced the Defective Module SN:210122278 with new Module SN 614401 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | Machine is operational Module B2/ SN: 210122382 encounter temp related errors and reportrd for further assessment |  | 900-0650RS | DR 13738 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488381240122409869/6488381240122409869_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6488381240122409869/6488381240122409869_signature_31.png
| 2026-03-09 | 2026-03-05 | Southern Philippines Medical Center | Bajada Davao City | 11 | Davao del Sur | RHIVDA | NO | REPAIR MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32281 | UNOPS | spmcrhivda@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1461 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1368 ASSAY: Xpert SARS-CoV-2, COUNT: 19 | 2026-02-11 | GX-IV 10C 4D | 110023506 | Module B3 SN:210122278 / For Service Required Due for Preventive Maintenance | Replaced the Defective Module SN:210122278 with new Module SN 614401 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | Machine is operational Module B2/ SN: 210122382 encounter temp related errors and reportrd for further assessment |  | 900-0650RS | DR 13738 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488381110124851723/6488381110124851723_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6488381110124851723/6488381110124851723_signature_31.png
| 2026-03-09 | 2026-03-04 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32280 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 27 | ASSAY: Xpert CT/NG, COUNT: 12 ASSAY: Xpert HBV VL, COUNT: 472 ASSAY: Xpert HIV-1 Viral Load, COUNT: 39 ASSAY: Xpert HPV, COUNT: 371 ASSAY: Xpert MTB/RIF Ultra, COUNT: 496 ASSAY: Xpert HCV VL, COUNT: 37 ASSAY: Xpert HIV-1 Viral Load, COUNT: 330 |  | GX4 L 6C | 807155 | As per end user they encounter temp error 1002 | Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed troubleshooting | C1,C2 operational C3 for observation C4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488329813614661280/6488329813614661280_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6488329813614661280/6488329813614661280_signature_31.png
| 2026-03-09 | 2026-03-04 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | YES | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32279 | PBSP | roseannparnes@yahoo.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2167 |  | GX16 | 803087 | For module replacement C4 714498 - 210031101 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed module replacement | C4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488328953617647265/6488328953617647265_signature_28.png | Rose Ann Parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/6488328953617647265/6488328953617647265_signature_31.png
| 2026-03-09 | 2026-03-06 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32278 | MACARE-HBDC | roseannpatnes@yahoo.com | 25 |  |  | GX4 10C | 120007667 | For installation | Delivered and installed unit Performed functionality test Complete accessories Upload ADF files Test run unit | Performed installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6488328363613193080/6488328363613193080_signature_28.png | Rose Ann Parnes-Le | https://www.jotform.com/uploads/macarewecare/222927978372067/6488328363613193080/6488328363613193080_signature_31.png
| 2026-03-07 | 2026-03-05 | Batac City Health Office | Batac, Ilocos Norte | 1 | Ilocos Norte | TB DOTS | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32277 | PBSP | bataccholaboratory@gmail.com | 30 |  | 2026-03-05 | GXIV-4-D-6C | 805272 | For Gatewayboard Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Gatewayboard 805272 >>> 910767956 - Programmed New Gatewayboard with Old SN 910767 >>> 805272 - Generated | Performed Gatewayboard Replacement | - Unit is Partially Operational - Module D2 is currently Disabled; Already reported in Jotform |  | Gatewayboard Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6486506532325882015/6486506532325882015_signature_28.png | JOY E. RACACHO | https://www.jotform.com/uploads/macarewecare/222927978372067/6486506532325882015/6486506532325882015_signature_31.png
| 2026-03-07 | 2026-03-05 | Mariano Marcos Memorial Hospital and Medical Center | Batac, Ilocos Norte | 1 | Ilocos Norte | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32276 | PBSP | mmmh.molecularlab@gmail.com | 30 |  | 2026-03-05 | GXIV-4-D-6C | 831908 | For XpertCheck | #NAME? | Performed XpertCheck | Unit is Fully Operational | All Modules passed XpertCheck | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6486504172321728570/6486504172321728570_signature_28.png | Allan Fellizar | https://www.jotform.com/uploads/macarewecare/222927978372067/6486504172321728570/6486504172321728570_signature_31.png
| 2026-03-07 | 2026-03-05 | Bacarra Rural Health Unit | Bacarra, Ilocos Norte | 1 | Ilocos Norte | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32275 | PBSP | davidangelogamidobarrozo@gmail.com | 25 |  | 2026-03-05 | GXIV-4-D-6C | 807430 | For XpertCheck | #NAME? | Performed XpertCheck | Unit is Fully Operational | All Modules passed XpertCheck | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6486503002322505619/6486503002322505619_signature_28.png | David Angelo Barrozo | https://www.jotform.com/uploads/macarewecare/222927978372067/6486503002322505619/6486503002322505619_signature_31.png
| 2026-03-07 | 2026-03-03 | Infanta Rural Health Unit | Infanta, Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32274 | PBSP | charissealmandres8@gmail.com | 28 |  | 2026-03-03 | GXIV-4-D-6C | 804726 | For XpertCheck | #NAME? | Performed XpertCheck | Unit is Fully Operational | All Modules passed XpertCheck | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6486501122321683043/6486501122321683043_signature_28.png | Charisse A. Alferez | https://www.jotform.com/uploads/macarewecare/222927978372067/6486501122321683043/6486501122321683043_signature_31.png
| 2026-03-07 | 2026-03-06 | UPD The Marine Science Institute | P. Velasquez St., Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Algal Lab | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 32273 | COMMERCIAL | scgonzaga@up.edu.ph | 25 |  |  | Stemi 305 Trino with Axiocam ERc5s | 3947017112 | Enduser stated equipment not powering-on | None, suggest for monitoring | Check-up and Cleaning | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6486400973713930554/6486400973713930554_signature_28.png | Shienna Mae Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6486400973713930554/6486400973713930554_signature_31.png
| 2026-03-06 | 2026-03-06 | Los Banos Rural Health Unit | Los Banos Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32272 | PBSP | rheamarie019@gmail.com | 25 |  |  | GXIV | 804716 | Unit for preventive maintenance. For module replacement. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. | Def module 210091372, New Module 701588 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6485925950424652318/6485925950424652318_signature_28.png | Rhea Marie Dorado | https://www.jotform.com/uploads/macarewecare/222927978372067/6485925950424652318/6485925950424652318_signature_31.png
| 2026-03-06 | 2026-03-06 | Magalang Super Health Center | 13th st. Guanzon Sbd. Sta. Cruz, Magalang Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32271 | PBSP | bevlozano29@gmail.com | 25 |  |  | GXIV-D 6C | 846044 | B3: 905016 for Module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed Module replacement Gathered reports Performed functionality test | Unit is operational B3: 210114859 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6485847949812589190/6485847949812589190_signature_28.png | Beverly Lozano | https://www.jotform.com/uploads/macarewecare/222927978372067/6485847949812589190/6485847949812589190_signature_31.png
| 2026-03-06 | 2026-03-06 | Vicente Mendiola Center for Health Infirmary | Poblacion Naga Cebu | 7 | Cebu | Tb Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 32270 | PBSP | geccaprocianos26@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 800 |  | Gx4 | 838564 | CPU problem boot system problem | Pull out cpu for hardisk testing. | #NAME? | CPU is for pull out. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6485824398028127483/6485824398028127483_signature_28.png | GENE CARMEL T. PROCIANOS | https://www.jotform.com/uploads/macarewecare/222927978372067/6485824398028127483/6485824398028127483_signature_31.png
| 2026-03-06 | 2026-03-06 | Angat Rural Health Unit | Angat, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32269 | PBSP | eileenrowenaicban@yahoo.com | 28 |  |  | GXIV | 814161 | Module B4 door for replacement | Replaced module door | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6485822620275797784/6485822620275797784_signature_28.png | Eileen Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/6485822620275797784/6485822620275797784_signature_31.png
| 2026-03-06 | 2026-03-06 | Ilagan City Health Office | Ilagan City, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32268 | PBSP | cho1ilagan.laboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2442 |  | GXIV-6C | 808635 | Genexpert PC restarts while running test | Troubleshoot PC Cleaned Ram Replaced Bios Battery | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6485784590027581879/6485784590027581879_signature_28.png | Hannah Jean Sanggo | https://www.jotform.com/uploads/macarewecare/222927978372067/6485784590027581879/6485784590027581879_signature_31.png
| 2026-03-06 | 2026-03-06 | Lubao Rural Health Unit | Lubao Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32267 | PBSP | nathaliecoleencovar@gmail.com | 25 |  |  | GXIV-D | 812834 | GatewayBoard replacement Defective gateway board sn:812834 | Preventive maintenance GatewayBoard replacement | Performed surface cleaning Gathered reports Perfomed GatewayBoard replacement Performed functionaity test | Unit is operational New gateway board:910767928 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6485751289816258035/6485751289816258035_signature_28.png | Nathalie Covar | https://www.jotform.com/uploads/macarewecare/222927978372067/6485751289816258035/6485751289816258035_signature_31.png
| 2026-03-06 | 2026-03-06 | Gamu Rural Health Unit | Gamu, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32266 | PBSP | Gamugeneralclinicallaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1225 |  | GXIV-6C | 814373 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Performed Genexpert Assessment | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6485694600029783648/6485694600029783648_signature_28.png | Precious Alma Maravilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6485694600029783648/6485694600029783648_signature_31.png
| 2026-03-06 | 2026-03-06 | Dinalupihan Rural Health Unit | Brgy. Culo Dinalupihan Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32265 | PBSP | villanuevarevienel@gmail.com | 25 |  |  | GXIV-D 6C | 807561 | B4:720070 for replacement | Preventive maintenance Module replacement | Performed surface cleaning Gatahered Pre IQR Perforemd Module replacement Gathered post IQR Performed functionality test | Unit is operational New Module SN: 661112 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6485667099817583370/6485667099817583370_signature_28.png | Revienel Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6485667099817583370/6485667099817583370_signature_31.png
| 2026-03-05 | 2026-03-05 | Cabagan Rural Health Unit | Poblacion Centro, Cabagan, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32264 | PBSP | rhuCLcabagan2022@gmail.com | 26 |  |  | GXIV-6C | 805104 | Due for Calibration | Performex Xpertcheck (Module 807263/B2 - Service Required) (Module B1, B3 & B4 - Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6485054020026968852/6485054020026968852_signature_28.png | Rosemarie Peca-at | https://www.jotform.com/uploads/macarewecare/222927978372067/6485054020026968852/6485054020026968852_signature_31.png
| 2026-03-05 | 2026-03-05 | Western Visayas Medical Center | Ilo-ilo City | 6 | Iloilo | TB Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32263 | PBSP | meobordo.pbsp@gmail.com | 23 |  | 2026-03-05 | GX16 | 839316 | Due for Xpertcheck. | . Save System Log Report . Save Installation Qualifications Report . Save Xpertcal file. | Perform Xpertcheck for all modules. | A1, A3, A4 - Fail for Xpertcheck C1, C3 - Fail for Xpertcheck D4 - fail for Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484998686614558352/6484998686614558352_signature_28.png | Melnia Andrea Obordo | https://www.jotform.com/uploads/macarewecare/222927978372067/6484998686614558352/6484998686614558352_signature_31.png
| 2026-03-05 | 2026-03-05 | Cagayan Valley Center for Health Development | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32262 | PBSP | mctorio.pbsp@gmail.com | 18 |  |  | GXIV-6C | 834434 | Due for Maintenance | Cleaned Icore Cleaned Cartridge Bay Cleaned Plunger Rod Cleaned Filters | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484923670028667154/6484923670028667154_signature_28.png | Mikaela Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/6484923670028667154/6484923670028667154_signature_31.png
| 2026-03-05 | 2026-03-05 | Cagayan Valley Center for Health Development | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32261 | PBSP | mctorio.pbsp@gmail.com | 18 |  |  | GXIV-10C | 120000178 | Due for Maintenance | Cleaned Icore Cleaned Cartridge Bay Cleaned Filters Cleaned Plunger Rod | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484923553795074021/6484923553795074021_signature_28.png | Mikaela Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/6484923553795074021/6484923553795074021_signature_31.png
| 2026-03-05 | 2026-03-05 | Cagayan Valley Center for Health Development | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32260 | PBSP | mctorio.pbsp@gmail.com | 17 |  |  | GXIV-6C | 808628 | Module C2 Failed Xpertcheck Gateway board for Replacement | Module SN: 210063113 Replaced with Module SN: 210011539 Replaced Gateway Board Programmed Gateway Board | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484919363799453207/6484919363799453207_signature_28.png | Mikaela Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/6484919363799453207/6484919363799453207_signature_31.png
| 2026-03-05 | 2026-03-05 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32259 | UNOPS | spmcrhivda@gmail.com | 27.8 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Xpert MTB-XDR, COUNT: 17 ASSAY: Xpert SARS-CoV-2, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1352 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2011 |  | GX-IV 10C 4D | 110023505 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484904062702897301/6484904062702897301_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6484904062702897301/6484904062702897301_signature_31.png
| 2026-03-05 | 2026-03-05 | Southern Philippines Medical Center | Bajada Davao City | 11 | Davao del Sur | RHIVDA | NO | REPAIR MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32258 | UNOPS | spmcrhivda@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1461 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1368 ASSAY: Xpert SARS-CoV-2, COUNT: 19 | 2026-02-11 | GX-IV 10C 4D | 110023506 | Module B3 SN:210122278 / For Service Required Due for Preventive Maintenance | Replaced the Defective Module SN:210122278 with new Module SN 614401 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | Machine is operational Module B2/ SN: 210122382 encounter temp related errors and reportrd for further assessment |  | 900-0650RS | DR 13738 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484903852705329252/6484903852705329252_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6484903852705329252/6484903852705329252_signature_31.png
| 2026-03-05 | 2026-03-05 | Botolan Rural Health Unit | Baton Lapoc Botolan, Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32257 | PBSP | abby.peboyjr@gmail.com | 29 |  |  | GXIV-D 6C | 805070 | B4:640579 for module replacement | Preventive maintenance Module replacement | Performance preventive maintenance Performed module replacement Performed functionality test | Unit is operational Defective module SN: 645079 New Module SN: 705270 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484892638815520786/6484892638815520786_signature_28.png | Abigail V. Jaring | https://www.jotform.com/uploads/macarewecare/222927978372067/6484892638815520786/6484892638815520786_signature_31.png
| 2026-03-05 | 2026-03-05 | Southern Philippines Medical Center | Bajada Davao City | 11 | Davao del Sur | RHIVDA | NO | REPAIR MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32256 | UNOPS | spmcrhivda@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1461 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1368 ASSAY: Xpert SARS-CoV-2, COUNT: 19 | 2026-02-11 | GX-IV 10C 4D | 110023506 | Module B3 SN:210122278 / For Service Required UPS Not working Due for Preventive Maintenance | Replaced the Defective Module SN:210122278 with new Module SN 614401 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replaced the Defective UPS SN: AT1K231111ST0024 with new SN: 310026737FA3728000436 | Perform Preventive Maintenance, UPS and Module Replacement | Machine is operational Module B2/ SN: 210122382 encounter temp related errors and reported for further assessment |  | 900-0650RS | DR 13738 | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/6484882042706960277/6484882042706960277_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6484882042706960277/6484882042706960277_signature_31.png
| 2026-03-05 | 2026-03-05 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32255 | UNOPS | spmcrhivda@gmail.com | 27.8 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Xpert MTB-XDR, COUNT: 17 ASSAY: Xpert SARS-CoV-2, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1352 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2011 |  | GX-IV 10C 4D | 110023505 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484882042707985119/6484882042707985119_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6484882042707985119/6484882042707985119_signature_31.png
| 2026-03-05 | 2026-03-05 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32254 | PBSP | spmcrhivda@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 237 ASSAY: Xpert HIV-1 Viral Load, COUNT: 140 |  | GX-IV 6C 4D | 808298 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules functional Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484877272705896892/6484877272705896892_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6484877272705896892/6484877272705896892_signature_31.png
| 2026-03-05 | 2026-03-05 | Passi City Health Office | Poblacion Passi City | 6 | Iloilo | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32253 | PBSP | passicholab@gmail.com | 23 |  | 2026-03-05 | GX IV 10C | 110013965 | Module B1 Hardware fail serial number 210067154 | . Save Installation Qualifications Report. . Save System Log Report . Save assay statistics Report | Replace module B1 with new module serial number 721585 | All modules are functional okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484870716614440610/6484870716614440610_signature_28.png | Kaycee Rhee Pormilos | https://www.jotform.com/uploads/macarewecare/222927978372067/6484870716614440610/6484870716614440610_signature_31.png
| 2026-03-05 | 2026-03-05 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32252 | MACARE-HBDC | ruellamorena@gmail.com | 29 |  |  | GXIV-6C | 110008457 | Due for Maintenance Modules High Temperature | Plugged in the Fan connector | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484837919287260955/6484837919287260955_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/6484837919287260955/6484837919287260955_signature_31.png
| 2026-03-05 | 2026-03-05 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32251 | PBSP | ruellamorena@gmail.com | 28 |  |  | GX XVI-6C | 842894 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Performed Machine assessment | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484837180023967203/6484837180023967203_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/6484837180023967203/6484837180023967203_signature_31.png
| 2026-03-04 | 2026-03-04 | Fatima Rural Health Unit | Brgy. FATIMA, General Santos City | 12 | South Cotabato | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32249 | PBSP | dglemao9087@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 71 |  | GxIV-6C | 805007 | Defective CPU | Perform Genexpert machine CPU replacement Perform Genexpert Annual Xpertcheck Perform Genexpert machine cleaning and maintenance Data collection | Genexpert machine replacement Genexpert machine Xpertcheck Calibration Genexpert machine Preventive maintenance | 3 operational modules Xpertcheck result Module 1 and 3 -Passed Module 2 skipped and retest required Module 4 requiring for Service |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484250977322697534/6484250977322697534_signature_28.png | Dante Ariel A. Glemao | https://www.jotform.com/uploads/macarewecare/222927978372067/6484250977322697534/6484250977322697534_signature_31.png
| 2026-03-04 | 2026-03-04 | Dr.Jorge P. Royeca Hospital | General Santos City | 12 | South Cotabato | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32250 | PBSP | elsigh76@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 455 ASSAY: Xpert HIV-1 Viral Load, COUNT: 188 | 2026-03-04 | GxIV-6C | 838130 | Module B3 defective module | Replace defective module B3 SN:806717 with new module SN:676837 Perform Genexpert machine Cleaning and maintenance Fan and filter cleaning Module cleaning Plunger rod and Optics cleaning Data collection | Genexpert machine module repair Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert 6-color module SN:676837 | 319226 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484250977329355519/6484250977329355519_signature_28.png | Elsie R. Dacua | https://www.jotform.com/uploads/macarewecare/222927978372067/6484250977329355519/6484250977329355519_signature_31.png
| 2026-03-04 | 2026-03-04 | St lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiologu | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 32248 | COMMERCIAL | megapo@stlukes.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 225 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2233 |  | GX4 6C | 839234 | As per end user they encounter 5007 in module B2. B1,B3,B4 confirmed for replacement | Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed troubleshooting | B2 for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484200431592240115/6484200431592240115_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6484200431592240115/6484200431592240115_signature_31.png
| 2026-03-04 | 2026-03-04 | Bagumbong Health Center | Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32247 | DOH | jenilyn_duenas@yahoo.com | 25 |  |  | GXIV | 831907 | Module A2 (210045625) for replacement | Replaced module A2 Replaced filter | - replaced module A2 by 210036060 - checked and cleaned cartridge bay, plunger rod, filter - replaced filter - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484153689198052364/6484153689198052364_signature_28.png | Bea Francisco | https://www.jotform.com/uploads/macarewecare/222927978372067/6484153689198052364/6484153689198052364_signature_31.png
| 2026-03-04 | 2026-03-04 | Bondoc Peninsula District Hospital | Catanauan, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32246 | PBSP | bpdhrtdl@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1366 | 2026-03-04 | GX-IV-D-6C | 805339 | Module B4 for replacement | Replaced module B4 | Performed maintenance Replaced module B4 Performed self test Generated reports | Unit is operational. All modules operational. |  | Old: B4/210068502 New: B4/643575 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484149491503342670/6484149491503342670_signature_28.png | Trinity carisse puyal | https://www.jotform.com/uploads/macarewecare/222927978372067/6484149491503342670/6484149491503342670_signature_31.png
| 2026-03-04 | 2026-03-04 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32245 | PBSP | vincentmilestone20@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2793 ASSAY: Xpert HIV-1 Viral Load, COUNT: 182 ASSAY: Xpert MTB-XDR, COUNT: 63 ASSAY: Xpert SARS-CoV-2, COUNT: 1107 |  | GX-IV 6C 4D | 805063 | For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Install XPERTCHECK Version 1.5B Run XpertCheck on 4 Modules | Perform Preventive Maintenance and XpertCheck Calibration | Modules B1 and B2 Retry and Passed Moduels B3 and B4 Passed All Modules Passed the Calibration Machine is Operational and Functional | Pull out extra XPC Calibration Cartidge |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484130951286114313/6484130951286114313_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/6484130951286114313/6484130951286114313_signature_31.png
| 2026-03-04 | 2026-03-02 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular | YES | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32244 | PBSP | vincentmilestone20@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2793 ASSAY: Xpert MTB-XDR, COUNT: 63 ASSAY: Xpert HIV-1 Viral Load, COUNT: 182 |  | GX-IV 6C 4D | 834440 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484108521286774120/6484108521286774120_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/6484108521286774120/6484108521286774120_signature_31.png
| 2026-03-04 | 2026-03-02 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Lab | YES | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32243 | PBSP | vincentmilestone20@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2793 ASSAY: Xpert MTB-XDR, COUNT: 63 ASSAY: Xpert HIV-1 Viral Load, COUNT: 182 ASSAY: Xpert SARS-CoV-2, COUNT: 1107 |  | GX-IV 10C 4D | 110020276 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484108081281674765/6484108081281674765_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/6484108081281674765/6484108081281674765_signature_31.png
| 2026-03-04 | 2026-03-04 | Margosatubig Regional Hospital | Margosatubig, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32242 | PBSP | clairemorales62@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 48 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2166 ASSAY: Xpert MTB-XDR, COUNT: 32 ASSAY: Others, COUNT: 3 | 2026-03-04 | GX 4 - 10c | 110019600 | Module 4 hardware failed. | Preventive Maintenence done as per PM checklist Document #D20843 Replaced defective module 210099163 with new module 210018099 as per Document #301-0569 rev A. Decontamination done as per Documen# 07203 | Preventive Maintenence, Replaced module 4 with new module. | All modules are available and operational. Genexpert machine is ready to use. | None | 900-0650RS | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484080154124184309/6484080154124184309_signature_28.png | Claire Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/6484080154124184309/6484080154124184309_signature_31.png
| 2026-03-04 | 2026-03-04 | Davao Regional Medical Center - BloodBank | Apokon, Tagum City | 11 | Davao del Norte | Blood Bank | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 32241 | COMMERCIAL | bloodbank@drmc.gov.ph | 25.6 |  |  | MCS+ | 21M389SPG | For Preventive Maintenance, and Calibration | Clean Surface Area Check the roller assay pumps Check and set the optimal values for Bowl Optics, SPM & DPM Sensors, Weigher and Line Sensors Run Cycle Test for functionality Test | Perform Preventive Maintenance, Calibration | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484076755161463587/6484076755161463587_signature_28.png | Mae Ann Leyrita | https://www.jotform.com/uploads/macarewecare/222927978372067/6484076755161463587/6484076755161463587_signature_31.png
| 2026-03-04 | 2026-03-04 | Limay Rural Health Unit | Pagcor Bldg. Limay Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32240 | PBSP | nelsonchavez40771@gmail.com | 25 |  |  | GXIV-D 6C | 805124 | Gateway board replacement GatewayBoard SN: 805124 | Repair/ gateway board replacement | Performed surface cleaning Gathered reports Performed Gateway board replacement Changed Gateway board SN FROM 910767922 to 805124 Performed functionality test | Unit is operational Gateway board defective SN: 805124 to New SN 910767922 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6484048261917893786/6484048261917893786_signature_28.png | Maria Lorena Ginauli | https://www.jotform.com/uploads/macarewecare/222927978372067/6484048261917893786/6484048261917893786_signature_31.png
| 2026-03-04 | 2026-03-04 | San Francisco Rural Health Unit | San Francisco, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32239 | DOH | annlouhpalma20@gmail.com | 24 |  | 2026-03-04 | GX-IV-D-6C | 831698 | Module A3 for replacement | Replaced module A3 | Performed maintenance Replaced module A3 Performed self test Generated reports | Unit os operational. All modules operational. |  | Old: A3/615114 New: A3/705087 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6483993473015008717/6483993473015008717_signature_28.png | Ann Lorraine Palma | https://www.jotform.com/uploads/macarewecare/222927978372067/6483993473015008717/6483993473015008717_signature_31.png
| 2026-03-04 | 2026-03-04 | San Isidro Rural Health Unit | San Isidro Northern Samar | 8 | Northern Samar | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32238 | UNOPS | hsdelamitas59@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 484 | 2026-03-04 | GXIV - 10C | 110023515 | Module B4 Defective SN:210121815 | - Replaced defective module B4/210121815 new Module B4/886791 - Archived Data - Data Collection - Performed functionality test | #NAME? | Unit is Operational |  | 1 Genexpert 10C module SN: 886791 | 13753 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6483976768711672922/6483976768711672922_signature_28.png | Hazel G. Salinas | https://www.jotform.com/uploads/macarewecare/222927978372067/6483976768711672922/6483976768711672922_signature_31.png
| 2026-03-04 | 2026-03-04 | Dr. Jose Rizal Memorial Hospital | Dapitan | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32237 | PBSP | calumbamendel@gmail.com | 26 |  |  | Gx 4 | 110020389 | Defective Module B3 | Install new module B3 SN 210106257 Gather data IQR, System log, assay stats | Install new module B3 SN 210106257 Gather data IQR, System log, assay stats | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6483956072928694418/6483956072928694418_signature_28.png | Mendel Q. CALUMBA | https://www.jotform.com/uploads/macarewecare/222927978372067/6483956072928694418/6483956072928694418_signature_31.png
| 2026-03-03 | 2026-03-03 | Manila Medical Center | U.N. Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32236 | CONSORTIUM | microbiologylab.manilamed@gmail.com | 25 |  |  | GXIV | 835540 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6483310136611466651/6483310136611466651_signature_28.png | Frances Levina Velarde | https://www.jotform.com/uploads/macarewecare/222927978372067/6483310136611466651/6483310136611466651_signature_31.png
| 2026-03-03 | 2026-03-03 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32235 | MACARE-HBDC | rhwclab@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 417 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3278 |  | GX-IV 6C 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log Archived Data test results | Perform Preventive Maintenance | Only Three Modules are operational Module B3 SN 866834 encounter multiple temperature related errors | Module B3 subjected for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6483231207163102650/6483231207163102650_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6483231207163102650/6483231207163102650_signature_31.png
| 2026-03-03 | 2026-03-03 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32234 | DOH | rhwclab@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 417 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3278 |  | GX-IV 6C 4D | 838262 | For Preventive Maintenance and Onsite Assistance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Troubleshooting | All. Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6483227897167673014/6483227897167673014_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6483227897167673014/6483227897167673014_signature_31.png
| 2026-03-03 | 2026-03-03 | Taguig Social Hygiene Clinic Drop-in Center | East Service Rd. Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32233 | MACARE-HBDC | nikko.eugenio26@gmail.com | 22 |  |  | GX-IV-D-10C | 120007666 | For installation | Installed the GeneXpert System | Installed the GeneXpert System -GeneXpert Unit -Workstation -Printer -Barcode scanner -UPS Functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6483221349146035784/6483221349146035784_signature_28.png | Nikko Eugenio | https://www.jotform.com/uploads/macarewecare/222927978372067/6483221349146035784/6483221349146035784_signature_31.png
| 2026-03-03 | 2026-03-03 | Taguig Social Hygiene Clinic | Col P Cruz, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32232 | MACARE-HBDC | tshcdic@gmail.com | 24 |  |  | GX-IV-D-10C | 120007669 | For installation | Installed the GeneXpert System | Installed the GeneXpert System: -GeneXpert Unit -Workstation -Printer -UPS -Barcode scanner Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6483212809146647410/6483212809146647410_signature_28.png | Lazo, Stephanie Lorraine | https://www.jotform.com/uploads/macarewecare/222927978372067/6483212809146647410/6483212809146647410_signature_31.png
| 2026-03-03 | 2026-03-03 | Salvacion Oppus-Yñiguez Memorial Provincial Hospital | Maasin City, Southern Leyte | 8 | Southern Leyte | GeneXpert Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32231 | PBSP | montillacherry2@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 777 ASSAY: Xpert MTB-XDR, COUNT: 38 |  | 10C GXIV-4D | 110013959 | XpertCheck Due | #NAME? | #NAME? | Machine is functional. Module B1 for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6483063828599479213/6483063828599479213_signature_28.png | Cherry Oja | https://www.jotform.com/uploads/macarewecare/222927978372067/6483063828599479213/6483063828599479213_signature_31.png
| 2026-03-03 | 2026-03-02 | Mindanao Central sanitarium General Hospital | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module and gateway board replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32230 | PBSP | ybioamie@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1236 | 2026-03-02 | Gx 4 - 6c | 821214 | Module 1, Module 2 and gateway board due for replacement. | Replaced module A1, A2 and gateway board. Preventive Maintenance. Data collection. Defective modules decontamination. | Preventive Maintenence as per PM checklist Document #D20843. Replaced defective modules 909139 and 744189 with new modules SN 865399 and 805020 respectively as per Document #301-0569 Rev A. Replaced Gateway board 700-3183 Rev F with new gateway board 700-6140S, Rev C as per Document #301-0569 Rev A. Decontamination done as per Document #07203. | All modules are operational and ready to use. | New modules SN (865399), (805020) and gateway board SN 910765525. | 900-0491rs, 700-6140s | 0054096 / 0055023 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482774035307583384/6482774035307583384_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/6482774035307583384/6482774035307583384_signature_31.png
| 2026-03-02 | 2026-02-20 | Estevez Memorial Hospital Inc. | Rizal st, Legazpi | 5 | Albay | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32229 | COMMERCIAL | rllaneta666@gmail.com | 25 |  |  | GX4-10C | 110017942 | Gx result loss communication error (all modules) | Checked UPS, CPU connection to power source Check power supply of UPS | Performed Troubleshooting | Unit is operational Defective UPS | UPS drained battery(request for replacement) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482701947701203243/6482701947701203243_signature_28.png | Roldan Llaneta | https://www.jotform.com/uploads/macarewecare/222927978372067/6482701947701203243/6482701947701203243_signature_31.png
| 2026-03-02 | 2026-02-19 | Tinambac Rural Health Unit | Poblacion, Tinambac | 5 | Camarines Sur | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32228 | PBSP | allancaballero623@gmail.com | 25 |  |  | GX4-6C | 808285 | Gx for Calibration Gx for Preventive Maintenance Date & Time reset | XpertCheck Result: B1/210100891 - service required B2/651554 - passed B3/699314 - passed B4/652128 - passed Cleaned cartridge bays, plunger rod, icore Replaced CMOS Bios Battery reset/adjusted date and time Generate Data | Performed XpertCheck Performed Preventive Maintenance Replaced Bios Battery | Unit is Operational | B1/210100891 - service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482698467708241563/6482698467708241563_signature_28.png | Allan F. Caballero | https://www.jotform.com/uploads/macarewecare/222927978372067/6482698467708241563/6482698467708241563_signature_31.png
| 2026-03-02 | 2026-02-13 | Tabaco City Health Unit | Taihi, Tabaco | 5 | Albay | Laboratory | YES | Relocation | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32227 | PBSP | elyjoy17rmt@gmail.com | 24 |  |  | GX4-10C | 110019623 | Relocation of Gx set | Relocation of Gx set to new nearby lab(building is at the back of the previous facility) Reboxing and reinstalling of Gx set Generate Data | Unboxing and Reinstalling of Gx set | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482669457708322937/6482669457708322937_signature_28.png | Ely joy | https://www.jotform.com/uploads/macarewecare/222927978372067/6482669457708322937/6482669457708322937_signature_31.png
| 2026-03-02 | 2026-02-11 | Bicol Medical Center | Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32226 | MACARE-HBDC | btsc@bmc.doh.gov.ph | 25 |  |  | MCS+ | 15H027G | Unit for Preventive Maintenance | Cleaned Unit surface Cleaned centrifuge, valve, rotor pump Checked parameters | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482664457709915406/6482664457709915406_signature_28.png | Sergs Bryan S Bausas | https://www.jotform.com/uploads/macarewecare/222927978372067/6482664457709915406/6482664457709915406_signature_31.png
| 2026-03-02 | 2026-02-27 | Benguet Provincial Health Office | La Trinidad, Benguet | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32225 | PBSP | benguetphortdl@gmail.com | 30 |  | 2026-02-27 | GXIV-4-D-6C | 807621 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B1 210091156 >>> 760173 - Generated Reports | Performed Module Replacement | Unit is Fully Operational |  | GeneXpert Module 760173 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482587128818667076/6482587128818667076_signature_28.png | Tillber Baniaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6482587128818667076/6482587128818667076_signature_31.png
| 2026-03-02 | 2026-02-25 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32224 | UNOPS | romarmtamayo@gmail.com | 24 |  |  | GXIV-4-D-6C | 110023501 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482584738816335076/6482584738816335076_signature_28.png | James elisha b. Balolong | https://www.jotform.com/uploads/macarewecare/222927978372067/6482584738816335076/6482584738816335076_signature_31.png
| 2026-03-02 | 2026-02-26 | Mangatarem District Hospital | Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32223 | PBSP | sweetsellveras1@gmail.com | 30 |  |  | GXIV-4-D-6C | 807598 | For XpertCheck | ‐ Cleaned Cartridge Bays, Plunger Rods and I-CORE ‐ Performed XpertCheck ‐ Generated Reports | Performed XpertCheck | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482583928819255280/6482583928819255280_signature_28.png | Kimberly C. Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/6482583928819255280/6482583928819255280_signature_31.png
| 2026-03-02 | 2026-03-02 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32222 | PBSP | vincentmilestone20@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4255 ASSAY: Xpert SARS-CoV-2, COUNT: 819 ASSAY: Xpert HIV-1 Viral Load, COUNT: 536 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 6 |  | GX-XVI 16D | 838756 | For Preventive Maintenance LCD Monitor Problem | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Reset cable connection of the Monitor, isolate the problem by changing the display cable | Perform Preventive Maintenance and LCD Monitor Troubleshooting | Machine is Operational | Monitor LCD problem documented and sent to higher ups for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482500925165226526/6482500925165226526_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/6482500925165226526/6482500925165226526_signature_31.png
| 2026-03-02 | 2026-02-27 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32221 | UNOPS | acho.laboratory@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 202 ASSAY: Xpert MTB/RIF Ultra, COUNT: 61 ASSAY: Xpert MTB-XDR, COUNT: 187 ASSAY: Xpert HIV-1 Viral Load, COUNT: 669 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482484803615542450/6482484803615542450_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6482484803615542450/6482484803615542450_signature_31.png
| 2026-03-02 | 2026-02-12 | Juan Posadas Health Center | E. Rodriguez St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32220 | PBSP | lizjunio67@gmail.com | 25 |  |  | GXIV | 805060 | Unit for preventive maintenance. Unit for module replacement. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. | New module 210043557, Def module 761520 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482482563615699959/6482482563615699959_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/6482482563615699959/6482482563615699959_signature_31.png
| 2026-03-02 | 2026-02-19 | Masbong Health Center | Masambong, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32219 | PBSP | michellebonrostro@gmail.com | 25 |  |  | GXIV-10C | 120000173 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Modules 1, 3 and 4 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482481733611936285/6482481733611936285_signature_28.png | Michelle Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/6482481733611936285/6482481733611936285_signature_31.png
| 2026-03-02 | 2026-02-20 | Caloocan City Health Office | Sangandaan, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32218 | PBSP | caloocancentrallaboratory@gmail.com | 25 |  |  | GXIV | 805271 | Unit for preventive maintenance. For module replacement of module A2 (210000207). | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. | New module A2 (210026094) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482481423618161884/6482481423618161884_signature_28.png | Anna Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/6482481423618161884/6482481423618161884_signature_31.png
| 2026-03-02 | 2026-03-02 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | RHIVDA | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | vincentmilestone20@gmail.com |  | 32217 | SHELL FOUNDATION | nimezsheilamarie@gmail.com | 25.4 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1316 ASSAY: Xpert SARS-CoV-2, COUNT: 5 |  | GX-IV 10C 4D | 120000127 | For Preventive Maintenance and Module Replacement | Replaced the Defective Module SN: 220011059 with new Module SN: 721010 Replaced the Defective Module SN: 220011388 with New Module SN: 721596 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482461695167002860/6482461695167002860_signature_28.png | Sheila Nimez,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6482461695167002860/6482461695167002860_signature_31.png
| 2026-03-02 | 2026-02-27 | Quezon City Jail | Payatas Rd. QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32216 | PBSP | carleecruz0@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2156 |  | GX4D 6C | 803089 | Due for xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck and maintenance | A1 666882 ,A3 661626 failed xc A2,A4 passed xc |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482446323615506637/6482446323615506637_signature_28.png | Carlee Mae L Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6482446323615506637/6482446323615506637_signature_31.png
| 2026-03-02 | 2026-02-27 | Payatas B Health Center | Bulacan st.,Payatas QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32215 | MACARE-HBDC | racquelmrocero@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1904 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1904 |  | GX4D 6C | 831899 | Due for Xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed Xpertcheck and maintenance | All module passed in xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482443843617613375/6482443843617613375_signature_28.png | Racquel M. Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/6482443843617613375/6482443843617613375_signature_31.png
| 2026-03-02 | 2026-02-26 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32214 | MACARE-HBDC | eamcmolecularlab@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 1113 ASSAY: Xpert HBV VL, COUNT: 167 ASSAY: Xpert HPV, COUNT: 55 ASSAY: Xpert SARS-CoV-2, COUNT: 447 ASSAY: Xpert SARS-CoV-2, COUNT: 129 ASSAY: Xpert HIV-1 Viral Load, COUNT: 17 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 620 |  | GX4 D 10C | 110016781 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482443153618926783/6482443153618926783_signature_28.png | Molecular Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/6482443153618926783/6482443153618926783_signature_31.png
| 2026-03-02 | 2026-02-25 | Casimiro Ynares Sr .Memorial Hospital | Becco St.Rodriguez Rizal | 4A | Rizal | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32213 | DOH | caysmhlab@gmail.com | 26 |  |  | GX4 6C | 831642 | Booting problem of workstation HP Z240 2UA8471W3L | Checked and reset hdd cable Replaced bios battery Checked bios set up and configuration Restart multiple time but error still occur Cleaned memory slot and swap in onter slot. | Performed troubleshooting | Unoperational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482441993616670593/6482441993616670593_signature_28.png | Aileen A. Docto | https://www.jotform.com/uploads/macarewecare/222927978372067/6482441993616670593/6482441993616670593_signature_31.png
| 2026-03-02 | 2026-03-02 | Liloan RHU | Liloan | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32212 | PBSP | vickcabahugmt@gmail.com | 31 |  |  | Gx 4 | 835189 | N/A | Replace fan filter Gather data IQR, system log | Replace fan filter Gather data IQR, system log | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6482435390129065439/6482435390129065439_signature_28.png | Vick Cabahug | https://www.jotform.com/uploads/macarewecare/222927978372067/6482435390129065439/6482435390129065439_signature_31.png
| 2026-03-01 | 2026-02-25 | Roxas City Health Office | Roxas City, Capiz | 6 | Capiz | TB Dots | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32211 | DOH | ruby_camacho@ymail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 511 ASSAY: Xpert MTB-XDR, COUNT: 3 |  | 10C GXIV-4D | 110025134 | None | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6481148083812235802/6481148083812235802_signature_28.png | Ruby Lut | https://www.jotform.com/uploads/macarewecare/222927978372067/6481148083812235802/6481148083812235802_signature_31.png
| 2026-02-27 | 2026-02-27 | Mariveles Rural Health Unit | Mariveles Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32210 | DOH | angelicaroseconos@gmail.com | 25 |  |  | GXIV-D 6C | 831894 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational B1: 698746 Passed B2: 761029 Failed B3: 758206 Passed B4: 601578 Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479969538322815689/6479969538322815689_signature_28.png | Angelica Rose Conos | https://www.jotform.com/uploads/macarewecare/222927978372067/6479969538322815689/6479969538322815689_signature_31.png
| 2026-02-27 | 2026-02-26 | Cordova Primary Health Care | Poblacion, Cordova Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | melodyjoynoelle0428@gmail.com |  | 32209 | PBSP | melodyjoynoelle0428@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 200 |  | GX4 | 808638 | Printer cannot operate due to errors in loading printer driver. | Clean up printer and conduct preventive maintenance. Install new printer updated driver Generate IQR, System Log, and Gxx files. Test Print | Clean up printer and conduct preventive maintenance. Install new printer updated driver Generate IQR, System Log, and Gxx files. Test Print | Units is operational including printer. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY | Mark Efren Ramas | https://www.jotform.com/uploads/macarewecare/222927978372067/6479913625898742860/6479913625898742860_signature_28.png | Melody Arasan | https://www.jotform.com/uploads/macarewecare/222927978372067/6479913625898742860/6479913625898742860_signature_31.png
| 2026-02-27 | 2026-02-27 | Manila Doctors Hospital | United Nations Avenue, Ermita, Manila | NCR | NCR 1st - Manila | Molecular Pathology Laboratory | NO | MAINTENANCE | CEPHEID | Julius Marion Biteng |  | engineering@macare-medicals.com | mjrivera@maniladoctors.com.ph |  | 32208 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 20.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1254 ASSAY: Xpert CT/NG, COUNT: 57 ASSAY: Xpert HBV VL, COUNT: 97 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 ASSAY: Xpert HPV, COUNT: 33 ASSAY: Xpert HCV VL, COUNT: 18 |  | GX4 - 10C | 110001648 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Optics Checked filter if clogged Performed self-test Generate reports | All Modules are Operational | Fan Filter for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479859931711544404/6479859931711544404_signature_28.png | Alexa Beatrice Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/6479859931711544404/6479859931711544404_signature_31.png
| 2026-02-27 | 2026-02-27 | Tanay Rural Health Unit | Tanay, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32207 | PBSP | maryjunnette@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1444 ASSAY: Xpert HIV-1 Viral Load, COUNT: 79 |  | GX-IV-D-6C | 807593 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -filter Run XpertCheck Generated reports | Unit is operational. Module B4 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479852343817033584/6479852343817033584_signature_28.png | Mary Junnette Fulgueras | https://www.jotform.com/uploads/macarewecare/222927978372067/6479852343817033584/6479852343817033584_signature_31.png
| 2026-02-27 | 2026-02-27 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32206 | PBSP | bmcrhivda@gmail.com | 27 |  |  | GXIV | 805112 | Unit for calibration | Replaced filter Performed XpertCheck | #NAME? | Unit is operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479840323127671327/6479840323127671327_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6479840323127671327/6479840323127671327_signature_31.png
| 2026-02-27 | 2026-02-27 | Limay Rural Health Unit | Limay Bataan | 3 | Bataan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32205 | DOH | nelsonchavez40771@gmail.com | 25 |  |  | GXIV-D 6C | 805124 | Due for calibration B1: valve error | Preventive maintenance Xpertcheck Module B1 remove stuck cartridge | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Xpertcheck Performed functionality test Gathered reports | Unit is not fully operational B1: 687261 Xpertcheck didn't go through Valve error B2: 210021632 Failed B3: 210100613 Failed B4: 609681 Failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479755302111196426/6479755302111196426_signature_28.png | Nelson Chavez | https://www.jotform.com/uploads/macarewecare/222927978372067/6479755302111196426/6479755302111196426_signature_31.png
| 2026-02-27 | 2026-02-27 | Zamboanga City Health Office | Pettit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 32204 | PBSP | floresteresitaj25@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1976 | 2026-02-27 | GX4 - 6C | 818902 | None | Module 1, 2,3 and 4 undergo xpertcheck calibration and passed. Preventive Maintenance and data collection. | Xpertcheck calibration and Preventive Maintenence. | All modules are available and ready to use | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479748606412932274/6479748606412932274_signature_28.png | Teresita Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6479748606412932274/6479748606412932274_signature_31.png
| 2026-02-27 | 2026-02-27 | Sultan Kudarat Provincial Hospital | Isulan,Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | YES | CALIBRATION MAINTENANCE Xpertcheck | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32203 | PBSP | hazyljoy_05@yahoo.com | 27.8 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 27 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2184 ASSAY: Xpert MTB-XDR, COUNT: 46 | 2026-02-27 | GxIV-10C | 110013958 | Due for Xpertcheck | Perform Genexpert machine annual Xpertcheck Calibration Perform Genexpert machine Preventive maintenance Checking of modules status and errors Data reporting | PM Done as per checklist document #D20843 Xpertcheck Done as per Document # 301-1896 rev. D | 4 operational modules Xpertcheck Results: Module B1,B2 and B4 -PASSED Module B3 - Requiring Service |  | Xpertcheck cartridge CE-5 | 334779 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479745796719792821/6479745796719792821_signature_28.png | Hazyl Joy Balindua | https://www.jotform.com/uploads/macarewecare/222927978372067/6479745796719792821/6479745796719792821_signature_31.png
| 2026-02-27 | 2026-02-27 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32201 | PBSP | acho.laboratory@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 84 ASSAY: Xpert MTB/RIF Ultra, COUNT: 9947 ASSAY: Xpert HIV-1 Viral Load, COUNT: 442 |  | GX-XVI-D-6C | 842226 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Modules B2, C2, C4, D1 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479726616342396098/6479726616342396098_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6479726616342396098/6479726616342396098_signature_31.png
| 2026-02-27 | 2026-02-27 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32202 | PBSP | acho.laboratory@gmail.com | 24 |  |  | GX-IV-D-6C | 808629 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Run XpertCheck Generated reports Return the GeneXpert to the box | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479726616348246694/6479726616348246694_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6479726616348246694/6479726616348246694_signature_31.png
| 2026-02-27 | 2026-02-27 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32200 | PBSP | acho.laboratory@gmail.com | 24 |  |  | GX-IV-D-6C | 808629 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Run XpertCheck Generated reports Return the GeneXpert to the box | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479722366341697837/6479722366341697837_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6479722366341697837/6479722366341697837_signature_31.png
| 2026-02-27 | 2026-02-27 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32199 | PBSP | acho.laboratory@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 84 ASSAY: Xpert MTB/RIF Ultra, COUNT: 9947 ASSAY: Xpert HIV-1 Viral Load, COUNT: 442 |  | GX-XVI-D-6C | 842226 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Modules B2, C2, C4, D1 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479722366349555384/6479722366349555384_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6479722366349555384/6479722366349555384_signature_31.png
| 2026-02-27 | 2026-02-25 | Sultan Kudarat Provincial Hospital | Isulan,Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32198 | PBSP | hazyljoy_05@yahoo.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1283 ASSAY: Xpert MTB-XDR, COUNT: 46 ASSAY: Xpert HIV-1 Viral Load, COUNT: 28 |  | GxIV-10C | 110013958 | Due for XPC | Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Data collection | Genexpert machine Preventive maintenance | 3 operational modules Module 1 for troubleshooting due to temperature error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479700766711450930/6479700766711450930_signature_28.png | Hazyl Joy Balindua | https://www.jotform.com/uploads/macarewecare/222927978372067/6479700766711450930/6479700766711450930_signature_31.png
| 2026-02-27 | 2026-02-24 | Tulunan Rural Health Unit | Poblacion Tulunan, North Cotabato | 12 | Cotabato | Laboratory | YES | MAINTENANCE Xpertcheck | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32197 | PBSP | marshalawrenceangeles@gmail.com | 28.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 493 | 2026-02-24 | GxIV-6C | 807615 | Due for Xpertcheck | Perform annual Xpertcheck Perform Genexpert maintenance and cleaning Checking of modules status and errors Plunger rod and Optics cleaning Fan and filter cleaning Data collection | Genexpert machine Xpertcheck Calibration Genexpert machine Preventive maintenance | 4 operational modules Xpertcheck result Module B1 - requiring for Service Module B2,B3 and B4 - Passed |  | Xpertcheck CE-5 cartridge | 332624 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6479700766714996664/6479700766714996664_signature_28.png | Marsha Lawrence Porras-Angeles | https://www.jotform.com/uploads/macarewecare/222927978372067/6479700766714996664/6479700766714996664_signature_31.png
| 2026-02-26 | 2026-02-26 | San Rafael Rural Health Unit | San Rafael, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32196 | DOH | tatsdayrit25@gmail.com | 27 |  |  | GXIV | 831692 | Workstation for further troubleshooting | Replaced bios battery | #NAME? | Unit is not operational For follow up report |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478972910804981511/6478972910804981511_signature_28.png | Joana Manelle Dayrit | https://www.jotform.com/uploads/macarewecare/222927978372067/6478972910804981511/6478972910804981511_signature_31.png
| 2026-02-26 | 2026-02-26 | Ramos Rural Health Unit | Poblacion South Ramos, Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32195 | DOH | jadeeh8143@yahoo.com | 25 |  |  | GXIV-D 6C | 831721 | B4: 631022 for module replacement Due For calibration | Preventive maintenance Xpertcheck Repair/ module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Xpertcheck Performed functionality test Gathered reports | Unit is operational B1:7575544 Passed B2: 757555 Failed B3: 766472 Passed B4: 745405 Newly replaced module Defective SN:631022 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478963275794462651/6478963275794462651_signature_28.png | Emily Jade Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/6478963275794462651/6478963275794462651_signature_31.png
| 2026-02-26 | 2026-02-26 | Mauban Rural Health Unit | Mauban, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32194 | PBSP | rhumauban.laboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 875 |  | GX-IV-D-6C | 814375 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module B3/674744 not operational due to error 4017. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478949133416317277/6478949133416317277_signature_28.png | Rhandy Buraga | https://www.jotform.com/uploads/macarewecare/222927978372067/6478949133416317277/6478949133416317277_signature_31.png
| 2026-02-26 | 2026-02-26 | Mauban Rural Health Unit | Mauban, Quezon | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32193 | MACARE-HBDC | rhumauban.laboratory@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 699 ASSAY: Xpert MTB-XDR, COUNT: 13 |  | GX-IV-D-10C | 110015487 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Run XpertCheck Generated reports | Unit is operational. Module B3/210077858 service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478947023415562546/6478947023415562546_signature_28.png | Rhandy Buraga | https://www.jotform.com/uploads/macarewecare/222927978372067/6478947023415562546/6478947023415562546_signature_31.png
| 2026-02-26 | 2026-02-26 | Meycauayan City Health Unit VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32192 | DOH | ronamarieiris_25@yahoo.com | 27 |  |  | GXIV | 831904 | Unit for calibration | Performed XpertCheck Replaced filter | #NAME? | Unit is operational All modules passed calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478915032707243798/6478915032707243798_signature_28.png | Rona Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/6478915032707243798/6478915032707243798_signature_31.png
| 2026-02-26 | 2026-02-26 | Toledo City Health Office | Poblacion Toledo City | 7 | Cebu | TB Laboratory | NO | TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32191 | PBSP | tchcl6038@gmail.com | 23 |  |  | GX IV 6C | 808640 | Module B4 Always booting. | Save File's system log , installation qualifications report. | Module B4 transfer to Module B3 is functional but the of the ultrasonic horn are twisting. | Module B4 for observation only. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478859260323605739/6478859260323605739_signature_28.png | Mary June Elemino | https://www.jotform.com/uploads/macarewecare/222927978372067/6478859260323605739/6478859260323605739_signature_31.png
| 2026-02-26 | 2026-02-26 | Hinabangan Rural Health Unit | Hinabangan Samar | 8 | Samar | TB Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | Johnpaullabutap@gmail.cok |  |  | 32190 | PBSP | nabloruth58@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 275 | 2026-02-26 | GXIV - 2 SITES | 835099 | Modules not detected on Genexpert App | #NAME? | #NAME? | Unit is now operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478849975226862545/6478849975226862545_signature_28.png | Ruth C Nablo | https://www.jotform.com/uploads/macarewecare/222927978372067/6478849975226862545/6478849975226862545_signature_31.png
| 2026-02-26 | 2026-02-26 | Mandaue City Health Office | Mandaue City | 7 | Cebu | Laboratory | NO | REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32189 | PBSP | mendozarey42@gmail.com | 28 |  |  | Gx 4 | 838510 | Stucked up cartridge in module 4 Clean Fan Filter | Remove Stucked up cartridge in module 4 Clean Fan Filter Clean module 4 bay | Remove Stucked up cartridge in module 4 Clean Fan Filter Clean module 4 bay | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478849405556629093/6478849405556629093_signature_28.png | REY C. MENDOZA | https://www.jotform.com/uploads/macarewecare/222927978372067/6478849405556629093/6478849405556629093_signature_31.png
| 2026-02-26 | 2026-02-26 | San Leonardo Super Health Center | San Bartolome San Leonardo Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32188 | PBSP | rhu1sanleonardo@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 462 |  | GXIV-6C | 846042 | Due for Maintenance Due for Calibration | Performed Xpertcheck (Module B1 and B4 = Passed) (Mdoule B2(904939) and B3(904942) = Failed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478835798228760650/6478835798228760650_signature_28.png | Ligaya Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/6478835798228760650/6478835798228760650_signature_31.png
| 2026-02-26 | 2026-02-26 | Cuyapo Rural Health Unit | Cuyapo Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32187 | DOH | jolenerodico09171998@gmail.com | 26 |  |  | GXIV-D 6C | 831672 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational A1: 757044 PASSED A2: 682096 PASSED A3: 682096 PASSED A4: 884790 PASSED/ 1 Retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478825414012792375/6478825414012792375_signature_28.png | Jolene Rodico | https://www.jotform.com/uploads/macarewecare/222927978372067/6478825414012792375/6478825414012792375_signature_31.png
| 2026-02-26 | 2026-02-25 | Agdao District Health Center | Agdao, Davao City | 11 | Davao del Sur | TB-DOTS | YES | REPAIR MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32186 | PBSP | agdao.lab@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 873 |  | GX-IV 6C 4D | 847254 | For Module Replacement and Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Replaced the Defective Module SN:910215 with new Module SN 870986 | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478786153914531922/6478786153914531922_signature_28.png | Princess Joy R. Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/6478786153914531922/6478786153914531922_signature_31.png
| 2026-02-25 | 2026-02-25 | De La Salle Health Sciences Institute | Congressional Ave, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32184 | CONSORTIUM | lmmasangcay@my.dlshsi.edu.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 911 |  | GX16-D-6C | 842166 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478377150415423564/6478377150415423564_signature_28.png | Leonedy M. Masangcay | https://www.jotform.com/uploads/macarewecare/222927978372067/6478377150415423564/6478377150415423564_signature_31.png
| 2026-02-25 | 2026-02-25 | De La Salle Health Sciences Institute | Congressional Ave., Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32185 | PBSP | lmrodil@dlsmhsi.edu.ph | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1367 |  | GXIV-L-6C | 707514 | Module 3 With 5007 Error. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Performed Self-test. Generate Reports. | B1, B2 and B4 are Operational. B3 for Further Observation after Icore Cleaning. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478377150417634800/6478377150417634800_signature_28.png | Leigh Danica M. Rodil | https://www.jotform.com/uploads/macarewecare/222927978372067/6478377150417634800/6478377150417634800_signature_31.png
| 2026-02-25 | 2026-02-25 | De La Salle Health Sciences Institute | Congressional Ave., Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32183 | CONSORTIUM | lmrodil@dlsmhsi.edu.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 204 |  | GXIV-D-6C | 812380 | Module 4 with Booting Problem. | Troubleshooting. Preventive Maintenance | Reset Gateway Board. Swap Modules to Isolate Problem. Performed Self-test. Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Replaced Filter. Generate Reports. | B4 Booting Problem Resolved. All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478377150419952417/6478377150419952417_signature_28.png | Leigh Danica Rodil | https://www.jotform.com/uploads/macarewecare/222927978372067/6478377150419952417/6478377150419952417_signature_31.png
| 2026-02-25 | 2026-02-25 | Olongapo City Health Office | West Bajac-Bajac Olongapo City Health Office | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32182 | PBSP | olongapochotmlandrtdl@gmail.com | 25 |  |  | GXIV-D 6C | 804700 | Due for calibration | Preventative maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational D1: 629603 PASSED D2: 617387 FAILED/ 1 retry D3:672030 PASSED D4: 210053177 PASSED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478146214627012664/6478146214627012664_signature_28.png | Roselyn Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/6478146214627012664/6478146214627012664_signature_31.png
| 2026-02-25 | 2026-02-25 | University of The East Ramon Magsaysay Memorial Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32181 | CONSORTIUM | biomed@uerm.edu.ph | 25 |  | 2026-02-24 | GXIV | 847578 | Unit for calibration Unit for maintenance | Performed XpertCheck | #NAME? | Unit is operational B1: Passed B2: Service Required B3: Service Required B4: Service Required Modules B2, B3, B4 are recommended for replacement |  | XpertCheck Kit Lot: 1001480606 | 102196 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478136728325130754/6478136728325130754_signature_28.png | Osang Comendador | https://www.jotform.com/uploads/macarewecare/222927978372067/6478136728325130754/6478136728325130754_signature_31.png
| 2026-02-25 | 2026-02-25 | Olongapo City Health Office | West Bajac-Bajac Olongapo City Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32180 | PBSP | olongapochotmlandrtdl@gmail.com | 25 |  |  | GXIV-D 10C | 120000606 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational B1: 220013803 FAILED/ 1 retry B2: 220013798 PASSED B3: 220013768 PASSED B4: 220013766 PASSED/ 1 retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478112614625213759/6478112614625213759_signature_28.png | Roselyn Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/6478112614625213759/6478112614625213759_signature_31.png
| 2026-02-25 | 2026-02-25 | Morong Rural Health Unit | Morong, Rizal | 4A | Rizal | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32179 | DOH | moronglab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 627 |  | GX-IV-D-6C | 821213 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -icore -Filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478093272813435095/6478093272813435095_signature_28.png | Necitas V Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/6478093272813435095/6478093272813435095_signature_31.png
| 2026-02-25 | 2026-02-25 | Lantawan Rural Health Unit | Lantawan, Basilan | BARMM | Basilan | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32178 | PBSP | abdurrauflakibulsagala@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 570 | 2026-02-25 | GX 4 - 6c | 807460 | Module stucked cartridge and hardware failed. | Preventive Maintenence, troubleshoot module 3, Swapped location at module 4, error carried over by the module at new location error, confirming module related issue. | Preventive Maintenence and removed stucked cartridge in module 3. Troubleshoot module 3 error 4008. | Module 1, 2 and 4 is available and ready to use. While module 2 is on observation. | Module 3 sn (631762) is for replacement. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478052796911452866/6478052796911452866_signature_28.png | Abdurrauf Sagala | https://www.jotform.com/uploads/macarewecare/222927978372067/6478052796911452866/6478052796911452866_signature_31.png
| 2026-02-25 | 2026-02-25 | Guagua Rural Health Unit | Cerville Subd. Guagua Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32177 | DOH | abigailgsula@gmail.com | 25 |  |  | GXIV-D | 831912 | Due for Calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational A1: 210001116 Passed A2: 720724 Passed A3: 739422 Passed A4: 210008371 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6478004274624096597/6478004274624096597_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/6478004274624096597/6478004274624096597_signature_31.png
| 2026-02-25 | 2026-02-24 | Carles Primary Care Facility | Carles, Iloilo | 6 | Iloilo | Laboratory | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32176 | UNOPS | vangieyuro3@gmail.com | 25 |  | 2026-02-11 | 10C GXIV-4D | 110023526 | Module B2 and B3 failed XPC | -Replace defective module B2/210122299 with SN: 643939 -Replace defective module B3/21012296 with SN: 892742 | #NAME? | Machine is fuctional. |  | Certified pre-owned GeneXpert 10-Color Module SN. 643939 and SN. 892742 | DR no. 13736 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477941365616576480/6477941365616576480_signature_28.png | Gilda Marcelo | https://www.jotform.com/uploads/macarewecare/222927978372067/6477941365616576480/6477941365616576480_signature_31.png
| 2026-02-25 | 2026-02-25 | Angeles Rural Health Unit 3 | Fajardo St. Lourdes Sur, Angeles City Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32175 | PBSP | femiepangilinan@yahoo.com | 25 |  |  | GXIV-D | 807569 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational B1: 210004953 PASSED B2: 670856 PASSED B3: 647576 For reporting Temperature error B4:875099 PASSED | B3: 647576 For reporting Temperature error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477909575123939925/6477909575123939925_signature_28.png | Femie Pangilinan | https://www.jotform.com/uploads/macarewecare/222927978372067/6477909575123939925/6477909575123939925_signature_31.png
| 2026-02-25 | 2026-02-25 | ILOILO PROVINCIAL HOSPITAL | POTOTAN, ILOILO | 6 | Iloilo | Laboratory | NO | MAINTENANCE Transfer | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 32174 | DOH | ipg.molab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1628 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GX-IV | 838289 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477898595899511020/6477898595899511020_signature_28.png | Vannessa Bianca Principe, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6477898595899511020/6477898595899511020_signature_31.png
| 2026-02-24 | 2026-02-20 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32173 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477254220113227126/6477254220113227126_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6477254220113227126/6477254220113227126_signature_31.png
| 2026-02-24 | 2026-02-24 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32170 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 582 ASSAY: Xpert MTB-XDR, COUNT: 24 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B4 under observation due to erro 2005. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622452328164/6477253622452328164_signature_28.png | Esty Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622452328164/6477253622452328164_signature_31.png
| 2026-02-24 | 2026-02-24 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32172 | PBSP | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 814 |  | GX-IV-D-10C | 110020361 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622457451179/6477253622457451179_signature_28.png | Esty Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622457451179/6477253622457451179_signature_31.png
| 2026-02-24 | 2026-02-24 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32169 | DOH | estyaragon2019@gmail.com | 22 |  |  | GX-IV-D-6C | 831588 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622457813283/6477253622457813283_signature_28.png | Esty Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622457813283/6477253622457813283_signature_31.png
| 2026-02-24 | 2026-02-24 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32171 | PBSP | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1564 |  | GX-IV-D-6C | 805211 | For XpertCheck No boot device found | Run XpertCheck Troubleshoot CPU | Replaced BIOS battery Change boot order and settings Performed maintenance Cleaned: -Cartridge bay -plunger rod -icore -filter Generated reports Run XpertCheck | Unit is operational. Modules 1,3 and 4 passed. Module 2 not calibrated due to it cannot be detected. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622458421446/6477253622458421446_signature_28.png | Esty Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622458421446/6477253622458421446_signature_31.png
| 2026-02-24 | 2026-02-24 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32168 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 843 ASSAY: Xpert MTB-XDR, COUNT: 3 |  | GX-IV-D-6C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -icore -filter Generated reports | Unit os operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622459996888/6477253622459996888_signature_28.png | Esty Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6477253622459996888/6477253622459996888_signature_31.png
| 2026-02-24 | 2026-02-24 | Bagac Rural Health Unit | Bagac Bataan | 3 | Bataan | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32167 | PBSP | jonathanricalde_17@yahoo.com.ph | 25 |  |  | GXII-D | 835173 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational B1: 788614 Passed B2: 788615 Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477233062715313443/6477233062715313443_signature_28.png | Jonathan Ricalde | https://www.jotform.com/uploads/macarewecare/222927978372067/6477233062715313443/6477233062715313443_signature_31.png
| 2026-02-24 | 2026-02-24 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 32166 | COMMERCIAL | cduhbloodbank2@gmail.com | 19 |  |  | MCS+ | 14F014 | Damage rubber of centrifuge cover. | Replacement of new cover of centrifuge. | Replacement of new cover of centrifuge. | Unit is operational. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477216855892092385/6477216855892092385_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6477216855892092385/6477216855892092385_signature_31.png
| 2026-02-24 | 2026-02-18 | Badian Rural Health Center( Primary Health Care) | Poblacion, Badian Cebu | 7 | Cebu | Tb laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | avenzr@gmail.com |  | 32165 | PBSP | avenzr@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1500 |  | GX4 | 807588 | Due for Calibration of GeneXpert machine. | Perform calibration in Module 1, 3 and 4. Performed preventive maintenance, clean up Icore, ultrasonic horn, filter and fan. Generate data IQR, System log, and Gxx files. | Perform calibration in Module 1, 3 and 4. Performed preventive maintenance, clean up Icore, ultrasonic horn, filter and fan. Generate data IQR, System log, and Gxx files. | All modules passed Module 1, 2 and 3 except Module 2 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477214805899475044/6477214805899475044_signature_28.png | Avenizer Absin | https://www.jotform.com/uploads/macarewecare/222927978372067/6477214805899475044/6477214805899475044_signature_31.png
| 2026-02-24 | 2026-02-24 | Kolambugan Rural Health Unit | Kolambugan, Lanao del Norte | 10 | Lanao del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 32164 | PBSP | yojecinaj@gmail.com | 25 |  |  | GX4 6 colors | 815272 | Defective Gateway Board Defective Module | None | Replaced Module SN: 610218 with SN: 866080 Replaced Gateway Board Preventive Maintenance | All modules and units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477201983442391712/6477201983442391712_signature_28.png | Janice Joy R. Siso | https://www.jotform.com/uploads/macarewecare/222927978372067/6477201983442391712/6477201983442391712_signature_31.png
| 2026-02-24 | 2026-02-19 | Cebu TB Reference Lab | South Road Property - Cebu | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32163 | PBSP | asmantiquilla.pbsp@gmail.com | 28 |  |  | Gx4 | 120000180 | Due Date for Xpertcheck Calibration Gather data IQR, System log Calibration Data | Due Date for Xpertcheck Calibration Gather data IQR, System log Calibration Data | Due Date for Xpertcheck Calibration Gather data IQR, System log Calibration Data | Xpertcheck Calibration Module 2 passed Module 1,3 &4 Service Require |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477197652324924881/6477197652324924881_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6477197652324924881/6477197652324924881_signature_31.png
| 2026-02-24 | 2026-02-19 | Cebu TB Reference Lab | South Road Property - Cebu | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32162 | PBSP | asmantiquilla.pbsp@gmail.com | 28 |  |  | Gx4 | 120000180 | Due Date for Xpertcheck Calibration Gather data IQR, System log Calibration Data | Due Date for Xpertcheck Calibration Gather data IQR, System log Calibration Data | Due Date for Xpertcheck Calibration Gather data IQR, System log Calibration Data | Xpertcheck Calibration Module 2 passed Module 1,3 &4 Service Require |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477197442224319525/6477197442224319525_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6477197442224319525/6477197442224319525_signature_31.png
| 2026-02-24 | 2026-02-20 | Cebu City Jail Male Dorm | Upper Kalunasan | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 32161 | PBSP | jasperperanco@gmail.com | 28 |  |  | Gx4 | 805214 | Defective modules B3: 912371 B4: 21012208 | New module Installed: B3-805754 and B4-865164 Gather data IQR,System log Clean Fan and Fan Filter | New module Installed: B3-805754 and B4-865164 Gather data IQR,System log Clean Fan and Fan Filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477194822516322180/6477194822516322180_signature_28.png | Jasper Peranco | https://www.jotform.com/uploads/macarewecare/222927978372067/6477194822516322180/6477194822516322180_signature_31.png
| 2026-02-24 | 2026-02-24 | Zamboanga city Health office | Pettit Baracks, Zamboanga city | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32160 | PBSP | floresteresitaj25@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1961 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1961 | 2026-02-24 | GX 4 - 6c | 807599 | None | Preventive Maintenence, Site inspection and checked module and room temperatures. | Preventive Maintenence, Site inspection and checked module and room temperatures. | All modules are available and ready to use. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477150248716676277/6477150248716676277_signature_28.png | Teresita Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6477150248716676277/6477150248716676277_signature_31.png
| 2026-02-24 | 2026-02-24 | Samal Super Health Center | Samal Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32159 | DOH | samalrhulab@gmail.com | 25 |  |  | GXIV-D 6C | 831729 | Due for Calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational A1:674018 PASSED A2:757532 PASSED A3:757533 PASSED A4:784697 Failed/ 1 retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6477073849617569261/6477073849617569261_signature_28.png | Maria Del Mar Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/6477073849617569261/6477073849617569261_signature_31.png
| 2026-02-23 | 2026-02-18 | Lambayong Rural Health Unit | Lambayong, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32157 | PBSP | jeanalonzo73@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 952 | 2026-02-18 | GxIV-6C | 838159 | Module B1 SN: 749488 defective module | Replaced defective module B1 SN:749488 with new module SN:875091 Perform Genexpert maintenance and cleaning Module cleaning Fan and filter cleaning Plunger rod and Optics cleaning Checking of modules status and errors Data collection | Genexpert machine module repair Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert 6-color module SN:875091 | 331870 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476548737323184603/6476548737323184603_signature_28.png | Jean Alonzo | https://www.jotform.com/uploads/macarewecare/222927978372067/6476548737323184603/6476548737323184603_signature_31.png
| 2026-02-23 | 2026-02-16 | Tboli Rural Health Unit | Tboli, South Cotabato | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32158 | PBSP | tbolirhulab@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 763 | 2026-02-16 | GxIV-6C | 815432 | Module B3 requiring service | Replace defective B3 SN:696323 with new module SN:872017 Perform Genexpert Cleaning and maintenance Module cleaning Fan and filter cleaning Plunger rod and Optics cleaning Data collection | Genexpert machine Preventive maintenance Genexpert machine module repair | 4 operational modules |  | Genexpert 6 color module SN:872017 | 331856 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476548737324519135/6476548737324519135_signature_28.png | Judessa San Juan | https://www.jotform.com/uploads/macarewecare/222927978372067/6476548737324519135/6476548737324519135_signature_31.png
| 2026-02-23 | 2026-02-23 | Philippine Tuberculosis Society, Inc. - Roxas City | Roxas City, Capiz | 6 | Capiz | TB Dots | NO | INSTALLATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32156 | MACARE-HBDC | pinkmansiontreatmenthub@gmail.com | 26 |  | 2026-01-12 | 10C GXIV-4D | 120007067 | For installation | #NAME? | -Installed New GeneXpert Unit with SN: 120007067 | Machine is functional |  |  | DR no. 13679 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476470597263444562/6476470597263444562_signature_28.png | Michael Alajada | https://www.jotform.com/uploads/macarewecare/222927978372067/6476470597263444562/6476470597263444562_signature_31.png
| 2026-02-23 | 2026-02-23 | St Lukes Extension Clinic | J. Bocobo St, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 32155 | MACARE-HBDC | laboratory@slec.ph | 22 |  |  | Infinity 48s | 809643 | Conveyor belt not working during operation | Adjusted damaged conveyor belt to avoid stuck up | Performed basic troubleshooting | Unit is operational, conveyor belt temporarily fixed but shows signs of wear and tear damage and needs replacing |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476463387817881258/6476463387817881258_signature_28.png | Ma. Sarah Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6476463387817881258/6476463387817881258_signature_31.png
| 2026-02-23 | 2026-02-23 | St Lukes Extension Clinic | J. Bocobo St, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 32154 | MACARE-HBDC | laboratory@slec.ph | 22 |  |  | Infinity 48s | 809643 | Conveyor belt not working during operation | Adjusted damaged conveyor belt to avoid stuck up | Performed basic troubleshooting | Unit is operational, conveyor belt temporarily fixed but shows signs of wear and tear damage and needs replacing |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476463367818975582/6476463367818975582_signature_28.png | Ma. Sarah Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6476463367818975582/6476463367818975582_signature_31.png
| 2026-02-23 | 2026-02-23 | Sen. Gerardo M. Roxas Memorial District Hospital | Dao, Capiz | 6 | Capiz | TB Dots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32153 | PBSP | bermendirowena@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1474 |  | 6C GXIV-4D | 815280 | Module 4 failed XPC | None | #NAME? | Machine is function. Module B4 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476323517268696367/6476323517268696367_signature_28.png | Rowena L. Bermendi | https://www.jotform.com/uploads/macarewecare/222927978372067/6476323517268696367/6476323517268696367_signature_31.png
| 2026-02-23 | 2026-02-11 | Bicol Medical Center | Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32152 | MACARE-HBDC | btsc@bmc.doh.gov.ph | 25 |  |  | MCS+ | 15H031G | Unit is for Preventive Maintenance | Cleaned Unit surface Cleaned centrifuge, valve, rotor pump Checked parameters | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476250387703332373/6476250387703332373_signature_28.png | Sergs Bryan S Bausas | https://www.jotform.com/uploads/macarewecare/222927978372067/6476250387703332373/6476250387703332373_signature_31.png
| 2026-02-23 | 2026-02-11 | Bicol Medical Center | Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32151 | MACARE-HBDC | btsc@bmc.doh.gov.ph | 25 |  |  | MCS+ | 16K520SPG | Protocol caed not compatible | Cleaned Unit Cleaned surface of manifold sensor Cleaned valve, rotor pump, centrifuge Checked parameter | Performed Preventive Maintenance Performed Troubleshooting | Unit is operational Unit is for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476250057702122416/6476250057702122416_signature_28.png | Sergs Bryan S Bausas | https://www.jotform.com/uploads/macarewecare/222927978372067/6476250057702122416/6476250057702122416_signature_31.png
| 2026-02-23 | 2026-02-23 | Tayasan Rural Health Unit | Tayasan, Negros Oriental | 7 | Negros Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 32150 | UNOPS | porshcabz151@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 556 | 2026-02-11 | GX4 6c | 110023502 | Modules 210122342/B1 and 210122340/B4 failed recent Xpertcheck. | . | Replaced B1 and B4 with new modules with SN 747039 and 800797 respectively. | 4 modules operational. |  | Genexpert modules 6-colors | DR 13737 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476229143817288453/6476229143817288453_signature_28.png | Portia Cabanting | https://www.jotform.com/uploads/macarewecare/222927978372067/6476229143817288453/6476229143817288453_signature_31.png
| 2026-02-23 | 2026-01-15 | Cainta Rural Health Unit | Karangalan Dr., Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32149 | DOH | ballesterrichard18@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1107 ASSAY: Others, COUNT: 148 |  | GX-IV-D-6C | 831623 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476186373616634938/6476186373616634938_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/6476186373616634938/6476186373616634938_signature_31.png
| 2026-02-23 | 2026-01-15 | Cainta Rural Health Unit | Karangalan Dr., Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32148 | PBSP | ballesterrichard18@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 881 |  | GX-IV-D-6C | 804687 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module B3 has error 2008 and 2005. Under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476186273617229493/6476186273617229493_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/6476186273617229493/6476186273617229493_signature_31.png
| 2026-02-23 | 2026-01-15 | Cainta Rural Health Unit | Karangalan Dr., Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32147 | PBSP | ballesterrichard18@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 137 ASSAY: Xpert MTB/RIF Ultra, COUNT: 945 ASSAY: Others, COUNT: 20 |  | GX-IV-D-6C | 838182 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module B3 has error 1001 2014. Under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476186143613100022/6476186143613100022_signature_28.png | Richard A. Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/6476186143613100022/6476186143613100022_signature_31.png
| 2026-02-23 | 2026-02-19 | Jala-Jala Rural Health Unit | Jala-Jala Municipal Compound, Jala-Jala, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32146 | DOH | jjrhu@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 318 |  | GX-IV-D-6C | 820184 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476185963617608023/6476185963617608023_signature_28.png | Loisita Libertad | https://www.jotform.com/uploads/macarewecare/222927978372067/6476185963617608023/6476185963617608023_signature_31.png
| 2026-02-23 | 2026-02-19 | Teresa Rural Health Unit | Teresa, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32145 | DOH | rosbenildatamayo@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1192 |  | GX-IV-D-6C | 820710 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476185813618925219/6476185813618925219_signature_28.png | Rosbenilda Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6476185813618925219/6476185813618925219_signature_31.png
| 2026-02-23 | 2026-01-16 | Pasig City General Hospital | Eusebio, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32144 | PBSP | pcghlaboratory@gmail.com | 24 |  |  | GX-IV-D-6C | 805010 | Black screen on cursor only ate displayed in the monitor | Troubleshoot CPU | Checked connections Reseat cables and components Tried startup repair | Unit is not operational. Possible corrupted operating system. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6476185513614848780/6476185513614848780_signature_28.png | Anne Patricia L. Lipana | https://www.jotform.com/uploads/macarewecare/222927978372067/6476185513614848780/6476185513614848780_signature_31.png
| 2026-02-20 | 2026-02-20 | Ospital ng Lungsod ng San Jose Del Monte | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32143 | PBSP | olsjdmrtdlgenexpert@gmail.com | 28 |  |  | GXIV | 807624 | Unit for calibration | Replaced filter Performed XpertCheck | #NAME? | Unit is operational B1: passed B2: passed B3: passed B4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6473857606324909461/6473857606324909461_signature_28.png | Mary Ann Celestino | https://www.jotform.com/uploads/macarewecare/222927978372067/6473857606324909461/6473857606324909461_signature_31.png
| 2026-02-20 | 2026-02-20 | PRC Benguet | La Trinidad, Benguet | CAR | Benguet | Blood Bank | NO | REPAIR | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com | lykabagano@yahoo.com |  | 32142 | MACARE-HBDC | benguet@redcross.org.ph | 28 |  | 2026-02-20 | MCS + | 17H363SPG | Pump problem during UPP procedure. | Performed troubleshooting. Performed function trst. | Checked and clean the pump rotors.. Performed functuonality. Performed simulation run. | Unit is operational. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6473840789614646291/6473840789614646291_signature_28.png | LYKAINE KATE D. BAGANO | https://www.jotform.com/uploads/macarewecare/222927978372067/6473840789614646291/6473840789614646291_signature_31.png
| 2026-02-20 | 2026-02-20 | Makati City NTP Office (From West Rembo Health Center | Morong St, Makati City | NCR | NCR 4th - Makati | NTP officr | NO | Check up | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32141 | DOH | kapid5sepe@yahoo.com | 22 |  |  | GX-IV-D-6C | 831712 | For check-up | Checked system functionality | Check system functionality Checked CPU Reseat internal connections Performed hard drive check | CPU is cannot boot to Windows. Drained CMOS battery. Possible hard drive failure. Unit is not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6473834862565811666/6473834862565811666_signature_28.png | Eduardo O. Sepe | https://www.jotform.com/uploads/macarewecare/222927978372067/6473834862565811666/6473834862565811666_signature_31.png
| 2026-02-20 | 2026-02-20 | Makati City NTP Office (From East Rembo Health Center) | Morong St, Makati City | NCR | NCR 4th - Makati | NTP Office | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32140 | PBSP | kapid5sepe@yahoo.com | 24 |  |  | GX-IV-D-6C | 807591 | For check-up | Checked GeneXpert system | Check GeneXpert system Performed maintenance | Unit is operational. CMOS battery for replacement. The rear stand of GeneXpert is damaged/cut. UPS is not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6473834082561263576/6473834082561263576_signature_28.png | Eduardo O. Sepe | https://www.jotform.com/uploads/macarewecare/222927978372067/6473834082561263576/6473834082561263576_signature_31.png
| 2026-02-20 | 2026-02-20 | Bangkal Health Center | Rodriguez Ave, Bangkal, Makati City | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32138 | DOH | mhdbangkalsatlab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4546 ASSAY: Xpert MTB-XDR, COUNT: 43 |  | GX-IV-D-6C | 831875 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6473804547492172918/6473804547492172918_signature_28.png | Danica R. Adolfo | https://www.jotform.com/uploads/macarewecare/222927978372067/6473804547492172918/6473804547492172918_signature_31.png
| 2026-02-20 | 2026-02-20 | Bangkal Health Center | Rodriguez Ave, Bangkal, Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32139 | PBSP | mhdbangkalsatlab@gmail.com | 24 |  |  | GX-IV-D-10C | 120000499 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6473804547499284808/6473804547499284808_signature_28.png | Danica Adolfo | https://www.jotform.com/uploads/macarewecare/222927978372067/6473804547499284808/6473804547499284808_signature_31.png
| 2026-02-20 | 2026-02-20 | Bulakan Rural Health Unit | Bulakan, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32137 | DOH | ventanilla_amiefaith@yahoo.com | 27 |  |  | GXIV | 831641 | Unit for calibration | Replaced filter Performed XpertCheck | #NAME? | Unit is operational A1: A2: A3: A4: |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6473772254221120320/6473772254221120320_signature_28.png | Amie Faith Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/6473772254221120320/6473772254221120320_signature_31.png
| 2026-02-20 | 2026-02-20 | Kalinga Health Hub | Boni Avenue,Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32136 | MACARE-HBDC | kalingahealth.manda@innovationsch.org | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 619 | 2026-02-18 | GX4D 6C | 110008420 | Due for xpertcheck and maintenance B3 and B4 hw failed | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Restart system Gathered reports Test run unit | Performed xpertcheck Performed Maintenance | All module failed in xpertcheck B1 220000585 B2 220000750 B3 210035689 B4 210035694 For reporting |  | Xpertcheck kits | 102168 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6473737910819704263/6473737910819704263_signature_28.png | Chin Lyn Corpin | https://www.jotform.com/uploads/macarewecare/222927978372067/6473737910819704263/6473737910819704263_signature_31.png
| 2026-02-20 | 2026-02-20 | Kalinga Health Hub | Boni Avenue Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32135 | MACARE-HBDC | kalingahealth.manda@innovationsch.org | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1000 | 2026-02-18 | GX4D 6C | 110008415 | Due for xpertcheck and maintenance B3 hw failed error 1002 | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Restart system Gathered reports Test run unit | Performed maintenance Performed xpertcheck | B1 220000619,B2 220000649 failed xpertcheck B3 210035636 continous error 1002 B4 passed xpertcheck |  | Xpertcheck kits | 102168 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6473736400818282853/6473736400818282853_signature_28.png | Chinlyn corpin | https://www.jotform.com/uploads/macarewecare/222927978372067/6473736400818282853/6473736400818282853_signature_31.png
| 2026-02-19 | 2026-02-19 | Sta. Ana Hospital | Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32134 | MACARE-HBDC | sahrhivdalab@gmail.com | 27 |  |  | GXIV | 847726 | Unit for calibration | Perfomed XpertCheck Replaced filter | #NAME? | Unit is operational C1: passed C2: passed C3: passed C4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472934854327192306/6472934854327192306_signature_28.png | Bon Angelo Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6472934854327192306/6472934854327192306_signature_31.png
| 2026-02-19 | 2026-02-19 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | leoomarvinreyesjacinto@gmail.com |  | 32133 | PBSP | mandaluyongcityhealthlab@gmail.com | 23 |  |  | GXIV | 831596 | Unit for calibration | Performed XpertCheck Replaced filter | #NAME? | Unit is operational C1: Service Required C2: Retry C3: Service Required C4: Retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472876114328330275/6472876114328330275_signature_28.png | Leo Jacinto | https://www.jotform.com/uploads/macarewecare/222927978372067/6472876114328330275/6472876114328330275_signature_31.png
| 2026-02-19 | 2026-02-06 | Bagong Silang Ph. 9 Health Center | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32132 | PBSP | bsphase9hclaboratory@gmail.com | 25 |  |  | GXIV | 831677 | Unit for maintenance Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: Passed B2: Service Required B3: Passed B4: Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472874454327193618/6472874454327193618_signature_28.png | Johanna Mae Ortiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6472874454327193618/6472874454327193618_signature_31.png
| 2026-02-19 | 2026-02-19 | Veterans Memorial Medical Center | North avenue, Diliman. Quezon City | NCR | NCR 2nd - Quezon City | Pathology - Molecular Laboratory | NO | ON-SITE ASSISTANCE Perform Run Test (MTB/Rif Ultra) for compliance on VMMC inspectors | CEPHEID | Julius Marion Biteng |  | techsupport@macare-medicals.com | ralphcayetanormt@gmail.com |  | 32131 | MACARE-HBDC | vmmc.moleclab@gmail.com | 25 |  |  | GX-4 | 120007530/120007529 | N/A ko | N/A | Perform Run Test (MTB/Rif Ultra) for compliance on VMMC inspectors | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472791785425859401/6472791785425859401_signature_28.png | Cayetano, Ralph Joseph Pangilinan | https://www.jotform.com/uploads/macarewecare/222927978372067/6472791785425859401/6472791785425859401_signature_31.png
| 2026-02-19 | 2026-02-18 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32130 | COMMERCIAL | biomed@mahmc.com.ph | 22 | ASSAY: Xpert HBV VL, COUNT: 16 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 182 ASSAY: Xpert HIV-1 Viral Load, COUNT: 26 |  | GX-IV-D-6C | 110005265 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -Cartridge bay -plunger rod -Icore -Filter Run XpertCheck Generated reports | Unit is operational. Module B1 passed. Module B2, B3 and B4 skipped and retry, recommend to replace modules. | Modules B2, B3 and B4 failed in XpertCheck last April 11, 2025. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472704921329118407/6472704921329118407_signature_28.png | Gehru bedayos | https://www.jotform.com/uploads/macarewecare/222927978372067/6472704921329118407/6472704921329118407_signature_31.png
| 2026-02-19 | 2026-02-18 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32129 | MACARE-HBDC | avabichuela3216@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 414 ASSAY: Others, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 86 ASSAY: Xpert SARS-CoV-2, COUNT: 7 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 11 |  | GX16D 6C | 838655 | Due for maintenance C2 encounter temp error 4009 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Restart system Gathered reports Test run unit | Performed maintenance | C2 hw failed for reporting B2,B3,C3 failed xc last nov 2025 8 modules in operational 4 modules disabled no xpertcheck perform. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472686003614329999/6472686003614329999_signature_28.png | Riza Angelica Cacayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6472686003614329999/6472686003614329999_signature_31.png
| 2026-02-19 | 2026-02-18 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32128 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  |  | MCS + | 11H204 | No power, busted fuse | Checked unit Checked voltage out ok 215 volts Checked power cord continuity ok Replaced fuse Performed functionality test Performed simulation running | Performed troubleshooting and repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472684453618177294/6472684453618177294_signature_28.png | Josephine Montilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6472684453618177294/6472684453618177294_signature_31.png
| 2026-02-18 | 2026-02-18 | Hi-Precision Diagnostics | D. Tuazon St., Quezon City | NCR | NCR 2nd - Quezon City | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32127 | CONSORTIUM | jose@hi-precision.com.ph | 25 |  |  | GXIV | 835601 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472074299676141843/6472074299676141843_signature_28.png | Andrei Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/6472074299676141843/6472074299676141843_signature_31.png
| 2026-02-18 | 2026-02-18 | Hi-Precision Diagnostics | D. Tuazon St. Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32126 | CONSORTIUM | jose@hi-precision.com.ph | 25 |  |  | GXIV | 835602 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472073829679305433/6472073829679305433_signature_28.png | Andrei Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/6472073829679305433/6472073829679305433_signature_31.png
| 2026-02-18 | 2026-02-18 | Marilao Rural Health Unit | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32125 | DOH | jmjenilo27@gmail.com | 27 |  |  | GXIV | 831897 | Unit for maintenance Unit for calibration | Performed XpertCheck Replaced filter | #NAME? | Unit is operational A1: passed A2: service required A3: passed A4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472071158108119341/6472071158108119341_signature_28.png | Christopher M. Dilag | https://www.jotform.com/uploads/macarewecare/222927978372067/6472071158108119341/6472071158108119341_signature_31.png
| 2026-02-18 | 2026-02-06 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32124 | MACARE-HBDC | dolly.banares@redcross.org | 25 |  |  | MCS+ | 16K521SPG | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472040589672720634/6472040589672720634_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6472040589672720634/6472040589672720634_signature_31.png
| 2026-02-18 | 2026-02-18 | Atang dela Rama Health Center | Gagalangin, Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32123 | DOH | monskie_22@yahoo.com | 25 |  |  | GXIV | 831644 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472024639677134280/6472024639677134280_signature_28.png | Lenie I. Malicdem | https://www.jotform.com/uploads/macarewecare/222927978372067/6472024639677134280/6472024639677134280_signature_31.png
| 2026-02-18 | 2026-02-13 | Benguet General Hospital | La Trinidad, Benguet | CAR | Benguet | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32122 | MACARE-HBDC | bevlapisac@gmail.com | 28 |  | 2026-02-10 | GXIV-4-D-10C | 120007537 | For Installation of GeneXpert System | #NAME? | Performed Installation of GeneXpert System | Unit is Fully Operational |  | GeneXpert System | DR 13729 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472023284028233455/6472023284028233455_signature_28.png | Beverlee lapisac | https://www.jotform.com/uploads/macarewecare/222927978372067/6472023284028233455/6472023284028233455_signature_31.png
| 2026-02-18 | 2026-02-10 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32121 | DOH | seansimonpaloyo@gmail.com | 22 |  |  | GXIV-4-D-6C | 838473 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472022104021213377/6472022104021213377_signature_28.png | Manolo Paloyo | https://www.jotform.com/uploads/macarewecare/222927978372067/6472022104021213377/6472022104021213377_signature_31.png
| 2026-02-18 | 2026-02-06 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32120 | DOH | dianeruiz147@gmail.com | 28 |  |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472017554026447623/6472017554026447623_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6472017554026447623/6472017554026447623_signature_31.png
| 2026-02-18 | 2026-02-06 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32119 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6472016414024557273/6472016414024557273_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6472016414024557273/6472016414024557273_signature_31.png
| 2026-02-18 | 2026-02-18 | Marulas 3S Health Station | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32118 | DOH | melissamartin0828@gmail.com | 27 |  |  | GXIV | 831696 | Unit for calibration | Perfomed XpertCheck Replaced filter | #NAME? | Unit is operational A1: passed A2: passed A3: retry A4: retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6471991028109782425/6471991028109782425_signature_28.png | Melissa Martin | https://www.jotform.com/uploads/macarewecare/222927978372067/6471991028109782425/6471991028109782425_signature_31.png
| 2026-02-17 | 2026-02-17 | Oriental Mindoro Provincial Hospital | Oriental Mindoro | 4B | Oriental Mindoro | Molecular laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32117 | PBSP | omphlabsarscov2@gmail.com | 25 |  |  | GXIV-10C | 120000741 | Unit for preventive maintenance Unit for Xpertcheck. | Preventive maintenance and xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6471244359476485216/6471244359476485216_signature_28.png | Morre Elmido | https://www.jotform.com/uploads/macarewecare/222927978372067/6471244359476485216/6471244359476485216_signature_31.png
| 2026-02-16 | 2025-12-15 | Trece Martires City Health Office | Governor's Drive, City Health Office Building, Brgy San Agustin, Trece Martires, Cavite | 4A | Cavite | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 32116 | MACARE-HBDC | loricavivian@yahoo.com | 24 |  | 2025-12-15 | GXIV-D-10C | 110027247 | Unit is due for Installation. | Installation. | Installed GeneXpert System. Checked Functionality. Generate Reports. | All Modules are Good. UPS and Printer are good. |  | GeneXpert System | SI 101586 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470402903611842103/6470402903611842103_signature_28.png | Vivian D. Lorica | https://www.jotform.com/uploads/macarewecare/222927978372067/6470402903611842103/6470402903611842103_signature_31.png
| 2026-02-16 | 2026-02-10 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | jniebres@asianhospital.com |  | 32115 | COMMERCIAL | bloodbank@asianhospital.com | 25 |  |  | MCS+ | 01F116 | Prompting Fluid Detected Error. | Troubleshooting. | Cleaned Centrifuge Chamber. Checked Spill Detector Sensor. Checked A/D Values. Performed Simulation Test. | Unit is Operational. | Recommendation for Assitance of Technical During Procedure. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470402763613372868/6470402763613372868_signature_28.png | Roi Jordan Ocay | https://www.jotform.com/uploads/macarewecare/222927978372067/6470402763613372868/6470402763613372868_signature_31.png
| 2026-02-16 | 2026-02-12 | General Trias City Health Office | Pinagtipunan, General Trias, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE XpertCheck | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | chogentri.lab@gmail.com |  | 32114 | PBSP | mackchico00@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4436 ASSAY: Xpert HIV-1 Viral Load, COUNT: 36 |  | GXIV-D-6C | 804691 | Unit is due for XpertCheck. | Preventive Maintenance. XpertCheck. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Replaced Filter. Performed Self-test. Archived Teat Results. Generate Reports. | All Modules Passed XpertCheck. | XpertCheck Details : 1001480607 (2027-03-21) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470402553618038158/6470402553618038158_signature_28.png | MACK ALFRED L. CHICO | https://www.jotform.com/uploads/macarewecare/222927978372067/6470402553618038158/6470402553618038158_signature_31.png
| 2026-02-16 | 2026-02-16 | Pulupandan Rural Health Unit | Pulupandan, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 32113 | WHO | jackieorcajada@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 514 |  | GX4 6C R2 | 847255 | Due for PM and Xpertcheck | Preventive Maintenance - cleaned optic slots, module bays, plunger rods, filter. | Performed Xpertcheck Pulled out 2 pcs UPS batteries for recharging (12V 7AH) | All 4 modules Passed the Xpertcheck; 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470342095215021313/6470342095215021313_signature_28.png | JACKIE LOU O. ORCAJADA | https://www.jotform.com/uploads/macarewecare/222927978372067/6470342095215021313/6470342095215021313_signature_31.png
| 2026-02-16 | 2026-02-13 | Ilog RHU and Birthing Center | Dancalan, Ilog, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 32112 | PBSP | villavicencioemyrose@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 299 |  | GX4 6C | 807455 | Defective module A2/671781 and A4/885751. | Preventive maintenance - cleaned optics, module bays, plunger rods, filter; archive tests | - Replaced modules A2 and A4 with new module SN 801603 and 801599 respectively. - Performed Xpertcheck in all 4 modules | A2, A4 - Passed Xpertcheck A1, A3 - Requiring Service, For report to Jotform |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470167491013617916/6470167491013617916_signature_28.png | Emy Rose V. Garanganao | https://www.jotform.com/uploads/macarewecare/222927978372067/6470167491013617916/6470167491013617916_signature_31.png
| 2026-02-16 | 2026-02-13 | Veterans Memorial Medical Center | North Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | ralphcayetanormt@gmail.com |  | 32111 | MACARE-HBDC | vmmc.moleclab@gmail.com | 25 |  |  | GX4 10C | 120007529 | For installation | Installed unit Upload ADF Files Complete accessories Test run unit | Performed installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470165103619206503/6470165103619206503_signature_28.png | RALPH JOSEPH P. CAYETANO | https://www.jotform.com/uploads/macarewecare/222927978372067/6470165103619206503/6470165103619206503_signature_31.png
| 2026-02-16 | 2026-02-13 | Veterans Memorial Medical Center | North Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | ralphcayetanormt@gmail.com |  | 32110 | MACARE-HBDC | vmmc.moleclab@gmail.com | 25 |  |  | GX4 10 C | 120007530 | For installation | Installed unit Upload ADF files Complete accessories Test run unit | Performed installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470164573613516095/6470164573613516095_signature_28.png | RALPH JOSEPH P. CAYETANO | https://www.jotform.com/uploads/macarewecare/222927978372067/6470164573613516095/6470164573613516095_signature_31.png
| 2026-02-16 | 2026-02-14 | Philippine Red Cross | Port Area Manila | NCR | NCR 1st - Manila | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32109 | MACARE-HBDC | dolly.banares@redcross.org.ph | 27 |  |  | MCS + | 16K521SPG | No power | Checked voltage output of ups Checked and test power cord Test fuse continuity Replaced fuse Performed fuctionality test Performed simulation running | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470163563617725784/6470163563617725784_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6470163563617725784/6470163563617725784_signature_31.png
| 2026-02-16 | 2026-02-12 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32108 | DOH | wcpatron11x4@gmail.com | 25 |  |  | GXIV | 812534 | Unit for prefentive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470138133619335195/6470138133619335195_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/6470138133619335195/6470138133619335195_signature_31.png
| 2026-02-16 | 2026-02-13 | Philippine Red Cross Benguet | Lan Trinidad, Benguet | CAR | Benguet | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32107 | MACARE-HBDC | benguet@redcross.org.ph | 20 |  | 2026-02-11 | MCS+ | 17H636 | Unit for calibration | Adjusted centrifuge full scale value Replaced rotor roller rubbers | #NAME? | Unit is operational Calibration parameters are within acceptable range |  | MCS+ Urethane Roller Assy | 13739 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470124013617309755/6470124013617309755_signature_28.png | Lykaine Kate D. Bagano | https://www.jotform.com/uploads/macarewecare/222927978372067/6470124013617309755/6470124013617309755_signature_31.png
| 2026-02-16 | 2026-02-13 | Philippine Red Cross Benguet | Lan Trinidad, Benguet | CAR | Benguet | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32106 | MACARE-HBDC | benguet@redcross.org.ph | 20 |  | 2026-02-11 | MCS+ | 17H636 | Unit for calibration | Adjusted centrifuge full scale value Replaced rotor roller rubbers | #NAME? | Unit is operational Calibration parameters are within acceptable range |  | MCS+ Urethane Roller Assy | 13739 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470123483615230952/6470123483615230952_signature_28.png | Lykaine Kate D. Bagano | https://www.jotform.com/uploads/macarewecare/222927978372067/6470123483615230952/6470123483615230952_signature_31.png
| 2026-02-16 | 2026-02-13 | Far Eastern University - NRMF | Regalado, Fairview, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed.hhfi@healthway.com.ph |  | 32105 | CONSORTIUM | mdflorentin.hhfi@gmail.com | 23 | ASSAY: Xpert HBV VL, COUNT: 12 ASSAY: Xpert HPV, COUNT: 53 ASSAY: Xpert MTB/RIF Ultra, COUNT: 421 |  | GX-IV-D-6C | 840975 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -Icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6470099523615023358/6470099523615023358_signature_28.png | Maria Janella D. Florentin | https://www.jotform.com/uploads/macarewecare/222927978372067/6470099523615023358/6470099523615023358_signature_31.png
| 2026-02-15 | 2026-02-12 | Cotabato Provincial Hospital | Brgy.Amas, Kidapawan City, Cotabato | 12 | Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32103 | PBSP | januarybing1206@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 699 ASSAY: Xpert MTB-XDR, COUNT: 36 ASSAY: Xpert HIV-1 Viral Load, COUNT: 238 |  | GxIV-10C | 120000483 | Due for Xpertcheck | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance | PM done as per PM checklist document #:D20843 Xpertcheck done as per document #: 301-1896 rev. D | 4 operational modules Xpertcheck result All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6469734967324370546/6469734967324370546_signature_28.png | Hermosilla Enero | https://www.jotform.com/uploads/macarewecare/222927978372067/6469734967324370546/6469734967324370546_signature_31.png
| 2026-02-15 | 2026-02-11 | Koronadal City Health Office | Koronadal City, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32104 | PBSP | calvokenneth12@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2621 |  | GxIV-6C | 707529 | Module 4 H/W failed | Repair Gx Fan problem Perform Genexpert Cleaning and maintenance Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6469734967328484553/6469734967328484553_signature_28.png | Kenneth Paolo T. Calvo | https://www.jotform.com/uploads/macarewecare/222927978372067/6469734967328484553/6469734967328484553_signature_31.png
| 2026-02-13 | 2026-02-13 | Eversly childs sanitarium | Jagobiao, Mandaue | 7 | Cebu | Laboratory | NO | MAINTENANCE Site Visit | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32102 | PBSP | alimurungjane@gmail.com | 28 |  |  | Gx4 | 805270 | N/A | Site Visit Clean Fan Filter Gather data IQR, System log and Assay Stats | Site Visit Clean Fan Filter Gather data IQR, System log and Assay Stats | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6467764842124761576/6467764842124761576_signature_28.png | Juanita alimurung | https://www.jotform.com/uploads/macarewecare/222927978372067/6467764842124761576/6467764842124761576_signature_31.png
| 2026-02-13 | 2026-02-13 | Narra MHO | Narra, Palawan | 4B | Palawan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 32101 | PBSP | esg110478@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1812 | 2026-02-13 | GXIV | 807454 | For module replacement. Def SN: 685799 | Perforned module replacement and maintwnance. | Perfotmed module replacement. New Mod SN:663292 Performed maintenance and cleaning. | All 4 modules operational. |  | Def module SN: 685799 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6467760609254547746/6467760609254547746_signature_28.png | Elizabeth S. Gapulao | https://www.jotform.com/uploads/macarewecare/222927978372067/6467760609254547746/6467760609254547746_signature_31.png
| 2026-02-13 | 2026-02-13 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | TROUBLESHOOTING Data Collection | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32100 | DOH | vsmmdeptpatho.genexpert@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3560 |  | 6C GXIV-4D | 838285 | Error 5007 | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6467708494614422629/6467708494614422629_signature_28.png | Giovanni Heyrana | https://www.jotform.com/uploads/macarewecare/222927978372067/6467708494614422629/6467708494614422629_signature_31.png
| 2026-02-13 | 2026-02-13 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | TROUBLESHOOTING Data Collection | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32099 | DOH | vsmmdeptpatho.genexpert@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3560 |  | 6C GXIV-4D | 838285 | Error 5007 | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6467707024617715454/6467707024617715454_signature_28.png | Giovanni Heyrana | https://www.jotform.com/uploads/macarewecare/222927978372067/6467707024617715454/6467707024617715454_signature_31.png
| 2026-02-13 | 2026-02-11 | Graman Medical and Maternity Hospital | Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32098 | COMMERCIAL | gramanmolecularlab@gmail.com | 27 |  |  | GXIV | 11008530 | Printer problem | Removed colored cartridge | #NAME? | Unit is operational Printer is working normally Suspected defective colored cartridge (printer unable to detect) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6467588759586908486/6467588759586908486_signature_28.png | Maria Elizabeth S. de Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/6467588759586908486/6467588759586908486_signature_31.png
| 2026-02-13 | 2026-02-12 | Taguig City Health Office | Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32097 | PBSP | ginabuendumawal@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5025 |  | GX-IV-D-6C | 814372 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6467544183612333967/6467544183612333967_signature_28.png | Gina Dumawal | https://www.jotform.com/uploads/macarewecare/222927978372067/6467544183612333967/6467544183612333967_signature_31.png
| 2026-02-12 | 2026-02-11 | Cadiz City Health Office | Poblacion,Cadiz City | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32096 | PBSP | zhevebacomo@gmail.com | 23 |  |  | GXIV 6-C | 807614 | . Defective Module Serial Number: 861159 / C4 | .Gathered Report ( Pre.and Post Installation Qualifications Report & System Log Report. . Done replace defective module with SN:792502 | . Replace defective module C4 / SN : 861159. | All modules are functional okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6467107522812948005/6467107522812948005_signature_28.png | Zherlyn Eve Bacomo | https://www.jotform.com/uploads/macarewecare/222927978372067/6467107522812948005/6467107522812948005_signature_31.png
| 2026-02-12 | 2026-02-12 | Biopath Garden Heights | E. Rodriguez QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32095 | COMMERCIAL | biopathqc123@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 43 ASSAY: Others, COUNT: 2 ASSAY: Xpert HPV, COUNT: 106 ASSAY: Xpert MTB/RIF Ultra, COUNT: 40 ASSAY: Others, COUNT: 9 ASSAY: Xpert SARS-CoV-2, COUNT: 6 |  | Gx16 D 6C | 839306 | Due for maintenance UPS unable to back up due to battery worn out | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Checked battery output voltage Test run unit | Performed maintenance Performed troubleshooting | A1,A4,B2,B4,C2,D2 failed Xpertcheck A3 hw failed A2,B1,B3,C1,C3,C4,D1,D3,D4 operational | For replacement of battery 6pcs |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466997854211847019/6466997854211847019_signature_28.png | Rana Valerie Milanes | https://www.jotform.com/uploads/macarewecare/222927978372067/6466997854211847019/6466997854211847019_signature_31.png
| 2026-02-12 | 2026-02-12 | Northern Palawan Provincial Hospital | Taytay, Palawan | 4B | Palawan | Laboratory | NO | CALIBRATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Harold Toco |  | harolduno0905@gmail.com |  |  | 32094 | PBSP | npphclinlab2022@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 197 | 2026-02-12 | GXIV | 807671 | For Xpertcheck calibration. | Performed xpertcheck calibration. | Performed maintenance and cleaning. Run xpertcheck calibration. | 3 modules operational. 1 modul with H/W failed. |  | Xpertcheck kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466930293024638948/6466930293024638948_signature_28.png | Mark Jake Paduga | https://www.jotform.com/uploads/macarewecare/222927978372067/6466930293024638948/6466930293024638948_signature_31.png
| 2026-02-12 | 2026-02-12 | Davao de Oro Provincial Hospital - Montevista | Bangkerohan Sur, Montevista | 11 | Davao de Oro | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32093 | PBSP | cecillejagurilles1122@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1401 |  | GX-IV 6C 4D | 808624 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck on Modules 1, 2 and 3 only | Perform Preventive Maintenance and XpertCheck Calibration | Module 1 Retry and Skipped Module 2 Pass Module 3 Pass Module 4 Excluded due to H/W Failure Only 2 Modules Passed the Calibration | Module 1 for under observation, Only 3 XPC Cartidge Available on the site. Module 4 reported to higherups for assessment. And High ambient temperature, enduser was advised to achieve the optimal operating temperature |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466872732517004999/6466872732517004999_signature_28.png | Cecille A. Jagurilles,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6466872732517004999/6466872732517004999_signature_31.png
| 2026-02-12 | 2026-02-11 | Cordova Primary Health Care | Poblacion, Cordova Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | melodyjoynoelle0428@gmail.com |  | 32092 | PBSP | melodyjoynoelle0428@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 29 |  | GX4 | 808638 | -Cartridge stucked in Module 2 SN # 882343. - Module 2 error in position in ultrasonic horn. | - Removal of cartridge stucked in Module 2 SN # 882343. - Conduct preventive maintenance in all modules. - Clean up all plunger, Icore and cartridge bay. - Clean up air filter and dust inside the machine. | Removal of cartridge stucked in Module 2 SN # 882343. | - All modules are operational except module 2 SN: 882343 ultrasonic horn out of position. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466844631248323056/6466844631248323056_signature_28.png | Melody Arasan | https://www.jotform.com/uploads/macarewecare/222927978372067/6466844631248323056/6466844631248323056_signature_31.png
| 2026-02-12 | 2026-02-12 | Taguig Pateros District Hospital | Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32091 | DOH | ierlanegail@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3387 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1295 |  | GX-IV-D-6C | 831638 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module A1 and A2 waiting for module replacement. Module A3 has error 1001, 1017, 2005 and 2008. Reported in jotform. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466789301211925468/6466789301211925468_signature_28.png | Erlanegail Isla | https://www.jotform.com/uploads/macarewecare/222927978372067/6466789301211925468/6466789301211925468_signature_31.png
| 2026-02-12 | 2026-02-11 | Santa Lucia Main Health Center | Diablo street, snta Lucia district, pagadian city | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32090 | PBSP | baguioelsie@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 352 | 2026-02-11 | GX 4 - 6c | 808626 | Due xpertcheck calibration | -Pm done as per PM checklist document #D20843. -Xpertcheck done as per Document # 301-1896, Rev. D. | -Pm done as per PM checklist document #D20843. -Xpertcheck done as per Document # 301-1896, Rev. D. | All modules passed the xpertcheck calibration. | All modules passed. | N/a | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466708065811609312/6466708065811609312_signature_28.png | Elsie Baguio | https://www.jotform.com/uploads/macarewecare/222927978372067/6466708065811609312/6466708065811609312_signature_31.png
| 2026-02-12 | 2026-02-12 | Molave Rural Health unit | Molave, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32089 | DOH | aldrinacain@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1044 | 2026-02-12 | GX 4 - 6C | 838233 | System unit unable to turn on. | Preventive maintenance, conducted troubleshooting, replaced cmos batteries and rearrange RAM location. | Preventive maintenance, conducted troubleshooting, replaced cmos batteries and rearrange RAM location. | All modules are operational. System unit is operational and ready to use. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466707145813362251/6466707145813362251_signature_28.png | Aldrin Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/6466707145813362251/6466707145813362251_signature_31.png
| 2026-02-12 | 2026-02-11 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | YES | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 32088 | PBSP | ajunio16@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1847 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 29 |  | GX4 6C | 803109 | For module replacement B1 210000965 - 210072558 B2 642784 - 210073114 B3 hw failed B4 service required xc performed 8/22/25 | Performed module replacement Performed decon Gathered reports Cleaned plunger rod,filter,icore,cartridge bay Test run unit | Performed module replacement | B1,B2 operational B3,B4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466668943615456869/6466668943615456869_signature_28.png | Armando V. Junio Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6466668943615456869/6466668943615456869_signature_31.png
| 2026-02-12 | 2026-02-11 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | THERMO FISHER SCIENTIFIC | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32087 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 26 |  |  | Myspin 12 | 75004081 | Due for maintenance | Cleaned unit Checked wear and tear parts Performed functionality test Checked program parameters Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466668273616142523/6466668273616142523_signature_28.png | Jessica Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6466668273616142523/6466668273616142523_signature_31.png
| 2026-02-11 | 2026-02-11 | St. Luke's Medical Center - BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | CI | YES | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 32085 | COMMERCIAL | advancedmoleculardx@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 48 ASSAY: Xpert CT/NG, COUNT: 177 ASSAY: Xpert HBV VL, COUNT: 129 ASSAY: Xpert HPV, COUNT: 348 ASSAY: Xpert MTB/RIF Ultra, COUNT: 816 ASSAY: Xpert HCV VL, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 160 | 2026-02-05 | GX-IV-D-6C | 839231 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -icore -Filter Run XpertCheck Generated reports | Unit is operational. Module A2 service required, recommend to replace. |  | Item used: XpertCheck kit | SI# 102106 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/6466062184115673867/6466062184115673867_signature_28.png | Vincent Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/6466062184115673867/6466062184115673867_signature_31.png
| 2026-02-11 | 2026-02-11 | St. Luke's Medical Center - BGC | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | CI | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 32086 | COMMERCIAL | advancedmoleculardx@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 150 ASSAY: Xpert CT/NG, COUNT: 575 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 170 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 ASSAY: Xpert HPV, COUNT: 831 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2130 ASSAY: Xpert MTB-XDR, COUNT: 34 ASSAY: Xpert HCV VL, COUNT: 42 ASSAY: Xpert HIV-1 Viral Load, COUNT: 292 | 2026-02-05 | GX-IV-D-10C | 120002383 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -icore -Filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  | Item used: XpertCheck kit | SI# 102106 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/6466062184116921697/6466062184116921697_signature_28.png | Vincent Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/6466062184116921697/6466062184116921697_signature_31.png
| 2026-02-11 | 2026-02-11 | Cabatuan Rural Health Unit | Cabatuan, Iloilo | 6 | Iloilo | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32084 | PBSP | padrillanleonisa@gmail.com | 27 |  |  | 6C GXIV-4D | 805079 | Due for XPC | #NAME? | #NAME? | Machine is functional. Module 2 and 4 for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466059529916185783/6466059529916185783_signature_28.png | Leonisa Padrillan | https://www.jotform.com/uploads/macarewecare/222927978372067/6466059529916185783/6466059529916185783_signature_31.png
| 2026-02-11 | 2026-02-11 | Dinagat District Hospital | Escolta, Dinagat, Dinagat Islands | 13 | Dinagat Islands | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 32083 | PBSP | idhona@yahoo.com | 27 |  |  | GX4 10C | 110013918 | None | None | XpertCheck Calibration Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466030520517867355/6466030520517867355_signature_28.png | Idhona N. Tillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6466030520517867355/6466030520517867355_signature_31.png
| 2026-02-11 | 2026-02-11 | Digos City Health Office | Lapu Lapu St, Digos City | 11 | Davao del Sur | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32082 | PBSP | enzodino.lpb@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1321 | 2026-01-28 | GX-IV 6C 4D | 807464 | Defective Gateway Board Error at Module 1 Due for Preventive Maintenance | Replaced the Defective Gateway Board SN 3010538 with new Gateway Board SN 910767915 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Gateway Board Replacement Perform Module Troubleshooting Perform Preventive Maintenance | Module 1 Disabled due to Hardware Failure/ Error 4004 Only Three Modules are operational | Damaged Module Door; Broke by NCR Atleos Engineer | PN 700-6140S | Monarch DR NO 0054999 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466002339019954600/6466002339019954600_signature_28.png | Lorenzo Paolo Butaslac | https://www.jotform.com/uploads/macarewecare/222927978372067/6466002339019954600/6466002339019954600_signature_31.png
| 2026-02-11 | 2026-02-11 | Snta lucia Main Health Center | Damon street, santa Lucia, pagadian city | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32081 | PBSP | baguioelsie@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 352 | 2026-02-11 | GX 4 - 6c | 808626 | Due Xpertcheck calibration. | -PM done as per PM checklist document #D20843 -Xpertcheck done as per Document #301-1896, Rev. D. | -PM done as per PM checklist document #D20843 -Xpertcheck done as per Document #301-1896, Rev. D. | All modules passed the calibration. All modules are operational. UPS broken and unable to be repaired. | UPS ( un interruptible power supply) must be purchased and installed as soon as possible.( ASAP) | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6466000354673156556/6466000354673156556_signature_28.png | Elsie Baguio | https://www.jotform.com/uploads/macarewecare/222927978372067/6466000354673156556/6466000354673156556_signature_31.png
| 2026-02-11 | 2026-02-11 | Western Visayas Medical Center | Iloilo City | 6 | Iloilo | Sub National Laboratory | NO | ON-SITE ASSISTANCE Transfer GeneXpert Unit | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32080 | UNOPS | zialjsparcia1116@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1734 |  | 10C GXIV-4D | 110023529 | #NAME? | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6465998130168071815/6465998130168071815_signature_28.png | Zialjs Jayme | https://www.jotform.com/uploads/macarewecare/222927978372067/6465998130168071815/6465998130168071815_signature_31.png
| 2026-02-11 | 2026-02-11 | Manapla Rural Health Unit | Barangay 1A Manapla | 6 | Negros Occidental | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32079 | PBSP | magalonaruderic@gmail.com | 23 |  |  | GXIV 4 6C | 805093 | Module A2,A3 Hardware fail. | . Save System log file . Save Installation Qualifications Report. . Save assay statistics Report . Save GXC file. | Perform preventive maintenance for all modules. Clean filter. | Module A1,A4 okay to use. Module A2,A3 not okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6465994830516745612/6465994830516745612_signature_28.png | Ruderic f. Magalona | https://www.jotform.com/uploads/macarewecare/222927978372067/6465994830516745612/6465994830516745612_signature_31.png
| 2026-02-11 | 2026-02-11 | Calaanan Health Center | Nadurata St. Caloocan City | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32078 | DOH | rosaleslorina0917@gmail.com | 25 |  |  | GXIV | 831892 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module 1, 2 and 3 Service Requires. Module 4 passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6465956221873953635/6465956221873953635_signature_28.png | Lorina Rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/6465956221873953635/6465956221873953635_signature_31.png
| 2026-02-11 | 2026-02-11 | Philippine Tuberculosis Society, Inc. - Iloilo | Iloilo City | 6 | Iloilo | TB Dots | NO | INSTALLATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32077 | MACARE-HBDC | rdureza.iccd@philtbsociety.org | 24 |  |  | 10C GXIV-4D | 120006949 | None | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6465869194763044297/6465869194763044297_signature_28.png | Rhoda Anne C. Dureza | https://www.jotform.com/uploads/macarewecare/222927978372067/6465869194763044297/6465869194763044297_signature_31.png
| 2026-02-10 | 2026-02-10 | Cebu TB Reference Lab | SRP-CEBU | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32076 | PBSP | asmantiquilla.pbsp@gmail.com | 21 |  |  | Gx4 | 120000133 | Due date for xpertcheck calibration | Gather data IQR, System log, Assay stats & Calibration data Clean Fan Filter | Xpertcheck Calibration Gather data IQR, System log, Assay stats & Calibration data Clean Fan Filter | Module 1,2 & 3 passed and module 4 service require during xpert check |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6465199208018006245/6465199208018006245_signature_28.png | Allan S. Mantiquillla | https://www.jotform.com/uploads/macarewecare/222927978372067/6465199208018006245/6465199208018006245_signature_31.png
| 2026-02-10 | 2026-02-10 | Metro Health Specialists Hospital Inc | Sorsogon City | 5 | Sorsogon | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32075 | COMMERCIAL | mhshimolecularpathology@gmail.com | 26 | ASSAY: Xpert HBV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 43 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX4-10C | 120002716 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6465183595316891901/6465183595316891901_signature_28.png | Shawn Manalansan | https://www.jotform.com/uploads/macarewecare/222927978372067/6465183595316891901/6465183595316891901_signature_31.png
| 2026-02-10 | 2026-02-06 | Casimiro Ynares Sr. Memorial Hospital | Becco Street, Rodriguez, 1860 Rizal | 4A | Rizal | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32074 | PBSP | caysmhlab@gmail.com | 27 |  |  | GX4 6C | 831642 | Due for xpertcheck Unable to perform xpertcheck due to workstation booting problem | Checked workstation Reset all connection Reset memory After troubleshooting is performed issue still occur | Performed troubleshooting | Unit unoperational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6465133513614251878/6465133513614251878_signature_28.png | Marjorie Ann Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6465133513614251878/6465133513614251878_signature_31.png
| 2026-02-10 | 2026-02-10 | Esperanza Rural Health Unit | Esperanza, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32073 | PBSP | nilosubido722@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 573 |  | GxlV - 6C | 807405 | Due for XPC | Pm Done as per PM checklist doc. # D20843 Xpertcheck Done as per document# 301-1896 Checking of modules status and Errors Data collection | Genexpert machine Xpertcheck Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6465114161618840403/6465114161618840403_signature_28.png | Nilo P. Subido | https://www.jotform.com/uploads/macarewecare/222927978372067/6465114161618840403/6465114161618840403_signature_31.png
| 2026-02-10 | 2026-02-10 | Medi Linx Laboratory Inc (CLDH) | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32072 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 |  |  | GXIV-D 6C | 839232 | CPU not booting up CPU SN: CPH0420F06989 | CPU troubleshooting | Performed CPU troubleshooting Performed functionality test | Unit is operationl | CPU under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464991315367305063/6464991315367305063_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6464991315367305063/6464991315367305063_signature_31.png
| 2026-02-09 | 2026-02-06 | Sorsogon Provincial Hospital | Sorsogon City | 5 | Sorsogon | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32071 | PBSP | carranzadoloiras21@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1541 ASSAY: Xpert MTB-XDR, COUNT: 75 |  | GX4-10C | 110020680 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464443540901002467/6464443540901002467_signature_28.png | Maria Sheila C. Doloiras | https://www.jotform.com/uploads/macarewecare/222927978372067/6464443540901002467/6464443540901002467_signature_31.png
| 2026-02-09 | 2026-02-05 | Estevez Memorial Hospital, Inc. | Rizal st, Legazpi | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32070 | COMMERCIAL | labestevez@gmail.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 6 ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 201 ASSAY: Xpert MTB/RIF Ultra, COUNT: 175 ASSAY: Xpert MTB-XDR, COUNT: 7 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 2 |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, fan filter Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464442480908724300/6464442480908724300_signature_28.png | Estevez Memorial Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6464442480908724300/6464442480908724300_signature_31.png
| 2026-02-09 | 2026-02-04 | Bicol Regional Hospital and Medical Center | Sagpon, Daraga | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32069 | MACARE-HBDC | brtth.covid19lab@gmail.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 278 ASSAY: Xpert HPV, COUNT: 73 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 ASSAY: Xpert HPV, COUNT: 7 |  | GX4-6C | 110026767 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464441730902112797/6464441730902112797_signature_28.png | Jennifer Vitasa | https://www.jotform.com/uploads/macarewecare/222927978372067/6464441730902112797/6464441730902112797_signature_31.png
| 2026-02-09 | 2026-02-06 | Philippine Red Cross QC | Kalayaan Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32068 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 27 |  |  | MCS + | 02A227 | Due for maintenance | Cleaned unit Checked rollers Checked and line sensor value Checked and test keypad and lcd display. Checked program parameters Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464359264212857492/6464359264212857492_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/6464359264212857492/6464359264212857492_signature_31.png
| 2026-02-09 | 2026-02-05 | East Avenue QC | East Avenue QC | NCR | NCR 2nd - Quezon City | Blood donor room | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 32067 | MACARE-HBDC | paulinobernaldo1419@gmail.com | 26 |  |  | MCS+ | 15H044 | Due for maintenance and calibration | Cleaned unit Checked rollers Checked line sensor Checked and test keypad and lcd display. Checked program parameters Performed calibration Performed simulation running Performed functionality test | Performed calibration and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464358534212493159/6464358534212493159_signature_28.png | Christopher Paul Bernaldo | https://www.jotform.com/uploads/macarewecare/222927978372067/6464358534212493159/6464358534212493159_signature_31.png
| 2026-02-09 | 2026-02-09 | Danao City Health Office | Danao City, Cebu | 7 | Cebu | Laboratory | YES | INSTALLATION TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32066 | PBSP | chockenriquez@gmail.com | 25 |  |  | Gx 4 | 808630 | Defective CPU from Bogo CHO genexpert unit | Install personal PC from end user Install Genexpert Software version 5.1 Install assay MTB ultra and HIV viral load | Install personal PC from end user Install Genexpert Software version 5.1 Install assay MTB ultra and HIV viral load | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464227748724462247/6464227748724462247_signature_28.png | Nicanor Enriquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6464227748724462247/6464227748724462247_signature_31.png
| 2026-02-09 | 2026-02-09 | Soccsksargen General Hospital | Surallah, South Cotabato | 12 | South Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32065 | PBSP | irisrano@ymail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1889 ASSAY: Xpert MTB-XDR, COUNT: 19 |  | GxIV-10C | 120000172 | None | Perform Genexpert maintenance and cleaning Module cleaning Plunger rod and optics cleaning Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464177647426245816/6464177647426245816_signature_28.png | Iris Rano | https://www.jotform.com/uploads/macarewecare/222927978372067/6464177647426245816/6464177647426245816_signature_31.png
| 2026-02-09 | 2026-02-06 | Danao City Health Office | Danao City | 7 | Cebu | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32064 | PBSP | chockenriquez@gmail.com | 28 |  |  | Gx4 | 808630 | CPU problem PC won't turn On | Replace new CMOS Battery Reset and Clean RAM | Reset and Clean RAM Replace new CMOS Battery | PC won't turn on |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464159538722477357/6464159538722477357_signature_28.png | Nicanor Enriquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6464159538722477357/6464159538722477357_signature_31.png
| 2026-02-09 | 2026-02-09 | St. Louis Molecular Diagnostic Laboratory | Brgy. New Carmen Tacurong City, Sultan Kudarat | 12 | Sultan Kudarat | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32063 | COMMERCIAL | stlouismoleculardiagnosticslab@gmail.com | 22.8 | ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 29 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GX-IV 10C 4D | 110024270 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6464155786444129923/6464155786444129923_signature_28.png | Shane Delposo | https://www.jotform.com/uploads/macarewecare/222927978372067/6464155786444129923/6464155786444129923_signature_31.png
| 2026-02-06 | 2026-02-06 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32062 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 16K521SPG | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461753362128600643/6461753362128600643_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6461753362128600643/6461753362128600643_signature_31.png
| 2026-02-06 | 2026-02-06 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32061 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461753062129615758/6461753062129615758_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6461753062129615758/6461753062129615758_signature_31.png
| 2026-02-06 | 2026-02-06 | Jose B. Lingad Memorial General Hospital | City of San Fernando Pampanga | 3 | Pampanga | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32060 | PBSP | jblmgh.mtblab@gmail.com | 25 |  |  | GXIV-D 6C | 804763 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational All modules passed calibration C1:693690- Passed C2:721058- Passed C3: 691412- Passed C4:691294- Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461718874329045704/6461718874329045704_signature_28.png | Earvin Justin Carlos | https://www.jotform.com/uploads/macarewecare/222927978372067/6461718874329045704/6461718874329045704_signature_31.png
| 2026-02-06 | 2026-02-06 | Pililla Rural Health Unit | Pililla, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32059 | DOH | tbdotspililla@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 981 |  | GX-IV-D-6C | 821293 | For XpertCheck | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461685539629771596/6461685539629771596_signature_28.png | Ayen Olaya | https://www.jotform.com/uploads/macarewecare/222927978372067/6461685539629771596/6461685539629771596_signature_31.png
| 2026-02-06 | 2026-02-05 | Cainta City Health Office | Karangalan Dr, Cainta, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32057 | DOH | ballesterrichard18@gmail.com | 22 | ASSAY: Others, COUNT: 57 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1112 ASSAY: Xpert HIV-1 Viral Load, COUNT: 146 |  | GX-IV-D-6C | 831623 | For XpertCheck | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -icore -Filter Run XpertCheck Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461673689622393207/6461673689622393207_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/6461673689622393207/6461673689622393207_signature_31.png
| 2026-02-06 | 2026-02-05 | Cainta City Health Office | Karangalan Dr, Cainta, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32058 | PBSP | ballesterrichard18@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 893 |  | GX-IV-D-6C | 804687 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461673689628093868/6461673689628093868_signature_28.png | Richard 6 | https://www.jotform.com/uploads/macarewecare/222927978372067/6461673689628093868/6461673689628093868_signature_31.png
| 2026-02-06 | 2026-02-06 | Panabo City Health Office | Poblacion, Panabo City Davao del Norte | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32056 | DOH | nadineeust67@gmail.com | 24.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2190 |  | GX-IV 6C 4D | 838295 | For Preventive Maintenance Barcode reader error | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Check the Barcode if it is operational | Perform Preventive Maintenance and Troubleshooting | All Modules are operational Barcode reader is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461635238453965958/6461635238453965958_signature_28.png | Nadine Gemma Eustaquio,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6461635238453965958/6461635238453965958_signature_31.png
| 2026-02-06 | 2026-02-06 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 32055 | DOH | riazarizzamae@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 684 |  | GXIV-6C | 831637 | Due for Calibration | Performed Calibration (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461618334127284726/6461618334127284726_signature_28.png | Rizza Mae Lagazo | https://www.jotform.com/uploads/macarewecare/222927978372067/6461618334127284726/6461618334127284726_signature_31.png
| 2026-02-06 | 2026-02-06 | Danao City Health Office | Danao City | 7 | Cebu | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32054 | PBSP | chockenriquez@gmail.com | 28 |  |  | Gx4 | 808630 | CPU problem PC won't turn On | Replace new CMOS Battery Reset and Clean RAM | Reset and Clean RAM Replace new CMOS Battery | PC won't turn on |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461582283586213824/6461582283586213824_signature_28.png | Nicanor Enriquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6461582283586213824/6461582283586213824_signature_31.png
| 2026-02-06 | 2026-02-04 | Cebu Tb Reference Lab-SRP | SRP-Cebu | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32053 | COMMERCIAL | asmantiquilla.pbsp@gmail.com | 28 |  |  | Gx 4 | 120000133 | Due Date for XPC | Site Visit/PM Gather data IQR, System log | Site Visit/PM Gather data IQR, System log | Systems operational Waiting Approval of Cephied to perform XPC calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6461581143581799211/6461581143581799211_signature_28.png | Allan Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6461581143581799211/6461581143581799211_signature_31.png
| 2026-02-05 | 2026-02-05 | Baliuag District Hospital | Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32052 | PBSP | reyeshannahxx@gmail.com | 25 |  |  | GXIV | 812770 | Unit for maintenance | Replaced filter | #NAME? | Unit is operational Unit is dur for calibration (28 Feb 2026) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6460846458304561985/6460846458304561985_signature_28.png | Hannah Reyes - Fajardo | https://www.jotform.com/uploads/macarewecare/222927978372067/6460846458304561985/6460846458304561985_signature_31.png
| 2026-02-05 | 2026-02-05 | San Rafael Rural Health Unit | San Rafael, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32051 | DOH | tatsdayrit25@gmail.com | 25 |  |  | GXIV | 832692 | Unit for calibration | Checked unit Workstation | #NAME? | Unit is not operational Unable to perform XpertCheck due to workstation problem |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6460825615425615345/6460825615425615345_signature_28.png | Joana Manelle Dayrit | https://www.jotform.com/uploads/macarewecare/222927978372067/6460825615425615345/6460825615425615345_signature_31.png
| 2026-02-05 | 2026-02-05 | Tuburan rural health unit | Poblacion Tuburan Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32050 | PBSP | romero_rissa@yahoo.com | 23 |  | 2026-02-05 | GXIV 6 C | 815485 | Site Visit | . Save Installation Qualifications Report . Save System Log file . Save Assay Statistics Report | Perform preventive maintenance. | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6460762297113238340/6460762297113238340_signature_28.png | Rissa R. Diamante | https://www.jotform.com/uploads/macarewecare/222927978372067/6460762297113238340/6460762297113238340_signature_31.png
| 2026-02-05 | 2026-02-05 | Tuburan rural health unit | Poblacion Tuburan Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32049 | PBSP | romero_rissa@yahoo.com | 23 |  | 2026-02-05 | GXIV 6 C | 815485 | Site Visit | . Save Installation Qualifications Report . Save System Log file . Save Assay Statistics Report | Perform preventive maintenance. | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6460762277119460164/6460762277119460164_signature_28.png | Rissa R. Diamante | https://www.jotform.com/uploads/macarewecare/222927978372067/6460762277119460164/6460762277119460164_signature_31.png
| 2026-02-05 | 2026-02-04 | Tropical Disease Foundation | Amorsolo st., Makati | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | krcueno@tdf.org.ph |  | 32048 | COMMERCIAL | wlsalarda@tdf.org.ph | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 196 ASSAY: Xpert MTB-XDR, COUNT: 33 |  | GX4 10C | 110026840 | Due for maintenance | Operational | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6460751151324289825/6460751151324289825_signature_28.png | Wyeth Kevin Salarda | https://www.jotform.com/uploads/macarewecare/222927978372067/6460751151324289825/6460751151324289825_signature_31.png
| 2026-02-05 | 2026-02-05 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32047 | MACARE-HBDC | lcpmicrolab@gmail.com | 26 |  |  | GX4 D(10C) | 110027386 | For installation | Delivered and installed unit Checked voltage output Checked accessories complete Upload ADF Generate IQR Test run unit ok | Performed installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6460750661326812986/6460750661326812986_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6460750661326812986/6460750661326812986_signature_31.png
| 2026-02-05 | 2026-02-04 | MMC Medilinx | Amorsolo st Makati | NCR | NCR 4th - Makati | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32046 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 275 ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Others, COUNT: 2 ASSAY: Xpert CT/NG, COUNT: 716 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 108 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 ASSAY: Xpert HPV, COUNT: 418 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3827 ASSAY: Others, COUNT: 20 ASSAY: Xpert HCV VL, COUNT: 193 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1744 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 12 |  | GX16 D 10C | 110026448 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | A1 disabled 15 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6460749811327573927/6460749811327573927_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6460749811327573927/6460749811327573927_signature_31.png
| 2026-02-05 | 2026-02-05 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32045 | PBSP | cdh.laboratorydept@gmail.com | 25 |  |  | GXIV-D 6C | 807562 | Due for Calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Xpertcheck Performed functionality test Gathered reports | Unit is operational Passed Calibration C1:787181 C2:210011996 C4:692832 Failed Calibration C3: 210068087 | C3: 210068087 for reporting and replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6460722913412673032/6460722913412673032_signature_28.png | Maria Rowena T. Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6460722913412673032/6460722913412673032_signature_31.png
| 2026-02-05 | 2026-02-05 | Mindanao central sanitarium | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | TROUBLESHOOTING Genexpert and system unit assessment | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 32044 | PBSP | ybioamie@gmail.com | 22 |  | 2026-02-05 | GX 4 | 821214 | Cpu unable to boot up properly, monitor lcd broken, 2 modules are down. | Genexpert machine and system unit assessment. Cpu and monitor already been troubleshoot and must be replaced immediately. | Genexpert machine and system unit assessment. Troubleshooted system unit due to unable to start up properly. | Broken monitor with serial number CNK831071F and a broken system unit with a serial number 2UA8471RTK for replacement and has been reported and confirmed by the end user. 2 new modules are already in site with a serial number of 865399 and 805020 that needs to be serviced. | Monitor and cpu are on the way as per end user, while 2 new modules are already in the laboratory but not have been replaced yet. | N/A | N/A | Gian Carlo Aizon | https://www.jotform.com/uploads/macarewecare/222927978372067/6460662933186850314/6460662933186850314_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/6460662933186850314/6460662933186850314_signature_31.png
| 2026-02-04 | 2026-02-04 | Malita Rural Health Unit | Poblacion, Malita Davao Occidental | 11 | Davao Occidental | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | erguerrostella@gmail.com |  | 32043 | PBSP | rhumalitalaboratory@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 881 |  | GX-IV 6C 4D | 805258 | Errors at Four Modules Due For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Troubleshooting | Module 1 and 2 - Disabled Only Module 3 and 4 are working and for under observation | Data and other documentation recorded and send to higherups for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6460013190772539270/6460013190772539270_signature_28.png | Maria Stella E. Solitana | https://www.jotform.com/uploads/macarewecare/222927978372067/6460013190772539270/6460013190772539270_signature_31.png
| 2026-02-04 | 2026-02-04 | University of The East Ramon Magsaysay Memorial Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com | uermbloodbank@gmail.com |  | 32042 | MACARE-HBDC | biomed@uerm.edu.ph | 25 |  |  | MCS+ | 13C270 | Unit for maintenance | NA | #NAME? | Unit is operational Calibration parameters are within normal range |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459927493231399435/6459927493231399435_signature_28.png | Ruel | https://www.jotform.com/uploads/macarewecare/222927978372067/6459927493231399435/6459927493231399435_signature_31.png
| 2026-02-04 | 2026-02-12 | Medi Linx Laboratory Inc. ( CLDH) | Hospital Drive Brgy San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32041 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 |  |  | GXII-D | 830841 | Due for Monthly PMS | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module C2:62096 fir module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459923546415929983/6459923546415929983_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6459923546415929983/6459923546415929983_signature_31.png
| 2026-02-04 | 2026-02-04 | Medi Linx Laboratory Inc (CLDH) | Hospital drive Brgy. San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 32040 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 |  |  | GXIV-D 6C | 839232 | Due for Monthly PMS | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module B4:738203 for Module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459923016414324690/6459923016414324690_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6459923016414324690/6459923016414324690_signature_31.png
| 2026-02-04 | 2026-02-04 | Toledo City Health Office | Poblacion Toledo City,Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32039 | PBSP | mljlongakit@gmail.com | 24.4 |  | 338220 | GX4 6C | 808640 | Always booting module B4 653904. | .Save Installation Qualifications Report . Save system log file. . Save Assay Statistics Report | Perform self test not functioning always booting. | Module B1,2,3 are functional .Not ok to use module B4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459915554513489821/6459915554513489821_signature_28.png | Lorie Mae Diapen-Longakit | https://www.jotform.com/uploads/macarewecare/222927978372067/6459915554513489821/6459915554513489821_signature_31.png
| 2026-02-04 | 2026-02-04 | Isabela city health office | Isabela, Basilan | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Genexpert and system unit assessment. | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | hkdrdsales@gmail.com |  | 32038 | PBSP | reynantemontano@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1210 | 2026-02-04 | GX 4 | 807595 | Module 3 failed the xpertcheck last April 13, 2025 and not yet been replaced. | Preventive Maintenence, assessment of Genexpert Machine and the system unit. | Preventive Maintenence, assessment of Genexpert Machine and the system unit. | Module 3 pending for replacement due to xpertcheck result last April 2025. Xpertcheck calibration due date is April 2026. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459900689864310063/6459900689864310063_signature_28.png | Reynante Montaño | https://www.jotform.com/uploads/macarewecare/222927978372067/6459900689864310063/6459900689864310063_signature_31.png
| 2026-02-04 | 2026-02-04 | Jaro Rural Health Unit | Jaro Leyte | 8 | Leyte | TB Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32037 | PBSP | tbdotslaboratoryjaroleyte@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 75 |  | GXIV 2 | 834817 | CPU not turning ON | #NAME? | #NAME? | Unit is now operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459893324536596517/6459893324536596517_signature_28.png | Jocelyn Pagpaguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/6459893324536596517/6459893324536596517_signature_31.png
| 2026-02-04 | 2026-02-04 | Basilan General Hospital | Basilan General Hospital | 9 | Basilan | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | hkdrdsales@gmail.com |  | 32036 | PBSP | basgenlab@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 513 | 2026-02-04 | GX 4 | 707515 | Module 1 confirmed h/w failed and cannot be use. | Preventive Maintenence for genexpert machine, condition assessment for genexpert machine and system unit including ups. | Preventive Maintenence for genexpert machine, condition assessment for genexpert machine and system unit including ups. Checked module temperatures. | Module 1 confirmed h/w failed and cannot be use. Over all assessment Genexpert machine all four modules need to undergo xpertcheck calibration. | Requires calibration | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459886469867162453/6459886469867162453_signature_28.png | Ahmadil-kubaysi m. Sawadjaan | https://www.jotform.com/uploads/macarewecare/222927978372067/6459886469867162453/6459886469867162453_signature_31.png
| 2026-02-04 | 2026-02-03 | St Luke's Medical Center | E.Rodriguez | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 32035 | COMMERCIAL | megapo@stlukes.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 223 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2240 |  | GX4 6C | 839234 | As per end user they encounter intermittent temp error 1002 in B1 module | Checked unit Checked fan ok operational Checked filter Gathered reports | Performed troubleshooting | B3,B4 disabled B1 for reporting advise end user not to use b1 module due to temp error. B2 in use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459882847121949016/6459882847121949016_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6459882847121949016/6459882847121949016_signature_31.png
| 2026-02-04 | 2026-02-04 | Toledo City Health Office | Poblacion Toledo City,Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32034 | PBSP | mljlongakit@gmail.com | 24.4 |  | 338220 | GX4 6C | 808640 | Always booting module B4 653904. | .Save Installation Qualifications Report . Save system log file. . Save Assay Statistics Report | Perform self test not functioning always booting. | Module B1,2,3 are functional .Not ok to use module B4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459842054517312835/6459842054517312835_signature_28.png | Lorie Mae Diapen-Longakit | https://www.jotform.com/uploads/macarewecare/222927978372067/6459842054517312835/6459842054517312835_signature_31.png
| 2026-02-04 | 2026-02-03 | Aloguinsan Rural Health Unit | Poblacion Aloguinsan Cebu | 7 | Cebu | TB Laboratory | NO | INSTALLATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32033 | PBSP | aloguinsanrhulaboratory@gmail.com | 23 |  |  | GX4 6C | 808620 | GeneXpert machine installation. | . Save system log file . Save Installation Qualifications Report. . Save Assay Statistics Report | Install GeneXpert Machine Install CPU Install LCD Install UPS | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459835704511010203/6459835704511010203_signature_28.png | Jennifer Sayson | https://www.jotform.com/uploads/macarewecare/222927978372067/6459835704511010203/6459835704511010203_signature_31.png
| 2026-02-04 | 2026-02-03 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | jadepelaez@gmail.com |  | 32030 | PBSP | molecularlab@rmc.doh.gov.ph | 24 | ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4783 ASSAY: Xpert HIV-1 Viral Load, COUNT: 182 |  | GX-IV-D-6C | 814374 | For maintenance Error 5007 on modules B2 and B3 | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Modules B2 and B3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459835553619221905/6459835553619221905_signature_28.png | Justine Lorenz V. Falzado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6459835553619221905/6459835553619221905_signature_31.png
| 2026-02-04 | 2026-02-03 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | jadepelaez@gmail.com |  | 32031 | COMMERCIAL | rmcbloodbank@gmail.com | 24 |  |  | MCS+ | 15H045G | For maintenance | Performed maintenance | Run simulation Verified calibration parameters Checked rotor pumps Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459835553619794623/6459835553619794623_signature_28.png | Ferdinand B. Barroga, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6459835553619794623/6459835553619794623_signature_31.png
| 2026-02-04 | 2026-02-03 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | jadepelaez@gmail.com |  | 32032 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 7 ASSAY: Xpert HPV, COUNT: 58 ASSAY: Xpert MTB/RIF Ultra, COUNT: 317 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 445 |  | GX-XVI-D-6C | 838627 | For maintenance Error 2034 and 4014 on module C3 | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module C3 not operational. For reporting. Modules A2, A4, B1, B3, C4, D4 not operational. Reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459835563613550847/6459835563613550847_signature_28.png | Justine Lorenz Falzado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6459835563613550847/6459835563613550847_signature_31.png
| 2026-02-04 | 2026-02-04 | Juan Posadas Health Center | Rodriguez St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 32029 | PBSP | lizjunio67@gmail.com | 25 |  |  | GXIV | 805060 | Unit for preventive maintenance. Loss communication error on module 4 | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. Performed troubleshooting on module 4. | Unit is operational. Module 4 has broken ultrasonic horn cable. Module 4 for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459789155338126613/6459789155338126613_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/6459789155338126613/6459789155338126613_signature_31.png
| 2026-02-04 | 2026-02-04 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | TB Dots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 32028 | DOH | vsmmcdeptpatho.genexpert@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2675 |  | 6C GXIV-4D | 820206 | Module A2 and A4 encountered error 5007 | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459744288022993051/6459744288022993051_signature_28.png | Leah Karen Noynay | https://www.jotform.com/uploads/macarewecare/222927978372067/6459744288022993051/6459744288022993051_signature_31.png
| 2026-02-03 | 2026-01-29 | Tampakan Rural Health Unit | Poblacion Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 32027 | DOH | jmedtech_lab@yahoo.com | 27.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 879 |  | GxlV-6C | 838544 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and optics cleaning Fan and filter cleaning Checking of module cleaning Data collection | Perform Genexpert machine preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459327107327319386/6459327107327319386_signature_28.png | Jovi Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/6459327107327319386/6459327107327319386_signature_31.png
| 2026-02-03 | 2026-02-03 | Alang Alang Rural Health Unit | Alang- Alang Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 32026 | WHO | vinagatela@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 665 |  | GXIV | 847249 | Error 2123: Lost Communication | TROUBLESHOOTING Check Module Connections - Properly connected Check Environmental Conditions Check Network and Security Settings - Disabled both private and public network Check System Performance Check Gateway Board Revision Check Ultrasonic Horn Cable - No problem encountered Check power supply - 24VDC output Check the F1 fuse on the module - Connected - Preformed self test - Preformed Preventive Maintenance - Temporary change UPS due to low voltage output - Change CMOS battery | #NAME? | Unit is Operational | Low voltage output UPS |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459123731617976701/6459123731617976701_signature_28.png | Vina Milagros G. Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6459123731617976701/6459123731617976701_signature_31.png
| 2026-02-03 | 2026-02-03 | Danao City Health Office | Danao City, Cebu | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 32025 | PBSP | chockenriquez@gmail.com | 27 |  |  | Gx4 | 808630 | N/A | Gather data IQR | Gather data IQR Check machine capability | Systems operational Minimal slow CPU speed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459062213518196640/6459062213518196640_signature_28.png | Nicanor Enriquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6459062213518196640/6459062213518196640_signature_31.png
| 2026-02-03 | 2026-02-03 | Davao Reproductive Health and Wellness Center | Jacinto St,. Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32024 | MACARE-HBDC | rhwclab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 405 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3245 |  | GX-IV 6C 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Module 3 encounter multiple temperature related errors module is currently disable. Only Three Modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459051041416182426/6459051041416182426_signature_28.png | Kent ferolin | https://www.jotform.com/uploads/macarewecare/222927978372067/6459051041416182426/6459051041416182426_signature_31.png
| 2026-02-03 | 2026-02-03 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 32023 | DOH | rhwclab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 405 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3245 |  | GX-IV 6C 4D | 838262 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459050491418046750/6459050491418046750_signature_28.png | Kent ferolin | https://www.jotform.com/uploads/macarewecare/222927978372067/6459050491418046750/6459050491418046750_signature_31.png
| 2026-02-03 | 2026-02-03 | Likhaan Center For Women’s Health | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 32022 | MACARE-HBDC | clatras@likhaan.org | 25 |  |  | GXIV 10C | 110025174 | Unit for maintenance | Na | #NAME? | Unit is operational No hardware related errors found |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459048718022139716/6459048718022139716_signature_28.png | Jocelyn Latras | https://www.jotform.com/uploads/macarewecare/222927978372067/6459048718022139716/6459048718022139716_signature_31.png
| 2026-02-03 | 2026-02-03 | Aloguinsan Rural Health Unit | Poblacion Aloguinsan Cebu | 7 | Cebu | TB Laboratory | NO | INSTALLATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 32021 | PBSP | aloguinsanrhulaboratory@gmail.com | 23 |  |  | GX4 6C | 808620 | GeneXpert machine installation. | . Save system log file . Save Installation Qualifications Report. . Save Assay Statistics Report | Install GeneXpert Machine Install CPU Install LCD Install UPS | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6459038544615993868/6459038544615993868_signature_28.png | Jennifer Sayson | https://www.jotform.com/uploads/macarewecare/222927978372067/6459038544615993868/6459038544615993868_signature_31.png
| 2026-02-03 | 2026-01-29 | Libacao Rural Health Unit | Libacao, Aklan | 6 | Aklan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 32020 | DOH | haladeeb18@gmail.com | 26 |  |  | GX4 6C | 838540 | #NAME? | . | - Replaced batteries for UPS - 2 pcs. 12V 7AH - Replaced LCD with brand new LCD monitor with SN H22V8VL658A250716622 | - 4 modules operational - damaged USB cable for Barcode Scanner |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458951220123969696/6458951220123969696_signature_28.png | Maria Hala Deeb Zaulda | https://www.jotform.com/uploads/macarewecare/222927978372067/6458951220123969696/6458951220123969696_signature_31.png
| 2026-02-02 | 2026-01-28 | Tabaco City Health Unit | Poblacion, Tabaco City | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32019 | PBSP | elyjoy17rmt@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4980 ASSAY: Xpert MTB-XDR, COUNT: 60 |  | GX4-10C | 110019623 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is operational All modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458524796013892064/6458524796013892064_signature_28.png | Ely Joy Hilario-Cao | https://www.jotform.com/uploads/macarewecare/222927978372067/6458524796013892064/6458524796013892064_signature_31.png
| 2026-02-02 | 2026-01-20 | Masbate City Health Office | City Hall Compound, Masbate City | 5 | Masbate | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 32018 | UNOPS | deosenovilla@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1862 ASSAY: Xpert MTB-XDR, COUNT: 52 |  | GX4-10C | 110023512 | Gx for Calibration Gx for Preventive Maintenance | XpertCheck Result: A1/210121669 - passed A2/210122064 - passed A3/210117513 - passed A4/210120091 - passed Cleaned cartridge bays, plunger rod Generate Data | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational | All 4-modules passed XpertCheck Calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458508706019220091/6458508706019220091_signature_28.png | Deosenia Tabo | https://www.jotform.com/uploads/macarewecare/222927978372067/6458508706019220091/6458508706019220091_signature_31.png
| 2026-02-02 | 2026-02-02 | Manuel J. Santos Hospital | 554 Montilla Blvd, Butuan City | 13 | Agusan del Norte | Molecular Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 32017 | DOH | biomed@mjsantos.com.ph | 23 |  |  | GX 4 6 Colors | 831292 | 2 modules found service required in XpertCheck Calibration, 2 modules was found an error in temperature. All modules are no longer working and in need of replacement. | None | Troubleshooting | All module are no longer functional. Other peripherals are working. | 4 modules ( 1,2,3 and 4) are in need of module replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458300391473442314/6458300391473442314_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6458300391473442314/6458300391473442314_signature_31.png
| 2026-02-02 | 2026-01-30 | Best Diagnostic Corp. | Masikap st,QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 32016 | COMMERCIAL | arvinrealr@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 11 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 2 |  | GX4 6C | 110008529 | As per end user they encounter workstation booting and due for maintenance | Reset workstation Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed troubleshooting and performed maintenance | Oprational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458230763613395890/6458230763613395890_signature_28.png | Best Diagnostic Corp | https://www.jotform.com/uploads/macarewecare/222927978372067/6458230763613395890/6458230763613395890_signature_31.png
| 2026-02-02 | 2026-01-30 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32015 | DOH | serly_reyes@yahoo.com | 28 |  |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458115652116027458/6458115652116027458_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6458115652116027458/6458115652116027458_signature_31.png
| 2026-02-02 | 2026-01-30 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32014 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXIV-4-D-6C | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458114532117329423/6458114532117329423_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6458114532117329423/6458114532117329423_signature_31.png
| 2026-02-02 | 2026-01-29 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32013 | DOH | arleneabad32@gmail.com | 28 |  |  | GXIV-4-D-6C | 838288 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458113712118101922/6458113712118101922_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/6458113712118101922/6458113712118101922_signature_31.png
| 2026-02-02 | 2026-01-29 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32012 | DOH | rhupozlab@gmail.com | 30 |  |  | GXIV-4-D-6C | 838467 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458112112114843801/6458112112114843801_signature_28.png | Azellamie V Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/6458112112114843801/6458112112114843801_signature_31.png
| 2026-02-02 | 2026-01-28 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32011 | DOH | kath101278@gmail.com | 26 |  |  | GXIV-4-D-6C | 838286 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458102002118980943/6458102002118980943_signature_28.png | Katherine Guarin | https://www.jotform.com/uploads/macarewecare/222927978372067/6458102002118980943/6458102002118980943_signature_31.png
| 2026-02-02 | 2026-01-27 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32010 | DOH | allan.vinluan2013@gmail.com | 30 |  |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458101012113944746/6458101012113944746_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6458101012113944746/6458101012113944746_signature_31.png
| 2026-02-02 | 2026-01-23 | Pangasinan Provincial Hospital | Bulingit, San Carlos City | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32009 | PBSP | judzfm@gmail.com | 28 |  |  | GXIV-4-D-10C | 120000484 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458098692115099257/6458098692115099257_signature_28.png | Judy Ann Mabunga | https://www.jotform.com/uploads/macarewecare/222927978372067/6458098692115099257/6458098692115099257_signature_31.png
| 2026-02-02 | 2026-01-22 | Region 1 Medical Center | Bonuan Binloc, Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32008 | PBSP | romarmtamayo@gmail.com | 23 |  |  | GX16 | 842230 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A2 and D3 Disabled | For Observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458095532117432125/6458095532117432125_signature_28.png | Mark Lawrence B. Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6458095532117432125/6458095532117432125_signature_31.png
| 2026-02-02 | 2026-01-22 | Region 1 Medical Center | Bonuan Binloc, Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 32007 | UNOPS | romarmtamayo@gmail.com | 24 |  | 2026-01-06 | GXIV-4-D-10C | 110023501 | For Preventive Maintenance and XpertCheck | #NAME? | Performed Preventive Maintenance and Xper | Unit is Fully Operational All Modules passed XpertCheck |  | XpertCheck Kits | DR 13646 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458094022113412548/6458094022113412548_signature_28.png | Mark lawrence B. Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6458094022113412548/6458094022113412548_signature_31.png
| 2026-02-02 | 2026-01-30 | Batis Health Center | F. Manalo St. San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32006 | DOH | almacenaio16@gmail.com | 24 | ASSAY: Others, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1767 ASSAY: Xpert HIV-1 Viral Load, COUNT: 210 |  | GX-IV-D-6C | 831598 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -filter Generated reports | Unit is operational. Module A3 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6458054073611689021/6458054073611689021_signature_28.png | Rosario T. Almacen | https://www.jotform.com/uploads/macarewecare/222927978372067/6458054073611689021/6458054073611689021_signature_31.png
| 2026-02-02 | 2026-01-30 | Banlat Health Center | Felicia st.,San Antonio homes | NCR | NCR 2nd - Quezon City | Lab | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 32005 | PBSP | ramosfely38@gmail.com | 28 |  |  | GX4 6C | 804688 | Workstation no power | Check workstation Replaced power cord and checked all connection Checked power supply. | Performed teoubleshooting | Unit unoperational Defective workstation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6457891934217595921/6457891934217595921_signature_28.png | Felicidad De Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/6457891934217595921/6457891934217595921_signature_31.png
| 2026-02-01 | 2026-01-30 | Philippine Red Cross - Cagayan de Oro Chapter | Capitol Grounds, Cagayan de Oro City | 11 | Misamis Oriental | Blood Bank | NO | REPAIR MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 32004 | COMMERCIAL | mrbc2@redcross.org.ph | 25 |  |  | MCS+ | 02H080 | Error at pump motors | Check and replace fhe roller pumps and assay Check the parameters of line sensor, SPM and DPM Values, weigher and cuff values Run cycle test | Troubleshooting the machine | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6457491258518638799/6457491258518638799_signature_28.png | Sheloah Emmanuelle Seriña | https://www.jotform.com/uploads/macarewecare/222927978372067/6457491258518638799/6457491258518638799_signature_31.png
| 2026-02-01 | 2026-01-30 | Philippine Red Cross - Cagayan de Oro Chapter | Capitol Grounds, Cagayan de Oro City | 11 | Misamis Oriental | Blood Bank | NO | REPAIR MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 32003 | COMMERCIAL | mrbc2@redcross.org.ph | 25 |  |  | MCS+ | 02H080 | Error at pump motors | Check and replace fhe roller pumps and assay Check the parameters of line sensor, SPM and DPM Values, weigher and cuff values Run cycle test | Troubleshooting the machine | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6457491228514316148/6457491228514316148_signature_28.png | Sheloah Emmanuelle Seriña | https://www.jotform.com/uploads/macarewecare/222927978372067/6457491228514316148/6457491228514316148_signature_31.png
| 2026-01-30 | 2026-01-29 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32002 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 594 ASSAY: Xpert MTB-XDR, COUNT: 31 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609907211682753/6455609907211682753_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609907211682753/6455609907211682753_signature_31.png
| 2026-01-30 | 2026-01-29 | Marikina City Health Office | Shoe Avenue, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32001 | PBSP | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 830 |  | GX-IV-D-10C | 110020361 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609907213800468/6455609907213800468_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609907213800468/6455609907213800468_signature_31.png
| 2026-01-30 | 2026-01-29 | Marikina City Health Office | Shoe Ave., Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31998 | PBSP | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1633 |  | GX-IV-D-6C | 805211 | Module A2 not detected | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module A2 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609897212682412/6455609897212682412_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609897212682412/6455609897212682412_signature_31.png
| 2026-01-30 | 2026-01-29 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31999 | DOH | estyaragon2019@gmail.com | 24 |  |  | GX-IV-D-6C | 831588 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609897219889245/6455609897219889245_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609897219889245/6455609897219889245_signature_31.png
| 2026-01-30 | 2026-01-29 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 32000 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 864 ASSAY: Xpert MTB-XDR, COUNT: 2 |  | GX-IV-D-10C | 110016107 | for maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609907213826738/6455609907213826738_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6455609907213826738/6455609907213826738_signature_31.png
| 2026-01-29 | 2026-01-16 | Tanza Wellness Center | Anterio Soriano Highway, Tanza, Cavite | 4A | Cavite | Social Hygiene Clinic | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31997 | MACARE-HBDC | tanzarhu.clinicallab@gmail.com | 27 |  |  | GXIV-D-10C | 120003582 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Replaced Filter. Performed Self-test. Installed LGU's Printer. Install HIV XC Assay. Generate Reports. | 4 Modules are Good. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6455009225152530274/6455009225152530274_signature_28.png | MERRIAM JOY G. SIBBALUCA | https://www.jotform.com/uploads/macarewecare/222927978372067/6455009225152530274/6455009225152530274_signature_31.png
| 2026-01-29 | 2026-01-28 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | jniebres@asianhospital.com |  | 31996 | COMMERCIAL | agalang@asianhospital.com | 25 |  |  | GXIV-D-10C | 110014996 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | B2 is Disabled for Module Replacement. B1, B3 and B4 are Good. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6455009225156450630/6455009225156450630_signature_28.png | Alyssa R Galang | https://www.jotform.com/uploads/macarewecare/222927978372067/6455009225156450630/6455009225156450630_signature_31.png
| 2026-01-29 | 2026-01-29 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Apheresis room | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31995 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 28 |  |  | MCS + | 15H030G | Noisy sound came frm transfer pump during procedure | Checked rollers Checked unit Checked program parameters | Performed troubleshooting | For replacement of urethane rollers |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6454787335477504029/6454787335477504029_signature_28.png | Xyrich Viem Lucas | https://www.jotform.com/uploads/macarewecare/222927978372067/6454787335477504029/6454787335477504029_signature_31.png
| 2026-01-29 | 2026-01-29 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Apheresis room | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31994 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 28 |  |  | MCS + | 15H029G | Noisy sound came from transfer pump during procedure | Checked rollers Checked program parameters Checked unit | Performed troubleshooting | For replacement of urethane rollers |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6454786335473234696/6454786335473234696_signature_28.png | Xyrich Viem Lucas | https://www.jotform.com/uploads/macarewecare/222927978372067/6454786335473234696/6454786335473234696_signature_31.png
| 2026-01-29 | 2026-01-28 | St.Luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 31993 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 24 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 468 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 ASSAY: Xpert HPV, COUNT: 380 ASSAY: Xpert MTB/RIF Ultra, COUNT: 481 ASSAY: Xpert HCV VL, COUNT: 37 ASSAY: Xpert HIV-1 Viral Load, COUNT: 334 |  | GX4 6C | 807155 | Due for maintenance C4 disabled due to error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | C1,C2,C3 operational C4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6454785435472513846/6454785435472513846_signature_28.png | Roiland Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6454785435472513846/6454785435472513846_signature_31.png
| 2026-01-29 | 2026-01-28 | St. Luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 31992 | COMMERCIAL | megapo@stlukes.com.ph | 22 | ASSAY: Xpert CT/NG, COUNT: 217 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2270 |  | GX4 6C | 839234 | Due for maintenance B3 and B4 reported already | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Restart system Gathered reports Test run unit | Performed maintenance | B1 and B2 in use B3 and B4 reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6454784995472048795/6454784995472048795_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6454784995472048795/6454784995472048795_signature_31.png
| 2026-01-29 | 2026-01-27 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | YES | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31991 | MACARE-HBDC | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 65 ASSAY: Xpert HIV-1 Viral Load, COUNT: 29 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 61 |  | GX4 6 color | 838837 | All module encounter errors | Gathered reports Cleaned filter | Performed troubleshooting | Unit unoperational B1 879215 failed xpertcheck last 8/15/24 B2 882381 temp error 4009,1002 B3 602135 temp error 1002,4009,2022 B4 780860 temp error 4009 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6454784645472244939/6454784645472244939_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6454784645472244939/6454784645472244939_signature_31.png
| 2026-01-29 | 2026-01-29 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | gehbedayos@gmail.com |  | 31990 | COMMERCIAL | mahmci_labcmt@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 16 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 183 ASSAY: Xpert HIV-1 Viral Load, COUNT: 15 |  | GX-IV-D-6C | 110005265 | Error 5007 on B1 and B3 | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module B1, B2 and B3 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6454751992715089818/6454751992715089818_signature_28.png | Fatima Busar | https://www.jotform.com/uploads/macarewecare/222927978372067/6454751992715089818/6454751992715089818_signature_31.png
| 2026-01-29 | 2026-01-29 | Philippine Red Cross - Cagayan de Oro | Cagayan de Oro City | 10 | Misamis Oriental | Blood Center | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31989 | COMMERCIAL | mrbc2@redcross.org.ph | 27 |  |  | MCS+ | 02H080 | UPS can't back up power during power interruption. | Replacement and Troubleshooting. | Replacing rollers for a new one. Replacing Battery Supply in UPS with new one. Replacing worn out cuff. | The unit is functional and working. | For replacement of Stand Roller. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6454644039676646245/6454644039676646245_signature_28.png | Sheloah Emmanuelle Seriña | https://www.jotform.com/uploads/macarewecare/222927978372067/6454644039676646245/6454644039676646245_signature_31.png
| 2026-01-29 | 2026-01-29 | Altavas Rural Health Unit | Altavas, Aklan | 6 | Aklan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 31988 | DOH | joricarmt29@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 676 | 2025-12-22 | GX4 | 838458 | Module B4/861443 for replacement | Preventive Maintenance | -Replaxed B4 with new module SN 673337 | 4 modules operational |  | Genexpert module | 13626 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6454559535205533879/6454559535205533879_signature_28.png | Jorica Lynn Duronio | https://www.jotform.com/uploads/macarewecare/222927978372067/6454559535205533879/6454559535205533879_signature_31.png
| 2026-01-28 | 2026-01-28 | Philippine Red Cross Davao City Chapter | Roxas Avenue, Davao City | 11 | Davao del Sur | Bloodbank | NO | REPAIR CALIBRATION | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 31987 | COMMERCIAL | davaocitybloodservices@gmail.com | 24 |  |  | MCS+ | 02J021 | Machine not working, safety board issue | Replaced the Defective Safety Board SN:MM050902198 with service safety board SN: SP081700879G Perform Preventive Maintenance and Calibration; check the sensors and recalibrate to operational values Perform cycle test to ensure the machine is functional and operational | Perform Safety Board Replacement Perform Calibration | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6453851676517375157/6453851676517375157_signature_28.png | Mars Panes | https://www.jotform.com/uploads/macarewecare/222927978372067/6453851676517375157/6453851676517375157_signature_31.png
| 2026-01-28 | 2026-01-28 | Saguiaran Rural Health Health | Saguiaran Lanao Del Sur | 10 | Lanao del Sur | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31986 | DOH | aiylacalandada01@gmail.com | 28 |  |  | Gx 4 | 838298 | Defective Module 2 Sn 808278 | Gather data IQR, System log and Assay Statistics Replace new Fan Filter Clean Fan | New module Installation Sn 731829 Gather data IQR, System log and Assay Statistics Replace new Fan Filter Clean Fan | Systems operational Module 3 temperature error related Under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6453776256275930385/6453776256275930385_signature_28.png | Asmalia A hadji Salic | https://www.jotform.com/uploads/macarewecare/222927978372067/6453776256275930385/6453776256275930385_signature_31.png
| 2026-01-28 | 2026-01-28 | Kidapawan City Health Office | Kidapawan City, Cotabato | 12 | Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31985 | DOH | kidcho@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2616 |  | GxIV-6C | 838545 | None | Perform Genexpert cleaning and maintenance Module Cleaning Plunger rod and optics cleaning Fan and Filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6453747008659958949/6453747008659958949_signature_28.png | Kidapawan City Health Office Primary Care Facility | https://www.jotform.com/uploads/macarewecare/222927978372067/6453747008659958949/6453747008659958949_signature_31.png
| 2026-01-28 | 2026-01-27 | San Joaquin Super Health Center | Elizco Rd, San Joaquin, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31984 | MACARE-HBDC | peachmango318@gmail.com | 26 |  |  | GX-IV-D-10C | 110026463 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6453684763613874822/6453684763613874822_signature_28.png | Arlene Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6453684763613874822/6453684763613874822_signature_31.png
| 2026-01-28 | 2026-01-27 | San Joaquin Super Health Center | Elizco Rd, San Joaquin, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31983 | MACARE-HBDC | peachmango318@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2067 |  | GX-IV-D-10C | 110024245 | For relocation | Installed the GeneXpert System | Installed the GeneXpert System: -GeneXpert unit -workstation -UPS -Printer -Barcode Scanner with stand Performed maintenance | Unit is operational. Unit from Pinagbuhatan TB Microscopy Laboratory |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6453684763615673533/6453684763615673533_signature_28.png | Arlene Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6453684763615673533/6453684763615673533_signature_31.png
| 2026-01-27 | 2026-01-27 | Cotabato Provincial Hospital | Brgy. Amas, Kidapawan City, North Cotabato | 12 | Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31982 | MACARE-HBDC | januarybing1206@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 87 ASSAY: Xpert MTB-XDR, COUNT: 9 ASSAY: Xpert HIV-1 Viral Load, COUNT: 206 |  | GxlV- 10C | 110026514 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and Filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6453130508653839261/6453130508653839261_signature_28.png | Hermosilla Enero | https://www.jotform.com/uploads/macarewecare/222927978372067/6453130508653839261/6453130508653839261_signature_31.png
| 2026-01-27 | 2026-01-27 | Cotabato Provincial Hospital | Brgy. Amas, Kidapawan City, North Cotabato | 12 | Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31981 | MACARE-HBDC | januarybing1206@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 87 ASSAY: Xpert MTB-XDR, COUNT: 9 ASSAY: Xpert HIV-1 Viral Load, COUNT: 206 |  | GxlV- 10C | 110026514 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and Filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6453130408659905489/6453130408659905489_signature_28.png | Hermosilla Enero | https://www.jotform.com/uploads/macarewecare/222927978372067/6453130408659905489/6453130408659905489_signature_31.png
| 2026-01-27 | 2026-01-27 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31980 | COMMERCIAL | cduhbloodbank2@gmail.com | 19 |  |  | MCS+ | 14F014 | Centrifuge chuck problem due to leak of blood during the procedure. | Clean up the blood residue in the centrifuge chuck. | Clean up the blood residue in the centrifuge chuck. | Unit is operational. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6453109580122102003/6453109580122102003_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6453109580122102003/6453109580122102003_signature_31.png
| 2026-01-27 | 2026-01-27 | Malilipot Rural Health Unit | Brgy. 4, Poblacion, Malilipot | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31979 | PBSP | faith.buban2124@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3046 |  | GX4-6C | 808280 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, fan filter Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6453065911716504427/6453065911716504427_signature_28.png | Fatima ana cabiles | https://www.jotform.com/uploads/macarewecare/222927978372067/6453065911716504427/6453065911716504427_signature_31.png
| 2026-01-27 | 2026-01-26 | Macare Medicals Inc | 1005 Aurora Grand, Aurora blvd | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | OTHERS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com | Randox | 31978 | COMMERCIAL | florescanoel77@gmail.com | 26 |  | 2026-01-26 | Evidence Multistat | 052-23-0712 | Unable to initialize | Replaced and installed thermal system PCB s/n 23013212 Checked all connection and connectors Performed initialization | Performed repair | Unit operational | Returned def parts s/n 23028209 | Thermal system PCB s/n 23013212 | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6452984393618008718/6452984393618008718_signature_28.png | Macare Medicals Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/6452984393618008718/6452984393618008718_signature_31.png
| 2026-01-27 | 2026-01-27 | Don Jose S. Monfort Medical Center | Barotac Nuevo, Iloilo | 6 | Iloilo | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31975 | DOH | judyboglosa35@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2042 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GX4 6C | 838501 | Module B4 HW Fail, Temperature-related errors | PM, cleaned filter, module bays, optics, pkunger rod | Installed repaired UPS | -UPS operational; - 3 modules operational- B1, B2, B3 - module B4 non-operational, for report to Cepheid |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6452922128902097813/6452922128902097813_signature_28.png | judy boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6452922128902097813/6452922128902097813_signature_31.png
| 2026-01-27 | 2026-01-23 | Nagpayong Super Health Center | Nagpayong, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31973 | DOH | mnoreend@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1417 |  | GX-IV-D-6C | 831905 | For maintenance Installation of UPS | Performed maintenance | Installed UPS: Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6452919021211763069/6452919021211763069_signature_28.png | Ma Noreen De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/6452919021211763069/6452919021211763069_signature_31.png
| 2026-01-27 | 2026-01-23 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | Packing | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31974 | MACARE-HBDC | peachmango318@gmail.com | 24 |  |  | GX-IV-D-10C | 110024245 | For pull out | Packed the GeneXpert system | Packed the GeneXpert System: -GeneXpert unit -UPS -Workstation -Printer -Keyboard and mouse -Barcode Scanner with stand -Monitor | Unit ready for pull out. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6452919021217957599/6452919021217957599_signature_28.png | Arlen Delgado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6452919021217957599/6452919021217957599_signature_31.png
| 2026-01-26 | 2026-01-21 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31972 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GX-II | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6452284422117750459/6452284422117750459_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6452284422117750459/6452284422117750459_signature_31.png
| 2026-01-23 | 2026-01-23 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | INSTALLATION REPAIR CALIBRATION | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31971 | COMMERCIAL | cduhbloodbank2@gmail.com | 19 |  |  | MCS+ | 16L508SPG | - SN: 10M104 Unit pull-out - Installation of demo unit machine SN: 16L508SPG | - SN: 10M104 Unit pull-out - Installation of demo unit machine SN: 16L508SPG (New Unit of Pneumatic Assembly, Line Sensor, CPU Board of 10M104) - Calibrate and Test Run the demo unit | - SN: 10M104 Unit pull-out - Installation of demo unit machine SN: 16L508SPG (New Unit of Pneumatic Assembly, Line Sensor, CPU Board of 10M104) - Calibrate and Test Run the demo unit | Unit is operational. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449713248426912348/6449713248426912348_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6449713248426912348/6449713248426912348_signature_31.png
| 2026-01-23 | 2026-01-23 | Perpetual Help Medical Center | Las Pinas | NCR | NCR 4th - Las Piñas | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31970 | COMMERCIAL | phmclp.mdl@gmail.com | 23 |  |  | GXIV | 831316 | Module A4(784666) for replacement | Replaced module | - checked and cleaned cartridge bay, plunger rod, filter - replaced module by 220049150 - generated reports - performed self tezt | Unit is operational Installed new module on A4 Modules A1, A2, A3 are recommended for replacement (Existing Cepheid case numbers) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449573124738053225/6449573124738053225_signature_28.png | Donn Dave Clerigo | https://www.jotform.com/uploads/macarewecare/222927978372067/6449573124738053225/6449573124738053225_signature_31.png
| 2026-01-23 | 2026-01-23 | Iligan Society of Internists | Iligan City | 10 | Lanao del Norte | TB Dots | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31969 | PBSP | dayolaermajessica@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2949 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 |  | 6C GXIV-4D | 805090 | Frequent H/W failures | #NAME? | #NAME? | Machine is functional. | For reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449496894253155279/6449496894253155279_signature_28.png | Erma Jessica Dayola | https://www.jotform.com/uploads/macarewecare/222927978372067/6449496894253155279/6449496894253155279_signature_31.png
| 2026-01-23 | 2026-01-23 | Puerto Princesa CHO | Puerto Princesa, Palawa | 4B | Palawan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 31968 | UNOPS | jocal0324@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 289 |  | GXIV | 110023510 | For Xpertcheck caliration and maintenance. | Peroformed xpertcheck calibration. Performed preventive maintenance.. | Performed xpertcheck and preventive maintenance. | All 4 modules PASSED the xpertcheck calibration. All 4 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449394558518968529/6449394558518968529_signature_28.png | Jocelyn L. Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/6449394558518968529/6449394558518968529_signature_31.png
| 2026-01-23 | 2026-01-22 | Sumilang Super Health Center | Dr. García St. Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31967 | PBSP | maricordevilla@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1160 ASSAY: Xpert MTB-XDR, COUNT: 163 |  | GX-IV-D-10C | 1200000495 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module B4 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393600917115091/6449393600917115091_signature_28.png | Maria Corazon Obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393600917115091/6449393600917115091_signature_31.png
| 2026-01-23 | 2026-01-21 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 31962 | COMMERCIAL | jovincelim5@gmail.com | 24 |  |  | MCS+ | 08M461 | For maintenance DPM occasionally have unstable pressure Red valve error from time to time | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Check valves Functionality test | Unit is operational. | Recommend to calibrate the unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540912064153/6449393540912064153_signature_28.png | Josh Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540912064153/6449393540912064153_signature_31.png
| 2026-01-23 | 2026-01-21 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31963 | COMMERCIAL | renalyn.cabaltea@medilinx.com.ph | 22 | ASSAY: Xpert HBV VL, COUNT: 16 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1188 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540912149355/6449393540912149355_signature_28.png | Renalyn Cabaltea | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540912149355/6449393540912149355_signature_31.png
| 2026-01-23 | 2026-01-22 | Pasig Treatment Hub - Sumilang | Dr. Garcia St. Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31964 | MACARE-HBDC | pathmedtech@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 886 |  | MCS+ | 110025128 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540914880842/6449393540914880842_signature_28.png | Jhozza Rei Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540914880842/6449393540914880842_signature_31.png
| 2026-01-23 | 2026-01-22 | Pasig Treatment Hub - CHAMP | CHAMP bldg., Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | dansumulong.pasigcho@gmail.com |  | 31965 | MACARE-HBDC | pathmedtech@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 794 ASSAY: Others, COUNT: 19 |  | GX-IV-D-10C | 110025127 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540919408892/6449393540919408892_signature_28.png | Dan Sumulong | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540919408892/6449393540919408892_signature_31.png
| 2026-01-23 | 2026-01-22 | Sumilang Super Health Center | Dr. Garcia St. Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31966 | PBSP | maricordevilla@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1185 |  | GX-IV-D-6C | 814163 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module B4 not operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540919858264/6449393540919858264_signature_28.png | Maria Corazon Obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/6449393540919858264/6449393540919858264_signature_31.png
| 2026-01-22 | 2026-01-22 | Tamparan District Hospital | Tamparan, Lanao del Sur | BARMM | Lanao del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31961 | IOM | pmamowalas@yahoo.com | 27 |  | 2026-01-07 | 10C GXIV-4D | 110023525 | #NAME? | #NAME? | #NAME? | Machine is functional |  |  | DR no. 5471 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6448853235257134779/6448853235257134779_signature_28.png | Pangilamun Mamowalas Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6448853235257134779/6448853235257134779_signature_31.png
| 2026-01-22 | 2026-01-22 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31960 | COMMERCIAL | cduhbloodbank2@gmail.com | 19 |  |  | MCS+ | 10M104 | SPM TEST FAILED. | Calibrate SPM to standard pressure from offset to Fullscale. Troubleshoot SPM switch. Need to pull out from the end-user. | Calibrate SPM to standard pressure from offset to Fullscale. Troubleshoot SPM switch. | Unit is not operational due to SPM is defective. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6448757298428828722/6448757298428828722_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6448757298428828722/6448757298428828722_signature_31.png
| 2026-01-22 | 2026-01-22 | Palawan PHO-Molecular Diagnostics Laboratory | Compound South National HW Irawan, Puerto Princesa City | 4B | Palawan | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 31959 | UNOPS | phomblpalawan@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 224 | 2026-01-07 | GXIV | 110023509 | For Xpertcheck. For maintenance. | Performed PM and calibration. | Performed unit maintenance. Run xpertcheck. | All 4 modules Passed. All4 modules operational. | NA | Xpertcheck kit | DR Num: 13650 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6448686044022988566/6448686044022988566_signature_28.png | Cyrus Cecilio A. Caabay RMT MPHA | https://www.jotform.com/uploads/macarewecare/222927978372067/6448686044022988566/6448686044022988566_signature_31.png
| 2026-01-22 | 2026-01-22 | National Center For Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31958 | PBSP | iancristrono@gnail.com | 27 |  |  | GXIV | 808625 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6448685738687050600/6448685738687050600_signature_28.png | Ian Cris Trono | https://www.jotform.com/uploads/macarewecare/222927978372067/6448685738687050600/6448685738687050600_signature_31.png
| 2026-01-22 | 2026-01-22 | National Center For Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31957 | UNOPS | iancristrono@gmail.com | 24 |  |  | GXIV-10C | 120003480 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6448685358689141306/6448685358689141306_signature_28.png | Ian Cris Trono | https://www.jotform.com/uploads/macarewecare/222927978372067/6448685358689141306/6448685358689141306_signature_31.png
| 2026-01-21 | 2026-01-21 | Cebu Doctors' University Hospital | Capitol Site | 7 | Cebu | PCR | NO | INSTALLATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 31956 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 22 |  |  | GX4 10C | 110026843 | . Installation of GeneXpert Machine System. . Installation of CPU . Installation of UPS . Installation LCD | Save Installation Qualifications Report Save System Log Report | Installation of GeneXpert Machine System. . Installation of CPU . Installation of UPS . Installation LCD | All are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447871187998225447/6447871187998225447_signature_28.png | Colleen Allison Libi | https://www.jotform.com/uploads/macarewecare/222927978372067/6447871187998225447/6447871187998225447_signature_31.png
| 2026-01-21 | 2026-01-21 | Philippine Red Cross Tower | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31955 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 01J038 | Unit for maintenance | Adjusted line sensor offset value | - checked rotor roller pumps - checked calibration parameters - adjusted line sensor offset value 2800 to 3000 - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447862299918231671/6447862299918231671_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6447862299918231671/6447862299918231671_signature_31.png
| 2026-01-21 | 2026-01-21 | Ciudad medical zamboanga | Mayor, vitaliano agan avenue, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31954 | COMMERCIAL | thairaphoebegrey13@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Xpert HBV VL, COUNT: 63 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 513 ASSAY: Xpert HCV VL, COUNT: 12 | 2026-01-21 | GX 2 | 834128 | None | Preventive maintenance, checked module status, checked system logs for recent errors, checked module and room temperatures. | Preventive maintenance, checked module status, checked system logs for recent errors, checked module and room temperatures. | All modules are available and ready to use | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447814291112535661/6447814291112535661_signature_28.png | Shereemer A. Cheong | https://www.jotform.com/uploads/macarewecare/222927978372067/6447814291112535661/6447814291112535661_signature_31.png
| 2026-01-21 | 2026-01-20 | Pasig Treatment Hub - Manggahan | Manggahan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31953 | MACARE-HBDC | pasigtreatmentmanggahan@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 27 |  | GX-IV-D-6C | 821648 | Can't import ADF | Import ADF | Import ADF for HIV Viral Load XC Scan cartridge Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Modules A2 and A4 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447708553511436484/6447708553511436484_signature_28.png | Katherine Godoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6447708553511436484/6447708553511436484_signature_31.png
| 2026-01-21 | 2026-01-20 | Bagong Ilog Health Center I | Sgt. L Pascua, Bagong Ilog, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31952 | MACARE-HBDC | agasangevelyn7@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 295 |  | GX-IV-D-10C | 110026461 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit os operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447708553515938846/6447708553515938846_signature_28.png | Evelyn Agasang | https://www.jotform.com/uploads/macarewecare/222927978372067/6447708553515938846/6447708553515938846_signature_31.png
| 2026-01-20 | 2026-01-16 | Sta. Maria Rural Health Unit | Sta. Maria, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31951 | UNOPS | jomarcaberto24@gmail.com | 27 |  | 2026-01-16 | GXIV-4-D-10C | 110023500 | For Preventive Maintenance and XpertCheck | #NAME? | Performed Preventive Maintenance and XpertCheck | Unit is Fully Operational All modules passed XpertCheck |  | XpertCheck Kits | DR 13654 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447135042112829720/6447135042112829720_signature_28.png | Jomar Caberto | https://www.jotform.com/uploads/macarewecare/222927978372067/6447135042112829720/6447135042112829720_signature_31.png
| 2026-01-20 | 2026-01-16 | Abra Provincial Hospital | Bangued, Abra | CAR | Abra | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31950 | PBSP | aphpathologydepartment@gmail.com | 26 |  |  | GXIV-4-D-10C | 110019619 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447133712117233236/6447133712117233236_signature_28.png | Riverlie Valera | https://www.jotform.com/uploads/macarewecare/222927978372067/6447133712117233236/6447133712117233236_signature_31.png
| 2026-01-20 | 2026-01-16 | Abra Provincial Hospital | Bangued, Abra | CAR | Abra | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31949 | UNOPS | aphpathologydepartment@gmail.com | 28 |  | 2026-01-06 | GXIV-4-D-10C | 110023499 | For Preventive Maintenance and XpertCheck | #NAME? | Performed Preventive Maintenance and XpertCheck | Unit is Fully Operational All modules passed XpertCheck |  | XpertCheck Kits | DR 13640 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447131342119420314/6447131342119420314_signature_28.png | Riverlie Valera | https://www.jotform.com/uploads/macarewecare/222927978372067/6447131342119420314/6447131342119420314_signature_31.png
| 2026-01-20 | 2026-01-15 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31948 | UNOPS | benguetphortdl@gmail.com | 28 |  | 2026-01-06 | GXIV-4-D-10C | 110023498 | For Preventive Maintenance and XpertCheck | #NAME? | Performed Preventive Maintenance and XpertCheck | #NAME? |  | XpertCheck Kits | DR 13642 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447129872114061833/6447129872114061833_signature_28.png | Tillber Baniaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6447129872114061833/6447129872114061833_signature_31.png
| 2026-01-20 | 2026-01-15 | Benguet Provincial Health Office | La Trinidad, Benguet | 1 | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31947 | PBSP | benguetphortdl@gmail.com | 29 |  |  | GXIV-4-D-6C | 807621 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance and | Module 1 failed XpertCheck on June 4, 2025 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447127822111960681/6447127822111960681_signature_28.png | Tillber Baniaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6447127822111960681/6447127822111960681_signature_31.png
| 2026-01-20 | 2026-01-13 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31946 | DOH | gbmmbg@yahoo.com | 29 |  |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational Module B1 for Observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447126472112269256/6447126472112269256_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/6447126472112269256/6447126472112269256_signature_31.png
| 2026-01-20 | 2026-01-08 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31945 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXIV-4-D-6C | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447125512114774689/6447125512114774689_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6447125512114774689/6447125512114774689_signature_31.png
| 2026-01-20 | 2026-01-20 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31944 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  |  | GXIV | 11008530 | Modules B1(220000626), B2(220000643), B3(210036044) for replacement | Upgraded GxDx software Replaced modules | - checked and cleaned catridge bay, plunger rod, filter - database back up and tests archiving - upgraded GxDx version 6.4 to 6.5 - replaced modules New modules SN: B1: 210088521 (10C) B2: 220049203 B3: 220049128 - performed self test | Unit is operational All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6447062002222602640/6447062002222602640_signature_28.png | Maria Elizabeth S. de Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/6447062002222602640/6447062002222602640_signature_31.png
| 2026-01-20 | 2026-01-20 | Sarangani Islands Rural Health Unit | Mabila, Sarangani Islands Davao Occidental | 11 | Davao Occidental | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | lovitosquennie26@gmail.com |  | 31943 | DOH | lovitosquennie26@gmail.com | 29.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 320 | 2025-11-14 | GX-IV 6C 4D | 837614 | System Unit is not working | Replaced the Defective System Unit SN: CPH0520F08269 With new System Unit SN: PC1135EB Install GX Software GX 6.4 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform System Unit replacement Perform Preventive Maintenance | Successfully install the new system unit and All Modules are operational | No ACU Available | fc-it-desktop | DR 13497 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6446891796885571626/6446891796885571626_signature_28.png | Quennie Lovitos,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6446891796885571626/6446891796885571626_signature_31.png
| 2026-01-19 | 2026-01-13 | Bicol South Luzon Sub National Reference Laboratory | Sagpon, Daraga | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31942 | PBSP | rbalzaga.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 954 ASSAY: Xpert MTB-XDR, COUNT: 141 |  | GX4-10C | 110016954 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6446429882216413430/6446429882216413430_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6446429882216413430/6446429882216413430_signature_31.png
| 2026-01-19 | 2026-01-13 | Bicol South Luzon Sub National Reference Laboratory | Sagpon, Daraga | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31941 | PBSP | rbalzaga.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2595 ASSAY: Xpert HIV-1 Viral Load, COUNT: 871 |  | GX4-6C | 808636 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6446428302218317864/6446428302218317864_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6446428302218317864/6446428302218317864_signature_31.png
| 2026-01-19 | 2026-01-13 | Bicol South Luzon Sub National Reference Laboratory | Sagpon, Daraga | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31940 | PBSP | rbalzaga.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1669 ASSAY: Xpert HIV-1 Viral Load, COUNT: 569 |  | GX4-6C | 807579 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6446426462219611299/6446426462219611299_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6446426462219611299/6446426462219611299_signature_31.png
| 2026-01-19 | 2026-01-14 | Legazpi City Health Office | Capitol, Legazpi | 5 | Albay | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31939 | WHO | chanjanzhel@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1746 |  | GX4-6C | 846354 | Gx for Preventive Maintenance Date and Time resets | Replace CPU Bios Battery for date and time(resolved) Cleaned cartridge bays, plunger rod, fan filter Generate Data | Performed Preventive Maintenance Performed Troubleshooting | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6446424532214624653/6446424532214624653_signature_28.png | Janzhel Marie Chan | https://www.jotform.com/uploads/macarewecare/222927978372067/6446424532214624653/6446424532214624653_signature_31.png
| 2026-01-19 | 2026-01-16 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31938 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 08H188 | Due for maintenance Urethane rollers yellowish color | Cleaned unit Checked rollers Checked and test tube sensor Checked and test keypad and lcd display. Checked program parameters Performed functionality test Performed simulation running | Performed maintenance | Unit still can use but need to replace urethane rollers 6pcs |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6446097833614737065/6446097833614737065_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/6446097833614737065/6446097833614737065_signature_31.png
| 2026-01-19 | 2026-01-15 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Apheresis donor room | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31937 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 26 |  |  | MCS+ | 13C269 | As per enduser they encounter unusual sound came frm transfer pump | Checked unit Checked if worn out rollers Checked program parameters Performed functionality test Performed simulation running | Performed troubleshooting | Unit operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6446097333613684421/6446097333613684421_signature_28.png | PBC | https://www.jotform.com/uploads/macarewecare/222927978372067/6446097333613684421/6446097333613684421_signature_31.png
| 2026-01-18 | 2026-01-14 | Lake Sebu Rural Health Unit | Lake Sebu, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31936 | DOH | jennypingkian23@yahoo.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 539 | 2026-01-14 | GxIV-6C | 838569 | None | Replace defective module B1 SN:705884 with new Module SN: 716372 Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine module repair Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert 6-color module SN:716372 | DR#: 13627 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6445506187324926267/6445506187324926267_signature_28.png | Jenny A Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/6445506187324926267/6445506187324926267_signature_31.png
| 2026-01-18 | 2026-01-13 | South Cotabato Provincial Hospital | Koronadal City, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31935 | MACARE-HBDC | anamariermt22@gmail.com | 29.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3000 ASSAY: Xpert MTB-XDR, COUNT: 48 ASSAY: Xpert HIV-1 Viral Load, COUNT: 345 |  | GX- XVl | 839314 | Instrument letter interchange | Fixed module instrument letters back to position Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and problem Data collection | Genexpert module troubleshooting Genexpert machine Preventive maintenance | 11 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6445506187328161513/6445506187328161513_signature_28.png | Ana Marie Grace Egasan | https://www.jotform.com/uploads/macarewecare/222927978372067/6445506187328161513/6445506187328161513_signature_31.png
| 2026-01-18 | 2026-01-15 | MEDI LINX LABORATORY INC. / St.Elizabeth Hospital Inc | General Santos City, South Cotabato | 12 | South Cotabato | Bacteriology | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31934 | COMMERCIAL | sofia.maclang@medilinx.com.ph | 23.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 93 |  | GxIV-6C | 831313 | Module 2 and 4 for module replacement | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 2 operational modules (module 1 and 3) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6445506187329473369/6445506187329473369_signature_28.png | Richei Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/6445506187329473369/6445506187329473369_signature_31.png
| 2026-01-17 | 2026-01-16 | Tayasan Rural Health Unit | Tayasan, Negros Oriental | 7 | Negros Oriental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com | gfmselga@gmail.com |  | 31933 | UNOPS | tayasanrhu2023@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 418 | 2026-01-07 | GX4 10C | 110023502 | Due for Calibration | Clean optics, filter, plunger rods, module bays; archive tests, backup | .Perform Xpertcheck | - B2, B3 Passed Xpertcheck; - B1, B4 Requiring Service (modules will be replaced by Cepheid) |  | Xpertcheck Kit | 5470 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6444618972238510702/6444618972238510702_signature_28.png | Welyn C. Mahinay | https://www.jotform.com/uploads/macarewecare/222927978372067/6444618972238510702/6444618972238510702_signature_31.png
| 2026-01-16 | 2026-01-16 | Bataan General Hospital and Medical Center | Balanga,Bataan | 3 | Bataan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31932 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 25 |  |  | GXIV-D 10C | 120003587 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6443623804398827566/6443623804398827566_signature_28.png | Rosemin Taguinod | https://www.jotform.com/uploads/macarewecare/222927978372067/6443623804398827566/6443623804398827566_signature_31.png
| 2026-01-16 | 2026-01-16 | St Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31931 | MACARE-HBDC | laboratory@slec.ph | 22 |  | 2026-01-15 | Infinity 48s | 809643 | Modules E4, F2, G2, G3 for replacement after XpertCheck | Replaced defective modules (695334 > 642045, 692341 > 608480, 695066 > 800987, 694263 > 210059833) | Replaced defective modules Performed functionality test Performed maintenance cleaning Generated reports | Unit is operational |  | Certified pre-owned GeneXpert module (x4) Modified ejector (x4) Assy Door GeneXpert module (x4) | DR 13681 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6443596922319836015/6443596922319836015_signature_28.png | Ma. Sarah Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6443596922319836015/6443596922319836015_signature_31.png
| 2026-01-16 | 2026-01-16 | Philippine Red Cross - Davao City Chapter | Roxas Avenue, Davao City | 11 | Davao del Sur | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 31930 | COMMERCIAL | davaocitybloodservices@gmail.com | 24.5 |  |  | MCS+ | 02J021 | Error during self test | Check line voltages and power Reset the connection, Safety Board, Valve Board, Cpu Board and Motor Board | Perform error troubleshooting | Found a burned part at safety board Machine not operational | Documented the faulty board and reported to higher ups for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6443461977266364065/6443461977266364065_signature_28.png | Joy Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/6443461977266364065/6443461977266364065_signature_31.png
| 2026-01-16 | 2026-01-16 | Macare Medicals Inc | 1005 Aurora Grand,Aurora QC | NCR | NCR 2nd - Quezon City | Macare in house | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 31929 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | MCS + | 02H078 | For return and installation of display cable, display cable stand by parts from NCH. Display cable originally came from 02H078 | Returned and installed display cable | Performed repair | Unit unoperational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6443434643615480561/6443434643615480561_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6443434643615480561/6443434643615480561_signature_31.png
| 2026-01-16 | 2026-01-16 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Mabini, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31928 | UNOPS | molecularlab@drpjgmrmc.doh.gov.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2854 ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert MTB-XDR, COUNT: 84 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 446 | 2026-01-06 | GXIV-10C | 110023508 | Due for Calibration | Performed Xpertcheck (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Xpertcheck Kit | 13645 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6443390640618084042/6443390640618084042_signature_28.png | Kenneth Delos Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6443390640618084042/6443390640618084042_signature_31.png
| 2026-01-15 | 2026-01-15 | Silang Rural Health Unit | New Municipal Compound, Silang, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31927 | DOH | silangrhu@gmail.com | 23 |  |  | GXIV-D-6C | 820708 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | B1 for Observation after Icore Cleaning. B2-B4 in Good Condition. | Last XpertCheck - 11/11/25 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442741538317164901/6442741538317164901_signature_28.png | Evalyn Martin | https://www.jotform.com/uploads/macarewecare/222927978372067/6442741538317164901/6442741538317164901_signature_31.png
| 2026-01-15 | 2026-01-15 | Santa Rosa City Health Office 1 | Santa Rosa, Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31926 | UNOPS | santarosacho1clinlab@gmail.com | 25 |  | 2026-01-07 | GXIV-10C | 110023497 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  | Xpertcheck Kit | DR#13655 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442586622914999999/6442586622914999999_signature_28.png | Alona Alvaro | https://www.jotform.com/uploads/macarewecare/222927978372067/6442586622914999999/6442586622914999999_signature_31.png
| 2026-01-15 | 2026-01-15 | Santa Rosa City Health Office 1 | Santa Rosa, Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31925 | PBSP | santarosacho1clinlab@gmail.com | 25 |  |  | GXIV | 804839 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442585822915277274/6442585822915277274_signature_28.png | Alona Alvaro | https://www.jotform.com/uploads/macarewecare/222927978372067/6442585822915277274/6442585822915277274_signature_31.png
| 2026-01-15 | 2026-01-15 | Santa Rosa City Health Office 1 | Santa Rosa, Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31924 | UNOPS | santarosacho1clinlab@gmail.com | 25 |  |  | GXIV | 838519 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442585152916035916/6442585152916035916_signature_28.png | Alona Alvaro | https://www.jotform.com/uploads/macarewecare/222927978372067/6442585152916035916/6442585152916035916_signature_31.png
| 2026-01-15 | 2026-01-15 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31923 | UNOPS | acho.laboratory@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 838 ASSAY: Xpert MTB/RIF Ultra, COUNT: 63 ASSAY: Xpert MTB-XDR, COUNT: 174 | 2026-01-06 | GX-IV-D-10C | 110023496 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Run XpertCheck Generated reports | Unit is operational. All modules passed. All modules operational. |  | Lot no.: 1001490130 | DR# 13641 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442572703396172904/6442572703396172904_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6442572703396172904/6442572703396172904_signature_31.png
| 2026-01-15 | 2026-01-13 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco Ave, Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31920 | MACARE-HBDC | peachmango318@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2000 |  | GX-IV-D-10C | 110024245 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442562343392768773/6442562343392768773_signature_28.png | Arlen Delgado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6442562343392768773/6442562343392768773_signature_31.png
| 2026-01-15 | 2026-01-13 | Klinika Eastwood | CityWalk 2, Eastwood, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31919 | MACARE-HBDC | klinikaeastwood@gmail.com | 20 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 158 |  | GX-IV-D-10C | 110020698 | Database problem when opening GeneXpert software | Restart SQL configuration | Restart SQL Database Configuration Performed maintenance Cleaned: -Cartridge bay -plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442562343393353479/6442562343393353479_signature_28.png | Robelyn B. Dacanay | https://www.jotform.com/uploads/macarewecare/222927978372067/6442562343393353479/6442562343393353479_signature_31.png
| 2026-01-15 | 2026-01-14 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31921 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 57 ASSAY: Xpert CT/NG, COUNT: 74 ASSAY: Xpert HBV VL, COUNT: 617 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1147 ASSAY: Xpert HPV, COUNT: 415 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6753 ASSAY: Xpert MTB-XDR, COUNT: 6 ASSAY: Others, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 202 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1133 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -plunger rod -Icore -Filter Generated reports Checked setting for MTB XDR upload to LIS | Unit is operational. Module D1 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442562343393778139/6442562343393778139_signature_28.png | Xchosa Kay C. Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/6442562343393778139/6442562343393778139_signature_31.png
| 2026-01-15 | 2026-01-14 | The Medical City | Ortigas ave, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31922 | COMMERCIAL | bloodbank@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. | Recommend to calibrate the unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442562343396800977/6442562343396800977_signature_28.png | Roxanne Paulina Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/6442562343396800977/6442562343396800977_signature_31.png
| 2026-01-15 | 2026-01-14 | Ospital Ng Muntinlupa | Civic Dr, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | CALIBRATION MAINTENANCE XpertCheck | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | osmunclinlab@gmail.com |  | 31918 | UNOPS | osmunmicro@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 728 | 2026-01-07 | GXIV-D-10C | 110023494 | Unit is due for XpertCheck. | XpertCheck. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter If Clogged. Archived Test Results. Performed XpertCheck. Performed Self-Test. Generate Reports. | B1 - Pass B2 - Pass B3 - Pass B4 - Pass |  | XpertCheck Kit. | DR 13649 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449367421425038/6442449367421425038_signature_28.png | Jesse Jasper Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449367421425038/6442449367421425038_signature_31.png
| 2026-01-15 | 2026-01-13 | Jose Reyes Memorial Medical Center | Rizal Ave, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31917 | MACARE-HBDC | nelsonmarksalvador17@yahoo.com | 21 | ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 641 |  | GXIV-D-10C | 110014099 | Unit is due for Maintenance | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Changed Filter. Performed Self-test. Generate Reports. | B3 with Temperature Errors. B1 Disabled for Errors. B2 and B4 are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449357421770502/6442449357421770502_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449357421770502/6442449357421770502_signature_31.png
| 2026-01-15 | 2026-01-13 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31916 | MACARE-HBDC | nelsonmarksalvador17@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 876 ASSAY: Xpert MTB/RIF Ultra, COUNT: 472 ASSAY: Xpert SARS-CoV-2, COUNT: 82 |  | GX16-D-10C | 110013856 | Unit is due for Maintenance. | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Replaced Filter. Performed Self-test. Generate Reports | Modules A2, B2, B3, B4 were disabled with temperature Errors. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449297422259745/6442449297422259745_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449297422259745/6442449297422259745_signature_31.png
| 2026-01-15 | 2025-12-23 | Dasmariñas City Health Office 2 | DBB-E, Sta.Cruz, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31913 | MACARE-HBDC | hazelmilan0777@gmail.com | 25 |  | 2025-12-19 | GXIV-D-10C | 110027240 | Unit is due for Installation. | Installation. | Install GeneXpert System. Generate Reports. | Unit is Operational. UPS and Printer is Good. |  | GeneXpert System | DR 13611 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449297427838508/6442449297427838508_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449297427838508/6442449297427838508_signature_31.png
| 2026-01-15 | 2026-01-06 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | bloodbank@asianhospital.com |  | 31914 | COMMERCIAL | kcaangay@asianhospital.com | 25 |  |  | MCS+ | 10B129 | Reported Valve Problem | Troubleshoot. | Checked Valve Pressure. Checked Diagnostic Parameters. Performed Simulation Test. | Unit is Operational. Valves are Okay. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449297429626992/6442449297429626992_signature_28.png | Keano Riggs Caañgay | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449297429626992/6442449297429626992_signature_31.png
| 2026-01-15 | 2026-01-06 | Asian Hospital and Medical Center | Civic Dr, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | bloodbank@asianhospital.com |  | 31915 | COMMERCIAL | kcaangay@asianhospital.com | 25 |  | 2026-01-05 | MCS+ | 01F116 | Unit has Defective Compressor Assay. | Part Replacement. | Replaced Compressor Assay. Checked Diagnostic Parameters. Performed Simulation Run. | Unit is Operational. |  | Compressor Assay. | DR 13637 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449297429818486/6442449297429818486_signature_28.png | Keano Riggs Caañgay | https://www.jotform.com/uploads/macarewecare/222927978372067/6442449297429818486/6442449297429818486_signature_31.png
| 2026-01-14 | 2026-01-14 | Pamplona Rural Health Unit | Centro, Pamplona, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31912 | UNOPS | rhupamplonalaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2863 | 2026-01-07 | GXIV-10C | 110023503 | Due for Calibration | Performed Xpertcheck (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test Performed Calibration | Unit is Operational |  | Xpertcheck Kit | 13651 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441828194781585670/6441828194781585670_signature_28.png | Wicki Zena Gerona | https://www.jotform.com/uploads/macarewecare/222927978372067/6441828194781585670/6441828194781585670_signature_31.png
| 2026-01-14 | 2026-01-14 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | daguilon.intp@gmail.com |  | 31911 | MACARE-HBDC | synch28@yahoo.com | 25 |  |  | GXIV-10C | 110001650 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441813573414078181/6441813573414078181_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6441813573414078181/6441813573414078181_signature_31.png
| 2026-01-14 | 2026-01-14 | Jordan Primary Care Facility | Jordan, Guimaras | 6 | Guimaras | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31910 | UNOPS | badette1815@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 422 | 2026-01-14 | GXIV | 110023528 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | all module passed | Xpert Check kit | 5485 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441812923877303395/6441812923877303395_signature_28.png | Bernadette Galila | https://www.jotform.com/uploads/macarewecare/222927978372067/6441812923877303395/6441812923877303395_signature_31.png
| 2026-01-14 | 2026-01-14 | Tondo Foreshore Suoer Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | daguilon.intp@gmail.com |  | 31909 | MACARE-HBDC | synch28@yahoo.com | 25 |  |  | GX16-10C | 110016343 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441811293411227332/6441811293411227332_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6441811293411227332/6441811293411227332_signature_31.png
| 2026-01-14 | 2026-01-14 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31908 | UNOPS | jomarilitonjua@gmail.com | 27 |  | 2026-01-07 | GXIV-10C | 110023495 | Unit for calibration | Performed calibration | #NAME? | Unit is operational All modules passed calibration |  | XpertCheck Kit Lot: 1001495845 | 13653 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441804529516273393/6441804529516273393_signature_28.png | Jomari Litonjua | https://www.jotform.com/uploads/macarewecare/222927978372067/6441804529516273393/6441804529516273393_signature_31.png
| 2026-01-14 | 2026-01-14 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31907 | PBSP | arizzamendoza@gmail.com | 28 |  |  | GXIV-10C | 110012766 | Unit for maintenance | NA | #NAME? | Unit is operational Unit is due for calibration (last calibration date: 6-25-24) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441799909518339859/6441799909518339859_signature_28.png | Arizza Sheena Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/6441799909518339859/6441799909518339859_signature_31.png
| 2026-01-14 | 2026-01-14 | Guimaras Provincial Health Office | Jordan, Guimaras | 6 | Guimaras | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31906 | UNOPS | vjangel31@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 334 | 2026-01-14 | GXIV | 110023540 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | 1 Xpert Check kit | 5484 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441736327614588305/6441736327614588305_signature_28.png | VJ Angeli L. Dalisay | https://www.jotform.com/uploads/macarewecare/222927978372067/6441736327614588305/6441736327614588305_signature_31.png
| 2026-01-14 | 2026-01-14 | Mabuhay Rural Health Unit | Mabuhay, Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31905 | UNOPS | sumondongcrmt@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 850 | 2025-12-11 | GX 4 - 10c | 110023517 | None | Preventive maintenance and Xpertcheck calibration. All modules passed the xpertcheck calibration. | Preventive maintenance and Xpertcheck calibration. | All modules passed and ready to use. | None | XPERTCHECK-CE-5 | 13582 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441726479029766097/6441726479029766097_signature_28.png | Charliza M. Sumondong | https://www.jotform.com/uploads/macarewecare/222927978372067/6441726479029766097/6441726479029766097_signature_31.png
| 2026-01-14 | 2026-01-14 | Sibunag Rural Health Unit | Sibunag, Guimaras | 6 | Guimaras | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31904 | UNOPS | rhodatrance1994@gmail.com | 25 |  | 2026-01-14 | GXIV | 110023538 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | 1 Xpert Check kit | 5487 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441704723823540124/6441704723823540124_signature_28.png | Ma Rhoda Lyn Villarta | https://www.jotform.com/uploads/macarewecare/222927978372067/6441704723823540124/6441704723823540124_signature_31.png
| 2026-01-14 | 2026-01-14 | San Lorenzo Rural Health Unit | San Lorenzo, Guimaras | 6 | Guimaras | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31903 | UNOPS | ellenpearlf@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 863 | 2026-01-14 | GXIV | 110023537 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | 1 Xpert Check kit | 5486 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441672477673250204/6441672477673250204_signature_28.png | ELLEN PEARL F. TALABAN | https://www.jotform.com/uploads/macarewecare/222927978372067/6441672477673250204/6441672477673250204_signature_31.png
| 2026-01-14 | 2026-01-14 | Buenavista Emergency Hospital | Buenavista, Guimaras | 6 | Guimaras | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31902 | UNOPS | janicedelacruz.0909@gmail.com | 28 |  |  | Gx4 | 110023539 | Duedate for xpercheck calibration | Xpertcheck calibration Gather data IQR, syatem log and Calibration info Clean Fan Filter | Xpertcheck calibration Gather data IQR, syatem log and Calibration info Clean Fan Filter | Systems operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441639217675744962/6441639217675744962_signature_28.png | Janice de cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6441639217675744962/6441639217675744962_signature_31.png
| 2026-01-13 | 2026-01-13 | Philippine Tuberculosis Society Inc | E. Rodriguez QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31901 | MACARE-HBDC | sannymisorena@gmail.com | 26 |  |  | GX4 10C | 110027248 | For installation | Installed unit Complete accesories Upload ADF Performed functionality test | Performed installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441044626116076944/6441044626116076944_signature_28.png | Sanny M. Isorena | https://www.jotform.com/uploads/macarewecare/222927978372067/6441044626116076944/6441044626116076944_signature_31.png
| 2026-01-13 | 2026-01-13 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | NRL | NO | INSTALLATION | OTHERS | Noel Floresca |  | florescanoel77@gmail.com |  | Fujifilm | 31900 | MACARE-HBDC | hazelcguevarra04@gmail.com | 26 |  | 2026-01-12 | Limusave MT 7500 | 37020508 | For installation | Delivered and installed unit Performed functionality test Test run unit | Performed installation | Operational | Cooling station s/n 2410016A, thermostation s/n 36920311,1kva ups s/n AT1000181128ST0106 |  | 13665 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441044126114199045/6441044126114199045_signature_28.png | Hazel Guevarra | https://www.jotform.com/uploads/macarewecare/222927978372067/6441044126114199045/6441044126114199045_signature_31.png
| 2026-01-13 | 2026-01-07 | Medi Linx Laboratory Inc. (CLDH) | Hospital Drive Brgy San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31899 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 |  |  | GXII-D | 830841 | Due for Monthly PMS | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module C2: 620096 for module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441021104397008076/6441021104397008076_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6441021104397008076/6441021104397008076_signature_31.png
| 2026-01-13 | 2025-11-12 | Medi Linx Laboratory Inc.(CLDH) | Hospital Dr. Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31898 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 |  | GXIV-D | 839232 | Due for monthly preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unut is operational except module B4: 738203 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6441020744396912682/6441020744396912682_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6441020744396912682/6441020744396912682_signature_31.png
| 2026-01-13 | 2026-01-13 | Cauayan City Health Office | Cauayan City, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31897 | UNOPS | cho1lab14@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1875 | 2026-01-06 | GXIV-10C | 110023504 | Due for Calibration | Performed Xpertcheck (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | 13644 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440998785321469569/6440998785321469569_signature_28.png | Pamela Uy | https://www.jotform.com/uploads/macarewecare/222927978372067/6440998785321469569/6440998785321469569_signature_31.png
| 2026-01-13 | 2026-01-13 | Mambajao Municipal Health Office PCF | Lakas, Mambajao, Camiguin | 10 | Camiguin | TB-DOTS RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31896 | UNOPS | mirasolflores76@gmail.com | 30 |  | 2026-01-07 | Gx 4 10c | 110023519 | UPS don't work. Cant power ON | Troubleshooting UPS | XpertCheck Calibration and Maintenance | All modules are functional Units are functional except for UPS |  | XpertCheck CE-IVD 5 Cartridges/Kit | 13658 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440994710425280689/6440994710425280689_signature_28.png | Mirasol Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6440994710425280689/6440994710425280689_signature_31.png
| 2026-01-13 | 2026-01-13 | Dumarao Primary Care Facility | Dumarao | 6 | Capiz | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31895 | UNOPS | doregoshielamarie@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 725 | 2026-01-13 | GXIV | 110023543 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | all module passed | 1 Xpert Check kit | 5477 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440959066367139744/6440959066367139744_signature_28.png | Shiela Marie D. Dorego | https://www.jotform.com/uploads/macarewecare/222927978372067/6440959066367139744/6440959066367139744_signature_31.png
| 2026-01-13 | 2026-01-13 | Dumangas Municipal Health Health Office | Burgos Regidor, dumangas | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31894 | UNOPS | marilyn_dusaran@yahoo.com | 28 |  |  | Gx4 | 110023532 | Due date for xpertcheck calibration | Xpertcheck calibration Gather data, IQR, system log, assay stats & Xpertcheck calibration | Xpertcheck calibration Gather data, IQR, system log, assay stats & Xpertcheck calibration | Systems operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440950864419781473/6440950864419781473_signature_28.png | Marilyn P. Dusaran | https://www.jotform.com/uploads/macarewecare/222927978372067/6440950864419781473/6440950864419781473_signature_31.png
| 2026-01-13 | 2026-01-13 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31893 | PBSP | bmcgenexpert@gmail.com | 27 |  |  | GXIV | 805113 | Unit for maintenance | NA | #NAME? | Unit is operational Unit is due for calibration (last calibration date: 1-2-25) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440930161519960433/6440930161519960433_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6440930161519960433/6440930161519960433_signature_31.png
| 2026-01-13 | 2026-01-13 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31892 | DOH | bmcgenexpert@gmail.com | 27 |  |  | GXIV | 834623 | Unit for maintenance Monitor problem | Reset VGA cable | #NAME? | Unit is operational Unit is due for calibration (last calibration date: 2-22-24) Monitor is recommended for repair due to intermittent freezing |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440929901513104046/6440929901513104046_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6440929901513104046/6440929901513104046_signature_31.png
| 2026-01-13 | 2026-01-13 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31891 | UNOPS | bmcgenexpert@gmail.com | 25 |  | 2026-01-06 | GXIV-10C | 110023507 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational A1: passed A2: passed A3: passed A4: passed |  | XpertCheck Kit Lot: 1001495845 | 13643 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440929571519408942/6440929571519408942_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6440929571519408942/6440929571519408942_signature_31.png
| 2026-01-13 | 2026-01-13 | Maayon Rural Health Unit | Maayon, Capiz | 6 | Capiz | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31890 | UNOPS | pamskii_28@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 403 | 2026-01-13 | GXIV | 110023541 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | all module passed | 1 Xpert Check kit | 5474 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440864385843533104/6440864385843533104_signature_28.png | pamela b. cerujano | https://www.jotform.com/uploads/macarewecare/222927978372067/6440864385843533104/6440864385843533104_signature_31.png
| 2026-01-13 | 2026-01-13 | Southern Philippines Medical Center - RHIVDA | JP Laurel Avenue | 11 | Davao del Sur | RHIVDA | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31889 | UNOPS | spmcrhivda@gmail.com | 26.5 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2040 ASSAY: Xpert MTB-XDR, COUNT: 14 ASSAY: Xpert SARS-CoV-2, COUNT: 8 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1398 | 2026-01-07 | GX-IV 10C 4D | 110023505 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Icore lenses and changed fan filters Install XPERTCHECK 1.5B Run XpertCheck on four modules Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and XpertCheck Calibration | All Modules passed the Calibration Machine is Operational |  | XPERTCHECK CE 5 | DR 13657 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440832491777931443/6440832491777931443_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6440832491777931443/6440832491777931443_signature_31.png
| 2026-01-13 | 2026-01-13 | Southern Philippines Medical Center - RHIVDA | JP Laurel Avenue, Davao City | 11 | Davao del Sur | RHIVDA | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31888 | UNOPS | spmcrhivda@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1267 ASSAY: Xpert SARS-CoV-2, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1379 | 2026-01-07 | GX-IV 10C 4D | 110023506 | For Preventive Maintenance and XpertCheck Calibration UPS Not working on battery and line mode | Clean Surface Area, Cartridge Bay, Plunger Rod, Icore Lenses and Changed the Fan Filters Install XPERTCHECK 1.5B Run XpertCheck on four modules Generate IQR, Assay Stat and System Log Check the UPS Line Voltages, Battery Voltages, reset its connection | Perform Preventive Maintenance and XpertCheck Perform UPS Troubleshooting | Module Three for Service Required Three Modules only Passed on Calibration Machine is Operational UPS still not working | Reported to Jotform for further assessment | XPERTCHECK CE 5 | DR 13657 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440832111778349658/6440832111778349658_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6440832111778349658/6440832111778349658_signature_31.png
| 2026-01-13 | 2026-01-13 | Jamindan Rural Health Unit | Jamindan, Capiz | 6 | Capiz | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31887 | UNOPS | louiseannaries.laa@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 539 | 2026-01-13 | GXIV | 110023543 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | 1 Xpert Check kit | 5478 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440792264277297625/6440792264277297625_signature_28.png | Louise Ann B. Aries | https://www.jotform.com/uploads/macarewecare/222927978372067/6440792264277297625/6440792264277297625_signature_31.png
| 2026-01-13 | 2026-01-13 | Western Visayas Medical Center | Mandurriao, Iloilo City | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31886 | UNOPS | wvmcrhivdacrcl@gmail.com | 20 |  |  | Gx4 | 110023529 | Duedate for xpertcheck calibration | Xpertcheck calibration Gather data IQR, system log, assay stats & calibration info Clean fan filter | Xpertcheck calibration Gather data IQR, system log, assay stats & calibration info Clean fan filter | Systems operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440750232023956087/6440750232023956087_signature_28.png | Harry Al Villaruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6440750232023956087/6440750232023956087_signature_31.png
| 2026-01-12 | 2026-01-08 | Isulan Rural Health Unit | Isulan, Sultan Kudarat | BARMM | Sultan Kudarat | Tb-lab | NO | CALIBRATION MAINTENANCE Xpertcheck | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31885 | UNOPS | namcymoscoso18@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1166 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 ASSAY: Xpert MTB-XDR, COUNT: 1 | 2026-01-08 | GxlV-10C | 110023520 | None | Perform Genexpert annual Xpertcheck calibration Perform Genexpert machine cleaning and maintenance Modules cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of modules status and problem Data collection | Genexpert machine Xpertcheck Calibration Genexpert machine Preventive Maintenance | 4 operational modules Xpertcheck calibration results -All modules Passed |  | Xpertcheck cartridge CE-5 | DR#:13581 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440297413567339908/6440297413567339908_signature_28.png | Nancy Moscoso | https://www.jotform.com/uploads/macarewecare/222927978372067/6440297413567339908/6440297413567339908_signature_31.png
| 2026-01-12 | 2026-01-12 | Leon Rural Health Unit | LEON, ILOILO | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31884 | UNOPS | camiquejolyn10@gmail.com | 28 |  |  | Gx4 | 110023533 | Due date for xpertcheck calibration | Xpertcheck calibration Clean fan filter Gather data IQR, system log, assay stats and Calibration info | Xpertcheck calibration Clean fan filter Gather data IQR, system log, assay stats and Calibration info | Systems operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440156500416202585/6440156500416202585_signature_28.png | Jolyn C. Maloco | https://www.jotform.com/uploads/macarewecare/222927978372067/6440156500416202585/6440156500416202585_signature_31.png
| 2026-01-12 | 2026-01-12 | Leon Rural Health Unit | LEON, ILOILO | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31883 | UNOPS | camiquejolyn10@gmail.com | 28 |  |  | Gx4 | 110023533 | Due date for xpertcheck calibration | Xpertcheck calibration Clean fan filter Gather data IQR, system log, assay stats and Calibration info | Xpertcheck calibration Clean fan filter Gather data IQR, system log, assay stats and Calibration info | Systems operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440156150411447150/6440156150411447150_signature_28.png | Jolyn C. Maloco | https://www.jotform.com/uploads/macarewecare/222927978372067/6440156150411447150/6440156150411447150_signature_31.png
| 2026-01-12 | 2026-01-12 | Carles Primary Care Facility | Carles, Iloilo | 6 | Iloilo | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31882 | UNOPS | vangieyuro3@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 491 | 2026-01-12 | GXIV | 110023526 | Due for Calibration | #NAME? | #NAME? | #NAME? |  | 1 Xpert Check kit | 5479 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440127945184919479/6440127945184919479_signature_28.png | Vangie C.Yuro | https://www.jotform.com/uploads/macarewecare/222927978372067/6440127945184919479/6440127945184919479_signature_31.png
| 2026-01-12 | 2026-01-12 | Oton Municipal Health Office Diagnostic Center | Oton, Iloilo | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31881 | UNOPS | gracie05191980@gmail.com | 28 |  |  | Gx4 | 110023530 | Duedate for xpertcheck calibration | Xpertcheck calibration Gather data IQR, system log and calibration info Clean fan filter | Xpertcheck calibration Gather data IQR, system log and calibration info Clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440107163914291802/6440107163914291802_signature_28.png | Mary Grace Ong-Lumasag | https://www.jotform.com/uploads/macarewecare/222927978372067/6440107163914291802/6440107163914291802_signature_31.png
| 2026-01-12 | 2026-01-12 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31880 | COMMERCIAL | cduhbloodbank2@gmail.com | 0 |  |  | MCS+ | 14F014 | Spillage in the bowl entry port during procedure. | Check centrifuge chuck and clean up. Test run the centrifuge | Check centrifuge chuck and clean up. Test run the centrifuge | Unit is operational. Recommended for use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440068118831985248/6440068118831985248_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6440068118831985248/6440068118831985248_signature_31.png
| 2026-01-12 | 2026-01-12 | Concepcion Rural Health Unit | Concepcion, Iloilo | 6 | Iloilo | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31879 | UNOPS | jescelmaymesias@gmail.com | 27 |  | 2026-01-12 | GXIV | 110023527 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | Module B3 Retry and Passed | 1 Xpert Check kit | 5490 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6440059759117290413/6440059759117290413_signature_28.png | Jescel May Mesias Melgar | https://www.jotform.com/uploads/macarewecare/222927978372067/6440059759117290413/6440059759117290413_signature_31.png
| 2026-01-12 | 2026-01-12 | Tubungan rural health unit | Tubungan, Iloilo | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31878 | UNOPS | sheryltacadao@gmail.com | 28 |  |  | Gx4 | 110023534 | Due date for xpertcheck calibration | Xpertcheck calibration Clean fan filter Collect data IQR, system log,assay stats and calibration info | Xpertcheck calibration Clean fan filter Collect data IQR, system log,assay stats and calibration info | Sytems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6439987958472667303/6439987958472667303_signature_28.png | SHERYL TACADAO | https://www.jotform.com/uploads/macarewecare/222927978372067/6439987958472667303/6439987958472667303_signature_31.png
| 2026-01-12 | 2026-01-12 | San Rafael Rural Health Unit | San Rafael Iloilo | 6 | Iloilo | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31877 | UNOPS | apilmayfaith@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 285 | 2026-01-12 | GXIV | 110023536 | Due for Calibration | #NAME? | #NAME? | #NAME? | All module passed | 1 Xpert Check kit | 5481 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6439979869024105227/6439979869024105227_signature_28.png | May Faith Apil | https://www.jotform.com/uploads/macarewecare/222927978372067/6439979869024105227/6439979869024105227_signature_31.png
| 2026-01-12 | 2026-01-09 | Mandaue City Health Office | 838508 | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31876 | DOH | mendozarey42@gmail.com | 27 |  |  | 6C GXIV-4D | 838508 | Assay test not found | -Import assay definition file of Xpert HIV-1 Viral Load XC | None | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6439952948832899914/6439952948832899914_signature_28.png | Rey C. Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/6439952948832899914/6439952948832899914_signature_31.png
| 2026-01-12 | 2026-01-12 | Macare In House | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 31875 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | MCS + | 14K081G | For replacement of cpu board | Replaced and installed cpu board Performed calibration. Performed functionality test | Performed repair | Operational CPU board s/n SP082100803G |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6439950733612849079/6439950733612849079_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6439950733612849079/6439950733612849079_signature_31.png
| 2026-01-12 | 2026-01-12 | Macare in house | 1005 Aurora Grand,QC | NCR | NCR 2nd - Quezon City | Engineering | NO | Machine testing | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 31874 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | MCS + | 14K081G | For testing During testing CRC bios error appear on display | Performed cpu board reset Check battery voltage 1.2v Reset unit | Performed troubleshooting | Unit unoperational For replacement of CPU board Def C091428352G |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6439948133616920590/6439948133616920590_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6439948133616920590/6439948133616920590_signature_31.png
| 2026-01-12 | 2026-01-12 | Calinog Rural Health Unit | Calinog iloilo | 6 | Iloilo | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31873 | UNOPS | icjssca@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 660 | 2026-01-12 | GXIV | 110023535 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed for calibration | 1 Xpert Check kit | 5474 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6439903090323163460/6439903090323163460_signature_28.png | Ic Jessica Coronado | https://www.jotform.com/uploads/macarewecare/222927978372067/6439903090323163460/6439903090323163460_signature_31.png
| 2026-01-09 | 2026-01-09 | Makati Medical Center | Amorsolo st.,Makati City | NCR | NCR 4th - Makati | CRM | YES | MAINTENANCE | OTHERS | Noel Floresca |  | florescanoel77@gmail.com | BIOMED@makatimed.net.ph | Biotec | 31872 | COMMERCIAL | AlcorDeeond.Padilla@makatimed.net.ph | 29 |  | 2026-01-08 | Clinimacs | 797 | Due for maintenance | Cleaned unit Test keypad Checked bearings Checked and test clamp Checked program parameters Performed functionality | Performed maintenance | Operational |  |  | 101675 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437543808699990080/6437543808699990080_signature_28.png | Alcor Deeond F. Padilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6437543808699990080/6437543808699990080_signature_31.png
| 2026-01-09 | 2026-01-09 | Makati Medical Center | Amorsolo St., Makati City | NCR | NCR 4th - Makati | Bloodbnak | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | BIOMED@makatimed.net.ph |  | 31871 | COMMERCIAL | Blood.Bank@makatimed.net.ph | 26 |  | 2026-01-08 | MCS+ | 97H330 | Due for maintenance and calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Performed calibration dpm,spm,centrifuge,weigher Replaced 6pcs urethane rollers Performed functionality test Performed simulation running | Performed maintenance and clibration | Operational | Surrender waste material | Urethane rollers | 101674 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437510878691729626/6437510878691729626_signature_28.png | Allysa Grace C. Batronel | https://www.jotform.com/uploads/macarewecare/222927978372067/6437510878691729626/6437510878691729626_signature_31.png
| 2026-01-09 | 2026-01-09 | Iriga City Health Office | Sta Cruz Sur, Iriga | 5 | Camarines Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31870 | UNOPS | mongealyzza@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1520 ASSAY: Xpert MTB-XDR, COUNT: 70 |  | GX4-10C | 110023511 | Gx for Calibration | XpertCheck Result: B1/210117798 - passed B2/210121970 - passed B3/210119501 - passed B4/210120184 - passed Cleaned cartridge bays, plunger rod Generate Data | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437505595326701843/6437505595326701843_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/6437505595326701843/6437505595326701843_signature_31.png
| 2026-01-09 | 2026-01-09 | Ospital ng Makati | Sampaguita St. Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 31869 | MACARE-HBDC | jinx01gs@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 680 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437465553614080177/6437465553614080177_signature_28.png | Cristina Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6437465553614080177/6437465553614080177_signature_31.png
| 2026-01-09 | 2026-01-08 | Marikina City Health Office | Shoe Avenue, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31867 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 537 ASSAY: Xpert MTB-XDR, COUNT: 26 ASSAY: Xpert HIV-1 Viral Load, COUNT: 505 |  | GX-IV-D-10C | 110024073 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437465553616124724/6437465553616124724_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6437465553616124724/6437465553616124724_signature_31.png
| 2026-01-09 | 2026-01-08 | Marikina City Health Office | Shoe Avenue, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | estyaragon2019@gmail.com |  | 31868 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 527 ASSAY: Xpert MTB-XDR, COUNT: 34 ASSAY: Xpert HIV-1 Viral Load, COUNT: 278 |  | GX-IV-D-10C | 110024072 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437465553619711695/6437465553619711695_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6437465553619711695/6437465553619711695_signature_31.png
| 2026-01-09 | 2026-01-09 | Manila Central University Hospital | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | lab_mcuhospital@yahoo.com |  | 31866 | CONSORTIUM | purchasing@mcufdtmf.org | 25 |  |  | GXIV-10C | 110018008 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437465401103037486/6437465401103037486_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6437465401103037486/6437465401103037486_signature_31.png
| 2026-01-09 | 2026-01-08 | Bicol Regional Hospital and Medical Center | Sagpon, Daraga | 5 | Albay | laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31865 | MACARE-HBDC | brtth.covid19lab@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 ASSAY: Xpert HPV, COUNT: 7 ASSAY: Xpert HPV, COUNT: 56 ASSAY: Xpert HIV-1 Viral Load, COUNT: 193 |  | GX4-10C | 110026767 | End user request of removing basic HPV, HIV assay | Removed HIV, HPV assay Cleaned cartridge bays, plunger rod Generate Data | Removed Assay Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437454735321390172/6437454735321390172_signature_28.png | Regina Coeli Naldo | https://www.jotform.com/uploads/macarewecare/222927978372067/6437454735321390172/6437454735321390172_signature_31.png
| 2026-01-09 | 2026-01-07 | Estevez Memorial Hospital Inc. | Rizal st., Legazpi City | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31864 | COMMERCIAL | labestevez@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 6 ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 226 ASSAY: Xpert MTB/RIF Ultra, COUNT: 175 ASSAY: Xpert MTB-XDR, COUNT: 10 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 3 |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays Plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437449045321403278/6437449045321403278_signature_28.png | Daria, Mary Grace M. | https://www.jotform.com/uploads/macarewecare/222927978372067/6437449045321403278/6437449045321403278_signature_31.png
| 2026-01-09 | 2026-01-09 | Southern Philippines Medical Center | JP Laurel Avenue, Bajada | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31863 | UNOPS | spmcrhivda@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1379 ASSAY: Xpert SARS-CoV-2, COUNT: 72 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1267 |  | GX-IV 10C 4D | 110023506 | Stuck Cartidge | Remove the Stuck Cartidge Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Cartidge Removal | Successfully Removed The Cartidge All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6437342341112300737/6437342341112300737_signature_28.png | Christelle Clarisse Sienes RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6437342341112300737/6437342341112300737_signature_31.png
| 2026-01-08 | 2026-01-08 | Cebu City Health Office | Cebu city, cebu | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31862 | UNOPS | laureanjo@yahoo.com | 30 |  |  | Gx4 | 110023513 | N/A | Expertcheck calibration Clean fan filter Gather data, IQR, system log assay stats & calibration info | Expertcheck calibration Clean fan filter Gather data, IQR, system log assay stats & calibration info | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436653864349328271/6436653864349328271_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6436653864349328271/6436653864349328271_signature_31.png
| 2026-01-08 | 2026-01-08 | Padre Burgos Rural Health Unit | Pandre Burgos Southern Leyte | 8 | Southern Leyte | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31861 | UNOPS | keishaelezze14@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 314 | 2026-01-08 | GXIV | 110023514 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | 1 Xpert Check kit | 13575 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436614675595184637/6436614675595184637_signature_28.png | KEISHA ELEZZE C. BATALON | https://www.jotform.com/uploads/macarewecare/222927978372067/6436614675595184637/6436614675595184637_signature_31.png
| 2026-01-08 | 2026-01-07 | Mandaue City Health Office | Mandaue City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31860 | PBSP | mendozarey42@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4830 ASSAY: Xpert HCV VL, COUNT: 110 |  | 6C GXIV-4D | 815275 | None | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436612675223982992/6436612675223982992_signature_28.png | REY C. MENDOZA | https://www.jotform.com/uploads/macarewecare/222927978372067/6436612675223982992/6436612675223982992_signature_31.png
| 2026-01-08 | 2026-01-08 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31859 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2358 |  | 6C GXIV-4D | 838561 | None | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436612675226037422/6436612675226037422_signature_28.png | Hannah Jane Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/6436612675226037422/6436612675226037422_signature_31.png
| 2026-01-08 | 2026-01-08 | Calamba rural health unit | Southwestern poblacion, Calamba, misamis occidental | 10 | Misamis Occidental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31858 | UNOPS | palubonibn@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 928 | 2025-12-11 | GX 4 - 10c | 110023518 | None | Preventive maintenance and xpertcheck calibration for all modules. | Preventive maintenance and xpertcheck calibration for all modules. | All modules passed the calibration and ready to use. | None | XPERTCHECK-CE-5 | 13583 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436540463855448634/6436540463855448634_signature_28.png | Mary-ibn E. Palubon | https://www.jotform.com/uploads/macarewecare/222927978372067/6436540463855448634/6436540463855448634_signature_31.png
| 2026-01-08 | 2026-01-08 | Metro Antipolo Hospital and Medical Center | Marcos highway, Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31857 | COMMERCIAL | mahmci_labcmt@yahoo.com | 20 | ASSAY: Xpert HBV VL, COUNT: 16 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit os operational. Module B2 not operational. Disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436520237817552205/6436520237817552205_signature_28.png | Fatima Busar | https://www.jotform.com/uploads/macarewecare/222927978372067/6436520237817552205/6436520237817552205_signature_31.png
| 2026-01-08 | 2026-01-07 | Santolan Super Health Center | Evangelista Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31856 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1574 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436520227811611046/6436520227811611046_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6436520227811611046/6436520227811611046_signature_31.png
| 2026-01-08 | 2026-01-07 | Rosario Super Health Center | Evangelista Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31855 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1150 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436520227818521339/6436520227818521339_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6436520227818521339/6436520227818521339_signature_31.png
| 2026-01-08 | 2026-01-07 | San Remigio Rural Health Unit | Poblacion San Remigio | 7 | Cebu | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 31854 | WHO | rodacanetecamay@gmail.com | 23 |  |  | GXIV-4 | 847333 | LCD Defective due to earthquake with serial number. | Need to replace new LCD monitor. | . Save System Log file. . Save Installation Qualifications Reoort. . Save Assay Statistics Report. | LCD not ok to use GeneXpert machine functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436514814229448529/6436514814229448529_signature_28.png | Roda Camay | https://www.jotform.com/uploads/macarewecare/222927978372067/6436514814229448529/6436514814229448529_signature_31.png
| 2026-01-07 | 2026-01-07 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | INSTALLATION | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | gsd-biomed@nkti.gov.ph |  | 31853 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 |  |  | GX16D 10C | 110027030 | For installation | Performed installation Upload ADF files Performed functionality test | Performed delivery and installation | Operational Complete accessories |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6436006019418612078/6436006019418612078_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6436006019418612078/6436006019418612078_signature_31.png
| 2026-01-07 | 2026-01-07 | Medi Linx Laboratory Inc (CLDH) | Hospital Drive Brgy San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31852 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 |  |  | GXIV-D | 839232 | Due for Monthly PMS | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module B4: 738203 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6435987144394545791/6435987144394545791_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6435987144394545791/6435987144394545791_signature_31.png
| 2026-01-07 | 2026-01-07 | San Juan District Health Center | Purok 3, Brgy. San Juan, Surigao City | 13 | Surigao del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31851 | UNOPS | sanjuandistricthealthcenter@yahoo.com | 28 |  | 2025-12-10 | GX 4 10c | 110023522 | None | None | XpertCheck Calibration and Maintenance | All modules and units are functional. |  | XpertCheck, CE-IVD 5 Cartridges/Kit | 13573 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6435807517515053669/6435807517515053669_signature_28.png | Anarose Ogdol | https://www.jotform.com/uploads/macarewecare/222927978372067/6435807517515053669/6435807517515053669_signature_31.png
| 2026-01-07 | 2026-01-07 | Philippine Red Cross Tower | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31850 | MACARE-HBDC | marnelli.manuel@gmail.com | 27 |  |  | MCS+ | 08B049 | Unit for maintenance | Set offset values of cuff and weigher | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6435753918125246072/6435753918125246072_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6435753918125246072/6435753918125246072_signature_31.png
| 2026-01-07 | 2026-01-07 | Makilala Rural Health Unit | Poblacion, Makilala North Cotabato | 11 | Cotabato | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31849 | UNOPS | abegailweill17@gmail.com | 23.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 481 ASSAY: Xpert MTB-XDR, COUNT: 2 | 2025-12-10 | GX-IV 10C 4D | 110023521 | For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install XPERTCHECK version 1.5B Run XpertCheck on four modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration All Modules are operational |  | XPERTCHECK CE 5 | DR NO 13572 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6435741183529839106/6435741183529839106_signature_28.png | Abegail W. Publico,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6435741183529839106/6435741183529839106_signature_31.png
| 2026-01-07 | 2026-01-07 | Dipolog City Health Office | Dipolog city | 9 | Zamboanga del Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31848 | UNOPS | mandelzoelugasan@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 ASSAY: Xpert HIV-1 Viral Load, COUNT: 85 | 2025-12-09 | GX 4 - 10c | 110023516 | None | Preventive maintenance, xpertcheck calibration. | Preventive maintenance, xpertcheck calibration. | All modules passed the expertcheck calibration and ready to use. | None | XPERTCHECK-CE-5 | 13571 | Gian Carlo Aizon | https://www.jotform.com/uploads/macarewecare/222927978372067/6435661224718789696/6435661224718789696_signature_28.png | Mandel Zoe B. Lugasan | https://www.jotform.com/uploads/macarewecare/222927978372067/6435661224718789696/6435661224718789696_signature_31.png
| 2026-01-07 | 2026-01-07 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31847 | DOH | gmpelaez144@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1245 |  | GxIV-6C | 838528 | None | Perform Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6435629559023457745/6435629559023457745_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6435629559023457745/6435629559023457745_signature_31.png
| 2026-01-07 | 2026-01-07 | Cotabato City Health Office | Cotabato City | BARMM | Maguindanao | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31846 | UNOPS | gmpelaez144@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1245 ASSAY: Xpert MTB-XDR, COUNT: 5 | 2026-01-07 | GxIV-10C | 110023524 | None | Perform annual Xpertcheck Calibration Perform Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Data collection | Genexpert machine Xpertcheck Calibration Genexpert machine Preventive maintenance | 4 operational modules Xpertcheck results -All modules passed |  | Xpertcheck cartridge CE-5 Lot#:1001436469 | DR#: 13574 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6435627369029376157/6435627369029376157_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6435627369029376157/6435627369029376157_signature_31.png
| 2026-01-07 | 2026-01-07 | Socorro Rural Health Unit | Brgy. Navarro, Socorro, Surigao del Norte | 13 | Surigao del Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31845 | UNOPS | dynamite.tux@gmail.com | 29 |  | 2025-12-11 | GX 4 10c | 110023523 | None | None | XpertCheck Calibration and Preventive Maintenance | All modules and units are functional. |  | XpertCheck, CE-IVD 5 Cartridges/Kit | 13580 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6435590462934457669/6435590462934457669_signature_28.png | Janikka Chluie Guma | https://www.jotform.com/uploads/macarewecare/222927978372067/6435590462934457669/6435590462934457669_signature_31.png
| 2026-01-06 | 2026-01-06 | Cotabato Regional Medical Center-Blood Center | Cotabato City | BARMM | Maguindanao | Blood Center | NO | REPAIR MAINTENANCE | HAEMONETICS | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31844 | MACARE-HBDC | bloodcenter@ro12.doh.gov.ph | 26 |  |  | MCS+ | 15H042G | Pump Rollers problem | Replace 3 defective pump rollers with new pump rollers Perform Heamonetics Preventive maintenance | Heamonetics parts repair Heamonetics machine Preventive maintenance | Heamonetics Apheresis machine operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434948331656356338/6434948331656356338_signature_28.png | Rhea Lim-Macion,RMT,MSMLS | https://www.jotform.com/uploads/macarewecare/222927978372067/6434948331656356338/6434948331656356338_signature_31.png
| 2026-01-06 | 2026-01-06 | Chinese General Hospital | Blumentritt St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 31843 | CONSORTIUM | oncologycgh@gmail.com | 25 |  |  | GXIV | 835604 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434939304028672410/6434939304028672410_signature_28.png | Kendrick Gregorio | https://www.jotform.com/uploads/macarewecare/222927978372067/6434939304028672410/6434939304028672410_signature_31.png
| 2026-01-06 | 2026-01-06 | Chinese General Hospital | Blumentritt St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 31842 | CONSORTIUM | oncologycgh@gmail.com | 25 |  |  | GXIV-19C | 120002026 | For relocation from Bacteriology to Molecular Laboratory. Unit for preventive maintenance. | Relocation and preventive maintenance. | Performed relocation and installed the GeneXpert System. Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434938874028542438/6434938874028542438_signature_28.png | Kendrick Gregorio | https://www.jotform.com/uploads/macarewecare/222927978372067/6434938874028542438/6434938874028542438_signature_31.png
| 2026-01-06 | 2026-01-06 | Metro Health Specialists Hospital Inc | Sorsogon City | 5 | Sorsogon | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31841 | COMMERCIAL | mhshimolecularpathology@gmail.com | 23 | ASSAY: Xpert HBV VL, COUNT: 9 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 32 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX4-10C | 120002716 | GeneXpert not detected | Check and adjusted network configuration to Cepheid settings Cleaned cartridge bays, plunger Generate Data | Performed Preventive Maintenance Performed Troubleshooting | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434901369024077563/6434901369024077563_signature_28.png | AILEEN LOPEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/6434901369024077563/6434901369024077563_signature_31.png
| 2026-01-06 | 2026-01-05 | St. Luke's Medical Center -BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | CI | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31838 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 158 ASSAY: Xpert CT/NG, COUNT: 597 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 150 ASSAY: Xpert HIV-1 Viral Load, COUNT: 348 ASSAY: Xpert HPV, COUNT: 872 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2285 ASSAY: Xpert MTB-XDR, COUNT: 30 ASSAY: Xpert HCV VL, COUNT: 42 |  | GX-IV-D-10C | 120002383 | Broken module B3 door | Replaced module B3 door | Performed maintenance Replaced module B3 door Functionality test Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434899523618750598/6434899523618750598_signature_28.png | Jezelle Nanett Pinic | https://www.jotform.com/uploads/macarewecare/222927978372067/6434899523618750598/6434899523618750598_signature_31.png
| 2026-01-06 | 2026-01-05 | St. Luke's Medical Center -BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | CI | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31839 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 44 ASSAY: Xpert CT/NG, COUNT: 170 ASSAY: Xpert HBV VL, COUNT: 93 ASSAY: Xpert HPV, COUNT: 331 ASSAY: Xpert MTB/RIF Ultra, COUNT: 724 ASSAY: Xpert HCV VL, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 166 |  | GX-IV-D-6C | 839232 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434899523619030031/6434899523619030031_signature_28.png | Jezelle Nanett Pinic | https://www.jotform.com/uploads/macarewecare/222927978372067/6434899523619030031/6434899523619030031_signature_31.png
| 2026-01-06 | 2026-01-05 | St. Luke's Medical Center - BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31840 | COMMERCIAL | jmtentica@stlukes.com.ph | 22 |  |  | MCS+ | 09C125 | For maintenance | Performed maintenance | Performed maintenance: Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434899533616607262/6434899533616607262_signature_28.png | Jen Michelle Entica | https://www.jotform.com/uploads/macarewecare/222927978372067/6434899533616607262/6434899533616607262_signature_31.png
| 2026-01-06 | 2026-01-06 | Candon City Health Office | Brgy. Calaosan Candon, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31837 | DOH | candoncholaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 731 | 2025-12-22 | GXIV-D | 838273 | B4: 807695 for module replacement | Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational B4: 669777 new module Defective module: 807695 |  | Genexpert module | 13629 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434880246544080262/6434880246544080262_signature_28.png | Michelle Bucalen | https://www.jotform.com/uploads/macarewecare/222927978372067/6434880246544080262/6434880246544080262_signature_31.png
| 2026-01-06 | 2026-01-06 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31836 | UNOPS | laureanjo@yahoo.com | 28 |  |  | Gx4 | 110023513 | N/A | Site visit Clean fan filter Gather data IQR, system log | Site visit Clean fan filter Gather data IQR, system log | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434826393525071542/6434826393525071542_signature_28.png | Laurean jo cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6434826393525071542/6434826393525071542_signature_31.png
| 2026-01-06 | 2026-01-06 | San Isidro Rural Health Unit | San Isidro Northern Samar | 8 | Northern Samar | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31835 | UNOPS | hsdelamitas59@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 392 | 2026-01-06 | GX 10C | 110023515 | Dur for Calibration | #NAME? | #NAME? | Unit is Operational |  | Xpert check Kit | 13570 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434754923481447074/6434754923481447074_signature_28.png | Hazel G. Salinas | https://www.jotform.com/uploads/macarewecare/222927978372067/6434754923481447074/6434754923481447074_signature_31.png
| 2026-01-05 | 2025-12-17 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31834 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434253504817938921/6434253504817938921_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6434253504817938921/6434253504817938921_signature_31.png
| 2026-01-05 | 2026-01-05 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31833 | MACARE-HBDC | laboratory@loveyourself.ph | 26 |  | 2025-12-29 | GXIV-10C | 120002891 | Unit for maintenance All four modules for replacement | Replaced modules | - replaced modules A1: 220027774 A2: 220027708 A3: 220027747 A4: 220027712 by 676870, 210047402, 210047888, 887290 respectively - generated reports - performed self test | Unit is operational |  | GeneXpert Modules | 13634 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6434044763824186681/6434044763824186681_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6434044763824186681/6434044763824186681_signature_31.png
| 2026-01-05 | 2025-12-26 | Bicol Regional Hospital and Medical Center | Sagpon, Daraga | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31832 | MACARE-HBDC | brtth.covid19lab@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 ASSAY: Xpert HPV, COUNT: 7 ASSAY: Others, COUNT: 182 ASSAY: Others, COUNT: 56 |  | GX4-6C | 110026767 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6433884666814686920/6433884666814686920_signature_28.png | Regina naldo | https://www.jotform.com/uploads/macarewecare/222927978372067/6433884666814686920/6433884666814686920_signature_31.png
| 2026-01-04 | 2026-01-04 | St lukes Medical Center Global City | Global City Taguig | NCR | NCR 4th - Taguig | Cellular Immunology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomedreports@stlukes.com.ph |  | 31831 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 158 ASSAY: Xpert CT/NG, COUNT: 598 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 146 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 ASSAY: Xpert HPV, COUNT: 874 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2290 ASSAY: Xpert MTB-XDR, COUNT: 30 ASSAY: Xpert HCV VL, COUNT: 42 ASSAY: Xpert HIV-1 Viral Load, COUNT: 328 |  | GX4 10C | 120002383 | As per end user they encounter Stuck cartridge in module B3 Sticky cartridge bay Damage door of b3 module | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Removed stuck cartridge Test run unit | Performed troubleshooting | B3 disable need to replaced door |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6433378476413645111/6433378476413645111_signature_28.png | JOSHUA ALLEN V. AYAAY/ VINCENT-RAQUIZA DY | https://www.jotform.com/uploads/macarewecare/222927978372067/6433378476413645111/6433378476413645111_signature_31.png
| 2026-01-02 | 2025-12-23 | Calamba Medical Center | Crossing, Asian Highway, Calamba, Laguna | 4A | Laguna | Bloodbank | NO | INSTALLATION | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedtech@cmc.ph |  | 31830 | COMMERCIAL | katpayosrmt@gmail.com | 24 |  |  | MCS+ | 19G058SPG | For installation | Installed the MCS+ System | Installed the MCS+ System Verified calibration parameters Replaced rollers Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6431277373528847874/6431277373528847874_signature_28.png | Katrina Librero | https://www.jotform.com/uploads/macarewecare/222927978372067/6431277373528847874/6431277373528847874_signature_31.png
| 2025-12-29 | 2025-12-29 | East Avenue Medical Center | East Ave. QC | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 31829 | MACARE-HBDC | blooddonorsection@gmail.com | 25 |  | 2025-12-29 | MCS + | 15H044 | Centrifuge lock unable to close. | Performed centrifuge lock mechanism troubleshooting. Performed functionality test. Observe performance during actual UPP procedure. | Cheched the centrifuge lock mechanism. Apply lubrication oil on lock mechanism. | Unit is operational |  | N/A | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6428042632114947390/6428042632114947390_signature_28.png | PAULA BETTINA. BARTOLOME | https://www.jotform.com/uploads/macarewecare/222927978372067/6428042632114947390/6428042632114947390_signature_31.png
| 2025-12-29 | 2025-12-29 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | lorieanneaquino@yahoo.com |  | 31828 | COMMERCIAL | laboratory@cmc.ph | 25 |  |  | MCS+ | 19G058SPG | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6427899453318206090/6427899453318206090_signature_28.png | Lorie Anne Hidalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/6427899453318206090/6427899453318206090_signature_31.png
| 2025-12-29 | 2025-12-29 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Microbiology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | aivee.alcantara@medilinx.com.ph |  | 31827 | COMMERCIAL | franchescadeguzman.medilinx@outlook.com | 25 |  |  | GXII | 812916 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6427890393313832606/6427890393313832606_signature_28.png | Franchesca Renee de Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/6427890393313832606/6427890393313832606_signature_31.png
| 2025-12-27 | 2025-12-15 | St.Louis Molecular Diagnostic Laboratory | Tacurong City, Sultan Kudarat | 12 | Sultan Kudarat | Molecular laboratory | NO | MAINTENANCE Cartridge delivery | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31825 | COMMERCIAL | stlouismoleculardiagnosticslab@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 10 ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GxIV-10C | 110024270 | None | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of modules status and errors Data collection | Genexpert machine Preventive maintenance Cartridge delivery | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6426521687322238408/6426521687322238408_signature_28.png | Lyra Huyong | https://www.jotform.com/uploads/macarewecare/222927978372067/6426521687322238408/6426521687322238408_signature_31.png
| 2025-12-27 | 2025-12-18 | Malungon Rural Health Unit | Malungon, Sarangani | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE CPU replacement | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31826 | DOH | madameerrmie@gmail.com | 26.4 |  |  | GxIV-6C | 838546 | Defective CPU | Replace defective CPU SN:MJOBKE21 with new CPU SN:PC1135E9 Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of modules status and errors Data collection | CPU replacement Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6426521687328256372/6426521687328256372_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/6426521687328256372/6426521687328256372_signature_31.png
| 2025-12-27 | 2025-12-27 | Vinzons RHU | Vinzons, Cam Norte | 5 | Camarines Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  | harolduno0905@gmail.com |  |  | 31824 | DOH | bardonamabel@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1222 | 2025-12-19 | GXIV | 838576 | Module B4 for module replacement. | Replaced module B4: 675496. Good Mod SN: 868967 | Performed module replacement. Performed maintenance and cleaning.. | All 4 modules opwrational |  | Good module SN: 868967 | DR NUMBER: 13616 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6426288779355640236/6426288779355640236_signature_28.png | Amabel A. Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/6426288779355640236/6426288779355640236_signature_31.png
| 2025-12-26 | 2025-12-26 | Pili Rural Health Unit | Pili, Cam Sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  | harolduno0905@gmail.com |  |  | 31823 | MACARE-HBDC | mariechristineorcine@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1519 | 2025-12-19 | GXIV | 838408 | Module B4: 662296 for module replacement. | Peplaced module B4 with module SN: 677390. | Perforned module replacement on module B4. Performed maintenance and cleaning. Perforned functionality test. | All 4 modules operational. |  | Defective mod SN: 662296 Good Module SN: 677390 | DR number: 13617 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6425237040894301009/6425237040894301009_signature_28.png | Marie Christine O. Nobleza | https://www.jotform.com/uploads/macarewecare/222927978372067/6425237040894301009/6425237040894301009_signature_31.png
| 2025-12-25 | 2025-12-25 | Asian Hospital and Medical Center | Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31822 | COMMERCIAL | bloodbank@asianhospital.com | 28 |  |  | MCS+ | 01F116 | Valve errors | Checked valve compressor | 0 | Unit is not operational Valve compressor for further troubleshooting/repair |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6424471283684731972/6424471283684731972_signature_28.png | Maritoni Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/6424471283684731972/6424471283684731972_signature_31.png
| 2025-12-24 | 2025-12-24 | Easy Avenue Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31821 | COMMERCIAL | paulinobernaldo1419@gmail.com | 28 |  |  | MCS+ | 15H055 | Centrifuge not locking | Adjusted centrifuge lock ball | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6423625864915403838/6423625864915403838_signature_28.png | Christopher Bernaldo | https://www.jotform.com/uploads/macarewecare/222927978372067/6423625864915403838/6423625864915403838_signature_31.png
| 2025-12-24 | 2025-12-23 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Bleeding room | YES | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31820 | MACARE-HBDC | paulinobernaldo1419@gmail.com | 26 |  |  | MCS + | 15H044G | As per end user they encounter of malfunction to lock the latch cover. | Check latch cover Cleaned latch cover mechanism Checked adjustment height Reset cable Test lock and unlocking of latch cover. | Performed troubleshooting | Unit unoperational After troubleshooting is performed error still occur For further troubleshooting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6423622539416538657/6423622539416538657_signature_28.png | Christopher paul Bernaldo | https://www.jotform.com/uploads/macarewecare/222927978372067/6423622539416538657/6423622539416538657_signature_31.png
| 2025-12-23 | 2025-12-23 | Manila Doctors Hospital | U.N. Ave., | NCR | NCR 1st - Manila | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | xcaguas@maniladoctorscomph.onmicrosoft.com |  | 31819 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 25 |  | 2025-12-19 | GXIV-10C | 110001648 | Unit for preventive maintenance. For module replacement of module B1. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module 10C (210075055) in replacement to defective module (210010344). | SI# 101646 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6422662986336487363/6422662986336487363_signature_28.png | Neriz Magne Burdan | https://www.jotform.com/uploads/macarewecare/222927978372067/6422662986336487363/6422662986336487363_signature_31.png
| 2025-12-22 | 2025-12-22 | Calamba Medical Center | Crossing, Asian Highway, Calamba, Laguna | 4A | Laguna | Bloodbank | NO | INSTALLATION | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | ssacdalan@cmc.ph |  | 31818 | COMMERCIAL | MLibiano@cmc.ph | 24 |  |  | MCS+ | 19G058SPG | For installation | Installed the MCS+ | Installed the MCS+ System -MCS+ unit -UPS -Stand Functionality test Actual running | Unit is not operational. The rollers in the blood pump cracks and came off during actual procedure. Other rollers have signs of damage also. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421964066316011028/6421964066316011028_signature_28.png | Maria Anthonette Libiani | https://www.jotform.com/uploads/macarewecare/222927978372067/6421964066316011028/6421964066316011028_signature_31.png
| 2025-12-22 | 2025-12-22 | National Center For Mental Health | Mandaluyong City | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31817 | UNOPS | iancristrono@gmail.com | 24 |  |  | GXIV-10C | 120003480 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421904915212684015/6421904915212684015_signature_28.png | Ian Cris Trono | https://www.jotform.com/uploads/macarewecare/222927978372067/6421904915212684015/6421904915212684015_signature_31.png
| 2025-12-22 | 2025-11-25 | Capitol Medical Center | Quezon Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31816 | COMMERCIAL | princeosena@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 611 |  | GXIV-D-10C | 110024553 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Cleaned Filter. Performed Self-test. Checked System's Functionality. Generate Reports. | B1 with 2014 Error for Observation. B2 and B3 in Good Condition. B4 for Replacement (Service Required). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421811743619192758/6421811743619192758_signature_28.png | Prince Henry Osena | https://www.jotform.com/uploads/macarewecare/222927978372067/6421811743619192758/6421811743619192758_signature_31.png
| 2025-12-22 | 2025-12-15 | Trece Martires City Health Office | Governor's Drive, City Health Office Building, Brgy San Agustin, Trece Martires, Cavite | 4A | Cavite | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31815 | MACARE-HBDC | loricavivian@yahoo.com | 24 |  | 2025-12-15 | GXIV-D-10C | 110027247 | Unit is due for Installation. | Installation. | Installed GeneXpert System. Checked Functionality. Generate Reports. | All Modules are Good. UPS and Printer are good. |  | GeneXpert System | SI 101586 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421810403615843541/6421810403615843541_signature_28.png | Vivian D. Lorica | https://www.jotform.com/uploads/macarewecare/222927978372067/6421810403615843541/6421810403615843541_signature_31.png
| 2025-12-22 | 2025-12-17 | Muntinlupa Reproductive Health and Wellness Center | Laguerta Health Center, Tunasan, Muntinlupa | NCR | NCR 4th - Muntinlupa | Social Hygiene Clinic | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31814 | MACARE-HBDC | petelerica@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 508 ASSAY: Others, COUNT: 16 | 2025-12-04 | GXIV-D-10C | 110025706 | Unit is due for XpertCheck. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality Generate Report. | B1 - Pass B2 - Pass B3 - Pass B4 - Pass |  | XpertCheck Kit. | DR 13563 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421808963611879920/6421808963611879920_signature_28.png | Erica Monique Petel | https://www.jotform.com/uploads/macarewecare/222927978372067/6421808963611879920/6421808963611879920_signature_31.png
| 2025-12-22 | 2025-12-18 | Capitol Medical Center | Quezon Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31813 | COMMERCIAL | princeosenarmt@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 608 | 2025-12-16 | GXIV-D-10C | 110024553 | B4 is Service Required on XpertCheck. | Module Replacement. | Replaced B4. Checked System's Functionality. Performed Self-test. Generate Reports. | All Modules are Operational. UPS is good. | B4: 210128226 -> 694287 | GeneXpert Module. | DR 13598 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421808963616823660/6421808963616823660_signature_28.png | Prince Henry Osena | https://www.jotform.com/uploads/macarewecare/222927978372067/6421808963616823660/6421808963616823660_signature_31.png
| 2025-12-22 | 2025-12-16 | Dasmariñas City Health Office 3 | Burmas St., Central A, San Marino, Salawag, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31811 | MACARE-HBDC | ajcmalcaraz@gmail.com | 26 |  | 2025-12-15 | GXIV-D-10C | 110027249 | Unit is due for Installation. | Installation. | Install GeneXpert System. Checked Funtionality. Generate Reports. | GX System is Good. |  | GeneXpert System | SI 101601 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421808963614304950/6421808963614304950_signature_28.png | ALYSSA JANE ALCARAZ,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6421808963614304950/6421808963614304950_signature_31.png
| 2025-12-22 | 2025-12-17 | UST Hospital | Dapitan St, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31812 | COMMERCIAL | bloodbank@usthospital.com.ph | 25 |  |  | MCS+ | 02F069 | Unit is due for Maintenance. | Preventive Maintenance. | Checked Parameters and Sensors. Cleaned Pump Rollers. Checked Centrifuge, Bowl Optics and Liquid Detectors. Performed Simulation. Checked UPS Function. | MCS+ is Good. UPS is Operational (With Backup). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421808963618514686/6421808963618514686_signature_28.png | Mylene Balboa | https://www.jotform.com/uploads/macarewecare/222927978372067/6421808963618514686/6421808963618514686_signature_31.png
| 2025-12-22 | 2025-12-19 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Blood donor room | YES | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | nch.jeff22@gmail.com |  | 31810 | MACARE-HBDC | michaeliza1205@gmail.com | 28 |  | 2025-12-16 | MCS + | 17H632SPG | Defective display cable and line sensor | Delivered and installed display cable and line sensor Adjust line sensor value 3071 Checked program parameters Performed functionality test Performed simulation running Endorse waste material display cable and line sensor | Performed repair | Unit operational |  | Display cable and line sensor | 101617/13597 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421725883618889254/6421725883618889254_signature_28.png | Michaeliza Bravante | https://www.jotform.com/uploads/macarewecare/222927978372067/6421725883618889254/6421725883618889254_signature_31.png
| 2025-12-22 | 2025-12-18 | MMC Medilinx | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31809 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 246 ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 634 ASSAY: Xpert HBV VL, COUNT: 102 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 ASSAY: Xpert HPV, COUNT: 363 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3149 ASSAY: Others, COUNT: 16 ASSAY: Xpert HCV VL, COUNT: 162 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1442 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 12 | 2025-12-16 | GX XV1 D 10C | 110026448 | For module replacement A1 210052199 - 794510 For maintenance | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed module replacement and maintenance | Operational |  | Genexpert 10 color module | 13604 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421725313618612534/6421725313618612534_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6421725313618612534/6421725313618612534_signature_31.png
| 2025-12-22 | 2025-12-17 | Philippine Heart Center | East avenue qc | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31808 | MACARE-HBDC | phc_bb@yahoo.com | 28 |  | 2025-11-26 | MCS+ | 11H204 | Unresponsive keyboard Defective display cable Roller worn out | Delivered and installed dispaly cable and roller Checked program parameters Performed functionality test Performed simulation running | Performed repair | Unit operational |  | Dispaly cable and rollers | 1352713529 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6421724613611516568/6421724613611516568_signature_28.png | Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6421724613611516568/6421724613611516568_signature_31.png
| 2025-12-19 | 2025-12-19 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31807 | UNOPS | jomarilitonjua@gmail.com | 23 |  |  | GXIV-10C | 110023495 | Unit for maintenance | NA | #NAME? | Unit is operational Unit is due for calibration (01–16-26) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6419232793423943134/6419232793423943134_signature_28.png | Maria Cecilia Vasquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6419232793423943134/6419232793423943134_signature_31.png
| 2025-12-19 | 2025-12-19 | Tagum City Health Office | Mabini Street, Tagum | 11 | Davao del Norte | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31806 | DOH | cristiarafernandez@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3317 | 2025-11-19 | GX-IV 6C 4D | 838553 | UPS Not Working Broken LCD Monitor display | Replaced the Defective Monitor SN: CK943032R with new Monitor SN: H22V8VL658A250922828 Replaced the Defective UPS SN: AT1000180910ST0672 with new UPS SN: AD1K240612ST0022 Generate IQR, Assay Stat and System Log | Replacement of UPS and Monitor | Monitor and UPS is Operational All Modules are functional |  | FC-IT-UPS-ACCU1K (UPS) FC-E-Monitor | DR 13515 AND DR13481 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6419227002109035959/6419227002109035959_signature_28.png | Cristina Tiara A. Fernandez,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6419227002109035959/6419227002109035959_signature_31.png
| 2025-12-19 | 2025-09-25 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31805 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 11 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 217 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -Icore -Filter Generated reports | Unit is operational. Module B2 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6419178653616943257/6419178653616943257_signature_28.png | Fatima Busar | https://www.jotform.com/uploads/macarewecare/222927978372067/6419178653616943257/6419178653616943257_signature_31.png
| 2025-12-19 | 2025-12-18 | Philippine Red Cross Laguna | Sta. Cruz, Laguna | 4A | Laguna | Bloodbank | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31804 | MACARE-HBDC | lagunablood@redcross.org.ph | 24 |  |  | MCS+ | 02J046 | UPS has no backup | replaced UPS Performed calibration | Replaced UPS Performed maintenance Performed calibration Checked rotor pumps Verified calibration parameters Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6419148493613779496/6419148493613779496_signature_28.png | Zyrine Villacentino | https://www.jotform.com/uploads/macarewecare/222927978372067/6419148493613779496/6419148493613779496_signature_31.png
| 2025-12-19 | 2025-12-17 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Microbiology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31803 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 56 ASSAY: Xpert CT/NG, COUNT: 75 ASSAY: Xpert HBV VL, COUNT: 628 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1121 ASSAY: Xpert HPV, COUNT: 423 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6740 ASSAY: Xpert MTB-XDR, COUNT: 6 ASSAY: Others, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 196 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1226 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge -Plunger rod -icore -Filter Generated reports | Unit is operational. Module D1 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6419148493619774416/6419148493619774416_signature_28.png | Ashley Espiel | https://www.jotform.com/uploads/macarewecare/222927978372067/6419148493619774416/6419148493619774416_signature_31.png
| 2025-12-17 | 2025-12-17 | Quezon City Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31802 | UNOPS | molecularlab.cityhealth@quezoncity.gov.ph | 28 |  |  | GXIV-10C | 120003581 | Unit for maintenance | Na | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6417658194196514927/6417658194196514927_signature_28.png | Ma Cyrell Arcueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6417658194196514927/6417658194196514927_signature_31.png
| 2025-12-17 | 2025-12-17 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31801 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | Dead pixels in display | Replaced display panel Performed calibration | Replaced display panel Performed calibration Replaced rotor pumps Verified calibration parameters Run simulation functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6417640566492173520/6417640566492173520_signature_28.png | Febbie Lou A Birang | https://www.jotform.com/uploads/macarewecare/222927978372067/6417640566492173520/6417640566492173520_signature_31.png
| 2025-12-17 | 2025-12-17 | Medi Linx Laboratory Inc. ( CLDH) | Hospital Drive Brgy San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31800 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 205 |  | GXII-D | 830841 | Due for Monthly PMS | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module C2: 620096 for module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6417565452341985652/6417565452341985652_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6417565452341985652/6417565452341985652_signature_31.png
| 2025-12-17 | 2025-12-17 | Medi Linx Laboratory Inc. ( CLDH) | Hospital Drive Brgy San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31799 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 205 |  | GXIV-D | 839232 | Due for Monthly PMS | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module B4: 738203 for module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6417565022344964915/6417565022344964915_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6417565022344964915/6417565022344964915_signature_31.png
| 2025-12-17 | 2025-12-16 | Cardinal Santos Medical Center | Wilson St. Greenhills San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 31798 | COMMERCIAL | kenzasarmiento5@gmail.com | 24 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Functionality test Cleaned instrument surface | Unit is operational. | Recommended to calibrate the unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6417433773611958500/6417433773611958500_signature_28.png | Kenza Sarmiento | https://www.jotform.com/uploads/macarewecare/222927978372067/6417433773611958500/6417433773611958500_signature_31.png
| 2025-12-16 | 2025-12-16 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31797 | UNOPS | bmcgenexpert@gmail.com | 28 |  |  | GXIV-10C | 110023507 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6416763219913007965/6416763219913007965_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6416763219913007965/6416763219913007965_signature_31.png
| 2025-12-16 | 2025-12-16 | Cardinal Santos Medical Center | Wilson St. San Juan | NCR | NCR 2nd - San Juan | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31796 | COMMERCIAL | janelvelasquez.medilinx@outlook.com | 22 | ASSAY: Xpert HBV VL, COUNT: 26 ASSAY: Xpert HPV, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1222 |  | GX-IV-D-10C | 110003714 | Error 5007 | Performed maintenance. | Performed maintenance Cleaned: -Cartridge bay -plunger rod -Icore -Filter Generated reports | Unit is operational. Modules under observation. Modules B1, B2, and B3 service required on XpertCheck performed last Oct. 17, 2025. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6416758934212267432/6416758934212267432_signature_28.png | Janel Velasquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6416758934212267432/6416758934212267432_signature_31.png
| 2025-12-16 | 2025-12-15 | St Luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 31795 | COMMERCIAL | megapo@stlukes.com.ph | 27 | ASSAY: Xpert CT/NG, COUNT: 204 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2308 |  | GX4D 6C | 839234 | As per enduser they encounter power fluctuation,ups unable to backup supply due to defective ups battery | Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed troubleshooting | B1 and B2 operational B3 and B4 reported already for replacement UPS already replaced the battery of in house personnel | UPS operational |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6416642363615085879/6416642363615085879_signature_28.png | Ma. Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6416642363615085879/6416642363615085879_signature_31.png
| 2025-12-16 | 2025-12-16 | San Joaquin Super Health Center | Elizco Rd. San Joaquin, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31794 | MACARE-HBDC | sanjoaquintml@gmail.com | 30 |  |  | GX-IV-D-10C | 110026463 | For Installation | Installed the GeneXpert System | Installed the GeneXpert System: -GeneXpert unit -Workstation -Printer -UPS Installed assay Generated IQR Functionality test | Unit is operational. | Recommended to install airconditioning unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6416623527225435036/6416623527225435036_signature_28.png | Ma. Noreen L De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/6416623527225435036/6416623527225435036_signature_31.png
| 2025-12-16 | 2025-11-12 | Medi Linx Laboratory Inc. ( CLDH) | Hospital Dr. Brgy. San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31793 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 |  | GXII-D | 830841 | Duefor monthly preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module C2: 620096 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6416612404397447356/6416612404397447356_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6416612404397447356/6416612404397447356_signature_31.png
| 2025-12-16 | 2025-12-16 | Southern Philippines Medical Center - RHIVDA | JP Laurel Avenue, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31792 | UNOPS | spmcrhivda@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1899 ASSAY: Xpert MTB-XDR, COUNT: 13 ASSAY: Xpert SARS-CoV-2, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1441 ASSAY: Xpert CT/NG, COUNT: 6 |  | GX-IV 10C 4D | 110023505 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6416599595619421063/6416599595619421063_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6416599595619421063/6416599595619421063_signature_31.png
| 2025-12-16 | 2025-12-16 | Southern Philippines Medical Center - RHIVDA | JP Laurel Avenue, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31791 | UNOPS | spmcrhivda@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1169 ASSAY: Xpert SARS-CoV-2, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1345 |  | GX-IV 10C 4D | 110023506 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6416599125613458108/6416599125613458108_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6416599125613458108/6416599125613458108_signature_31.png
| 2025-12-15 | 2025-12-15 | Philippine Red Cross | Intramuros, Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31790 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  | 2025-12-15 | MCS+ | 16K521SPG | The SPM Test Failed error. | Performed Repair. Replaced PCB Assembly SPM. Calibrated all parameters. Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Repair and Preventive Maintenance. | Unit is now operational. |  | PCS ASSY SPM (HY33107) in replacement to defective (SP091608347G) | DR#13587 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6415942012126459713/6415942012126459713_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6415942012126459713/6415942012126459713_signature_31.png
| 2025-12-15 | 2025-12-12 | Chong Hua Hospital- Fuente | Fuente Osmena, Capitol Site, Cebu City | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31789 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 40 |  |  | MCS+ | 11J219 | Red cell spillage during procedure. | Adjust line sensor sensitivity. | Adjust line sensor sensitivity. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6415922660611884747/6415922660611884747_signature_28.png | Chong Hua Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6415922660611884747/6415922660611884747_signature_31.png
| 2025-12-15 | 2025-11-27 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31788 | COMMERCIAL | cduhbloodbank2@gmail.com | 40 |  |  | MCS+ | 10M104 | Worn-out rubber roller of draw pump and transfer pump. | Replacement of transfer pump (white) and draw pump (red) | Replacement of transfer pump (white) and draw pump (red) | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6415862860617750131/6415862860617750131_signature_28.png | CDUH BLOODBANK | https://www.jotform.com/uploads/macarewecare/222927978372067/6415862860617750131/6415862860617750131_signature_31.png
| 2025-12-15 | 2025-11-27 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31787 | COMMERCIAL | cduhbloodbank2@gmail.com | 40 |  |  | MCS+ | 14F014 | Worn-out rubber roller of draw pump and transfer pump. | Replacement of transfer pump (white) and draw pump (red) | Replacement of transfer pump (white) and draw pump (red) | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6415861060614453186/6415861060614453186_signature_28.png | CDUH BLOODBANK | https://www.jotform.com/uploads/macarewecare/222927978372067/6415861060614453186/6415861060614453186_signature_31.png
| 2025-12-15 | 2025-12-15 | Region II Trauma and Medical Center | Magsaysay, Bayombong, Nueva Vizcaya | 2 | Nueva Vizcaya | Bloodbank | NO | INSTALLATION | HAEMONETICS | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31786 | MACARE-HBDC | r2tmc.bb@gmail.com | 25 |  |  | MCS+ | 12D166 | Due for installation | Installed MCS+ Installed UPS | Assembled MCS+ Stand | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6415819967796252765/6415819967796252765_signature_28.png | Lilian Dumangeng | https://www.jotform.com/uploads/macarewecare/222927978372067/6415819967796252765/6415819967796252765_signature_31.png
| 2025-12-09 | 2025-10-27 | University of the East Ramon Magsaysay Memorial Medical Center | 64 Aurora Blvd, Quezon City | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedteam@uerm.edu.ph |  | 31785 | COMMERCIAL | bloodbank@uerm.edu.ph | 24 |  |  | MCS+ | 13C270 | Loose display panel | Tighten display panel | Tighten display panel Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6410720093616722320/6410720093616722320_signature_28.png | Maxyme Desiderio | https://www.jotform.com/uploads/macarewecare/222927978372067/6410720093616722320/6410720093616722320_signature_31.png
| 2025-12-09 | 2025-10-28 | Pinagbuhatan TB Microscopy Lab | Urbano Velasco Ave, Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31784 | MACARE-HBDC | peachmango318@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2006 |  | GX-IV-D-10C | 110024245 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6410719983619719063/6410719983619719063_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6410719983619719063/6410719983619719063_signature_31.png
| 2025-12-09 | 2025-11-25 | Pinagbuhatan TB Microscopy Laboratory | Evangelista Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31783 | MACARE-HBDC | peachmango318@gmail.com | 25 | ASSAY: Xpert MTB-XDR, COUNT: 2019 |  | GX-IV-D-10C | 110024245 | Error 2014 on module B3 | Performed maintenance Checked module B3 | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Checked module B3 Generated reports | Unit is operational. Module B3 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6410719823614940145/6410719823614940145_signature_28.png | Arlen Delgado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6410719823614940145/6410719823614940145_signature_31.png
| 2025-12-09 | 2025-11-25 | Rosario Super Health Center | Rosario, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | consalaneljaine@gmail.com |  | 31782 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1094 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6410719663618345505/6410719663618345505_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6410719663618345505/6410719663618345505_signature_31.png
| 2025-12-09 | 2025-11-27 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31781 | UNOPS | acho.laboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 17 ASSAY: Xpert MTB-XDR, COUNT: 176 ASSAY: Xpert HIV-1 Viral Load, COUNT: 865 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6410719423614281870/6410719423614281870_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6410719423614281870/6410719423614281870_signature_31.png
| 2025-12-09 | 2025-12-03 | Manila Doctors Hospital | U.N. Ave, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | xcaguas@maniladoctorscomph.onmicrosoft.com |  | 31780 | COMMERCIAL | molecular.lab@maniladoctors.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 68 ASSAY: Xpert HBV VL, COUNT: 109 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1240 ASSAY: Xpert HCV VL, COUNT: 15 ASSAY: Xpert HIV-1 Viral Load, COUNT: 188 |  | GXIV-D-10C | 110001648 | Unit is due for Maintenance. B2 and B4 with 5007 Errors. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, And Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | B1 is Disabled. (Service Required) B2 and B4 for Observation. B3 is Good. | B2 and B4's 5007 Errors Reported to Cepheid. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6410539613616965952/6410539613616965952_signature_28.png | Earl John Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6410539613616965952/6410539613616965952_signature_31.png
| 2025-12-09 | 2025-12-04 | Marantao Rural Health Unit | Marawi City | BARMM | Lanao del Sur | TB Laboratory | NO | INSTALLATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 31779 | DOH | salhamcyo@gmail.com | 23 |  |  | GXIV-4D | 838549 | CPU defective. | Save System Log file Save installation qualification report Save GXX file. | Replace new CPU with serial number. PC1132FH | CPU are functional and machine are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6410525040616407620/6410525040616407620_signature_28.png | Norsaleah M. Ayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6410525040616407620/6410525040616407620_signature_31.png
| 2025-12-07 | 2025-12-06 | Philippine Red Cross Port Area | Port Area Manila | NCR | NCR 1st - Manila | Blood bank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31778 | MACARE-HBDC | dolly.banares@redcross.org.ph | 28 |  |  | MCS + | 16K521SPG | As per end user they encounter spm error | Checked unit Checked calibration parameters of spm Performed visual check in spm Performed simulation running Performed functionality test | Performed troubleshooting | During simulation running no error encountered Unit for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6408832389413102494/6408832389413102494_signature_28.png | Dolly banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6408832389413102494/6408832389413102494_signature_31.png
| 2025-12-07 | 2025-11-25 | St Luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 31777 | COMMERCIAL | megapo@stlukes.com.ph | 26 | ASSAY: Xpert CT/NG, COUNT: 202 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2343 |  | GX4 D | 839234 | As per end user they encounter error 5007 in module 2 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Restart system Test run unit | Performed troubleshooting and maintenance | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6408831449413860364/6408831449413860364_signature_28.png | Ma. Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6408831449413860364/6408831449413860364_signature_31.png
| 2025-12-07 | 2025-11-25 | St Luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 31776 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 23 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 479 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 ASSAY: Xpert HPV, COUNT: 385 ASSAY: Xpert MTB/RIF Ultra, COUNT: 416 ASSAY: Xpert HCV VL, COUNT: 42 ASSAY: Xpert HIV-1 Viral Load, COUNT: 331 |  | GX4 L | 807155 | Due for maintenance As per end user they encounter loss communication during running in module 3 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Reset gateway board Replaced cable at module 3 Test run unit | Performed maintenance and troubleshooting | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6408830619416190185/6408830619416190185_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6408830619416190185/6408830619416190185_signature_31.png
| 2025-12-06 | 2025-12-04 | Saguiaran Rural Health Unit | Marawe City | 8 | Lanao del Sur | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 31775 | DOH | jdnoel.eisi@gmail.com | 23 |  |  | GXIV 4 | 838298 | LCD monitor defective Module B2 defective | . Save Installation Qualifications Report . Save System Log file . Save GXX file. | Replace old lcd monitor to new lcd monitor with serial number H22V8VL658A250922 | LCD monitor functional ok to use. B2 module not ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6407948027712973240/6407948027712973240_signature_28.png | Asmalia A. Hadji Salic | https://www.jotform.com/uploads/macarewecare/222927978372067/6407948027712973240/6407948027712973240_signature_31.png
| 2025-12-03 | 2025-12-03 | San Lorenzo Municipal Health Office | San Lorenzo, Guimaras | 6 | Guimaras | Laboratory | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31774 | UNOPS | ellenpearlf@gmail.com | 26 |  |  | Gx 4 | 110023537 | Stucked up carthridge | Removal of stucked up carthridge | Removal of stucked up carthridge Replace new fan filter Clean carthedige bay Gather IQR, System log and assay statistics | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6405369549223419514/6405369549223419514_signature_28.png | Ellen Pearl Talaban | https://www.jotform.com/uploads/macarewecare/222927978372067/6405369549223419514/6405369549223419514_signature_31.png
| 2025-12-02 | 2025-12-02 | Marine Science Institute | UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Algal Lab | YES | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 31773 | COMMERCIAL | rtaguilar1@up.edu.ph | 25 |  | 2025-11-28 | Primostar 1 | 3198009534 | None | None | Installed microscope Installed 100x Objective Placed Blue Filter | Unit is operational |  |  | SI 101499 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6404680323322427948/6404680323322427948_signature_28.png | Ronel Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/6404680323322427948/6404680323322427948_signature_31.png
| 2025-12-01 | 2025-12-01 | Biopath Clinical Diagnostics, Inc. - Aurora Grand | 11/F Aurora Grand, 1005 Aurora Blvd., Brgy. Duyan-duyan, Project 3, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Dwyer | 31772 | COMMERCIAL | rmlao@biopath.ph | 22 |  |  | 2300-60PA | Pre-Analysis, Analysis, Post-PCR | Readings on magnehelic devices are minimal | Relocation of (+) and (-) ports of the magnehelic | Checked air seals Verified air flow readings versus those in mechanical contractor's report | Magnehelic gauges are operational Magnehelic readings: Pre-Analysis -6.5Pa Analysis -2Pa Post-Analysis -2.5Pa |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6403756683619517270/6403756683619517270_signature_28.png | Rhenee Lao-Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6403756683619517270/6403756683619517270_signature_31.png
| 2025-11-30 | 2025-11-24 | Region 1 Medical Center | Binloc, Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31771 | UNOPS | sorianocmarygrace.17@gmail.com | 25 |  |  | GXIV-4-D-10C | 110023501 | For Installation | #NAME? | Performed Installation | Unit is Fully Operational | GeneXpert Installed is from Infanta RHU (UNOPS) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6402732922028416356/6402732922028416356_signature_28.png | Mary Grace C. Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/6402732922028416356/6402732922028416356_signature_31.png
| 2025-11-30 | 2025-11-27 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31770 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6402731792022345853/6402731792022345853_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6402731792022345853/6402731792022345853_signature_31.png
| 2025-11-30 | 2025-11-27 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31769 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6402730752027348195/6402730752027348195_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6402730752027348195/6402730752027348195_signature_31.png
| 2025-11-28 | 2025-11-28 | Tropical Disease Foundation | Amorsolo St. Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | aageronimo@tdf.org.ph |  | 31768 | MACARE-HBDC | krcueno@tdf.org.ph | 22 |  |  | GX-IV-D-6C | 812560 | For relocation | Transferred GeneXpert unit to other room | Transferred GeneXpert unit to other room Performed maintenance Generated reports Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6401196964688214924/6401196964688214924_signature_28.png | Karl Francis R. Cueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6401196964688214924/6401196964688214924_signature_31.png
| 2025-11-28 | 2025-11-28 | South Cotabato Provincial Hospital | Koronadal City, South Cotabato | 12 | South Cotabato | Tb-lab | NO | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31767 | MACARE-HBDC | anamariermt22@gmail.com | 28.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2873 ASSAY: Xpert MTB-XDR, COUNT: 49 ASSAY: Xpert HIV-1 Viral Load, COUNT: 368 |  | Gx-XVl | 839314 | Module C2,D2&D3 H/W Failed due to multiple temperature errors | Transfer of Genexpert machine Gx16 from Dr. Arthuro P. Pinggoy Medical Center to South Cotabato Provincial Hospital Perform Genexpert machine Gx16 Re-installation Checking of module status and problem Data collection | Genexpert machine relocation and Re-installation | 12 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6401187271417350036/6401187271417350036_signature_28.png | Ana Marie Grace Egasan | https://www.jotform.com/uploads/macarewecare/222927978372067/6401187271417350036/6401187271417350036_signature_31.png
| 2025-11-28 | 2025-11-28 | Tropical Disease Foundation | Amorsolo St. Makati City | NCR | NCR 4th - Makati | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | aageronimo@tdf.org.ph |  | 31766 | MACARE-HBDC | krcueno@tdf.org.ph | 22 |  |  | GX-IV-D-10C | 110026840 | For installation | Installed the GeneXpert System | Installed the GeneXpert system -GeneXpert unit -Workstation -Barcode Scanner Generated IQR Functionality test | Unit is operational. All modules operational. | Printer is delivered but not installed. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6401181634681556082/6401181634681556082_signature_28.png | Karl Francis Cueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6401181634681556082/6401181634681556082_signature_31.png
| 2025-11-28 | 2025-11-27 | Santolan Superhealth Center | Evangelista Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31765 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1480 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6401163574686808809/6401163574686808809_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6401163574686808809/6401163574686808809_signature_31.png
| 2025-11-28 | 2025-11-28 | Sta. Rosa Community Hospital | Olango Island, Lapu-lapu Island | 7 | Cebu | Laboratory | NO | INSTALLATION CALIBRATION | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31764 | DOH | ariestumz@yahoo.com | 27 |  |  | Gx 4 | 838534 | Defective monitor hp sn CNK9430328 outdated time and date in CPU | Replace new monitor Nvision Sn H22V8VL658A250922989 Install new CMOS battery | Replace new monitor Nvision Sn H22V8VL658A250922989 Install new CMOS battery | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6401088547023049496/6401088547023049496_signature_28.png | Ma.Anunciacian G.Ahat | https://www.jotform.com/uploads/macarewecare/222927978372067/6401088547023049496/6401088547023049496_signature_31.png
| 2025-11-27 | 2025-11-27 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | jniebres@asianhospital.com |  | 31762 | COMMERCIAL | bloodbank@asianhospital.com | 25 |  |  | MCS+ | 10B129 | Unit is due for Maintenance. Busted Fuse. | Preventive Maintenance. | Cleaned Surface Area of the Unit. Checked Sensors' Values. Checked and Cleaned Pumps' Rollers. Run Simulation Test. Replaced 2 2.5A New Fuse. | Unit is Good and Operational. Cart Roller Worn-out |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6400280775362403677/6400280775362403677_signature_28.png | Keano Riggs Caañgay | https://www.jotform.com/uploads/macarewecare/222927978372067/6400280775362403677/6400280775362403677_signature_31.png
| 2025-11-27 | 2025-11-27 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | jniebres@asianhospital.com |  | 31763 | COMMERCIAL | bloodbank@asianhospital.com | 25 |  |  | MCS+ | 01F116 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Surface Area of the Unit. Checked Sensors' Values. Checked and Cleaned Pumps' Rollers. Run Simulation Test. | Unit is in Good Condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6400280775369833411/6400280775369833411_signature_28.png | Keano Riggs Caañgay | https://www.jotform.com/uploads/macarewecare/222927978372067/6400280775369833411/6400280775369833411_signature_31.png
| 2025-11-26 | 2025-11-26 | Baganga Rural Health Unit | Poblacion, Baganga | 11 | Davao Oriental | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31761 | DOH | rhubaganga@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 454 |  | GX-IV 6C 4D | 838395 | UPS not working during battery mode | Replaced the Defective UPS Battery with new Batteries | Perform Battery Replacement | Successfully changed the UPS battery, machine is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6399416622753414992/6399416622753414992_signature_28.png | Liza N. Narciso,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6399416622753414992/6399416622753414992_signature_31.png
| 2025-11-26 | 2025-11-20 | Infanta Rural Health Unit | Infanta, Pangasinan | 1 | Pangasinan | Laboratory | NO | Packing of GX for Transfer R1MC | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31760 | UNOPS | charissealmandres8@gmail.com | 26 |  |  | GXIV-4-D-10C | 110023501 | For Packing and pullout of GeneXpert System | Packed GeneXpert System for Pullout | Packed GeneXpert System for Pullout | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6399281612026650558/6399281612026650558_signature_28.png | Charisse A. Alferez | https://www.jotform.com/uploads/macarewecare/222927978372067/6399281612026650558/6399281612026650558_signature_31.png
| 2025-11-26 | 2025-11-19 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31759 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is Disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6399280162024308135/6399280162024308135_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6399280162024308135/6399280162024308135_signature_31.png
| 2025-11-26 | 2025-11-19 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31758 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is Disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6399279252025283185/6399279252025283185_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6399279252025283185/6399279252025283185_signature_31.png
| 2025-11-25 | 2025-11-25 | Rosario Super Health Center | Rosario, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | consalaneljaine@gmail.com |  | 31757 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1094 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6398684049283808158/6398684049283808158_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6398684049283808158/6398684049283808158_signature_31.png
| 2025-11-25 | 2025-11-25 | Pinagbuhatan TB Microscopy Laboratory | Evangelista Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31756 | MACARE-HBDC | peachmango318@gmail.com | 25 | ASSAY: Xpert MTB-XDR, COUNT: 2019 |  | GX-IV-D-10C | 110024245 | Error 2014 on module B3 | Performed maintenance Checked module B3 | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Checked module B3 Generated reports | Unit is operational. Module B3 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6398684049287700573/6398684049287700573_signature_28.png | Arlen Delgado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6398684049287700573/6398684049287700573_signature_31.png
| 2025-11-25 | 2025-11-25 | Vicente Sotto Memorial Medical Center | Cebu city | 7 | Cebu | Blood bank | YES | REPAIR | HAEMONETICS | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 31755 | COMMERCIAL | vsmmcbloodservices@gmail.com | 23 |  |  | Mcs+ | 16L508SPG | CPU board | Check all part of the machine. | Swap unit haemonetics machine. Sr. 02A090 | Ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6398431793029985547/6398431793029985547_signature_28.png | Charlotte Mae Ygoña | https://www.jotform.com/uploads/macarewecare/222927978372067/6398431793029985547/6398431793029985547_signature_31.png
| 2025-11-24 | 2025-11-24 | Macare in house | Aurora grand,1005 Aurora blvd.QC | NCR | NCR 2nd - Quezon City | Macare in house | NO | Inspection | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 31754 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | MCS + | 19G058SPG | Brand new unit for inspection | Checked unit Performed visual inspection | Performed inspection | Unit findings crack in the cover |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6397982056513543706/6397982056513543706_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6397982056513543706/6397982056513543706_signature_31.png
| 2025-11-24 | 2025-11-18 | Medi Linx Laboratory Inc./ St. Elizabeth Hospital Inc. | General Santos City | 12 | South Cotabato | Microbiology | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | sofia.maclang@medilinx.com.ph |  | 31753 | COMMERCIAL | marylou.villaruel@medilinx.com.ph | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 18 |  | GxIV-6C | 831313 | Module 2 and 4 for module replacement | Perform Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 2 operational modules ( module 1 and 3) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6397973757325792376/6397973757325792376_signature_28.png | Richei Grace Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/6397973757325792376/6397973757325792376_signature_31.png
| 2025-11-24 | 2025-11-24 | Biopath Clinical Diagnostics Inc. | Aurora Grand,1005 Aurora blvd QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | LW Scientific | 31752 | COMMERCIAL | florescanoel77@gmail.com | 26 |  |  | E8 Portafuge | 1101816 | Due for maintenance and calibration | Cleaned unit Checked rotor and motor Checked wear and tear parts Performed functionality test Performed calibration Test run unit | Performed maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6397753743618046036/6397753743618046036_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6397753743618046036/6397753743618046036_signature_31.png
| 2025-11-24 | 2025-11-24 | Biopath Clinical Diagnostics Inc. | Aurora Grand,1005 Aurora blvd QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | LW Scientific | 31751 | COMMERCIAL | florescanoel77@gmail.com | 26 |  |  | Microscope | 118129 | Due for maintenance and calibration | Cleaned unit Cleaned eyepiece and objective Checked mechanical parts Checked coarse and fine adjustment Performed functionality test Test run unit | Performed maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6397750553619062203/6397750553619062203_signature_28.png | ' | https://www.jotform.com/uploads/macarewecare/222927978372067/6397750553619062203/6397750553619062203_signature_31.png
| 2025-11-24 | 2025-11-24 | Biopath Clinical Diagnostics Inc. | Aurora Grand,1005 Aurora blvd QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | LW Scientific | 31750 | COMMERCIAL | florescanoel77@gmail.com | 26 |  |  | Revalation III | 153394 | Due for maintenance and calibration | Cleaned unit Cleaned eyepiece and objective Checked mechanical parts Checked coarse and fine adjustment Performed functionality test Test run unit | Performed maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6397748003617692732/6397748003617692732_signature_28.png | ' | https://www.jotform.com/uploads/macarewecare/222927978372067/6397748003617692732/6397748003617692732_signature_31.png
| 2025-11-24 | 2025-11-20 | Estevez Memorial Hospital, Inc. | Don Juan st., Brgy Cabangan, Legazpi | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31749 | COMMERCIAL | reylovendino83@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 230 ASSAY: Xpert MTB/RIF Ultra, COUNT: 208 ASSAY: Xpert MTB-XDR, COUNT: 11 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 4 |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6397615074519414661/6397615074519414661_signature_28.png | Rey A. Lovendino | https://www.jotform.com/uploads/macarewecare/222927978372067/6397615074519414661/6397615074519414661_signature_31.png
| 2025-11-24 | 2025-11-24 | VICENTE SOTTO MEMORIAL MEDICAL CENTER | Cebu City | 7 | Cebu | Blood Bank | NO | ON-SITE ASSISTANCE Pull out MCS+ | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31748 | MACARE-HBDC | vsmmcbloodservices@gmail.con | 22 |  | 2025-11-24 | MCS+ | 14K081G | -PULL OUT MCS+ WITH SN: 02A090 FOR TROUBLESHOOTING | #NAME? | -INSTALL SERVICE UNIT SN 14K081G -PERFORM FUNCTIONALITY TEST,PASSED | EQUIPMENT IS OPERATIONAL |  |  | 5462 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6397612908975324905/6397612908975324905_signature_28.png | Kristine Jane Agasin | https://www.jotform.com/uploads/macarewecare/222927978372067/6397612908975324905/6397612908975324905_signature_31.png
| 2025-11-23 | 2025-11-23 | St. Lukes Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | biomed.qc@stlukes.com.ph |  | 31747 | MACARE-HBDC | megapo@stlukes.com.ph | 27 |  |  | GXIV | 839234 | Error 5007 on module B2 | Generated reports | - checked and scanned system logs for hardware errors (None found) - checked and scanned test reports for errors (5007—Sample related) - generated reports | Unit is operational No hardware related errors found on modules B1, B2 Unit is overdue of calibration (Mar. 2023 last cal) Unit is advised for calibration Error 5007 to be forwarded to technical specialist team for further validation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6396925269022469873/6396925269022469873_signature_28.png | Cherry Ann P. Pereda | https://www.jotform.com/uploads/macarewecare/222927978372067/6396925269022469873/6396925269022469873_signature_31.png
| 2025-11-22 | 2025-11-22 | Southern Philippines Medical Center | JP Laurel Avenue, Bajada,Davao City | 11 | Davao del Sur | ICU-2 | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 31746 | COMMERCIAL | icu2estuya@gmail.com | 22.5 |  |  | MCS+ | 15HO34G | No power on Machine | Check for line voltages, reset connection, check fuses | Perform Machine Troubleshooting | Machine successfully power on, unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6396184206211817991/6396184206211817991_signature_28.png | Shiela Nicolas | https://www.jotform.com/uploads/macarewecare/222927978372067/6396184206211817991/6396184206211817991_signature_31.png
| 2025-11-22 | 2025-11-22 | Cardinal Santos Medical Center | San Juan | NCR | NCR 2nd - San Juan | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31745 | MACARE-HBDC | mjoylagrana@gmail.com | 27 |  |  | MCS+ | 08M461 | Centrifuge lock problem | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6396106948118631565/6396106948118631565_signature_28.png | Maryjoy Lagrana | https://www.jotform.com/uploads/macarewecare/222927978372067/6396106948118631565/6396106948118631565_signature_31.png
| 2025-11-22 | 2025-11-21 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Apheresis room | NO | INSTALLATION | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | pbcmbd@doh.gov.ph |  | 31744 | MACARE-HBDC | strobecatcher@gmail.com | 26 |  |  | MCS + | 13C269 | For installation | Delivered and installed unit With 1kva ups,transport case,and protocol card,cuff Performed functionality test | Performed installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395940527119350267/6395940527119350267_signature_28.png | Ivan Ritz P. Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6395940527119350267/6395940527119350267_signature_31.png
| 2025-11-22 | 2025-11-21 | Philippine Blood Center | Queson Aveneu QC | NCR | NCR 2nd - Quezon City | Apheresis Room | NO | INSTALLATION | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | pbcmbd@doh.gov.ph |  | 31743 | MACARE-HBDC | strobecatcher@gmail.com | 26 |  |  | MCS + | 16L505SPG | For installation | Delivered and installed unit With 1kva ups,transport case and protocol card,cuff Performed functionality test | Performed installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395940007117144297/6395940007117144297_signature_28.png | Ivan Ritz P. Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6395940007117144297/6395940007117144297_signature_31.png
| 2025-11-22 | 2025-11-21 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31742 | COMMERCIAL | monet.laxina@yahoo.com.ph | 22 |  |  | GX4D | 804205 | Due for maintenance D1,D2 reported failed xpertcheck last 11/15/24 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Unit not in use D2,D3,D4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395939447113751313/6395939447113751313_signature_28.png | Fiel ballena | https://www.jotform.com/uploads/macarewecare/222927978372067/6395939447113751313/6395939447113751313_signature_31.png
| 2025-11-21 | 2026-03-10 | Victor R. Potenciano Medical Center | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar | . | rcb.macare@gmail.com |  |  | 31741 | CONSORTIUM | rcb.macare@gmail.com | 26 |  |  | GXIV | 835600 | Unit for maintenance | Cleaned clogged filter | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395168489514580975/6395168489514580975_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6395168489514580975/6395168489514580975_signature_31.png
| 2025-11-21 | 2025-11-21 | Ospital ng Makati | Sampaguita St. Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 31740 | MACARE-HBDC | jinx01gs@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 632 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395158419804737703/6395158419804737703_signature_28.png | Cristina M. Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6395158419804737703/6395158419804737703_signature_31.png
| 2025-11-21 | 2025-11-21 | Wester Visayas Sanitarium and General Hospital | Sta. Barbara, Iloilo | 6 | Iloilo | TB Dots Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31739 | DOH | ruthiedequilla@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1615 | 2025-10-17 | 6C GXIV-4D | 838393 | Module B1 and B3 defective | -Replace defective module B1/860474 with new module SN: 914872 -Replace defective module B3/860966 with new module SN: 684758 | #NAME? | Machine is functional. Module B2 for reporting |  | 2 Certified pre-owned GeneXpert module SN. 684758 and 914872 | DR no. 13415 and 13418 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395156208218930944/6395156208218930944_signature_28.png | Ruth T. Dequilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6395156208218930944/6395156208218930944_signature_31.png
| 2025-11-21 | 2025-11-21 | Macare in house | Aurora grand 1005 Aurora blvd | NCR | NCR 2nd - Quezon City | Macare 8n house | NO | REPAIR | OTHERS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com | HP monitor | 31738 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | HP pro display | CNK831071H | No power | Repair monitor Checked power supply | Performed repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395074082152114751/6395074082152114751_signature_28.png | Macare in House | https://www.jotform.com/uploads/macarewecare/222927978372067/6395074082152114751/6395074082152114751_signature_31.png
| 2025-11-21 | 2025-11-20 | U.P Diliman | Diliman QC | NCR | NCR 2nd - Quezon City | IB-SCIARM | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  | florescanoel77@gmail.com |  | Airinspace | 31737 | MACARE-HBDC | amcorachea@up.edu.ph | 37 |  |  | T2006-S | 2006-TS-833 | No power Busted fuse | Replaced fuse Test power entry module ok functional | Performed troubleshooting | Unit unoperational For further troubleshooting possible faulty main board After replacing fuse unit still unopertional due to short circuit |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395054992152824584/6395054992152824584_signature_28.png | Allen Joy Corachea | https://www.jotform.com/uploads/macarewecare/222927978372067/6395054992152824584/6395054992152824584_signature_31.png
| 2025-11-21 | 2025-11-19 | Tropical Disease Foundation | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31736 | COMMERCIAL | krcueno@tdf.org.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2087 ASSAY: Xpert HIV-1 Viral Load, COUNT: 863 |  | GX4 D 6C | 812560 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395054622152210352/6395054622152210352_signature_28.png | Karl Francis R. Cueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6395054622152210352/6395054622152210352_signature_31.png
| 2025-11-21 | 2025-11-21 | Jones Rural Health Unit | Jones, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31735 | DOH | oroprecy19@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 684 | 2025-11-06 | GXIV-6C | 838484 | Module 4 for Replacement | Performed Module Replacement B4(613605 - 619545) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Module | 13472 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395022725588679254/6395022725588679254_signature_28.png | Babylene Mamauag | https://www.jotform.com/uploads/macarewecare/222927978372067/6395022725588679254/6395022725588679254_signature_31.png
| 2025-11-21 | 2025-11-21 | Mariano Marcos Memorial Hospital And Medical Center | Batac, Ilocos Norte | 1 | Ilocos Norte | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 31734 | MACARE-HBDC | mmmhmc.molecularlab@gmail.com | 26 |  | 2025-11-12 | GXIV | 839147 | For unit installation. | Performed unit ibstallation. | Perforned unit setup and installation. | Unit is operational. |  | None | DR 13490 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6395011360522938363/6395011360522938363_signature_28.png | Allan Fellizar | https://www.jotform.com/uploads/macarewecare/222927978372067/6395011360522938363/6395011360522938363_signature_31.png
| 2025-11-20 | 2025-11-20 | Perpetual Soccour Hospital | Cebu City | 7 | Cebu | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 31733 | COMMERCIAL | debbiebaga@gmail.com | 28 |  |  | Mcs+ | 11H307 | Platelet spilling during processing | Adjusting line sensor range | Troubleshooting Line Sensor Check rubber pump quality | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6394309663518366214/6394309663518366214_signature_28.png | Debbie Joyce A. Baga | https://www.jotform.com/uploads/macarewecare/222927978372067/6394309663518366214/6394309663518366214_signature_31.png
| 2025-11-20 | 2025-11-20 | Iloilo Provincial Hospital | Pototan, Iloilo | 6 | Iloilo | TB Dots | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31732 | DOH | djpragale92@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1560 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2025-10-01 | 6C GXIV-4D | 838289 | Monitor Defective | #NAME? | #NAME? | Machine is functional |  | Nvision H22V8 21.5" LED Monitor SN. H22V8VL658A250716588 | DR no. 13374 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6394265193123461742/6394265193123461742_signature_28.png | Rachel Pragale | https://www.jotform.com/uploads/macarewecare/222927978372067/6394265193123461742/6394265193123461742_signature_31.png
| 2025-11-20 | 2025-11-20 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31731 | MACARE-HBDC | laboratory@loveyourself.ph | 27 |  |  | GXIV | 844726 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6394262429747968071/6394262429747968071_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6394262429747968071/6394262429747968071_signature_31.png
| 2025-11-20 | 2025-11-20 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31730 | MACARE-HBDC | lanoratory@loveyourself.ph | 27 |  |  | GXIV | 807956 | Unit for maintenance | #NAME? | NA | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6394261989742986886/6394261989742986886_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6394261989742986886/6394261989742986886_signature_31.png
| 2025-11-20 | 2025-11-20 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31729 | MACARE-HBDC | laboratory@loveyourself.ph | 26 |  |  | GXIV | 807112 | Unit for installation Unit for maintenance | Installed unit | - installed unit along with 2 other GX units via switch - checked and cleaned cartridge bay, plunger rod, filter, icore slot - generated reports - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6394261619746775891/6394261619746775891_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6394261619746775891/6394261619746775891_signature_31.png
| 2025-11-20 | 2025-11-20 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31728 | UNOPS | laboratory@loveyourself.ph | 27 |  | 2025-11-13 | GXIV-10C | 120002891 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules failed calibration All modules for reporting |  | XpertCheck Kit Lot: 1001436469 Exp. Date: 5-10-2026 | 13493 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6394261199746157053/6394261199746157053_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6394261199746157053/6394261199746157053_signature_31.png
| 2025-11-20 | 2025-11-20 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Microbiology | NO | Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | renalyn.cabaltea@medilinx.com.ph |  | 31727 | COMMERCIAL | ellaine.gatchalian@medilinx.com.ph | 22 | ASSAY: Xpert HBV VL, COUNT: 31 ASSAY: Xpert HPV, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1186 |  | GX-IV-D-10C | 110003714 | For relocation | Transferred the GeneXpert System | Transferred the GeneXpert System to its original location. -GeneXpert unit -Workstation -barcode Scanner -Printer -UPS Performed maintenance | Unit is operational. | No additional testing needed. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6394255513618680185/6394255513618680185_signature_28.png | Ellaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/6394255513618680185/6394255513618680185_signature_31.png
| 2025-11-20 | 2025-11-20 | Calamba rural health unit | Calamba misamis occidental | 9 | Misamis Occidental | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31726 | UNOPS | palubonibn@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 885 | 2025-11-20 | GX 4 | 110023518 | None | Preventive maintenance | Preventive maintenance | All modules are available angld ready to use. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6394162888014213528/6394162888014213528_signature_28.png | MARY-ibn E. PALUBON | https://www.jotform.com/uploads/macarewecare/222927978372067/6394162888014213528/6394162888014213528_signature_31.png
| 2025-11-19 | 2025-11-19 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31725 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | For maintenance Checked display Run simulation Verified calibration parameters Checked rotor pumps Functionality test | Unit is operational | Display panel for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6393480613612256442/6393480613612256442_signature_28.png | Jhan Edisel B. Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6393480613612256442/6393480613612256442_signature_31.png
| 2025-11-19 | 2025-11-19 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31724 | COMMERCIAL | rpmercado92@gmail.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6393480413616396370/6393480413616396370_signature_28.png | Roxanne Paulina Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/6393480413616396370/6393480413616396370_signature_31.png
| 2025-11-19 | 2025-11-19 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31723 | COMMERCIAL | moleculardx@themedicalcity.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 56 ASSAY: Xpert CT/NG, COUNT: 75 ASSAY: Xpert HBV VL, COUNT: 624 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1055 ASSAY: Xpert HPV, COUNT: 427 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6731 ASSAY: Xpert MTB-XDR, COUNT: 5 ASSAY: Others, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 184 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1311 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module D1/210082300 not operational. For module replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6393480173619218886/6393480173619218886_signature_28.png | Ashley Espiel | https://www.jotform.com/uploads/macarewecare/222927978372067/6393480173619218886/6393480173619218886_signature_31.png
| 2025-11-19 | 2025-11-19 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31722 | COMMERCIAL | moleculardx@themedicalcity.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 56 ASSAY: Xpert CT/NG, COUNT: 75 ASSAY: Xpert HBV VL, COUNT: 624 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1055 ASSAY: Xpert HPV, COUNT: 427 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6731 ASSAY: Xpert MTB-XDR, COUNT: 5 ASSAY: Others, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 184 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1311 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module D1/210082300 not operational. For module replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6393479823616729936/6393479823616729936_signature_28.png | Ashley Espiel | https://www.jotform.com/uploads/macarewecare/222927978372067/6393479823616729936/6393479823616729936_signature_31.png
| 2025-11-19 | 2025-11-19 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31720 | COMMERCIAL | rpmercado92@gmail.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6393479823616380753/6393479823616380753_signature_28.png | Roxanne Paulina Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/6393479823616380753/6393479823616380753_signature_31.png
| 2025-11-19 | 2025-11-19 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31721 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | For maintenance Checked display Run simulation Verified calibration parameters Checked rotor pumps Functionality test | Unit is operational | Display panel for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6393479823617230031/6393479823617230031_signature_28.png | Jhan Edisel B. Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6393479823617230031/6393479823617230031_signature_31.png
| 2025-11-19 | 2025-11-19 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | biomed@uerm.edu.ph |  | 31719 | CONSORTIUM | lpcomendador@uerm.edu.ph | 26 |  |  | GXIV | 847578 | Unit for maintenance | NA | #NAME? | Unit is operational Module B3 has temperature errors since 2024 Unit is overdue of calibration (March 21, 2025) Unit is advised to be calibrated |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6393435770716992390/6393435770716992390_signature_28.png | Osang Comendador | https://www.jotform.com/uploads/macarewecare/222927978372067/6393435770716992390/6393435770716992390_signature_31.png
| 2025-11-19 | 2025-11-19 | Dipolog City Health Office | Circumferential road biasong dipolog city | 9 | Zamboanga del Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31718 | UNOPS | jz.ebron@yahoo.com.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 830 | 2025-11-19 | GX 4 | 110023516 | None | Preventive Maintenance | Preventive Maintenance | All modules are available and ready to use | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6393415590018400779/6393415590018400779_signature_28.png | James Zebedee C. Ebron | https://www.jotform.com/uploads/macarewecare/222927978372067/6393415590018400779/6393415590018400779_signature_31.png
| 2025-11-18 | 2025-11-18 | Biopath Clinical Diagnostics Inc | Aurora Grand,1005 Aurora blvd QC | NCR | NCR 2nd - Quezon City | Lab | NO | CALIBRATION MAINTENANCE | OTHERS | Noel Floresca |  | florescanoel77@gmail.com |  | Randox | 31717 | COMMERCIAL | rmlao@biopath.ph | 26 |  |  | Evidence Investigator | E1-24820 | Due for maintenance and calibration | Cleaned unit Checked and test camera alignment Checked wear and tear parts Performed functionality test Performed calibration | Performed maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392907875425120311/6392907875425120311_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6392907875425120311/6392907875425120311_signature_31.png
| 2025-11-18 | 2025-11-18 | Biopath Clinical Diagnostics Inc | Aurora Grand ,1005 Aurora Blvd QC | NCR | NCR 2nd - Quezon City | Lab | NO | CALIBRATION MAINTENANCE | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | Randox | 31716 | COMMERCIAL | rmlao@biopath.ph | 26 |  |  | Evidence Multistat | 052-23-0711 | Due for maintenance and calibration | Cleaned unit Checked and test lcd display Checked motors Checked all connection Performed functionality test Performed calibration | Performed maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392905375423965207/6392905375423965207_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6392905375423965207/6392905375423965207_signature_31.png
| 2025-11-18 | 2025-11-12 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | molecularlab@phc.gov.ph |  | 31715 | MACARE-HBDC | avabichuela3216@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 347 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 162 ASSAY: Xpert SARS-CoV-2, COUNT: 7 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 4 | 2025-11-07 | GX16D 6C | 838655 | Due for maintenance and xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and xpertcheck | A1,A2,A3,A4,B1,B4,C1,C2,C4 Passed in xpertcheck. B2,B3,C3 Failed in xpertcheck D1,D2,D3,D4 xpertcheck not done insufficient cartridges |  | Xpertcheck kits | 13482 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392900885427285364/6392900885427285364_signature_28.png | Philippine Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6392900885427285364/6392900885427285364_signature_31.png
| 2025-11-18 | 2025-11-13 | De Los Santos Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular lab | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@dlsmc.ph |  | 31714 | COMMERCIAL | mndionisio@dlsmc.ph | 38 | ASSAY: Xpert HBV VL, COUNT: 49 ASSAY: Xpert HPV, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 556 |  | GX4 (10C) | 110020700 | Due for maintenance As per end user they encounter error 5007 in module 3 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B3 for reporting B1,B2,B4 in use Last xpertcheck performed 3/8/2024 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392900025427931588/6392900025427931588_signature_28.png | De Los Santos Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6392900025427931588/6392900025427931588_signature_31.png
| 2025-11-18 | 2025-11-18 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Microbiology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | renalyn.cabaltea@medilinx.com.ph |  | 31713 | COMMERCIAL | ellaine.gatchalian@medilinx.com.ph | 22 | ASSAY: Xpert HBV VL, COUNT: 33 ASSAY: Xpert HPV, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1191 |  | GX-IV-D-10C | 110003714 | For temporary relocation | Relocated the GeneXpert system | Transferred the GeneXpert System to other room Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. | No additional testing required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392574758394078916/6392574758394078916_signature_28.png | Ellaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/6392574758394078916/6392574758394078916_signature_31.png
| 2025-11-18 | 2025-11-05 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31712 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is Disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392543702027264333/6392543702027264333_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6392543702027264333/6392543702027264333_signature_31.png
| 2025-11-18 | 2025-11-18 | Sta Cruz Rural Health Unit | Poblacion, Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31711 | DOH | chendalagan72@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 539 |  | GX-IV 6C 4D | 838505 | UPS not working during battery mode | Replace the defective batteries with new batteries Check the line voltages | Perform UPS battery replacement | Successfully changed the battery, UPS is working during battery mode All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392522566325324989/6392522566325324989_signature_28.png | Dorothy Dalagan,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6392522566325324989/6392522566325324989_signature_31.png
| 2025-11-18 | 2025-11-18 | Teodoro B. Galagar District Hospital | Jagna, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31710 | DOH | venusaton4@gmail.com | 27 |  |  | Gx 4 | 838292 | Defective Module B3 | Gather data IQR, System log and assay stats Clean Fan and Fan Filter New installed module B3 sn 913026 | New installed module B3 sn 913026 | Systems operational There are temperature error related occured in module B1 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392479936914891614/6392479936914891614_signature_28.png | Venus Aton | https://www.jotform.com/uploads/macarewecare/222927978372067/6392479936914891614/6392479936914891614_signature_31.png
| 2025-11-18 | 2025-11-18 | Macare Medicals | 1005 Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Accupower | 31709 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  |  | AT1000 | AT1000190402ST0400 | No power | None | Troubleshooting using good battery Test back-up mode | Unit not operational. For decommissioning |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392474463619208708/6392474463619208708_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6392474463619208708/6392474463619208708_signature_31.png
| 2025-11-18 | 2025-11-18 | Macare Medicals | 1005 Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Accupower | 31708 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  |  | AT1000 | AT1000211227ST0141 | No power | None | Troubleshooting using good battery Test back-up mode | Unit not operational. For decommissioning |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392472973618896051/6392472973618896051_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6392472973618896051/6392472973618896051_signature_31.png
| 2025-11-18 | 2025-11-18 | Macare Medicals | 1005 Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Accupower | 31707 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  |  | AT1000 | AT1000180709ST0052 | No power | None | Troubleshooting using good battery Test back-up mode | Unit not operational. For decommissioning |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392471833615889536/6392471833615889536_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6392471833615889536/6392471833615889536_signature_31.png
| 2025-11-18 | 2025-11-18 | Macare Medicals | 1005 Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Accupower | 31706 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  |  | AT1000 | AT1000180709ST0077 | No power | None | Troubleshooting using good battery Test back-up mode | Unit not operational. For decommissioning |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6392470433616966534/6392470433616966534_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6392470433616966534/6392470433616966534_signature_31.png
| 2025-11-17 | 2025-11-12 | Calabanga Rural Health Unit | Poblacion, Calabanga | 5 | Camarines Sur | Laboratory | YES | REPAIR | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31705 | DOH | martyarc11@yahoo.com | 25 |  |  | GX4-6C | 838398 | CPU not powering on | Repaired CPU Data collected | Performed Troubleshooting | Unit is Operational All Gx 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6391577563411789691/6391577563411789691_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/6391577563411789691/6391577563411789691_signature_31.png
| 2025-11-17 | 2025-07-31 | Antipolo City Health Office | M. Santos St, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31704 | UNOPS | acho.laboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert MTB-XDR, COUNT: 98 ASSAY: Xpert HIV-1 Viral Load, COUNT: 725 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Wll modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565883614762155/6391565883614762155_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565883614762155/6391565883614762155_signature_31.png
| 2025-11-17 | 2025-07-31 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31703 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 251 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B2 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565723611647181/6391565723611647181_signature_28.png | Fatima Busar | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565723611647181/6391565723611647181_signature_31.png
| 2025-11-17 | 2025-10-28 | Pinagbuhatan TB Microscopy Lab | Urbano Velasco Ave, Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31702 | MACARE-HBDC | peachmango318@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2006 |  | GX-IV-D-10C | 110024245 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565493613451010/6391565493613451010_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565493613451010/6391565493613451010_signature_31.png
| 2025-11-17 | 2025-10-27 | University of the East Ramon Magsaysay Memorial Medical Center | 64 Aurora Blvd, Quezon City | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedteam@uerm.edu.ph |  | 31701 | COMMERCIAL | bloodbank@uerm.edu.ph | 24 |  |  | MCS+ | 13C270 | Loose display panel | Tighten display panel | Tighten display panel Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565293616359453/6391565293616359453_signature_28.png | Maxyme Desiderio | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565293616359453/6391565293616359453_signature_31.png
| 2025-11-17 | 2025-09-25 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31700 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 11 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 217 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -plunger rod -Icore -Filter Generated reports | Unit is operational. Module B2 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565103615681802/6391565103615681802_signature_28.png | Fatima Busar | https://www.jotform.com/uploads/macarewecare/222927978372067/6391565103615681802/6391565103615681802_signature_31.png
| 2025-11-17 | 2025-11-14 | Pasig Treatment Hub - Sumilang | Dr. García St, Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31699 | MACARE-HBDC | marcabapo@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 596 |  | GX-IV-D-10C | 110025128 | Error 2014 on module A2 | Performed maintenance | Performed maintenance Checked module Cleaned: -Cartridge bay -Plunger rod -Icore -filter Generated reports | Unit is operational. All modules operational. Module A2 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6391562163612341742/6391562163612341742_signature_28.png | Marc francis R. Abapo, RN | https://www.jotform.com/uploads/macarewecare/222927978372067/6391562163612341742/6391562163612341742_signature_31.png
| 2025-11-17 | 2025-11-14 | Pasig Treatment Hub - CHAMP | Champ building, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | pathmedtech@gmail.com |  | 31698 | MACARE-HBDC | pasigtreatment@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 646 | 2025-11-14 | GX-IV-D-10C | 110025127 | Module B3 for replacement. | Replaced module B3 | Performed maintenance Replaced module B3 Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  | Old: B3/210129340 New: B3/210001991 | DR# 13494 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6391562163615301777/6391562163615301777_signature_28.png | Pasig Treatment Hub | https://www.jotform.com/uploads/macarewecare/222927978372067/6391562163615301777/6391562163615301777_signature_31.png
| 2025-11-17 | 2025-11-17 | . | 1005,Aurora Blvd. | NCR | NCR 2nd - Quezon City | . | NO | CALIBRATION MAINTENANCE | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | . | 31697 | COMMERCIAL | florescanoel77@gmail.com | 26 |  |  | . | . | . | . | . | . |  |  | G 8013 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6391502823613588379/6391502823613588379_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6391502823613588379/6391502823613588379_signature_31.png
| 2025-11-14 | 2025-11-14 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31696 | COMMERCIAL | gramanmolecularlab@gmail.com | 22 |  |  | GXIV | 110008530 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Modules B1, B2, B3 awaiting for module allocation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6389001497919255270/6389001497919255270_signature_28.png | Maria Elizabeth S. de Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/6389001497919255270/6389001497919255270_signature_31.png
| 2025-11-14 | 2025-11-14 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31693 | COMMERCIAL | agalang@asianhospital.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 10 ASSAY: Xpert CT/NG, COUNT: 67 ASSAY: Xpert HBV VL, COUNT: 186 ASSAY: Xpert HPV, COUNT: 210 ASSAY: Xpert HIV-1 Viral Load, COUNT: 167 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1205 ASSAY: Others, COUNT: 92 |  | GXIV-D-10C | 110014996 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Cleaned Filter. Performed Self-test. Generate Reports. | Modules 1,3 and 4 are in good condition. Module 2 with 2014 error for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388963191782086866/6388963191782086866_signature_28.png | Alyssa Galang | https://www.jotform.com/uploads/macarewecare/222927978372067/6388963191782086866/6388963191782086866_signature_31.png
| 2025-11-14 | 2025-10-24 | Romblon District Hospital | Romblon,Romblon | 4B | Romblon | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31694 | DOH | rdhlab5500@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 636 |  | GXIV-D-6C | 838573 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Cleaned Filter. Performed Self-test. Generate Reports. | All Modules are Operational. GX System is Okay. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388963191785295524/6388963191785295524_signature_28.png | SHEILAH MAE R. PUNAY | https://www.jotform.com/uploads/macarewecare/222927978372067/6388963191785295524/6388963191785295524_signature_31.png
| 2025-11-14 | 2025-11-07 | Dasmariñas City Health Office II | DBB-E , Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31695 | DOH | hazelmilan0777@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3427 | 2025-11-06 | GXIV-D-6C | 838535 | B2 with Temperature Error. Unit is due for Maintenance. | Module Replacement. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Replaced Module 2. Performed Self-test. Generate Reports. | Module 2 Newly Replaced. B1, B3, B4 - Good Condition. | B2: Old(638084) -> New(897828) | GeneXpert Module. | DR 13473 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388963191788138417/6388963191788138417_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/6388963191788138417/6388963191788138417_signature_31.png
| 2025-11-13 | 2025-11-13 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank Service Unit (Sub-National) | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31692 | COMMERCIAL | vsmmcbloodservices@gmail.com | 24 |  |  | MCS+ | 02A090 | Red valve out of position problem persisting during platelet procedure. | For replacement pneumatic assembly unit of the MCS+. | Troubleshoot and check pneumatic assembly unit. | Machine is under repair for replacement of pneumatic assembly unit. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388454330616500042/6388454330616500042_signature_28.png | Charlotte Mae Ygoña, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6388454330616500042/6388454330616500042_signature_31.png
| 2025-11-13 | 2025-11-13 | Philippine Heart Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31691 | MACARE-HBDC | ronacarmelamdl@gmail.com | 28 |  |  | MCS+ | 13B243 | Unit for check up AC Pump error | Temporary replaced AC Pump rotor roller | - checked rotor roller pumps (worn out AC pump rubbers) - checked calibration parameters (adjusted bowl offset from 200 to 25) - performed self test | Unit is operational Rotor roller rubbers for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388362597915583723/6388362597915583723_signature_28.png | Rona Carmela M. De Luna | https://www.jotform.com/uploads/macarewecare/222927978372067/6388362597915583723/6388362597915583723_signature_31.png
| 2025-11-13 | 2025-11-12 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Microbiology | NO | MAINTENANCE | OTHERS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph | Fujifilm | 31690 | COMMERCIAL | alegalit@stlukes.com.ph | 22 |  |  | Limusave MT-7500 | 47020656 | For maintenance | Performed maintenance | Performed maintenance Cleaned instrument surface Checked filter Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388256403611008081/6388256403611008081_signature_28.png | Antoniela Leih Galit | https://www.jotform.com/uploads/macarewecare/222927978372067/6388256403611008081/6388256403611008081_signature_31.png
| 2025-11-13 | 2025-11-12 | St. Luke's Medical Center - BGC | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | CI | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31687 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 37 ASSAY: Xpert CT/NG, COUNT: 170 ASSAY: Xpert HBV VL, COUNT: 55 ASSAY: Xpert HPV, COUNT: 310 ASSAY: Xpert MTB/RIF Ultra, COUNT: 663 ASSAY: Xpert HCV VL, COUNT: 14 ASSAY: Xpert HIV-1 Viral Load, COUNT: 142 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. Module A2/866696 has error 1001. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388256403615573626/6388256403615573626_signature_28.png | Joshua Allen V. Aya-Ay/ Vincent-Raquiza Dy Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6388256403615573626/6388256403615573626_signature_31.png
| 2025-11-13 | 2025-11-12 | St. Luke's Medical Center - BGC | Rizal Dr, BGC -; Taguig City | NCR | NCR 4th - Taguig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31689 | COMMERCIAL | bloodbank.bgc@stlukes.com.ph | 22 |  |  | MCS+ | 09C125 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388256403618544346/6388256403618544346_signature_28.png | Aliah Yvonne Carandang | https://www.jotform.com/uploads/macarewecare/222927978372067/6388256403618544346/6388256403618544346_signature_31.png
| 2025-11-13 | 2025-11-12 | St. Luke's Medical Center - BGC | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | CI | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31688 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 168 ASSAY: Xpert CT/NG, COUNT: 599 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 125 ASSAY: Xpert HPV, COUNT: 904 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2297 ASSAY: Xpert MTB-XDR, COUNT: 20 ASSAY: Xpert HCV VL, COUNT: 39 ASSAY: Xpert HIV-1 Viral Load, COUNT: 330 |  | GX-IV-D-6C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388256403619993561/6388256403619993561_signature_28.png | Joshua Allen V. Aya-Ay/ Vincent-Raquiza Dy Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6388256403619993561/6388256403619993561_signature_31.png
| 2025-11-13 | 2025-11-13 | Manila Central University | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | lab_mcuhospital@yahoo.com |  | 31686 | CONSORTIUM | purchasing@mcufdtmf.org | 25 |  | 2025-11-05 | GXIV-10C | 110018008 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed the calibration |  | XpertCheck Kit Lot: 1001480606 Expiry Date: 3-21-27 | 101336 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388239017913461769/6388239017913461769_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6388239017913461769/6388239017913461769_signature_31.png
| 2025-11-13 | 2025-11-13 | Laoang Rural Health Unit | Laoang northern Samar | 8 | Northern Samar | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31685 | DOH | rosarioadarayan26@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1093 | 2025-11-13 | GXIV | 838202 | Module B1/ 807336 Defective | #NAME? | -Module B1 replacement SN: 877387 - Preventive Maintenance | Unit is Operational - Module B4 defective reported to PBSP |  |  | 13417 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388175780524954340/6388175780524954340_signature_28.png | Rosario B.Adarayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6388175780524954340/6388175780524954340_signature_31.png
| 2025-11-13 | 2025-11-13 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31684 | DOH | glennmacauba22@gmail.com | 25 |  | 2025-11-06 | GXIV | 838225 | Unit for preventive maintenance For module replacement of module B1. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B1 (210004539) in replacement to defective module B1 (739631). | DR#13474 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388160508511621984/6388160508511621984_signature_28.png | Glenn A. Macauba | https://www.jotform.com/uploads/macarewecare/222927978372067/6388160508511621984/6388160508511621984_signature_31.png
| 2025-11-13 | 2025-11-13 | SILLIMAN UNIVERSITY MEDICAL CENTER FOUNDATION, INC. | Dumaguete City | 7 | Negros Oriental | Molecular | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com | bangalando23@gmail.com |  | 31683 | COMMERCIAL | sumcmoleclab@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 12 ASSAY: Xpert C. difficile BT, COUNT: 12 ASSAY: Xpert CT/NG, COUNT: 75 ASSAY: Xpert HBV VL, COUNT: 183 ASSAY: Xpert HPV, COUNT: 269 ASSAY: Xpert MTB/RIF Ultra, COUNT: 659 ASSAY: Xpert SARS-CoV-2, COUNT: 12 ASSAY: Xpert HCV VL, COUNT: 66 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 |  | GX4 6C | 821471 | E2126, Communication Lost Errors, Module Reset for A1 and A4 | . | #NAME? | A1 A3, A4 - operational A2 - for replacement, module for purchase |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6388097393571908182/6388097393571908182_signature_28.png | SUMCFI MOLECULAR LABORATORY | https://www.jotform.com/uploads/macarewecare/222927978372067/6388097393571908182/6388097393571908182_signature_31.png
| 2025-11-12 | 2025-11-12 | Manila Doctors Hospital | U. n. Ave., Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | xcaguas@maniladoctorscomph.onmicrosoft.com |  | 31682 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 25 |  |  | GXIV-10C | 110001648 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6387403599119145103/6387403599119145103_signature_28.png | Kaye martelino | https://www.jotform.com/uploads/macarewecare/222927978372067/6387403599119145103/6387403599119145103_signature_31.png
| 2025-11-12 | 2025-11-12 | Medi Linx Laboratory Inc. ( CLDH) | Hospital Dr. Brgy. San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31681 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 |  | GXII-D | 830841 | Duefor monthly preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module C2: 620096 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6387306446541977121/6387306446541977121_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6387306446541977121/6387306446541977121_signature_31.png
| 2025-11-12 | 2025-11-12 | Medi Linx Laboratory Inc.(CLDH) | Hospital Dr. Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31680 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 |  | GXIV-D | 839232 | Due for monthly preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unut is operational except module B4: 738203 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6387306026549351706/6387306026549351706_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6387306026549351706/6387306026549351706_signature_31.png
| 2025-11-12 | 2025-11-06 | Isulan Rural Health Unit | Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31679 | UNOPS | isulanrhuclin.lab@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1182 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 ASSAY: Xpert MTB-XDR, COUNT: 1 |  | GxIV-10C | 110023520 | None | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6387106217329189628/6387106217329189628_signature_28.png | February jen apiado | https://www.jotform.com/uploads/macarewecare/222927978372067/6387106217329189628/6387106217329189628_signature_31.png
| 2025-11-11 | 2025-11-07 | Philippine Red Cross QC | Kalayaan Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31678 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 26 |  |  | MCS + | 02A227 | Due for calibration and maintenance | Cleaned unit Checked rollers Checked and test tube sensor Checked and test keypad and lcd display. Checked program parameters Performed functionality test | Performed calibration and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6386404063617028517/6386404063617028517_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/6386404063617028517/6386404063617028517_signature_31.png
| 2025-11-11 | 2025-11-07 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Apheresis rm | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | pbcmbd@doh.gov.ph |  | 31677 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15H030G | Due for calibration and maintenance Urethane rollers slight worn out | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed calibration and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6386403473614310846/6386403473614310846_signature_28.png | Ivan Ritz P. Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6386403473614310846/6386403473614310846_signature_31.png
| 2025-11-11 | 2025-11-07 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Apheresis rm | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | pbcmbd@doh.gov.ph |  | 31676 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15H029G | Due for Calibration and Maintenance Urethane rollers slight worn out | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed calibration Performed functionality test Performed simulation running | Performed calibration and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6386403003617629483/6386403003617629483_signature_28.png | Ivan Ritz P. Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6386403003617629483/6386403003617629483_signature_31.png
| 2025-11-11 | 2025-11-07 | Maluso rural health unit | Maluso, basilan | BARMM | Basilan | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 31675 | DOH | sharminahussin@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 597 | 2025-10-17 | GX 4 | 838835 | Module 1 temperature errors. | Replaced defective module 863146 with new module 761436. Conducted preventive maitenance. | Replaced module 1 and conducted preventive maintenance. | All modules are working and ready to use. | None | 900-041RS | 13416 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6386399097165893097/6386399097165893097_signature_28.png | Mudzna D.Salihuddin | https://www.jotform.com/uploads/macarewecare/222927978372067/6386399097165893097/6386399097165893097_signature_31.png
| 2025-11-11 | 2025-11-07 | Marikina City Health Office | Shoe Avenue, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31674 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 468 ASSAY: Xpert MTB-XDR, COUNT: 25 ASSAY: Xpert HIV-1 Viral Load, COUNT: 267 |  | GX-IV-D-10C | 110024072 | for maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6386345463612486860/6386345463612486860_signature_28.png | Monica Sarmiento | https://www.jotform.com/uploads/macarewecare/222927978372067/6386345463612486860/6386345463612486860_signature_31.png
| 2025-11-11 | 2025-11-07 | Marikina City Health Office | Shoe Avenue, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31673 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 465 ASSAY: Xpert MTB-XDR, COUNT: 25 ASSAY: Xpert HIV-1 Viral Load, COUNT: 452 |  | GX-IV-D-10C | 110024073 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6386345463614142484/6386345463614142484_signature_28.png | Monica Sarmiento | https://www.jotform.com/uploads/macarewecare/222927978372067/6386345463614142484/6386345463614142484_signature_31.png
| 2025-11-11 | 2025-11-07 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31672 | COMMERCIAL | almuetebianca1@mail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 12 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module B2 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6386345463614688742/6386345463614688742_signature_28.png | Bianca Almuete | https://www.jotform.com/uploads/macarewecare/222927978372067/6386345463614688742/6386345463614688742_signature_31.png
| 2025-11-07 | 2025-11-07 | St. Luke’s Extension Clinic | Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Patrick Los Baños |  | plosbanos@macare.medicals.com |  |  | 31671 | MACARE-HBDC | laboratory@slec.ph | 23 |  | 2025-11-06 | Gx Infinity | 809643 | -> Unit for xpertcheck | -> None | -> Performed preventive maintenance cleaning -> Performed xpertcheck | -> unit is operational -> modules E3, F3, F4 passed xpertcheck -> modules E4, F2, G2, G3 failed xpertcheck |  | -> xpertcheck x2 | DR13470 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6383165598483577782/6383165598483577782_signature_28.png | MA. SARAH R. SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/6383165598483577782/6383165598483577782_signature_31.png
| 2025-11-07 | 2025-11-07 | St Lukes Extension Clinic | J Bocobo St, Ermita | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31670 | MACARE-HBDC | laboratory@slec.ph | 24 |  |  | GXIV | 815686 | Unit for XpertCheck | None | Performed preventive maintenance cleaning Performed XpertCheck | Unit is operational, all modules passed XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6382951670937940978/6382951670937940978_signature_28.png | Lexanne Krysshia Par | https://www.jotform.com/uploads/macarewecare/222927978372067/6382951670937940978/6382951670937940978_signature_31.png
| 2025-11-07 | 2025-11-07 | St Lukes Extension Clinic | J Bocobo St., Ermita | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31669 | MACARE-HBDC | laboratory@slec.ph | 24 |  |  | GXIV | 808092 | Unit for XpertCheck | None | Performed preventive maintenance cleaning Performed XpertCheck | Unit is operational, all modules passed XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6382950590939240113/6382950590939240113_signature_28.png | Lexanne Krysshia Par | https://www.jotform.com/uploads/macarewecare/222927978372067/6382950590939240113/6382950590939240113_signature_31.png
| 2025-11-07 | 2025-11-07 | St. Luke’s Extension Clinic | Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Patrick Los Baños |  | plosbanos@macare.medicals.com |  |  | 31668 | MACARE-HBDC | laboratory@slec.ph | 24 |  |  | GX IV | 110016786 | Unit for xpertcheck | none | -> performed preventive maintenance cleaning -> performed xpertcheck | unit is operational, all modules passed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6382950278488237308/6382950278488237308_signature_28.png | Lexanne Krysshia Par | https://www.jotform.com/uploads/macarewecare/222927978372067/6382950278488237308/6382950278488237308_signature_31.png
| 2025-11-07 | 2025-11-06 | Biopath QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31667 | COMMERCIAL | rmlao@biopath.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 38 ASSAY: Others, COUNT: 1 ASSAY: Xpert HPV, COUNT: 75 ASSAY: Xpert MTB/RIF Ultra, COUNT: 17 ASSAY: Others, COUNT: 7 ASSAY: Xpert SARS-CoV-2, COUNT: 16 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX16 D | 839306 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | A1,A4,B2,B4,C2,D2 failed xpertcheck disabled A3 hardware failure disabled A2,B1,B3,C1,C3,C4,D1,D3,D4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6382846042625508406/6382846042625508406_signature_28.png | Rhenee Lao Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6382846042625508406/6382846042625508406_signature_31.png
| 2025-11-07 | 2025-11-06 | Best Diagnostic Corp. | Masikap st., QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31666 | COMMERCIAL | bestdiagnosticlaboratory@gmail.com | 19 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 2 |  | GX4D 6C | 110008529 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational Last xpertcheck 2/7/2024 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6382845392629015456/6382845392629015456_signature_28.png | Best Diagnostic Corp | https://www.jotform.com/uploads/macarewecare/222927978372067/6382845392629015456/6382845392629015456_signature_31.png
| 2025-11-06 | 2025-11-06 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31665 | UNOPS | bmcgenexpert@gmail.com | 28 |  |  | GXIV-10C | 110023507 | Unit for maintenance | Na | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6382162144408474580/6382162144408474580_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6382162144408474580/6382162144408474580_signature_31.png
| 2025-11-06 | 2025-10-30 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31664 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS+ | 08H188 | Due for maintenance Urethane rollers need to be replace slight worn out | Cleaned unit Checked rollers Checked and test tube sensor Checked and test keypad and lcd display. Checked program parameters Performed functionality test | Performed maintenance | For quotation urethane roller Unit still can be used |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6382084883313821465/6382084883313821465_signature_28.png | Kristia Mae Bigornia | https://www.jotform.com/uploads/macarewecare/222927978372067/6382084883313821465/6382084883313821465_signature_31.png
| 2025-11-06 | 2025-10-30 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 31663 | UNOPS | eamcmolecularlab@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 1100 ASSAY: Xpert HBV VL, COUNT: 68 ASSAY: Xpert HPV, COUNT: 82 ASSAY: Xpert SARS-CoV-2, COUNT: 658 ASSAY: Xpert SARS-CoV-2, COUNT: 127 ASSAY: Xpert HIV-1 Viral Load, COUNT: 11 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 518 |  | GX4D 10C | 110016781 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Back files Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6382084383314599115/6382084383314599115_signature_28.png | Christelle Kimberly Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6382084383314599115/6382084383314599115_signature_31.png
| 2025-11-05 | 2025-11-05 | Governor Generoso Municipal Health Office | Poblacion, Governor Generoso | 11 | Davao Oriental | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31662 | DOH | josephmagallanes72@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 459 | 2025-10-15 | GX-IV 6C 4D | 838490 | Module B2 encounter multiple errors For Preventive Maintenance | Replace the defective Module SN: 615764 with new module SN:912461 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | GeneXpert Module 6C | DR NO. 13406 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6381348532912919917/6381348532912919917_signature_28.png | Charlene Magallanes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6381348532912919917/6381348532912919917_signature_31.png
| 2025-11-05 | 2025-11-05 | Chinese General Hospital | Blumentritt, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 31661 | COMMERCIAL | cghmcbloodbank2016@gmail.com | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. Unit for calibration. | Preventive maintenance and calibration. | Performed Calibration. Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6381255464119007467/6381255464119007467_signature_28.png | Matromanrey M. Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/6381255464119007467/6381255464119007467_signature_31.png
| 2025-11-05 | 2025-11-05 | Chinese General Hospital | Blumentritt, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31660 | COMMERCIAL | cghoncolab@gmail.com | 25 |  |  | GXIV | 835607 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6381231433917188431/6381231433917188431_signature_28.png | SANDRA ABIGAIL CLARIDAD | https://www.jotform.com/uploads/macarewecare/222927978372067/6381231433917188431/6381231433917188431_signature_31.png
| 2025-11-05 | 2025-11-05 | Chinese General Hospital | Blumentritt, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Bacteriology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31659 | COMMERCIAL | cgh_clinicalpathology@yahoo.com | 25 |  |  | GXIV-10C | 120002026 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. For schedule of transfer the system to Molecular laboratory. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6381221523919363255/6381221523919363255_signature_28.png | Maria Odessa R. Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/6381221523919363255/6381221523919363255_signature_31.png
| 2025-11-05 | 2025-11-05 | Manila Med | UN Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31658 | COMMERCIAL | jrcabasug.manilamed@gmail.com | 25 |  |  | GXIV | 835540 | Unit for preventive maintenance. For assay definition file loading. | Preventive maintenance and ADF loading. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Loaded ADFs. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6381168928226539594/6381168928226539594_signature_28.png | Jericho Nolan Cabasug | https://www.jotform.com/uploads/macarewecare/222927978372067/6381168928226539594/6381168928226539594_signature_31.png
| 2025-11-04 | 2025-11-03 | Bagumbayan Rural Health Unit | Bagumbayan,Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31657 | DOH | marizalenoid@yahoo.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1449 | 2025-11-03 | GxIV-6C | 838261 | Module 2 H/W failed due to temperature errors | Replace defective module B2 SN:624012 with new Module SN: 897368 Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine module repair Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert 6 color module SN:897368 | DR #: 13419 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6380745557325385085/6380745557325385085_signature_28.png | Ma. RIZA R. DIONELA, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6380745557325385085/6380745557325385085_signature_31.png
| 2025-11-04 | 2025-10-29 | Bicol Regional Hospital and Medical Center | Sagpon, Daraga | 5 | Albay | Laboratory | NO | REPAIR | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31656 | MACARE-HBDC | labandadrinah@gmail.com | 19 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 ASSAY: Xpert HPV, COUNT: 3 ASSAY: Others, COUNT: 24 |  | GX4-10C | 110026767 | B4/210138841 - ERROR 4014 | Replaced module: B4 - 210138841>>210024476 Generate Data | Replaced B4/210138841 | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6380419697915552818/6380419697915552818_signature_28.png | Drinah L. Labanda | https://www.jotform.com/uploads/macarewecare/222927978372067/6380419697915552818/6380419697915552818_signature_31.png
| 2025-11-03 | 2025-11-03 | Southern Philippines Medical Center - ICU | JP Laurel Avenue, Bajada, Davao City | 11 | Davao del Sur | ICU | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 31655 | COMMERCIAL | spmcimdept@gmail.com | 25.4 |  |  | MCS+ | 15H034G | Machine not working, enduser can't perform TPE | Check for Line Voltages, reset connection, double check the values and operability of the UPS and MCS+ Machine | Perform Machine Troubleshooting | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6379739091977923937/6379739091977923937_signature_28.png | Shiela Cres Nicolas | https://www.jotform.com/uploads/macarewecare/222927978372067/6379739091977923937/6379739091977923937_signature_31.png
| 2025-11-03 | 2025-10-30 | Philippine Red Cross - Laguna | Sta. Cruz, Laguna | 4A | Laguna | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31654 | COMMERCIAL | lagunablood@redcross.org.ph | 22 |  |  | MCS+ | 02J046 | For maintenance | Performed maintenance | Performed maintenance Run simulation Verified calibration parameters Adjusted cuff value Functionality test checked UPS | Unit is operational. UPS has no backup power. For replacement. UPS SN: 20163711301 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6379565363613611447/6379565363613611447_signature_28.png | ZYRINE T VILLACENTINO | https://www.jotform.com/uploads/macarewecare/222927978372067/6379565363613611447/6379565363613611447_signature_31.png
| 2025-11-03 | 2025-10-29 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31653 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXIV-4-D-6C | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is Disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6379401588325564215/6379401588325564215_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6379401588325564215/6379401588325564215_signature_31.png
| 2025-10-30 | 2025-10-30 | NKTI | East Ave. QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | THERMO FISHER SCIENTIFIC | Harold Toco |  | htoco@macare-medicals.com |  |  | 31652 | COMMERCIAL | janettequirino@yahoo.com | 25 |  | 2025-10-30 | Indiko Plus | 864000000213 | Worn out cuvette dispenser belt. | Replaced cuvette dispenser belt. Performed water blank and extra wash Run calibrator and control. | Replaced cuvette dispenser belt. | Unit is operational. |  | Cuvette dispenser Belt | DR num: 13455 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6376172392217218660/6376172392217218660_signature_28.png | Janette V. Quirino | https://www.jotform.com/uploads/macarewecare/222927978372067/6376172392217218660/6376172392217218660_signature_31.png
| 2025-10-30 | 2025-10-30 | Davao Reproductive Health and Wellness Center | Jacinto Street, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31651 | MACARE-HBDC | rhwclab@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 400 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2732 |  | GX-IV 6C 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Three Operational Modules only Module 3 encounter multiple temperature related errors | Error sent to office for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6376109536601113007/6376109536601113007_signature_28.png | RHEA NIERVO | https://www.jotform.com/uploads/macarewecare/222927978372067/6376109536601113007/6376109536601113007_signature_31.png
| 2025-10-29 | 2025-10-29 | St. Luke’s Extension Clinic | Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31650 | MACARE-HBDC | laboratory@slec.ph | 21 |  | 2025-10-28 | GeneXpert Infinity | 809643 | -> conveyer needs parts replacement -> 8 modules for module replacement after XpertCheck | -> repaired conveyer -> calibrated conveyer sensors -> replaced 8 modules (220001426 > 692061, 710723 > 651705, 701127 > 690277, 792461 > 734837, 647149 > 875016, 652279 > 910590, 611037 > 912929, 695487 > 210062409) -> performed functionality test | -> repaired conveyer -> replaced 8 modules | Unit is operational |  | -> GeneXpert module x8 -> modified ejector for GeneXpert Infinity module x8 -> assy assembly door GeneXpert module x8 -> coupling (part for the infinity system) | DR#13443 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6375332286598590168/6375332286598590168_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6375332286598590168/6375332286598590168_signature_31.png
| 2025-10-29 | 2025-10-29 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 31649 | COMMERCIAL | csmclab@csmc.ph | 22 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Run simulation Functionality test Slight adjustment of DPM and SPM | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6375306093617356288/6375306093617356288_signature_28.png | Angel Lynn Grace Pagulayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6375306093617356288/6375306093617356288_signature_31.png
| 2025-10-29 | 2025-10-28 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Blood donation room | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | nch.jeff22@gmail.com |  | 31648 | MACARE-HBDC | michaeliza1205@gmail.com | 26 |  |  | MCS + | 17H632SPG | Due for maintenance Line sensor malfunction of potentiometer unable to adjust to desired value 3071 line sensor potentionmeter can adjust only to 2614 Left pole loose rod | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Line sensor assy need to replace for quotation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6375058663616799770/6375058663616799770_signature_28.png | Michaeliza Bravante | https://www.jotform.com/uploads/macarewecare/222927978372067/6375058663616799770/6375058663616799770_signature_31.png
| 2025-10-28 | 2025-10-23 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31647 | DOH | sherryannemarrero@gmail.com | 22 |  | 2025-10-17 | GXIV-4-D-6C | 838281 | For Module Replacement | - Cleaned Cartridge Bays,Plunger Rods and I-CORE - Replaced Defective Module B3 716721 >>>695483 - Generated Reports | Performed Module Replacement | Unit is Fully Operational |  | GeneXpert Module 695483 | DR 13414 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6374424135717830796/6374424135717830796_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6374424135717830796/6374424135717830796_signature_31.png
| 2025-10-28 | 2025-10-21 | Ilocos Training and Regional Medical Center | San Fernando City, La Union | 1 | La Union | Molecular Laboratory | NO | CALIBRATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31646 | MACARE-HBDC | itrmc.mbl@gmail.com | 29 |  | 2025-10-09 | GXIV-4-D-10C | 120003504 | For Calibration | #NAME? | Performed Calibration | All Modules Passed XpertCheck |  | XpertCheck Kits | DR 13392 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6374422585716769930/6374422585716769930_signature_28.png | Dominga Patacsil | https://www.jotform.com/uploads/macarewecare/222927978372067/6374422585716769930/6374422585716769930_signature_31.png
| 2025-10-28 | 2025-10-28 | Cotabato City Health Office | Cotabato City | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31645 | UNOPS | gmpelaez144@gmail.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1273 |  | GxIV-10C | 110023524 | None | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6374309724555870637/6374309724555870637_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6374309724555870637/6374309724555870637_signature_31.png
| 2025-10-28 | 2025-10-28 | Cotabato City Health Office | Cotabato City | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31644 | UNOPS | gmpelaez144@gmail.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1273 |  | GxIV-10C | 110023524 | None | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6374309394556952682/6374309394556952682_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6374309394556952682/6374309394556952682_signature_31.png
| 2025-10-28 | 2025-10-28 | Cotabato City Health Office | Cotabato City | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31643 | UNOPS | gmpelaez144@gmail.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1273 |  | GxIV-10C | 110023524 | None | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6374308284551027032/6374308284551027032_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6374308284551027032/6374308284551027032_signature_31.png
| 2025-10-28 | 2025-10-24 | Vinzons Rural Health Unit | Poblacion, Vinzons | 5 | Camarines Norte | Laboratory | YES | REPAIR | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31642 | DOH | jasminlottidp@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1205 |  | GX4-6C | 838576 | Module B2/862906 - temperature Error | Replaced module B2: 862906>>610183 Collected Data | Replaced B2/862906 | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6373926556412017699/6373926556412017699_signature_28.png | Jasmin Lotti A. Dela Peña | https://www.jotform.com/uploads/macarewecare/222927978372067/6373926556412017699/6373926556412017699_signature_31.png
| 2025-10-28 | 2025-10-23 | Naga City General Hospital | Balatasan, Naga City | 5 | Camarines Sur | Clinical Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | paulinefederis@gmail.com |  | 31641 | MACARE-HBDC | ncgh.clinicallaboratory@gmail.com | 25 |  |  | GX4-10C | 120003510 | Gx for Calibration Gx for Preventive Maintenance | XpertCheck result: B1/220030041 - passed B2/220030085 - passed B3/220030032 - passed B4/220030017 - passed Cleaned cartridge bays, plunger rod Generate Data | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6373926246416845978/6373926246416845978_signature_28.png | Franz Pauline M. Federis | https://www.jotform.com/uploads/macarewecare/222927978372067/6373926246416845978/6373926246416845978_signature_31.png
| 2025-10-27 | 2025-10-27 | Cotabato Regional Medical Center - Blood Center | Cotabato City | BARMM | Maguindanao | Blood Center | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | hkdrdsales@gmail.com |  | 31640 | MACARE-HBDC | bloodcenter@ro12.doh.gov.ph | 48 |  |  | MCS+ | 15HO42G | Bloody Plasma collection during procedure | Perform Preventive maintenance and cleaning Checking of Line Sensor Value and status Checking Optics bowl value and status Checking of pump rollers | Perform MCS+ Heamonetics problem troubleshooting | MCS+ Heamonetics is operational/working All pump and sensors is in normal range value |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6373618311546830028/6373618311546830028_signature_28.png | Rhea Lim-Macion, RMT,MSMLS | https://www.jotform.com/uploads/macarewecare/222927978372067/6373618311546830028/6373618311546830028_signature_31.png
| 2025-10-27 | 2025-10-27 | Tagum City Health Office | Poblacion, Tagum | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31639 | DOH | cristiarafernandez@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3143 | 2025-10-01 | GX-IV 6C 4D | 838553 | UPS not working | Replaced the Defective UPS SN: AT1000190402ST0483 with new UPS SN: AT1000180910ST0672 Check for line connectivity, reset connection | Perform UPS and UPS Troubleshooting | UPS Still not working. Documented the new UPS reported to office for assessment |  |  | DR 13372 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6373554177912854780/6373554177912854780_signature_28.png | Cristina Tiara A. Fernande,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6373554177912854780/6373554177912854780_signature_31.png
| 2025-10-27 | 2025-10-27 | Philippine Red Cross - Ormoc | Ormoc City | 8 | Leyte | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Paul Labutap |  | johnpaullabutap@gmail.com | ormocbloodservice@redcross.org.ph |  | 31638 | COMMERCIAL | cochingvenus@gmail.com | 28 |  |  | MCS+ 9000 | 15H043G | Due for Calibration | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6373472212217664828/6373472212217664828_signature_28.png | Venus D.Coching | https://www.jotform.com/uploads/macarewecare/222927978372067/6373472212217664828/6373472212217664828_signature_31.png
| 2025-10-27 | 2025-10-24 | MMC Medilinx | Amorsolo st., Makati | NCR | NCR 4th - Makati | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31637 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 188 ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 511 ASSAY: Xpert HBV VL, COUNT: 93 ASSAY: Xpert HIV-1 Viral Load, COUNT: 5 ASSAY: Xpert HPV, COUNT: 293 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2444 ASSAY: Others, COUNT: 14 ASSAY: Xpert HCV VL, COUNT: 132 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1215 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 12 |  | GX16D 10C | 110026448 | As per end user they encounter error 2014 A1 module | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Restart system Gathered reports Test run unit | Maintenance and troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6373380103617508945/6373380103617508945_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6373380103617508945/6373380103617508945_signature_31.png
| 2025-10-27 | 2025-10-24 | MMC Medilinx | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31636 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 21 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 119 ASSAY: Xpert CT/NG, COUNT: 329 ASSAY: Xpert HBV VL, COUNT: 48 ASSAY: Xpert HPV, COUNT: 251 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1680 ASSAY: Others, COUNT: 5 ASSAY: Xpert HCV VL, COUNT: 69 ASSAY: Xpert HIV-1 Viral Load, COUNT: 974 |  | GX4D 10C | 110001647 | Due for maintenance B2,B3 failed xpertcheck reported already for module replacement B1 reported already for module replacement. | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | B1,B2,B3 reported for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6373379443614780057/6373379443614780057_signature_28.png | Aiza G. Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6373379443614780057/6373379443614780057_signature_31.png
| 2025-10-24 | 2025-10-23 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | efms.biomed@rmc.doh.gov.ph |  | 31635 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 24 | ASSAY: Xpert HPV, COUNT: 58 ASSAY: Xpert MTB/RIF Ultra, COUNT: 178 ASSAY: Xpert SARS-CoV-2, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 61 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 445 | 2025-10-17 | GX-XVI-D-6C | 838627 | Module A1/220037444 for replacement | Replaced module A1 | Replaced module A1 Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Modules A2/860940, B3/702534, D4/661946, not operational, reported already. Modules C4/730100 error 4017, A4/647085 and B1/701496 not detected, for reporting. |  | Old: A1/220037444 New: A1/220038018 | DR# 13421 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6370810433614758546/6370810433614758546_signature_28.png | Justine Lorenz Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/6370810433614758546/6370810433614758546_signature_31.png
| 2025-10-24 | 2025-10-24 | Cauayan City Health Office | Cauayan City, Isabela | 2 | Isabela | Microbiology Room | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31634 | UNOPS | cho1lab14@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1821 |  | GXIV-10C | 110023504 | Due for Maintenance | Performed Maintenance Cleaned Icore Cleaned Cartridge Bay Cleaned Plunger Rod | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6370805680771130414/6370805680771130414_signature_28.png | Stephen P. Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6370805680771130414/6370805680771130414_signature_31.png
| 2025-10-23 | 2025-10-23 | SACCL | Quricada St., Sta. Cruz Manila | NCR | NCR 1st - Manila | HIV Training Center | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31633 | MACARE-HBDC | patriciaantonette.pomar@slh.doh.gov.ph | 25 |  |  | GXIV | 821257 | For checking of UPS. | Troubleshooting. | Performed checking and troubleshooting of UPS. Tried on different power outlets. | Unit is operational. UPS is not working and also not charging and is for replacement. AT1000 (sn: AT1000211227ST0139). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6370120531951045146/6370120531951045146_signature_28.png | Patricia Antonette Pomar | https://www.jotform.com/uploads/macarewecare/222927978372067/6370120531951045146/6370120531951045146_signature_31.png
| 2025-10-23 | 2025-10-23 | SACCL | Quiricada St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | HIV Training Center | NO | Pullout from demo | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31632 | MACARE-HBDC | norwingayle.rmt@gmail.com | 25 |  |  | GXIV | 839147 | For pullout of GeneXpert system (demo). | Pullout | Performed pullout of GeneXpert system. Checked all items included on the delivery receipt. Generated reports. | Unit is operational and in good condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6370119651956456730/6370119651956456730_signature_28.png | NORWIN GAYLE D MARTIN | https://www.jotform.com/uploads/macarewecare/222927978372067/6370119651956456730/6370119651956456730_signature_31.png
| 2025-10-23 | 2025-10-23 | Bolinao Rural Health Unit | Brgy. Germinal, Bolinao, Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31631 | DOH | nylemae14@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 793 | 2025-10-15 | GXIV-D | 838278 | Replacement of back-up CPU | Back-up CPU replacement Preventive maintenance | Performed surface cleaning Performed preventive maintenance Archived data and transfer to new CPU Setup of new CPU Performed functionality test | Unit is operational | Back up CPU:2UA5201PHX | LENOVO CPU SN:CPH0520F08323 | 13407 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6370094360711365821/6370094360711365821_signature_28.png | Alma sicat | https://www.jotform.com/uploads/macarewecare/222927978372067/6370094360711365821/6370094360711365821_signature_31.png
| 2025-10-23 | 2025-10-23 | Philippine Red Cross Zamboanga | Rizal street, Pettit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE | HAEMONETICS | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31630 | COMMERCIAL | zamboangacity@redcross.org.ph | 23 |  | 2025-10-23 | MCS + | 01M256 | Anticoagulant pump rotor rubber destroyed. | Replaced anticoagulant rotor rubber replacement and Preventive Maintenence. Used 2 rotor rubbers. Available rotor rubber on hand 2pcs. | Replaced anticoagulant rotor rubbers and Preventive Maintenence. Undergone cycle test using diagnostic card. Checked the other 2 rotor for prevention. All rotor rubber from the three rotors are perfectly fine. | Haemonetics machine ready to use. | Used 2 rotor rubbers. Available rotor rubbers on hand 2pcs. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6370078756023187878/6370078756023187878_signature_28.png | Nimfa A.Linao | https://www.jotform.com/uploads/macarewecare/222927978372067/6370078756023187878/6370078756023187878_signature_31.png
| 2025-10-23 | 2025-10-23 | Philippine Red Cross - Ormoc | Ormoc City | 8 | Leyte | Blood Bank | NO | MAINTENANCE | HAEMONETICS | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31629 | COMMERCIAL | shaina.rico@redcross.org.ph | 24 |  | 2025-10-23 | REF 9000 | 15H043G | n/a | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6370026299684093535/6370026299684093535_signature_28.png | shaina mae rico,rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/6370026299684093535/6370026299684093535_signature_31.png
| 2025-10-23 | 2025-10-23 | Tuao Rural Health Unit | Centro 2, Tuao, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31628 | DOH | annrollaine@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 762 |  | GXIV-6C | 838465 | Module 3 and 4 Errors | Performed Module Replacement B3(746175 - 799537) B4(861441 - 680003) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6370003184116359747/6370003184116359747_signature_28.png | Cling Pamittan | https://www.jotform.com/uploads/macarewecare/222927978372067/6370003184116359747/6370003184116359747_signature_31.png
| 2025-10-23 | 2025-10-23 | Laguindingan Rural Health Unit | Zone 10 Poblacion, Laguindingan | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31627 | DOH | chellesambaan@gmail.com | 28 |  |  | GX-4 6 Colors | 838575 | UPS won't work. Doesn't provide power. | UPS Battery Replacement | Replacing UPS Battery and Preventive Maintenance | All units and modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6369999344115715674/6369999344115715674_signature_28.png | Rochelle Sambaan | https://www.jotform.com/uploads/macarewecare/222927978372067/6369999344115715674/6369999344115715674_signature_31.png
| 2025-10-22 | 2025-10-22 | Sacred Heart Medical Center | McArthur Highway, Sto. Domingo, Angeles City | 3 | Pampanga | Blood Bank | YES | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Mark Joseph Venzon |  | mark10venzon@gmail.com | laboratory@acsacredheart.com |  | 31626 | MACARE-HBDC | bloodbank@acsacredheart.com | 25 |  | 2025-10-21 | MCS+ | 09D157 | Roller for replacement Due for calibration | Replacement of roller Calibration | Performed surface cleaning Performed replacement of roller Checked and configure parameters Performed functionality test | Machine is fully operational UPS for Replacement ( atleast 1Kva) |  | MCS+ Urethane Roller Assy | 101246 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6369422937585470523/6369422937585470523_signature_28.png | Maria Cecilia Tinio | https://www.jotform.com/uploads/macarewecare/222927978372067/6369422937585470523/6369422937585470523_signature_31.png
| 2025-10-22 | 2025-10-22 | Pili Rural Health Unit | Poblacion, Pili | 5 | Camarines Sur | Laboratory | YES | REPAIR | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31625 | DOH | lfsanao724@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1499 |  | GX4-6C | 838408 | CPU not powering on B1/688686 temperature Error | Replaced module B1 - 688686>>630528 Repaired CPU(resolved) | Replaced B1/688686 Repaired CPU | Unit is operational Module B1, B2, B3 operational CPU operational | Module B4/662296 pending for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6369293961716681643/6369293961716681643_signature_28.png | Liezl Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/6369293961716681643/6369293961716681643_signature_31.png
| 2025-10-22 | 2025-10-22 | Makilala Rural Health Unit | Poblacion, Makilala | 12 | Cotabato | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31624 | UNOPS | abegailweill17@gmail.com | 23.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 505 ASSAY: Xpert MTB-XDR, COUNT: 2 |  | GX-IV 6C 10D | 110023521 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6369190588443622180/6369190588443622180_signature_28.png | Shayne s. Belandres | https://www.jotform.com/uploads/macarewecare/222927978372067/6369190588443622180/6369190588443622180_signature_31.png
| 2025-10-22 | 2025-10-22 | Pamplona Rural Health Unit | Pamplona, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31623 | UNOPS | wickizena@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2549 |  | GXIV-10C | 110023503 | Due for Maintenance | Performed Maintenance Cleaned Icore Cleaned Cartridge Bay Cleaned Plunger Rod Cleaned Filter | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6369140199323085714/6369140199323085714_signature_28.png | Wicki Zena Gerona | https://www.jotform.com/uploads/macarewecare/222927978372067/6369140199323085714/6369140199323085714_signature_31.png
| 2025-10-22 | 2025-10-22 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31622 | COMMERCIAL | cchdserology@gmail.com | 26 |  |  | Gx 4 | 829419 | Defective module A3 | Collecting data IQR, system log and assay stats | Collecting data IQR, system log and assay stats | Defective module 2 and 3 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6369120668121696019/6369120668121696019_signature_28.png | Owen Joshua L. Briones | https://www.jotform.com/uploads/macarewecare/222927978372067/6369120668121696019/6369120668121696019_signature_31.png
| 2025-10-22 | 2025-10-22 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31621 | UNOPS | laureanjo@yahoo.com | 25 |  |  | Gx 4 | 110023513 | N/A | Preventive Maintenance Clean Module Bay Gather data IQR, system log and assys statistics | Preventive Maintenance Clean Module Bay Gather data IQR, system log and assys statistics | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6369109338124212931/6369109338124212931_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6369109338124212931/6369109338124212931_signature_31.png
| 2025-10-22 | 2025-10-16 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31620 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational. Module A1 is Disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6369074557182155335/6369074557182155335_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6369074557182155335/6369074557182155335_signature_31.png
| 2025-10-21 | 2025-10-21 | St. Luke's Medical Center | E. Rodriguez Ave., Quezon City | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed.qc@stlukes.com.ph |  | 31619 | COMMERCIAL | megapo@stlukes.com.ph | 25 |  |  | GXIV | 839234 | Unit for preventive maintenance. For generating of reports. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived and generated test results. Generated reports. | Unit is operational. Module B2 has 1 error 5007, but after running today, a result was given with reading. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6368370626918322570/6368370626918322570_signature_28.png | liza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6368370626918322570/6368370626918322570_signature_31.png
| 2025-10-21 | 2025-10-21 | Bagong Ilog Health Center I | Sgt. Pascua St, Bagong Ilog, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31618 | MACARE-HBDC | agasangevelyn7@gmail.com | 24 |  |  | GX-IV-D-10C | 110026461 | For installation | Installed the GeneXpert system | Installed the GeneXpert System: -GeneXpert Unit -Workstation -UPS -Barcode Scanner with stand -Printer Generated IQR | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6368283319515877320/6368283319515877320_signature_28.png | Evelyn Agasang | https://www.jotform.com/uploads/macarewecare/222927978372067/6368283319515877320/6368283319515877320_signature_31.png
| 2025-10-21 | 2025-10-16 | Lubang Rural Health Unit | Lubang, Occidental Mindoro | 4B | Occidental Mindoro | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31617 | DOH | lalaine.batalla0821@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 232 |  | GX-XVI-D-6C | 838228 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -Plunger rod -Icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6368272338154769603/6368272338154769603_signature_28.png | Lalaine V. Batalla | https://www.jotform.com/uploads/macarewecare/222927978372067/6368272338154769603/6368272338154769603_signature_31.png
| 2025-10-20 | 2025-10-15 | San Fernando Rural Health Unit | Poblacion, San Fernando | 5 | Camarines Sur | Laboratory | YES | Endorsement letter | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31616 | DOH | josephus_garfin@yahoo.com | 30 |  |  | GX4-6C | 838496 | Endorsement letter | Provide endorsement letter for transition of service | Provide endorsement letter | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6367786733319449411/6367786733319449411_signature_28.png | Josephus Garfin | https://www.jotform.com/uploads/macarewecare/222927978372067/6367786733319449411/6367786733319449411_signature_31.png
| 2025-10-18 | 2025-10-17 | Cardinal Santos Medical Center (Medilinx) | Wilson St., San Juan | NCR | NCR 2nd - San Juan | Microbiology | NO | CALIBRATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31615 | COMMERCIAL | renalyn.cabaltea@medilinx.com.ph | 20 | ASSAY: Xpert HBV VL, COUNT: 45 ASSAY: Xpert HPV, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1164 |  | GX4D (10C) | 110003714 | Due for xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck | B1,B2,B3,failed in xpertcheck for replacement B4 retry lack of kits |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6365592428668107628/6365592428668107628_signature_28.png | Renalyn Cabaltea | https://www.jotform.com/uploads/macarewecare/222927978372067/6365592428668107628/6365592428668107628_signature_31.png
| 2025-10-17 | 2025-10-17 | Medical Center Manila | U.N. Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31614 | COMMERCIAL | microbiology.manilamed@gmail.com | 25 |  |  | GXIV | 835540 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6364943935547912770/6364943935547912770_signature_28.png | FRANCES B. VELARDE | https://www.jotform.com/uploads/macarewecare/222927978372067/6364943935547912770/6364943935547912770_signature_31.png
| 2025-10-17 | 2025-10-16 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | INSTALLATION | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31613 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  | 2025-10-15 | MCS + | 13B243 | For back up unit and installation | Installed unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed installation | Unit operational |  | MCS + and transport case | 13409 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6364940350317577424/6364940350317577424_signature_28.png | Philippine Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6364940350317577424/6364940350317577424_signature_31.png
| 2025-10-17 | 2025-10-16 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31611 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 16K521SPG | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6364936135545086016/6364936135545086016_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6364936135545086016/6364936135545086016_signature_31.png
| 2025-10-17 | 2025-10-16 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31612 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. | Preventive maintenance. | Performed calibration. Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6364936145548780508/6364936145548780508_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6364936145548780508/6364936145548780508_signature_31.png
| 2025-10-17 | 2025-10-17 | Pasig Treatment Hub | Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31610 | UNOPS | pathmedtech@gmail.com | 28 |  | 2025-10-09 | GXIV-10C | 110025127 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: passed B2: passed B3: failed B4: passed |  | XpertCheck Kit | 13389 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6364867936919823273/6364867936919823273_signature_28.png | Jhozza Rei M. Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/6364867936919823273/6364867936919823273_signature_31.png
| 2025-10-17 | 2025-10-15 | Malungon Rural Health Unit | Malungon, Sarangani | 12 | Sarangani | Tb-Lab | NO | INSTALLATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31608 | DOH | madameerrmie@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1239 |  | GxIV-6C | 838546 | CPU problem | Perform Genexpert temporary CPU installation Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking module status and errors Data collection | Installed and set up temporary CPU Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6364667018421021347/6364667018421021347_signature_28.png | KOBE SHANE L. BACULIO | https://www.jotform.com/uploads/macarewecare/222927978372067/6364667018421021347/6364667018421021347_signature_31.png
| 2025-10-17 | 2025-10-16 | Calumpang Rural Health Unit | Brgy. Calumpang General Santos City | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31609 | DOH | jan19.cam@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1520 |  | GxIV-6C | 838504 | None | Perform Genexpert maintenance and cleaning Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6364667018423042970/6364667018423042970_signature_28.png | Margareth Zaldivar | https://www.jotform.com/uploads/macarewecare/222927978372067/6364667018423042970/6364667018423042970_signature_31.png
| 2025-10-16 | 2025-10-16 | National Center for Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31607 | UNOPS | iancristrono@gmail.com | 28 |  | 2025-10-09 | GXIV-10C | 120003480 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed | w/ Jesh | XpertCheck Kit | 13391 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6364021400751592170/6364021400751592170_signature_28.png | Ian Cris Trono | https://www.jotform.com/uploads/macarewecare/222927978372067/6364021400751592170/6364021400751592170_signature_31.png
| 2025-10-16 | 2025-10-16 | Mandaluyong Social Hygiene Clinic | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31606 | CONSORTIUM | mandaluyongshc.laboratory@gmail.com | 25 |  | 2025-10-09 | GXIV-10C | 110025111 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: B2: B3: B4: | w/ Jesh | XpertCheck Kit | 13393 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363948431752980170/6363948431752980170_signature_28.png | Ron Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/6363948431752980170/6363948431752980170_signature_31.png
| 2025-10-16 | 2025-10-16 | San Lazaro Hospital | Quiricada St., Sta.Cruz, Manila | NCR | NCR 1st - Manila | SACCL | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31605 | MACARE-HBDC | melanie.perez@slh.doh.gov.ph | 26.5 | ASSAY: Xpert HBV VL, COUNT: 127 ASSAY: Xpert HIV-1 Viral Load, COUNT: 64 |  | GXIV-L-6C | 821257 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Checked Printer and UPS Functionality. Performed Self-test. Generate Reports. | Modules 1,2 and 4 are Operational. Module 3 is disabled ( with 4017 error). Black Cartridge of Printer for Replacement. UPS is Good (with Backup). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363923933634063574/6363923933634063574_signature_28.png | Melanie S. Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/6363923933634063574/6363923933634063574_signature_31.png
| 2025-10-15 | 2025-10-15 | Medi Linx Laboratory Inc. (CLDH) | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31604 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 |  | GXII-D | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module C2 | C2:620096 for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363189560817397262/6363189560817397262_signature_28.png | Alyssa Sytu | https://www.jotform.com/uploads/macarewecare/222927978372067/6363189560817397262/6363189560817397262_signature_31.png
| 2025-10-15 | 2025-10-15 | Medi Linx Laboratory Inc. (CLDH) | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31603 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 |  | GXIV-D | 839232 | Due for Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module B4 | B4: 738203 for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363188940813558963/6363188940813558963_signature_28.png | Alyssa Sytu | https://www.jotform.com/uploads/macarewecare/222927978372067/6363188940813558963/6363188940813558963_signature_31.png
| 2025-10-15 | 2025-10-15 | San Isidro Rural Health Unit | San Isidro Northern Samar | 8 | Northern Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31602 | UNOPS | monsampayan24@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 361 | 2025-10-15 | GXIV | 110023515 | n/a | Data Collection | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363061006127293262/6363061006127293262_signature_28.png | Mon Edcelle Q. Sampayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6363061006127293262/6363061006127293262_signature_31.png
| 2025-10-15 | 2025-10-08 | Sta. Maria Rural Health Unit | Sta. Maria, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31601 | UNOPS | jomarcaberto24@gmail.com | 30 |  |  | GXIV-4-D-6C | 110023500 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363016776156601102/6363016776156601102_signature_28.png | Jomar Caberto | https://www.jotform.com/uploads/macarewecare/222927978372067/6363016776156601102/6363016776156601102_signature_31.png
| 2025-10-15 | 2025-10-08 | Sinait District Hospital | Sinait,Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31600 | DOH | isdhsinaitlaboratory@gmail.com | 30 |  |  | GXIV-4-D-6C | 838277 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363015226153976570/6363015226153976570_signature_28.png | Claire Marie Baguyo-Diculen | https://www.jotform.com/uploads/macarewecare/222927978372067/6363015226153976570/6363015226153976570_signature_31.png
| 2025-10-15 | 2025-10-08 | Laoag City Health Office | Laoag City, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31599 | DOH | denver.mahor@gmail.com | 30 |  |  | GXIV-4-D-6C | 838416 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363013516155874983/6363013516155874983_signature_28.png | MARY STEPHANIE R. SUGUITAN | https://www.jotform.com/uploads/macarewecare/222927978372067/6363013516155874983/6363013516155874983_signature_31.png
| 2025-10-15 | 2025-10-07 | Candon City Health Office | Candon,Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31598 | DOH | candoncholaboratory@gmail.com | 30 |  |  | GXIV-4-D-6C | 838273 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Partially Operational Module 4 for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363012586159995182/6363012586159995182_signature_28.png | Michelle Bucalen | https://www.jotform.com/uploads/macarewecare/222927978372067/6363012586159995182/6363012586159995182_signature_31.png
| 2025-10-15 | 2025-10-07 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | La Union | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31597 | DOH | piaramuj@gmail.com | 29 |  |  | GXIV-4-D-6C | 838270 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363011616157198558/6363011616157198558_signature_28.png | Pia Jumar T. Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/6363011616157198558/6363011616157198558_signature_31.png
| 2025-10-15 | 2025-10-07 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31596 | DOH | seansimonpaloyo@gmail.com | 28 |  | 2025-09-26 | GXIV-4-D-6C | 838473 | For UPS Replacement | #NAME? | Replaced Defective UPS | Unit is Fully Operational |  |  | DR 13360 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6363010316152765933/6363010316152765933_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6363010316152765933/6363010316152765933_signature_31.png
| 2025-10-14 | 2025-10-14 | Biopath Clinical Diagnostics | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | OTHERS | Jovy Noel |  | jdnoel.eisi@gmail.com |  | Esco | 31595 | COMMERCIAL | BiopathCAP@biopath.ph | 23 |  |  | PCR-3A3 | 2020-154133 | Due for preventive maintenance. | Save file. | #NAME? | PCR functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6362309104213196178/6362309104213196178_signature_28.png | Mary Joy Lacsima | https://www.jotform.com/uploads/macarewecare/222927978372067/6362309104213196178/6362309104213196178_signature_31.png
| 2025-10-14 | 2025-06-06 | Biopath Clinical Diagnostics | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | OTHERS | Jovy Noel |  | jdnoel.eisi@gmail.com |  | Esco | 31594 | COMMERCIAL | BiopathCAP@biopath.ph | 22 |  |  | PCR-3A3 | 2020-154133 | Due for Preventive Maintenance. | #NAME? | Due for Preventive Maintenance. | Functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6362305944212605318/6362305944212605318_signature_28.png | Mary Joy Lacsima | https://www.jotform.com/uploads/macarewecare/222927978372067/6362305944212605318/6362305944212605318_signature_31.png
| 2025-10-14 | 2025-03-07 | Biopath Clinical Diagnostics | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | OTHERS | Jovy Noel |  | jdnoel.eisi@gmail.com |  | Esco | 31593 | COMMERCIAL | BiopathCAP@biopath.ph | 23 |  |  | PCR-3A3 | 2020-154133 | Due for Quarterly Maintenance | Save files. | #NAME? | PCR Cabinet are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6362298424217074265/6362298424217074265_signature_28.png | Mary Joy Lacsima | https://www.jotform.com/uploads/macarewecare/222927978372067/6362298424217074265/6362298424217074265_signature_31.png
| 2025-10-14 | 2025-10-10 | Pio Duran Rural Health Unit | Caratagan, Pio Duran | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31592 | DOH | mitzijc@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 456 |  | GX4-6C | 838264 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6362290177006725488/6362290177006725488_signature_28.png | Karen Pepan̈o Nieva | https://www.jotform.com/uploads/macarewecare/222927978372067/6362290177006725488/6362290177006725488_signature_31.png
| 2025-10-14 | 2025-10-09 | Iriga City Health Office | Sta. Cruz Sur, Iriga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31591 | DOH | mongealyzza@yahoo.com.ph | 25 |  |  | GX4-6C | 838391 | GX4-6C for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data Provide endorsement letter | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6362289827009546402/6362289827009546402_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/6362289827009546402/6362289827009546402_signature_31.png
| 2025-10-14 | 2025-10-09 | Iriga City City Health Office | Sta. Cruz Sur | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31590 | UNOPS | mongealyzza@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1461 ASSAY: Xpert MTB-XDR, COUNT: 58 |  | GX4-10C | 110023511 | GX4-10C for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6362289017007990593/6362289017007990593_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/6362289017007990593/6362289017007990593_signature_31.png
| 2025-10-14 | 2025-10-14 | Pasig Treatment Hub | Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31589 | MACARE-HBDC | marcabapo@gmail.com | 28 |  | 2025-10-09 | GXIV-10C | 110025128 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed the calibration |  | XpertCheck Kit | 13389 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6362186372515860759/6362186372515860759_signature_28.png | Marc francis Abapo | https://www.jotform.com/uploads/macarewecare/222927978372067/6362186372515860759/6362186372515860759_signature_31.png
| 2025-10-14 | 2025-10-14 | Bolinao RHU | Bolinao Pangasinan | 1 | Pangasinan | Lab | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 31588 | DOH | florescanoel77@gmail.com | 26 |  |  | GX IV D | 838278 | Defective workstation | Replaced and installed mother board,SSD,PSU Performed functionality test | Performed repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6362160213619174550/6362160213619174550_signature_28.png | In house repair | https://www.jotform.com/uploads/macarewecare/222927978372067/6362160213619174550/6362160213619174550_signature_31.png
| 2025-10-13 | 2025-10-09 | Polomolok Rural Health Unit | Polomolok, South Cotabato | 12 | South Cotabato | Tb-lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31587 | DOH | jmmirasolrmt@gmail.com | 24.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2170 |  | GxlV-6C | 838265 | None | Perform Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of modules status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6361687515321148638/6361687515321148638_signature_28.png | June Michele Mirasol | https://www.jotform.com/uploads/macarewecare/222927978372067/6361687515321148638/6361687515321148638_signature_31.png
| 2025-10-13 | 2025-10-01 | Capitol Medical Center | Scout Magbanua Street, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31586 | COMMERCIAL | princeosena@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 576 | 2025-09-30 | GXIV-D-10C | 110024553 | Unit is due for XpertCheck. Unit is due for Preventive Maintenance. | Preventive Maintenance. XpertCheck. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | XpertCheck Result: B1 - Pass B2- Pass B3- Pass B4- Service Required | B4/210128226 for Module Replacement | XpertCheck Kit | SI# 101121 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6361472343613183591/6361472343613183591_signature_28.png | Prince Henry Oseña | https://www.jotform.com/uploads/macarewecare/222927978372067/6361472343613183591/6361472343613183591_signature_31.png
| 2025-10-13 | 2025-10-13 | Jose Abad Santos Rural Health Unit | Poblacion, Jose Abad Santos | 11 | Davao Occidental | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31585 | DOH | alcalend47777@gmail.com | 27.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 468 |  | GX-IV 6C 4D | 838460 | For Preventive Maintenance UPS not working System Unit Not working | Change UPS Batteries Change Cmos Battery, reset network connection Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance System Unit Troubleshooting UPS Battery Replacement | All Modules are operational Successfully replaced UPS Batteries System Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6361289763023557581/6361289763023557581_signature_28.png | Dranreb John T. Alcalen | https://www.jotform.com/uploads/macarewecare/222927978372067/6361289763023557581/6361289763023557581_signature_31.png
| 2025-10-12 | 2025-10-12 | Philippine Heart Center | Matalino St., Quezon City | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31584 | COMMERCIAL | phc_bb@yahoo.com | 22 |  |  | MCS+ | 11H204 | No display | Troubleshoot the unit | Checked connections Reseat display cable and CPU board Check display panel | Unit is not operational. For possible display cable replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6360625409301302160/6360625409301302160_signature_28.png | Cathlyn Pua | https://www.jotform.com/uploads/macarewecare/222927978372067/6360625409301302160/6360625409301302160_signature_31.png
| 2025-10-10 | 2025-10-10 | Jolo sanitarium and General Hospital | Upper san Raymondo, sulu Sanitarium | BARMM | Sulu | Laboratory | NO | MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31583 | DOH | shirleymedtech22@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2025-10-10 | GX 4 | 110015481 | None | Preventive Maintenence and data collection. | Preventive Maintenence and data collection. | Preventive Maintenence and data collection. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358958462119414432/6358958462119414432_signature_28.png | Shirley y. Julkanain | https://www.jotform.com/uploads/macarewecare/222927978372067/6358958462119414432/6358958462119414432_signature_31.png
| 2025-10-10 | 2025-10-10 | Rizal Medical Center | Pasig Blvd. Pasig City | NCR | NCR 2nd - Pasig | Molecular Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | princeszvillaran13@gmail.com |  | 31582 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 | ASSAY: Xpert HPV, COUNT: 57 ASSAY: Xpert MTB/RIF Ultra, COUNT: 179 ASSAY: Xpert SARS-CoV-2, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 55 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 445 | 2025-10-09 | GX-XVI-D-6C | 838627 | Network switch power supply replacement | Replaced network switch power supply | Performed maintenance Replaced network switch power supply Functionality test Checked other modules | Unit is operational. Module A1/220037444 not detected ,for replacement (under warranty). Module A2/860940, B3/702534, D4/661946 reported already for replacement. Module C4/730100 error 4017 for reporting. Module A4/647085 not detected, for reporting. Module B1/701496 not detected, for reporting. UPS has no backup power. Recommend to replace. |  | Power supply. | SI# 101188, DR#13394 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358926188743300759/6358926188743300759_signature_28.png | Benita Gracia N. Viloria, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6358926188743300759/6358926188743300759_signature_31.png
| 2025-10-10 | 2025-10-10 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | princeszvillaran13@gmail.com |  | 31581 | COMMERCIAL | rmcbloodbank@gmail.com | 22 |  |  | MCS+ | 15H045G | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358926188746584267/6358926188746584267_signature_28.png | Ferdinand B. Barroga, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6358926188746584267/6358926188746584267_signature_31.png
| 2025-10-10 | 2025-10-10 | Lung Center Of The Philippines | Quezon Ave., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | YES | REPAIR TROUBLESHOOTING | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 31580 | MACARE-HBDC | lcpbloodbank@gmail.com | 25 |  | 2025-10-08 | MCS + | 02D166 | For troibleshooting due to display ang initialization issue.. UPS unable to provide back up power. | Performed troubleshooting using good CPU and Battery for UPS. System initialization was completed after the swapping of CPU.. UPS provide back up after swapping of battery. | Swapped CPU. Swapped UPS battery. | Unit not operational.. Recommended for CPU replacement. UPS recommended for battery replacement. | Pulled out the CPU and battery. | CPU and Battery for troubleshooting and problem issolation. | DR NUMBER: 13384 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358909580411465942/6358909580411465942_signature_28.png | Janella Madulid | https://www.jotform.com/uploads/macarewecare/222927978372067/6358909580411465942/6358909580411465942_signature_31.png
| 2025-10-10 | 2025-10-10 | St. Lukes Medical Center Extension Clinic | Bocobo Street Ermita Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.con |  |  | 31579 | MACARE-HBDC | laboratory@slec.ph | 22 |  | 2025-10-01 | GeneXpert Infinity | 809643 | Modules A3 and B2 for replacement | Replaced modules A3 (754001 > 698176)and B2 (702345 > 866654) | Replaced defective modules | Unit is operational |  | Genexpert Module, Assembly door, Modified ejector for Genexpert Infinity (x2) | DR 13370 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358894097312664801/6358894097312664801_signature_28.png | Jasmine Rule | https://www.jotform.com/uploads/macarewecare/222927978372067/6358894097312664801/6358894097312664801_signature_31.png
| 2025-10-10 | 2025-10-10 | Maimbung District Hospital | Maimbung sulu | BARMM | Sulu | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31578 | DOH | mdhclinlab23@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 458 | 2025-10-10 | GX 4 | 838838 | None | Preventive Maintenence and Data collection. | Preventive Maintenence and data collection | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358848428117679374/6358848428117679374_signature_28.png | Amerkhan d adji, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6358848428117679374/6358848428117679374_signature_31.png
| 2025-10-10 | 2025-10-10 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31577 | DOH | ccmclaboratory01@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2358 |  | 6C GXIV-4D | 838561 | Due for PM | -Perform preventive maintenance; plunger rod, cartridge bay, fan filter, optic slot | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358717415027174913/6358717415027174913_signature_28.png | Hannah Jane Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/6358717415027174913/6358717415027174913_signature_31.png
| 2025-10-10 | 2025-10-09 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | gsd-biomed@nkti.gov.ph |  | 31576 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1725 ASSAY: Xpert CT/NG, COUNT: 24 ASSAY: Others, COUNT: 115 ASSAY: Xpert HBV VL, COUNT: 914 ASSAY: Xpert HPV, COUNT: 127 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2683 ASSAY: Xpert SARS-CoV-2, COUNT: 840 ASSAY: Xpert HCV VL, COUNT: 169 | 2025-10-09 | GX16D | 110005587 | Due for Maintenance and Xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and xpertcheck | A1,A2,C2,D2 reported already for replacement A3,B3 failed XC last 9/30/24 reported for replacement A4,B2,C1,C4,D1,D4 Passed in Xpertcheck B1,B4,C3,D3 failed in Xpertcheck |  | Xpertcheck kits | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358688483613599721/6358688483613599721_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6358688483613599721/6358688483613599721_signature_31.png
| 2025-10-10 | 2025-10-09 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | gsd-biomed@nkti.gov.ph |  | 31575 | MACARE-HBDC | molecularlabnkti@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 84 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Others, COUNT: 8 ASSAY: Xpert HBV VL, COUNT: 49 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 93 ASSAY: Xpert SARS-CoV-2, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 11 | 2025-10-09 | GX4D | 838524 | Due for maintenance and xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Restart worstation and genexpert Gathered reports Test run unit | Performed maintenance and xpertcheck | All module passed in xpert check |  | Xpertcheck kits. | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358684393613025668/6358684393613025668_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6358684393613025668/6358684393613025668_signature_31.png
| 2025-10-09 | 2025-10-09 | Marantao Rural Health Unit | Bandara-ingud, Camalig Marantao Lanao del Sur | 8 | Lanao del Sur | TB Laboratory | YES | INSTALLATION TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 31574 | DOH | sorayajmmuti@gmail.com | 24 |  |  | GXIV-4 | 847303 | Monitor problem vertical line with serial number CNK94301W2. | Cannot save installation report and other's due to defective CPU problem no power ON. | Replace old monitor with new monitor with serial number H22V8VL658A250716069 | New monitor are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358188910636125748/6358188910636125748_signature_28.png | Soraya Jamila M. Muti R.M.T., A.M.T | https://www.jotform.com/uploads/macarewecare/222927978372067/6358188910636125748/6358188910636125748_signature_31.png
| 2025-10-09 | 2025-10-09 | University of the East Ramon Magsaysay Memorial Medical Center, Inc. | Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com | biomed@uerm.edu.ph |  | 31573 | MACARE-HBDC | uermbloodbank@gmail.com | 28 |  | 2025-10-09 | MCS+ | 13C270 | Unit for maintenance Unit for calibration | -> set calibration parameters to accepted value range -> replaced rotor motor pump rubbers | -> replaced roller motor pump rubbers -> checked calibration parameters (set bowl optics offset value and whole scale value; set DPM and SPM whole scale values) -> performed self test | unit is operational | w/ Jesh | MCS+ Urethane Roller Assy | 13385 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358039528176179345/6358039528176179345_signature_28.png | Maxyme Erika B. Desiderio | https://www.jotform.com/uploads/macarewecare/222927978372067/6358039528176179345/6358039528176179345_signature_31.png
| 2025-10-09 | 2025-10-09 | Hi Precision Diagnostics | D. Tuazon St. Quezon City | NCR | NCR 2nd - Quezon City | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31572 | COMMERCIAL | jose@hi-precision.com.ph | 25 |  |  | GXIV | 835602 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. 2 modules, A2(692016) and A3 (693796) are service required based on the Xpertcheck result by Cepheid PH last June 2025. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358025374342688540/6358025374342688540_signature_28.png | Joselito T. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/6358025374342688540/6358025374342688540_signature_31.png
| 2025-10-09 | 2025-10-09 | Hi Precision Diagnostics | D. Tuazon St. Quezon City | NCR | NCR 2nd - Quezon City | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31571 | COMMERCIAL | jose@hi-precision.com.ph | 25 |  |  | GXIV | 835601 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. 2 modules, B2 (787789) and B3 (891796) are service required based on the Xpertcheck result by Cepheid PH last June 2025. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6358025374342880401/6358025374342880401_signature_28.png | Joselito T. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/6358025374342880401/6358025374342880401_signature_31.png
| 2025-10-09 | 2025-10-09 | Iriga City City Health Office | Sta. Cruz Sur | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31570 | UNOPS | mongealyzza@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1461 ASSAY: Xpert MTB-XDR, COUNT: 58 |  | GX4-10C | 110023511 | GX4-10C for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6357985885327718727/6357985885327718727_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/6357985885327718727/6357985885327718727_signature_31.png
| 2025-10-09 | 2025-10-09 | Padada Rural Health Unit | Poblacion | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31569 | DOH | zhakie.marvz@gmail.com | 612 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 612 |  | GX-IV 6C 4D | 838511 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6357968534828128656/6357968534828128656_signature_28.png | Marvielyn Corpin Lusica | https://www.jotform.com/uploads/macarewecare/222927978372067/6357968534828128656/6357968534828128656_signature_31.png
| 2025-10-09 | 2025-10-09 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31568 | DOH | burauenlab@yahoo.com | 26 |  |  | GXIV | 838551 | n/a | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6357951924418027144/6357951924418027144_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6357951924418027144/6357951924418027144_signature_31.png
| 2025-10-08 | 2025-10-08 | Diit District Health Center | Diit, Tacloban Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31567 | DOH | dalepenalosa189@gmail.com | 25 |  |  | GXIV | 838421 | #NAME? | Data Collection | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6357290097426555021/6357290097426555021_signature_28.png | dalepenalosa189@gmail.com | https://www.jotform.com/uploads/macarewecare/222927978372067/6357290097426555021/6357290097426555021_signature_31.png
| 2025-10-08 | 2025-10-08 | Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31566 | DOH | miniforestlaboratory@gmail.com | 23.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 585 |  | GX-IV 6C 4D | 838507 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6357084504826468832/6357084504826468832_signature_28.png | Lolijell Torralba,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6357084504826468832/6357084504826468832_signature_31.png
| 2025-10-07 | 2025-10-02 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31565 | DOH | celinejean16@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1777 |  | GxIV-6C | 838570 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status/problem Data reporting | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6356517485324116264/6356517485324116264_signature_28.png | Celine Jean villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6356517485324116264/6356517485324116264_signature_31.png
| 2025-10-07 | 2025-10-01 | Malungon Rural Health Unit | Malungon, Sarangani | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31564 | DOH | madameerrmie@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2384 |  | GxIV-6C | 838546 | Genexpert computer unit /CPU defective | Perform CPU troubleshooting Perform Genexpert machine Cleaning and maintenance CPU problem reporting | CPU troubleshooting Preventive maintenance | CPU motherboard problem CPU request for replacement due to Motherboard problem Genexpert none operational due to computer/CPU problem |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6356516155321910774/6356516155321910774_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/6356516155321910774/6356516155321910774_signature_31.png
| 2025-10-07 | 2025-09-30 | Tampakan Rural Health Unit | Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31563 | DOH | lotisglegada@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 866 |  | GxIV-6C | 838544 | None | Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status/errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6356515215327388530/6356515215327388530_signature_28.png | Lotis G. Legada | https://www.jotform.com/uploads/macarewecare/222927978372067/6356515215327388530/6356515215327388530_signature_31.png
| 2025-10-07 | 2025-10-07 | MEDI LINX Laboratory Inc. / St. Elizabeth Hospital Inc. | General Santos City | 12 | South Cotabato | Microbiology | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | therese.cariaga@medilinx.com.ph |  | 31562 | COMMERCIAL | marylou.villaruel@medilinx.com.ph | 23.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 50 | 2025-10-07 | GxIV-6C | 831313 | All modules requesting for Module replacement due to H/W failed | Replace defective module A1 SN:620845 with new module SN:220045650 Replace defective module A3 SN:713143 with new module SN:220045603 Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status/problem Data reporting | Module repair Preventive maintenance | - 2 operational modules ( A1 and A3) - Module A2 and A4 requesting for module replacement |  | Genexpert 6-color module SN: 220045650 SN: 220045603 | 13369 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6356498245324387232/6356498245324387232_signature_28.png | Maricris Sales | https://www.jotform.com/uploads/macarewecare/222927978372067/6356498245324387232/6356498245324387232_signature_31.png
| 2025-10-07 | 2025-10-07 | Chinese General Hospital | Blumentrit, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Bloodbank | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 31561 | MACARE-HBDC | cghmcbloodbank2016@gmail.com | 25 |  |  | MCS+ | 08M463 | Fluid detected on centrifuge. | Troubleshooting and cleaning of sensors. | Performed troubleshooting of MCS+. Performed cleaning of fluid sensor. Performed simulation running and diagnostic setting. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6356250985187605282/6356250985187605282_signature_28.png | Jscqueline Zatzoso | https://www.jotform.com/uploads/macarewecare/222927978372067/6356250985187605282/6356250985187605282_signature_31.png
| 2025-10-07 | 2025-10-07 | Maco Rural Health Unit | Poblacion Maco | 11 | Davao de Oro | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31560 | DOH | edlienkai26@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 806 |  | GX-IV 6C 4D | 838500 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6356233753785652837/6356233753785652837_signature_28.png | La perla Belacho | https://www.jotform.com/uploads/macarewecare/222927978372067/6356233753785652837/6356233753785652837_signature_31.png
| 2025-10-07 | 2025-10-07 | Libon Rural Health Unit | Libon, Albay | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31559 | DOH | liorinlirio24@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 620 |  | GX4-6C | 838397 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data Provide letter of transition of service | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6356196513422823625/6356196513422823625_signature_28.png | Lirio llorin | https://www.jotform.com/uploads/macarewecare/222927978372067/6356196513422823625/6356196513422823625_signature_31.png
| 2025-10-07 | 2025-10-07 | TAYASAN RURAL HEALTH UNIT | National Hiway, Tayasan, Negros Oriental | 7 | Negros Oriental | Laboratory | NO | ON-SITE ASSISTANCE Unit transfer to main laboratory building | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 31558 | UNOPS | tayasanrhu2023@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 385 |  | GX4 | 110023502 | Unit for transfer to main laboratory building | . | Assist transfer and installation of Genexpert system and accessories. | All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6356129050813684541/6356129050813684541_signature_28.png | Portia Cabanting RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6356129050813684541/6356129050813684541_signature_31.png
| 2025-10-06 | 2025-10-02 | Naujan Municipal Health Office | Naujan, Oriental Mindoro | 4B | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31557 | DOH | ellainehernandez27@gmail.com | 25 |  |  | GXIV | 838226 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6355462723616033014/6355462723616033014_signature_28.png | Ellaine Hernandez-Villas | https://www.jotform.com/uploads/macarewecare/222927978372067/6355462723616033014/6355462723616033014_signature_31.png
| 2025-10-05 | 2025-10-03 | Metro Health Specialists Hospital Inc. | Sorsogon City | 5 | Sorsogon | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31556 | COMMERCIAL | mhshimolecularpathology@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 7 |  | GX4-10C | 120002716 | Gx for Preventive | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6354607106123042736/6354607106123042736_signature_28.png | Shawn Manalansan | https://www.jotform.com/uploads/macarewecare/222927978372067/6354607106123042736/6354607106123042736_signature_31.png
| 2025-10-05 | 2025-10-01 | Bicol Regional Hospital and Medical Center | Sagpon, Daraga | 5 | Albay | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31555 | MACARE-HBDC | brtth.covid19@gmail.com | 20 |  |  | GX4-10C | 110026767 | B4/210138841 - 4014 error | Restarts Unit Data collection | Performed Troubleshooting | Unit is Operational Modulw B1, B2, B3 Operational | Reported Defective Module, request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6354606816124876261/6354606816124876261_signature_28.png | Jennifer Vitasa | https://www.jotform.com/uploads/macarewecare/222927978372067/6354606816124876261/6354606816124876261_signature_31.png
| 2025-10-05 | 2025-09-30 | Calabanga Rural Health | Poblacion, Calabanga | 5 | Camarines Sur | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31554 | DOH | martyarc11@yahoo.com | 23 |  |  | GX4-6C | 838398 | CPU restarting | Cleaned RAM checked/swapped power cables Resets cable connections from power source in CPU | Performed Troubleshooting | Unit is non operational CPU not powering on |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6354604856128684200/6354604856128684200_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/6354604856128684200/6354604856128684200_signature_31.png
| 2025-10-03 | 2025-10-03 | Our Lady of Mt. Carmel Medical Center | San Fernando City Pampanga | 3 | Pampanga | Blood Bank | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31553 | MACARE-HBDC | olmcmc.bloodbank@gmail.com | 25 |  | 2025-10-02 | MCS+ | 15G043 | Due for calibration Roller for replacement | Preventive maintenance Calibration Repair | Performed surface cleaning Checked parameters Performed Calibration Replacement of rollers Performed functionality test | Unit is operational |  | 6 pump rotor rollers | 101141 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6352754473867298640/6352754473867298640_signature_28.png | Genmerch Manabat | https://www.jotform.com/uploads/macarewecare/222927978372067/6352754473867298640/6352754473867298640_signature_31.png
| 2025-10-03 | 2025-10-03 | Mt. Carmel | San Fernando City Pampanga | 3 | Pampanga | Blood Bank | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31552 | MACARE-HBDC | olmcmc.bloodbank@gmail.com | 25 |  | 2025-10-02 | MCS+ | 15G043 | Due for calibration Roller for replacement | Preventive maintenance Calibration Repair | Performed surface cleaning Checked parameters Performed Calibration Replacement of rollers Performed functionality test | Unit is operational |  | 6 pump rotor rollers | 101141 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6352754453869006730/6352754453869006730_signature_28.png | Genmerch Manabat | https://www.jotform.com/uploads/macarewecare/222927978372067/6352754453869006730/6352754453869006730_signature_31.png
| 2025-10-03 | 2025-10-02 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31551 | DOH | allan.vinluan2013@gmail.com | 30 |  |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6352713844228665752/6352713844228665752_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6352713844228665752/6352713844228665752_signature_31.png
| 2025-10-03 | 2025-10-02 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31550 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is Disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6352712124222552367/6352712124222552367_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6352712124222552367/6352712124222552367_signature_31.png
| 2025-10-03 | 2025-10-02 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31549 | DOH | serly_reyes@yahoo.com | 30 |  |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6352706864225382187/6352706864225382187_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6352706864225382187/6352706864225382187_signature_31.png
| 2025-10-03 | 2025-09-30 | Ilocos Training and Regional Medical Center | San Fernando, La Union | 1 | La Union | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31548 | MACARE-HBDC | itrmc.mbl@gmail.com | 28 |  |  | GXIV-4-D-10C | 120003504 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6352705044229567081/6352705044229567081_signature_28.png | Dominga Patacsil | https://www.jotform.com/uploads/macarewecare/222927978372067/6352705044229567081/6352705044229567081_signature_31.png
| 2025-10-03 | 2025-09-30 | Naguilian District Hospital | Naguilian, La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31547 | DOH | marjelynrayray01@gmail.com | 30 |  |  | GXIV-4-D-6C | 838426 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational -High temperature due to nonfunctional Air-conditioning unit. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6352703984224979401/6352703984224979401_signature_28.png | Justine Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6352703984224979401/6352703984224979401_signature_31.png
| 2025-10-02 | 2025-10-02 | Sta. Cruz Rural Health Unit | Poblacion, Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31546 | DOH | chendalagan72@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 479 |  | GX-IV 6C 4D | 838505 | UPS battery not working For Preventive Maintenance | Reset connection of the UPS, check for line voltages Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance And UPS Troubleshooting | All Modules are operational Faulty UPS Battery | Reported to Jotform for assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6351904650428503549/6351904650428503549_signature_28.png | Dorothy Dalagan,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6351904650428503549/6351904650428503549_signature_31.png
| 2025-10-02 | 2025-10-02 | Pinamalayan Rural Health Unit | Pinamalayan, Oriental Mindoro | 4B | Oriental Mindoro | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31545 | DOH | medtechdave22@gmail.com | 25 |  |  | GXIV | 838230 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6351831170116658829/6351831170116658829_signature_28.png | Kris Allen A. Mabansag | https://www.jotform.com/uploads/macarewecare/222927978372067/6351831170116658829/6351831170116658829_signature_31.png
| 2025-10-02 | 2025-10-02 | Southern Philippines Medical Center | Davao City | 11 | Davao del Sur | Lab | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 31544 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | MCS + | 15H034G | Stuck up mechanical chuck | Checked centrifuge assy Cleaned motor Removed and cleaned mechanical chuck Adjust screw Reset connector Checked motor and alignment Performed functionality test Test run | Performed troubleshooting and repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6351756263612933964/6351756263612933964_signature_28.png | Southern Philippines Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6351756263612933964/6351756263612933964_signature_31.png
| 2025-10-01 | 2025-10-01 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31543 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 54 ASSAY: Xpert CT/NG, COUNT: 69 ASSAY: Xpert HBV VL, COUNT: 618 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1020 ASSAY: Xpert HPV, COUNT: 427 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6613 ASSAY: Xpert MTB-XDR, COUNT: 6 ASSAY: Others, COUNT: 6 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 177 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1421 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. Module D1 has error 1001, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6351100303555350038/6351100303555350038_signature_28.png | Xchosa Kay C. Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/6351100303555350038/6351100303555350038_signature_31.png
| 2025-10-01 | 2025-10-01 | Chinese General Hospital | Blumentrit, Sta. Cruz Manila | NCR | NCR 1st - Manila | Bacteriology | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 31542 | MACARE-HBDC | cgh_clinicalpathology@yahoo.com | 25 |  | 2025-10-01 | GXIV-10C | 120002026 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed. |  | Xpertcheck Kit | SI# 101134 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6351080213972372238/6351080213972372238_signature_28.png | Russell Panem | https://www.jotform.com/uploads/macarewecare/222927978372067/6351080213972372238/6351080213972372238_signature_31.png
| 2025-10-01 | 2025-10-01 | Chinese General Hospital | Blumentrit, Sta Cruz Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 31541 | COMMERCIAL | oncologycgh@gmail.com | 25 |  | 2025-10-01 | GXIV | 835604 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module B1, B3 and B4 passed Xpertcheck. Module B2 (690208) service required. |  | Xpertcheck kit | SI# 101133 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6351080203976717015/6351080203976717015_signature_28.png | Sab claridad | https://www.jotform.com/uploads/macarewecare/222927978372067/6351080203976717015/6351080203976717015_signature_31.png
| 2025-10-01 | 2025-10-01 | Nasipit Rural Health Unit | Brgy. Talisay, Nasipit, ADN | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31540 | DOH | peterdolotina29@gmail.com | 24 |  |  | GX-4 6 Colors | 837611 | UPS doesn't work. Unable to power ON CPU. | None | Replacing UPS battery. CPU Troubleshooting, CMOS Battery Replacement. | All modules and units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6351080013927391028/6351080013927391028_signature_28.png | Peter C Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/6351080013927391028/6351080013927391028_signature_31.png
| 2025-10-01 | 2025-10-01 | Tagum City Health Office | Mabini Street, Tagum City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | tagumcholab.mabini@gmail.com |  | 31539 | DOH | cristiarafernandez@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2947 |  | GX-IV 6C 4D | 838553 | Monitor Problem, Flickering Issue For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Reset power chord, change display cable. Still issue persist | Perform Preventive Maintenance Perform Monitor troubleshooting | All Modules are operational Monitor still having flickering issue | Monitor issue reported to Jotform for assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6351059656224420289/6351059656224420289_signature_28.png | Cristina Tiara A. Fernandez,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6351059656224420289/6351059656224420289_signature_31.png
| 2025-10-01 | 2025-10-01 | St Luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 31538 | COMMERCIAL | megapo@stlukes.com.ph | 26 | ASSAY: Xpert CT/NG, COUNT: 206 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2453 |  | GX4 D 6C | 839234 | For maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Restart workstation Gathered reports Test run unit | Performed maintenance | B1 and B2 operational B3 and B4 reported already for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6350963683411226622/6350963683411226622_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6350963683411226622/6350963683411226622_signature_31.png
| 2025-10-01 | 2025-10-01 | St Luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 31537 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 451 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 ASSAY: Xpert HPV, COUNT: 392 ASSAY: Xpert MTB/RIF Ultra, COUNT: 305 ASSAY: Xpert HCV VL, COUNT: 36 ASSAY: Xpert HIV-1 Viral Load, COUNT: 286 |  | GX4 L 6C | 807155 | For maintenance C4 disabled | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Restart workstation Gathered reports Test run unit | Performed maintenance | Operational C1,C2,C3 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6350963063417038939/6350963063417038939_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6350963063417038939/6350963063417038939_signature_31.png
| 2025-09-30 | 2025-09-25 | Liloy rural health unit | Liloy Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31536 | DOH | nova.timusan79@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1222 | 2025-09-25 | GX 4 | 838536 | Module 1 temperature error. | Preventive Maintenence and data collection. Check module and room temperatures. Module 1 under observation. | Preventive Maintenence and data collection. | All modules are available. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6350342373571269092/6350342373571269092_signature_28.png | Nova L Timusan | https://www.jotform.com/uploads/macarewecare/222927978372067/6350342373571269092/6350342373571269092_signature_31.png
| 2025-09-30 | 2025-09-30 | Vicente Sotto Memorial Medical Center | B. Rodriguez Street, Cebu City | 7 | Cebu | Blood Bank | NO | INSTALLATION | HAEMONETICS | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31535 | COMMERCIAL | charlimcmxci03@gmail.com | 28 |  | 2025-09-30 | REF 9000 | 16L508SPG | new Installation | #NAME? | #NAME? | Unit is in good condition |  |  | 5450 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6350306806129533557/6350306806129533557_signature_28.png | Charlotte Mae M. Ygoña | https://www.jotform.com/uploads/macarewecare/222927978372067/6350306806129533557/6350306806129533557_signature_31.png
| 2025-09-30 | 2025-09-29 | Tropical Disease Foundation | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31534 | COMMERCIAL | aageronimo@tdf.org.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2241 ASSAY: Xpert HIV-1 Viral Load, COUNT: 633 | 2025-09-22 | GX4 D 6C | 812560 | Due for maintenance and xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and xpertcheck | C1,C2,C3,C4 passed in xpertcheck |  | Xpertcheck kits | 101059 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6350130973611267126/6350130973611267126_signature_28.png | Tropical Disease Foundation, Inc. | https://www.jotform.com/uploads/macarewecare/222927978372067/6350130973611267126/6350130973611267126_signature_31.png
| 2025-09-30 | 2025-09-24 | Mary Mediatrix Medical Center | J.P. Laurel Highway, Lipa, Batangas | 4A | Batangas | Molecular | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | biomedmgrmi.mmmc2@gmail.com |  | 31531 | COMMERCIAL | mmmcpcrlab@yahoo.com | 25 |  | 2025-09-22 | GXIV-D-10C | 120002889 | Unit is due for Intallation. | Installation. | Performed Site Assessment. Checked Outlet Voltage and Room Temperature. Install GeneXpert System ( with UPS and Printer). Checked Funtionality. Performed Self-test. | GeneXpert System is Operational. |  | GeneXpert IV Desktop | DR 13349 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6350054853615370591/6350054853615370591_signature_28.png | Trisha Mae M. Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/6350054853615370591/6350054853615370591_signature_31.png
| 2025-09-30 | 2025-09-25 | UST Hospital | Dapitan St, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31533 | COMMERCIAL | ustimp@yahoo.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1006 |  | GXIV-D-6C | 110000711 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | B1 - With Sample Related Error. B2,B3,B4 - Good. | GX Unit For XpertCheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6350054853616776708/6350054853616776708_signature_28.png | Maria Liova Sherleen D. Calaunan | https://www.jotform.com/uploads/macarewecare/222927978372067/6350054853616776708/6350054853616776708_signature_31.png
| 2025-09-30 | 2025-09-17 | Tanza Wellness Center | Antero Soriano Highway, Tanza, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31532 | MACARE-HBDC | tanzarhu.clinicallab@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 156 | 2025-09-02 | GXIV-D-10C | 120003582 | Unit is due for XpertCheck. | XpertCheck. Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed XpertCheck. Performed Self-test. Generate Reports. | XpertCheck Result: B1 -Pass B2- Pass B3- Pass B4- Pass |  | XpertCheck Kit | DR# 13299 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6350054853617113872/6350054853617113872_signature_28.png | Merriam Joy G. Sibbaluca | https://www.jotform.com/uploads/macarewecare/222927978372067/6350054853617113872/6350054853617113872_signature_31.png
| 2025-09-29 | 2025-09-23 | Lake Sebu Rural Health Unit | Lake Sebu, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31530 | DOH | jennypingkian23@yahoo.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 467 |  | GxIV-6C | 838569 | Module 1 under observation due to temperature error | Perform troubleshooting at module 1 due to multiple temperature error Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Data collection | Genexpert module 1 troubleshooting Preventive maintenance | 4 operational modules Module 1 under observation due to temperature errors |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6349639315323876601/6349639315323876601_signature_28.png | Jenny A. Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/6349639315323876601/6349639315323876601_signature_31.png
| 2025-09-29 | 2025-09-24 | Bicol Regional Hospital and Medical Center | Sagpon, Daraga | 5 | Albay | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31529 | MACARE-HBDC | brtth.covid19lab@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 52 |  | GX4-10C | 110026767 | B4/210138841 - 2008 Error (3rd time this September) | Cleaned cartridge bays, plunger rod, Collected Data reported to Cepheid | Performed Troubleshooting | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6349606464017749349/6349606464017749349_signature_28.png | Drinah Labanda | https://www.jotform.com/uploads/macarewecare/222927978372067/6349606464017749349/6349606464017749349_signature_31.png
| 2025-09-29 | 2025-09-23 | Bicol Regional Hospital & Medical Center | Sagpon, Daraga | 5 | Albay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31528 | MACARE-HBDC | brtth.covid19lab@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 44 |  | GX4-10C | 110026767 | B4/110026767 - 4014 error | Cleaned Icore, self test module B4 | Performed Troubleshooting | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6349604194016290796/6349604194016290796_signature_28.png | Drinah Labanda | https://www.jotform.com/uploads/macarewecare/222927978372067/6349604194016290796/6349604194016290796_signature_31.png
| 2025-09-29 | 2025-09-29 | Chong Hua Hospital | Feunte, Cebu City | 7 | Cebu | Bloodbank | NO | REPAIR MAINTENANCE TROUBLESHOOTING | HAEMONETICS | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31527 | COMMERCIAL | ronahofquila@gmail.com | 29 |  |  | 9000 | 11J219 | #NAME? | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6349262135664329321/6349262135664329321_signature_28.png | ronah may ofquila | https://www.jotform.com/uploads/macarewecare/222927978372067/6349262135664329321/6349262135664329321_signature_31.png
| 2025-09-29 | 2025-09-24 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31526 | DOH | gbmmbg@yahoo.com | 25 |  |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6349257234223584823/6349257234223584823_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/6349257234223584823/6349257234223584823_signature_31.png
| 2025-09-29 | 2025-09-25 | Loveyour Self | Cebu City | 7 | Cebu | TB Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 31525 | COMMERCIAL | preforzado@loveyourself.ph | 23 |  |  | GXIV-4D | 814150 | . Hardware Fail module 4 with serial number 627679. | Save Installation Qualification Reports Save System Log Report Save Gxx file | Replace old module and install new module with new serial number 210071653 | Module A1 and A4 ok to use. Module A2 and A3 not ok to use,hardware fail |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6349238854218267538/6349238854218267538_signature_28.png | Paulo Crispo Reforzado | https://www.jotform.com/uploads/macarewecare/222927978372067/6349238854218267538/6349238854218267538_signature_31.png
| 2025-09-26 | 2025-09-26 | Tomas Claudio Health Center | Camus St, Poblacion Davao | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31524 | DOH | arleen_p_alconera@yahoo.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1109 |  | GX-IV 6C 4D | 838492 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Generate IQR, Assay Stat, and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6346752350227141579/6346752350227141579_signature_28.png | Arleen Alconera,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6346752350227141579/6346752350227141579_signature_31.png
| 2025-09-26 | 2025-09-26 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | INSTALLATION REPAIR | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31523 | DOH | bmccrcl2022@gmail.com | 29 |  | 2025-08-28 | GX 4 6 Colors | 838403 | Defective LCD Monitor | Monitor Replacement | Replacing Defective Monitor with SN: CNK94301W3 with SN: H22V8VL658A250716279 Preventive Maintenance | All modules and units are functional. |  | NVISION H22V8 LED MONITOR | 13285 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6346733650973767238/6346733650973767238_signature_28.png | Kristine Joy Pelletero | https://www.jotform.com/uploads/macarewecare/222927978372067/6346733650973767238/6346733650973767238_signature_31.png
| 2025-09-26 | 2025-09-25 | Talacogon District Hospital | Talacogon, Agusan del Sur | 13 | Agusan del Sur | RTDL | NO | INSTALLATION REPAIR | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31522 | DOH | shyra_tortor@gmail.com | 24 |  | 2025-09-16 | GX-4 6 Colors | 838275 | Defective Monitor | Monitor Replacement | Replacing LCD Monitor w/ SN: CNK9430C2P to SN: H22V8VL658A250613059 (NVISION H22V8) | All modules and units are functional |  | NVISION H22V8 21.5" LED MONITOR | 13340 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6346725080972546153/6346725080972546153_signature_28.png | Shyra Erika Tortor | https://www.jotform.com/uploads/macarewecare/222927978372067/6346725080972546153/6346725080972546153_signature_31.png
| 2025-09-26 | 2025-09-26 | Corazon Locsin Hospital | Bacolod City | 6 | Negros Occidental | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31521 | MACARE-HBDC | alainjuayang@yahoo.com | 26 |  |  | GXIV 10C | 110020743 | n/a | #NAME? | Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6346692581817745384/6346692581817745384_signature_28.png | alain juayang | https://www.jotform.com/uploads/macarewecare/222927978372067/6346692581817745384/6346692581817745384_signature_31.png
| 2025-09-26 | 2025-09-24 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31520 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 02D166 | No display UPS unable to backup electricity supply | Reset CPU board Reset display cable Checked cable connectors Checked dispaly cable Checked ups | Performed troubleshooting | Unit unoperational Possible defective display cable For further assessment | For replacement of 2pcs ups battery |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6346558243617840175/6346558243617840175_signature_28.png | Janella Bertiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6346558243617840175/6346558243617840175_signature_31.png
| 2025-09-26 | 2025-09-25 | FEU NRMF | Dahlia st Fairview QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | delosrios.mcq@healthway.com.ph |  | 31519 | COMMERCIAL | kmtejoso@feu-nrmf.edu.ph | 23 | ASSAY: Xpert HPV, COUNT: 42 ASSAY: Xpert MTB/RIF Ultra, COUNT: 415 |  | GX4D 6C | 840975 | Due for maintenance Stuck cartridge in A3 module | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Fixed stuck cartridge Gathered reports Test run unit | Performed maintenance Performed troubleshooting | Operational Last xpertcheck 3/7/2024 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6346557473615435195/6346557473615435195_signature_28.png | Kenneth Anthony M. Tejoso | https://www.jotform.com/uploads/macarewecare/222927978372067/6346557473615435195/6346557473615435195_signature_31.png
| 2025-09-25 | 2025-09-25 | Lung Center Of The Philippines | Quezon Ave QC | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 31518 | COMMERCIAL | lcpbloodbank@gmail.com | 28 |  |  | MCS + | 02D166 | No display during initialization. Buttons not responsive. | Performed troubleshooting. | Reset all connectors. Restart the system. Checked display cable. | Unit is not operational. For further troubleshooting and repair. Display cable related issue. | For problem isolation using known good display cable. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6345936806017491932/6345936806017491932_signature_28.png | Janella Madulid | https://www.jotform.com/uploads/macarewecare/222927978372067/6345936806017491932/6345936806017491932_signature_31.png
| 2025-09-25 | 2025-09-25 | Klinika Project 7 | Bansalangin, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31517 | PBSP | freyaoriana@yahoo.com | 27 |  |  | GXIV-10C | 845446 | Unit for check up Unit for maintenance | NA | #NAME? | Unit is operational Temperature Errors found on Modules B1, B3. Modules B1, B3 is recommended for repair/service |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6345935606128169400/6345935606128169400_signature_28.png | Alyanna Faye Zerrudo | https://www.jotform.com/uploads/macarewecare/222927978372067/6345935606128169400/6345935606128169400_signature_31.png
| 2025-09-25 | 2025-09-25 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31516 | DOH | buhanginlaboratory@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 915 |  | GX-IV 6C 4D | 838232 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6345868066716163268/6345868066716163268_signature_28.png | Catherine B. Deocampo,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6345868066716163268/6345868066716163268_signature_31.png
| 2025-09-24 | 2025-09-24 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31515 | UNOPS | bmcgenexpert@gmail.com | 28 |  |  | GXIV-10C | 110023507 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6345032816129507189/6345032816129507189_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6345032816129507189/6345032816129507189_signature_31.png
| 2025-09-24 | 2025-09-24 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31514 | DOH | luzvimindanavidad2@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 928 | 2025-09-24 | 6C GXIV-4D | 838297 | None | -Perform preventive maintenance; plunger rod, fan filter, cartridge bay, optic slot | #NAME? | Machine is functional. P |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6345011298562032965/6345011298562032965_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6345011298562032965/6345011298562032965_signature_31.png
| 2025-09-24 | 2025-09-24 | Monkayo Rural Health Unit | Poblacion, Monkayo | 11 | Davao de Oro | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31513 | DOH | kendralois730@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1066 |  | GX-IV 6C 4D | 838487 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational Damaged Fan filter | Reported to JotForm |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6345010150714276628/6345010150714276628_signature_28.png | Karen S. Anuta,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6345010150714276628/6345010150714276628_signature_31.png
| 2025-09-24 | 2025-09-24 | Don Jose Monfort Medical Center | Tabucan, Barotac Nuevo, Iloilo | 6 | Iloilo | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31512 | DOH | judyboglosa35@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1909 |  | GXIV | 838501 | n/a | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6344991529416927233/6344991529416927233_signature_28.png | Judy Boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6344991529416927233/6344991529416927233_signature_31.png
| 2025-09-24 | 2025-09-24 | Iloilo Provincial Hospital | Pototan, Iloilo | 6 | Iloilo | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31511 | DOH | djpragale92@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1609 | 2025-09-24 | GXIV | 838289 | LCD monitor SN/ CNK943032K (Defective) | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6344923809417971585/6344923809417971585_signature_28.png | Rachel A Pragale | https://www.jotform.com/uploads/macarewecare/222927978372067/6344923809417971585/6344923809417971585_signature_31.png
| 2025-09-23 | 2025-09-23 | St Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31510 | MACARE-HBDC | laboratory@slec.ph | 23 |  |  | Infinity 48s | 809643 | Conveyor not working | Performed additional troubleshooting | Performed additional troubleshooting steps recommended by Cepheid Tech Support | Unit is operational, automation still down pending parts replacement of conveyor |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6344245723497958680/6344245723497958680_signature_28.png | MA. SARAH R. SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/6344245723497958680/6344245723497958680_signature_31.png
| 2025-09-23 | 2025-09-23 | Iloilo City Health Office | Iloilo City | 6 | Iloilo | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31509 | DOH | mc-molecular@wvsu.edu.ph | 27 |  |  | GXIV | 838290 | n/a | Collect Data | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6344185186612483557/6344185186612483557_signature_28.png | Karla Danielle T. Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6344185186612483557/6344185186612483557_signature_31.png
| 2025-09-23 | 2025-09-23 | Iloilo City Health Office | Iloilo City | 6 | Iloilo | Molecular and Genome Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31508 | DOH | mc-molecular@wvsu.edu.ph | 27 |  |  | GXIV | 838290 | n/a | Collect Data | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6344185186612483557/6344185186612483557_signature_28.png | Karla Danielle T. Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6344185186612483557/6344185186612483557_signature_31.png
| 2025-09-23 | 2025-09-23 | West Visayas State University Medical Center | Jaro , Iloilo City | 6 | Iloilo | Molecular and Genome Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31507 | MACARE-HBDC | mc-molecular@wvsu.edu.ph | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 128 ASSAY: Xpert HPV, COUNT: 53 |  | GXIV | 110021716 | n/a | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6344169436618427004/6344169436618427004_signature_28.png | Karla Danielle T. Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6344169436618427004/6344169436618427004_signature_31.png
| 2025-09-23 | 2025-09-23 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com | biomed@uerm.edu.ph |  | 31506 | MACARE-HBDC | uermbloodbank@gmail.com | 25 |  |  | MCS+ | 13C270 | Unit for maintenance Unit shutting down during autoloading pumps | NA | #NAME? | Unit is operational Rotor roller pumps for replacement Unit for calibration next visit |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6344143306126115065/6344143306126115065_signature_28.png | Ruel | https://www.jotform.com/uploads/macarewecare/222927978372067/6344143306126115065/6344143306126115065_signature_31.png
| 2025-09-23 | 2025-09-23 | Vicente Mendiola Center for Health | Poblacion Naga City Cebu | 7 | Cebu | Tb Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 31505 | DOH | geccaprocianos26@gmail.com | 23 |  |  | GXIV-4 | 838564 | Due for preventive maintenance. | Save Installation Qualification report. Save Assay statistics report. Save Systemog report. Save GXX file. | Perform Preventive maintenance. Clean all modules 4. Clean plunger rod. Clean surface near Genexpert Machine. | All are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6344024727716924947/6344024727716924947_signature_28.png | GENE CARMEL T. PROCIANOS, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6344024727716924947/6344024727716924947_signature_31.png
| 2025-09-23 | 2025-09-21 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Tb laboratory | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31504 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | GX4 | 11J219 | No Display in the panel board. | Reset cpu board in the machine. Performed self test. | Reset cpu board in the machine. Performed self test. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343995155489434132/6343995155489434132_signature_28.png | Chong Hua Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/6343995155489434132/6343995155489434132_signature_31.png
| 2025-09-23 | 2025-09-15 | Minglanilla Rural Health Unit | Pider 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31503 | DOH | edhmaruy@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1480 |  | GX4 | 838512 | Installation of Assay HCV and HBV Due for Monthly Maintenance | Performed Preventive in all modules Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. Require end-user to separate the phlegm and the saliva of the sputum | Performed Preventive in all modules Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. Require end-user to separate the phlegm and the saliva of the sputum | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343993195484299453/6343993195484299453_signature_28.png | Ehd Mar Owen Uy Jambongana | https://www.jotform.com/uploads/macarewecare/222927978372067/6343993195484299453/6343993195484299453_signature_31.png
| 2025-09-23 | 2025-09-23 | Western Visayas Sanitarium | Sta. Barbara, Iloilo | 6 | Iloilo | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31502 | DOH | rugiedequilla@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1589 |  | GXIV | 838393 | LCD monitor | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343990943674616078/6343990943674616078_signature_28.png | Ruth Dequilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6343990943674616078/6343990943674616078_signature_31.png
| 2025-09-22 | 2025-09-17 | Antipas Rural Health Unit | Antipas, North Cotabato | 12 | Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31499 | DOH | codillaalvin06@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 576 |  | GxIV-6C | 838259 | None | Perform Genexpert machine preventive maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status/problem Data reporting | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343600535321087902/6343600535321087902_signature_28.png | Yvan Kris V. Namo | https://www.jotform.com/uploads/macarewecare/222927978372067/6343600535321087902/6343600535321087902_signature_31.png
| 2025-09-22 | 2025-09-16 | Kidapawan City Health Office | Kidapawan City | 12 | Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31501 | DOH | malufranz5@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2448 |  | GxIV-6C | 838545 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of modules status/problem Data reporting | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343600535323842576/6343600535323842576_signature_28.png | Ma. Luisa M. Francisco | https://www.jotform.com/uploads/macarewecare/222927978372067/6343600535323842576/6343600535323842576_signature_31.png
| 2025-09-22 | 2025-09-17 | Cotabato Provincial Hospital | Brgy. Amas, Kidapawan City, North Cotabato | 12 | Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31498 | MACARE-HBDC | januarybing1206@gmail.com | 23.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 52 ASSAY: Xpert HIV-1 Viral Load, COUNT: 95 ASSAY: Xpert MTB-XDR, COUNT: 4 |  | GxIV-6C | 110026514 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and Filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343600535324946694/6343600535324946694_signature_28.png | Hermosilla Enero | https://www.jotform.com/uploads/macarewecare/222927978372067/6343600535324946694/6343600535324946694_signature_31.png
| 2025-09-22 | 2025-09-18 | St. Louis Molecular Diagnostic Laboratory | Tacurong City, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE New Assay installation | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31500 | COMMERCIAL | stlouismoleculardiagnosticslab@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 ASSAY: Xpert HPV, COUNT: 1 |  | GxIV-10C | 110024270 | None | Installed HIV Viral load 1 Assay Genexpert Cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343600535328486316/6343600535328486316_signature_28.png | DE PEDRO, JOHANA SHANE | https://www.jotform.com/uploads/macarewecare/222927978372067/6343600535328486316/6343600535328486316_signature_31.png
| 2025-09-22 | 2025-08-20 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31497 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Partially Operational. Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343263308325780019/6343263308325780019_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6343263308325780019/6343263308325780019_signature_31.png
| 2025-09-22 | 2025-08-19 | Mangatarem Rural Health Unit | Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31496 | DOH | slasheiji1988new@gmail.com | 25 |  | 2025-08-14 | GXIV-4-D-6C | 838284 | For Module Replacement | - Cleaned Cartridge Bays,Plunger Rods and I-CORE - Replaced Defective Module B1 806111 >>>656962 - Generated Reports | Performed Module Replacement | Unit is Fully Operational | Work with Sir Toco | GeneXpert Module 656962 | DR 13259 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343262308324712496/6343262308324712496_signature_28.png | Ken Lee Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/6343262308324712496/6343262308324712496_signature_31.png
| 2025-09-22 | 2025-09-18 | Vinzons Rural Health Unit | Poblacion, Vinzons | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31495 | DOH | bardonamabel@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1193 |  | GX4-6C | 838576 | Gx for Preventive Maintenance B4/675496 temperature Error | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance Checked system log | Unit is Operational B2/862906, B4/675496 temperature Error(1001,2014) | B4/675496 request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343190379518519957/6343190379518519957_signature_28.png | Amabel A. Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/6343190379518519957/6343190379518519957_signature_31.png
| 2025-09-22 | 2025-09-18 | Labo District Hospital | Talobatib, Labo | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31494 | DOH | ldh.lab17@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1372 |  | GX4-6C | 838207 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343187889519904963/6343187889519904963_signature_28.png | Analyn C. Consuelo | https://www.jotform.com/uploads/macarewecare/222927978372067/6343187889519904963/6343187889519904963_signature_31.png
| 2025-09-22 | 2025-09-16 | Estevez Memorial Hospital, Inc. | Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City, 4500 Albay | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31493 | COMMERCIAL | reylovendino83@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 6 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 244 ASSAY: Xpert MTB/RIF Ultra, COUNT: 263 ASSAY: Xpert MTB-XDR, COUNT: 14 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 3 |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, unit surface Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343186449513859082/6343186449513859082_signature_28.png | Rey A. Lovendino | https://www.jotform.com/uploads/macarewecare/222927978372067/6343186449513859082/6343186449513859082_signature_31.png
| 2025-09-22 | 2025-09-22 | Nueva Valencia Rural Health Unit | Nueva Valencia, Guimaras | 6 | Guimaras | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31492 | DOH | galletomaeann@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 493 |  | GXIV | 838291 | #NAME? | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343186236319555351/6343186236319555351_signature_28.png | Mae Ann Galleto | https://www.jotform.com/uploads/macarewecare/222927978372067/6343186236319555351/6343186236319555351_signature_31.png
| 2025-09-22 | 2025-09-15 | Bicol Regional Hospital and Medical Center | Sagpon, Daraga | 5 | Albay | Laboratory | NO | Training | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31491 | MACARE-HBDC | brtth.covid19lab@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX4-10C | 110026767 | Preventive Maintenance training for biomed engr compliance | GeneXpert System Operation and Basic Troubleshooting Training conducted | Performed Preventive Maintenance training | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343185909512349738/6343185909512349738_signature_28.png | Drinah Labanda | https://www.jotform.com/uploads/macarewecare/222927978372067/6343185909512349738/6343185909512349738_signature_31.png
| 2025-09-22 | 2025-09-19 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31490 | DOH | thisisformyfamily@yahoo.com | 27 |  |  | GXIV-4-D-6C | 838286 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343176868327208876/6343176868327208876_signature_28.png | SHIRLEY G. CAYABYAB | https://www.jotform.com/uploads/macarewecare/222927978372067/6343176868327208876/6343176868327208876_signature_31.png
| 2025-09-22 | 2025-09-16 | Sta. Barbara Rural Health Unit | Sta. Barbara Rural Health Unit | 1 | Pangasinan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31489 | DOH | dianeruiz147@gmail.com | 29 |  |  | GXIV-4-D-6C | 838438 | For CPU Troubleshooting | #NAME? | Performed CPU Troubleshooting | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343175868322671614/6343175868322671614_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6343175868322671614/6343175868322671614_signature_31.png
| 2025-09-22 | 2025-09-12 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31488 | DOH | seansimonpaloyo@gmail.com | 24 |  |  | GXIV-4-D-6C | 838473 | For UPS Troubleshooting | #NAME? | Performed UPS Troubleshooting | Unit is Operational. UPS for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343174888321112528/6343174888321112528_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6343174888321112528/6343174888321112528_signature_31.png
| 2025-09-22 | 2025-09-11 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31487 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational. Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343172568323547613/6343172568323547613_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6343172568323547613/6343172568323547613_signature_31.png
| 2025-09-22 | 2025-09-04 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31486 | DOH | allan.vinluan2013@gmail.com | 30 |  |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343171188322784211/6343171188322784211_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6343171188322784211/6343171188322784211_signature_31.png
| 2025-09-22 | 2025-09-04 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31485 | DOH | seansimonpaloyo@gmail.com | 26 |  |  | GXIV-4-D-6C | 838473 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343168148324227585/6343168148324227585_signature_28.png | Manolo Paloyo | https://www.jotform.com/uploads/macarewecare/222927978372067/6343168148324227585/6343168148324227585_signature_31.png
| 2025-09-22 | 2025-09-03 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31484 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 |  |  | GXIV-4-D-6C | 820303 | For Preventive Maintenance and Troubleshooting | #NAME? | Performed Preventive Maintenance and Troubleshooting | Unit is Partially Operational. Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343167038325655740/6343167038325655740_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6343167038325655740/6343167038325655740_signature_31.png
| 2025-09-22 | 2025-08-29 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31483 | DOH | gbmmbg@yahoo.com | 26 |  |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343165738327919887/6343165738327919887_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/6343165738327919887/6343165738327919887_signature_31.png
| 2025-09-22 | 2025-08-28 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31482 | DOH | arleneabad32@gmail.com | 28 |  |  | GXIV-4-D-6C | 838288 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343163958323661163/6343163958323661163_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/6343163958323661163/6343163958323661163_signature_31.png
| 2025-09-22 | 2025-08-28 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31481 | DOH | rhupozorrubio@yahoo.com | 29 |  |  | GXIV-4-D-6C | 838467 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343158648321598147/6343158648321598147_signature_28.png | Azellamie Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/6343158648321598147/6343158648321598147_signature_31.png
| 2025-09-22 | 2025-08-27 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31480 | DOH | serly_reyes@yahoo.com | 28 |  | 2025-08-13 | GXIV-4-D-6C | 838516 | #NAME? | - Replaced Defective Module AT1000190123ST0250 >>> AT1000190123ST0182 - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Generated Reports | Performed UPS Replacement | Unit is Fully Operational |  | UPS AT1000190123ST0190 | DR 13248 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343157128327931039/6343157128327931039_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6343157128327931039/6343157128327931039_signature_31.png
| 2025-09-22 | 2025-09-18 | St. Luke's Medical Center - BGC | Rizal Drive, BGC, Taguig City | NCR | NCR 4th - Taguig | Microbiology | NO | INSTALLATION | OTHERS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph | Fujifilm | 31479 | COMMERCIAL | ltortega@stlukes.com.ph | 24 |  |  | Limusave MT-7500 | 47020656 | For installation | Installed the unit | Installed the unit: Limusave MT-7500 Thermostation TS-70/20 -Sn: 47020388 Cooling station -Sn: 2410015A Functionality test Change date and time settings | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343148773613465553/6343148773613465553_signature_28.png | Lorna T. Ortega | https://www.jotform.com/uploads/macarewecare/222927978372067/6343148773613465553/6343148773613465553_signature_31.png
| 2025-09-22 | 2025-09-22 | Sibunag Rural Health Unit | Sibunag , Guimaras | 6 | Guimaras | TB Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31478 | UNOPS | rhodatrance1994@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 101 | 2025-09-22 | GXIV - 10color | 110023538 | n/a | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6343125961485577002/6343125961485577002_signature_28.png | Ma Rhoda Lyn Villarta | https://www.jotform.com/uploads/macarewecare/222927978372067/6343125961485577002/6343125961485577002_signature_31.png
| 2025-09-19 | 2025-09-19 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31477 | MACARE-HBDC | raedionisio@gmail.com | 26 | ASSAY: Xpert HBV VL, COUNT: 203 ASSAY: Xpert HPV, COUNT: 109 ASSAY: Xpert HCV VL, COUNT: 117 ASSAY: Xpert HIV-1 Viral Load, COUNT: 497 | 2026-05-10 | GXIV-D 10C | 120003587 | Due for Calibration | Preventive maintenance Xpertcheck | Performed Preventive maintenance Update and installation of xpertcheck 1 5B Performed xpertcheck Performed functionality test | Unit is operational | All modules passed | 5 xpertcheck cartridges | XPERTCHECK CE 5 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6340716888215159668/6340716888215159668_signature_28.png | Reanne Dionisio | https://www.jotform.com/uploads/macarewecare/222927978372067/6340716888215159668/6340716888215159668_signature_31.png
| 2025-09-19 | 2025-09-19 | Manuel J. Santos Hospital | 554 Montilla Blvd. Butuan City | 13 | Agusan del Norte | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31476 | DOH | molecularlab@mjsantos.com.ph | 22 |  |  | GX 4 6 Colors | 831292 | Module 1 and 2 failed last year's XpertCheck, Module 4 was found temperature error and found for replacement. | None | XpertCheck Calibration Preventive Maintenance Data Collection | 3 modules for replacement. 1 module passed in xpc calibration. All units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6340694679227771535/6340694679227771535_signature_28.png | Daniella Andrei Plaza | https://www.jotform.com/uploads/macarewecare/222927978372067/6340694679227771535/6340694679227771535_signature_31.png
| 2025-09-19 | 2025-09-19 | Likhaan Center For Women's Health | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31475 | MACARE-HBDC | office@likhaan.org | 27 |  | 2025-09-10 | GXIV-10C | 110025174 | Module B1 (210128780) for replacement | Replaced B1 | - replaced B1 by 706221 - generated reports - performed self test | Unit is operational All modules are functional |  | GeneXpert Module SN: 706221 | 13328 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6340691458117966514/6340691458117966514_signature_28.png | Michelle Armayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6340691458117966514/6340691458117966514_signature_31.png
| 2025-09-19 | 2025-09-19 | Balamban Rural Health Unit | Poblacion Balamban Cebu | 7 | Cebu | TB Laboratory | NO | Install LCD | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 31474 | DOH | climaco.mildred@yahoo.com | 23 |  |  | GXIV-4D | 838231 | Install new LCD with serial number V2360HDJKX20230962817 | Save Installation qualification reports. Save system Log file. Save GXX file. | Install new LCD with serial number V2360HDJKX20230962817 | Functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6340662672716736009/6340662672716736009_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6340662672716736009/6340662672716736009_signature_31.png
| 2025-09-19 | 2025-09-18 | Balamban Rural Health Unit | Poblacion Balamban Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE LCD | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 31473 | DOH | climaco.mildred@yahoo.com | 23 |  |  | GXIV-4D | 838231 | Pull-out LCD with serial number ZV0A1910029805 | Save System Log. Save Installation Qualification. Save Gxx file. | Pull-out LCD with serial number ZV0A1910029805 | Functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6340658262713002705/6340658262713002705_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6340658262713002705/6340658262713002705_signature_31.png
| 2025-09-19 | 2025-09-19 | Quezon City Molecular Diagnostics Laboratory | C.P.Garcia St.,Teacher Village East QC | NCR | NCR 2nd - Quezon City | Molecular lab | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31472 | MACARE-HBDC | molecularlab.cityhealth@quezoncity.gov.ph | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 180 | 2025-09-02 | GX4 D 10C | 120003581 | Due for Maintenance | Performed xpertcheck in all module Cleaned unit Cleaned and disinfect plunger rod,cartridge bay Cleaned filter and icore Gathered reports Test run unit | Performed maintenance | Operational all module passed |  | Xpertcheck kits | 0.133 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6340633075777913364/6340633075777913364_signature_28.png | Kaneesha Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/6340633075777913364/6340633075777913364_signature_31.png
| 2025-09-19 | 2025-09-19 | Cebu TB Reference Lab - SRP | SRP - Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31471 | DOH | asmantiquilla.pbsp@gmail.com | 26 |  |  | Gx4 | 838263 | Defective LCD monitor HP ProDisplay | New monitor Installed N-Vision Sn H22V8VL658A250716492 | New monitor Installed N-Vision Sn H22V8VL658A250716492 Defective monitor bring back to cebu office | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6340575516718846334/6340575516718846334_signature_28.png | Allan S.Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6340575516718846334/6340575516718846334_signature_31.png
| 2025-09-18 | 2025-09-16 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31470 | DOH | sarah_yangzon@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 634 | 2025-09-16 | GXIV | 838406 | #NAME? | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6340099356021416794/6340099356021416794_signature_28.png | sarah yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/6340099356021416794/6340099356021416794_signature_31.png
| 2025-09-18 | 2025-09-18 | Malay Primary Care Facility | Caticlan, Malay, Aklan | 6 | Aklan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31469 | DOH | mhomalayclinicallaboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | 6C GXIV-4D | 838457 | CMOS batt need to be replaced | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6340050532243951171/6340050532243951171_signature_28.png | May Soledad | https://www.jotform.com/uploads/macarewecare/222927978372067/6340050532243951171/6340050532243951171_signature_31.png
| 2025-09-18 | 2025-09-18 | Sarangani Islands Rural Health Unit | Poblacion, Mabila, Sarangani | 11 | Davao Occidental | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31468 | DOH | lovitosquennie26@gmail.com | 29.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 287 |  | GX-IV 6C 4D | 837614 | UPS Battery not working during battery Mode System Unit Force Restarting Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Change UPS Batteries Change Cmos Battery Clean and reset RAM and its slot Check for irregularities at line voltage | Perform Preventive Maintenance Perform UPS Battery Replacement Perform System Unit Troubleshooting | All Modules are operational UPS Battery Successfully Installed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6339944787218286503/6339944787218286503_signature_28.png | QUENNIE C LOVITOS | https://www.jotform.com/uploads/macarewecare/222927978372067/6339944787218286503/6339944787218286503_signature_31.png
| 2025-09-18 | 2025-09-18 | St Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31467 | MACARE-HBDC | laboratory@slec.ph | 21 |  |  | Infinity 48s | 809643 | For Xpertcheck | None | Performed additional Xpertcheck with SW Version 1.5b | Unit is operational. Modules A2/220001426, B1/710723, C2/701127, D4/734837, E1/695487 require servicing. Modules E3, E4, F2,F3, F4, G2, G3 needs retrying with additional Xpertcheck cartridges |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6339890049119344803/6339890049119344803_signature_28.png | Joan Sigrid Goco | https://www.jotform.com/uploads/macarewecare/222927978372067/6339890049119344803/6339890049119344803_signature_31.png
| 2025-09-18 | 2025-09-18 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31466 | COMMERCIAL | gramanmolecularlab@gmail.com | 27 |  |  | GXIV | 110008530 | Unit for maintenance Unit for troubleshooting | NA | - checked and cleaned cartridge bay, plunger rod, filter, icore slot - generated reports - generated .gxx file for futher troubleshooting - performed self test | Modules B1, B2, B3 are recommended for replacement Modules failed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6339840123428126340/6339840123428126340_signature_28.png | Maria Elizabeth De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/6339840123428126340/6339840123428126340_signature_31.png
| 2025-09-18 | 2025-09-17 | Bicol Medical Center | Concepcion Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | REPAIR MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31465 | MACARE-HBDC | judebuenaventurarn@gmail.com | 25 |  |  | MCS+ | 16K520SPG | Display cable temporary borrowed from another machine | Replaced Display Cable of Unit Reset all display cable connections Cleaned Unit Cleaned Centrifuge, valves, rotor pumps, rollers Performed dry run | Repaired machine | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6339827709111165413/6339827709111165413_signature_28.png | Jude Buenaventura | https://www.jotform.com/uploads/macarewecare/222927978372067/6339827709111165413/6339827709111165413_signature_31.png
| 2025-09-18 | 2025-09-17 | Bicol Medical Center | Concepcion Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31464 | MACARE-HBDC | judebuenaventurarn@gmail.com | 25 |  |  | MCS+ | 15H027G | CRC Bios failed | Replaced CRC BIOS Battery Checked/reset display cable connections Calibration of machine after BIOS battery replacement Checked/adjusted technical parameters Cleaned Unit Cleaned Centrifuge, valves, rollers, rotor pump Calibrate MCS+ Machine | Performed Troubleshooting repaired machine Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6339827329113025094/6339827329113025094_signature_28.png | Jude Buenaventura | https://www.jotform.com/uploads/macarewecare/222927978372067/6339827329113025094/6339827329113025094_signature_31.png
| 2025-09-17 | 2025-09-17 | Medi Linx Laboratory Inc (CLDH) | Hospital Drive Brgy San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31463 | COMMERCIAL | mea.delcruz@medilinx.com.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 |  | GXII-D | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module C2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338984834516680876/6338984834516680876_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6338984834516680876/6338984834516680876_signature_31.png
| 2025-09-17 | 2025-09-17 | Medi Linx Laboratory Inc ( CLDH) | Hospital Drive Brgy San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31462 | COMMERCIAL | mea.delacruz@medilincx.com.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 |  | GXIV-D | 739232 | Duefor Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module B4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338984344519679466/6338984344519679466_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6338984344519679466/6338984344519679466_signature_31.png
| 2025-09-17 | 2025-09-17 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31461 | DOH | suzzettekho34@gmail.com | 30 |  |  | GX 4 6 Colors | 838554 | UPS not working. Need Battery Replacement. | UPS Battery Replacement | UPS Battery Replacement Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338910768282997522/6338910768282997522_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/6338910768282997522/6338910768282997522_signature_31.png
| 2025-09-17 | 2025-09-17 | Dr. Rafael S. Tumbokon Memorial Hospital | Kalibo, Aklan | 6 | Aklan | Rhivdha | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31460 | MACARE-HBDC | drstmhcrcl@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1548 ASSAY: Xpert MTB-XDR, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 447 |  | 10C GXIV-4D | 110020745 | None | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338872743024641261/6338872743024641261_signature_28.png | Lynar Ann C. Neptuno | https://www.jotform.com/uploads/macarewecare/222927978372067/6338872743024641261/6338872743024641261_signature_31.png
| 2025-09-17 | 2025-09-16 | Ospital ng Makati | Sampaguita St Makati City | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 31459 | MACARE-HBDC | jinx01gs@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 598 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338855013613133289/6338855013613133289_signature_28.png | Cristina M. Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6338855013613133289/6338855013613133289_signature_31.png
| 2025-09-16 | 2025-09-16 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31458 | UNOPS | jomarilitonjua@gmail.com | 27 |  |  | GXIV-10C | 110023495 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338079603896572680/6338079603896572680_signature_28.png | Cecille Vasquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6338079603896572680/6338079603896572680_signature_31.png
| 2025-09-16 | 2025-09-16 | Altavas Rural Health Unit | Altvas, Aklan | 6 | Aklan | Laboratory | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31457 | DOH | nycolebenedicto@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 574 |  | 6C GXIV-4D | 838458 | None | -Perform preventive maintenance; plunger rod, cartridge bay, fan filter, optic slot | #NAME? | Machine is functional. Module B4 experiencing hardware fail for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338077496868231016/6338077496868231016_signature_28.png | Ma. Nycole S. Benedicto | https://www.jotform.com/uploads/macarewecare/222927978372067/6338077496868231016/6338077496868231016_signature_31.png
| 2025-09-16 | 2025-09-16 | Mandaue City Health Office | Mandaue City, Cebu | 7 | Cebu | Laboratory | YES | ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31456 | DOH | mendozarey42@gmail.com | 28 |  |  | Gx 4 | 838510 | Everyday adjust the time and date of the computer Computer CMOS battery is defective | Replace new CMOS battery for Computer | Replace new CMOS battery for Computer Gather data IQR system log | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338075461414589651/6338075461414589651_signature_28.png | REY C. MENDOZA | https://www.jotform.com/uploads/macarewecare/222927978372067/6338075461414589651/6338075461414589651_signature_31.png
| 2025-09-16 | 2025-09-16 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31455 | MACARE-HBDC | treatmenthub.qcgh@gmail.com | 27 |  |  | GXIV-10C | 110012766 | Unit for maintenance | Na | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338071520364916872/6338071520364916872_signature_28.png | Arizza Sheena Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/6338071520364916872/6338071520364916872_signature_31.png
| 2025-09-16 | 2025-09-16 | Talomo South Health Center | Puan, Libby Road Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | dyrosemarie71@gmail.com |  | 31454 | DOH | dy_rosemarie71@gmail.com | 28.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1316 |  | GX-IV 6C 4D | 838401 | UPS not working dueing battery mode For Preventive Maintenance | Replaced the Defective UPS Battery with new Battery Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance Perform UPS Battery Replacement | All Modules are operational Successfully replaced the UPS Batteries |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6338051791214800526/6338051791214800526_signature_28.png | Rosemarie Dy,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6338051791214800526/6338051791214800526_signature_31.png
| 2025-09-16 | 2025-09-16 | Libacao Rural Health Unit | Libacao, Aklan | 6 | Aklan | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31453 | DOH | haladeeb18@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 419 |  | 6C GXIV-4D | 838540 | None | -Perform preventive maintenance; plunger rod, cartridge bay, fan filter, optic slot | #NAME? | Machine is functional. UPS does not work properly. | UPS batteries for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337981077443292856/6337981077443292856_signature_28.png | Maria Hala Deeb E. Zaulda | https://www.jotform.com/uploads/macarewecare/222927978372067/6337981077443292856/6337981077443292856_signature_31.png
| 2025-09-15 | 2025-09-11 | Cotabato City Health Office | Cotabato City | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31452 | UNOPS | ohslabcbo@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1246 |  | GxIV-10C | 110023524 | None | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337571135329720297/6337571135329720297_signature_28.png | Al-Hakim Muhammad | https://www.jotform.com/uploads/macarewecare/222927978372067/6337571135329720297/6337571135329720297_signature_31.png
| 2025-09-15 | 2025-09-11 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31451 | DOH | ohslabcbo@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1246 |  | GxIV-6C | 838528 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337569885323046554/6337569885323046554_signature_28.png | Al-hakim Muhammad | https://www.jotform.com/uploads/macarewecare/222927978372067/6337569885323046554/6337569885323046554_signature_31.png
| 2025-09-15 | 2025-09-10 | Sultan Mastura Rural Health Unit | Sultan Mastura,Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31450 | DOH | salehshahinaz@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 507 |  | GxIV-6C | 837613 | None | Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status/problem Data reporting | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337569075328404245/6337569075328404245_signature_28.png | Shahinaz Saleh | https://www.jotform.com/uploads/macarewecare/222927978372067/6337569075328404245/6337569075328404245_signature_31.png
| 2025-09-15 | 2025-09-11 | Macare In House | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com | Accupower | 31449 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  | 2025-09-12 | 1kva ups | AT1000190123ST0389 | Defective battery 2pcs | Replaced and installed battery s/n 540428,540429 | Performed repair | Operational |  | Battery 12v 7ah | G8111 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337288093614527614/6337288093614527614_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6337288093614527614/6337288093614527614_signature_31.png
| 2025-09-15 | 2025-09-11 | Macare in House | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | Accupower | 31448 | MACARE-HBDC | florescanoel77@gmail.com | 28 |  | 2025-09-12 | 1kva ups | AT1000190123ST0279 | Defective battery 2pcs | Replaced and installed battery s/n 548676,548677 | Performed repair | Operational |  | Battery 12v 7ah | G8111 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337287663617464192/6337287663617464192_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6337287663617464192/6337287663617464192_signature_31.png
| 2025-09-15 | 2025-09-15 | Manila Central University Hospital | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | lab_mcuhospital@yahoo.com |  | 31447 | COMMERCIAL | purchasing@mcufdtmf.org | 27 |  |  | GXIV-10C | 110018008 | Unit for maintenance | NA | #NAME? | Unit is operational. Unit is due for calibration (3-6-25) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337251052906839453/6337251052906839453_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6337251052906839453/6337251052906839453_signature_31.png
| 2025-09-15 | 2025-09-12 | St. Luke's Medical Center - BGC | Rizal Dr. BGC Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31445 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 36 ASSAY: Xpert CT/NG, COUNT: 155 ASSAY: Xpert HBV VL, COUNT: 24 ASSAY: Xpert HPV, COUNT: 231 ASSAY: Xpert MTB/RIF Ultra, COUNT: 504 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 159 |  | GX-IV-D-10C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337180243612727819/6337180243612727819_signature_28.png | Joshua Aya-ay | https://www.jotform.com/uploads/macarewecare/222927978372067/6337180243612727819/6337180243612727819_signature_31.png
| 2025-09-15 | 2025-09-12 | St. Luke's Medical Center - BGC | Rizal Dr. BGC Taguig City | NCR | NCR 4th - Taguig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31444 | COMMERCIAL | jkcidiesca@stlukes.com.ph | 22 |  |  | MCS+ | 09C125 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337180243618400917/6337180243618400917_signature_28.png | Kenneth Idiesca | https://www.jotform.com/uploads/macarewecare/222927978372067/6337180243618400917/6337180243618400917_signature_31.png
| 2025-09-15 | 2025-09-12 | St. Luke's Medical Center - BGC | Rizal Dr. BGC Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31446 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 186 ASSAY: Xpert CT/NG, COUNT: 664 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 83 ASSAY: Xpert HPV, COUNT: 951 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2375 ASSAY: Xpert MTB-XDR, COUNT: 10 ASSAY: Xpert HCV VL, COUNT: 39 ASSAY: Xpert HIV-1 Viral Load, COUNT: 356 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337180243618846156/6337180243618846156_signature_28.png | Joshua Aya-ay | https://www.jotform.com/uploads/macarewecare/222927978372067/6337180243618846156/6337180243618846156_signature_31.png
| 2025-09-15 | 2025-08-29 | General Trias Social Hygiene Clinic | Hospital Road, General Trias, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | chogentri.lab@gmail.com |  | 31443 | MACARE-HBDC | mackchico00@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1149 |  | GXIV-D-10C | 110016785 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337162273616445162/6337162273616445162_signature_28.png | Mack Alfred L. Chico | https://www.jotform.com/uploads/macarewecare/222927978372067/6337162273616445162/6337162273616445162_signature_31.png
| 2025-09-15 | 2025-09-05 | Asian Hospital and Medical Center | Civic Drive, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | INSTALLATION | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31442 | COMMERCIAL | bloodbank@asianhospital.com | 25 |  | 2025-08-28 | MCS+ | 01F116 | Unit is Repaired and for Installation. | Installation. | Repair and Cleaning of Centrifuge. Install MCS+ Unit. Performed Simulation Test. Pulled-out Service Unit. | Unit is Operational. | Pulled-out Service Unit: 12D166 | MCS+ | DR 13290 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337162273612603757/6337162273612603757_signature_28.png | Roi Jordan A. Ocay | https://www.jotform.com/uploads/macarewecare/222927978372067/6337162273612603757/6337162273612603757_signature_31.png
| 2025-09-15 | 2025-09-11 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31441 | DOH | calambatbdots@gmail.com | 25 |  | 2025-09-02 | GXIV | 838225 | Unit for preventive maintenance. For installation of new ups. | Preventive maintenance and installation of UPS. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed installation of UPS. Generated reports. | Unit is operational. For reporting of module B1 (temperature error). |  | AT1000 UPS (AT1000190123ST0100) in replacement to defective AT1000 (AT1000190123ST0256). | DR#13308 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337162193617895951/6337162193617895951_signature_28.png | Michelle N Pagkalinawan | https://www.jotform.com/uploads/macarewecare/222927978372067/6337162193617895951/6337162193617895951_signature_31.png
| 2025-09-15 | 2025-09-12 | Bicol Medical Center | Concepcion Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com | btsc@bmc.doh.gov.ph |  | 31440 | MACARE-HBDC | josefreyes0715@gmail.com | 25 |  |  | MCS+ | 16K520SPG | Display buttons not working | Reset display board cable connections Swapped display cable Cleaned Unit surface Cleaned, centrifuge, valves, rotor pumps rollers Checked system parameters | Performed Preventive Maintenance Performed Troubleshooting | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6337097376902127874/6337097376902127874_signature_28.png | Josef D. Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6337097376902127874/6337097376902127874_signature_31.png
| 2025-09-12 | 2025-09-12 | National Center for Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31439 | MACARE-HBDC | iancristrono@gmail.com | 25 |  |  | GXIV-10C | 120003480 | Unit for relocation | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334680962909252296/6334680962909252296_signature_28.png | Ian Cris Tono | https://www.jotform.com/uploads/macarewecare/222927978372067/6334680962909252296/6334680962909252296_signature_31.png
| 2025-09-12 | 2025-09-12 | St. Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.con |  |  | 31438 | MACARE-HBDC | laboratory@slec.ph | 21 |  | 2025-09-10 | Infinity 48s | 809643 | Unit for Xpertcheck | None | #NAME? | Unit is operational, 3 out of 6 modules passed. Modules B4/647149, G1/652279, G4/611037 require replacement. Other modules showing up as not available for Xpertcheck. |  | Xpertcheck kit (x8) | DR# 13330 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334672417014059493/6334672417014059493_signature_28.png | Joan Sigrid Goco | https://www.jotform.com/uploads/macarewecare/222927978372067/6334672417014059493/6334672417014059493_signature_31.png
| 2025-09-12 | 2025-09-12 | Panabo City Health Office | Poblacion, Panabo | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31437 | DOH | nadineeust67@gmail.com | 22.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1877 |  | GX-IV 6C 4D | 838295 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334657836882174631/6334657836882174631_signature_28.png | Nadine Gemma Eustaquio,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6334657836882174631/6334657836882174631_signature_31.png
| 2025-09-12 | 2025-09-11 | Santolan Super Health Center | Evangelista Ave Santolan Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31436 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1541 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281881414467/6334478281881414467_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281881414467/6334478281881414467_signature_31.png
| 2025-09-12 | 2025-09-11 | Rosario Super Health Center | Dr. Sixto Antonio Avenue Rosario Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31434 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1340 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281885796101/6334478281885796101_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281885796101/6334478281885796101_signature_31.png
| 2025-09-12 | 2025-09-10 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE Check-up | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31435 | COMMERCIAL | pedelacruz@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | Red blood spillage | Check sensors | Performed maintenance Checked rotor pumps Verified calibration parameters Cleaned and adjust line sensor Cleaned bowl optics Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281887924019/6334478281887924019_signature_28.png | Perlita De La Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281887924019/6334478281887924019_signature_31.png
| 2025-09-12 | 2025-09-10 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31433 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281887924864/6334478281887924864_signature_28.png | Febbie Lou Birang | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281887924864/6334478281887924864_signature_31.png
| 2025-09-12 | 2025-09-10 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31432 | COMMERCIAL | moleculardx@themedicalcity.com | 22 |  |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281889466310/6334478281889466310_signature_28.png | Xchosa Kay C. Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/6334478281889466310/6334478281889466310_signature_31.png
| 2025-09-11 | 2025-09-10 | Pili Rural Health Unit | Poblacion, Pili | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31431 | DOH | lfsanao724@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1528 |  | GX4-6C | 838408 | Gx for Preventive Maintenance CPU hard to open | Cleaned cartridge bays, plunger rod Checked/Cleaned CPU RAM,replugged Power supply cable connections, Swapped power cord Collected Dat | Performed Preventive Maintenance Performed Troubleshooting | Unit is non operational(CPU hard to start) Module B4 662296 - temp error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334077571711781620/6334077571711781620_signature_28.png | Liezl Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/6334077571711781620/6334077571711781620_signature_31.png
| 2025-09-11 | 2025-09-10 | Calabanga Rural Health Unit | Poblacion, Calabanga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31430 | DOH | martyarc11@yahoo.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1114 |  | GX4-6C | 838398 | Gx for Preventive Maintenance CPU kept restarting | Cleaned cartridge bays, plunger rod Cleaned CPU RAM, replugged power supply cable connection, swapped power cord Collect Data | Performed Preventive Maintenance Performed Troubleshooting | Unit is non Operational (CPU kept restarting even after troubleshooting) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334074121717403954/6334074121717403954_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/6334074121717403954/6334074121717403954_signature_31.png
| 2025-09-11 | 2025-09-09 | Bicol Regional Hospital and Medical Center | Sagpon, Legazpi | 5 | Albay | Laboratory | NO | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31429 | MACARE-HBDC | labandadrinah@gmail.com | 20 |  |  | GX4-10C | 110026767 | GX4-10C for Installation | Installed Gx4-10C(110026767), CPU(CPH0725F60007), Monitor(CN-01HRFC- FCC00-47P- CARX-A00), Printer(CN56658GY4), Barcode Scanner(24303010551949) UPS(AT1K240612ST0012) Generate Data | Performed GX4-10C and parts Installation | Unit is Operational All 4-modules and parts are fully functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334068941713209146/6334068941713209146_signature_28.png | Drinah Labanda | https://www.jotform.com/uploads/macarewecare/222927978372067/6334068941713209146/6334068941713209146_signature_31.png
| 2025-09-11 | 2025-09-05 | Virac Rural Health Unit | Virac | 5 | Catanduanes | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31428 | DOH | rhusiv22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2788 |  | GX4-6C | 838563 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334062751715756031/6334062751715756031_signature_28.png | Ramil A. Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/6334062751715756031/6334062751715756031_signature_31.png
| 2025-09-11 | 2025-09-03 | Metro Health Specialists Hospital Inc. | Sorsogon City | 5 | Sorsogon | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31427 | COMMERCIAL | mhshilab@gmail.com | 25 |  |  | GX4-10C | 120002716 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6334060911715730945/6334060911715730945_signature_28.png | Aileen Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6334060911715730945/6334060911715730945_signature_31.png
| 2025-09-11 | 2025-09-11 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Microbiology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31426 | COMMERCIAL | aivee.alcantara@medilinx.com.ph | 25 |  |  | GXII | 812916 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333810334024837957/6333810334024837957_signature_28.png | Aivee Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/6333810334024837957/6333810334024837957_signature_31.png
| 2025-09-11 | 2025-09-11 | St. Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.con |  |  | 31425 | MACARE-HBDC | laboratory@slec.ph | 21 |  |  | Infinity 48s | 809643 | Conveyor belt not working | switched unit to manual mode for now | #NAME? | Unit is operational, running possible in manual mode only. Conveyor requires part replacement due to wear and tear damage. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333799454013894332/6333799454013894332_signature_28.png | MA. SARAH R. SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/6333799454013894332/6333799454013894332_signature_31.png
| 2025-09-11 | 2025-09-11 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31424 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV-10C | 120002891 | Unit for maintenance | Na | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333754170594478124/6333754170594478124_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6333754170594478124/6333754170594478124_signature_31.png
| 2025-09-11 | 2025-09-11 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31423 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 807956 | Unit for maintenance | NA | #NAME? | Unit is operational Module 1&4 for reporting/follow up |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333753770596968960/6333753770596968960_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6333753770596968960/6333753770596968960_signature_31.png
| 2025-09-11 | 2025-09-11 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31422 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 844726 | Unit for installation | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333752362905719852/6333752362905719852_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6333752362905719852/6333752362905719852_signature_31.png
| 2025-09-11 | 2025-09-11 | Solano Rural Health Unit | Solano, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | roxasmatthew659@yahoo.com.ph |  | 31421 | DOH | rhusolanolab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 |  | GXIV-6C | 838201 | Due for Maintenance | Cleaned Icore Replaced Filter Cleaned Cartridge Bay Cleaned Plunger Rod | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333742979519393134/6333742979519393134_signature_28.png | Vanessa Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6333742979519393134/6333742979519393134_signature_31.png
| 2025-09-11 | 2025-09-11 | San Pablo Colleges Medical Center | 110024552 | 4A | Laguna | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31420 | CONSORTIUM | spcmc.molecularlab@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 222 | 2025-08-28 | GXIV-10C | 110024552 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck Kit | SI# 100898 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333742527443321362/6333742527443321362_signature_28.png | RICHARD E. TEMPLO | https://www.jotform.com/uploads/macarewecare/222927978372067/6333742527443321362/6333742527443321362_signature_31.png
| 2025-09-11 | 2025-09-11 | Talakag Rural Health Unit | Talakag, Bukidnon | 10 | Bukidnon | Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31419 | DOH | dianaobera26@gmail.com | 27 |  |  | GX-4 6 Colors | 838566 | None | None | Preventive Maintenance Signing of documents | All modules and units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333721772417413608/6333721772417413608_signature_28.png | Diana Obera | https://www.jotform.com/uploads/macarewecare/222927978372067/6333721772417413608/6333721772417413608_signature_31.png
| 2025-09-11 | 2025-09-11 | Diit District Health Center | Diit Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31418 | DOH | dalepenalosa189@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1083 | 2025-09-11 | GXIV | 838421 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333714280154801675/6333714280154801675_signature_28.png | dalepenalosa189@gmail.com | https://www.jotform.com/uploads/macarewecare/222927978372067/6333714280154801675/6333714280154801675_signature_31.png
| 2025-09-11 | 2025-09-11 | Ramon Rural Health Unit | Ramon, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31417 | DOH | rhuramonlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 661 |  | GXIV-6C | 838480 | Due for Maintenance | Cleaned Icore Cleaned Cartridge Bay Cleaned Filters Cleaned Plundger Rod | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333675119124236287/6333675119124236287_signature_28.png | Carmina Jozenas J. Vizcarra | https://www.jotform.com/uploads/macarewecare/222927978372067/6333675119124236287/6333675119124236287_signature_31.png
| 2025-09-11 | 2025-09-11 | Southern Philippines Medical Center - IM | Bajada, Davao City | 11 | Davao del Sur | IM | NO | Pullout | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 31416 | COMMERCIAL | spmcimdept@gmail.com | 24.5 |  |  | MCS+ | 15HO34G | For pullout defective Centrifuge | Pull out the Defective Centrifuge to be sent back to Office for diagnosis and repair | Perform Pull-out of the Defective Centrifuge | Successfully pullout the Centrifuge |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333674037017499479/6333674037017499479_signature_28.png | Alan S. Togoto | https://www.jotform.com/uploads/macarewecare/222927978372067/6333674037017499479/6333674037017499479_signature_31.png
| 2025-09-11 | 2025-09-11 | Southern Philippines Medical Center | Bajada St, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31415 | UNOPS | spmcrhivda@gmail.com | 26.8 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1634 ASSAY: Xpert MTB-XDR, COUNT: 10 ASSAY: Xpert SARS-CoV-2, COUNT: 12 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1320 |  | GX-IV 10C 4D | 110023505 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333635167012831651/6333635167012831651_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6333635167012831651/6333635167012831651_signature_31.png
| 2025-09-11 | 2025-09-11 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31414 | UNOPS | spmcrhivda@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 621 ASSAY: Xpert SARS-CoV-2, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 858 |  | GX-IV 10C 4D | 110023506 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6333634287011238251/6333634287011238251_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6333634287011238251/6333634287011238251_signature_31.png
| 2025-09-10 | 2025-09-10 | Cagayan Valley Regional Blood Center | Tuguegarao City, Cagayan | 2 | Cagayan | Apheresis Room | NO | REPAIR MAINTENANCE | HAEMONETICS | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31413 | MACARE-HBDC | nvbspchdii2022@gmail.com | 25 |  | 2025-08-13 | MCS+ | 19G146SPG | Rollers worn out | Replaced Rollers | Performed Maintenance Cheked Parameters Performed Test Run | Unit is Operational |  | Rollers | 13246 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332979116979641711/6332979116979641711_signature_28.png | Katherine I. Baclig | https://www.jotform.com/uploads/macarewecare/222927978372067/6332979116979641711/6332979116979641711_signature_31.png
| 2025-09-10 | 2025-09-10 | Chinese General Hospital | Blumentritt, Sta. Cruz Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31412 | COMMERCIAL | cghmcbloodbank2016@gmail.com | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. For checking of purple valve. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Checked valve manifold. Checked all valve's tubings. Performed toggle valve. Performed simulation running. Performed multiple restarting of MCS for monitoring. | Unit is operational. No errors on purple valve after start up test. No error on toggle valve. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332949871475771494/6332949871475771494_signature_28.png | Joan Escolano | https://www.jotform.com/uploads/macarewecare/222927978372067/6332949871475771494/6332949871475771494_signature_31.png
| 2025-09-10 | 2025-09-10 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31411 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2354 |  | 6C GXIV-4D | 838561 | Due for monthly PM | Perform preventive maintenance; plunger rod, fan filter, optic slot, cartridge bay | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332932756682688325/6332932756682688325_signature_28.png | Sheila Mae B. Lariosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6332932756682688325/6332932756682688325_signature_31.png
| 2025-09-10 | 2025-09-10 | Cagayan Valley Medical Center | Carig, Tuguegarao City, Cagayan | 2 | Cagayan | BloodBank | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31410 | MACARE-HBDC | iegparayno@gmail.com | 25 |  | 2025-09-02 | MCS+ | 15H033G | Self Test Problem Bios Battery | Troubleshoot Bios Battery Replaced Bios Battery Configure Technical Parameters Performed Calibration | Performed Maintenance Gathered Reports Performed Test Run 2x ( Procedure Complete) | Unit is Operational |  | Bios Battery | 13307 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332916185916502587/6332916185916502587_signature_28.png | Reggie Parayno | https://www.jotform.com/uploads/macarewecare/222927978372067/6332916185916502587/6332916185916502587_signature_31.png
| 2025-09-10 | 2025-09-10 | Matanao Rural Health Unit | Poblacion, Matanao | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31409 | DOH | pamsustento@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 481 |  | GX-IV 6C 4D | 838456 | UPS not working during battery mode For Preventive Maintenance | Replace the defective battery of the UPS Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and UPS Repair | UPS Batteries Successfully Replaced All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332915296025922498/6332915296025922498_signature_28.png | Dominador s. Castro III | https://www.jotform.com/uploads/macarewecare/222927978372067/6332915296025922498/6332915296025922498_signature_31.png
| 2025-09-10 | 2025-09-10 | The Medical City Clark | Clark Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | tmci_laboratory@themedicalcity.com |  | 31408 | COMMERCIAL | jnasuncion@themedicalcity.com | 25 | ASSAY: Xpert CT/NG, COUNT: 14 ASSAY: Xpert HBV VL, COUNT: 54 ASSAY: Xpert HPV, COUNT: 34 ASSAY: Xpert MTB/RIF Ultra, COUNT: 101 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2027-03-21 | GXIV-D | 839235 | Due for Preventive maintenance and Calibration | Calibration | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Calibration Performed functionality test Gathered reports | Unit is operational all modules passed calibration |  | Xpertcheck kits | Xpertcheck-CE-5 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332912046324274734/6332912046324274734_signature_28.png | James Arthur Asuncion | https://www.jotform.com/uploads/macarewecare/222927978372067/6332912046324274734/6332912046324274734_signature_31.png
| 2025-09-10 | 2025-09-10 | Baybay City Health Office | Baybay City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31407 | DOH | sajecopooh20@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2025-09-10 | GXIV | 838555 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332842284537281476/6332842284537281476_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/6332842284537281476/6332842284537281476_signature_31.png
| 2025-09-10 | 2025-09-10 | CDO Polymedic Medical Plaza | Cagayan de Oro City | 10 | Misamis Oriental | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31406 | COMMERCIAL | cpmpmolecularlab@gmail.com | 23 |  |  | GX 4 10 Colors | 110025567 | None | None | Preventive Maintenance Archiving Tests | All modules and units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332841497129352370/6332841497129352370_signature_28.png | Mary Lou B. Llatuna | https://www.jotform.com/uploads/macarewecare/222927978372067/6332841497129352370/6332841497129352370_signature_31.png
| 2025-09-10 | 2025-09-05 | Bantayan Rural Health Unit | Bantayan Island | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31405 | DOH | mcdelle916bausin@gmai.com | 28 |  |  | Gx 4 | 838509 | N/A | Site Visit Preventive Maintenance Gather data IQR, system log Clean Fan Filter | Site Visit Preventive Maintenance Gather data IQR, system log Clean Fan Filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332834707752299847/6332834707752299847_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/6332834707752299847/6332834707752299847_signature_31.png
| 2025-09-10 | 2025-09-05 | Bantayan Rural Health Unit | Bantayan Island | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31404 | DOH | mcdelle916bausin@gmai.com | 28 |  |  | Gx 4 | 838509 | N/A | Site Visit Preventive Maintenance Gather data IQR, system log Clean Fan Filter | Site Visit Preventive Maintenance Gather data IQR, system log Clean Fan Filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332834397752270215/6332834397752270215_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/6332834397752270215/6332834397752270215_signature_31.png
| 2025-09-09 | 2025-09-09 | Cardinal Santos Medical Center | Wilson St San Juan Greenhills | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE Check-up | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | renalyn.cabaltea@medilinx.com.ph |  | 31403 | COMMERCIAL | janeldvelasquez@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 66 ASSAY: Xpert HPV, COUNT: 39 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1155 |  | GX-IV-D-10C | 110003714 | Error 5007 on module B4 | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports Generated.gxx file | Unit is operational. Module B3 disabled. Module B4 under observation. GeneXpert unit due for XpertCheck. | Liver mass used in B4 with error 5007 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332273775803479025/6332273775803479025_signature_28.png | Janel | https://www.jotform.com/uploads/macarewecare/222927978372067/6332273775803479025/6332273775803479025_signature_31.png
| 2025-09-09 | 2025-09-09 | Tagoloan Rural Health Unit | Poblacion, Tagoloan | 10 | Misamis Oriental | Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31402 | DOH | venzbacoy@gmail.com | 28 |  |  | GX-4 6 Colors | 838574 | None | None | Preventive Maintenance and Signing Documents | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332153057123288510/6332153057123288510_signature_28.png | Venz Michael Bacoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6332153057123288510/6332153057123288510_signature_31.png
| 2025-09-09 | 2025-09-09 | Cotabato Regional Medical Center | Cotabato City | BARMM | Maguindanao | Molecular lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31401 | DOH | crmcmolecularlab@gmail.com | 24.1 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 3169 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 704 ASSAY: Xpert SARS-CoV-2, COUNT: 289 |  | GxlV-6color | 837662 | None | Import HIV viral load XC Assay Perform genexpert cleaning and maintenance Module cleaning Plunger rod and optics cleaning Fan and filter cleaning Checking of module status and problem Data reporting | Genexpert Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6332129758021306856/6332129758021306856_signature_28.png | Darianne Mae Tubilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6332129758021306856/6332129758021306856_signature_31.png
| 2025-09-05 | 2025-09-05 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | Relocation | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31400 | COMMERCIAL | monetlaxina@yahoo.com.ph | 26 |  |  | GX4L | 804205 | For relocation | Relocate unit Set up unit | Performed relocation | Unit not in use D2,D3,D4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6328628393611895360/6328628393611895360_signature_28.png | Simonette Laxina | https://www.jotform.com/uploads/macarewecare/222927978372067/6328628393611895360/6328628393611895360_signature_31.png
| 2025-09-05 | 2025-09-04 | MMC Medilinx | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31399 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 150 ASSAY: Others, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 391 ASSAY: Xpert HBV VL, COUNT: 76 ASSAY: Xpert HPV, COUNT: 219 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1747 ASSAY: Others, COUNT: 9 ASSAY: Xpert HCV VL, COUNT: 98 ASSAY: Xpert HIV-1 Viral Load, COUNT: 951 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 12 |  | GXXVI D | 110026448 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6328589503613590257/6328589503613590257_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6328589503613590257/6328589503613590257_signature_31.png
| 2025-09-05 | 2025-09-02 | Best Diagnostic Corp. | 94 Masikap st.,Diliman QC | NCR | NCR 2nd - Quezon City | Lab | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31398 | COMMERCIAL | laboratory@bestdiagnotic.com.ph | 30 | ASSAY: Xpert HBV VL, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX4 D | 110008529 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6328588863611433613/6328588863611433613_signature_28.png | Rose Marie Balisi | https://www.jotform.com/uploads/macarewecare/222927978372067/6328588863611433613/6328588863611433613_signature_31.png
| 2025-09-05 | 2025-09-04 | Marikina City Health Office | Shoe Avenue Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31397 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 350 ASSAY: Xpert MTB-XDR, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 226 |  | GX-IV-D-10C | 110024072 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6328530843244549500/6328530843244549500_signature_28.png | Esty Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6328530843244549500/6328530843244549500_signature_31.png
| 2025-09-05 | 2025-09-04 | Marikina City Health Office | Shoe Avenue, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31396 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 343 ASSAY: Xpert MTB-XDR, COUNT: 22 ASSAY: Xpert HIV-1 Viral Load, COUNT: 330 |  | GX-IV-D-10C | 110024073 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6328530843246154409/6328530843246154409_signature_28.png | Esty Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6328530843246154409/6328530843246154409_signature_31.png
| 2025-09-04 | 2025-09-04 | Madrid District Hospital | Madrid, Surigao del Sur | 13 | Surigao del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31395 | DOH | madridlaboratory@gmail.com | 28 |  |  | GX-4 6 Colors | 838439 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  | Faizar Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/6327815505416922005/6327815505416922005_signature_28.png | Rochelle Pueblas | https://www.jotform.com/uploads/macarewecare/222927978372067/6327815505416922005/6327815505416922005_signature_31.png
| 2025-09-04 | 2025-09-04 | Poro Rural Health Unit | Poro, Camotes, Cebu | 7 | Cebu | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31394 | DOH | avilacharmaine6@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 478 | 2025-08-14 | 6C GXIV-4D | 838279 | Module B2 for replacement | -Replace defective module B2/808037 with module SN: 644627 | #NAME? | Machine is functional. |  | 900-0491RS Certified pre-owned GeneXpert module SN. 644627 | DR no. 13258 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6327732183024744453/6327732183024744453_signature_28.png | Charmaine A. Apat | https://www.jotform.com/uploads/macarewecare/222927978372067/6327732183024744453/6327732183024744453_signature_31.png
| 2025-09-04 | 2025-09-04 | Dinas Rural Health Unit | Dinas, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement and UPS Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31393 | DOH | bernantereynacres@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 806 | 2025-08-14 | GX 4 | 838533 | UPS unable to work properly | Replaced defective module 862294 with new module 210059928. Replaced defective ups SN AT1000190123ST0358 with new UPS SN AT1000190123ST0280. Preventive Maintenance, data collection. | Replaced defective module 862294 with new module 210059928. Replaced defective ups SN AT1000190123ST0358 with new UPS SN AT1000190123ST0280. Preventive Maintenance, data collection. | All modules are working properly while UPS is ready to use and needs to be charged atleast 8hrs. | Newly replaced UPS must be charge for atleast 8hrs | 900-0491RS FC-IT-UPS-ACCU1K | 13274/ 13261 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6327730967442417221/6327730967442417221_signature_28.png | Reyna Cres Bernante | https://www.jotform.com/uploads/macarewecare/222927978372067/6327730967442417221/6327730967442417221_signature_31.png
| 2025-09-04 | 2025-09-04 | Antipolo City Health Office | M. Santos St. Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31392 | UNOPS | acho.laboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5 ASSAY: Xpert MTB-XDR, COUNT: 120 ASSAY: Xpert HIV-1 Viral Load, COUNT: 757 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6327657185807460715/6327657185807460715_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6327657185807460715/6327657185807460715_signature_31.png
| 2025-09-03 | 2025-09-02 | Bicol Medical Center | Concepcion Pequeña, Naga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com | btsc@bmc.doh.gov.ph |  | 31391 | MACARE-HBDC | ma.charitydelosreyes@yahoo.com | 25 |  |  | MCS+ | 16K520SPG | Machine display button not working | Reset display cables Swapped display board for testing Checked display cable connections Reset machine | Performed Troubleshooting | Unit is non operational Display buttons not working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6326943705325313348/6326943705325313348_signature_28.png | Ma. Charity de los Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6326943705325313348/6326943705325313348_signature_31.png
| 2025-09-03 | 2025-09-03 | Macare In House | 1005 Aurora Grand | NCR | NCR 2nd - Quezon City | Macare in house | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com | UPS | 31390 | MACARE-HBDC | florescanoel77@gmail.com | 28 |  |  | Accupower 1kva | . | Drain battery 0 volts | Checked and test voltage output of battery | Performed troubleshooting | Unit for replacement of battery 4pcs for UPS s/n AT1000190123ST0389, AT1000190123ST0279 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6326936093619981694/6326936093619981694_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6326936093619981694/6326936093619981694_signature_31.png
| 2025-09-03 | 2025-09-03 | Alicia Rural Health Unit | Alicia Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31389 | DOH | castrojore29@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 187 | 2025-06-29 | GX 4 | 838276 | Monitor lcd malfunctioning. | Replaced defective monitor CNK9430C28 with properly functioning monitor CN48180550. Preventive maintenence. | Replaced defective monitor CNK9430C28 with properly functioning monitor CN48180550. Preventive maintenence. | Genexpert Machine is working fine and ready to use. | None | FC-E-MON8TOR | 13214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6326880779311999831/6326880779311999831_signature_28.png | Joremay Panganoron | https://www.jotform.com/uploads/macarewecare/222927978372067/6326880779311999831/6326880779311999831_signature_31.png
| 2025-09-03 | 2025-09-03 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31388 | DOH | burauenlab@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 567 |  | GXIV | 838551 | n/a | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6326873776484709572/6326873776484709572_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6326873776484709572/6326873776484709572_signature_31.png
| 2025-09-03 | 2025-09-03 | Baganga Rural Health Unit | Poblacion, Baganga | 11 | Davao Oriental | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31387 | DOH | rhubaganga@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 427 |  | GX-IV 6C 4D | 838395 | Due for Preventive Maintenance UPS Problem | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Check Line Voltages Check Battery Voltages | Perform Preventive Maintenance Perform UPS Troubleshooting | All Modules are operational UPS not working during battery mode |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6326853099308801729/6326853099308801729_signature_28.png | Liza Narciso,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6326853099308801729/6326853099308801729_signature_31.png
| 2025-09-03 | 2025-09-03 | Cebu TB reference Lab | Cebu City - SRP | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31386 | DOH | asmantiquilla.pbsp@gmail.com | 26 |  |  | Gx4 | 838263 | Defective LCD Monitor | Gather data IQR, System log Clean Fan Filter Calibration Info report Site has own spare LCD monitor to replace the defective LCD monitor for the meantime | Gather data IQR, System log Clean Fan Filter Calibration Info report Site has own spare LCD monitor to replace the defective LCD monitor for the meantime | Sytems operational All modules passed during xpertcheck calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6326782511152109553/6326782511152109553_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6326782511152109553/6326782511152109553_signature_31.png
| 2025-09-03 | 2025-09-02 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31385 | UNOPS | eamcmolecularlab@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 1165 ASSAY: Xpert HBV VL, COUNT: 22 ASSAY: Xpert HPV, COUNT: 102 ASSAY: Xpert SARS-CoV-2, COUNT: 708 ASSAY: Xpert SARS-CoV-2, COUNT: 48 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 547 |  | GX4 10C | 110016781 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6326687253612993178/6326687253612993178_signature_28.png | East Avenue Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6326687253612993178/6326687253612993178_signature_31.png
| 2025-09-02 | 2025-09-02 | Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31384 | DOH | miniforestlaboratory@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 557 |  | GX-IV 6C 4D | 838507 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6325974016511786263/6325974016511786263_signature_28.png | Lolijell Torralba,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6325974016511786263/6325974016511786263_signature_31.png
| 2025-09-02 | 2025-09-02 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31383 | DOH | laureanjo@yahoo.com | 28 |  |  | Gx 4 | 838527 | Due date for xpertcheck calibration | Clean fan Data gathered system log, IQR Xpertcheck calibration | Clean fan Data gathered system log, IQR Xpertcheck calibration | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6325951375519327826/6325951375519327826_signature_28.png | Laurean jo cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6325951375519327826/6325951375519327826_signature_31.png
| 2025-09-02 | 2025-08-28 | Cagayancillo Rural Health Unit | Brgy. Calsada I, Cagayancillo, Palawan | 4B | Palawan | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31382 | MACARE-HBDC | kmricablanca@gmail.com | 24 |  |  | GX-IV-D-10C | 110026465 | For installation | Installed the GeneXpert System | Installed the GeneXpert System -GeneXpert Unit -Workstation -Printer -Barcode Scanner w/ stand -UPS Functionality test | Unit is operational. All modules operational. | Repack the GeneXpert System |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6325846093612585623/6325846093612585623_signature_28.png | Karen Mae Argamaso | https://www.jotform.com/uploads/macarewecare/222927978372067/6325846093612585623/6325846093612585623_signature_31.png
| 2025-09-02 | 2025-08-29 | Puerto Princesa City Health Office | Puerto Princesa, Palawan | 4B | Palawan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31381 | UNOPS | jocal0324@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 254 |  | GX-IV-D-10C | 110023510 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -Plunger rod -Icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6325846093613572739/6325846093613572739_signature_28.png | Jocelyn L Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/6325846093613572739/6325846093613572739_signature_31.png
| 2025-09-02 | 2025-08-29 | Puerto Princesa City Health Office | Puerto Princesa, Palawan | 4B | Palawan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31380 | MACARE-HBDC | jocal0324@gmqip.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2142 |  | GX-XVI-D-10C | 110021518 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6325846093615241054/6325846093615241054_signature_28.png | Jocelyn L. Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/6325846093615241054/6325846093615241054_signature_31.png
| 2025-08-29 | 2025-08-29 | Claveria Rural Health Unit | Poblacion, Claveria | 5 | Masbate | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31379 | DOH | mgrmartinez757@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 628 |  | GX4-6C | 838572 | CPU not powering on | Repaired CPU - CPH0520F08309 Cleaned cartridge bays, plunger rod Generate Data | Repaired CPU Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6322638060746258792/6322638060746258792_signature_28.png | Mary Grace R.Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/6322638060746258792/6322638060746258792_signature_31.png
| 2025-08-29 | 2025-08-28 | Pio Duran Rural Health Unit | Caratagan, Pio Duran | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31378 | DOH | mitzijc@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 471 |  | GX4-6C | 838264 | Gx for Preventive Maintenance | Cleaned cartridge bays and plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6322637530745647392/6322637530745647392_signature_28.png | Joselyn Oscuro | https://www.jotform.com/uploads/macarewecare/222927978372067/6322637530745647392/6322637530745647392_signature_31.png
| 2025-08-29 | 2025-08-29 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31377 | DOH | calambatndots@gmail.com | 25 |  |  | GXIV | 838225 | Unit for preventive maintenance. For checking of Ups. | Preventive maintenance and troubleshootong. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. Performed inspection of Ups. | Unit is operational. UPS AT1000 (sn: AT1000190123ST0256) not backing up and is for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6322581353112442617/6322581353112442617_signature_28.png | Lily P.Banalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6322581353112442617/6322581353112442617_signature_31.png
| 2025-08-28 | 2025-08-28 | Palawan Molecular Diagnostic Laboratory | Irawan, Puerto Princesa, Palawan | 4B | Palawan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31376 | UNOPS | maejg.gualdrapa@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 534 |  | GX-IV-D-10C | 110023509 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6321794791276684692/6321794791276684692_signature_28.png | MYKEANNE MAEJ'G GUALDRAPA | https://www.jotform.com/uploads/macarewecare/222927978372067/6321794791276684692/6321794791276684692_signature_31.png
| 2025-08-28 | 2025-08-28 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31375 | DOH | lornsabaigar@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 740 |  | GXIV | 838409 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6321661875419489579/6321661875419489579_signature_28.png | Ma. Bella P Lozada | https://www.jotform.com/uploads/macarewecare/222927978372067/6321661875419489579/6321661875419489579_signature_31.png
| 2025-08-28 | 2025-08-28 | Sta Cruz Rural Health Unit | Poblacion Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31374 | DOH | chendalagan72@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 498 |  | GX-IV 6C 4D | 838505 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6321659818227944499/6321659818227944499_signature_28.png | Dorothy Dalagan,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6321659818227944499/6321659818227944499_signature_31.png
| 2025-08-28 | 2025-08-27 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 31373 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  | 2025-08-22 | MCS | 20168 | Defective display cable Unresponsive keypad | Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Replaced and installed display cable Performed functionality test Performed simulation running | Performed repair | Operational |  |  | 13276 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6321531203612762460/6321531203612762460_signature_28.png | Lorrabelle Salada | https://www.jotform.com/uploads/macarewecare/222927978372067/6321531203612762460/6321531203612762460_signature_31.png
| 2025-08-27 | 2025-08-27 | Macare In House | Aurora Grand | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com | engineering@macare-medicals.com |  | 31372 | COMMERCIAL | htoco@macare-medicals.com | 25 |  | 2025-08-27 | MCS+ | 01F116 | 1. Stucked up centrifuge mechanical chuck. 2. Always fluid detected in the centrifuge. | Performed centrifuge overhaul and repair. Performed centrifuge cleaning and unit maintenance. Verify unit calibration parameters. Performed functionality test and simulation run. | 1. Overhaul/cleaned the centrifuge assembly.. 2. Repair mechancal chuck. Apply lubricant solutionn. 3. Cleaned fluid detector. | Unit is operational. For return to facolity | Repaired unit, for return. | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320869433611231471/6320869433611231471_signature_28.png | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/6320869433611231471/6320869433611231471_signature_31.png
| 2025-08-27 | 2025-08-27 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31371 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3191 |  | 6C GXIV-4D | 838285 | Due for PM | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320834538137435903/6320834538137435903_signature_28.png | Fran Marie Uberas | https://www.jotform.com/uploads/macarewecare/222927978372067/6320834538137435903/6320834538137435903_signature_31.png
| 2025-08-27 | 2025-08-27 | Pilar Rural Health Unit | Calongay, Pilar | 5 | Sorsogon | Lbc | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31370 | DOH | manalangmarkramir@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1871 |  | GX4-6C | 838266 | Gx for Preventive Maintenance Auto adjust frequent appearance on Display | Replaced Display Cable(VGA-VGA) Checked/adjust display resolution Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance Performed Troubleshooting (monitor) | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320821693692124062/6320821693692124062_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/6320821693692124062/6320821693692124062_signature_31.png
| 2025-08-27 | 2025-08-27 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31369 | DOH | rhwclab@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 389 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2284 |  | GX-IV 6C 4D | 838262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320810127029991088/6320810127029991088_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6320810127029991088/6320810127029991088_signature_31.png
| 2025-08-27 | 2025-08-22 | Ragay Rural Health Unit | Poblacion, Ragay | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31368 | DOH | phylgomez0115@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 727 |  | GX4-6C | 838560 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational | Requesting for extra Gx filter |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320783583695185515/6320783583695185515_signature_28.png | Phyl Frances C.Gomez | https://www.jotform.com/uploads/macarewecare/222927978372067/6320783583695185515/6320783583695185515_signature_31.png
| 2025-08-27 | 2025-08-27 | Roxas Rural Health Unit | Roxas, Palawan | 4B | Palawan | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31367 | MACARE-HBDC | benjpaul.1994@gmail.com | 24 |  |  | GX-IV-D-10C | 110026464 | For installation | Installed the GeneXpert system | Installed the GeneXpert system -GeneXpert Unit -Workstation- -Barcode scanner with Stand -Printer -UPS Functionality test | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320776458128976400/6320776458128976400_signature_28.png | Paul John Benjamin | https://www.jotform.com/uploads/macarewecare/222927978372067/6320776458128976400/6320776458128976400_signature_31.png
| 2025-08-27 | 2025-08-27 | Diit District Health Center | Diit, Tacloban City | 8 | Leyte | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 31366 | DOH | cholaboratory23@gmail.com | 25 |  | 2025-08-27 | GXIV | 838421 | #NAME? | - Module replacement B3/210005379 into B3/731306 - Preventive Maintenance - Functionality test | #NAME? | Unit is Operational |  | 1 Genexpert Module 6 color SN/731306 | 13260 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320739655442930646/6320739655442930646_signature_28.png | Ma. Gennette Dolly D. Garapan | https://www.jotform.com/uploads/macarewecare/222927978372067/6320739655442930646/6320739655442930646_signature_31.png
| 2025-08-27 | 2025-08-20 | St. Luke's Medical Center - BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31365 | COMMERCIAL | jaato@st.lukes.com.ph | 20 |  |  | MCS+ | 09C125 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Functionality test Simulation run | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663448126377833/6320663448126377833_signature_28.png | JAQUILYN ATO | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663448126377833/6320663448126377833_signature_31.png
| 2025-08-27 | 2025-08-20 | St. Luke's Medical Center - BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31363 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 39 ASSAY: Xpert CT/NG, COUNT: 154 ASSAY: Xpert HBV VL, COUNT: 16 ASSAY: Xpert HPV, COUNT: 227 ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 143 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663438121775177/6320663438121775177_signature_28.png | Jordan Purpura | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663438121775177/6320663438121775177_signature_31.png
| 2025-08-27 | 2025-08-22 | Atimonan Rural Health Unit | Atimonan, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31364 | DOH | AMCL2008@yahoo.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 610 | 2025-08-14 | GX-IV-D-6C | 838206 | Module B1 for replacement | Replaced module B1 | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Replaced module B1 Generated reports | Unit is operational. All modules operational |  | Old: B1/807287 New: B1/861246 | DR# 13262 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663438123909370/6320663438123909370_signature_28.png | Marife Villamiel | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663438123909370/6320663438123909370_signature_31.png
| 2025-08-27 | 2025-08-20 | St. Luke's Medical Center - BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | CI | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31362 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 187 ASSAY: Xpert CT/NG, COUNT: 670 ASSAY: Xpert HBV VL, COUNT: 67 ASSAY: Xpert HPV, COUNT: 956 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2405 ASSAY: Xpert MTB-XDR, COUNT: 10 ASSAY: Xpert HCV VL, COUNT: 44 ASSAY: Xpert HIV-1 Viral Load, COUNT: 335 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663438126033992/6320663438126033992_signature_28.png | Jordan Purpura | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663438126033992/6320663438126033992_signature_31.png
| 2025-08-27 | 2025-08-19 | Santolan Super Health Center | Evangelista Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31361 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1496 |  | GX-IV-D-10C | 110024247 | Stuck cartridge on module B1 | Removed stuck cartridge | Manual removal of stuck cartridge Performed maintenance Cleaned: -Cartridge bay -Plunger rod -Icore -Filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663438127210622/6320663438127210622_signature_28.png | Maricel m. Cayabyab | https://www.jotform.com/uploads/macarewecare/222927978372067/6320663438127210622/6320663438127210622_signature_31.png
| 2025-08-27 | 2025-08-20 | Columbio Rural Health Unit | Columbio, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31360 | DOH | rochelle.toledo19@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 |  | GxIV-6C | 838556 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6320610215327306614/6320610215327306614_signature_28.png | Rochelle S. Toledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6320610215327306614/6320610215327306614_signature_31.png
| 2025-08-26 | 2025-08-24 | Vicente Sotto Memorial Medical Center | Pide 1, Minglanilla ,Cebu | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31359 | COMMERCIAL | vsmmcbloodservices@gmail.com | 24 |  |  | MCS+ | 02A090 | Cuff pressure problem. | Pull out the cuff and restart the machine Check the cuff solenoid. Instruct end-user how to resolve problem | Pull out the cuff and restart the machine Check the cuff solenoid. Instruct end-user how to resolve problem | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319966667754123242/6319966667754123242_signature_28.png | Charlotte Mae Ygoña, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6319966667754123242/6319966667754123242_signature_31.png
| 2025-08-26 | 2025-08-21 | Chong Hua Hospital- Fuente MAB | FUNTE OSMENA, CEBU | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | chh_microbiology@chonghua.com.ph |  | 31358 | COMMERCIAL | chhc_laboratory@chonghua.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2936 ASSAY: Xpert BCR-ABL Ultra, COUNT: 70 ASSAY: Xpert CT/NG, COUNT: 368 ASSAY: Xpert HBV VL, COUNT: 460 ASSAY: Xpert HBV VL, COUNT: 201 ASSAY: Xpert MTB-XDR, COUNT: 16 ASSAY: Xpert HCV VL, COUNT: 241 ASSAY: Xpert HIV-1 Viral Load, COUNT: 709 |  | GX4 | 110020777 | Conduct Preventive Maintenance, and Performed calibration in all modules Clean Icore, Plunger and Filter Generate IQR, System log, Gxx Files and Assay Statistics | Conduct Preventive Maintenance, and Performed calibration in all modules Clean Icore, Plunger and Filter Generate IQR, System log, Gxx Files and Assay Statistics | Conduct Preventive Maintenance, and Performed calibration in all modules Clean Icore, Plunger and Filter Generate IQR, System log, Gxx Files and Assay Statistics | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319964257758038382/6319964257758038382_signature_28.png | Chong Hua Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/6319964257758038382/6319964257758038382_signature_31.png
| 2025-08-26 | 2025-08-20 | Minglanilla Rural Health Unit | Pide 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | ramasmarkefren@gmail.com |  | 31357 | COMMERCIAL | edhmaruy@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1480 |  | GX4 | 838512 | Repeated error of 2008 in different Modules. | Require end-user to review the proper procedure in sample process Performed Preventive in all modules Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. Separate the phlegm and the saliva of the sputum All 2008 errors eliminated as of today. | Performed Preventive in all modules Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. Require end-user to separate the phlegm and the saliva of the sputum | All modules are operational. | Pide 1, Minglanilla ,Cebu |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319951087759773608/6319951087759773608_signature_28.png | Ehd Mar Owen Uy Jambongana | https://www.jotform.com/uploads/macarewecare/222927978372067/6319951087759773608/6319951087759773608_signature_31.png
| 2025-08-26 | 2025-08-26 | Padada Rural Health Unit | Poblacion, Padada | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31356 | DOH | eejjgamboa@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 616 |  | GX-IV 6C 4D | 838511 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319908848711308336/6319908848711308336_signature_28.png | Erica Jane O. Gamboa | https://www.jotform.com/uploads/macarewecare/222927978372067/6319908848711308336/6319908848711308336_signature_31.png
| 2025-08-26 | 2025-08-20 | Mandaue City Health Office | Mandaue City, Cebu | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31355 | DOH | mendozarey42@gmail.com | 28 |  |  | Gx4 | 838510 | Repetitive change time of Computer | Install HVB assay Clean Fan Filter Gather data IQR, System log | Install HVB assay Clean Fan Filter Gather data IQR, System log | Systems operational, need replacement for CMOS battery |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319904657753861362/6319904657753861362_signature_28.png | REY C. MENDOZA | https://www.jotform.com/uploads/macarewecare/222927978372067/6319904657753861362/6319904657753861362_signature_31.png
| 2025-08-26 | 2025-08-15 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31354 | DOH | thisisformyfamily@yahoo.com | 26 |  | 2025-08-13 | GXIV-4-D-6C | 838286 | For UPS Replacement | - Replaced Defective UPS AT1000190123ST0279 >>> AT1000190402ST0370 - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Filter - Generated Reports | Performed UPS Replacement | Unit is Fully Operational |  | UPS AT1000190402ST0370 | DR 13249 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319833549247431441/6319833549247431441_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/6319833549247431441/6319833549247431441_signature_31.png
| 2025-08-26 | 2025-08-13 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31353 | COMMERCIAL | annabellelcatbagan@gmail.com | 28 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational, Module 1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319831889249262564/6319831889249262564_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6319831889249262564/6319831889249262564_signature_31.png
| 2025-08-26 | 2025-08-12 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31352 | DOH | seansimonpaloyo@gmail.com | 26 |  |  | GXIV-4-D-6C | 838473 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319829969245496299/6319829969245496299_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6319829969245496299/6319829969245496299_signature_31.png
| 2025-08-26 | 2025-08-07 | Rosales Rural Health Unit | Resales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31351 | DOH | allan.vinluan2013@gmail.com | 28 |  |  | GXIV-4-D-6C | 838472 | #NAME? | #NAME? | #NAME? | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319827069245211001/6319827069245211001_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6319827069245211001/6319827069245211001_signature_31.png
| 2025-08-26 | 2025-08-07 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31350 | DOH | seansimonpaloyo@gmail.com | 23 |  |  | GXIV-4-D-6C | 838473 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319825229249070007/6319825229249070007_signature_28.png | Manolo C. Paloyo | https://www.jotform.com/uploads/macarewecare/222927978372067/6319825229249070007/6319825229249070007_signature_31.png
| 2025-08-26 | 2025-08-14 | Bataan General Hospital and Medica Center | Balanga Btaan | 3 | Bataan | PCR Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31349 | MACARE-HBDC | raedionisio@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 159 ASSAY: Xpert HPV, COUNT: 109 ASSAY: Xpert HCV VL, COUNT: 94 ASSAY: Xpert HIV-1 Viral Load, COUNT: 436 |  | GXIV-D 10C | 120003587 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319823167978425067/6319823167978425067_signature_28.png | Raeann Dionisio | https://www.jotform.com/uploads/macarewecare/222927978372067/6319823167978425067/6319823167978425067_signature_31.png
| 2025-08-26 | 2025-08-26 | Aurora Grand | 1005 Aurora Blvd., Brgy. Duyan-duyan, Proj. 3, Quezon City | NCR | NCR 2nd - Quezon City | Building Administration | NO | TROUBLESHOOTING | OTHERS | Carlos Leniel Origenes |  | carlos.peniel.origenes@gmail.com |  | DCM | 31348 | COMMERCIAL | rmagpayo@macare-medicals.com | 24 |  |  | SL-2K-ACE | 191203070004 | Not backing up | None | Troubleshoot possible battery or board problem | Unit for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6319797923616317841/6319797923616317841_signature_28.png | Ramil Magpayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6319797923616317841/6319797923616317841_signature_31.png
| 2025-08-22 | 2025-08-22 | Quezon City Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31347 | MACARE-HBDC | molecularlab.cityhealth@quezoncity.gov.ph | 29 |  |  | GXIV-10C | 110020696 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6316554294306363826/6316554294306363826_signature_28.png | Cyrell Arcueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6316554294306363826/6316554294306363826_signature_31.png
| 2025-08-22 | 2025-08-22 | Quezon City Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31346 | MACARE-HBDC | molecularlab.cityhealth@quezoncity.gov.ph | 25 |  |  | GXIV-10C | 120003581 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6316554231705651833/6316554231705651833_signature_28.png | Cyrell Arcueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6316554231705651833/6316554231705651833_signature_31.png
| 2025-08-22 | 2025-08-22 | Santa Rosa City Health Office 1 | Santa Rosa City Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31345 | UNOPS | santarosacho1clinlqb@gmail.com | 25 |  |  | GXIV-10C | 110023497 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6316535550718498730/6316535550718498730_signature_28.png | Marivic Batallones | https://www.jotform.com/uploads/macarewecare/222927978372067/6316535550718498730/6316535550718498730_signature_31.png
| 2025-08-22 | 2025-08-22 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31344 | DOH | bmccrcl2022@gmail.com | 27 |  |  | GX-4 6 Colors | 838403 | Monitor sometimes goes OSD lockout and automatically access menu setting. | Reported to Jotform | Preventive Maintenance and Troublehooting | All modules are functional. Monitor can be used as for now. CPU and other accessories are working. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6316509694514556886/6316509694514556886_signature_28.png | Joechita Doysabas | https://www.jotform.com/uploads/macarewecare/222927978372067/6316509694514556886/6316509694514556886_signature_31.png
| 2025-08-22 | 2025-08-22 | Likhaan Center For Women’s Health | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | Check Up | MEDGYN | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31343 | MACARE-HBDC | clatras@likhaan.org | 25 |  |  | Cryogun | 11922, 11918, 11112, 11917 | Unit for check up | Tightened seals | #NAME? | Units are operational O-rings and washers are recommended for replacement since units were stuck for a long time to prevent leaks |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6316501034301843926/6316501034301843926_signature_28.png | Jocelyn Latras | https://www.jotform.com/uploads/macarewecare/222927978372067/6316501034301843926/6316501034301843926_signature_31.png
| 2025-08-22 | 2025-08-22 | Likhaan Center For Women's Health | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31342 | MACARE-HBDC | clatras@likhaan.org | 25 |  |  | GXII-10C | 110025174 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6316500971702859108/6316500971702859108_signature_28.png | Jocelyn Latras | https://www.jotform.com/uploads/macarewecare/222927978372067/6316500971702859108/6316500971702859108_signature_31.png
| 2025-08-22 | 2025-08-12 | Minglanilla Rural Health Unit | Pide 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | ramasmarkefren@gmail.com |  | 31341 | DOH | edhmaruy@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1402 |  | GX4 | 838512 | Repeated error of 2008 in different Modules. | Require end-user to review the proper procedure in sample process Performed Preventive in all modules Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. | Performed Preventive in all modules Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6316349317759671088/6316349317759671088_signature_28.png | Ehd Mar Owen Uy Jambongana | https://www.jotform.com/uploads/macarewecare/222927978372067/6316349317759671088/6316349317759671088_signature_31.png
| 2025-08-22 | 2025-05-08 | Quezon City Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31340 | MACARE-HBDC | molecularIab.cityhealth@quezoncity.gov.ph | 27 |  |  | GXIV-10C | 120003581 | Unit for maintenance | Na | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6316343932127633565/6316343932127633565_signature_28.png | Cyrell Arcueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6316343932127633565/6316343932127633565_signature_31.png
| 2025-08-22 | 2025-08-22 | St Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31339 | MACARE-HBDC | laboratory@slec.ph | 21 |  |  | Infinity 48s | 809643 | Module latching error on modules A3 (754001) and B2 (702345) | Cleared automation error | #NAME? | Unit is operational, modules are under observation and for reporting to Cepheid TS |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6316342142422311397/6316342142422311397_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6316342142422311397/6316342142422311397_signature_31.png
| 2025-08-20 | 2025-08-20 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Blood donor section | YES | REPAIR CALIBRATION | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 31338 | MACARE-HBDC | hanielzlee@gmail.com | 28 |  | 2025-08-12 | MCS + | 15H044 | Due for calibration For replacement of line sensor assy | Cleaned unit Checked program parameters Checked and adjust line sensor value Replaced and installed line sensor assy Checked dpm,spm,centrifuge,weigher Replaced and installed rollers Performed functionality test Performed calibration Performed simulation running | Performed maintenance and calibration. Performed repair | Operational |  | Urethane roller Line sensor assy | 100791/13275/13241 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6314844890669268112/6314844890669268112_signature_28.png | EAMC | https://www.jotform.com/uploads/macarewecare/222927978372067/6314844890669268112/6314844890669268112_signature_31.png
| 2025-08-20 | 2025-08-20 | Baybay City Health Office | Baybay Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 31337 | DOH | sajecopooh20@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1401 |  | GXIV | 838555 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6314756854535958212/6314756854535958212_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/6314756854535958212/6314756854535958212_signature_31.png
| 2025-08-20 | 2025-08-19 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 31336 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 23 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 433 ASSAY: Xpert HPV, COUNT: 411 ASSAY: Xpert MTB/RIF Ultra, COUNT: 262 ASSAY: Xpert HCV VL, COUNT: 36 ASSAY: Xpert HIV-1 Viral Load, COUNT: 254 |  | GX4 L | 807155 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | C4 disabled due to error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6314571903614195459/6314571903614195459_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6314571903614195459/6314571903614195459_signature_31.png
| 2025-08-20 | 2025-08-19 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | YES | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 31335 | COMMERCIAL | megapo@stlukes.com.ph | 26 | ASSAY: Xpert CT/NG, COUNT: 196 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2541 |  | GX4 D | 839234 | Due for maintenance Broken uniframe module B3 866517, B4 866800 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Checked broken uniframe B3,B4 Test run unit | Performed maintenance | B3,B4 disabled due to failed xpertcheck last march 6,2023 and march 14,2023 and broken uniframe B1,B2 operational | For replacement of module b3,b4 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6314571633614248352/6314571633614248352_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6314571633614248352/6314571633614248352_signature_31.png
| 2025-08-19 | 2025-08-19 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31334 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  |  | GXIV | 110008530 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313941712313998080/6313941712313998080_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6313941712313998080/6313941712313998080_signature_31.png
| 2025-08-19 | 2025-08-19 | Tagum City Health Office | Mabini St, Tagum City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | cristiarafernandez@gmail.com |  | 31333 | DOH | tagumcholab.mabini@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2814 |  | GX-IV 6C 4D | 838553 | For Preventive Maintenance UPS Troubleshooting | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Check Line Voltages of UPS, reset connection | Perform Preventive Maintenance and UPS Troubleshooting | No power on UPS in battery or powerline mode All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313916987313205855/6313916987313205855_signature_28.png | Cristina Tiara A. Fernandez,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6313916987313205855/6313916987313205855_signature_31.png
| 2025-08-19 | 2025-08-19 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31332 | COMMERCIAL | owenjoshuabriones@yahoo.com | 27 |  |  | Gx4 | 829419 | Defective module 2 | Hvb assay installaton Replace new genexpert filter Collect data IQR, System log | Hvb assay installaton Replace new genexpert filter Collect data IQR, System log | Defective module 2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313866538745950653/6313866538745950653_signature_28.png | Owen joshua l. Briones | https://www.jotform.com/uploads/macarewecare/222927978372067/6313866538745950653/6313866538745950653_signature_31.png
| 2025-08-19 | 2025-08-15 | San Fernando Rural Health Unit | Poblacion, San Fernando | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31331 | DOH | josephus_garfin@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 977 |  | G | 838496 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational All 4 modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313759472121337408/6313759472121337408_signature_28.png | Josephus Garfin | https://www.jotform.com/uploads/macarewecare/222927978372067/6313759472121337408/6313759472121337408_signature_31.png
| 2025-08-19 | 2025-08-14 | San Jose Municipal Health Office | San Jose | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31330 | DOH | jhuneric2008@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 956 |  | GX4-6C | 838389 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313759132126773190/6313759132126773190_signature_28.png | Rosmar V. Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6313759132126773190/6313759132126773190_signature_31.png
| 2025-08-19 | 2025-08-14 | Tigaon Municipal Health Office | Tigaon | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31329 | DOH | jingimperial24@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1218 |  | GX4-6C | 838558 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313757492126769781/6313757492126769781_signature_28.png | Maggie Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6313757492126769781/6313757492126769781_signature_31.png
| 2025-08-18 | 2025-08-13 | San Juan District Health Center | Brgy. San Juan, Surigao City | 13 | Surigao del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31328 | UNOPS | sanjuandistricthealthcenter@yahoo.com | 27 |  |  | GX4 10 Colors | 110023522 | None | Preventive maintenance | Preventive Maintenanc Data Collection | All module and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313240588814522373/6313240588814522373_signature_28.png | Anarose Ogdol | https://www.jotform.com/uploads/macarewecare/222927978372067/6313240588814522373/6313240588814522373_signature_31.png
| 2025-08-18 | 2025-08-18 | Bagumbayan Rural Health Unit | Bagumbayan, Sultan Kudarat | 13 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31327 | DOH | bambi.vic6@gmail.com | 28.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1474 |  | GxIV-6C | 838261 | Module 2&3 H/W failed | Perform module 2 and 3 troubleshooting due to temperature errors Checking of module status Problem reporting Data collection | Genexpert machine module errors troubleshooting | 3 operational modules Module 2 disabled | Module 2 requesting for module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313045529282405789/6313045529282405789_signature_28.png | Videlyne Carlet | https://www.jotform.com/uploads/macarewecare/222927978372067/6313045529282405789/6313045529282405789_signature_31.png
| 2025-08-18 | 2025-08-14 | Polomolok Rural Health Unit | Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31326 | DOH | jmmirasolrmtrmt@gmail.com | 27.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2147 |  | GxIV-6C | 838265 | None | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313012209282647275/6313012209282647275_signature_28.png | Rommel Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/6313012209282647275/6313012209282647275_signature_31.png
| 2025-08-18 | 2025-08-14 | Polomolok Rural Health Unit | Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31325 | DOH | jmmirasolrmtrmt@gmail.com | 27.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2147 |  | GxIV-6C | 838265 | None | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6313012039281949752/6313012039281949752_signature_28.png | Rommel Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/6313012039281949752/6313012039281949752_signature_31.png
| 2025-08-18 | 2025-08-14 | Pasig Treatment Hub - Klinika Pasig | Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31322 | MACARE-HBDC | pasigtreatmentsumilang@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 412 |  | GX-IV-D-10C | 110025128 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312948453611913171/6312948453611913171_signature_28.png | Joseph Victor Cainto | https://www.jotform.com/uploads/macarewecare/222927978372067/6312948453611913171/6312948453611913171_signature_31.png
| 2025-08-18 | 2025-08-13 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 31324 | COMMERCIAL | samantha.gallardo@medilinx.com.ph | 22 | ASSAY: Xpert HBV VL, COUNT: 76 ASSAY: Xpert HPV, COUNT: 41 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1119 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B3 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312948453611947254/6312948453611947254_signature_28.png | Samantha Victoria A. Gallardo | https://www.jotform.com/uploads/macarewecare/222927978372067/6312948453611947254/6312948453611947254_signature_31.png
| 2025-08-18 | 2025-08-14 | Pasig Treatment Hub - CHAMP | Champ Building, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31323 | MACARE-HBDC | pasigtreatment@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 388 |  | GX-IV-D-10C | 110025127 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312948453618681622/6312948453618681622_signature_28.png | Abner D. Godes | https://www.jotform.com/uploads/macarewecare/222927978372067/6312948453618681622/6312948453618681622_signature_31.png
| 2025-08-18 | 2025-08-13 | Cardinal Santos Medical Center | Wilson St, Greenhills, San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 31321 | COMMERCIAL | CSMCLAB@CSMC.PH | 22 |  |  | MCS+ | 08M461 | For maintenance SPM low during start of cycle | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312948453611326725/6312948453611326725_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/6312948453611326725/6312948453611326725_signature_31.png
| 2025-08-18 | 2025-08-14 | Biopath QC | E.Rodriguez | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31320 | COMMERCIAL | rmlao@biopath.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 39 ASSAY: Others, COUNT: 1 ASSAY: Xpert HPV, COUNT: 51 ASSAY: Others, COUNT: 6 ASSAY: Xpert SARS-CoV-2, COUNT: 10 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX16D | 839306 | A1,A3,A4,B2,B4,C2,D2 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Due for maintenance | A1,A4,B2,B4,C2,D2 failed xpertcheck disabled A3 hardware failure |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312926983618102797/6312926983618102797_signature_28.png | Rhenee Lao-Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6312926983618102797/6312926983618102797_signature_31.png
| 2025-08-18 | 2025-08-15 | Rosario Rural Health Unit | Rosario Municipal Compound, Rosario, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31319 | DOH | arlenepenaarmijo@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1139 |  | GXIV-D-6C | 838537 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | 4 Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916533613646587/6312916533613646587_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916533613646587/6312916533613646587_signature_31.png
| 2025-08-18 | 2025-08-07 | Muntinlupa Reproductive Health and Wellness Center | Centennial Road, Tunasan, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31318 | MACARE-HBDC | petelerica@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 345 |  | GXIV-D-10C | 110025706 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test and Functionality Test. Generate Reports. | 4 Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916483619814351/6312916483619814351_signature_28.png | Erica Petel | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916483619814351/6312916483619814351_signature_31.png
| 2025-08-18 | 2025-08-05 | University of Santo Tomas Hospital | Lacson Ave, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31316 | MACARE-HBDC | bloodbank@usthospital.com.ph | 25 |  |  | MCS+ | 02F069 | Unit is due for Calibration. | Calibration. | Calibrate A/D and Cebtrifuge Values. Checked Sensors. Cleaned Pumps and Sensors. Performed Simulation Test. | Unit is Good and Operational l. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916473613301179/6312916473613301179_signature_28.png | Mylene Balboa | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916473613301179/6312916473613301179_signature_31.png
| 2025-08-18 | 2025-08-14 | Asian Hospital and Medical Center | Civic Drive, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | Pull-out | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | bloodbank@asiahospital.com |  | 31317 | COMMERCIAL | kcaangay@asianhospital.com | 25 |  | 2025-08-13 | MCS+ | 01F116 | Spill Detected Problem. | Troubleshoot. | Pulled-out MCS+ Unit. Install Service Unit. | Pulled-out Defective Unit for Repair and Further Troubleshooting. Installed Service Unit. | Service Unit S/N: 12D166 | MCS+ Apheresis System. | DR 13252 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916473613772528/6312916473613772528_signature_28.png | Keano Riggs Caañgay | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916473613772528/6312916473613772528_signature_31.png
| 2025-08-18 | 2025-08-14 | Ospital ng Muntinlupa | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31314 | UNOPS | osmunmicro@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 744 |  | GXIV-D-10C | 110023494 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | 4 Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916463613737207/6312916463613737207_signature_28.png | Jesse Jasper Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916463613737207/6312916463613737207_signature_31.png
| 2025-08-18 | 2025-08-08 | Dasmariñas City Health Office II | DBB-E, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31315 | DOH | hazelmilan0777@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3145 |  | GXIV-D-6C | 838535 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Replaced Filter. Performed Self-test. Generate Reports. Checked System's Functionality. | 4 Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916473618782657/6312916473618782657_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/6312916473618782657/6312916473618782657_signature_31.png
| 2025-08-15 | 2025-08-15 | Kairos Laboratories and Diagnostics Inc. | Mabini St. Ermita Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31313 | CONSORTIUM | kairos.jamartinez@gmail.com | 25 |  |  | GX-IV 10C (2 modules) | 110022622 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6310476895336220624/6310476895336220624_signature_28.png | Jude Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/6310476895336220624/6310476895336220624_signature_31.png
| 2025-08-14 | 2025-08-08 | Argao Primary Health Care (RHU II) | Argao, Poblacion Cebu City | 7 | Cebu | Tb laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31312 | DOH | victoria.jumamil@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 920 |  | GX4 | 838283 | #NAME? | Preventive Maintenance of the Genexpert. Check UPS voltage and charged the UPS, check power switch. Check back up voltage. Short time back up battery. | Preventive Maintenance of the Genexpert. Check UPS voltage and charged the UPS, check power switch. Check back up voltage. Generate data gxx, IQR, System log and Assay Statistics | Genepert unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309692547758002875/6309692547758002875_signature_28.png | Victoria Jumamil, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6309692547758002875/6309692547758002875_signature_31.png
| 2025-08-14 | 2025-08-14 | Western Visayas Sanitarium and General Hospital | Sta. Barbara Iloilo | 6 | Iloilo | TB DOTS | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31311 | DOH | ruthiedequilla@gmail.com | 23 |  |  | GX-IV | 838393 | -MODULE 1 H/W FAILED | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL, MODULE B1 FOR REPORTING |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309658392323396489/6309658392323396489_signature_28.png | Ruth Dequila, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6309658392323396489/6309658392323396489_signature_31.png
| 2025-08-14 | 2025-08-14 | Tigaon Municipal Health Office | Caraycayon, Tigaon | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31310 | DOH | jingimperial24@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1218 |  | GX4-6C | 838558 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309625526415488116/6309625526415488116_signature_28.png | Maria Jesusa P. Imperial RMT II | https://www.jotform.com/uploads/macarewecare/222927978372067/6309625526415488116/6309625526415488116_signature_31.png
| 2025-08-14 | 2025-08-14 | Tigaon Municipal Health Office | Caraycayon, Tigaon | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31309 | DOH | jingimperial24@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1218 |  | GX4-6C | 838558 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is operational All 4-modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309624576412503661/6309624576412503661_signature_28.png | Maria Jesusa P. Imperial RMT II | https://www.jotform.com/uploads/macarewecare/222927978372067/6309624576412503661/6309624576412503661_signature_31.png
| 2025-08-14 | 2025-08-14 | Maco Rural Health Unit | Poblacion, Maco | 11 | Davao de Oro | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31308 | DOH | edlienkai26@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 921 |  | GX-IV 6C 4D | 838500 | Due for Preventive Maintenance and Data Gathering | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test | Perform Preventive Maintenance and Data Gathering | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309596724786011479/6309596724786011479_signature_28.png | Mary Jean Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/6309596724786011479/6309596724786011479_signature_31.png
| 2025-08-14 | 2025-08-14 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31307 | MACARE-HBDC | phc_bb@yahoo.com | 28 |  |  | MCS | 20168 | Unresponsive keypad | Checked display cable Reset cable Reser machine After troubleshooting is performed still the machine keypad is unresponsive Possible defective parts display cable | Performed troubleshooting | Unoperational For replacement of display cable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309594375212601556/6309594375212601556_signature_28.png | Philippine Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6309594375212601556/6309594375212601556_signature_31.png
| 2025-08-14 | 2025-08-14 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31306 | DOH | burauenlab@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 538 |  | GXIV | 838551 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309574141217249309/6309574141217249309_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6309574141217249309/6309574141217249309_signature_31.png
| 2025-08-14 | 2025-08-14 | Chinese General Hospital | Blumentritt, Sta. Cruz Manila | NCR | NCR 1st - Manila | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31305 | MACARE-HBDC | cghmcbloodbank2016@gmail.com | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309561691515306706/6309561691515306706_signature_28.png | Matromanrey M. Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/6309561691515306706/6309561691515306706_signature_31.png
| 2025-08-14 | 2025-08-14 | Calubian Rural Health Unit | Calubian, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31304 | DOH | mjcabardo6@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 703 |  | 6C GXIV-4D | 838399 | Due for PM | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309554444542354960/6309554444542354960_signature_28.png | Mighty Joy Cabardo | https://www.jotform.com/uploads/macarewecare/222927978372067/6309554444542354960/6309554444542354960_signature_31.png
| 2025-08-14 | 2025-08-13 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE Data Gather | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31303 | DOH | luzvimindanavidad2@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 932 |  | 6C GXIV-4D | 838296 | None | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309545764549282723/6309545764549282723_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6309545764549282723/6309545764549282723_signature_31.png
| 2025-08-14 | 2025-08-14 | Chinese General Hospital | Blumentritt, Sta. Cruz Manila | NCR | NCR 1st - Manila | Bacteriology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 31302 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1075 ASSAY: Xpert HIV-1 Viral Load, COUNT: 125 |  | GXIV-10C | 120002026 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309537261519894340/6309537261519894340_signature_28.png | Hershelyn Bale | https://www.jotform.com/uploads/macarewecare/222927978372067/6309537261519894340/6309537261519894340_signature_31.png
| 2025-08-14 | 2025-08-13 | Delos Santos Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@dlsmc.ph |  | 31301 | CONSORTIUM | jdvelarde@dlsmc.ph | 36 | ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 41 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 |  | GX4D 10C | 110020700 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309287892297674682/6309287892297674682_signature_28.png | JEANELA D. VELARDE | https://www.jotform.com/uploads/macarewecare/222927978372067/6309287892297674682/6309287892297674682_signature_31.png
| 2025-08-14 | 2025-08-13 | Iriga City Health Office | Sta. Cruz Sur, Iriga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31300 | UNOPS | hannahalevon1@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1459 |  | GX4-10C | 110023511 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational All 4 modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309120167219843899/6309120167219843899_signature_28.png | Hannah Kirsten Novela | https://www.jotform.com/uploads/macarewecare/222927978372067/6309120167219843899/6309120167219843899_signature_31.png
| 2025-08-14 | 2025-08-13 | Iriga City Health Office | Sta. Cruz Sur, Iriga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31299 | DOH | hannahalevon1@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1261 |  | GX4-6C | 838391 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational All 4 modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309119307218251903/6309119307218251903_signature_28.png | Hannah Kirsten Novela | https://www.jotform.com/uploads/macarewecare/222927978372067/6309119307218251903/6309119307218251903_signature_31.png
| 2025-08-14 | 2025-08-12 | Estevez Memorial Hospital, Inc. | Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | emh@estevezhospital.com |  | 31298 | COMMERCIAL | reylovendino83@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 6 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 258 ASSAY: Xpert MTB/RIF Ultra, COUNT: 276 ASSAY: Xpert MTB-XDR, COUNT: 14 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 4 |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore, fan filter Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6309117167213827247/6309117167213827247_signature_28.png | Rey Lovendino | https://www.jotform.com/uploads/macarewecare/222927978372067/6309117167213827247/6309117167213827247_signature_31.png
| 2025-08-13 | 2025-08-13 | The Medical City Clark | Clark City Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | tmcclarkbiomed@themedicalcity.com |  | 31297 | COMMERCIAL | tmci_laboratory@themedicalcity.com | 22 | ASSAY: Xpert CT/NG, COUNT: 13 ASSAY: Xpert HBV VL, COUNT: 56 ASSAY: Xpert HPV, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 107 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GXIV-D 6C | 839235 | Due for Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except Module B4: 866908 Temperature Error (4012) | B4 for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6308747437317201044/6308747437317201044_signature_28.png | Jessa Mae Remorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6308747437317201044/6308747437317201044_signature_31.png
| 2025-08-13 | 2025-08-13 | DON JOSE S MONFORT MEDICAL CENTER | Tabucan Barotac Nuevo | 6 | Iloilo | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31296 | DOH | judyboglosa35@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1959 |  | GX-IV | 838501 | #NAME? | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE : CLEAN FAN FILTER, PLUNGER ROD, AND CARTRIDGE BAY -INSPECT UPS, CONFIRMED DEFECTIVE BATTERY. DEF SN: 1000190123ST0097 | EQUIPMENT IS OPERATIONAL, UPS FOR PULL-OUT |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6308608950693400787/6308608950693400787_signature_28.png | judy boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6308608950693400787/6308608950693400787_signature_31.png
| 2025-08-13 | 2025-08-11 | MMC Medilinx | Amorsolo st., Makati | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31295 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 127 ASSAY: Others, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 348 ASSAY: Xpert HBV VL, COUNT: 66 ASSAY: Xpert HPV, COUNT: 173 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1466 ASSAY: Others, COUNT: 9 ASSAY: Xpert HCV VL, COUNT: 90 ASSAY: Xpert HIV-1 Viral Load, COUNT: 832 |  | GX16 D | 110026448 | End user encounter A1 2014 temp error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | A1 for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6308535233613807602/6308535233613807602_signature_28.png | Aiza G. Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6308535233613807602/6308535233613807602_signature_31.png
| 2025-08-12 | 2025-08-12 | St. Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31294 | MACARE-HBDC | laboratory@slec.ph | 21 |  |  | Infinity 48s | 809643 | Unit for checking after fire incident in laboratory with fire extinguisher residue on unit | #NAME? | #NAME? | Unit is operational, no errors encountered during observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307968021819001712/6307968021819001712_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6307968021819001712/6307968021819001712_signature_31.png
| 2025-08-12 | 2025-08-12 | Bagumbayan Rural Health Unit | Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31293 | DOH | marizalenoid@yahoo.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1488 |  | GxIV-6C | 838261 | None | Perform Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter Cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307911315324383143/6307911315324383143_signature_28.png | MA. RIZA R. DIONELA | https://www.jotform.com/uploads/macarewecare/222927978372067/6307911315324383143/6307911315324383143_signature_31.png
| 2025-08-12 | 2025-08-12 | Diit District Health Center | Diit, Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31292 | DOH | dalepenalosa189@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1090 |  | GXIV | 838421 | N/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307873998961899852/6307873998961899852_signature_28.png | Dale D. Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6307873998961899852/6307873998961899852_signature_31.png
| 2025-08-12 | 2025-08-12 | Talomo South Health Center | Puan,Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | dyrosemarie71@gmail.com |  | 31291 | DOH | dy_rosemarie71@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1289 |  | GX-IV 6C 4D | 838401 | Due for Preventive Maintenance UPS Troubleshooting | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Check Line Voltages, reset connection of UPS | Perform Preventive Maintenance and UPS Troubleshooting | All Modules are operational No battery load AT UPS |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307790360619716550/6307790360619716550_signature_28.png | Rhodora Mae I. Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/6307790360619716550/6307790360619716550_signature_31.png
| 2025-08-12 | 2025-08-12 | WESTERN VISAYAS MEDICAL CENTER | Mandurriao, Iloilo | 6 | Iloilo | rHIVda | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31290 | UNOPS | wvmcrhivdacrcl@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1422 |  | GX-IV | 110023529 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307739554205802441/6307739554205802441_signature_28.png | Harry Al Villaruz RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6307739554205802441/6307739554205802441_signature_31.png
| 2025-08-12 | 2025-08-04 | Escalante City Health Office | Poblacion Escalante City | 6 | Negros Occidental | TB Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 31289 | DOH | jirehforester@gmail.com | 23 |  |  | GXIV-4D | 838523 | Installation of CPU with serial number CPH05020F08317 | Repair Power supply of the cpu. Clean the electronic board and wire of the cpu. | Save Installation Qualification Report. Save System Log Report. Save GXX file. Save Assay statistics report. | CPU functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307735955313369133/6307735955313369133_signature_28.png | Jireh Cablos | https://www.jotform.com/uploads/macarewecare/222927978372067/6307735955313369133/6307735955313369133_signature_31.png
| 2025-08-11 | 2025-08-05 | Malungon Rural Health Unit | Malungon,Sarangani | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31288 | DOH | Madameerrmie@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2306 |  | GxIV-6C | 838546 | None | Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and filter cleaning Checking of module status and erros Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307267885327590728/6307267885327590728_signature_28.png | Ermie D Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/6307267885327590728/6307267885327590728_signature_31.png
| 2025-08-11 | 2025-08-06 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Tb | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31287 | DOH | jan19.cam@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1423 |  | GxIV-6C | 838504 | 4 operational modules | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics rod Fan and filter cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307267165321027285/6307267165321027285_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/6307267165321027285/6307267165321027285_signature_31.png
| 2025-08-11 | 2025-08-07 | RHU- Magsaysay | Kakilgan St. Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31286 | DOH | jkakilgan12@gmail.com | 25 |  |  | GX-4 6 Colors | 838557 | None | None | Preventive Maintenance Signing documents | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307241551216409792/6307241551216409792_signature_28.png | Jesusa Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/6307241551216409792/6307241551216409792_signature_31.png
| 2025-08-11 | 2025-08-06 | Ramon magsaysay rural health unit | Ramon magsaysay, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31285 | DOH | rhuramonmagsaysay7@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 | 2025-08-06 | GX 4 | 838433 | None | Preventive Maintenence and data collection | Preventive Maintenence and data collection | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307086293938423057/6307086293938423057_signature_28.png | Gina Borinaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6307086293938423057/6307086293938423057_signature_31.png
| 2025-08-11 | 2025-08-07 | Midsalip Rural health unit | Midsalip, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31284 | DOH | genaliner@rocketmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 409 | 2025-08-07 | GX 4 | 837664 | Only module that gives excessive errors of 1001 and 4012. | Preventive Maintenence and data collection. Forwarded data collection to ceiphied for further assessment if for replacement. | Preventive Maintenence and data collection. | All units except module 3 that gives excessive temperature related errors are OK and ready to use. | Only module that gives excessive error of 1001 and 4012. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307085883932139849/6307085883932139849_signature_28.png | Genaline A. Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/6307085883932139849/6307085883932139849_signature_31.png
| 2025-08-11 | 2025-08-06 | Molave Rural Health unit | Molave, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31283 | DOH | aldrinacain@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 906 | 2025-08-06 | Gx 4 | 838233 | None | Preventive Maintenence and data collection. | Preventive Maintenence and data collection. | All modules are available and ready to use. | All errors are regarding error 2008. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307076533932103377/6307076533932103377_signature_28.png | Aldrin Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/6307076533932103377/6307076533932103377_signature_31.png
| 2025-08-11 | 2025-08-07 | Santolan Super Health Center | Evangelista Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31282 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1556 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6307034813618436451/6307034813618436451_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6307034813618436451/6307034813618436451_signature_31.png
| 2025-08-11 | 2025-08-11 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | Iloilo City | 6 | Iloilo | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31281 | UNOPS | mc-molecular@wvsu.edu.ph | 25 | ASSAY: Xpert HBV VL, COUNT: 103 ASSAY: Xpert HPV, COUNT: 42 ASSAY: Xpert MTB-XDR, COUNT: 3 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 54 |  | GX-IV 10C | 110021716 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6306886692326142060/6306886692326142060_signature_28.png | Karla Danielle T. Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6306886692326142060/6306886692326142060_signature_31.png
| 2025-08-11 | 2025-08-07 | Philippine Blood Center | Queson Avenue QC | NCR | NCR 2nd - Quezon City | Blood donation rm | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31280 | MACARE-HBDC | pbcmbd@doh.gov.ph | 28 |  |  | MCS+ | 15H030G | Due for maintenance Roller yellowish color need for replacement | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6306824643616354546/6306824643616354546_signature_28.png | Clariza Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/6306824643616354546/6306824643616354546_signature_31.png
| 2025-08-11 | 2025-08-07 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Blood donation rm | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31279 | MACARE-HBDC | pbcmbd@doh.gov.ph | 27 |  |  | MCS + | 15H029G | Due for maintenance Roller yellowish color need for replacement | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operarional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6306822713617457154/6306822713617457154_signature_28.png | Clariza Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/6306822713617457154/6306822713617457154_signature_31.png
| 2025-08-11 | 2025-08-07 | Philippine Red Cross QC | Kalayaan Ave. QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31278 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 27 |  | 2025-08-05 | MCS + | 02A227 | For replacement of urethane rollers | Delivered and installed urethane rollers Checked program parameters Performed functionality test Test run unit | Performed repair | Operational |  | Urethane rollers | 13229 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6306820503618005351/6306820503618005351_signature_28.png | PRC QUEZON CITY | https://www.jotform.com/uploads/macarewecare/222927978372067/6306820503618005351/6306820503618005351_signature_31.png
| 2025-08-11 | 2025-08-06 | St luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 31277 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 27 |  | 2025-08-05 | GX4L | 807155 | UPS battery unable to charge | Delivered and replaced 2pcs Battery Performed functionality test Test run unit | Performed repair ups drain battery S/N AT1000190402ST0319 | Operational C4 disabled due to 5007 error |  | Ups battery s/n 540426,540427 | 13228 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6306819873617862167/6306819873617862167_signature_28.png | Roiland Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6306819873617862167/6306819873617862167_signature_31.png
| 2025-08-11 | 2025-08-05 | Batasan Social Hygiene Clinic | Batasan Rd QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31276 | MACARE-HBDC | nicateves0708@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 26 ASSAY: Xpert HBV VL, COUNT: 49 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1071 |  | GX4D 10C | 110020697 | Due for maintenance End user encounter intermittent module not detected during starting the system | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Reset all cables Reset gatewayboard Gathered reports Test run unit | Performed maintenance | For observation | During functionality test no error encounter |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6306818753617631040/6306818753617631040_signature_28.png | Monique Teves | https://www.jotform.com/uploads/macarewecare/222927978372067/6306818753617631040/6306818753617631040_signature_31.png
| 2025-08-08 | 2025-08-08 | RIVERSIDE MEDICAL CENTER (MEDI LINX LABORATORY INC.) | Bacolod City | 6 | Negros Occidental | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31275 | COMMERCIAL | angelalorraineg@gmail.com | 20 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 19 ASSAY: Xpert HBV VL, COUNT: 132 ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF Ultra, COUNT: 369 ASSAY: Xpert HCV VL, COUNT: 53 | 2025-07-30 | GX-IV | 110005268 | FOR MODULE B3 & MODULE B4 REPLACEMENT | #NAME? | -PERFORM PREV. MAINTENANCE PROCEDURE AND CLEANING OF GX WORKSTATION -PERFORM MODULE REPLACEMENT: MODULE B3 DEF SN: 210028770 WITH NEW SN: 220038702 MODULE B4 DEF SN: 210029714 WITH NEW SN: 220038722 | EQUIPMENT IS OPERATIONAL | MODULE B1 & B2 DISABLED | GENEXPERT MODULE 6-COLOR | 13215 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304492659322158719/6304492659322158719_signature_28.png | Angela Lorraine Gamboa | https://www.jotform.com/uploads/macarewecare/222927978372067/6304492659322158719/6304492659322158719_signature_31.png
| 2025-08-08 | 2025-08-08 | Nueva Vizcaya Provincial Hospital | Bambang, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31274 | DOH | dumlaojoy4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1322 | 2025-07-30 | GXIV-6C | 838272 | AVR not powering on | Replaced AVR | Performed Maintenance Gathered Reports Performed Functionalty Test | Unit is Operational |  | AVR | 13216 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304461232028022371/6304461232028022371_signature_28.png | Gina Jalimao | https://www.jotform.com/uploads/macarewecare/222927978372067/6304461232028022371/6304461232028022371_signature_31.png
| 2025-08-08 | 2025-08-08 | Manila Doctors Hospital | U.N. Ave., Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | xcaguas@maniladoctorscomph.onmicrosoft.com |  | 31273 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 25 |  |  | GXIV-10C | 110001648 | Unit for preventive maintenance. For Xpertcheck of module B1. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck of module B1. Generated reports. | Unit is operational. Module B1, Service Required. Module B1 is for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304451559394057883/6304451559394057883_signature_28.png | Franchesca Agripa | https://www.jotform.com/uploads/macarewecare/222927978372067/6304451559394057883/6304451559394057883_signature_31.png
| 2025-08-08 | 2025-08-06 | Batangas City Health Office | Batangas City, Batangas | 3 | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31272 | MACARE-HBDC | kaimoral68@yyahoo.com | 25 |  | 2025-07-29 | GCIV-10C | 110024615 | Unit for preventive maintenance. For module replacement of module B1 and B3. | Repair and preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX Module B1 (692749) and B3 (655724) in replacement to defective B1 (210128840) and B3 (210128781). | DR#13213 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304430831422482006/6304430831422482006_signature_28.png | Catherine M.Moral | https://www.jotform.com/uploads/macarewecare/222927978372067/6304430831422482006/6304430831422482006_signature_31.png
| 2025-08-08 | 2025-08-08 | Project 7 Social Hygiene Clinic | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31271 | MACARE-HBDC | michelenaval@gmail.com | 25 |  |  | GXIV-10C | 120000577 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational B4 is not operational B4 has valve errors B4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304423533326157545/6304423533326157545_signature_28.png | Michele Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6304423533326157545/6304423533326157545_signature_31.png
| 2025-08-08 | 2025-08-08 | Balamban Rural Health Unit | Poblacion Cebu | 7 | Cebu | TB Laboratory | NO | TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 31270 | DOH | climaco.mildred@yahoo.com | 23 |  |  | GXIV-4D | 838231 | Lost communication while running error code 2123 for all module's and stuck-up cartridge for all modules. | Change cable wire from CPU to Machine. Pull out stuck-up cartridge. | Save System Log file. Save Installation Qualification report. Save Assay Statisitcs report. | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304401354213234659/6304401354213234659_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6304401354213234659/6304401354213234659_signature_31.png
| 2025-08-08 | 2025-08-08 | Medi Linx Laboratory Inc. ( CLDH) | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31269 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 | 2025-08-07 | GXIV-D 6C | 839232 | Due for preventive maintenance | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed xpertcheck Performed functionality test Gathered reports | Unit is operational except for modules B4: 738203 Failed calibration Good module B1:742501 Passed Calibration B2:893512 Passed Calibration B3:893779 Passed Calibration | Module B4 for module replacement, B1 for observation | 4 xpertcheck Kits | 13233 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304388331983678185/6304388331983678185_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6304388331983678185/6304388331983678185_signature_31.png
| 2025-08-08 | 2025-08-08 | Medi Linx Laboratory Inc. (CLDH) | Hospital Drive Brgy. San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31268 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 | 2025-08-07 | GXII-D | 830841 | Due for preventive maintenance | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed xpertcheck Performed functionality test Gathered reports | Unit is operational Module C1: 210019432 passed calibration C2: 620096 H/W Failed ( Temperature Error/ 4012) | C2 for module replacement | 1 Xpertcheck Kit | 13233 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304387861987511695/6304387861987511695_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6304387861987511695/6304387861987511695_signature_31.png
| 2025-08-08 | 2025-01-03 | Central Luzon Doctor's Hospital | Hospital Drive Brgy San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31267 | COMMERCIAL | pjlsantos@cldh.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 |  | GXIV-D | 839232 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304358641988529177/6304358641988529177_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6304358641988529177/6304358641988529177_signature_31.png
| 2025-08-08 | 2025-01-03 | Central Luzon Doctor's Hospital | Hospital Drive Brgy San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 31266 | COMMERCIAL | pjlsantos@cldh.com.ph | 24 | ASSAY: Xpert MTB/RIF, COUNT: 193 |  | GXII-D | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304358361986318544/6304358361986318544_signature_28.png | Maria Rowena Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6304358361986318544/6304358361986318544_signature_31.png
| 2025-08-08 | 2025-04-24 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31265 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 27 | ASSAY: Xpert HBV VL, COUNT: 4 ASSAY: Xpert HPV, COUNT: 27 ASSAY: Xpert SARS-CoV-2, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1567 |  | GXIV-D 10C | 110017894 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304358181989056483/6304358181989056483_signature_28.png | Thea A. Dula | https://www.jotform.com/uploads/macarewecare/222927978372067/6304358181989056483/6304358181989056483_signature_31.png
| 2025-08-08 | 2025-04-24 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31264 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 27 | ASSAY: Xpert SARS-CoV-2, COUNT: 0 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 981 |  | GXIV-D 6C | 829390 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304357811987665424/6304357811987665424_signature_28.png | Thea A. Dula | https://www.jotform.com/uploads/macarewecare/222927978372067/6304357811987665424/6304357811987665424_signature_31.png
| 2025-08-08 | 2025-04-24 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31263 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 27 | ASSAY: Xpert HBV VL, COUNT: 57 ASSAY: Xpert HPV, COUNT: 52 ASSAY: Xpert HCV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 324 |  | GXIV-D 10C | 120003587 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit us operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304357501989881207/6304357501989881207_signature_28.png | Thea A. Dula | https://www.jotform.com/uploads/macarewecare/222927978372067/6304357501989881207/6304357501989881207_signature_31.png
| 2025-08-08 | 2025-05-15 | Jose B. Lingad Memorial and General Hospital | Dolores City of Sa Fernando Pampanga | 3 | Pampanga | RHIVDA | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | jblmgh.is@gmail.com |  | 31262 | MACARE-HBDC | jblmgh.is@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 297 ASSAY: Xpert HCV VL, COUNT: 80 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1891 | 2025-05-08 | GXIV-D 10C | 110013748 | Defective UPS | Preventive maintenance UPS replacement | Performed Preventive maintenance Replacement of Defective UPS Gathered reports Performed functionality test | Unit is operational |  | UPS | 13006 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304356321981506540/6304356321981506540_signature_28.png | CHARLENE MAE LAXAMANA | https://www.jotform.com/uploads/macarewecare/222927978372067/6304356321981506540/6304356321981506540_signature_31.png
| 2025-08-08 | 2024-11-15 | Central Luzon Doctor's Hospital | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 31261 | COMMERCIAL | pjlsantos@cldh.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 ASSAY: Xpert SARS-CoV-2, COUNT: 16 |  | GXIV-D 6C | 839232 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304356121982496964/6304356121982496964_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6304356121982496964/6304356121982496964_signature_31.png
| 2025-08-08 | 2024-11-15 | Central Luzon Doctor's Hospital | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31260 | COMMERCIAL | pjlsantos@cldh.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 ASSAY: Xpert SARS-CoV-2, COUNT: 16 |  | GXII-D 6C | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304355921986834700/6304355921986834700_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6304355921986834700/6304355921986834700_signature_31.png
| 2025-08-08 | 2025-08-07 | Dasol Community Hospital | Dasol Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31259 | DOH | jeremypulido0521@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV-D | 838287 | Due for maintenance CPU not booting | Preventive maintenance CPU repair | Performed surface cleaning Perform CPU troubleshooting Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304355341984878983/6304355341984878983_signature_28.png | Jeremy Pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/6304355341984878983/6304355341984878983_signature_31.png
| 2025-08-08 | 2025-08-07 | Infanta Rural Health Unit | Infanta Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31258 | UNOPS | charissealmandres8@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 210 |  | GXIV-D 10C | 110023501 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304355151984637398/6304355151984637398_signature_28.png | Charisse A. Alferez | https://www.jotform.com/uploads/macarewecare/222927978372067/6304355151984637398/6304355151984637398_signature_31.png
| 2025-08-08 | 2025-08-07 | Bolinao Rural Health Unit | Poblacion Bolinao Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31257 | DOH | nylemae14@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 742 |  | GXIV-D | 838278 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304354731989998839/6304354731989998839_signature_28.png | Emylyn Celzo | https://www.jotform.com/uploads/macarewecare/222927978372067/6304354731989998839/6304354731989998839_signature_31.png
| 2025-08-08 | 2025-08-08 | Zamboanga del sur medical center | Zamboanga del sur medical center | 8 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31256 | DOH | limboemillie@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2934 | 2025-08-08 | GX 4 | 838204 | None | Preventive Maintenence and data collection | Preventive Maintenence and data collection. | All modules are available and ready to use. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6304292076449489651/6304292076449489651_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/6304292076449489651/6304292076449489651_signature_31.png
| 2025-08-07 | 2025-08-07 | Bicol Medical Center | Concepcion Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com | btsc@bmc.doh.gov.ph |  | 31255 | MACARE-HBDC | mongeshiela@gmail.com | 25 |  |  | MCS+ | 15H027G | MCS+ for Preventive Maintenance | Cleaned Unit surface Cleaned, rotor pump, roller, valves, centrifuge | Performed Preventive Maintenance | Unit is non operational (CRC Bios Error) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303807859817477141/6303807859817477141_signature_28.png | Shiela P. Monge | https://www.jotform.com/uploads/macarewecare/222927978372067/6303807859817477141/6303807859817477141_signature_31.png
| 2025-08-07 | 2025-08-07 | Bicol Medical Center | Concepcion Pequeña, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com | btsc@bmc.doh.gov.ph |  | 31254 | MACARE-HBDC | mongeshiela@gmail.com | 25 |  |  | MCS+ | 15H031G | MCS+ for Preventive Maintenance | Cleaned surface of unit Cleaned rotor pumps, rollers, valve, centrifuge Checked system parameters | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303805899815452894/6303805899815452894_signature_28.png | Shiela P. Monge | https://www.jotform.com/uploads/macarewecare/222927978372067/6303805899815452894/6303805899815452894_signature_31.png
| 2025-08-07 | 2025-08-06 | Bicol Medical Center | Concepcion, Pequeña, Naga | 5 | Camarines Sur | Laboratory | YES | INSTALLATION | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com | btsc@bmc.doh.gov.ph |  | 31253 | MACARE-HBDC | mongeshiela@gmail.com | 25 |  |  | MCS+ | 16K520SPG | Installed Unit | Installed MCS+(16K520SPG)tie-up pull out MCS+(11J221)tie-up Checked parameters of newly installed unit based from Calibration certificate | Installed tie-up MCS+ unit | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303802619813309633/6303802619813309633_signature_28.png | Shiela P. Monge | https://www.jotform.com/uploads/macarewecare/222927978372067/6303802619813309633/6303802619813309633_signature_31.png
| 2025-08-07 | 2025-08-05 | Metro Health Specialists Hospital Inc. | Sorsogon City | 5 | Sorsogon | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | aileenrmt@gmail.com |  | 31252 | COMMERCIAL | mhshilab@gmail.com | 20 |  |  | GX4-10C | 120002716 | Gx scheduled fro Preventive Maintenance | Moved machine to new facility inside the hospital compound Setup/installed Gx machine Cleaned carridge bays, surface Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303796999817331522/6303796999817331522_signature_28.png | Aileen V. Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6303796999817331522/6303796999817331522_signature_31.png
| 2025-08-07 | 2025-08-07 | St. Louis Molecular Diagnostic Laboratory | New Carmen, Tacurong City, Sultan Kudarat | 12 | South Cotabato | Molecular laboratory | NO | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31251 | COMMERCIAL | stlouismoleculardiagnosticslab@gmail.com | 25.3 |  |  | GxIV-10C | 110024270 | None | Installed GxIV-10C SN:110024270 Installed Desktop SN:CPH0224F491 with mouse and keyboard Installed Monitor Dell SN:CN4AH427DG Installed Pinter HP 2337 SN: 23103010550471 Checking of modules status Data collection | Genexpert machine Installation | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303612413643646321/6303612413643646321_signature_28.png | Lyra Luz Huyong | https://www.jotform.com/uploads/macarewecare/222927978372067/6303612413643646321/6303612413643646321_signature_31.png
| 2025-08-07 | 2025-08-07 | Alicia Rural Health Unit | Alicia, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31250 | DOH | labalicia.rhu@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 737 | 2025-07-23 | GXIV-6C | 838521 | UPS not powering On | Replaced UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | UPS | 13202 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303609334617329905/6303609334617329905_signature_28.png | Jeslie Raymunda Valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/6303609334617329905/6303609334617329905_signature_31.png
| 2025-08-07 | 2025-08-07 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31249 | DOH | sanjoseantiquegxsite@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3192 ASSAY: Xpert MTB-XDR, COUNT: 51 ASSAY: Xpert HIV-1 Viral Load, COUNT: 394 |  | 6C GXIV-4D | 838412 | None | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303552581113552604/6303552581113552604_signature_28.png | Hubert Spencer Francis Ang | https://www.jotform.com/uploads/macarewecare/222927978372067/6303552581113552604/6303552581113552604_signature_31.png
| 2025-08-07 | 2025-08-07 | Cagayan Valley Regional Blood Center | Tuguegarao City, Cagayan | 2 | Cagayan | Apheresis Room | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31248 | MACARE-HBDC | nvbspchdii2022@gmail.com | 22 |  |  | MCS+ | 19G146SPG | Rollers worn out | Checked Rollers Status | Cleaned MCS+ | Unit is Operational | Rollers for Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303465125547049595/6303465125547049595_signature_28.png | Baby Jane Taguinod | https://www.jotform.com/uploads/macarewecare/222927978372067/6303465125547049595/6303465125547049595_signature_31.png
| 2025-08-07 | 2025-08-07 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31247 | UNOPS | spmcrhivda@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1622 ASSAY: Xpert MTB-XDR, COUNT: 8 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1315 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 12 ASSAY: Xpert CT/NG, COUNT: 6 |  | GX-IV 10C 4D | 110023505 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303425949964183148/6303425949964183148_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6303425949964183148/6303425949964183148_signature_31.png
| 2025-08-07 | 2025-08-07 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31246 | UNOPS | spmcrhivda@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 433 ASSAY: Xpert SARS-CoV-2, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 766 |  | GX-IV 10C 4D | 110023506 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303424279961589744/6303424279961589744_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6303424279961589744/6303424279961589744_signature_31.png
| 2025-08-07 | 2025-08-07 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 31245 | DOH | materesaaliman78@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 632 | 2025-08-07 | GXIV | 838406 | n/a | n/a | Preventive Maintenance Data Collection | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6303423920817024978/6303423920817024978_signature_28.png | Ma.Teresa E. Aliman | https://www.jotform.com/uploads/macarewecare/222927978372067/6303423920817024978/6303423920817024978_signature_31.png
| 2025-08-06 | 2025-08-06 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31244 | DOH | thisisformyfamily@yahoo.com | 28 |  |  | GXIV-4-D-6C | 838286 | UPS for Troubleshooting | #NAME? | Performed UPS Troubleshooting | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302695652863984413/6302695652863984413_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/6302695652863984413/6302695652863984413_signature_31.png
| 2025-08-06 | 2025-08-06 | Lasam District Hospital | Lasam, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31243 | DOH | lasamdh.laboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 498 | 2025-07-03 | GXIV-6C | 838530 | Monitor Problem | Replaced Monitor | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Monitor | 13163 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302688189517953724/6302688189517953724_signature_28.png | Kimberly Kate T. Nicolas | https://www.jotform.com/uploads/macarewecare/222927978372067/6302688189517953724/6302688189517953724_signature_31.png
| 2025-08-06 | 2025-08-06 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31242 | MACARE-HBDC | bloodcollection.tower@gmail.com | 25 |  |  | MCS+ | 01J038 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302670137026069868/6302670137026069868_signature_28.png | Ardel Brecinio | https://www.jotform.com/uploads/macarewecare/222927978372067/6302670137026069868/6302670137026069868_signature_31.png
| 2025-08-06 | 2025-08-06 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31241 | MACARE-HBDC | bloodcollection.tower@gmail.com | 25 |  |  | MCS+ | 08B049 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302669737023668813/6302669737023668813_signature_28.png | Ardel Brecinio | https://www.jotform.com/uploads/macarewecare/222927978372067/6302669737023668813/6302669737023668813_signature_31.png
| 2025-08-06 | 2025-08-06 | Laoang Rural Health Unit | Laoang Northern Samar | 8 | Northern Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31240 | DOH | markiansarmiento727@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1052 | 2025-08-06 | GXIV | 838202 | Module B1 temperature error | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302649173292122039/6302649173292122039_signature_28.png | mark ian sarmiento | https://www.jotform.com/uploads/macarewecare/222927978372067/6302649173292122039/6302649173292122039_signature_31.png
| 2025-08-06 | 2025-08-06 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31239 | MACARE-HBDC | iegparayno@gmail.com | 28 |  |  | MCS+ | 15H033G | Bios CRC Failed | Checked Bios Battery Voltage Reset CPU Card | Reset CPU Board | Unit is not Operational. For Cpu Board Repair(CRC Bios Replacement) | CPU Board For Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302593533716388615/6302593533716388615_signature_28.png | Reggie Parayno | https://www.jotform.com/uploads/macarewecare/222927978372067/6302593533716388615/6302593533716388615_signature_31.png
| 2025-08-06 | 2025-07-31 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31238 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 251 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B2 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302587943619873874/6302587943619873874_signature_28.png | Fatima Busar | https://www.jotform.com/uploads/macarewecare/222927978372067/6302587943619873874/6302587943619873874_signature_31.png
| 2025-08-06 | 2025-07-31 | Antipolo City Health Office | M. Santos St, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31237 | UNOPS | acho.laboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert MTB-XDR, COUNT: 98 ASSAY: Xpert HIV-1 Viral Load, COUNT: 725 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Wll modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302587123619932960/6302587123619932960_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6302587123619932960/6302587123619932960_signature_31.png
| 2025-08-06 | 2025-08-06 | Medi linx Laboratory Inc. | Hospital Drive Brgy. San Vicente, Tarlac City ty | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31236 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 |  | GXII-D | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302583679479183089/6302583679479183089_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6302583679479183089/6302583679479183089_signature_31.png
| 2025-08-06 | 2025-08-06 | Medi Linx Laboratory Inc. | Hospital drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31235 | COMMERCIAL | mea.delacruz@medilinx.com.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 |  | GXIV-D 6C | 839232 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except modules B1: 742501 B4: 738203 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302581959479535894/6302581959479535894_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6302581959479535894/6302581959479535894_signature_31.png
| 2025-08-06 | 2025-08-06 | San Roque Rural Health Unit | San Roque Northern Samar | 8 | Northern Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31234 | DOH | lynnjarito@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 534 |  | GXIV | 838498 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302566794211553899/6302566794211553899_signature_28.png | LYNNN JARITO | https://www.jotform.com/uploads/macarewecare/222927978372067/6302566794211553899/6302566794211553899_signature_31.png
| 2025-08-06 | 2025-08-05 | Klinika Eastwood | Citywalk2, Eastwood, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | klinikaeastwoodlaboratory@gmail.com |  | 31233 | MACARE-HBDC | clarabaptista2786@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 98 |  | GX-IV-D-10C | 110020698 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6302505883611309650/6302505883611309650_signature_28.png | Mary claire p baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/6302505883611309650/6302505883611309650_signature_31.png
| 2025-08-05 | 2025-08-05 | National Center For Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31232 | MACARE-HBDC | iancristrono@gmail.com | 25 |  |  | GXIV-10C | 120003480 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6301870918735098676/6301870918735098676_signature_28.png | Ian Cris Tono | https://www.jotform.com/uploads/macarewecare/222927978372067/6301870918735098676/6301870918735098676_signature_31.png
| 2025-08-05 | 2025-08-05 | Ospital ng Makati | Sampaguita St, Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 31231 | COMMERCIAL | jinx01gs@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 602 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6301827480122795591/6301827480122795591_signature_28.png | Cristina Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6301827480122795591/6301827480122795591_signature_31.png
| 2025-08-05 | 2025-08-05 | Dinas Rural Health Unit | Dinas, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31230 | PBSP | bernantereynacres@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 861 | 2025-08-05 | Gx 4 | 838533 | UPS unable to work properly. | Preventive Maintenence and data collection. Troubleshoot UPS. | Preventive Maintenence and data collection. Troubleshoot UPS. | UPS unable to work properly, module 2 disabled already reported to ceiphied by macare Medicals. | UPS batteries will be replaced ASAP. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6301725917012093170/6301725917012093170_signature_28.png | Reyna Cres Bernante-Lazaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6301725917012093170/6301725917012093170_signature_31.png
| 2025-08-05 | 2025-07-30 | Western Visayas Sanitarium and Gen. Hospital | Santa barbara, iloilo | 6 | Iloilo | Laboratory | NO | MAINTENANCE Collecting data | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31229 | DOH | ruthiedequilla@gmail.com | 24 |  |  | Gx4 | 838393 | Defective module 3 temperature error Module 1 first temperature error encounter Dirty fan filter | Gathering data IQR, System log, gxx files Clean fan filter | Gathering data IQR, System log, gxx files Clean fan filter | Module replacement module 3 Under observation module 1 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6301721043635254209/6301721043635254209_signature_28.png | Ruth Dequila | https://www.jotform.com/uploads/macarewecare/222927978372067/6301721043635254209/6301721043635254209_signature_31.png
| 2025-08-05 | 2025-08-05 | Tapaz Super Health Center | Tapaz, Capiz | 6 | Capiz | Laboratory | NO | MAINTENANCE Pack Up Machine | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31228 | DOH | chadmayrs@gmail.com | 28 |  |  | 6C GXIV-4D | 837615 | For transfer to Dumalag RHU | None | #NAME? | Machine is functional. | UPS is defective |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6301669486214650165/6301669486214650165_signature_28.png | Chad May Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6301669486214650165/6301669486214650165_signature_31.png
| 2025-08-04 | 2025-08-03 | Makati Medical Center (Medilinx) | Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31226 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 174 ASSAY: Xpert CT/NG, COUNT: 497 ASSAY: Xpert HBV VL, COUNT: 49 ASSAY: Xpert HPV, COUNT: 390 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2480 ASSAY: Others, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 94 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1496 |  | GX-IV-D-10C | 110001647 | For relocation | Transferred unit to other room | Transferred GeneXpert to other room Functionality test Performed maintenance Generated reports | Unit is operational. Module B1 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6301068583614444853/6301068583614444853_signature_28.png | Aiza G. Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6301068583614444853/6301068583614444853_signature_31.png
| 2025-08-04 | 2025-08-01 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | OTHERS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com | Fujifilm | 31225 | COMMERCIAL | yvdelacruz@themedicalcity.com | 24 |  |  | Limusave MT-7500 | 3702058 | For maintenance | Performed maintenance | Performed maintenance Checked fan Cleaned instrument surface Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6301068583614617078/6301068583614617078_signature_28.png | Yehleen Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6301068583614617078/6301068583614617078_signature_31.png
| 2025-08-04 | 2025-08-01 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31227 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: G2 ASSAY: Xpert CT/NG, COUNT: 53 ASSAY: Xpert HBV VL, COUNT: 642 ASSAY: Xpert HPV, COUNT: 382 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6518 ASSAY: Xpert MTB-XDR, COUNT: 5 ASSAY: Others, COUNT: 6 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 177 ASSAY: Xpert HIV-1 Viral Load, COUNT: 954 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1435 |  | GX-XVI-D-10C | 110015767 | Modules B2, B4, C1, C3 and D1 for retry. | Run XpertCheck on modules B2, B4, C1, C3 and D1. | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck on modules B2, B4, C1, C3 and D1 Generated reports | Unit is operational. Modules A1, A3, A4 and B1 service required but operational, under observation(07/01/2025). Module D1 passed but has previous error 1001, currently disabled. Other modules passed and operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6301068593612649551/6301068593612649551_signature_28.png | Kristine Chelle Parilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6301068593612649551/6301068593612649551_signature_31.png
| 2025-08-04 | 2025-07-29 | Capitol Medical Center | Quezon Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31224 | CONSORTIUM | princeosena@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 537 |  | GXIV-10C-D | 110024553 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked Sytem's Functionality. | 3 Modules are Operational. 1 Module Disabled with 5007 Error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300948543616694686/6300948543616694686_signature_28.png | Prince Henry Osena | https://www.jotform.com/uploads/macarewecare/222927978372067/6300948543616694686/6300948543616694686_signature_31.png
| 2025-08-04 | 2025-07-30 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Road, Las Piñas | NCR | NCR 4th - Las Piñas | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31223 | COMMERCIAL | phmclp.mdl@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 4 ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 25 ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 242 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GXIV-D-6C | 831316 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | 4 Modules for Replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300948213612482822/6300948213612482822_signature_28.png | Marissa Machado | https://www.jotform.com/uploads/macarewecare/222927978372067/6300948213612482822/6300948213612482822_signature_31.png
| 2025-08-04 | 2025-07-30 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | YES | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31222 | COMMERCIAL | agalang@asianhospital.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 10 ASSAY: Xpert CT/NG, COUNT: 41 ASSAY: Others, COUNT: 91 ASSAY: Xpert HBV VL, COUNT: 171 ASSAY: Xpert HPV, COUNT: 208 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1142 ASSAY: Xpert HIV-1 Viral Load, COUNT: 143 | 2025-07-18 | GXIV-10C-D | 110014996 | Unit is due for XpertCheck. | XpertCheck. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed XpertCheck. Performed Self-test. Checked System's Functionality. Generate Reports. | XpertCheck Result: B1 - Pass B2 - Pass B3 - Pass B4 - Pass |  | XpertCheck Kit | S.I. # 100650 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300948023619336932/6300948023619336932_signature_28.png | Alyssa Galang | https://www.jotform.com/uploads/macarewecare/222927978372067/6300948023619336932/6300948023619336932_signature_31.png
| 2025-08-04 | 2025-07-31 | Tanza Wellness Center | Anterio Soriano Highway, Tanza, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31221 | MACARE-HBDC | tanzarhu.clinicallab@gmail.com | 26 | ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 126 |  | GXIV-D-10C | 120003582 | Unit is due for Maintenance. Relocate GeneXpert System. | Preventive Maintenance. | Transfer and Install GX System. Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Fikter if Clogged. Performed Self-test. Generate Reports. | 4 Modules are Operational. Relocate from Social Hygiene Clinic to Laboratory. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300946883617271326/6300946883617271326_signature_28.png | MERRIAM JOY G. SIBBALUCA | https://www.jotform.com/uploads/macarewecare/222927978372067/6300946883617271326/6300946883617271326_signature_31.png
| 2025-08-04 | 2025-07-30 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31219 | COMMERCIAL | kcaangay@asianhospital.com | 25 |  |  | MCS+ | 01F116 | Fluid Detected Error. | Troubleshoot and Cleaning. | Cleaned Centrifuge and Spill Detector Sensor. Cleaned Surface Area and Pump Rotors. Checked A/D and Calibration Value. | Unit is not Operational. Recommendation for Replacement of Centrifuge Assembly. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300946883618047768/6300946883618047768_signature_28.png | Keano Riggs Caañgay | https://www.jotform.com/uploads/macarewecare/222927978372067/6300946883618047768/6300946883618047768_signature_31.png
| 2025-08-04 | 2025-07-30 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31220 | COMMERCIAL | kcaangay@asianhospital.com | 25 |  |  | MCS+ | 10B129 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Rollers, Centrifuge and Surface Area. Checked A/D and Diagnostic Values. Performed Simulation Test. | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300946883618764072/6300946883618764072_signature_28.png | Keano Riggs Caañgay | https://www.jotform.com/uploads/macarewecare/222927978372067/6300946883618764072/6300946883618764072_signature_31.png
| 2025-08-04 | 2025-08-01 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Blood donation | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 31218 | MACARE-HBDC | bloodbankeamc@gmail.com | 26 |  |  | MCS + | 15H044 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Test run | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300900783613422460/6300900783613422460_signature_28.png | EAMC | https://www.jotform.com/uploads/macarewecare/222927978372067/6300900783613422460/6300900783613422460_signature_31.png
| 2025-08-04 | 2025-08-01 | Best Diagnostic Corp. | Masikap st., QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31217 | COMMERCIAL | bestdiagnosticlaboratory@gmail.com | 23 | ASSAY: Xpert HBV VL, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX4 D | 110008529 | Due for maintenance Last xpertcheck feb 7,2024 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational Unit need to performed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300900423612768286/6300900423612768286_signature_28.png | Rose Marie Balisi | https://www.jotform.com/uploads/macarewecare/222927978372067/6300900423612768286/6300900423612768286_signature_31.png
| 2025-08-04 | 2025-08-01 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 31216 | MACARE-HBDC | eamcmolecularlab@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 1176 ASSAY: Xpert HPV, COUNT: 108 ASSAY: Xpert SARS-CoV-2, COUNT: 705 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 ASSAY: Xpert SARS-CoV-2, COUNT: 488 |  | GX4 D 10C | 110016781 | Due for maintenance B3 disabled encounter intermittent error 5007 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B3 for observation Enable B3 module B1,B2,B4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300900063619164959/6300900063619164959_signature_28.png | EAMC Molecular Lab | https://www.jotform.com/uploads/macarewecare/222927978372067/6300900063619164959/6300900063619164959_signature_31.png
| 2025-08-04 | 2025-07-31 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31215 | DOH | dianeruiz147@gmail.com | 28 |  |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300795873619511955/6300795873619511955_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6300795873619511955/6300795873619511955_signature_31.png
| 2025-08-04 | 2025-07-31 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31214 | COMMERCIAL | annabellelcatbagan@gmail.com | 28 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module A1 is disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300794853613593797/6300794853613593797_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6300794853613593797/6300794853613593797_signature_31.png
| 2025-08-04 | 2025-07-31 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31213 | DOH | serly_reyes@yahoo.com | 28 |  |  | GXIV-4-D-6C | 838516 | For UPS Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Partially Operational UPS for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300792773611813514/6300792773611813514_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6300792773611813514/6300792773611813514_signature_31.png
| 2025-08-04 | 2025-07-30 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31212 | DOH | gbmmbg@yahoo.com | 28 |  |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300791433615955251/6300791433615955251_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/6300791433615955251/6300791433615955251_signature_31.png
| 2025-08-04 | 2025-07-30 | Mangatarem Rural Health Unit | Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31211 | DOH | slasheiji1988new@gmail.com | 26 |  |  | GXIV-4-D-6C | 838284 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Pending Module 1 Replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300789153613466768/6300789153613466768_signature_28.png | Ken Lee Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/6300789153613466768/6300789153613466768_signature_31.png
| 2025-08-03 | 2025-08-01 | Libon Rural Health Unit | Velasco St., Zone 1, Libon | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31210 | DOH | zh.surge@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 835 |  | GX4-6C | 838397 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational All 4 modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6300038353529816564/6300038353529816564_signature_28.png | Cedric Ravago | https://www.jotform.com/uploads/macarewecare/222927978372067/6300038353529816564/6300038353529816564_signature_31.png
| 2025-08-01 | 2025-08-01 | St Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | TB Lab | NO | REPAIR | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31209 | MACARE-HBDC | laboratory@slec.ph | 25 |  | 2025-07-23 | GX XVI | 110022374 | Modules failed XpertCheck and need replacement | Replaced defective modules (210122077 > 210075235, 210121937 > 210024944 ) | #NAME? | Unit is operational |  | Genexpert 10-color module (x2) | DR# 13201 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6298426385428421395/6298426385428421395_signature_28.png | Amy allayban | https://www.jotform.com/uploads/macarewecare/222927978372067/6298426385428421395/6298426385428421395_signature_31.png
| 2025-08-01 | 2025-07-31 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | nch.jeff22@gmail.com |  | 31208 | MACARE-HBDC | zarabellehilario@yahoo.com | 25 |  |  | MCS | 17H632SPG | Noisy vibration when centrifuge is spinning | Checked centrifuge alignment Checked all screw connected ing centrifuge Aligned centrifuge Checked centrifuge balance | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6298306126474841449/6298306126474841449_signature_28.png | Zarabelle Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/6298306126474841449/6298306126474841449_signature_31.png
| 2025-08-01 | 2025-07-31 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Stem Cell | NO | MAINTENANCE | OTHERS | Noel Floresca |  | florescanoel77@gmail.com |  | Airinspace | 31207 | MACARE-HBDC | lcp.mdctl2024@gmail.com | 26 |  |  | Plasmair | 2006-T 1188 | Due for maintenance | Cleaned unit Cleaned filter remove dust Checked terminal and connection Checked condition of filters Performed functionality test Test run unit | Performed maintenance Dirty Pre filter | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6298305446473204509/6298305446473204509_signature_28.png | Jennifer Panuelos | https://www.jotform.com/uploads/macarewecare/222927978372067/6298305446473204509/6298305446473204509_signature_31.png
| 2025-08-01 | 2025-07-31 | Philippine Red Cross QC | Kalayaan Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31206 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 26 |  |  | MCS + | 02A227 | Due formaintenance Rollers slight yellowish color | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Unit operational For replacement of urethane rollers |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6298304866474746571/6298304866474746571_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/6298304866474746571/6298304866474746571_signature_31.png
| 2025-08-01 | 2025-08-01 | Panabo City Health Office | Poblacion, Panabo | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31205 | DOH | nadineeust67@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2003 |  | GX-IV 6C 4D | 838295 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6298295728445159367/6298295728445159367_signature_28.png | Nadine Gemma Eustaquio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6298295728445159367/6298295728445159367_signature_31.png
| 2025-08-01 | 2025-08-01 | Balamban Rural Health Unit | Poblacion Balamban | 7 | Cebu | TB laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 31204 | DOH | balambanrhu45@gmail.com | 23 |  |  | GXIV-4 | 838231 | Due for monthly maintenance. | Save installation qualification reports. Save System log reports. Save GXX file | Plunger rod maintenance. Modules cleaning. | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6298228127424196960/6298228127424196960_signature_28.png | Loraine Mae D. Tagpuno | https://www.jotform.com/uploads/macarewecare/222927978372067/6298228127424196960/6298228127424196960_signature_31.png
| 2025-07-31 | 2025-07-31 | Kapatagan Rural Health Unit | Kapatagan, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31203 | DOH | aganub2005@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 450 |  | 6C GXIV-4D | 838224 | None | None | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297598111654500936/6297598111654500936_signature_28.png | Angelita Paradero | https://www.jotform.com/uploads/macarewecare/222927978372067/6297598111654500936/6297598111654500936_signature_31.png
| 2025-07-31 | 2025-07-31 | Pilar Rural Health Unit | Calongay, Pilar | 5 | Sorsogon | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31202 | DOH | manalangmarkramir@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1883 |  | GX4-6C | 838266 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational All 4 modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297581373235538209/6297581373235538209_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/6297581373235538209/6297581373235538209_signature_31.png
| 2025-07-31 | 2025-07-31 | Lala Rural Health Unit | Lala, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31201 | DOH | villagonzalojoyv@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 815 |  | 6C GXIV-4D | 838229 | None | None | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297572537118490907/6297572537118490907_signature_28.png | Maria Glenda Villagonzalo | https://www.jotform.com/uploads/macarewecare/222927978372067/6297572537118490907/6297572537118490907_signature_31.png
| 2025-07-31 | 2025-07-31 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31200 | DOH | buhanginlaboratory@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1092 |  | GX-IV 6C 4D | 838232 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297502349448679650/6297502349448679650_signature_28.png | Catherine B. Deocampo,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6297502349448679650/6297502349448679650_signature_31.png
| 2025-07-31 | 2025-07-31 | Kauswagan Rural Health Unit | Kauswagan, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31199 | DOH | dextertinoy@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 317 |  | 6C GXIV-4D | 838402 | CPU CMOS battery failing | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297476905231978090/6297476905231978090_signature_28.png | Dexter Tinoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6297476905231978090/6297476905231978090_signature_31.png
| 2025-07-31 | 2025-07-30 | Diit District Health Center | Diit Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 31198 | DOH | dalepenalosa189@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV | 838421 | Module B1 error | #NAME? | #NAME? | Unit is Operational (Module B3 disabled) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297396180123239262/6297396180123239262_signature_28.png | Dale Penalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6297396180123239262/6297396180123239262_signature_31.png
| 2025-07-31 | 2025-07-30 | National Childrens Hospital | E.rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | nch.jeff22@gmail.com |  | 31197 | MACARE-HBDC | arellanojohneric09@gmail.com | 28 |  |  | MCS + | 17H632SPG | Encounter fluid detection Spillage dry dirt bottom of centrifuge | Checked unit Dismantle centrifuge Cleaned centrifuge Checked program parameters Performed functionality test Performed simulation running | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297389470116869343/6297389470116869343_signature_28.png | John Eric L. Arellano | https://www.jotform.com/uploads/macarewecare/222927978372067/6297389470116869343/6297389470116869343_signature_31.png
| 2025-07-31 | 2025-07-25 | Labuan General hospital | Labuan Zamboanga city | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31196 | DOH | irenecasee@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2849 | 2025-07-25 | GX 4 | 838568 | None | Preventive Maintenence and data collection | Preventive Maintenence and data collection. | All modules modules are available and ready to use | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297388836467431051/6297388836467431051_signature_28.png | Irene Casepe | https://www.jotform.com/uploads/macarewecare/222927978372067/6297388836467431051/6297388836467431051_signature_31.png
| 2025-07-31 | 2025-07-29 | St Lukes Medical Center QC | E.Rodriguez | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 31195 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 26 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 424 ASSAY: Xpert HPV, COUNT: 419 ASSAY: Xpert MTB/RIF Ultra, COUNT: 240 ASSAY: Xpert HCV VL, COUNT: 33 ASSAY: Xpert HIV-1 Viral Load, COUNT: 235 |  | GX4 L | 807155 | For maintenance UPS unable to store supply defective battery | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | C1,C2,C3 operational C4 disabled due to error 5007 For replacement of ups battery |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297388790115827611/6297388790115827611_signature_28.png | NICOLE JOYCE U. TY | https://www.jotform.com/uploads/macarewecare/222927978372067/6297388790115827611/6297388790115827611_signature_31.png
| 2025-07-31 | 2025-07-29 | St Lukes Medical Center QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 31194 | COMMERCIAL | megapo@stlukes.com.ph | 26 | ASSAY: Xpert CT/NG, COUNT: 198 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2582 |  | GX4D | 839234 | For maintenance B3 march 6,2023 and B4 march 14,2023 disabled due to failed xpertcheck last | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B1 and B2 in use B3,B4 disabled failed in xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297388240112602319/6297388240112602319_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6297388240112602319/6297388240112602319_signature_31.png
| 2025-07-31 | 2025-07-31 | Maluso Rural Health Unit | Maluso Basilan | BARMM | Basilan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31193 | DOH | sharminahussin@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 554 | 2025-07-31 | GX 4 | 838835 | Module 1 h/w failed. | Module 1 undergo troubleshooting by running some sample and provide data for the error. Preventive Maintenence and data collection. | Module 1 undergo troubleshooting by running some sample and provide data for the error. Preventive Maintenence and data collection. | Module 1 gave a proper result after running a sample test. | Advised to lower the room temperature more if possible. Average module temperatures are slightly higher that the required running temperature. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6297387596463661651/6297387596463661651_signature_28.png | Sharmina A. Hussin | https://www.jotform.com/uploads/macarewecare/222927978372067/6297387596463661651/6297387596463661651_signature_31.png
| 2025-07-30 | 2025-07-30 | Hinaplanon Health Center | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31192 | DOH | villacortaamy50@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 822 |  | 6C GXIV-4D | 838227 | Module B2 H/W failure | None | #NAME? | Module B2 H/W fail for reporting. Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6296691087479097231/6296691087479097231_signature_28.png | Amy Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6296691087479097231/6296691087479097231_signature_31.png
| 2025-07-30 | 2025-07-30 | Tomas Claudio Health Center | Camus St., Poblacion, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31191 | DOH | arleen_p_alconera@yahoo.com | 28.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 984 |  | GX-IV 6C 4D | 838492 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6296670899446837888/6296670899446837888_signature_28.png | Arleen Alconera,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6296670899446837888/6296670899446837888_signature_31.png
| 2025-07-30 | 2025-07-30 | Lugait Rural Health Unit | Lugait, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31190 | DOH | lugaitmaternity@gmail.com | 27 |  |  | 6C GXIV-4D | 838562 | CPU not turning on | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6296645929159785607/6296645929159785607_signature_28.png | ZOE PHIL CAGAS, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6296645929159785607/6296645929159785607_signature_31.png
| 2025-07-30 | 2025-07-30 | Tubungan Rural Health Unit | Tubungan, Iloilo | 6 | Iloilo | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31189 | UNOPS | sheryltacadao@gmail.com | 28 |  |  | Gx4 | 110023534 | Stucked-up carthridge module 1 and 3 | Removing stucked-up carthridge Clean Fan Filter IQR, System log data gathering | Removing stucked-up carthridge Clean Fan Filter IQR, System log data gathering | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6296526560418094332/6296526560418094332_signature_28.png | SHERYL TACADAO | https://www.jotform.com/uploads/macarewecare/222927978372067/6296526560418094332/6296526560418094332_signature_31.png
| 2025-07-30 | 2025-07-29 | Rosario Superhealth Center | Dr. Sixto Antonio Avenue, Rosario, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31187 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1334 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6296453253612445355/6296453253612445355_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6296453253612445355/6296453253612445355_signature_31.png
| 2025-07-30 | 2025-07-29 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco Ave, Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31188 | MACARE-HBDC | peachmango318@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1578 |  | GX-IV-D-10C | 110024245 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6296453253612661399/6296453253612661399_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6296453253612661399/6296453253612661399_signature_31.png
| 2025-07-29 | 2025-07-29 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | dolly.banares@redcross.org.ph |  | 31186 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  | 2025-07-25 | MCS+ | 12D063 | For installation of photo-optic assembly. | Repair | Performed installation of photo-optic assembly. Adjusted bowl optics fullscale value. Performed simulation running. | Unit is now operational and ready to use. |  | Photo-optic assembly | DR#13210 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6295898152121207373/6295898152121207373_signature_28.png | Lalaine Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6295898152121207373/6295898152121207373_signature_31.png
| 2025-07-29 | 2025-07-29 | Likhaan Center For Women's Health | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31185 | MACARE-HBDC | avega@likhaan.org | 25 |  | 2025-07-24 | GXIV-10C | 110025174 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: Failed B2: Pass |  | XpertCheck Kit | 13207 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6295796846317092607/6295796846317092607_signature_28.png | Arnold Vega | https://www.jotform.com/uploads/macarewecare/222927978372067/6295796846317092607/6295796846317092607_signature_31.png
| 2025-07-29 | 2025-07-29 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 31184 | DOH | claritomirasol@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 | 2025-07-29 | GXIV | 838548 | N/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6295737116018932796/6295737116018932796_signature_28.png | mirasol Clarito | https://www.jotform.com/uploads/macarewecare/222927978372067/6295737116018932796/6295737116018932796_signature_31.png
| 2025-07-29 | 2025-07-29 | Laguindingan Rural Health Unit | Laguindingan, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31183 | DOH | chellesambaan@gmail.com | 27 |  |  | 6C GXIV-4D | 838575 | None | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6295712269646249940/6295712269646249940_signature_28.png | Rochelle Sambaan | https://www.jotform.com/uploads/macarewecare/222927978372067/6295712269646249940/6295712269646249940_signature_31.png
| 2025-07-28 | 2025-07-28 | Corazon Locsin Montelibano Memorial Regional Hospital | Lacson St., Bacolod City | 6 | Negros Occidental | MOLECULAR LAB | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | mollab@clmmrh.doh.gov.ph |  | 31182 | UNOPS | alainjuayang@yahoo.com | 19 | ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert HBV VL, COUNT: 82 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1717 |  | GX-IV 10C | 110020743 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6294811677119266839/6294811677119266839_signature_28.png | Alain Juayang, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6294811677119266839/6294811677119266839_signature_31.png
| 2025-07-28 | 2025-07-23 | MMC (Medilinx) | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Molecular | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31181 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 115 ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 294 ASSAY: Xpert HBV VL, COUNT: 62 ASSAY: Xpert HPV, COUNT: 150 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1251 ASSAY: Others, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 76 ASSAY: Xpert HIV-1 Viral Load, COUNT: 731 |  | GX16D (10C) | 110026448 | For relocation End user encounter leakage of water in laboratory ceiling and the machine accidentally got wet | Performed relocation Cleaned interior and exterior of the machine Test run unit | Performed relocation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6294733043612286588/6294733043612286588_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6294733043612286588/6294733043612286588_signature_31.png
| 2025-07-26 | 2025-07-25 | St. Luke's Medical Center - BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | Cellular Immunology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31180 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 184 ASSAY: Xpert CT/NG, COUNT: 680 ASSAY: Xpert HBV VL, COUNT: 63 ASSAY: Xpert HPV, COUNT: 948 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2414 ASSAY: Xpert MTB-XDR, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 44 ASSAY: Xpert HIV-1 Viral Load, COUNT: 313 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6293003448438162700/6293003448438162700_signature_28.png | MA. MIKHAELA RAPADAS | https://www.jotform.com/uploads/macarewecare/222927978372067/6293003448438162700/6293003448438162700_signature_31.png
| 2025-07-26 | 2025-07-25 | St. Luke's Medical Center - BGC | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Cellular Immunology | YES | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31179 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 41 ASSAY: Xpert CT/NG, COUNT: 154 ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert HPV, COUNT: 232 ASSAY: Xpert MTB/RIF Ultra, COUNT: 479 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 130 | 2025-07-21 | GX-IV-D | 839231 | Module A3 for replacement | Replaced module A3 | Replaced module A3 Performed maintenance Generated reports | Unit is operational. All modules operational. |  | Old: A3/866793 New: A3/210058796 | SI# 100657 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6293003128437209010/6293003128437209010_signature_28.png | MA. MIKHAELA RAPADAS | https://www.jotform.com/uploads/macarewecare/222927978372067/6293003128437209010/6293003128437209010_signature_31.png
| 2025-07-26 | 2025-07-25 | St. Luke's Medical Center - BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | Cellular Immunology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31178 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 184 ASSAY: Xpert CT/NG, COUNT: 680 ASSAY: Xpert HBV VL, COUNT: 63 ASSAY: Xpert HPV, COUNT: 948 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2414 ASSAY: Xpert MTB-XDR, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 44 ASSAY: Xpert HIV-1 Viral Load, COUNT: 313 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6293002898436856075/6293002898436856075_signature_28.png | MA. MIKHAELA RAPADAS | https://www.jotform.com/uploads/macarewecare/222927978372067/6293002898436856075/6293002898436856075_signature_31.png
| 2025-07-26 | 2025-07-25 | St. Luke's Medical Center - BGC | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Cellular Immunology | YES | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31177 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 41 ASSAY: Xpert CT/NG, COUNT: 154 ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert HPV, COUNT: 232 ASSAY: Xpert MTB/RIF Ultra, COUNT: 479 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 130 | 2025-07-21 | GX-IV-D | 839231 | Module A3 for replacement | Replaced module A3 | Replaced module A3 Performed maintenance Generated reports | Unit is operational. All modules operational. |  | Old: A3/866793 New: A3/210058796 | SI# 100657 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6293002898437623709/6293002898437623709_signature_28.png | MA. MIKHAELA RAPADAS | https://www.jotform.com/uploads/macarewecare/222927978372067/6293002898437623709/6293002898437623709_signature_31.png
| 2025-07-25 | 2025-07-16 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31176 | DOH | sherryannemarrero@gmail.com | 24 |  |  | GXIV-4-D-6C | 838281 | #NAME? | #NAME? | #NAME? | Unit is Partially Operational Module B3 for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6292397323296422591/6292397323296422591_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6292397323296422591/6292397323296422591_signature_31.png
| 2025-07-25 | 2025-07-16 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31175 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 17 |  | 2025-07-16 | GXIV-4-D-10C | 110016112 | For Calibration | #NAME? | Performed Calibration | #NAME? |  | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6292396309874239149/6292396309874239149_signature_28.png | SYXERLYNE DANG-AY | https://www.jotform.com/uploads/macarewecare/222927978372067/6292396309874239149/6292396309874239149_signature_31.png
| 2025-07-24 | 2025-07-18 | Capitol Medical Center | Q. Ave, Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31174 | COMMERCIAL | laboratory@capitolmedical.com.ph | 25 |  |  | MCS+ | 07C002 | Centrifuge with Strange Noise. | Troubleshoot. | Cleaned Centrifuge. Removed Debris from Broken Kit. Performed Simulation Run. | Unit is Operational. Centrifuge is Okay. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6291541693617048064/6291541693617048064_signature_28.png | REYCART PAUL ANDOLOY | https://www.jotform.com/uploads/macarewecare/222927978372067/6291541693617048064/6291541693617048064_signature_31.png
| 2025-07-24 | 2025-07-24 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31173 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | For checking of MCS+. Centrifuge not closing. For checking of bowl optics value. | Troubleshooting and repair. | Performed troubleshooting and repair. Adjusted centrifuge sensor on desired height. Performed simulation running. Adjusted low and hi on photo-optics assembly for offset and fullscale. | Unable to adjust photo-optics assembly. For replacement of photo-optics assembly. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6291506882122931882/6291506882122931882_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6291506882122931882/6291506882122931882_signature_31.png
| 2025-07-24 | 2025-07-24 | Mini Forest Health Center (District C) | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31172 | DOH | miniforestlaboratory@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 556 |  | GX-IV 6C 4D | 838507 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6291496276211941906/6291496276211941906_signature_28.png | Lolijell O. Torralba,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6291496276211941906/6291496276211941906_signature_31.png
| 2025-07-24 | 2025-07-23 | Norala District Hospital | Norala, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31171 | DOH | noralaDHlaboratory@gmail.com | 23.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1202 |  | GxIV-6C | 838543 | None | Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of module status and problem Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6291455316865137042/6291455316865137042_signature_28.png | Reziel D. Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/6291455316865137042/6291455316865137042_signature_31.png
| 2025-07-24 | 2025-07-24 | Poro Rural Health Unit | Poro, Cebu | 7 | Cebu | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31170 | DOH | avilacharmaine6@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 463 |  | 6C GXIV-4D | 838279 | None | #NAME? | #NAME? | Machine is functional. Waiting for module B2 replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6291444023024175732/6291444023024175732_signature_28.png | Charmaine Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/6291444023024175732/6291444023024175732_signature_31.png
| 2025-07-24 | 2025-07-24 | Padre Burgos Rural Health Unit | Padres Burgos Southern Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31169 | UNOPS | keishaelezze14@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 399 | 2025-07-24 | GXIV | 110023514 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6291412062023265584/6291412062023265584_signature_28.png | KEISHA ELEZZE C. BATALON | https://www.jotform.com/uploads/macarewecare/222927978372067/6291412062023265584/6291412062023265584_signature_31.png
| 2025-07-24 | 2025-07-17 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | dolly.banares@redcross.org.ph |  | 31167 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for calibration. Unit for maintenance. | Calibration and preventive maintenance. | Performed calibration. Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6291399632124865523/6291399632124865523_signature_28.png | Lalaine Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6291399632124865523/6291399632124865523_signature_31.png
| 2025-07-24 | 2025-07-17 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | dolly.banares@redcross.org.ph |  | 31168 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 16K521SPG | Unit for calibration. Unit for preventive maintenance. | Preventive maintenance and calibration. | Performed calibration. Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6291399632126035586/6291399632126035586_signature_28.png | Lalaine Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6291399632126035586/6291399632126035586_signature_31.png
| 2025-07-24 | 2025-06-04 | Philippine Red Cross Benguet | La Trinidad, Benguet | CAR | Benguet | Blood Bank | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31166 | MACARE-HBDC | benguet@redcross.org.ph | 25 |  | 2025-06-03 | MCS+ | 17H636 | CPU Board for replacement | Replaced CPU board Calibrated unit | #NAME? | Unit is operational |  | MCS+ CPU Board Assembly SN: MM061004515 | 13068 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6291388713612433480/6291388713612433480_signature_28.png | LYKAINE KATE BAGANO | https://www.jotform.com/uploads/macarewecare/222927978372067/6291388713612433480/6291388713612433480_signature_31.png
| 2025-07-23 | 2025-07-23 | Matanao Rural Health Unit | Poblacion Matanao | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31165 | DOH | pamsustento@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 458 |  | GX-IV 6C 4D | 838456 | For Preventive Maintenance UPS Troubleshooting Data CSV Gathering | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Check Line Voltage, reset cable network connection | Perform Preventive Maintenance UPS Troubleshooting and Data CSV Gathering | All Modules are operational UPS no charge during battery mode | Reported to Jotform for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6290548190329671185/6290548190329671185_signature_28.png | Dominador s. Castro III | https://www.jotform.com/uploads/macarewecare/222927978372067/6290548190329671185/6290548190329671185_signature_31.png
| 2025-07-23 | 2025-07-22 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31164 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 62 ASSAY: Xpert CT/NG, COUNT: 53 ASSAY: Xpert HBV VL, COUNT: 641 ASSAY: Xpert HPV, COUNT: 377 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6488 ASSAY: Xpert MTB-XDR, COUNT: 5 ASSAY: Others, COUNT: 6 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 177 ASSAY: Xpert HIV-1 Viral Load, COUNT: 943 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1459 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6290509813616829940/6290509813616829940_signature_28.png | Kim San Sebastian | https://www.jotform.com/uploads/macarewecare/222927978372067/6290509813616829940/6290509813616829940_signature_31.png
| 2025-07-23 | 2025-07-23 | Cebu TB Reference Lab - SRP | SRP cebu city | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31163 | DOH | asmantiquilla.pbsp@gmail.com | 23 |  |  | Gx4 | 838263 | Monthly PM Clean fan filter Plunger rod maintenance | Monthly PM Clean fan filter Plunger rod maintenance | Monthly PM Clean fan filter Plunger rod maintenance | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6290461967252073797/6290461967252073797_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6290461967252073797/6290461967252073797_signature_31.png
| 2025-07-22 | 2025-07-22 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31162 | DOH | lornsabaigar@gmail.com | 28 |  |  | GXIV | 838409 | n/a | n/a | Monthly Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6289754526549713406/6289754526549713406_signature_28.png | lorna abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/6289754526549713406/6289754526549713406_signature_31.png
| 2025-07-21 | 2025-07-16 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31159 | DOH | renlizros11@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1629 |  | GxIV-6C | 838570 | None | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of module status and Errors Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6289103206862589038/6289103206862589038_signature_28.png | Renee Liz Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6289103206862589038/6289103206862589038_signature_31.png
| 2025-07-21 | 2025-07-15 | Sarangani Health Care Facility | Alabel, Sarangani Province | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31160 | DOH | ronayadao@yahoo.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 958 ASSAY: Xpert HIV-1 Viral Load, COUNT: 32 |  | GxIV-6C | 838542 | None | Perform Genexpert machine Cleaning and maintenance Module cleaning Optics and Plunger rod cleaning Fan filter cleaning Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6289103206865600944/6289103206865600944_signature_28.png | Ronora Yadao | https://www.jotform.com/uploads/macarewecare/222927978372067/6289103206865600944/6289103206865600944_signature_31.png
| 2025-07-21 | 2025-07-15 | Malapatan Rural Health Unit | Malapatan, Sarangani Province | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31161 | DOH | kamilahnichol@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 653 |  | GxIV-6C | 838419 | None | Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6289103206866246260/6289103206866246260_signature_28.png | Kamilah Nichol Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/6289103206866246260/6289103206866246260_signature_31.png
| 2025-07-21 | 2025-07-17 | Esperanza Rural Health Unit | Poblacion, Esperanza | 5 | Masbate | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31158 | DOH | floramarmoncada@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 689 |  | GX4-6C | 838268 | Monitor for replacement | Replaced Monitor(CN4552011X) Cleaned cartridge bays, plunger Generate Data | Replaced monitor | Unit is Operational | UPS for observation. battery drained |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288901102415300336/6288901102415300336_signature_28.png | Floramar Moncada | https://www.jotform.com/uploads/macarewecare/222927978372067/6288901102415300336/6288901102415300336_signature_31.png
| 2025-07-21 | 2025-07-21 | Mindanao central sanitarium | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 31157 | DOH | ybioamie@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2176 | 2025-07-21 | GX 4 | 821214 | None | Preventive Maintenence and data collection | Preventive Maintenence and Data collection. Re-orient end user regarding xpertcheck calibration. | Module 1 and 3 disabled due to excessive errors or 2032, 2120 and 1001 and 2120 respectively. | Module 1 already reported to cephied directly as per enduser. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288885363018739344/6288885363018739344_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/6288885363018739344/6288885363018739344_signature_31.png
| 2025-07-21 | 2025-07-18 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31156 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 02D166 | Due for maintenance Urethane roller need for replacement | Cleaned unit Checked rollers Checked and test tube sensor Checked and test keypad and lcd display. Checked program parameters Performed functionality test. | Performed maintenance | Urethane rollers need to be replaced |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288867453616434289/6288867453616434289_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/6288867453616434289/6288867453616434289_signature_31.png
| 2025-07-21 | 2025-07-18 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31155 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 08H188 | Due for maintenance Urethane rollers need for replacement | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Urethane roller need to replace |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288867023614566187/6288867023614566187_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/6288867023614566187/6288867023614566187_signature_31.png
| 2025-07-21 | 2025-07-21 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhlabresults2@yahoo.com |  | 31154 | COMMERCIAL | cduhlabresults2@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1104 |  | GX4 | 110008524 | Module 3 defective Serial Number: 210019546 | Replacement of New Module SR : 220038034 Generate IQR, Sys Log, Assay Statistics, Gxx files | Replacement of New Module SR : 220038034 Generate IQR, Sys Log, Assay Statistics, Gxx files | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288809073447104096/6288809073447104096_signature_28.png | Cebu Doctors Hospital University TB DOTS | https://www.jotform.com/uploads/macarewecare/222927978372067/6288809073447104096/6288809073447104096_signature_31.png
| 2025-07-21 | 2025-04-24 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31153 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 27 | ASSAY: Xpert HBV VL, COUNT: 57 ASSAY: Xpert HPV, COUNT: 52 ASSAY: Xpert HCV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 324 |  | GXIV-D 10C | 120003587 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit us operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288772241275693949/6288772241275693949_signature_28.png | Thea A. Dula | https://www.jotform.com/uploads/macarewecare/222927978372067/6288772241275693949/6288772241275693949_signature_31.png
| 2025-07-21 | 2025-04-24 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31152 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 27 | ASSAY: Xpert HBV VL, COUNT: 4 ASSAY: Xpert HPV, COUNT: 27 ASSAY: Xpert SARS-CoV-2, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1567 |  | GXIV-D 10C | 110017894 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288772021275370731/6288772021275370731_signature_28.png | Thea A. Dula | https://www.jotform.com/uploads/macarewecare/222927978372067/6288772021275370731/6288772021275370731_signature_31.png
| 2025-07-21 | 2025-04-24 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31151 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 27 | ASSAY: Xpert SARS-CoV-2, COUNT: 0 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 981 |  | GXIV-D 6C | 829390 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288771771273949325/6288771771273949325_signature_28.png | Thea A. Dula | https://www.jotform.com/uploads/macarewecare/222927978372067/6288771771273949325/6288771771273949325_signature_31.png
| 2025-07-21 | 2025-05-15 | Jose B. Lingad Memorial and General Hospital | Dolores City of Sa Fernando Pampanga | 3 | Pampanga | RHIVDA | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | jblmgh.is@gmail.com |  | 31150 | MACARE-HBDC | jblmgh.is@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 297 ASSAY: Xpert HCV VL, COUNT: 80 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1891 | 2025-05-08 | GXIV-D 10C | 110013748 | Defective UPS | Preventive maintenance UPS replacement | Performed Preventive maintenance Replacement of Defective UPS Gathered reports Performed functionality test | Unit is operational |  | UPS | 13006 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288771001273965080/6288771001273965080_signature_28.png | CHARLENE MAE LAXAMANA | https://www.jotform.com/uploads/macarewecare/222927978372067/6288771001273965080/6288771001273965080_signature_31.png
| 2025-07-21 | 2025-07-18 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | princeszvillaran13@gmail.com |  | 31149 | COMMERCIAL | rmcbloodbank@gmail.com | 22 |  |  | MCS+ | 15H045G | For calibration | Performed calibration | Performed maintenance Performed calibration Simulation run Checked rotor pumps Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6288698473617527440/6288698473617527440_signature_28.png | Mae Sarah Alojado | https://www.jotform.com/uploads/macarewecare/222927978372067/6288698473617527440/6288698473617527440_signature_31.png
| 2025-07-18 | 2025-07-18 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31148 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  | 2025-10-10 | GXIV-10C | 110008530 | Modules B1, B3 for XpertCheck | Performed XpertCheck | #NAME? | Unit is operational Pending module replacement for B2 B1: Failed B3: Failed | Overall Status: B1, B2, B3 Failed XC | XpertCheck Kit | 13180 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6286318915012252565/6286318915012252565_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6286318915012252565/6286318915012252565_signature_31.png
| 2025-07-18 | 2025-07-18 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31147 | MACARE-HBDC | phc_bb@yahoo.com | 27 |  | 2025-07-03 | MCS + | 20168 | For replacement of line sensor assy | Replaced and installed line sensor assy. Adjust line sensor value 3071 Performed functionality test Performed simulation running | Performed repair | Operational |  | Line sensor assy | 13164 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6286268614215390470/6286268614215390470_signature_28.png | Phil Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6286268614215390470/6286268614215390470_signature_31.png
| 2025-07-18 | 2025-07-17 | Medilinx (MMC) | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31146 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 110 ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 271 ASSAY: Xpert HBV VL, COUNT: 60 ASSAY: Xpert HPV, COUNT: 133 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1190 ASSAY: Others, COUNT: 6 ASSAY: Xpert HCV VL, COUNT: 74 ASSAY: Xpert HIV-1 Viral Load, COUNT: 706 |  | GX16D (10C) | 110026448 | C3 module encounter 4016 error. | Gathered report Reset gateway board Enable C3 module Cleaned plunger | Performed troubleshooting | C3 for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6286267824215446876/6286267824215446876_signature_28.png | Aiza G. Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6286267824215446876/6286267824215446876_signature_31.png
| 2025-07-18 | 2025-07-18 | Bureau of Plant Industry | Lucena City, Quezon Province | 4A | Quezon | Plant Quarantine | YES | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | pqsquezon4a@gmail.com |  | 31145 | COMMERCIAL | gmromano97@gmail.com | 25 |  | 2025-07-18 | Stemi 305 | 3946033491 | N/A | N/A | Installed Stemi 305 Performed functionality test | Unit is operational | SI 100647 | Spot Illuminator Stemi 305 SN: 3946033491 Stand SN: 3943038072 | DR13194 | Carlos Leniel Origenes | https://www.jotform.com/uploads/macarewecare/222927978372067/6286265455427393936/6286265455427393936_signature_28.png | Gilbert Romano | https://www.jotform.com/uploads/macarewecare/222927978372067/6286265455427393936/6286265455427393936_signature_31.png
| 2025-07-17 | 2025-07-16 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | bloodbank@asianhospital.com |  | 31144 | COMMERCIAL | kcaangay@asianhospital.com | 25 |  |  | MCS+ | 01F116 | Leak Detected At Centrifuge. | Troubleshoot. | Checked and Cleaned Centrifuge. Run Simulation Test. | Simulation Test is Okay. Unit is Operational. | Centrifuge for Observation after Cleaning. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6285467995517970917/6285467995517970917_signature_28.png | Keano Riggs Caangay | https://www.jotform.com/uploads/macarewecare/222927978372067/6285467995517970917/6285467995517970917_signature_31.png
| 2025-07-17 | 2025-07-17 | Davao Reproductive Health and Wellness Center | Jacinto Street, Poblacion District Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31143 | DOH | rhwclab@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 409 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2086 |  | GX-IV 6C 4D | 838262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6285395084313852306/6285395084313852306_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6285395084313852306/6285395084313852306_signature_31.png
| 2025-07-17 | 2025-07-17 | Baybay City Health Office | Baybay City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31142 | DOH | sajecopooh20@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1243 | 2025-07-17 | GXIV | 838555 | N/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6285357084533355746/6285357084533355746_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/6285357084533355746/6285357084533355746_signature_31.png
| 2025-07-17 | 2025-07-17 | Tampakan Rural Health Unit | Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31141 | DOH | lotisglegada@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 797 |  | Gx4-6color | 838544 | UPS - no back-up power | Replaced/rapair UPS battery Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Data collection | Perform UPS battery repair Perform Genexpert machine preventive maintenance | 4 functional module UPS back up power restored | Repair UPS back up power problem |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6285351409868566759/6285351409868566759_signature_28.png | Lotis G. Legada | https://www.jotform.com/uploads/macarewecare/222927978372067/6285351409868566759/6285351409868566759_signature_31.png
| 2025-07-16 | 2025-07-16 | Carmen Health Center | Brgy. Carmen, Cagayan de Oro | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31140 | DOH | juvymadarang@yahoo.com | 25 |  |  | GX-4 6 Colors | 838442 | UPS unable to back-up. | UPS Battery Replacement | Replacing UPS battery with the new one. Preventive Maintenance | All modules and units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6284565203377998633/6284565203377998633_signature_28.png | Juvy Madarang | https://www.jotform.com/uploads/macarewecare/222927978372067/6284565203377998633/6284565203377998633_signature_31.png
| 2025-07-16 | 2025-07-16 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | UPS | 31139 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  | 2025-07-03 | ACCUPOWER | , | For battery replacement | Replaced and installed battery AT1000190410ST0397 548680, 548679 NEW BATTERY AT1000190123ST0280 548669, 548670 NEW BATTERY AT1000181128ST0025 548678, 548666 NEW BATTERY AT1000190123ST0100 548667, 548668 NEW BATTERY AT1000190123ST0182 548665, 548664 NEW BATTERY AT1000190402ST0370 548683, 548675 NEW BATTERY | Performed repair | Operational |  | 12 pcs 12v,7AH Battery | G 8013 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6284479003613659314/6284479003613659314_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6284479003613659314/6284479003613659314_signature_31.png
| 2025-07-16 | 2025-07-16 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 31138 | DOH | burauenlab@yahoo.com | 28 |  | 2025-07-16 | GXIV | 838551 | n/a | n/a | Monthly Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6284386016226547085/6284386016226547085_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6284386016226547085/6284386016226547085_signature_31.png
| 2025-07-15 | 2025-07-15 | Diit District Health Center | Diit , Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31137 | DOH | dalepenalosa189@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1109 | 2025-07-15 | GXIV | 838421 | n/a | n/a | Monthly Preventive Maintenance | Unit is Operational - Module B3 (Defective) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6283649129485332720/6283649129485332720_signature_28.png | Dale D. Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6283649129485332720/6283649129485332720_signature_31.png
| 2025-07-15 | 2025-07-15 | Monkayo Rural Health Unit | Poblacion Monkayo | 11 | Davao de Oro | TB-DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31136 | DOH | kendralois730@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1230 |  | GX-IV 6C 4D | 838487 | System Unit CPU not working For Preventive Maintenance And Data CSV Gathering | Reset power connection, check line voltages and replace the CMOS Battery of the Motherboard Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform System Unit Troubleshooting, Preventive Maintenance and Data Gathering | System unit is working, All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6283644432328712142/6283644432328712142_signature_28.png | Karen S. Anuta,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6283644432328712142/6283644432328712142_signature_31.png
| 2025-07-15 | 2025-07-13 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | Chhc_bloodbank@chonghua.com.ph |  | 31135 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 11J219 | Display button problem. | Reset CPU board in the MCS+ Check UPS supply voltage Check line voltage | Reset CPU board in the MCS+ Check UPS supply voltage Check line voltage | Unit is not operational. | FUNTE OSMENA, CEBU |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6283634843446876764/6283634843446876764_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6283634843446876764/6283634843446876764_signature_31.png
| 2025-07-15 | 2025-07-12 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31134 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 11J219 | Display button problem. | Reset CPU board in the MCS+ | Reset CPU board in the MCS+ | Unit is operational. | FUNTE OSMENA, CEBU |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6283632963441179081/6283632963441179081_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6283632963441179081/6283632963441179081_signature_31.png
| 2025-07-14 | 2025-07-08 | Columbio Rural Health Unit | Columbio, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31133 | DOH | rochelle.toledo19@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 918 | 2025-07-08 | GxIV-6C | 838556 | Module B2 H/W failed | Replace defective module SN:862572 with new module SN:699053 Perform module B3 troubleshooting Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Data collection | Genexpert machine Preventive maintenance Genexpert machine module repair Genexpert machine module B3 troubleshooting | 4 functional modules |  | Genexpert 6-color module SN:862572 | DR #: 13111 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6283140456864069427/6283140456864069427_signature_28.png | Rochelle S Toledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6283140456864069427/6283140456864069427_signature_31.png
| 2025-07-14 | 2025-07-14 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | nch.jeff22@gmail.com |  | 31132 | MACARE-HBDC | michaeliza1205@gmail.com | 26 |  | 2025-07-10 | MCS + | 17H632SPG | Due for Maintenance and Calibration | Cleaned unit Checked program parameters Checked line sensor value Checked and replaced rollers Adjust dpm,spm,centrifuge,weigher value Performed calibration Performed functionality test Performed simulation running | Performed maintenance and calibration | Unit operational Indorse to the end user 6pcs urethane rollers waste material |  | 6 pcs urethane rollers | 100600/13173 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6282913609806854501/6282913609806854501_signature_28.png | Michaeliza Bravante | https://www.jotform.com/uploads/macarewecare/222927978372067/6282913609806854501/6282913609806854501_signature_31.png
| 2025-07-14 | 2025-07-09 | Cardinal Santos Medical Center | Wilson St San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 31128 | COMMERCIAL | csmclab@csmc.ph | 22 |  |  | MCS+ | 08M461 | For check-up/Maintenance | Performed maintenance | Performed maintenance Run simulation Adjust centrifuge cover Checked valves Functionality test Verified calibration parameters | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763612813281/6282619763612813281_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763612813281/6282619763612813281_signature_31.png
| 2025-07-14 | 2025-07-10 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31130 | MACARE-HBDC | estyaragon2019@gmail.com | 24 |  |  | GX-IV-D-10C | 110024073 | Error 2005 on module B4 | Performed maintenance Checked module B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763614152612/6282619763614152612_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763614152612/6282619763614152612_signature_31.png
| 2025-07-14 | 2025-07-10 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31127 | MACARE-HBDC | estyaragon2019@gmail.com | 22 |  |  | GX-IV-D-10C | 110024072 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763614378272/6282619763614378272_signature_28.png | Estrelita s. Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763614378272/6282619763614378272_signature_31.png
| 2025-07-14 | 2025-07-09 | Cardinal Santos Medical Center | Wilson St San Juan | NCR | NCR 2nd - San Juan | Microbiology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31131 | COMMERCIAL | fetalver.glenn@medilinx.com.ph | 22 |  |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763617207230/6282619763617207230_signature_28.png | Glenn Mark G. Fetalver | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763617207230/6282619763617207230_signature_31.png
| 2025-07-14 | 2025-07-08 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | INSTALLATION | OTHERS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com | Miltenyi Biotec | 31129 | COMMERCIAL | dms-ipmm@themedicalcity.com | 22 |  |  | gentleMACS Octo Dissociator with Heaters | 1576 | For installation | Installed GentleMACS Octo Dissociator System | Installed GentleMACS unit with 8 heaters Installed UPS Power on test Functionality test | Unit is operational. | Tie-up Unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763618894553/6282619763618894553_signature_28.png | Aldrich De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/6282619763618894553/6282619763618894553_signature_31.png
| 2025-07-11 | 2025-07-11 | Bago City Health Office | Bago City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 31126 | DOH | swabteambago@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2864 |  | GX4 6C | 838486 | System Date/Time problem | Preventive Maintenance | #NAME? | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6280193301522065537/6280193301522065537_signature_28.png | Glenly Casona | https://www.jotform.com/uploads/macarewecare/222927978372067/6280193301522065537/6280193301522065537_signature_31.png
| 2025-07-10 | 2025-07-09 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31125 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 11J219 | Failure to detect the lock of the centrifuge during procedure. Purple valve error during diagnostic test. | Replacement of CPU to its original parts due failure to detect the lock of the centrifuge. Replacement of Pneumatic Assembly due to purple valve error. | Replacement of CPU to its original parts due failure to detect the lock of the centrifuge. Replacement of Pneumatic Assembly due to purple valve error. | Unit is operational. |  |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6279603366132884609/6279603366132884609_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6279603366132884609/6279603366132884609_signature_31.png
| 2025-07-10 | 2025-07-10 | Makati Medical Center | Amorsolo st,Makati | NCR | NCR 4th - Makati | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | BIOMED@makatimed.net.ph |  | 31124 | COMMERCIAL | annark3ene@yahoo.com | 28 |  | 2025-07-07 | MCS + | 97H330 | Due for calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked value of,dpm,spm,centrifuge,weigher Replaced 6pcs rollers Performed functionality test Performed calibration Performed simulation running | Performed calibration | Operational |  | 6pcs Urethane roller | 100588 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6279384003492355330/6279384003492355330_signature_28.png | Mark Fil P Dalisay | https://www.jotform.com/uploads/macarewecare/222927978372067/6279384003492355330/6279384003492355330_signature_31.png
| 2025-07-10 | 2025-07-10 | Makati Medical Center | Amorsolo st Makati | NCR | NCR 4th - Makati | CRM | NO | MAINTENANCE | OTHERS | Noel Floresca |  | florescanoel77@gmail.com | BIOMED@makatimed.net.ph | Miltenyi | 31123 | COMMERCIAL | AlcorDeeond.Padilla@makatimed.net.ph | 29 |  |  | Clinimacs | 797 | Due for maintenance | Cleaned unit Test keypad Checked bearings Checked and test clamp Checked program parameters Performed functionality | Performed maintenancd | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6279381793496762065/6279381793496762065_signature_28.png | Alcor Deeond F. Padilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6279381793496762065/6279381793496762065_signature_31.png
| 2025-07-10 | 2025-07-10 | Sta Cruz Rural Health Unit | Poblacion Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31122 | DOH | chendalagan72@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 496 |  | GX-IV 6C 4D | 838505 | For Preventive Maintenance and Data CSV Gathering | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6279361053522720566/6279361053522720566_signature_28.png | Patrisha Dante,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6279361053522720566/6279361053522720566_signature_31.png
| 2025-07-10 | 2025-07-10 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31121 | DOH | ccmclaboratory01@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2213 |  | 6C GXIV-4D | 838561 | Due for monthly PM | -Perform preventive maintenance; plunger rod, optic slot, cartridge bay, fan filter | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6279162471884945044/6279162471884945044_signature_28.png | Hannah Jane Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/6279162471884945044/6279162471884945044_signature_31.png
| 2025-07-09 | 2025-07-09 | Manila Doctors Hospital | U.N. Ave, Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | xcaguas@maniladoctorscomph.onmicrosoft.com |  | 31120 | COMMERCIAL | molecular.lab@maniladoctors.com.ph | 25 |  |  | GXIV-10C | 110001648 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module B2, B3, and B4 passed Xpertcheck. Module B1 for retry. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6278536416857092555/6278536416857092555_signature_28.png | Neriz Magne Burdan | https://www.jotform.com/uploads/macarewecare/222927978372067/6278536416857092555/6278536416857092555_signature_31.png
| 2025-07-09 | 2025-07-08 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 31119 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 31 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 402 ASSAY: Xpert HPV, COUNT: 422 ASSAY: Xpert MTB/RIF Ultra, COUNT: 220 ASSAY: Xpert HCV VL, COUNT: 32 ASSAY: Xpert HIV-1 Viral Load, COUNT: 214 |  | GX4 L | 807155 | For relocation from molec to cellular immunology room | Performed re location Re installed and set up unit Test run unit | Relocation | C1,C2,C3 operational C4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6278340508255336073/6278340508255336073_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6278340508255336073/6278340508255336073_signature_31.png
| 2025-07-09 | 2025-07-09 | Socorro Rural Health Unit | Brgy. Navarro, Socorro, Surigao del Norte | 13 | Surigao del Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31118 | UNOPS | dynamite.tux@gmail.com | 30 |  | 2025-06-17 | GX-4 10 Colors | 110023523 | Module B3 not detected. | Module Replacement | Replacing module B3 with SN: 210120303 to SN: 210092507 Preventive Maintenance | All modules and units are functional. |  | Certified pre-owned GeneXpert 10- colors module | 13094 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6278326000019803040/6278326000019803040_signature_28.png | James P. Anastacio | https://www.jotform.com/uploads/macarewecare/222927978372067/6278326000019803040/6278326000019803040_signature_31.png
| 2025-07-08 | 2025-07-08 | Claver Rural Health Unit | Brgy. Tayaga, Claver | 13 | Surigao del Norte | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31117 | DOH | andoy_joyce5@yahoo.com | 29 |  | 2025-06-20 | GX 4 6 colors | 838514 | Module B1 found errors in temperature | Module Replacement | Replacing Module B1 with SN: 862373 to SN: 606244 Preventive Maintenance Data Collection | All modules and units are functional |  | Certified pre-owned GeneXpert Module | 13105 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6277656720509787494/6277656720509787494_signature_28.png | Joyce E. Andoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6277656720509787494/6277656720509787494_signature_31.png
| 2025-07-08 | 2025-07-08 | Mandaluyong Social Hygiene Clinic | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31116 | UNOPS | mandaluyongshc.laboratory@gmail.com | 24 |  |  | GXIV-10C | 110025111 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6277618154235106992/6277618154235106992_signature_28.png | Ron Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/6277618154235106992/6277618154235106992_signature_31.png
| 2025-07-08 | 2025-07-08 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 31115 | UNOPS | laureanjo@yahoo.com | 27 |  |  | Gx4 | 110023513 | N/A | Monthly PM Clean module plunger Clean Fan Filter | Monthly PM Clean module plunger Clean Fan Filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6277491314327037701/6277491314327037701_signature_28.png | Laurean jo cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6277491314327037701/6277491314327037701_signature_31.png
| 2025-07-07 | 2025-07-01 | Polomolok Rural Health Unit | Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31114 | DOH | jmmirasolrmtrmt@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2189 |  | GxIV-6C | 838265 | None | Perform Genexpert cleaning and maintenance Checking of module status and problem Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6277059506867348740/6277059506867348740_signature_28.png | JUNE MICHELE MIRASOL | https://www.jotform.com/uploads/macarewecare/222927978372067/6277059506867348740/6277059506867348740_signature_31.png
| 2025-07-07 | 2025-07-03 | Calumpang Rural Health Unit | Brgy. Calumpang General Santos City, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31113 | DOH | jan19.cam@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1318 |  | GxIV-6C | 838504 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6277055596861083451/6277055596861083451_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/6277055596861083451/6277055596861083451_signature_31.png
| 2025-07-07 | 2025-07-03 | Calumpang Rural Health Unit | Brgy. Calumpang General Santos City, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31112 | DOH | jan19.cam@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1318 |  | GxIV-6C | 838504 | None | Perform Genexpert machine cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of module status and problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6277055336861213456/6277055336861213456_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/6277055336861213456/6277055336861213456_signature_31.png
| 2025-07-07 | 2025-07-03 | Nazareth Health Center | Brgy. Nazareth, Cagayan de Oro | 10 | Misamis Oriental | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31111 | DOH | nanethbtolentino@gmail.com | 27 |  |  | GX-4 6 Colors | 838269 | Stuck Cartridge at Module B3 | Stuck Cartridge Removal | Removing Stuck Cartridge Manually. Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6276717065229341629/6276717065229341629_signature_28.png | Naneth B. Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/6276717065229341629/6276717065229341629_signature_31.png
| 2025-07-07 | 2025-07-03 | Cagayan de Oro City Health Office | Cagayan de Oro City | 10 | Misamis Oriental | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31110 | DOH | chocdodiagnosticlab@gmail.com | 27 |  |  | GX-4 6 Colors | 838441 | UPS not working | UPS Replacement | Replacing UPS and UPS Battery. Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6276714855224599546/6276714855224599546_signature_28.png | Chariza Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/6276714855224599546/6276714855224599546_signature_31.png
| 2025-07-07 | 2025-07-05 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31109 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 11J219 | Display and Panel board problem Draw, Prime and Return, Modify button not functional. | CPU board replacement Calibration of the MCS+ and functional test. | CPU board replacement Calibration of the MCS+ and functional test. | MCS+ unit is operational and under observation. | San Isidro, Talisay City Cebu |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6276664663443221808/6276664663443221808_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6276664663443221808/6276664663443221808_signature_31.png
| 2025-07-07 | 2025-07-04 | Sta Rosa Community District Hospital | Sta Rosa, Olango, Lapulapu City | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31108 | DOH | ariestumz@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 350 |  | GX4 | 838534 | Gather data CSV files, Gxx files, IQR, System log, Conduct Preventive Maintenance, clean up filter and Clean Icore | Gather data CSV files, Gxx files, IQR, System log, Conduct Preventive Maintenance, clean up filter and Clean Icore | Gather data CSV files, Gxx files, IQR, System log, Conduct Preventive Maintenance, clean up filter and Clean Icore | All modules are operational. | San Isidro, Talisay City Cebu |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6276662043444588316/6276662043444588316_signature_28.png | Ariston Tumulak, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6276662043444588316/6276662043444588316_signature_31.png
| 2025-07-04 | 2025-07-04 | ST. PAUL'S HOSPITAL OF ILOILO | Iloilo City | 6 | Iloilo | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31107 | COMMERCIAL | sphibloodbank69@gmail.com | 22 |  |  | MCS+ | 98B020 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6274286241226754541/6274286241226754541_signature_28.png | Mariejol G.Dureza | https://www.jotform.com/uploads/macarewecare/222927978372067/6274286241226754541/6274286241226754541_signature_31.png
| 2025-07-04 | 2025-07-04 | Davao Regional Medical Center Blood Bank | Apokon, Tagum City | 11 | Davao del Norte | Blood Bank | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 31106 | COMMERCIAL | bloodbank@drmc.gov.ph | 24.6 |  |  | Apheresis MCS+ | 21M389SPG | Protocol Card does not match thr disposable ID For Preventive Maintenance and Calibration | Reset Connection for Disposable Card and check for other concern Check for line sensor voltages Check for SPM, DPM Parameter Run Cycle Test for Functionality Test | Perform Troubleshooting, Calibration and Preventive Maintenance | Machine is Operational | Machine for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6274269704668362037/6274269704668362037_signature_28.png | Orlene Tricia L. Salomon | https://www.jotform.com/uploads/macarewecare/222927978372067/6274269704668362037/6274269704668362037_signature_31.png
| 2025-07-04 | 2025-07-04 | Calubian Rural Health Unit | Calubian, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31105 | DOH | mjcabardo6@gmail.com | 27 |  |  | 6C GXIV-4D | 838399 | None | -Perform preventive maintenance; plunger rod, fan filter, optic slot, cartridge bay | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6274229615516825018/6274229615516825018_signature_28.png | Might Joy Cabardo | https://www.jotform.com/uploads/macarewecare/222927978372067/6274229615516825018/6274229615516825018_signature_31.png
| 2025-07-04 | 2025-07-04 | PONTEVEDRA PRIMARY CARE FACILITY | Pontevedra, Capiz | 6 | Capiz | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31104 | DOH | rtabngjr@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 642 | 2025-06-20 | GX-IV | 838296 | FOR MODULE B2 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM B2 DEF SN:808282 MODULE REPLACEMENT WITH NEW SN : 688254 | EQUIPMENT IS OPERATIONAL |  | Certified pre-owned GeneXpert Module | 13108 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6274113679188492918/6274113679188492918_signature_28.png | Reynaldo Tabang Jr. , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6274113679188492918/6274113679188492918_signature_31.png
| 2025-07-03 | 2025-07-03 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31103 | DOH | luzvimindanavidad2@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1018 | 2025-06-20 | 6C GXIV-4D | 838297 | Module B4 for replacement | -Replace defective module B4/785115 with module SN: 693458 | #NAME? | All modules are functional. |  | 900-0491RS Certified pre-owned GeneXpert module SN. 693458 | DR no. 13110 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6273424299217111285/6273424299217111285_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6273424299217111285/6273424299217111285_signature_31.png
| 2025-07-03 | 2025-07-03 | Cong. Sineon Toribio Memorial Hospital | Carmen Bohol | 7 | Bohol | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 31102 | DOH | himaya.vale@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 650 | 2025-07-03 | GXIV | 838394 | #NAME? | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6273366274919387478/6273366274919387478_signature_28.png | himaya vale | https://www.jotform.com/uploads/macarewecare/222927978372067/6273366274919387478/6273366274919387478_signature_31.png
| 2025-07-03 | 2025-07-03 | Candijay Community Hospital | Candijay Bohol | 7 | Bohol | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 31101 | DOH | angelinaredulla@yahoo.com.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 308 | 2025-07-03 | GXIV | 838396 | n/a | #NAME? | #NAME? | unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6273299494981131109/6273299494981131109_signature_28.png | Angelina O.Redulla | https://www.jotform.com/uploads/macarewecare/222927978372067/6273299494981131109/6273299494981131109_signature_31.png
| 2025-07-03 | 2025-07-03 | Teodoro Galagar District Hospital | Jagna Bohol | 7 | Bohol | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31100 | DOH | venusaton4@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 779 | 2025-07-03 | GXIV | 838292 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6273183656717494774/6273183656717494774_signature_28.png | VENUS C. ATON | https://www.jotform.com/uploads/macarewecare/222927978372067/6273183656717494774/6273183656717494774_signature_31.png
| 2025-07-02 | 2025-07-02 | CDO Polymedic Medical Plaza | National Highway, Cagayan de Oro | 10 | Misamis Oriental | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31099 | COMMERCIAL | cpmpmolecularlab@gmail.com | 25 |  |  | GX-4 10 Colors | 110025567 | None | Installation and demo | Installation of New GeneXpert 4 modules 10 colors machine and peripherals. Set-up and demo of genexpert software. Installation of GeneXpert Assays. | All modules and units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272612629202423878/6272612629202423878_signature_28.png | Mary Lou Baluarte | https://www.jotform.com/uploads/macarewecare/222927978372067/6272612629202423878/6272612629202423878_signature_31.png
| 2025-07-02 | 2025-07-02 | WESTERN VISAYAS MEDICAL CENTER | Mandurriao, Iloilo City | 6 | Iloilo | rHIVda | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31098 | UNOPS | wvmcrhivdacrcl@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1211 |  | GX-IV 10C | 110023529 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272580937941901726/6272580937941901726_signature_28.png | Jessa Krista Riina Jalandoni | https://www.jotform.com/uploads/macarewecare/222927978372067/6272580937941901726/6272580937941901726_signature_31.png
| 2025-07-02 | 2025-07-01 | Cebu South Medical Center | San Isidro, Talisay City Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | csmcmicro@gmail.com |  | 31097 | DOH | csmcgenexpert@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1308 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert HBV VL, COUNT: 53 |  | GX4 | 838559 | Module 2 Serial Number : 838559 defective | Replacement of new module Serial Number: 838559 | Replacement of new module Serial Number: 838559 Generate IQR, Assay Stat, System Log, and Gxx files | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272494030227425343/6272494030227425343_signature_28.png | Kristel Sheen Miranda, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6272494030227425343/6272494030227425343_signature_31.png
| 2025-07-02 | 2025-07-02 | Tagum City Health Office | Poblacion, Tagum | 11 | Davao del Norte | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31096 | DOH | cristiarafernandez@gmail.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 3219 | 2025-06-20 | GX-IV 6C 4D | 838553 | Module 2 encounter multiple temperature related errors For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Replaced the Defective Module SN:760219 with new Module SN: 646566 | Perform Module Replacement and Preventive Maintenance | All Modules are operational |  | 900-0491rs | DR 13106 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272471815229676238/6272471815229676238_signature_28.png | Cristina Tiara A. Fernandez,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6272471815229676238/6272471815229676238_signature_31.png
| 2025-07-02 | 2025-07-02 | Cagayan Valley Regional Blood Center | Carig, Tuguegarao, Cagayan | 2 | Cagayan | Apheresis Room | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31095 | MACARE-HBDC | nvbspchdii2022@gmail.com | 26 |  |  | MCS+ | 19G146SPG | Due for Maintenance Due for Calibration | Performed Calibration Performef Maintenance Performef Functionality Test | Check Line Sensor Check Weigher Check DPM Check SPN Check Bowl Optics Adjust Line Sensor | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272451878035940296/6272451878035940296_signature_28.png | Baby Jane Taguinod | https://www.jotform.com/uploads/macarewecare/222927978372067/6272451878035940296/6272451878035940296_signature_31.png
| 2025-07-02 | 2025-07-02 | Cagayan Valley Regional Blood Center | Carig, Tuguegarao, Cagayan | 2 | Cagayan | Apheresis Room | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31094 | MACARE-HBDC | nvbspchdii2022@gmail.com | 26 |  |  | MCS+ | 19G146SPG | Due for Maintenance Due for Calibration | Performed Calibration Performef Maintenance Performef Functionality Test | Check Line Sensor Check Weigher Check DPM Check SPN Check Bowl Optics Adjust Line Sensor | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272451868039144428/6272451868039144428_signature_28.png | Baby Jane Taguinod | https://www.jotform.com/uploads/macarewecare/222927978372067/6272451868039144428/6272451868039144428_signature_31.png
| 2025-07-02 | 2025-07-01 | Escalante City Health Office | Escalante City | 7 | Negros Occidental | TB Laboratory | NO | TROUBLESHOOTING CPU Problem | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 31093 | DOH | jirehforester@gmail.com | 23 |  |  | GXIV-4D | 838523 | CPU cannot power ON. | Pull-out CPU with serial number CPH0520F08317. | Check all wire's at the back of the machine and inside the CPU. | Not okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272436435848438663/6272436435848438663_signature_28.png | JIREH M. CABUS | https://www.jotform.com/uploads/macarewecare/222927978372067/6272436435848438663/6272436435848438663_signature_31.png
| 2025-07-02 | 2025-07-02 | Cagayan Valley Medical Center | Carig, Tuguegarao, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | HAEMONETICS | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31092 | MACARE-HBDC | cvmcbloodbank888@gmail.com | 26 |  |  | MCS+ | 15HO33G | Due for Maintenance | Performed Maintenance | Checked Rollers Checked Machine Parameters Cleaned MCS+ | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272400081707352616/6272400081707352616_signature_28.png | Reggie Parayno | https://www.jotform.com/uploads/macarewecare/222927978372067/6272400081707352616/6272400081707352616_signature_31.png
| 2025-07-02 | 2025-07-02 | USWAG Iloilo City Molecular Laboratory | Plaza Libertad Legaspi Dela Rama 5000, Iloilo City | 6 | Iloilo | Laboratory | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31091 | MACARE-HBDC | uswagiloilomolecularlaboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2025-06-24 | GX-IV 10C | 110026513 | -NEW INSTALLATION GX SN: 110026513 GeneXpert Desktop SN: CPH1024F55093 | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | FOR SCHEDULE OF FINAL INSTALLATION(AFTER LABORATORY CONSTRUCTION) AND END-USER TRAINING | GeneXpert IV R2, 4-Modules, Desktop, 10-Color | 13118 | Joseph Acebedo | https://www.jotform.com/uploads/macarewecare/222927978372067/6272360717941514713/6272360717941514713_signature_28.png | ANNABELLE ROSARIO S. BLANCAFLOR,RMT,MPG | https://www.jotform.com/uploads/macarewecare/222927978372067/6272360717941514713/6272360717941514713_signature_31.png
| 2025-07-02 | 2025-07-01 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31088 | COMMERCIAL | bloodbank@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Verified calibration parameters Functionality test Checked rotor pumps | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272355133611433400/6272355133611433400_signature_28.png | Kristina Enteria | https://www.jotform.com/uploads/macarewecare/222927978372067/6272355133611433400/6272355133611433400_signature_31.png
| 2025-07-02 | 2025-07-01 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31090 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 22 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Cleaned centrifuge bowl Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272355133613510264/6272355133613510264_signature_28.png | Gracia Tina Angara | https://www.jotform.com/uploads/macarewecare/222927978372067/6272355133613510264/6272355133613510264_signature_31.png
| 2025-07-02 | 2025-07-01 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 31089 | COMMERCIAL | moleculardx@themedicalcity.com | 22 |  |  | GX-XVI-D-10C | 110015767 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. Modules A2, B3, C2, C4, D2, D3 and D4 passed. Modules A1, A3, A4, and B1 service required. Modules B2, B4, C1, C3, D1 for retry insufficient cartridge. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272355133618487028/6272355133618487028_signature_28.png | Xchosa Kay C. Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/6272355133618487028/6272355133618487028_signature_31.png
| 2025-07-02 | 2025-07-02 | Macare Medicals Inc | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Macare in house | NO | REPAIR | OTHERS | Noel Floresca |  | florescanoel77@gmail.com |  | Accupower | 31087 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  |  | 1kva ups | AT1000190123ST0396 | No power drain 2 pcs battery | Replaced and installed 2pcs battery S/N 548671,548672 | Performed repair and troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272277593618291360/6272277593618291360_signature_28.png | Macare Medicals Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/6272277593618291360/6272277593618291360_signature_31.png
| 2025-07-02 | 2025-07-02 | Macare In House | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | OTHERS | Noel Floresca |  | florescanoel77@gmail.com |  | Accupower | 31086 | MACARE-HBDC | nfloresca@macare-medicals.com | 27 |  |  | 1kva ups | AT1000190123ST0128 | No power drain 2 pcs battery | Replaced and installed 2pcs battery S/N 548674,548673 | Performed troubleshooting and repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272275333615407554/6272275333615407554_signature_28.png | Macare Medicals Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/6272275333615407554/6272275333615407554_signature_31.png
| 2025-07-02 | 2025-07-01 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31085 | COMMERCIAL | monet.laxina@yahoo.com.ph | 26 |  |  | GX4 L | 804205 | Due for maintenance D1,D2 reported failed xpertcheck last 11/15/24 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Unit not in use D2,D3,D4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272261093615364727/6272261093615364727_signature_28.png | Simonette R. Laxina | https://www.jotform.com/uploads/macarewecare/222927978372067/6272261093615364727/6272261093615364727_signature_31.png
| 2025-07-02 | 2025-07-01 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31084 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  |  | MCS+ | 11H204 | Due for Calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,centrifuge Calibrate dpm,spm,weigher Performed functionality test Performed simulation running | Performed calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272260743616000666/6272260743616000666_signature_28.png | PHC Blood Bank | https://www.jotform.com/uploads/macarewecare/222927978372067/6272260743616000666/6272260743616000666_signature_31.png
| 2025-07-02 | 2025-07-01 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31083 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  |  | MCS+ | 20168 | Due for calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,centrifuge, Calibrate weigher,DPM,SPM Performed calibration Performed functionality test Performed simulation running | Performed calibration | Need to replace line sensor assy Unit still can use | Line sensor unable to reach desire value of 3071. Line sensor value max 2690 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6272259873613376185/6272259873613376185_signature_28.png | PHC Blood Bank | https://www.jotform.com/uploads/macarewecare/222927978372067/6272259873613376185/6272259873613376185_signature_31.png
| 2025-07-01 | 2025-07-01 | Baggao Rural Health Unit | San Jose, Baggao, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31082 | DOH | baggaorhulaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1561 | 2025-07-01 | GXIV-6C | 838531 | Module 1 and 4 Error | Performed Module Replacement B1(862455 - 657480) B4(862508 - 662430) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexepet Modules | 13107 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271660680813657467/6271660680813657467_signature_28.png | Jeany Anne Galera | https://www.jotform.com/uploads/macarewecare/222927978372067/6271660680813657467/6271660680813657467_signature_31.png
| 2025-07-01 | 2025-07-01 | Quezon City Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31081 | MACARE-HBDC | molecularlab.cityhealth@quezoncity.gov.ph | 25 |  |  | GXIV-10C | 120003581 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271598768605907075/6271598768605907075_signature_28.png | Cyrell Arcueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6271598768605907075/6271598768605907075_signature_31.png
| 2025-07-01 | 2025-07-01 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Rd, Las Piñas | NCR | NCR 4th - Las Piñas | Blood Bank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | noel.feranil67@gmail.com |  | 31080 | COMMERCIAL | uphdmc_lab@yahoo.com | 25 |  | 2025-06-26 | MCS+ | 11J220 | Unit is due for Calibration. | Preventive Maintenance. | Replaced Pump Rollers. Cleaned Machine's Surface Area. Performed Simulation Test. Checked Sensors Values. Calibrate A/D and Centrifuge Values. | Unit is Operational. | Unit has No UPS. | Pump Rotor Rollers. | S.I. 100521 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271573982827983245/6271573982827983245_signature_28.png | Divine Grace Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/6271573982827983245/6271573982827983245_signature_31.png
| 2025-07-01 | 2025-07-01 | Alfonso Ponce Enrile Memorial District Hospital | Gonzaga, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31079 | DOH | apemdh_laboratory@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2102 |  | GXIV-6C | 838474 | Due for Maintenance | Cleaned Icore Cleaned Cartridge Bay Cleaned Plunger Rod Cleaned Filter | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271573313511615284/6271573313511615284_signature_28.png | Regina Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6271573313511615284/6271573313511615284_signature_31.png
| 2025-07-01 | 2025-07-01 | Alfonso Ponce Enrile Memorial District Hospital | Gonzaga, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 31078 | DOH | apemdh_laboratory@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2102 |  | GXIV-6C | 838474 | Due for Maintenance | Cleaned Icore Cleaned Cartridge Bay Cleaned Plunger Rod Cleaned Filter | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271573283511431040/6271573283511431040_signature_28.png | Regina Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6271573283511431040/6271573283511431040_signature_31.png
| 2025-07-01 | 2025-07-01 | Catigbian District Hospital | Catigbian Bohol | 7 | Bohol | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31077 | DOH | charissemaeclaros97@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 595 | 2025-07-01 | GXIV | 838494 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271569553916805709/6271569553916805709_signature_28.png | Charisse Mae Claros | https://www.jotform.com/uploads/macarewecare/222927978372067/6271569553916805709/6271569553916805709_signature_31.png
| 2025-07-01 | 2025-07-01 | Mini Forest Health Center (District C) | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31076 | DOH | miniforestlaboratory@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 532 | 2025-06-20 | GX-IV 6C 4D | 838507 | For Preventive Maintenance Module 1 encounter multiple temperature related errors | Replaced the Defective Module SN:900881 with new Module SN: 717401 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | 900-0491rs | DR 13113 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271554835228225251/6271554835228225251_signature_28.png | Jezzal Perez Carbonell | https://www.jotform.com/uploads/macarewecare/222927978372067/6271554835228225251/6271554835228225251_signature_31.png
| 2025-07-01 | 2025-07-01 | Garcia Memorial Provincial Hospital | Talibon Bohol | 7 | Bohol | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31075 | DOH | ybanezrosalina19@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1506 | 2025-07-01 | GXIV | 838203 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271468661492286455/6271468661492286455_signature_28.png | Rosalina Ybañez | https://www.jotform.com/uploads/macarewecare/222927978372067/6271468661492286455/6271468661492286455_signature_31.png
| 2025-06-30 | 2025-06-27 | Isulan Rural Health Unit | Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31074 | UNOPS | febjennambong12@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1145 ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 |  | GxIV-10C | 110023520 | None | Perform Genexpert Cleaning and maintenance Import MTB-XDR Assay Checking of module status CSV Data collection Data collection | Perform Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271019826865409628/6271019826865409628_signature_28.png | February jen apiado | https://www.jotform.com/uploads/macarewecare/222927978372067/6271019826865409628/6271019826865409628_signature_31.png
| 2025-06-30 | 2025-06-25 | Tampakan Rural Health Unit | Poblacion, Tampakan,South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31073 | DOH | jmedtech_lab@yahoo.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 776 |  | GxIV-6C | 838544 | UPS no backup power | UPS problem troubleshooting UPS Problem reporting Data collection | Perform Genexpert UPS troubleshooting Perform Genexpert maintenance | 4 operational modules | UPS no backup power, requesting for battery replacement/UPS replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271019816866562057/6271019816866562057_signature_28.png | Jovi G. Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/6271019816866562057/6271019816866562057_signature_31.png
| 2025-06-30 | 2025-06-27 | Claveria Rural Health Unit | Poblacion, Claveria | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31072 | DOH | mgrmrtinez757@gmail.com | 28 |  |  | GX4-6C | 838572 | UPS not working CPU not powering on | Replaced fuse of UPS & AVR(resolved) Cleaned CPU ram & replaced bios battery Replaced CPU power cord Cleaned cartridge bays | Performed Troubleshooting Performed Preventive Maintenance | Unit is non operational | CPU not powering on |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6271011197213427695/6271011197213427695_signature_28.png | Mary Grace R.Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/6271011197213427695/6271011197213427695_signature_31.png
| 2025-06-30 | 2025-06-26 | Vicente Sotto Memorial Medical Center | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31071 | COMMERCIAL | vsmmcbloodservices@gmail.com | 2090 |  |  | MCS+ | 02A090 | Bowl problem noisy during draw process. | Check bowl area bowl lock parts. Check supply voltage and the UPS. Check kit during error encounter. Re-orient end-user the correct and proper procedure during installation of the bowl. | Check bowl area bowl lock parts. Check supply voltage and the UPS. Check kit during error encounter. | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270651300224206181/6270651300224206181_signature_28.png | Vicente Sotto Memorial Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6270651300224206181/6270651300224206181_signature_31.png
| 2025-06-30 | 2025-06-26 | Pasig Treatment Hub - Champ | Caruncho Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | dansumulong.pasigcho@gmail.com |  | 31070 | MACARE-HBDC | pathmedtech@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 342 |  | GX-IV-D-10C | 1100251270 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615463616985787/6270615463616985787_signature_28.png | Pasig Treatment Hub | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615463616985787/6270615463616985787_signature_31.png
| 2025-06-30 | 2025-06-25 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | CI | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31069 | COMMERCIAL | javayaay@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 182 ASSAY: Xpert CT/NG, COUNT: 694 ASSAY: Xpert HBV VL, COUNT: 67 ASSAY: Xpert HPV, COUNT: 957 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2376 ASSAY: Xpert MTB-XDR, COUNT: 5 ASSAY: Xpert HCV VL, COUNT: 47 ASSAY: Xpert HIV-1 Viral Load, COUNT: 279 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615413616808432/6270615413616808432_signature_28.png | Joshua Aya-ay | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615413616808432/6270615413616808432_signature_31.png
| 2025-06-30 | 2025-06-25 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Blood bank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31067 | COMMERCIAL | ajmbalsomo@stlukes.com.ph | 22 |  |  | MCS+ | 09C125 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615403616061310/6270615403616061310_signature_28.png | Aldwin Balsomo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615403616061310/6270615403616061310_signature_31.png
| 2025-06-30 | 2025-06-24 | Rosario Super Health Center | Dr. Sixto Antonio Avenue, Rosario, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31066 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1335 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615403617439240/6270615403617439240_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615403617439240/6270615403617439240_signature_31.png
| 2025-06-30 | 2025-06-25 | St. Luke's Medical Center - BGC | Rizal Dr, BGC Taguig City | NCR | NCR 4th - Taguig | CI | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 31065 | COMMERCIAL | javayaay@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 39 ASSAY: Xpert CT/NG, COUNT: 147 ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert HPV, COUNT: 213 ASSAY: Xpert MTB/RIF Ultra, COUNT: 439 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 110 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A3 not operational. For replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615403617682201/6270615403617682201_signature_28.png | Joshua Aya-ay | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615403617682201/6270615403617682201_signature_31.png
| 2025-06-30 | 2025-06-26 | Pasig Treatment Hub - Klinika Pasig | Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31068 | MACARE-HBDC | pathmedtech@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 343 |  | GX-IV-D-10C | 110025128 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615403619108293/6270615403619108293_signature_28.png | Jhozza Rei Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/6270615403619108293/6270615403619108293_signature_31.png
| 2025-06-30 | 2025-06-23 | Estevez Memorial Hospital, Inc. | Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City, 4500 Albay | 5 | Albay | Lbc | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31064 | COMMERCIAL | reylovendino83@gmail.com | 25 | ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 7 ASSAY: Xpert MTB-XDR, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 297 ASSAY: Xpert HBV VL, COUNT: 323 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert BCR-ABL Ultra, COUNT: 9 |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270582887006549106/6270582887006549106_signature_28.png | Estevez Memorial Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6270582887006549106/6270582887006549106_signature_31.png
| 2025-06-30 | 2025-06-26 | Batasan Social Hygiene Clinic | Batasan Rd QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31063 | MACARE-HBDC | moniqueteves99@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 32 ASSAY: Xpert HBV VL, COUNT: 50 ASSAY: Xpert HIV-1 Viral Load, COUNT: 997 |  | GX4 D 10C | 110020697 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270533753614318444/6270533753614318444_signature_28.png | Monique Teves | https://www.jotform.com/uploads/macarewecare/222927978372067/6270533753614318444/6270533753614318444_signature_31.png
| 2025-06-30 | 2025-06-26 | Biopath qc | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31062 | COMMERCIAL | rmlao@biopath.ph | 32 | ASSAY: Xpert CT/NG, COUNT: 37 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 49 ASSAY: Others, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 10 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX16D | 839306 | A1 and A3 encounter temp error 4009,4012. UPS 2kva unable to back supply due to drain battery | Checked ups Performed module self test, error still encounter Replaced 2kva ups spare ups in the laboratory | Performed troubleshooting | A1,A3,A4,B2,B4,C2,D2 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6270533373617691637/6270533373617691637_signature_28.png | Rhenee Lao-Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6270533373617691637/6270533373617691637_signature_31.png
| 2025-06-29 | 2025-06-19 | Naguilian District Hospital | Naguilian, La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31061 | DOH | marjelynrayray01@gmail.com | 30 |  |  | GXIV-4-D-6C | 838426 | For Data Collection | #NAME? | Performed Data Collection | Unit is Fully Operational Room is currently not air-conditioned |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6269902919144555126/6269902919144555126_signature_28.png | Marjelyn Rayray | https://www.jotform.com/uploads/macarewecare/222927978372067/6269902919144555126/6269902919144555126_signature_31.png
| 2025-06-29 | 2025-06-19 | Ilocos Training and Regional Medical Center | San Fernando, La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31060 | MACARE-HBDC | itrmc.mbl@gmail.com | 28 |  |  | GXIV-4-D-10C | 120003504 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6269901859146203571/6269901859146203571_signature_28.png | Dominga Patacsil | https://www.jotform.com/uploads/macarewecare/222927978372067/6269901859146203571/6269901859146203571_signature_31.png
| 2025-06-29 | 2025-06-18 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31059 | DOH | thisisformyfamily@yahoo.com | 30 |  |  | GXIV-4-D-6C | 838286 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6269900369142137490/6269900369142137490_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/6269900369142137490/6269900369142137490_signature_31.png
| 2025-06-29 | 2025-06-18 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31058 | DOH | arleneabad32@gmail.com | 29 |  |  | GXIV-4-D-6C | 838288 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6269899339148459256/6269899339148459256_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/6269899339148459256/6269899339148459256_signature_31.png
| 2025-06-29 | 2025-06-18 | San Jacinto Rural Health Unit | San Jacinto, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 31057 | DOH | allen.indong@gmail.com | 30 |  |  | GXIV-4-D-6C | 838518 | For Data Collection | #NAME? | Performed Data Collection | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6269898009145874720/6269898009145874720_signature_28.png | Alexander Indong | https://www.jotform.com/uploads/macarewecare/222927978372067/6269898009145874720/6269898009145874720_signature_31.png
| 2025-06-27 | 2025-06-27 | NEGROS ORIENTAL PROVINCIAL HOSPITAL | Dumaguete City | 7 | Negros Oriental | Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31056 | DOH | rezelmarielagos@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3438 ASSAY: Xpert MTB-XDR, COUNT: 80 ASSAY: Xpert HCV VL, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 467 |  | GX-IV | 838525 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6268263368511391578/6268263368511391578_signature_28.png | Rezel Marie S. Lagos , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6268263368511391578/6268263368511391578_signature_31.png
| 2025-06-27 | 2025-06-27 | NEGROS ORIENTAL PROVINCIAL HOSPITAL | Dumaguete City | 7 | Negros Oriental | Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 31055 | DOH | rezelmarielagos@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3438 ASSAY: Xpert MTB-XDR, COUNT: 80 ASSAY: Xpert HCV VL, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 467 |  | GX-IV | 838525 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6268263338517337430/6268263338517337430_signature_28.png | Rezel Marie S. Lagos , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6268263338517337430/6268263338517337430_signature_31.png
| 2025-06-27 | 2025-06-27 | Magsaysay Rural Health Unit | Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31054 | DOH | jkakilgan12@gmail.com | 29 |  |  | GX-4 6 Colors | 838557 | None | None | Preventive Maintenance CSV Data Gathering | All module and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6268105050363299012/6268105050363299012_signature_28.png | Jesusa N. Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/6268105050363299012/6268105050363299012_signature_31.png
| 2025-06-26 | 2025-06-26 | St Lukes Extension Clinic | J. Bocobo St. | NCR | NCR 1st - Manila | TB Lab | NO | CALIBRATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 31053 | MACARE-HBDC | laboratory@slec.ph | 25 |  | 2025-06-10 | GX XVI 10C | 110022374 | Unit for XpertCheck | None | Performed maintenance cleaning, performed XpertCheck procedure | Unit olis operational, modules B4 and D3 for replacement. |  | XpertCheck kit (x4) | DR# 13080 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267363634415040795/6267363634415040795_signature_28.png | Ma.Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6267363634415040795/6267363634415040795_signature_31.png
| 2025-06-26 | 2025-06-26 | Tomas Claudio Health Center | Camus St, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31052 | DOH | arleen_p_alconera@yahoo.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 897 |  | GX-IV 6C 4D | 838492 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are Operational Fan Filter damaged; reported to Jotform for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267296594712128406/6267296594712128406_signature_28.png | Karen Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6267296594712128406/6267296594712128406_signature_31.png
| 2025-06-26 | 2025-06-26 | Philippine Institute of Volcanology and Seismology | C.P. Garcia Ave., UP Campus, Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Geology Geophysics Research and Development Division | YES | MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | jeannemyrtia@gmail.com |  | 31051 | COMMERCIAL | mariethess.quilalang@phivolcs.dost.gov.ph | 25 |  |  | Stemi 305 | 3947020962 | None | N/A | Performed surface cleaning Check mechanical part Checks lights Checked lenses and eyepieces Cleaned Eyepieces Checked functionality Performed Preventive Maintenance | Unit is operational | Under PO No. 2025-06-153W : SI No. 100483 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267239413616923224/6267239413616923224_signature_28.png | Marie Thess Quilalang-Gemal | https://www.jotform.com/uploads/macarewecare/222927978372067/6267239413616923224/6267239413616923224_signature_31.png
| 2025-06-26 | 2025-06-26 | Philippine Institute of Volcanology and Seismology | C.P. Garcia Ave., UP Campus, Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Geology Geophysics Research and Development Division | YES | MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | jeannemyrtia@gmail.com |  | 31050 | COMMERCIAL | mariethess.quilalang@phivolcs.dost.gov.ph | 25 |  |  | Stemi 508 | 3951004842 | Molds inside eyepieces | None. Molds were inaccessible. | Performed surface cleaning Check mechanical part Checks lights Checked lenses and eyepieces Cleaned Eyepieces Cleaned camera adapter Checked functionality Performed Preventive Maintenance | Unit is operation but eyepiece molds cannot be accessed because located between 2 lenses | Under PO No. 2025-06-153W : SI No. 100483 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267238373617246913/6267238373617246913_signature_28.png | Marie Thess Quilalang-Gemal | https://www.jotform.com/uploads/macarewecare/222927978372067/6267238373617246913/6267238373617246913_signature_31.png
| 2025-06-26 | 2025-06-26 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31049 | MACARE-HBDC | treatmenthub.qcgh@gmail.com | 26 |  |  | GXIV-10C | 110012766 | Unit for maintenance Reported temperature error on B2 | NA | #NAME? | Unit is operational Room ambient temp. 27 deg. C on average Module B2 is under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267223294425471149/6267223294425471149_signature_28.png | Arizza Sheena Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/6267223294425471149/6267223294425471149_signature_31.png
| 2025-06-26 | 2025-06-26 | Diit District Health Center | Tacloban City | 8 | Leyte | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 31048 | DOH | dalepenalosa189@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1093 | 2025-06-26 | GXIV | 838421 | Computer Keyboard defective | #NAME? | #NAME? | Unit is Operational (Module B3 for replacement) |  | 1 Keyboard SN:24JIU00 | 494324 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267217782795600246/6267217782795600246_signature_28.png | Dale D. Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6267217782795600246/6267217782795600246_signature_31.png
| 2025-06-26 | 2025-06-26 | Chinese General Hospital | Blumentritt St. Sta. Cruz, Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31047 | COMMERCIAL | cghmcbloodbank2016@gmail.com | 25 |  | 2025-06-13 | MCS+ | 08M463 | Unit for preventive maintenance. For replacement if rubber rollers. | Repair and preventive maintenance. | Performed Preventive Maintenance. Checked and replaced worn-out rubber rollers of rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  | 6pcs. Uretane rollers. | DR#13086 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267081196292780675/6267081196292780675_signature_28.png | Matromanrey M. Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/6267081196292780675/6267081196292780675_signature_31.png
| 2025-06-26 | 2025-06-25 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31046 | COMMERCIAL | cduhbloodbank2@gmail.com | 24 |  |  | MCS+ | 10M104 | Blood pump positioning error. | Replacement of pump motor Blood pump, transfer pump motor, Calibrate the machine | Replacement of pump motor Blood pump, transfer pump motor, | Unit is operational. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267065365074068162/6267065365074068162_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6267065365074068162/6267065365074068162_signature_31.png
| 2025-06-26 | 2025-06-25 | Delos Santos Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@dlsmc.ph |  | 31045 | CONSORTIUM | mndionisio@dlsmc.ph | 23 | ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 40 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 401 |  | GX4D 10C | 110020700 | Due for maintenance Last xpertcheck 3/8/24 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267061573614107849/6267061573614107849_signature_28.png | Margarita N. Dionisio | https://www.jotform.com/uploads/macarewecare/222927978372067/6267061573614107849/6267061573614107849_signature_31.png
| 2025-06-26 | 2025-06-24 | MMC Medilinx | Amorosolo st,Makati | NCR | NCR 4th - Makati | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 31044 | COMMERCIAL | janeshyrelenriquez.medilinx@outlook.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 89 ASSAY: Xpert CT/NG, COUNT: 214 ASSAY: Xpert HBV VL, COUNT: 51 ASSAY: Xpert HPV, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 917 ASSAY: Others, COUNT: 6 ASSAY: Xpert HCV VL, COUNT: 50 ASSAY: Xpert HIV-1 Viral Load, COUNT: 544 | 2025-06-17 | GX16 D (10C) | 110026448 | Due for maintenance For module replacement C1 210135631 - 611643 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed module replacement and maintenance | Operational All modules working |  | 900-0650rs | 13095 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6267057163611109495/6267057163611109495_signature_28.png | Medi Linx Laboratory Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/6267057163611109495/6267057163611109495_signature_31.png
| 2025-06-25 | 2025-06-25 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Microbiology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | aivee.alcantara@medilinx.com.ph |  | 31043 | COMMERCIAL | rosandhreikatherinejose@gmail.com | 25 |  |  | GXII | 812916 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6266459146291791778/6266459146291791778_signature_28.png | Rosandhrei Katherine Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/6266459146291791778/6266459146291791778_signature_31.png
| 2025-06-25 | 2025-06-25 | Dasol Community Hospital | Dasol Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 31042 | DOH | jeremypulido0521@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 578 | 2025-06-20 | GXIV-D 6C | 838287 | B3: 808251 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational B3:719205 new module |  | Genexpert module | 13112 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6266445953123678265/6266445953123678265_signature_28.png | Jeremy Pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/6266445953123678265/6266445953123678265_signature_31.png
| 2025-06-25 | 2025-06-25 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31041 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2993 |  | 6C GXIV-4D | 838285 | Due for monthly PM | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6266401563387329530/6266401563387329530_signature_28.png | Leah Karen Noynay | https://www.jotform.com/uploads/macarewecare/222927978372067/6266401563387329530/6266401563387329530_signature_31.png
| 2025-06-25 | 2025-06-25 | San Pablo Colleges Medical Center | Sna Pablo Laguna | 4A | Laguna | Molecular laboratory | NO | INSTALLATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 31040 | CONSORTIUM | spcmc.molecularlab@gmail.com | 25 |  |  | GXIV-10C | 11002552 | Unit for relocation to another room. Unit for preventive maintenance. | Relocation and preventive maintenance. | Performed relocation. Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6266391826297547043/6266391826297547043_signature_28.png | Leonybelle Ann Alvarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6266391826297547043/6266391826297547043_signature_31.png
| 2025-06-25 | 2025-06-25 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | REPAIR | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31039 | MACARE-HBDC | bmc.bloodbank@gmail.com | 25 |  |  | MCS+ | 15H031G | Buzzing sound on centrifuge during operational | Adjusted screw of centrifuge lid Tested if there is any buzzing sound(resolved) | Performed Troubleshooting | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6266381955758691044/6266381955758691044_signature_28.png | Jake S. Moreno | https://www.jotform.com/uploads/macarewecare/222927978372067/6266381955758691044/6266381955758691044_signature_31.png
| 2025-06-25 | 2025-06-25 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31038 | DOH | buhanginlaboratory@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1178 |  | GX-IV 6C 4D | 838232 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6266369054713812221/6266369054713812221_signature_28.png | Catherine B. Deocampo,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6266369054713812221/6266369054713812221_signature_31.png
| 2025-06-24 | 2025-06-24 | PRC Port Area | Bonifacio Drive, Port Area, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | REPAIR | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 31037 | COMMERCIAL | lalaine.ilagan@redcross.org.ph | 25 |  | 2025-06-13 | MCS+ | 12D063 | Worn out rotor rubber roller. | Replaced rotor rubber roller. | Replaced rotor rubber roller. Performed functionality test. | Unit is operational. | N/A | Rotor rubber roller | DR #: 13085 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6265609760224680368/6265609760224680368_signature_28.png | Lalaine ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6265609760224680368/6265609760224680368_signature_31.png
| 2025-06-24 | 2025-06-24 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31036 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  | 2025-06-09 | GXIV | 110008530 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: insufficient cartridge B2: failed B3: insufficient cartridge B4: passed |  | XpertCheck Kit | 100406 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6265573536318809589/6265573536318809589_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6265573536318809589/6265573536318809589_signature_31.png
| 2025-06-24 | 2025-06-22 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31035 | COMMERCIAL | cduhbloodbank2@gmail.com | 24 |  |  | MCS+ | 10M104 | Line sensor hard to detect the tubing. | Adjust line sensor and check the valve of the mcs+ machine. | Adjust line sensor and check the valve of the mcs+ machine. | Unit is operational. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6265557755074340606/6265557755074340606_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6265557755074340606/6265557755074340606_signature_31.png
| 2025-06-24 | 2025-06-21 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31034 | COMMERCIAL | cduhbloodbank2@gmail.com | 24 |  |  | MCS+ | 10M104 | Blood pump positioning error. | Replacement of pump Blood pump, transfer pump, | Replacement of pump Blood pump, transfer pump, | Unit is operational. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6265556145072901934/6265556145072901934_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6265556145072901934/6265556145072901934_signature_31.png
| 2025-06-24 | 2025-06-24 | Vicente Mendiola Center for Health | Naga City, Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 31033 | DOH | geccaprocianos26@gmail.com | 23 |  | 2025-06-24 | GXIV-4D | 838564 | Due for preventive maintenance. | . Save system log file. . Save installation qualification report. . Save gxx file . Save assay statistics report. | . Clean plunger rod. . Clean filter. . Clean cartrage bay. . Clean modules and wire's. . Clean CPU and LCD | All modules are functional okay to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6265536265846099885/6265536265846099885_signature_28.png | Gene Carmel T. Procianos | https://www.jotform.com/uploads/macarewecare/222927978372067/6265536265846099885/6265536265846099885_signature_31.png
| 2025-06-24 | 2025-06-24 | Governor Generoso Municipal Health Office | Poblacion, Governor Generoso | 11 | Davao Oriental | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31032 | DOH | josephmagallanes72@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 494 |  | GX-IV 6C 4D | 838490 | Stucked Cartidge For Preventive Maintenance Data Gathering | Removed Stucked Cartidge, Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Stucked Cartidge Removal, Preventive Maintenance and Data Gathering | All Modules are operational | For observation Module 2 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6265527613225677325/6265527613225677325_signature_28.png | Charlene V. Magallanes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6265527613225677325/6265527613225677325_signature_31.png
| 2025-06-24 | 2025-06-18 | Talacogon District Hospital | Purok Bayabas, San Isidro, Talacogon | 13 | Agusan del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31031 | DOH | shyratortor@gmail.com | 25 |  |  | GX4 6 Colors | 838275 | UPS not working/ No power | UPS Replacement | Replacing defective UPS to a working one. Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6265490325427847775/6265490325427847775_signature_28.png | Shyra Ericka T. Calamba | https://www.jotform.com/uploads/macarewecare/222927978372067/6265490325427847775/6265490325427847775_signature_31.png
| 2025-06-24 | 2025-06-24 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31030 | DOH | bmccrcl2022@gmail.com | 28 |  |  | GX-4 6 colors | 838403 | Defective LCD Monitor. Blue thick line found at the bottom of the screen. | Reported to Main Office | Data Collection Preventive Maintenance Troubleshooting of LCD Monitor | All modules are functional. Monitor can be used temporarily. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6265408221304273871/6265408221304273871_signature_28.png | Joechita Doysabas | https://www.jotform.com/uploads/macarewecare/222927978372067/6265408221304273871/6265408221304273871_signature_31.png
| 2025-06-24 | 2025-06-23 | Cotabato Provincial Hospital | Brgy.Amas, Kidapawan City | 12 | Cotabato | Laboratory | NO | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31029 | MACARE-HBDC | januarybing1206@gmail.com | 25.4 |  |  | GxIV-10C | 11002654 | None | Installed Genexpert machine :110026514 Installed CPU SN:*CPH1124F56416* with keyboard and mouse Installed Monitor Del SN:CN-0WMYC5FCC00-478-CR3X-A07 Installed UPS Accupower SN:AT1K240705ST0038 Installed Printer HP SN:CN4AH427D7 Installed Assays: MTB-ULTRA, MTB-XDR and HIV-1 Viral load Checking of module status Data Collection | Perform Genexpert machine Installation | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6265289285514871074/6265289285514871074_signature_28.png | Hermosilla Enero | https://www.jotform.com/uploads/macarewecare/222927978372067/6265289285514871074/6265289285514871074_signature_31.png
| 2025-06-23 | 2025-06-20 | Tanza Wellness Center | Municipal Compound, Tanza, Cavite | 4A | Cavite | Social Hygiene Clinic | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 31028 | MACARE-HBDC | tanzarhu.clinicallab@gmail.com | 28 | ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 102 |  | GXIV-D-10C | 120003582 | Unit is due for Maintenance. | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6264768873613182900/6264768873613182900_signature_28.png | MERRIAM JOY G. SIBBALUCA | https://www.jotform.com/uploads/macarewecare/222927978372067/6264768873613182900/6264768873613182900_signature_31.png
| 2025-06-23 | 2025-06-23 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 31027 | DOH | suzzettekho34@gmail.com | 27 |  |  | GX-4 6 colors | 838554 | None | None | Preventive Maintenance Data Gathering | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6264679659212620813/6264679659212620813_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/6264679659212620813/6264679659212620813_signature_31.png
| 2025-06-23 | 2025-06-18 | Klinika Eastwood | Citywalk 2, Eastwood, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31026 | MACARE-HBDC | klinikaeastwoodlaboratory@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 92 |  | GX-IV-D-10C | 110020698 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6264559333612588463/6264559333612588463_signature_28.png | mary claire baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/6264559333612588463/6264559333612588463_signature_31.png
| 2025-06-23 | 2025-06-19 | Atimonan Rural Health Unit | Atimonan, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 31025 | DOH | amcl2008@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 594 |  | GX-IV-D-6C | 838206 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B1 disabled, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6264559323616685629/6264559323616685629_signature_28.png | Sandy Cecille P. Velasco | https://www.jotform.com/uploads/macarewecare/222927978372067/6264559323616685629/6264559323616685629_signature_31.png
| 2025-06-22 | 2025-06-19 | Kidapawan City Health Office | Kidapawan City, North Cotabato b | 12 | Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31024 | DOH | kidcholaboratory@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2310 |  | GxIV-6C | 838545 | None | Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of module status and errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6264114456861254953/6264114456861254953_signature_28.png | Cunanan, Neil Everson P. | https://www.jotform.com/uploads/macarewecare/222927978372067/6264114456861254953/6264114456861254953_signature_31.png
| 2025-06-22 | 2025-06-18 | Dr. Amado Diaz Provincial Foundation Hospital | Midsayap, North Cotabato | 12 | Cotabato | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 31023 | MACARE-HBDC | charlenemaeamatac526@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1273 ASSAY: Xpert HIV-1 Viral Load, COUNT: 192 |  | GxIV-10C | 110024249 | None | Perform Genexpert re-installation to new laboratory Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking modules status and errors Data collection | Genexpert Preventive maintenance Genexpert machine re-installation | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6264109056865858903/6264109056865858903_signature_28.png | Charlene Mae A. Amatac | https://www.jotform.com/uploads/macarewecare/222927978372067/6264109056865858903/6264109056865858903_signature_31.png
| 2025-06-22 | 2025-06-18 | Virac Municipal Health Office | Virac | 5 | Catanduanes | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31022 | DOH | limarmariano50@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2505 |  | GX4-6C | 838563 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6263970736427759880/6263970736427759880_signature_28.png | Ramil A. Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/6263970736427759880/6263970736427759880_signature_31.png
| 2025-06-20 | 2025-06-20 | Laguindingan Rural Health Unit | Laguindingan, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE Data Gathering | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31021 | DOH | chellesambaan@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 777 |  | 6C GXIV-4D | 838575 | None | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6262149259648750150/6262149259648750150_signature_28.png | Rochelle Sambaan | https://www.jotform.com/uploads/macarewecare/222927978372067/6262149259648750150/6262149259648750150_signature_31.png
| 2025-06-20 | 2025-06-20 | Esperanza Rural Health Unit | Poblacion, Esperanza | 5 | Masbate | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31020 | DOH | floramarmoncada@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 696 |  | GX4-6C | 838268 | Gx for Preventive Maintenance CPU non operational (beeping code) | Cleaned dust inside CPU Cleaned RAM Changed RAM slot input Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance Performed Troubleshooting | Unit is operational CPU is operational | Grid appearance on scren of monitor |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6262108248924052015/6262108248924052015_signature_28.png | Floramar M. Antipolo | https://www.jotform.com/uploads/macarewecare/222927978372067/6262108248924052015/6262108248924052015_signature_31.png
| 2025-06-20 | 2025-06-20 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 31019 | MACARE-HBDC | galagao@loveyourself.ph | 27 |  |  | GXIV | 807956 | Missing assay definition file | Imported assay definition | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6262080940023152728/6262080940023152728_signature_28.png | Genesis Alagao | https://www.jotform.com/uploads/macarewecare/222927978372067/6262080940023152728/6262080940023152728_signature_31.png
| 2025-06-20 | 2025-06-20 | Hinaplanon Health Center | Brgy. Hinaplanon, Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | MAINTENANCE Data Gathering | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31018 | DOH | villacortaamy50@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 777 |  | 6C GXIV-4D | 838227 | None | None | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6262075015575521462/6262075015575521462_signature_28.png | Amy S. Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6262075015575521462/6262075015575521462_signature_31.png
| 2025-06-20 | 2025-06-20 | Southern Philippines Medical Center | Bajada Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31017 | UNOPS | spmcrhivda@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1601 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1379 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 12 ASSAY: Xpert MTB-XDR, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert CT/NG, COUNT: 9 |  | GX-IV 10C 4D | 110023505 | Stucked Cartidge on Module B4 For Preventive Maintenance | Removed the Stucked Cartidge form Module 4 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance Removal of Stucked Cartidge | Successfully removed the Stucked Cartidge All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261955059698205588/6261955059698205588_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6261955059698205588/6261955059698205588_signature_31.png
| 2025-06-20 | 2025-06-20 | Masbate City Health Office | Masbate City | 5 | Masbate | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31016 | UNOPS | desrmt2017@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1125 ASSAY: Xpert MTB-XDR, COUNT: 37 |  | GX4-10C | 110023512 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261934082289864446/6261934082289864446_signature_28.png | Desiree Mae D Dela Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6261934082289864446/6261934082289864446_signature_31.png
| 2025-06-20 | 2025-06-20 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 31015 | UNOPS | spmcrhivda@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 273 ASSAY: Xpert SARS-CoV-2, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 591 |  | GX-IV 10C 4D | 110023506 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261902509691934040/6261902509691934040_signature_28.png | Flordeliza Go,. RMT MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/6261902509691934040/6261902509691934040_signature_31.png
| 2025-06-20 | 2025-06-20 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | INSTALLATION | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31014 | COMMERCIAL | cduhbloodbank2@gmail.com | 24 |  |  | MCS+ | 10M104 | Installation of Demo Unit for the replacement of the defective Unit Serial No. 17H639SPG. | Installation of Demo Unit for the replacement of the defective Unit Serial No. 17H639SPG. | Installation of Demo Unit for the replacement of the defective Unit Serial No. 17H639SPG. | Unit is operational. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261864993446707502/6261864993446707502_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6261864993446707502/6261864993446707502_signature_31.png
| 2025-06-19 | 2025-06-19 | Tomas Cabili Health Center | Iligan City, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE Data Gather | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31013 | DOH | mariacielovenezuela@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1058 |  | 6C GXIV-4D | 838294 | None | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261309916475290558/6261309916475290558_signature_28.png | Maria Cielo Venezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/6261309916475290558/6261309916475290558_signature_31.png
| 2025-06-19 | 2025-06-19 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR Pull out the unit for repair | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31012 | COMMERCIAL | cduhbloodbank2@gmail.com | 24 |  |  | MCS+ | 17H639SPG | SPM is out of tolerance and failed test during start up. | Replacement of the motherboard of the CPU. | Replacement of the motherboard of the CPU. | Unit for pull out and replacement of the demo unit. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261302223446907966/6261302223446907966_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6261302223446907966/6261302223446907966_signature_31.png
| 2025-06-19 | 2025-06-19 | Kauswagan Rural Health Unit | Kauswagan, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE Data Gathering | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31011 | DOH | dextertinoy@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 330 |  | 6C GXIV-4D | 838420 | None | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261272643215551320/6261272643215551320_signature_28.png | Dexter Tinoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6261272643215551320/6261272643215551320_signature_31.png
| 2025-06-19 | 2025-06-19 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 31010 | COMMERCIAL | cduhbloodbank2@gmail.com | 24 |  |  | MCS+ | 17H639SPG | SPM is out of tolerance and failed test during start up. | Replacement of the motherboard of the CPU. | Replacement of the motherboard of the CPU. Calibrate the SPM unit. | Unit is not operational due to defective SPM unit. Unit recommended for repair in the office and provide temporary demo unit. | OSMENA BLVD. CAPITOL SITE CEBU CITY |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261272533446368819/6261272533446368819_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6261272533446368819/6261272533446368819_signature_31.png
| 2025-06-19 | 2025-06-12 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 31009 | COMMERCIAL | cduhbloodbank2@gmail.com | 24 |  |  | MCS+ | 17H639SPG | SPM is out of tolerance and failed test during start up. | Replacement of the motherboard of the CPU. | Replacement of the motherboard of the CPU. | Unit is operational and under observation. |  |  | MM-DD-YYYY |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261269473444513181/6261269473444513181_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6261269473444513181/6261269473444513181_signature_31.png
| 2025-06-19 | 2025-06-19 | East Avenue Medical Center | East Ave. QC | NCR | NCR 2nd - Quezon City | Blood Bank | YES | REPAIR | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com | bloodbankeamc@gmail.com |  | 31008 | MACARE-HBDC | eamc.biomed@gmail.com | 27 |  | 2025-06-18 | MCS+ | 15H044 G | Defective Line Sensor Assembly. | Replaced line sensor assembly. Performed functionality test. Run simulation test. | Replaced line Senson Assembly. Note: Service part only. | Unit is operational. | Service part only. | Line sensor assembly. | DR #: 13102 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261262475617133485/6261262475617133485_signature_28.png | Rock banda | https://www.jotform.com/uploads/macarewecare/222927978372067/6261262475617133485/6261262475617133485_signature_31.png
| 2025-06-19 | 2025-06-19 | Cawayan Rural Health Office | Cawayan | 5 | Masbate | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31007 | DOH | clenettedollison@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1337 |  | GX4-6C | 838267 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261259856334570644/6261259856334570644_signature_28.png | Clenette Nyza A. Dollison | https://www.jotform.com/uploads/macarewecare/222927978372067/6261259856334570644/6261259856334570644_signature_31.png
| 2025-06-19 | 2025-06-19 | Lugait Rural Health Unit | Lugait, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE Data Gathering | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 31006 | DOH | lugaitmaternity@gmail.com | 27 |  |  | 6C GXIV-4D | 838562 | None | None | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6261186203216139302/6261186203216139302_signature_28.png | Rodelyn Limocon | https://www.jotform.com/uploads/macarewecare/222927978372067/6261186203216139302/6261186203216139302_signature_31.png
| 2025-06-18 | 2025-06-13 | Tigaon Municipal Health Office | Caraycayon, Tigaon | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31005 | DOH | jingimperial24@gmail.com | 27 |  |  | GX4-6C | 838558 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260671356426284739/6260671356426284739_signature_28.png | Maria Jesusa P.Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6260671356426284739/6260671356426284739_signature_31.png
| 2025-06-18 | 2025-06-13 | San Jose Municipal Health Office | San Jose | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31004 | DOH | jhuneric2008@yahoo.com | 29 |  |  | GX4-6C | 838389 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260671066427713536/6260671066427713536_signature_28.png | Rosmar Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6260671066427713536/6260671066427713536_signature_31.png
| 2025-06-18 | 2025-06-11 | Vinzons Rural Health Unit | Poblacion, Vinzons | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31003 | DOH | bardonamabel@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1206 |  | GX4-6C | 838576 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational | B2/862906 - 2014, 4009 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260670516429894373/6260670516429894373_signature_28.png | Amabel A. Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/6260670516429894373/6260670516429894373_signature_31.png
| 2025-06-18 | 2025-06-11 | Labo District Hospital | Talobatib, Labo | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31002 | DOH | triciamaemariano@gmail.com | 24 |  |  | GX4-6C | 838207 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260669126424960205/6260669126424960205_signature_28.png | Tricia Mae Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/6260669126424960205/6260669126424960205_signature_31.png
| 2025-06-18 | 2025-06-11 | Calabanga Rural Health Unit | Calabanga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31001 | DOH | martyarc11@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 968 |  | GX4-6C | 838398 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260668756429145801/6260668756429145801_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/6260668756429145801/6260668756429145801_signature_31.png
| 2025-06-18 | 2025-06-10 | Pili Rural Health Unit | Poblacion, Pili | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 31000 | DOH | lfsanao724@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1298 |  | GX4-6C | 838408 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is operational | B1/688686 - 1002 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260665446422134367/6260665446422134367_signature_28.png | Liezl F. Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/6260665446422134367/6260665446422134367_signature_31.png
| 2025-06-18 | 2025-06-18 | Ramon Rural Health Unit | Ramon, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30999 | DOH | rhuramonlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 596 |  | GXIV-6C | 838480 | Due for Maintenance | Cleaned Icore Cleaned Cartridge Bay Cleaned Filter Cleaned Plunger Rod | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260417720889334092/6260417720889334092_signature_28.png | Janette Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/6260417720889334092/6260417720889334092_signature_31.png
| 2025-06-18 | 2025-06-18 | Labuan General hospital | Labuan Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30998 | DOH | irenecasee@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2577 | 2025-05-29 | GX 4 | 838568 | System unit unable to turn on. Module 2 h/w failed | Replaced defective module 862915 with new module 885595. Replaced RAM. | Replaced defective module 862915 with new module 885595. Replaced RAM. | All module are available and ready to use. System Unit is working properly. | None | 900-0491RS | 13058 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260382246418491544/6260382246418491544_signature_28.png | Irene Casepe | https://www.jotform.com/uploads/macarewecare/222927978372067/6260382246418491544/6260382246418491544_signature_31.png
| 2025-06-18 | 2025-06-18 | Cauayan City Health Office | Cauayan City, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30997 | UNOPS | cho1lab14@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1585 |  | GXIV-10C | 110023504 | Due for Maintenance | Cleaned Icore Cleaned Cartridge Bay Cleaned Plunger Rod Cleaned Filter | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260368601791435037/6260368601791435037_signature_28.png | Arlyn Daggao | https://www.jotform.com/uploads/macarewecare/222927978372067/6260368601791435037/6260368601791435037_signature_31.png
| 2025-06-18 | 2025-06-18 | Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30996 | DOH | miniforestlaboratory@gmail.com | 22.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 519 |  | GX-IV 6C 4D | 838507 | UPS Battery defective Data Gathering | Replaced the Defective UPS Battery with new UPS Battery Export CSV Files | Perform UPS Battery Replacement Data Gathering | Successfully replaced the UPS Batteries |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260337175693525779/6260337175693525779_signature_28.png | Lolijell Torralba,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6260337175693525779/6260337175693525779_signature_31.png
| 2025-06-18 | 2025-06-18 | Baybay City Health Office | Baybay Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30995 | DOH | sajecopooh20@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1229 | 2025-06-18 | GXIV | 838555 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260284684534776404/6260284684534776404_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/6260284684534776404/6260284684534776404_signature_31.png
| 2025-06-18 | 2025-06-18 | Nazareth Health Center | Brgy. Nazareth, Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE Data Gathering | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30994 | DOH | nanettebtolentino@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1210 |  | 6C GXIV-4D | 838269 | None | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260284013211578881/6260284013211578881_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6260284013211578881/6260284013211578881_signature_31.png
| 2025-06-18 | 2025-06-18 | DR. CATALINO NAVA PROVINCIAL HOSPITAL | Guimaras | 6 | Guimaras | Laboratory | NO | MAINTENANCE GX Transfer | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | dcgnphlab@gmail.com |  | 30993 | UNOPS | mylene210@yahoo.com | 28 | ASSAY: , COUNT: 0 |  | GX-IV | 110023540 | #NAME? | #NAME? | #NAME? | Equipment is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260272259317342349/6260272259317342349_signature_28.png | Mylene Publico, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6260272259317342349/6260272259317342349_signature_31.png
| 2025-06-18 | 2025-06-17 | Carmen Health Center | Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE Data Gather | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30992 | DOH | juvymadarang@yahoo.com | 28 |  |  | 6C GXIV-4D | 838442 | None | None | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6260198846019103550/6260198846019103550_signature_28.png | Juvy T. Madarang | https://www.jotform.com/uploads/macarewecare/222927978372067/6260198846019103550/6260198846019103550_signature_31.png
| 2025-06-17 | 2025-06-11 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30991 | DOH | gbmmbg@yahoo.com | 29 |  |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6259591702529477621/6259591702529477621_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/6259591702529477621/6259591702529477621_signature_31.png
| 2025-06-17 | 2025-06-11 | Mangatarem Rural Health Unit | Mangatarem, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30990 | DOH | slasheiji1988new@gmail.com | 28 |  |  | GXIV-4-D-6C | 838284 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6259589902524438872/6259589902524438872_signature_28.png | Ken Lee Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/6259589902524438872/6259589902524438872_signature_31.png
| 2025-06-17 | 2025-06-17 | Piat Municipal Health Unit | Piat, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30989 | DOH | sentrongsigla1999@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 274 |  | GXIV-6C | 838532 | Due for Maintenance | Performed Maintenance Performed Functionality Test | Gathered Reports Cleaned Icore Cleaned Cartridge Bay Cleaned Filter Cleaned plunger arod | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6259545678561012814/6259545678561012814_signature_28.png | Marie Atap Salas | https://www.jotform.com/uploads/macarewecare/222927978372067/6259545678561012814/6259545678561012814_signature_31.png
| 2025-06-17 | 2025-06-17 | Tuao Rural Health Unit | Tuao, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30988 | DOH | annrollaine@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 745 |  | GXIV-6C | 838465 | Due for Maintenance | Performed Maintenance Performed Functionality Test | Gathered Reports Cleaned Icore Cleaned Cartridge Bay Cleaned Plunger Rox Cleaned Filter | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6259510558605565317/6259510558605565317_signature_28.png | Rollaine Ann Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/6259510558605565317/6259510558605565317_signature_31.png
| 2025-06-17 | 2025-06-17 | Cagayan de Oro City Health Office | Cagayan de Oro City | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30987 | DOH | chocdotbdots@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1951 |  | 6C GXIV-4D | 838441 | None | None | #NAME? | Machine is functional. | Filter needs to be replaced. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6259504273213761972/6259504273213761972_signature_28.png | Darline Bollozos | https://www.jotform.com/uploads/macarewecare/222927978372067/6259504273213761972/6259504273213761972_signature_31.png
| 2025-06-17 | 2025-06-17 | Cagayan de Oro City Health Office | Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30986 | DOH | chocdotbdots@gmail.com | 26 |  |  | 6C GXIV-4D | 838435 | None | None | #NAME? | Machine is functional | Filter needs to be replaced |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6259504273218515541/6259504273218515541_signature_28.png | Darline Bollozos | https://www.jotform.com/uploads/macarewecare/222927978372067/6259504273218515541/6259504273218515541_signature_31.png
| 2025-06-17 | 2025-06-17 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30985 | MACARE-HBDC | rhwclab@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 393 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1923 |  | GX-IV 6C 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Only Three Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6259467757542177214/6259467757542177214_signature_28.png | Charlene Dueñas,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6259467757542177214/6259467757542177214_signature_31.png
| 2025-06-17 | 2025-06-17 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30984 | DOH | rhwclab@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 393 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1923 |  | GX-IV 6C 4D | 838262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6259465967549799286/6259465967549799286_signature_28.png | Charlene Dueñas,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6259465967549799286/6259465967549799286_signature_31.png
| 2025-06-16 | 2025-06-11 | St.Elizabeth Hospital Inc. | General Santos City | 12 | South Cotabato | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30983 | COMMERCIAL | vfbristol@sehi.ph | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 9 |  | GxIV-10C | 110003713 | Module 2,3, and 4 H/W failed (Hardwares failed) | Perform Genexpert cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of errors/problem Problem reporting Data reporting | Perform Genexpert machine Preventive maintenance | 1 operational module (module 1) Module 2,3 & 4 H/W failed | Module 2,3 & 4 requesting for module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6258910876867868365/6258910876867868365_signature_28.png | Virdelyn Bristol | https://www.jotform.com/uploads/macarewecare/222927978372067/6258910876867868365/6258910876867868365_signature_31.png
| 2025-06-16 | 2025-06-11 | MEDI LINX Laboratory Inc./ St.Elizabeth Hospital Inc. | General Santos City | 12 | South Cotabato | Micrology lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | cl_SEHImaintest2@medilinx.com.ph |  | 30982 | COMMERCIAL | sofia.maclang@medilinx.com.ph | 29.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 50 |  | GxIV-6C | 831313 | All modules H/W failed (Hardware failed) | Checking of modules status and problem Errors/problem reporting Data collection | Perform Genexpert machine Preventive maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6258907146864265021/6258907146864265021_signature_28.png | Therese Christia Diana Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6258907146864265021/6258907146864265021_signature_31.png
| 2025-06-16 | 2025-06-10 | Malungon Rural Health Unit | Poblacion Malungon, Sarangani | 12 | Sarangani | Tb-Lab | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30981 | DOH | madameerrmie@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1139 |  | GxIV-6C | 838546 | CPU hard start-up | Perform CPU problem troubleshooting Replace CMOS battery CPU Cleaning Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of errors/problem Data collection | CPU troubleshooting CPU repair Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6258903526864137039/6258903526864137039_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/6258903526864137039/6258903526864137039_signature_31.png
| 2025-06-16 | 2025-06-16 | Maddela District Hospital | Maddela, Quirino | 2 | Quirino | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30980 | DOH | mdhlaboratory00@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 380 |  | GXIV-6C | 838469 | Monitor no power | Replaced Monitor Adaptor | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6258724979113645310/6258724979113645310_signature_28.png | Kathleen Joy Budomo | https://www.jotform.com/uploads/macarewecare/222927978372067/6258724979113645310/6258724979113645310_signature_31.png
| 2025-06-16 | 2025-06-16 | Ragay Rural Health Unit | Poblacion-Ilawod, Ragay | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30979 | DOH | phylgomez0115@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 681 |  | GX4-6C | 838560 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6258676234428294536/6258676234428294536_signature_28.png | PHYL FRANCES C. GOMEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/6258676234428294536/6258676234428294536_signature_31.png
| 2025-06-16 | 2025-06-16 | Philippine Red Cross Zamboanga | Pettit Baracks Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30978 | COMMERCIAL | zamboangacity@redcross.org.ph | 25 |  | 2025-06-16 | MCS + | 01M256 | Anti coagulant rotor rubber broken. | Replaced anti coagulant rotor rubber and double check all rotor rubber in the machine. Undergone diagnostics test. | Replaced anti coagulant rotor rubber and double check all rotor rubber in the machine. Undergone diagnostics test. | Machine is ready to use and available. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6258564271557541893/6258564271557541893_signature_28.png | Pilar Janette Y. Domo | https://www.jotform.com/uploads/macarewecare/222927978372067/6258564271557541893/6258564271557541893_signature_31.png
| 2025-06-13 | 2025-06-11 | Ospital ng Makati | Sampaguita St Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 30977 | MACARE-HBDC | jinx01gs@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 500 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011745361072140/6256011745361072140_signature_28.png | Cristina Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011745361072140/6256011745361072140_signature_31.png
| 2025-06-13 | 2025-06-13 | Antipolo City Health Office | M. Santos St. Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30975 | UNOPS | acho.laboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert MTB-XDR, COUNT: 79 ASSAY: Xpert HIV-1 Viral Load, COUNT: 650 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011735362679597/6256011735362679597_signature_28.png | Christine S. Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011735362679597/6256011735362679597_signature_31.png
| 2025-06-13 | 2025-06-10 | Santolan Super Health Center | Evangelista Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30974 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1544 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011735362739593/6256011735362739593_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011735362739593/6256011735362739593_signature_31.png
| 2025-06-13 | 2025-06-13 | Metro Antipolo Hospital and Medical Center | Marcos Highway Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30976 | COMMERCIAL | mahmci_labcmt@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 17 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 267 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GX-IV-D-10C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011735364366265/6256011735364366265_signature_28.png | Grazie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011735364366265/6256011735364366265_signature_31.png
| 2025-06-13 | 2025-06-10 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30973 | MACARE-HBDC | peachmango318@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1238 |  | GX-IV-D-10C | 110024245 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011735368379544/6256011735368379544_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6256011735368379544/6256011735368379544_signature_31.png
| 2025-06-13 | 2025-05-28 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | agimang@asianhospital.com |  | 30971 | COMMERCIAL | bloodbank@asianhospital.com | 25 |  |  | MCS+ | 01F116 | Centrifuge Cover Problem. Centrifuge Leaked Detected. | Troubleshoot | Adjust Centrifuge Cover. Cleaned Centrifuge and Spill Detector Sensor. Cleanes Surface Unit. Checked Diagnostic Values. Performed Simulation Run. | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872094159391074/6255872094159391074_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872094159391074/6255872094159391074_signature_31.png
| 2025-06-13 | 2025-05-30 | Rosario Rural Health Unit | Scout Torillio, Rosario, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30972 | DOH | arlenepenaarmijo@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1181 | 2025-05-19 | GXIV-D-6C | 838537 | UPS is Defective. | Repair. | Replaced UPS. Checked System's Functionality. Performed Self-test. Generate Reports. | Unit is Operational. | AT1000190123ST0255 -> AT1K240705ST0030 | Accupower | DR 13028 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872104154460146/6255872104154460146_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872104154460146/6255872104154460146_signature_31.png
| 2025-06-13 | 2025-05-21 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30969 | MACARE-HBDC | cjasonmiguel@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 88 ASSAY: Xpert MTB/RIF Ultra, COUNT: 98 ASSAY: Others, COUNT: 10 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 666 |  | GXIV-D-10C | 110015482 | UPS Problem. | Troubleshoot | Checked UPS Funtionality. Checked UPS Voltage Output. Gather Data for Reporting. | UPS is Defective. (No Power) Door of Module 2 is Broken. | Def UPS: AT1000211227ST0044 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872014156085904/6255872014156085904_signature_28.png | Jason Miguel Carandang | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872014156085904/6255872014156085904_signature_31.png
| 2025-06-13 | 2025-05-23 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 30970 | COMMERCIAL | agimang@asianhospital.com | 25 |  |  | MCS+ | 10B129 | Donor Line Error Detected. | Troubleshoot. | Checked DL1 and DL2 Sensors. Reset Connections. Cleaned Surface Area. | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872014157627821/6255872014157627821_signature_28.png | Angel danica Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872014157627821/6255872014157627821_signature_31.png
| 2025-06-13 | 2025-05-21 | Muntinlupa Reproductive Health and Wellness Center | Centennial Road, Muntinlupa | NCR | NCR 4th - Muntinlupa | Social Hygiene Clinic | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30967 | MACARE-HBDC | petelerica@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 152 |  | GXIV-D-10C | 110025706 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872004152710571/6255872004152710571_signature_28.png | Erica Monique Petel | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872004152710571/6255872004152710571_signature_31.png
| 2025-06-13 | 2025-05-23 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 30968 | COMMERCIAL | agimang@asianhospital.com | 25 |  |  | MCS+ | 01F116 | Leak On Centrifuge. | Troubleshoot. | Cleaned Centrifuge. Checked Leak at Centrifuge Assembly. Reset Fluid Detector Sensor. Cleaned Surface Unit. | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872004158069846/6255872004158069846_signature_28.png | Angel Danica Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6255872004158069846/6255872004158069846_signature_31.png
| 2025-06-13 | 2025-05-28 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | agimang@asianhospital.com |  | 30966 | COMMERCIAL | bloodbank@asianhospital.com | 25 |  |  | MCS+ | 10B129 | Fluid Detected at Centrifuge. | Troubleshoot. | Cleaned Centrifuge and Spill Detector Sensor. Cleanes Surface Unit. Checked Diagnostic Values. Performed Simulation Run. | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6255871994153014213/6255871994153014213_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6255871994153014213/6255871994153014213_signature_31.png
| 2025-06-11 | 2025-06-11 | Panabo City Health Office | Poblacion, Panabo | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30965 | DOH | nadineeust67@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2084 |  | GX-IV 6C 4D | 838295 | For Preventive Maintenance and Data Gathering | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6254440782113973304/6254440782113973304_signature_28.png | Nadine Gemma Eustaquio,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6254440782113973304/6254440782113973304_signature_31.png
| 2025-06-11 | 2025-06-11 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY, CEBU | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30964 | DOH | mendozarey42@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4241 ASSAY: Xpert HCV VL, COUNT: 134 ASSAY: Xpert HIV-1 Viral Load, COUNT: 592 |  | GX-IV | 838510 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6254429315815364717/6254429315815364717_signature_28.png | Rey C Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6254429315815364717/6254429315815364717_signature_31.png
| 2025-06-11 | 2025-06-11 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30963 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2078 |  | 6C GXIV-4D | 838561 | Due for monthly PM | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6254284564111664403/6254284564111664403_signature_28.png | Hannah Jane Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/6254284564111664403/6254284564111664403_signature_31.png
| 2025-06-11 | 2025-06-11 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30962 | DOH | buraeunlab@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 515 | 2025-06-11 | GXIV | 838551 | Module B2/862862 Defective | #NAME? | #NAME? | Unit is Operational |  | 1 Unit Genexpert Module SN: 607663 | 13059 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6254205092026123862/6254205092026123862_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6254205092026123862/6254205092026123862_signature_31.png
| 2025-06-10 | 2025-06-10 | Teresita Jalandoni Provincial Hospital | Silay City, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30961 | DOH | aklit26@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2705 ASSAY: Xpert MTB-XDR, COUNT: 197 | 2025-05-29 | GX4 6c | 838464 | Defective module 210019325/A2. | Preventive maintenance - clean optics, filter, module bay areas, plunger rods; archive tests | Replaced module A2 with new module SN 710566 | 3 modules operational. module A1 reported to Application Specialist for evaluation of errors. |  | Genexpert module 6c | 13063 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6253550653167434894/6253550653167434894_signature_28.png | Ma Kriska Awacay | https://www.jotform.com/uploads/macarewecare/222927978372067/6253550653167434894/6253550653167434894_signature_31.png
| 2025-06-10 | 2025-06-10 | Likhaan Center For Women's Health | Ofelia St., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30960 | COMMERCIAL | clatras@likhaan.org | 28 |  |  | GXII-10C | 110025174 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6253401721522379813/6253401721522379813_signature_28.png | Jocelyn latras | https://www.jotform.com/uploads/macarewecare/222927978372067/6253401721522379813/6253401721522379813_signature_31.png
| 2025-06-10 | 2025-06-10 | Iriga City Health Office | Sta Cruz Sur | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30959 | DOH | mongealyzza@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1178 |  | GX4-6C | 838391 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6253393076126172341/6253393076126172341_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/6253393076126172341/6253393076126172341_signature_31.png
| 2025-06-10 | 2025-06-10 | Iriga City Health Office | Sta Cruz Sur | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30958 | UNOPS | mongealyzza@yahoo.com.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1389 |  | GX4-10C | 110023511 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6253392426124392300/6253392426124392300_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/6253392426124392300/6253392426124392300_signature_31.png
| 2025-06-10 | 2025-06-09 | Capitol Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30957 | MACARE-HBDC | laboratory@capitolmedical.com.ph | 25 |  | 2025-06-05 | MCS+ | 07C002 | Unit for calibration Rotor roller rubbers for replacement | Adjusted bowl optics' offset and full scale values Replaced rotor roller rubbers | - replaced rotor roller rubbers - checked calibration parameters - adjusted bowl optics' offset and full scale values to acceptable range - performed self test | Unit is operational Calibration certificate to follow |  | MCS+ Urethane Roller Assy | 100395 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6253379821521430028/6253379821521430028_signature_28.png | Sheree Bacalla | https://www.jotform.com/uploads/macarewecare/222927978372067/6253379821521430028/6253379821521430028_signature_31.png
| 2025-06-09 | 2025-06-04 | Bagumbayan Rural Health Unit | Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30956 | DOH | marizalenoid@yahoo.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1408 | 2025-06-04 | GxIV-6C | 838261 | Module 1 H/W failed | Replace defective module SN:882094 with new module SN:657462 Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of errors/problem Data collection | Perform Genexpert machine module repair Perform Genexpert preventive maintenance | 4 operational modules |  | Genexpert 6-color module SN:657462 | DR#:13061 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252864456861030576/6252864456861030576_signature_28.png | MA. RIZA R. DIONELA , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6252864456861030576/6252864456861030576_signature_31.png
| 2025-06-09 | 2025-06-03 | Marikina City Health Office | Shoe Avenue Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | estyaragon2019@gmail.com |  | 30955 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 22 |  |  | GX-IV-D-10C | 120000613/110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661593618365100/6252661593618365100_signature_28.png | Estrelita s. Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661593618365100/6252661593618365100_signature_31.png
| 2025-06-09 | 2025-05-29 | Lubang Rural Health | Lubang, Occidental Mindoro | 4B | Occidental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30951 | DOH | lalaine.batalla0821@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 244 | 2025-04-11 | GX-IV-D-6C | 838228 | For UPS replacement | Replaced defective UPS | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore Replaced defective UPS Generated reports | Unit is operational. |  | Old UPS SN:AT1000190123ST0182 New UPS SN: AT1K240323ST0044 | DR# 12924 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661583613132259/6252661583613132259_signature_28.png | Lalaine V.Batalla | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661583613132259/6252661583613132259_signature_31.png
| 2025-06-09 | 2025-06-04 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30953 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 53 ASSAY: Xpert CT/NG, COUNT: 55 ASSAY: Xpert HBV VL, COUNT: 643 ASSAY: Xpert HPV, COUNT: 384 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6432 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 ASSAY: Others, COUNT: 6 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 166 ASSAY: Xpert HIV-1 Viral Load, COUNT: 908 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1476 |  | GX-XVI-D-10C | 110015767 | For maintenance Error 2034 on module A2 | Checked module A2 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports Checked module A2. | Unit is operational. Module A2 under observation. Unit is due for XpertCheck. Last XpertCheck: March 20, 2024 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661583615287123/6252661583615287123_signature_28.png | Xchosa Kay C. Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661583615287123/6252661583615287123_signature_31.png
| 2025-06-09 | 2025-06-03 | Marikina City Health Office | Shoe Avenue Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | estyaragon2019@gmail.com |  | 30952 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 22 |  |  | GX-IV-D-10C | 110024072/110024073 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661583616921991/6252661583616921991_signature_28.png | Estrelita s. Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661583616921991/6252661583616921991_signature_31.png
| 2025-06-09 | 2025-06-04 | Cardinal Santos Medical Center | Wilson St. Greenhills San Juan | NCR | NCR 2nd - San Juan | Microbiology | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30954 | COMMERCIAL | gbuenaventurarmt@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 110 ASSAY: Xpert HPV, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1142 |  | GX-IV-D-10C | 110003714 | Error 5007/2008/2120/2124 | Checked modules and gateway board | Performed maintenance Checked modules Reseat gateway board Run reliability test Run actual sample Generated reports | Unit is operational. Module B3 disabled. Modules B1 and B4 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661593617632523/6252661593617632523_signature_28.png | Gabriella Buenaventura | https://www.jotform.com/uploads/macarewecare/222927978372067/6252661593617632523/6252661593617632523_signature_31.png
| 2025-06-09 | 2025-06-09 | Philippine Red Cross Zamboanga | Pettit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30950 | COMMERCIAL | zamboangacity@redcross.org.ph | 25 |  | 2025-06-09 | MCS + | 01M256 | Haemonetics machine unable to turn on. | Undergone diagnostic cycles, Preventive maintenence and check output voltage values of UPS. | Undergone diagnostic cycles, Preventive maintenence and check output voltage values of UPS. | Passed the diagnostic test. Completed 1 procedure after the test. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252611111122520997/6252611111122520997_signature_28.png | Romeo wee | https://www.jotform.com/uploads/macarewecare/222927978372067/6252611111122520997/6252611111122520997_signature_31.png
| 2025-06-09 | 2025-05-29 | MMC Medilinx | Amorsolo st,Makati | NCR | NCR 4th - Makati | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30949 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 64 ASSAY: Xpert CT/NG, COUNT: 155 ASSAY: Xpert HBV VL, COUNT: 40 ASSAY: Xpert HPV, COUNT: 74 ASSAY: Xpert MTB/RIF Ultra, COUNT: 612 ASSAY: Others, COUNT: 4 ASSAY: Xpert HCV VL, COUNT: 30 ASSAY: Xpert HIV-1 Viral Load, COUNT: 364 | 2025-05-26 | GX16D 10C | 110026448 | For module replacement A1 210135512-210052199,C2 210135699-628743 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed module replacement | 15 modules operational 1 module disabled reported already |  | Genexpert 10C module | 13045 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252596525945878351/6252596525945878351_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6252596525945878351/6252596525945878351_signature_31.png
| 2025-06-09 | 2025-06-09 | Mabinay Medicare Community Hospital | Mabinay, Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30948 | DOH | mmchlaboratory2017@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 552 | 2025-05-29 | GX4 6C | 838481 | Defective module 862088/B1 | Preventive maintenance - clean optics, filter, module bay areas, plunger rods; archive tests | Replaced module B1 with new module SN 884570. | 4 modules operational. |  | Genexpert module 6c | 13057 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252574908217589949/6252574908217589949_signature_28.png | Mabinay Medicare Community Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6252574908217589949/6252574908217589949_signature_31.png
| 2025-06-09 | 2025-06-05 | Laoag City Health Office | Laoag City, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30947 | DOH | denver.mahor@gmail.com | 27 |  |  | GXIV-4-D-6C | 838416 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252546512561538887/6252546512561538887_signature_28.png | Rodenver Mahor | https://www.jotform.com/uploads/macarewecare/222927978372067/6252546512561538887/6252546512561538887_signature_31.png
| 2025-06-09 | 2025-06-05 | Mariano Marcos State University | Batac, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30946 | MACARE-HBDC | mpolivar@mmsu.edu.ph | 26 |  |  | GXIV-4-D-10C | 110022474 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252544822562363842/6252544822562363842_signature_28.png | Marvielyn Olivar | https://www.jotform.com/uploads/macarewecare/222927978372067/6252544822562363842/6252544822562363842_signature_31.png
| 2025-06-09 | 2025-06-05 | Sinait District Hospital | Sinait, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30945 | DOH | isdhsinaitlaboratory@gmail.com | 29 |  |  | GXIV-4-D-6C | 838277 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252542732569456319/6252542732569456319_signature_28.png | Claire Marie Baguyo-Diculen | https://www.jotform.com/uploads/macarewecare/222927978372067/6252542732569456319/6252542732569456319_signature_31.png
| 2025-06-09 | 2025-06-04 | Abra Provincial Hospital | Bangued, Abra | CAR | Abra | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30944 | UNOPS | aphpathologydepartment@gmail.com | 28 |  |  | GXIV-4-D-6C | 110023499 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252540962569009447/6252540962569009447_signature_28.png | Riverlie Valera | https://www.jotform.com/uploads/macarewecare/222927978372067/6252540962569009447/6252540962569009447_signature_31.png
| 2025-06-09 | 2025-06-04 | Manabo Rural Health Unit | Manabo, Abra | CAR | Abra | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30943 | DOH | lenmayclairelumiwag15@gmail.com | 26 |  | 2025-05-29 | GXIV-4-D-6C | 838274 | For Module Replacement | #NAME? | Performed Module Replacement | Unit is Fully Operational |  | GeneXpert Module | DR 13066 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252538872561955274/6252538872561955274_signature_28.png | Ellen May Claire S. Lumiwag | https://www.jotform.com/uploads/macarewecare/222927978372067/6252538872561955274/6252538872561955274_signature_31.png
| 2025-06-09 | 2025-06-04 | Sta. Maria Rural Health Unit | Sta. Maria, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30942 | UNOPS | jomarcaberto24@gmail.com | 30 |  |  | GXIV-4-D-10C | 110023500 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252536912561366372/6252536912561366372_signature_28.png | Jomar Caberto | https://www.jotform.com/uploads/macarewecare/222927978372067/6252536912561366372/6252536912561366372_signature_31.png
| 2025-06-09 | 2025-06-04 | Candon City Health Office | Candon, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30941 | DOH | candoncholaboratory@gmail.com | 28 |  |  | GXIV-4-D-6C | 838273 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252536012567754371/6252536012567754371_signature_28.png | Michelle Bucalen | https://www.jotform.com/uploads/macarewecare/222927978372067/6252536012567754371/6252536012567754371_signature_31.png
| 2025-06-09 | 2025-06-04 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | La Union | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30940 | DOH | piaramuj@gmail.com | 28 |  |  | GXIV-4-D-6C | 838270 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252534782564909132/6252534782564909132_signature_28.png | Marivic Tinaza | https://www.jotform.com/uploads/macarewecare/222927978372067/6252534782564909132/6252534782564909132_signature_31.png
| 2025-06-09 | 2025-06-09 | Itogon Rural Health Unit | Itogon, Benguet | CAR | Benguet | TB DOTS | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30939 | DOH | ggdepayso@gmail.com | 28 |  | 2025-05-29 | GXIV-4-D-6C | 838282 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 626320 >>> 732593 - Generated Reports | Performed Module Replacement | Unit is Fully Operational | Work with Engineer Venzon | GeneXpert Module | DR 13064 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252514342567844489/6252514342567844489_signature_28.png | Gloria Depayso | https://www.jotform.com/uploads/macarewecare/222927978372067/6252514342567844489/6252514342567844489_signature_31.png
| 2025-06-09 | 2025-06-03 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30938 | DOH | rhupozorrubio@yahoo.com | 30 |  | 2025-05-29 | GXIV-4-D-6C | 838467 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B1 872240 >>> 692828 - Generated Reports | Performed Module Replacement | Unit is Operational | Work with Engineer Venzon | GeneXpert Module | 13062 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252499992568901633/6252499992568901633_signature_28.png | Caroline Nacis | https://www.jotform.com/uploads/macarewecare/222927978372067/6252499992568901633/6252499992568901633_signature_31.png
| 2025-06-09 | 2025-06-05 | Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30937 | DOH | miniforestlaboratory@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 514 |  | GX-IV 6C 4D | 838507 | For Preventive Maintenance UPS Problem | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Reset connection, and check for line voltages | Perform Preventive Maintenance Perform UPS Troubleshooting | Machine is Operational UPS Battery is not working, no battery backup during battery mode |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252488811514996974/6252488811514996974_signature_28.png | Lolijell Torralba,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6252488811514996974/6252488811514996974_signature_31.png
| 2025-06-09 | 2025-06-05 | Pio Duran Rural Health Unit | Caratagan | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30936 | DOH | mitzijc@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 |  | GX4-6C | 838264 | UPS defective | Replaced UPS(AT1K240705ST0037) Cleaned cartridge bays, plunger rod, icore Generate Data | Replaced UPS performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252411966813127582/6252411966813127582_signature_28.png | Mitzi Cedeño | https://www.jotform.com/uploads/macarewecare/222927978372067/6252411966813127582/6252411966813127582_signature_31.png
| 2025-06-09 | 2025-06-04 | Rapu-Rapu Rural Health Unit | Laboratory | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30935 | DOH | suzetteoguan1@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 538 |  | GX4-6C | 838565 | B2/730958 temperature Error | Replaced module : B2 - 730958 >> 627350 Cleaned cartridge bays, plunger rod, icore Generate Data | Replaced module B2/730958 Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252411416814173240/6252411416814173240_signature_28.png | Suzette C. Oguan | https://www.jotform.com/uploads/macarewecare/222927978372067/6252411416814173240/6252411416814173240_signature_31.png
| 2025-06-09 | 2025-06-03 | Libon Rural Health Unit | Poblacion, Libon | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30934 | DOH | libonmho@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 931 |  | GX4-6C | 838397 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6252410476814448624/6252410476814448624_signature_28.png | Joe Ed C. Sayson | https://www.jotform.com/uploads/macarewecare/222927978372067/6252410476814448624/6252410476814448624_signature_31.png
| 2025-06-06 | 2025-06-06 | Zamboanga del sur medical center | Zamboanga del sur medical center | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30933 | DOH | limboemilie@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2265 ASSAY: Xpert HIV-1 Viral Load, COUNT: 103 | 2025-06-06 | Gx 4 | 838204 | UPS not functioning properly | Replaced connectors of UPS batteries due to corrosion. | Repaied UPS. Data collection and Preventive maintenance | All modules are available and UPS is working good with output voltage of 220V. | UPS needs to be charged. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6249899693527581609/6249899693527581609_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/6249899693527581609/6249899693527581609_signature_31.png
| 2025-06-06 | 2025-06-05 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30932 | DOH | lornabaigar@gmail.com | 28 |  | 2025-06-05 | GXIV | 838409 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6249749774815409056/6249749774815409056_signature_28.png | Lorna Abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/6249749774815409056/6249749774815409056_signature_31.png
| 2025-06-05 | 2025-06-05 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30931 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 16K521SPG | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6249087082122556393/6249087082122556393_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6249087082122556393/6249087082122556393_signature_31.png
| 2025-06-05 | 2025-06-05 | Philippine Red Cross | Blood bank | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30930 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6249085652129017233/6249085652129017233_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6249085652129017233/6249085652129017233_signature_31.png
| 2025-06-05 | 2025-06-05 | Molave Rural Health unit | Molave Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30929 | DOH | aldrinacain@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 | 2025-06-05 | Gx 4 | 838233 | None | Preventive Maintenence and data collection | Preventive Maintenence and data collection | All modules are available and ready to use | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6249006592998798677/6249006592998798677_signature_28.png | Aldrin Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/6249006592998798677/6249006592998798677_signature_31.png
| 2025-06-04 | 2025-05-28 | Metro Health Specialists Hospital Inc. | Sorsogon City | 5 | Sorsogon | Laboratory | YES | INSTALLATION PM training | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | aileenrmt@gmail.com |  | 30928 | COMMERCIAL | mhshilab@gmail.com | 28 |  |  | GX4-10C | 120002716 | PM training Installation in molec lab | Performed Preventive Maintenance training Installed Gx4-10C, UPS, Monitor, CPU, Printer Generate Data | Performed PM training Installed unit in lab | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6248342446815412910/6248342446815412910_signature_28.png | Aileen V. Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6248342446815412910/6248342446815412910_signature_31.png
| 2025-06-04 | 2025-06-04 | Balamban Rural Health Unit | Poblacion Balamban Cebu | 7 | Cebu | TB Laboratory | NO | INSTALLATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 30927 | DOH | climaco.mildred@yahoo.com | 23 |  | 2025-06-04 | GXIV-4D | 838231 | Defective LCD monitor vertical problem not working properly with serial number CNK9430BTK | . Save file Installation qualification reports . Save file system log file . Save csv file. . Save gxx file. | Temporary install new LCD monitor with serial number ZV0A1910029805 | LCD monitor functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6248319375845647455/6248319375845647455_signature_28.png | climaco.mildred@yahoo.com | https://www.jotform.com/uploads/macarewecare/222927978372067/6248319375845647455/6248319375845647455_signature_31.png
| 2025-06-04 | 2025-06-04 | Dinas Rural Health Unit | Dinas Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30926 | DOH | bernantereynacres@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 739 | 2025-06-04 | GX 4 | 86229 | Module 2 error 4009 | Troubleshoot module 2. Conducted Preventive maintenence and data collection. | Troubleshoot module 2. Conducted Preventive maintenence and data collection. | All modules are available | Module 2 will be reported to cephied | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6248192652448777050/6248192652448777050_signature_28.png | Reyna cres Bernante | https://www.jotform.com/uploads/macarewecare/222927978372067/6248192652448777050/6248192652448777050_signature_31.png
| 2025-06-04 | 2025-06-04 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30925 | DOH | sajecopooh20@gmail.com | 26 |  |  | 6C GXIV-4D | 838555 | Date/Time CPU resets every time unit turns on | #NAME? | None | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6248191293967097389/6248191293967097389_signature_28.png | Sarah Jean G. Pasana | https://www.jotform.com/uploads/macarewecare/222927978372067/6248191293967097389/6248191293967097389_signature_31.png
| 2025-06-04 | 2025-06-03 | Ramon magsaysay rural health unit | Ramon magsaysay zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30924 | DOH | rhuramonmagsaysay7@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 536 | 2025-06-03 | GX 4 | 838443 | None | Preventive Maintenence and data collection | Preventive Maintenence a d data collection | All modules are available and ready to use | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6248183442449011637/6248183442449011637_signature_28.png | Gina Borinaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6248183442449011637/6248183442449011637_signature_31.png
| 2025-06-04 | 2025-06-03 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | REPAIR TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30923 | DOH | luzvimindanavidad2@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1028 |  | 6C GXIV-4D | 838297 | #NAME? | #NAME? | #NAME? | Machine is operational. Waiting for module B4 replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6248077283674859895/6248077283674859895_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6248077283674859895/6248077283674859895_signature_31.png
| 2025-06-03 | 2025-05-30 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30922 | DOH | allan.vinluan2013@gmail.com | 28 |  |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247666631129149687/6247666631129149687_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6247666631129149687/6247666631129149687_signature_31.png
| 2025-06-03 | 2025-05-30 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30921 | DOH | seansimonpaloyo@gmail.com | 25 |  |  | GXIV-4-D-6C | 838473 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247664241125400300/6247664241125400300_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6247664241125400300/6247664241125400300_signature_31.png
| 2025-06-03 | 2025-05-28 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30920 | UNOPS | frangipanyang29@gmail.com | 30 |  |  | GXIV-4-D-6C | 110023498 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247656731127690572/6247656731127690572_signature_28.png | Frannie Calixto | https://www.jotform.com/uploads/macarewecare/222927978372067/6247656731127690572/6247656731127690572_signature_31.png
| 2025-06-03 | 2025-05-28 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30919 | DOH | sherryannemarrero@gmail.com | 25 |  |  | GXIV-4-D-6C | 838281 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247655251122357155/6247655251122357155_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6247655251122357155/6247655251122357155_signature_31.png
| 2025-06-03 | 2025-05-28 | Itogon Rural Health Unit | Itogon, Benguet | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30918 | DOH | ggdepayso@gmail.com | 30 |  |  | GXIV-4-D-6C | 838282 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247653891128735100/6247653891128735100_signature_28.png | Gloria Depayso | https://www.jotform.com/uploads/macarewecare/222927978372067/6247653891128735100/6247653891128735100_signature_31.png
| 2025-06-03 | 2025-05-22 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30917 | DOH | dianeruiz147@gmail.com | 29 |  |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247651811125379489/6247651811125379489_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6247651811125379489/6247651811125379489_signature_31.png
| 2025-06-03 | 2025-05-22 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30916 | COMMERCIAL | annabellelcatbagan@gmail.com | 25 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247650071122139219/6247650071122139219_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6247650071122139219/6247650071122139219_signature_31.png
| 2025-06-03 | 2025-05-22 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30915 | DOH | serly_reyes@yahoo.com | 30 |  |  | GXIV-4-D-6C | 838516 | To Replace BIOS Battery | #NAME? | Replaced BIOS Battery | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247648811126740484/6247648811126740484_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6247648811126740484/6247648811126740484_signature_31.png
| 2025-06-03 | 2025-05-16 | San Carlos City Health Office | San Carlos City,Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30914 | DOH | gbmmbg@yahoo.com | 25 |  |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247646891125036425/6247646891125036425_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/6247646891125036425/6247646891125036425_signature_31.png
| 2025-06-03 | 2025-05-16 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30913 | COMMERCIAL | annabellelcatbagan@gmail.com | 25 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247645851123466720/6247645851123466720_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6247645851123466720/6247645851123466720_signature_31.png
| 2025-06-03 | 2025-05-16 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30912 | DOH | serly_reyes@yahoo.com | 30 |  |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247637911129326504/6247637911129326504_signature_28.png | Hazel Ganapin | https://www.jotform.com/uploads/macarewecare/222927978372067/6247637911129326504/6247637911129326504_signature_31.png
| 2025-06-03 | 2025-05-15 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30911 | DOH | dianeruiz147@gmail.com | 29 |  |  | GXIV-4-D-6C | 838438 | To Replace UPS | #NAME? | Replaced UPS | Unit is Fully Operational | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247636741128182004/6247636741128182004_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6247636741128182004/6247636741128182004_signature_31.png
| 2025-06-03 | 2025-05-15 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30910 | DOH | allan.vinluan2013@gmail.com | 30 |  |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247634741125991179/6247634741125991179_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6247634741125991179/6247634741125991179_signature_31.png
| 2025-06-03 | 2025-06-03 | Southern Philippines Medical Center | JP Laurel Avenue, Bajada, Davao City | 11 | Davao del Sur | Internal Medicine | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 30909 | COMMERCIAL | icu2estuya@gmail.com | 24.6 |  |  | MCS+ | 15HO34G | Error on spill detector | Remove and clean all possible debris, moisture, and spillage on the centrifuge and fluid sensors Check for the parameters of SPM, DPM, Line Sensor Check the roller assays and replaced rollers with presence of threadings Perform Cycle Test | Perform Troubleshooting and Diagnostics | Machine is functional Changed rollers on red pumps | Replaced rollers, subjected for request for roller assays |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247478346265787286/6247478346265787286_signature_28.png | Rowenna Entig | https://www.jotform.com/uploads/macarewecare/222927978372067/6247478346265787286/6247478346265787286_signature_31.png
| 2025-06-03 | 2025-06-03 | Balamban Rural Health Unit | Balamban Cebu | 7 | Cebu | TB Laboratory | NO | TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 30908 | DOH | climaco.mildred@yahoo.com | 23 |  | 2025-06-03 | GX4 | 838231 | LCD Defective. | Request for replacement of LCD monitor. | . Save vedio for Documentation . Save Installation qualification report . Save system log file report. | Not ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247431065849454838/6247431065849454838_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6247431065849454838/6247431065849454838_signature_31.png
| 2025-06-03 | 2025-06-03 | Diit District Health Center | Diit, Tacloban City | 8 | Leyte | TB Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30907 | DOH | dalepenalosa189@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1083 | 2025-06-03 | GXIV | 838421 | #NAME? | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247387139015652760/6247387139015652760_signature_28.png | Dale Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6247387139015652760/6247387139015652760_signature_31.png
| 2025-06-03 | 2025-06-03 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30906 | UNOPS | molecularlab@drpjgmrmc.doh.gov.ph | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2108 ASSAY: Xpert MTB-XDR, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 327 |  | GXIV-10C | 110023508 | Due for Maintenance | Cleaned Icore Cleaned Cartridge Bay Cleaned Filter Cleaned Plunger Rod | Performed Maintenance Gatherex Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247379799716453145/6247379799716453145_signature_28.png | Kenneth Delos Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6247379799716453145/6247379799716453145_signature_31.png
| 2025-06-03 | 2025-06-02 | Cebu CHO | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30905 | COMMERCIAL | cchdserology@gmail.com | 27 |  |  | Gx4 | 829419 | Defective module 2 | Collecting CSV data IQR and system log | Collecting CSV data IQR and system log | Desective module 2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247312024131473337/6247312024131473337_signature_28.png | Owen Joshua L. Briones | https://www.jotform.com/uploads/macarewecare/222927978372067/6247312024131473337/6247312024131473337_signature_31.png
| 2025-06-03 | 2025-06-03 | Quezon Rural Health Unit | Poblacion, Quezon, Bukidnon | 10 | Bukidnon | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30904 | DOH | quenixlopez@gmail.com | 27 |  |  | GX4 6 Colors | 838440 | Stuck Cartridge in Module B1 | Stuck Cartridge Removal Preventive Maintenance | Manually removing of Stuck Cartridge in Module B1 Preventive Maintenance | All modules and units are now functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247300304068726934/6247300304068726934_signature_28.png | Quenie Rose Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6247300304068726934/6247300304068726934_signature_31.png
| 2025-06-03 | 2025-06-03 | Talavera General Hospital | Maestrang Kikay, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30903 | MACARE-HBDC | pjgtehlab@gmail.com | 22 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 116 ASSAY: Xpert MTB/RIF Ultra, COUNT: 556 ASSAY: Xpert CT/NG, COUNT: 82 ASSAY: Xpert HCV VL, COUNT: 11 ASSAY: Xpert HIV-1 Viral Load, COUNT: 252 |  | GXIV-10C | 110013773 | Due for Calibration | Performed Calibration (All Modules Passed) | Performed Maintenance Gathered Reports Performed Xpertcheck Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6247269749717199957/6247269749717199957_signature_28.png | Majalene Olar | https://www.jotform.com/uploads/macarewecare/222927978372067/6247269749717199957/6247269749717199957_signature_31.png
| 2025-06-02 | 2025-06-02 | Southern Philippines Medical Center | JP Laurel Avenue, Bajada, Davao City | 11 | Davao del Sur | Heart Institute | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 30902 | COMMERCIAL | aepatriana@gmail.com | 24.5 |  |  | MCS+ | 15HO34G | Centrifuge Problem | Clean Surface Area, check parameters of the MCS+ Machine Removed any debris from the Centrifuge Perform Cycle Test for functionality | Perform Preventive Maintenance Centrifuge Troubleshooting | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6246832397311114347/6246832397311114347_signature_28.png | Angelie Patriana | https://www.jotform.com/uploads/macarewecare/222927978372067/6246832397311114347/6246832397311114347_signature_31.png
| 2025-06-02 | 2025-05-28 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30901 | MACARE-HBDC | honeyeichel@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 76 ASSAY: Xpert HPV, COUNT: 65 ASSAY: Xpert HCV VL, COUNT: 43 ASSAY: Xpert HIV-1 Viral Load, COUNT: 360 |  | GXIV-D 10C | 120003587 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6246767181277077148/6246767181277077148_signature_28.png | Honey Eicher Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/6246767181277077148/6246767181277077148_signature_31.png
| 2025-06-02 | 2025-06-02 | Argao Primary Health Care (RHU II) | Argao, Poblacion Cebu City | 7 | Cebu | Tb laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30900 | DOH | victoria.jumamil@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1425 |  | GX4 | 838283 | Cpu and mother board defective. | Replacement of the motherboard of the CPU. | Replacement of the motherboard of the CPU. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6246551613441513638/6246551613441513638_signature_28.png | Victoria Jumamil, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6246551613441513638/6246551613441513638_signature_31.png
| 2025-06-02 | 2025-05-30 | Argao Primary Health Care (RHU II) | Argao, Poblacion Cebu City | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30899 | DOH | victoria.jumamil@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1425 |  | GX4 | 838283 | Cpu problem immediate shut down and hard start up. | Replacement of the motherboard of the CPU. | Replacement of the motherboard of the CPU. | Defective CPU. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6246549490804493850/6246549490804493850_signature_28.png | Victoria Jumamil, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6246549490804493850/6246549490804493850_signature_31.png
| 2025-06-02 | 2025-05-31 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30898 | COMMERCIAL | cduhbloodbank2@gmail.com | 24 |  |  | MCS+ | 17H639SPG | AC pump, Transfer pump and Blood pump wear out rubber Error 482 SPM error | - Replacement AC pump, Transfer pump and Blood pump wear out rubber - Restart and replaced new protocol due to Error 482 SPM error | -Replacement AC pump, Transfer pump and Blood pump wear out rubber - Restart and replaced new protocol due to Error 482 SPM error | For observation for 1 week, as of now machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6246546130803678379/6246546130803678379_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6246546130803678379/6246546130803678379_signature_31.png
| 2025-06-02 | 2025-06-02 | Mariveles District Hospital | Mariveles Bataan | 3 | Batanes | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30897 | MACARE-HBDC | mdh.mdl2020@gmail.com | 30 | ASSAY: Xpert SARS-CoV-2, COUNT: 0 |  | GXIV-D 6C | 806377 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6246544435512192938/6246544435512192938_signature_28.png | Mariga Krista Monilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6246544435512192938/6246544435512192938_signature_31.png
| 2025-06-02 | 2025-06-02 | Zamboanga del sur medical center | Zamboanga del sur medical center | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30896 | DOH | limboemilie@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2047 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 | 2025-06-02 | GX 4 | 838204 | None | Preventive Maintenence and data collection | Preventive Maintenence and data collection | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6246529303437066145/6246529303437066145_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/6246529303437066145/6246529303437066145_signature_31.png
| 2025-06-01 | 2025-05-28 | Lake Sebu Rural Health Unit | Purok Rosas, Brgy. Poblacion, Lake Sebu South Cotabato | 12 | South Cotabato | Tb-Lab | NO | TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30895 | DOH | jennypingkian23@yahoo.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 451 |  | GxIV-6C | 838562 | Module 1 H/W failed | Perform Genexpert Cleaning and maintenance Perform XPC Calibration at module 1 for troubleshooting Reporting of errors/problem Problem reported to Jotform Data collection | Perform Genexpert module troubleshooting | XPC Results: Module 1 passed 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6245953266866331842/6245953266866331842_signature_28.png | Jenny A.Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/6245953266866331842/6245953266866331842_signature_31.png
| 2025-06-01 | 2025-05-30 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30894 | DOH | dorothygundaya@gmail.com | 27 |  |  | GX-4 6 colors | 838403 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6245892794587616067/6245892794587616067_signature_28.png | Dorothy Gundaya | https://www.jotform.com/uploads/macarewecare/222927978372067/6245892794587616067/6245892794587616067_signature_31.png
| 2025-05-30 | 2025-05-30 | National Children's Hospital | E. Rodriquez Sr. Ave., Q.C. | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 30893 | MACARE-HBDC | michaeliza1205@gmail.com | 26 |  | 2025-05-29 | MCS + | 17H632SPG | Not responsive Keypads | Replaced display cable. Performed keypad test. Run functionlity test. | Replaced Diplay cable. Service part only. | Unit is operational. |  | Display cable assy. | DR #: 13056 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243996010529314941/6243996010529314941_signature_28.png | Michaeliza Bravante | https://www.jotform.com/uploads/macarewecare/222927978372067/6243996010529314941/6243996010529314941_signature_31.png
| 2025-05-30 | 2025-05-30 | Diit District Health Center | Diit Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 30892 | DOH | dalepenalosa189@gmail.com | 19 |  | 2025-05-30 | GXIV | 838421 | Module B3 temperature error | - | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243937745526755454/6243937745526755454_signature_28.png | Dale Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6243937745526755454/6243937745526755454_signature_31.png
| 2025-05-30 | 2025-05-29 | Tuburan Rural health unit | Tuburan, Basilan | BARMM | Basilan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30891 | DOH | rjakilan2@gmail.com | 28 |  | 2025-05-28 | GX 4 | 838832 | System unit unable to turn on. | Replaced Power supply of system unit and conducted Preventive Maintenence. | Replaced Power supply of system unit and conducted Preventive Maintenence. | All modules are available and system unit is working properly. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243924940743098136/6243924940743098136_signature_28.png | Raiyana Hassan Jakilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6243924940743098136/6243924940743098136_signature_31.png
| 2025-05-30 | 2025-05-30 | Bulacan Medical Center | Malolos, Bulacan | NCR | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30890 | UNOPS | bmcrhivda@gmail.com | 25 |  |  | GXIV-10C | 110023507 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243920316316507172/6243920316316507172_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6243920316316507172/6243920316316507172_signature_31.png
| 2025-05-30 | 2025-05-30 | Nueva Vizcaya Provincial Hospital | Bambamg, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30889 | DOH | dumlaojoy4@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1158 |  | GXIV-6C | 838272 | Module 1 error 2008 AVR not working | Troubleshoot Genexpert Check AVR Fuse | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | AVR for Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243869613114888616/6243869613114888616_signature_28.png | Bernadette Dumlao | https://www.jotform.com/uploads/macarewecare/222927978372067/6243869613114888616/6243869613114888616_signature_31.png
| 2025-05-30 | 2025-05-30 | Calubian Rural Health Unit | Calubian, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30888 | DOH | mjcabardo06@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 703 |  | 6C GXIV-4D | 838399 | None | None | -Preventive Maintenance -Data Gathering -Run XPC on module 1 | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243856205871781981/6243856205871781981_signature_28.png | Might Joy Cabardo | https://www.jotform.com/uploads/macarewecare/222927978372067/6243856205871781981/6243856205871781981_signature_31.png
| 2025-05-30 | 2025-05-30 | Solano Rural Health Unit | Solano, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30887 | DOH | rhusolanolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 749 |  | GXIV-6C | 838201 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243833533119480130/6243833533119480130_signature_28.png | Roderick Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6243833533119480130/6243833533119480130_signature_31.png
| 2025-05-30 | 2025-05-30 | SILLIMAN UNIVERSITY MEDICAL CENTER | Dumaguete City | 7 | Negros Oriental | MOLECULAR LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | bangalando23@gmail.com |  | 30886 | COMMERCIAL | sumcmoleclab@gmail.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 9 ASSAY: Xpert C. difficile BT, COUNT: 17 ASSAY: Xpert CT/NG, COUNT: 34 ASSAY: Xpert HBV VL, COUNT: 185 ASSAY: Xpert HPV, COUNT: 260 ASSAY: Xpert MTB/RIF Ultra, COUNT: 651 ASSAY: Xpert SARS-CoV-2, COUNT: 22 ASSAY: Xpert HCV VL, COUNT: 22 |  | GX-IV | 821471 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | MODULE A2 AND MODULE A3 DISABLED |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243808606823943125/6243808606823943125_signature_28.png | Maryjune Bangalando | https://www.jotform.com/uploads/macarewecare/222927978372067/6243808606823943125/6243808606823943125_signature_31.png
| 2025-05-29 | 2025-05-29 | Loveyour self | Cebu city | 7 | Cebu | HIV laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30885 | COMMERCIAL | paulo14reforzado@gmail.com | 24 |  |  | GX4 | 814150 | Due for preventive maintenance. | . Save installation qualification report. . Save system Log file . Save csv file | Perform preventive maintenance for gx4 | Module A1, A2 ok to use module A3, A4 not ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243102015847378580/6243102015847378580_signature_28.png | Paulo Reforzado | https://www.jotform.com/uploads/macarewecare/222927978372067/6243102015847378580/6243102015847378580_signature_31.png
| 2025-05-29 | 2025-05-29 | National Center for Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30884 | UNOPS | iancristrono@gmail.com | 27 |  |  | GXIV-10C | 120003480 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243060406318234053/6243060406318234053_signature_28.png | Carlo Martin Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6243060406318234053/6243060406318234053_signature_31.png
| 2025-05-29 | 2025-05-29 | Alicia Rural Health Unit | Alicia, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30883 | DOH | labalicia.rhu@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 689 |  | GXIV-6C | 838521 | UPS not powering On | Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243057631791764383/6243057631791764383_signature_28.png | Jeslie Raymunda Valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/6243057631791764383/6243057631791764383_signature_31.png
| 2025-05-29 | 2025-05-29 | Buhangin Health Center | NHA, Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30882 | DOH | buhanginlaboratory@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1211 |  | GX-IV 6C 4D | 838232 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243055398702273952/6243055398702273952_signature_28.png | Catherine B. Deocampo,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6243055398702273952/6243055398702273952_signature_31.png
| 2025-05-29 | 2025-05-29 | Maluso Rural Health Unit | Maluso, Basilan | BARMM | Basilan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30881 | DOH | sharminahussin@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 597 | 2025-05-29 | GX 4 | 838835 | Module 1 error 4012. UPS unable to work properly. | Preventive Maintenence. Replaced 2pcs. UPS batteries with 12v 7AH. Module 1 undergone Xpertcheck calibration. | Preventive Maintenence. Replaced 2pcs. UPS batteries with 12v 7AH. Module 1 undergone Xpertcheck calibration. | UPS needs to be charge approximately 8hrs. All modules are available. module 1 passed Xpertcheck Calibration. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6243023024485842425/6243023024485842425_signature_28.png | Sharmina A. Hussin | https://www.jotform.com/uploads/macarewecare/222927978372067/6243023024485842425/6243023024485842425_signature_31.png
| 2025-05-29 | 2025-05-29 | Infanta Rural Health Unit | Infanta Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30880 | UNOPS | charissealmandres8@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 208 |  | GXIV-D 10C | 110023501 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242985026011968654/6242985026011968654_signature_28.png | Charisse Alferez | https://www.jotform.com/uploads/macarewecare/222927978372067/6242985026011968654/6242985026011968654_signature_31.png
| 2025-05-29 | 2025-05-29 | Tumauini Community Hospital | Tumauini, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30879 | MACARE-HBDC | kayecatembung2911@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 |  | GXIV-6C | 110013709 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242981526711783275/6242981526711783275_signature_28.png | Mylene Claravall | https://www.jotform.com/uploads/macarewecare/222927978372067/6242981526711783275/6242981526711783275_signature_31.png
| 2025-05-29 | 2025-05-29 | Delfin Albanao Memorial Hospital | Delfin Albanao, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30878 | DOH | valorozos@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 788 |  | GXIV-6C | 838208 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242961426711297860/6242961426711297860_signature_28.png | Stephen Valorozo | https://www.jotform.com/uploads/macarewecare/222927978372067/6242961426711297860/6242961426711297860_signature_31.png
| 2025-05-29 | 2025-05-29 | Dasol Community Hospital | Dasol Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30877 | DOH | jeremypulido0521@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 564 |  | GXIV-D 6C | 838287 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242955546015961358/6242955546015961358_signature_28.png | Jeremy Pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/6242955546015961358/6242955546015961358_signature_31.png
| 2025-05-29 | 2025-05-29 | Santo Tomas Rural Health Unit | Bolinao, Santo Tomas, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30876 | MACARE-HBDC | gloangelinebagain@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2025-05-26 | GXIV-10C | 110024248 | Module 4 Failed Calibration | Performed Module Replacement B4(210127605 - 210043769) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Module | 13042 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242934126719861133/6242934126719861133_signature_28.png | Glo Angelie Bagain | https://www.jotform.com/uploads/macarewecare/222927978372067/6242934126719861133/6242934126719861133_signature_31.png
| 2025-05-29 | 2025-05-29 | Bolinao Rural Health Unit | Bolinao Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30875 | DOH | nylemae14@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 717 |  | GXIV-D | 838278 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242881216012147552/6242881216012147552_signature_28.png | Kathleen May Carolino | https://www.jotform.com/uploads/macarewecare/222927978372067/6242881216012147552/6242881216012147552_signature_31.png
| 2025-05-28 | 2025-05-28 | Pamplona Rural Health Unit | Centro, Pamplona, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30874 | UNOPS | wickizena@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1268 |  | GXIV-10C | 110023503 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242218662513085762/6242218662513085762_signature_28.png | Jelina Keith Ifurung | https://www.jotform.com/uploads/macarewecare/222927978372067/6242218662513085762/6242218662513085762_signature_31.png
| 2025-05-28 | 2025-05-28 | Tomas Claudio Health Center | Camus Street, Poblacion District Davao | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30873 | DOH | arleen_p_alconera@yahoo.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 845 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-IV 6C 4D | 838492 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242212299779952258/6242212299779952258_signature_28.png | Karen C. Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6242212299779952258/6242212299779952258_signature_31.png
| 2025-05-28 | 2025-05-28 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | Tb Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30872 | DOH | luzvimindanavidad2@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 930 |  | 6C GXIV-4D | 838297 | None | None | #NAME? | Machine is functional. Waiting for module 4 replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242194379912146424/6242194379912146424_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6242194379912146424/6242194379912146424_signature_31.png
| 2025-05-28 | 2025-05-28 | Lasam District Hospital | IBJ, Lasam, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30871 | DOH | lasamdh.laboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 424 |  | GXIV-6C | 838530 | PC not powering on | Troubleshoot PC Replaced Bios Battery | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242127842105191613/6242127842105191613_signature_28.png | Verbell Manuela | https://www.jotform.com/uploads/macarewecare/222927978372067/6242127842105191613/6242127842105191613_signature_31.png
| 2025-05-28 | 2025-05-28 | RIVERSIDE MEDICAL CENTER, INC. | BS Aquino Dr, Bacolod City | 6 | Negros Occidental | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | rumary.manito@medilinx.com.ph |  | 30870 | COMMERCIAL | rjmanito@gmail.com | 20 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 25 ASSAY: Xpert HBV VL, COUNT: 165 ASSAY: Xpert HPV, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 ASSAY: Xpert HCV VL, COUNT: 58 |  | GX-IV | 110005268 | #NAME? | #NAME? | #NAME? | #NAME? | MODULE B3 AND MODULE B4 FOR REPLACEMENT |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6242071709599171582/6242071709599171582_signature_28.png | Rumary Joan Manito | https://www.jotform.com/uploads/macarewecare/222927978372067/6242071709599171582/6242071709599171582_signature_31.png
| 2025-05-28 | 2025-05-27 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Apheresis room | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 30869 | MACARE-HBDC | zarabellehilario@yahoo.com | 26 |  |  | MCS + | 17H632SPG | Unresponsive keypad Defective display cable | Checked unit Reset dispaly cable Performed system reset | Performed troubleshooting | After performed troubleshooting still the unit is unresponsive keypad when pressing the button. Need to replace display cable For quotation of display cable and line sensor |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6241999463614803608/6241999463614803608_signature_28.png | Zarabelle Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/6241999463614803608/6241999463614803608_signature_31.png
| 2025-05-28 | 2025-05-27 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | eamc.biomed@gmail.com |  | 30868 | MACARE-HBDC | blooddonorsection@gmail.com | 28 |  |  | MCS + | 15H044 | As per end user they encounter red valve error | Performed valve testing Performed simulation running | Performed troubleshooting | During testing no error encountered. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6241997903614746002/6241997903614746002_signature_28.png | Sheena Kate Pascua | https://www.jotform.com/uploads/macarewecare/222927978372067/6241997903614746002/6241997903614746002_signature_31.png
| 2025-05-27 | 2025-05-27 | Pilar Rural Health Unit | Calongay, Pilar | 5 | Sorsogon | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30867 | DOH | manalangmarkramir@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1815 |  | GX4-6C | 838266 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6241394159322526440/6241394159322526440_signature_28.png | Mark ramir c manalang jr | https://www.jotform.com/uploads/macarewecare/222927978372067/6241394159322526440/6241394159322526440_signature_31.png
| 2025-05-27 | 2025-05-27 | The Medical City Clark | Clark Pampanga | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | tmcclarkbiomed@themedicalcity.com |  | 30866 | COMMERCIAL | tmci_laboratory@themedicalcity.com | 25 | ASSAY: Xpert CT/NG, COUNT: 15 ASSAY: Xpert HBV VL, COUNT: 59 ASSAY: Xpert HPV, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 126 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GXIV-D 6C | 839235 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module 4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6241360098123353199/6241360098123353199_signature_28.png | Jessa Rae Remorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6241360098123353199/6241360098123353199_signature_31.png
| 2025-05-27 | 2025-05-27 | La Castellana RHU | La Castellana, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30865 | DOH | milouepascual@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 708 |  | GX4 6C | 838483 | Defective UPS - turns off abruptly during power interruptions. | Preventive Maintenance | - Replaced UPS batteries with 2 pcs. new batteries 12V 9AH. - Functional testing of UPS | - UPS with power backup when line power is removed - 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6241337313746213331/6241337313746213331_signature_28.png | Mira Luna Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/6241337313746213331/6241337313746213331_signature_31.png
| 2025-05-27 | 2025-05-27 | Ciudad Medical Zamboanga | Ciudad Medical Zamboanga, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30864 | DOH | thairaphoebegrey13@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 54 ASSAY: Xpert MTB/RIF Ultra, COUNT: 466 ASSAY: Xpert HCV VL, COUNT: 19 | 2025-05-27 | GX 4 | 834128 | None | Preventive Maintenence and data collection | Preventive Maintenence and data collection | Module 1 is available. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6241305740717533950/6241305740717533950_signature_28.png | Shereemer A.Cheong-Ynawat | https://www.jotform.com/uploads/macarewecare/222927978372067/6241305740717533950/6241305740717533950_signature_31.png
| 2025-05-27 | 2025-05-27 | Philippine Red Cross - Cagayan de Oro | Cagayan de Oro City | 10 | Misamis Oriental | Blood Center | NO | REPAIR | HAEMONETICS | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30863 | COMMERCIAL | mrbc2@redcross.org | 27 |  |  | MCS+ | 02H080 | White Spot found on LCD. Not clear/unreadable screen. | None | Replacing LCD panel board. Preventive Maintenance | Unit is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6241236195128933076/6241236195128933076_signature_28.png | Phoenix James Dahili | https://www.jotform.com/uploads/macarewecare/222927978372067/6241236195128933076/6241236195128933076_signature_31.png
| 2025-05-27 | 2025-05-27 | Jose B. Lingad Memorial and General Hospital | Dolores City of San Fernando Pampanga | 3 | Pampanga | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30862 | MACARE-HBDC | jblmghmpl@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 871 ASSAY: Xpert MTB-XDR, COUNT: 173 |  | GXIV-D 10C | 110012825 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6241209938028482525/6241209938028482525_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/6241209938028482525/6241209938028482525_signature_31.png
| 2025-05-27 | 2025-05-27 | Jose B. Lingad Memorial and General Hospital | Dolores City of San Fernando, Pampanga | 3 | Pampanga | RHIVDA | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30861 | MACARE-HBDC | jblmgh.is@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 304 ASSAY: Xpert HCV VL, COUNT: 77 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2000 |  | GXIV-D 10C | 110013748 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6241195298023561167/6241195298023561167_signature_28.png | Charlene Mae Laxamana | https://www.jotform.com/uploads/macarewecare/222927978372067/6241195298023561167/6241195298023561167_signature_31.png
| 2025-05-26 | 2025-05-23 | Best Diagnostic Corporation | Masikap ext QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30860 | COMMERCIAL | bestdiagnosticlaboratory@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 ASSAY: Xpert HCV VL, COUNT: 2 |  | GX4D | 110008529 | Due for maintenance Due for xpertcheck last xpertcheck 2/7/2024 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6240417833617590475/6240417833617590475_signature_28.png | Donrichmond Catibod | https://www.jotform.com/uploads/macarewecare/222927978372067/6240417833617590475/6240417833617590475_signature_31.png
| 2025-05-26 | 2025-05-22 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | eamc.biomed@gmail.com |  | 30859 | UNOPS | eamcmolecularlab@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 983 ASSAY: Xpert HPV, COUNT: 93 ASSAY: Xpert SARS-CoV-2, COUNT: 602 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 366 | 2025-05-07 | GX4 D 10C | 110016781 | Due for xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpert check | All module passed in xpertcheck |  | Xpertcheck | 12994 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6240417453619398322/6240417453619398322_signature_28.png | East Avenue Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6240417453619398322/6240417453619398322_signature_31.png
| 2025-05-26 | 2025-05-20 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30858 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 51 ASSAY: Xpert BCR-ABL Ultra, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1266 ASSAY: Xpert SARS-CoV-2, COUNT: 20 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1492 | 2025-05-07 | GX16D 10C | 110014171 | For re xpertcheck C2 210069271,D1 614214,D2 210078182,D3 210025172,D4 210069227 | Performed xpertcheck Cleaned plunger,i core Gathered reports | Performed re xpertcheck | After performed rexpert check the status of module C2,D1,D2,D3,D4 is retry in xpertcheck Disabled module A1,A2,A3,A4,B1,B2,B3,B4,C1,C2,D1,D2,D3,D4 |  | Xpertcheck | 12996 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6240417073618643390/6240417073618643390_signature_28.png | National Children's Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6240417073618643390/6240417073618643390_signature_31.png
| 2025-05-26 | 2025-05-23 | San Fernando Rural Health Unit | Poblacion, San Fernando | 5 | Camarines Sur | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30857 | DOH | rhusanfernando.2018@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 863 |  | GX4-6C | 838496 | GX4-6C for Preventive Maintenance Request for extra fan filtee | Replaced fan filter Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6240395316421043095/6240395316421043095_signature_28.png | Kim Jason Vega | https://www.jotform.com/uploads/macarewecare/222927978372067/6240395316421043095/6240395316421043095_signature_31.png
| 2025-05-26 | 2025-05-22 | Cotabato City Health Office | Poblacion, Cotabato | BARMM | Maguindanao | TB DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30856 | DOH | onehope834@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 864 |  | GX-IV 6C 4D | 838528 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6240385872118508276/6240385872118508276_signature_28.png | EMELYN A. CADIZ | https://www.jotform.com/uploads/macarewecare/222927978372067/6240385872118508276/6240385872118508276_signature_31.png
| 2025-05-26 | 2025-05-22 | Cotabato City Health Office | Poblacion, Cotabato | 11 | Maguindanao | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30855 | UNOPS | onehope834@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 864 |  | GX-IV 10C 4D | 110023524 | For Preventive Maintenance | All Modules are operational | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6240385702115872990/6240385702115872990_signature_28.png | Emelyn Cadiz,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6240385702115872990/6240385702115872990_signature_31.png
| 2025-05-23 | 2025-05-23 | Rizal Medical Center | Pasig Blvd. Pasig City | NCR | NCR 2nd - Pasig | Molecular | YES | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | princeszvillaran13@gmail.com |  | 30854 | MACARE-HBDC | molecularlab@gmail.com | 22 | ASSAY: Others, COUNT: 2 ASSAY: Xpert HPV, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 313 ASSAY: Xpert SARS-CoV-2, COUNT: 52 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 9 | 2025-05-08 | GX-XVI-D-6C | 838627 | For module replacement | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | Unit is operational. |  | New modules: 220037444 220038019 220038040 220038027 220038045 220038099 220038062 | SI# 100252 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237971937072490027/6237971937072490027_signature_28.png | Justine Lorenz V Falzado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6237971937072490027/6237971937072490027_signature_31.png
| 2025-05-23 | 2025-05-23 | St Luke's Medical Center QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 30853 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 346 ASSAY: Xpert HPV, COUNT: 426 ASSAY: Xpert MTB/RIF Ultra, COUNT: 188 ASSAY: Xpert HCV VL, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 157 |  | GX4L | 807155 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | C1,C2,C3,in use C4 disabled due to error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237933223323225891/6237933223323225891_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6237933223323225891/6237933223323225891_signature_31.png
| 2025-05-23 | 2025-05-23 | The Medical City - South Luzon | Santa Rosa, Laguna | 4A | Laguna | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30852 | COMMERCIAL | tmcsl-molecularlab@themedicalcity.com | 25 |  |  | GXIV | 839233 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. | Module B2, service required. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237931238412579599/6237931238412579599_signature_28.png | ARJEANNE ASTEJADA | https://www.jotform.com/uploads/macarewecare/222927978372067/6237931238412579599/6237931238412579599_signature_31.png
| 2025-05-23 | 2025-05-23 | St.Lukes Medical Center QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 30851 | COMMERCIAL | megapo@stlukes.com.ph | 29 | ASSAY: Xpert CT/NG, COUNT: 198 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2711 |  | GX4 D | 839234 | Due for maintenance Last xpertcheck march 2023 Failed xpertcheck B3 866517,B4 866800 last march 2023 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B3,B4 disabled B1,B2 in use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237898883328724524/6237898883328724524_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6237898883328724524/6237898883328724524_signature_31.png
| 2025-05-23 | 2025-05-23 | Project 7 Social Hygiene Clinic | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30850 | MACARE-HBDC | michelenaval@gmail.com | 25 |  |  | GXIV-10C | 120000577 | Unit for maintenance | NA | #NAME? | Unit is operatio al |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237837334025631403/6237837334025631403_signature_28.png | Michele Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6237837334025631403/6237837334025631403_signature_31.png
| 2025-05-23 | 2025-05-23 | San Pablo Colleges Medical Center | San Pablo Laguna | 4A | Laguna | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30849 | CONSORTIUM | spcmc.molecularlab@gmail.com | 25 |  |  | GXIV-10C | 110024552 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237803456616023201/6237803456616023201_signature_28.png | RICHARD E. TEMPLO | https://www.jotform.com/uploads/macarewecare/222927978372067/6237803456616023201/6237803456616023201_signature_31.png
| 2025-05-23 | 2025-05-22 | MMC Medilinx | Amorsolo st Makati | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30848 | COMMERCIAL | janeshyrelenriquez.medilinx@outlook.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 58 ASSAY: Xpert CT/NG, COUNT: 142 ASSAY: Xpert HBV VL, COUNT: 36 ASSAY: Xpert HPV, COUNT: 66 ASSAY: Xpert MTB/RIF Ultra, COUNT: 550 ASSAY: Others, COUNT: 4 ASSAY: Xpert HCV VL, COUNT: 24 ASSAY: Xpert HIV-1 Viral Load, COUNT: 312 | 2025-05-19 | GX16 D 10C | 110026448 | For module replacement C4 210135730 - 690272 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed module replacement | A1,C1,C2 disabled reported already 13 modules operational |  | Genexpert module 10C | 13029 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237748822676968402/6237748822676968402_signature_28.png | Jane Shyrel M. Enriquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6237748822676968402/6237748822676968402_signature_31.png
| 2025-05-22 | 2025-05-22 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30847 | MACARE-HBDC | bmc.bacteriology@gmaol.com | 27 | ASSAY: Xpert C. difficile BT, COUNT: 4 ASSAY: Others, COUNT: 14 ASSAY: Xpert HIV-1 Viral Load, COUNT: 143 ASSAY: Others, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 108 ASSAY: Xpert SARS-CoV-2, COUNT: 44 |  | GX4-6C | 110008458 | Date and time resets | Replaced Bios battery of CPU Cleaned cartridge bays, plunger rod Generate Data | Performed Troubleshooting Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237140658419101936/6237140658419101936_signature_28.png | Mary elaine tolidanes | https://www.jotform.com/uploads/macarewecare/222927978372067/6237140658419101936/6237140658419101936_signature_31.png
| 2025-05-22 | 2025-05-22 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | REPAIR | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com | bmc.bloodbank@gmail.com |  | 30846 | MACARE-HBDC | collaoleo17@gmail.com | 25 |  |  | MCS+ | 15H027G | Machine screen no image Defective display cable | Replaced display cable(functional) | Replaced display cable | Unit is non operational CRC Bios error | MCS+ CPU request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237096019121561470/6237096019121561470_signature_28.png | LEO C. COLLAO | https://www.jotform.com/uploads/macarewecare/222927978372067/6237096019121561470/6237096019121561470_signature_31.png
| 2025-05-22 | 2025-05-22 | Poro Rural Health Unit | Poro, Camotes Island, Cebu | 7 | Cebu | TB Dots | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30845 | DOH | avilacharmaine6@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 445 |  | 6C GXIV-4D | 838279 | #NAME? | #NAME? | #NAME? | Module B2 for replacement. Other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6237006618503659804/6237006618503659804_signature_28.png | Charmaine Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/6237006618503659804/6237006618503659804_signature_31.png
| 2025-05-22 | 2025-05-22 | Valenzuela Social Hygiene Clinic | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30844 | MACARE-HBDC | markanthony.manaois@gmail.com | 25 |  |  | GXIV-10C | 110020699 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236990842748639032/6236990842748639032_signature_28.png | Mark Manaois | https://www.jotform.com/uploads/macarewecare/222927978372067/6236990842748639032/6236990842748639032_signature_31.png
| 2025-05-22 | 2025-05-15 | Tuburan rural health unit | Tuburan basilan | BARMM | Basilan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30843 | DOH | rjakilan20@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2025-05-15 | GX 4 | 838832 | UPS unable to work properly. System unit unable to turn on. | Replaced 2 UPS batteries with 12V 7Ah. Pulled out System unit. Preventive Maintenance. | Replaced 2 UPS batteries with 12V 7Ah. Pulled out System unit. Preventive Maintenance. | System unit under repair. 2 new batteries with 12v 7ah in the repaired UPS. | System unit ongoing repair , pulled out at the moment. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236851010291621151/6236851010291621151_signature_28.png | Raiyana Hassan Jakilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6236851010291621151/6236851010291621151_signature_31.png
| 2025-05-21 | 2025-05-21 | Metro Health Specialists Hospital Inc. | Brgy. Cabid-an, East District, Sorsogon City | 5 | Sorsogon | Laboratory | YES | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30842 | COMMERCIAL | mhshilab@gmail.com | 29 |  |  | GX4-10C | 120002716 | Gx for Installation | Installed Gx4-10C (120002716), CPU (CPH0424F51338), Monitor (CN-0WMYC5-FCC00-426-C08X-A06), UPS (AT1K240705ST0031), Barcode Scanner (24068010552108), Printer (CN4AH427D0) Setup/adjusted Data and information for Gx parameters/setup Collect data | Installed GeneXpert 10C | Unit is Operational | Installed and rebox(laboratory under construction) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236269558508513443/6236269558508513443_signature_28.png | Aileen V. Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6236269558508513443/6236269558508513443_signature_31.png
| 2025-05-21 | 2025-05-15 | Marikina City Health Office | Sumulong highway Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30841 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 496 ASSAY: Xpert MTB-XDR, COUNT: 50 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236251313614684192/6236251313614684192_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6236251313614684192/6236251313614684192_signature_31.png
| 2025-05-21 | 2025-05-15 | Marikina City Health Office | Sumulong highway Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30840 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 9i6 ASSAY: Xpert MTB-XDR, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236251163613087048/6236251163613087048_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6236251163613087048/6236251163613087048_signature_31.png
| 2025-05-21 | 2025-05-15 | Marikina City Health Office | Sumulong highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30839 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 135 ASSAY: Xpert HIV-1 Viral Load, COUNT: 130 |  | GX-IV-D-10C | 110024072 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  | DR# 12995 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250983613231484/6236250983613231484_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250983613231484/6236250983613231484_signature_31.png
| 2025-05-21 | 2025-05-15 | Marikina City Health Office | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30838 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 137 ASSAY: Xpert MTB-XDR, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 174 |  | GX-IV-D-10C | 110024073 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  | DR# 12995 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250853617170037/6236250853617170037_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250853617170037/6236250853617170037_signature_31.png
| 2025-05-21 | 2025-05-14 | St. Luke's Medical Center - BGC | Rizal Dr BGC Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 30837 | COMMERCIAL | jmtentica@stlukes.com.ph | 22 |  |  | MCS+ | 09C125 | For calibration | Performed calibration | Performed maintenance Checked rotor pumps Run simulation Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250563617693705/6236250563617693705_signature_28.png | Jen Entica | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250563617693705/6236250563617693705_signature_31.png
| 2025-05-21 | 2025-05-14 | St. Luke's Medical Center - BGC | Rizal Dr. BGC Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 30836 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 33 ASSAY: Xpert CT/NG, COUNT: 139 ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert HPV, COUNT: 186 ASSAY: Xpert MTB/RIF Ultra, COUNT: 397 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 92 |  | GX-IV-D-6C | 839232 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A3 for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250273617857432/6236250273617857432_signature_28.png | Joshua Aya-ay | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250273617857432/6236250273617857432_signature_31.png
| 2025-05-21 | 2025-05-14 | St. Luke's Medical Center - BGC | Rizal Dr. BGC Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 30835 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 160 ASSAY: Xpert CT/NG, COUNT: 624 ASSAY: Xpert HBV VL, COUNT: 52 ASSAY: Xpert HPV, COUNT: 846 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2094 ASSAY: Xpert MTB-XDR, COUNT: 5 ASSAY: Xpert HCV VL, COUNT: 39 ASSAY: Xpert HIV-1 Viral Load, COUNT: 242 |  | GX-IV-D-10C | 120002483 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250113618238639/6236250113618238639_signature_28.png | Joshua Aya-ay | https://www.jotform.com/uploads/macarewecare/222927978372067/6236250113618238639/6236250113618238639_signature_31.png
| 2025-05-21 | 2025-03-18 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30834 | COMMERCIAL | mahmci_labcmt@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 21 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 282 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 |  | GX-IV-D-6C | 110005265 | Error on module B4. UPS no backup. | Checked UPA and connections. | Check UPS backup power. Check Connections Run reliability test Performed maintenance Generated reports Generated gxx file | Unit is operational. UPS has no backup power. | Recommended to replace UPS. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236249293619900735/6236249293619900735_signature_28.png | Graziela Marie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/6236249293619900735/6236249293619900735_signature_31.png
| 2025-05-21 | 2025-05-15 | Pasig Treatment Hub - Manggahan | Kaginhawaan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30833 | MACARE-HBDC | pathmedtech@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 74 |  | GX-IV-L-6C | 821648 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236249133614950690/6236249133614950690_signature_28.png | Katherine Godoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6236249133614950690/6236249133614950690_signature_31.png
| 2025-05-21 | 2025-05-20 | St. Luke's Medical Center - BGC | Rizal Dr BGC Taguig City | NCR | NCR 4th - Taguig | Blood bank | NO | Change time | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 30832 | COMMERCIAL | jmtentica@stlukes.com.ph | 22 |  |  | MCS+ | 09C125 | Change time | Adjust time of machine | Adjust time of machine to standard time Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236244223617647351/6236244223617647351_signature_28.png | Jen Michelle Entica | https://www.jotform.com/uploads/macarewecare/222927978372067/6236244223617647351/6236244223617647351_signature_31.png
| 2025-05-21 | 2025-05-21 | Cotabato Regional and Medical Center - Blood Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Blood Center | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Jesson Tupas |  | jessontupas.macaremed@gmail.com | hkdrdsales@gmail.com |  | 30831 | COMMERCIAL | bloodcenter@ro12.doh.gov.ph | 23.7 |  |  | MCS+ | 15H042G | Line sensor out of range and need to replace pump rollers | Install Line Sensor Assembly (service unit only), to be pulled out upon installation of new Line Sensor Assembly. And replaced pump rotors with rollers. Repair/replace Defective line sensor SN:SPO41504257G with service unit line sensor SN:SP121703214G Perform cycle test | Line Sensor and pump rollers repair | Fixed/repair line sensor and pump rotors with rollers Unit is Operational | Requesting for quotation for new line sensor assembly. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236235491818176056/6236235491818176056_signature_28.png | Rhea Lim-Macion, RMT, MSMLS | https://www.jotform.com/uploads/macarewecare/222927978372067/6236235491818176056/6236235491818176056_signature_31.png
| 2025-05-21 | 2025-05-21 | Cotabato Regional and Medical Center - Blood Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Blood Center | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Jesson Tupas |  | jessontupas.macaremed@gmail.com | hkdrdsales@gmail.com |  | 30830 | COMMERCIAL | bloodcenter@ro12.doh.gov.ph | 23.7 |  |  | MCS+ | 15HO42G | Line Sensor problem and pump rollers | Perform MCS+ Cleaning and maintenance Perform Calibration Checking of errors/problem | Preventive maintenance MCS+ Calibration | Line Sensor value out of range |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236224231819098409/6236224231819098409_signature_28.png | Rhea Lim-Macion, RMT, MSMLS | https://www.jotform.com/uploads/macarewecare/222927978372067/6236224231819098409/6236224231819098409_signature_31.png
| 2025-05-21 | 2025-05-21 | Chinese General Hospital | Blumentritt, Sta. Cruz Manila | NCR | NCR 1st - Manila | Bacteriology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 30829 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1595 ASSAY: Xpert HIV-1 Viral Load, COUNT: 147 |  | GXIV-10C | 120002026 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236218076437095265/6236218076437095265_signature_28.png | Jeri Elizr P. Berongoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6236218076437095265/6236218076437095265_signature_31.png
| 2025-05-21 | 2025-05-21 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30828 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2669 |  | 6C GXIV-4D | 838285 | Due for PM | -Perform preventive maintenance; fan filter, plunger rod, optic slot, cartridge bay | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236154123181600935/6236154123181600935_signature_28.png | Eldric Roel Sebastian Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/6236154123181600935/6236154123181600935_signature_31.png
| 2025-05-21 | 2025-05-21 | Central Luzon Doctir's Hospital | Hospital drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30827 | COMMERCIAL | pjlsantos@cldh.com.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 200 |  | GXIV-D | 839232 | Monthly preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational B1: 742501 Temperature Error B4: 738203 Temperature error B1 and B4 for module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236148156832866106/6236148156832866106_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6236148156832866106/6236148156832866106_signature_31.png
| 2025-05-21 | 2025-05-21 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30826 | UNOPS | jomarilitonjua@gmail.com | 25 |  |  | GXIV-10C | 110023495 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236147762125336212/6236147762125336212_signature_28.png | Cynthia Bernante | https://www.jotform.com/uploads/macarewecare/222927978372067/6236147762125336212/6236147762125336212_signature_31.png
| 2025-05-21 | 2025-05-21 | Central Luzon Doctor's Hospital | Hospital drive Brgy. Sa Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 30825 | COMMERCIAL | pjlsantos@cldh.com.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 200 |  | GXII-D | 830841 | Monthly preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236147696832540775/6236147696832540775_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6236147696832540775/6236147696832540775_signature_31.png
| 2025-05-21 | 2025-05-21 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30824 | MACARE-HBDC | treatmenthub.qcgh@gmail.com | 25 |  |  | GXIV-10C | 110012766 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236141512123307694/6236141512123307694_signature_28.png | Darnel F. Bustillos | https://www.jotform.com/uploads/macarewecare/222927978372067/6236141512123307694/6236141512123307694_signature_31.png
| 2025-05-21 | 2025-05-21 | Kairos Laboratories and Diagnostics, Inc. | A. Mabini St. Ermita Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | kairosermita.lab@gmail.com |  | 30823 | CONSORTIUM | kairos.djreingen@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 120 |  | GXIV (2 modules) | 110022622 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236074616436929103/6236074616436929103_signature_28.png | Denise Joyce L. Reiñgen | https://www.jotform.com/uploads/macarewecare/222927978372067/6236074616436929103/6236074616436929103_signature_31.png
| 2025-05-21 | 2025-05-21 | Baybay City Health Office | Baybay Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30822 | DOH | sajecopooh20@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1203 | 2025-05-21 | GXIV | 838555 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236064164537409769/6236064164537409769_signature_28.png | sarah jean gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/6236064164537409769/6236064164537409769_signature_31.png
| 2025-05-21 | 2025-05-21 | Cebu City health office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30821 | UNOPS | laureanjo@yahoo.com | 25 |  |  | Gx4 | 110023513 | No problem reported | Gather CSV files from 2024 to 2025 Gather data Clean fan filter | Gather CSV files from 2024 to 2025 Gather data Clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236055697972739897/6236055697972739897_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6236055697972739897/6236055697972739897_signature_31.png
| 2025-05-21 | 2025-05-21 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30820 | DOH | laureanjo@yahoo.com | 25 |  |  | Gx 4 | 838527 | No problem reported | Gather CSV files from 2019 to 2025 Clean fan filter Gather data | Gather CSV files from 2019 to 2025 Clean fan filter Gather data | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6236053817974192008/6236053817974192008_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6236053817974192008/6236053817974192008_signature_31.png
| 2025-05-20 | 2025-05-15 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30819 | DOH | suzzettekho@gmail.com | 27 |  |  | GX4-6 colors | 838554 | None | Maintenance | Preventive Maintenance | All module and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6235587665952981753/6235587665952981753_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/6235587665952981753/6235587665952981753_signature_31.png
| 2025-05-20 | 2025-05-20 | Pio Duran Rural Health Uniy | Caratagan, Pio Duran | 5 | Albay | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30818 | DOH | mitzijc@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 535 |  | GX4-6C | 838264 | UPS overheat | Checked fuse(busted) Checked UPS batteries(1battery busted) Cleaned cartridge bays, plunger rod Data Collection | Performed UPS Troubleshooting Performed Preventive Maintenance | Unit is Operational UPS defective | Requesting for UPS replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6235327333115177245/6235327333115177245_signature_28.png | Mitzi D. Cedeño | https://www.jotform.com/uploads/macarewecare/222927978372067/6235327333115177245/6235327333115177245_signature_31.png
| 2025-05-20 | 2025-05-20 | Graman Medical and Maternity Hospital | Malolos, Bulacan | NCR | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30817 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  |  | GXIV-10C | 110008530 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Unit is due for calibration (12.22.24) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6235251492212173598/6235251492212173598_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6235251492212173598/6235251492212173598_signature_31.png
| 2025-05-20 | 2025-05-20 | Socorro Rural Health Unit | Brgy. Navarro, Socorro | 13 | Surigao del Norte | RTDL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30816 | UNOPS | dynamite.tux@gmail.com | 28 |  |  | GX4-10 Colors | 110023523 | Issue found in Module B3. Module not detected in GX software. Orange Indicator. Error due to communication. | Troubleshooting. Sending pictures and documents to Main Office. Maintenance | Troubleshooting Module B3 including connection reset and swapping B3 to another location. Preventive Maintenance | 3 modules working, 1 module for replacement. Gx unit peripherals are working. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6235131630814885643/6235131630814885643_signature_28.png | James Anastacio | https://www.jotform.com/uploads/macarewecare/222927978372067/6235131630814885643/6235131630814885643_signature_31.png
| 2025-05-19 | 2025-05-15 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30815 | DOH | celinejean16@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1593 |  | GxIV-6C | 838570 | None | Perform Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Filter cleaning Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234719169811743295/6234719169811743295_signature_28.png | Celine Jean villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6234719169811743295/6234719169811743295_signature_31.png
| 2025-05-19 | 2025-05-19 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | Iloilo City | 6 | Iloilo | MOLECULAR LAB | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30814 | DOH | mc-molecular@wvsu.edu.ph | 28 | ASSAY: Xpert HBV VL, COUNT: 63 ASSAY: Xpert HPV, COUNT: 22 ASSAY: Xpert MTB-XDR, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 |  | GX-IV | 110021716 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234560886027992952/6234560886027992952_signature_28.png | Karla Danielle Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6234560886027992952/6234560886027992952_signature_31.png
| 2025-05-19 | 2025-05-16 | Pasig Treatment Hub - Klinika Pasig (Sumilang) | Dr. Garcia St. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30813 | MACARE-HBDC | pasigtreatment@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 297 |  | GX-IV-D-10C | 110025128 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234531493614763435/6234531493614763435_signature_28.png | Torate, Melanio M | https://www.jotform.com/uploads/macarewecare/222927978372067/6234531493614763435/6234531493614763435_signature_31.png
| 2025-05-19 | 2025-05-16 | Pasig Treatment Hub - CHAMP | Champ bldg. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | dansumulong.pasigcho@gmail.com |  | 30812 | MACARE-HBDC | pathmedtech@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 318 |  | GX-IV-D-10C | 110025127 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234531483613790594/6234531483613790594_signature_28.png | Dan nathaniel sumulong | https://www.jotform.com/uploads/macarewecare/222927978372067/6234531483613790594/6234531483613790594_signature_31.png
| 2025-05-19 | 2025-05-16 | Klinika Eastwood | 3rd Level CityWalk, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30811 | MACARE-HBDC | clarabaptista2786@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 64 |  | GX-IV-D-10C | 110020698 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234531483615642597/6234531483615642597_signature_28.png | Mary Claire P Baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/6234531483615642597/6234531483615642597_signature_31.png
| 2025-05-19 | 2025-05-16 | Rosario Rural Health Unit | Rosario Municipal Compound, Rosario, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30810 | DOH | arlenepenaarmijo@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1166 |  | GXIV-D-6C | 838537 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Archive Test Results. Checked System's Functionality. Generate Reports. | All Modules are Good. UPS is Defective. | S/N: AT1000190123ST0255 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233703619306519/6234233703619306519_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233703619306519/6234233703619306519_signature_31.png
| 2025-05-19 | 2025-05-15 | UST Hospital | Dapitan St, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30808 | MACARE-HBDC | bloodbank@usthospital.com.ph | 25 |  |  | MCS+ | 02F069 | Unit is due for Maintenance. | Preventive Maintenance. | Checked Diagnostic Values. Cleaned Pump Rollers. Cleaned Optics and Sensors. Checked Machine's Functionality. Performed Simulation Run. | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653615001488/6234233653615001488_signature_28.png | Mylene Balboa | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653615001488/6234233653615001488_signature_31.png
| 2025-05-19 | 2025-04-30 | Ospital Ng Muntinlupa | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30809 | UNOPS | jjpsanchez21@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 562 |  | GXIV-D-10C | 110023494 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653619561469/6234233653619561469_signature_28.png | Jesse Jasper Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653619561469/6234233653619561469_signature_31.png
| 2025-05-19 | 2025-05-14 | Capitol Medical Center | Q. Ave, Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | Check-up | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30806 | MACARE-HBDC | gfalcantara.rmt@gmail.com | 25 |  |  | MCS+ | 07C002 | For Checkup of Apheresis Unit. | Checkup | Checked Machine Functionality. Checked Rollers. Checked Sensor Values. Checked Optic Values. | Pump Rollers for Replacement. Unit is due for Calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653616564137/6234233653616564137_signature_28.png | Gretha Faye Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653616564137/6234233653616564137_signature_31.png
| 2025-05-19 | 2025-05-09 | Dasmariñas City Health Office II | DBB-E, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30805 | DOH | hazelmilan0777@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2376 |  | GXIV-D-6C | 838535 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. | Unit is due for XpertCheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653619051610/6234233653619051610_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653619051610/6234233653619051610_signature_31.png
| 2025-05-19 | 2025-05-14 | Capitol Medical Center | Q. Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30807 | CONSORTIUM | princeosena@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 402 |  | GXIV-D-10C | 110025453 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. GX System is Functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653619567702/6234233653619567702_signature_28.png | Prince Henry Osena | https://www.jotform.com/uploads/macarewecare/222927978372067/6234233653619567702/6234233653619567702_signature_31.png
| 2025-05-19 | 2025-05-13 | Batasan Social hygiene clinic | Batasan Rd QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30804 | MACARE-HBDC | moniqueteves99@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 32 ASSAY: Xpert HBV VL, COUNT: 50 ASSAY: Xpert HIV-1 Viral Load, COUNT: 855 |  | GX4 10C | 110020697 | Due for maintenance As per end user they encounter loss communication and unable to detect module | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit Reset lan cable | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6234218843619597629/6234218843619597629_signature_28.png | Monique Teves | https://www.jotform.com/uploads/macarewecare/222927978372067/6234218843619597629/6234218843619597629_signature_31.png
| 2025-05-17 | 2025-05-16 | Matanao Rural Health Unit | Poblacion Matanao | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30803 | DOH | pamsustento@gmail.com | 26.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 481 |  | GX-IV 6C 4D | 8383456 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are opeeational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6232992625517437877/6232992625517437877_signature_28.png | Dominador S. Castro III | https://www.jotform.com/uploads/macarewecare/222927978372067/6232992625517437877/6232992625517437877_signature_31.png
| 2025-05-16 | 2025-05-16 | Makati Medical Center | Amorsolo st.,Makati | NCR | NCR 4th - Makati | CRM | YES | MAINTENANCE | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | alohakae.jabat@makatimed.net.ph | Miltenyi | 30802 | COMMERCIAL | AlcorDeeond.Padilla@makatimed.net.ph | 28 |  | 2025-05-15 | Gentlemacs octo dissociator | 1261 | Due for maintenance | Cleaned unit and surface Checked wear and tear parts Test lcd display Checked program parameters Checked connectors Performed functionality test | Performed maintenance | Operational |  |  | 100283 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6231849719329946788/6231849719329946788_signature_28.png | Alcor Deeond F. Padilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6231849719329946788/6231849719329946788_signature_31.png
| 2025-05-16 | 2025-05-16 | La Castellana Rural Health Unit | La Castellana, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30801 | DOH | milouepascual@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 708 |  | GX4 6C | 838483 | Defective UPS, no power backup during power interruptions. | . | #NAME? | UPS batteries for replacement, 2 pcs 12 V 7AH. UPS SN AT100019042ST0082 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6231835052816523870/6231835052816523870_signature_28.png | Mira Luna Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/6231835052816523870/6231835052816523870_signature_31.png
| 2025-05-16 | 2025-05-15 | Jose B. Lingad Memorial and General Hospital | Dolores City of Sa Fernando Pampanga | 3 | Pampanga | RHIVDA | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | jblmgh.is@gmail.com |  | 30800 | MACARE-HBDC | jblmgh.is@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 297 ASSAY: Xpert HCV VL, COUNT: 80 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1891 | 2025-05-08 | GXIV-D 10C | 110013748 | Defective UPS | Preventive maintenance UPS replacement | Performed Preventive maintenance Replacement of Defective UPS Gathered reports Performed functionality test | Unit is operational |  | UPS | 13006 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6231804565813341699/6231804565813341699_signature_28.png | CHARLENE MAE LAXAMANA | https://www.jotform.com/uploads/macarewecare/222927978372067/6231804565813341699/6231804565813341699_signature_31.png
| 2025-05-16 | 2025-05-16 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com | bmc.bloodbank@gmail.com |  | 30799 | MACARE-HBDC | collaoleo17@gmail.com | 25 |  |  | MCS+ | 15H031G | Leaked fluid inside centrifuge basin | Cleaned Centrifuge Cleaned Unit Check the rpm of centrifuge Checked the stability/steadiness of centrifuge while rotating Checked parameters | Performed Troubleshooting Performed Preventive Maintenance | Unit is operational | No humming sound during centrifuge test, centrifuge placement is stable during test |  |  |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/6231745398117302103/6231745398117302103_signature_28.png | Leo Collao | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/6231745398117302103/6231745398117302103_signature_31.png
| 2025-05-15 | 2025-05-15 | Tigaon Municipal Health Office | Caraycayon | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30798 | DOH | jingimperial24@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX4-6C | 838558 | UPS defective | Replaced UPS(AT1K240705ST0033 cleaned cartridge bays, plunger rod | Replaced UPS Performed Preventive Maintenance | Unit is Operational UPS Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6231087920938609599/6231087920938609599_signature_28.png | Maria Jesusa P Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6231087920938609599/6231087920938609599_signature_31.png
| 2025-05-15 | 2025-05-15 | Luxecare Shaw | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30797 | MACARE-HBDC | luxecarelab@loveyourself.ph | 25 |  |  | GXIV | 807112 | Unit for installation Unit for maintenance | NA | #NAME? | Unit is operational Module D2 for reporting, D3 awaiting for allocation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6231033666506094353/6231033666506094353_signature_28.png | Edmar Tungpalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6231033666506094353/6231033666506094353_signature_31.png
| 2025-05-15 | 2025-05-15 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30796 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  | 2025-05-08 | GXIV | 807956 | Unit for maintenance Monitor for replacement | Replaced monitor | - replaced monitor by AOC 22B1HS SN: CN45320WN1 - checked and cleaned cartridge bay, plunger rod, filter - generated reports | Unit is operational Modules B1, B2, B3 are awaiting allocation |  | GeneXpert Monitor AOC 22B1HS 21.5" SN: CN45320WN1 | 13005 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230992756504574247/6230992756504574247_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6230992756504574247/6230992756504574247_signature_31.png
| 2025-05-15 | 2025-05-15 | Manila Doctors Hospital | United Nations Ave., Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | xcaguas@maniladoctorscomph.onmicrosoft.com |  | 30795 | COMMERCIAL | molecular.lab@maniladoctors.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 70 ASSAY: Xpert HBV VL, COUNT: 110 ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1219 ASSAY: Xpert HCV VL, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 112 |  | GXIV-10C | 110001648 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230991653519488034/6230991653519488034_signature_28.png | Alyssa Mae Jolo | https://www.jotform.com/uploads/macarewecare/222927978372067/6230991653519488034/6230991653519488034_signature_31.png
| 2025-05-15 | 2025-05-15 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30794 | COMMERCIAL | laboratory@loveyourself.com.ph | 25 |  |  | GXIV-10C | 120002891 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230991596501258393/6230991596501258393_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6230991596501258393/6230991596501258393_signature_31.png
| 2025-05-15 | 2025-05-15 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30793 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 807112 | Unit for transfer | NA | Packed unit for transfer | Unit is ready for transfer |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230990526509970856/6230990526509970856_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6230990526509970856/6230990526509970856_signature_31.png
| 2025-05-15 | 2025-05-15 | DUMARAO RURAL HEALTH UNIT | Dumarao, Capiz | 6 | Capiz | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30792 | UNOPS | doregoshielamarie@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 210 | 2025-04-29 | GX-IV | 110023542 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | Configured Desktop for GeneXpert Machine | 12975 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230972936025723877/6230972936025723877_signature_28.png | Shiela Marie Dorego | https://www.jotform.com/uploads/macarewecare/222927978372067/6230972936025723877/6230972936025723877_signature_31.png
| 2025-05-15 | 2025-05-15 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 30791 | COMMERCIAL | bloodbank@drmc.gov.ph | 25.8 |  |  | MCS+ | 21M389SPG | Due for Preventive Maintenance | Clean Surface Area Check parameters of SPM, DPM and Line Sensor Check Roller assay Run functionality test | Perform Preventive Maintenance | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230953137328699809/6230953137328699809_signature_28.png | Christina Porquiado,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6230953137328699809/6230953137328699809_signature_31.png
| 2025-05-15 | 2025-05-15 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | REPAIR | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30790 | DOH | cauayanmholab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 690 |  | GX4 6C | 838477 | Defective LCD monitor; defective CPU power supply, RAM. | . | #NAME? | System operational, 4 modules functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230837962235064573/6230837962235064573_signature_28.png | Doriza Joy Tizon | https://www.jotform.com/uploads/macarewecare/222927978372067/6230837962235064573/6230837962235064573_signature_31.png
| 2025-05-14 | 2025-04-29 | UP-ICCHD | Central Laboratories, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30789 | MACARE-HBDC | dcsantos9@up.edu.ph | 25 |  |  | GXIV | 808015 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230448391717889585/6230448391717889585_signature_28.png | Dana Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6230448391717889585/6230448391717889585_signature_31.png
| 2025-05-14 | 2025-05-14 | Philippine Red Cross | Intramuros, Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30788 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  | 2025-05-02 | MCS+ | 12D063 | Unit for preventive maintenance. For replacement of defective transfer pump motor assembly. | Preventive maintenance and repair. | Performed Preventive Maintenance. Performed installation of new transfer pump motor assembly Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  | New Motor Assembly (0221). Def. Motor Assembly (118) | DR#12982 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230447621718306266/6230447621718306266_signature_28.png | Lalaine ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6230447621718306266/6230447621718306266_signature_31.png
| 2025-05-14 | 2025-05-14 | Bacolod City Health Office | Bacolod City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30787 | DOH | recelenylrej@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6822 |  | GX4 6C | 837612 | Stuck cartridge in module B2. | Preventive Maintenance | Removed stuck cartridge in module B2 | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230413322239679725/6230413322239679725_signature_28.png | Recele Nylrej Geronga | https://www.jotform.com/uploads/macarewecare/222927978372067/6230413322239679725/6230413322239679725_signature_31.png
| 2025-05-14 | 2025-05-14 | Sta Cruz Rural Health Unit | Poblacion, Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30786 | DOH | chendalagan72@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 463 |  | GX-IV 6C 4D | 838505 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230040393513261215/6230040393513261215_signature_28.png | Dorothy A. Dalagan,. RMT. | https://www.jotform.com/uploads/macarewecare/222927978372067/6230040393513261215/6230040393513261215_signature_31.png
| 2025-05-14 | 2025-05-14 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30785 | DOH | Buraeunlab@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 494 | 2025-05-14 | GXIV | 838551 | n/a | #NAME? | #NAME? | Unit is Operational - Module B2 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6230010246786597921/6230010246786597921_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6230010246786597921/6230010246786597921_signature_31.png
| 2025-05-14 | 2025-05-08 | Mangatarem Rural Health Unit | Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30784 | DOH | slasheiji1988new@gmail.com | 28 |  |  | GXIV-4-D-6C | 838472 | To Replace BIOS battery | #NAME? | Replaced BIOS Battery | Unit is Partially Operational Module B1 For Replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229992041259671175/6229992041259671175_signature_28.png | Ken Lee Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/6229992041259671175/6229992041259671175_signature_31.png
| 2025-05-14 | 2025-05-06 | Sarangani Health Care Facility | Alabel, Sarangani | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30783 | DOH | ronayadao@yahoo.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 920 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GxIV-6C | 838542 | None | Perform Genexpert machine Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229969646863207786/6229969646863207786_signature_28.png | Ronora Yadao | https://www.jotform.com/uploads/macarewecare/222927978372067/6229969646863207786/6229969646863207786_signature_31.png
| 2025-05-14 | 2025-05-07 | Malapatan Rural Health Unit | Poblacion, Malapatan, Sarangani | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30782 | DOH | kamilahnichol@yahoo.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 622 |  | GxIV-6C | 838419 | None | Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229969636861547578/6229969636861547578_signature_28.png | Kamilah Nichol S. Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/6229969636861547578/6229969636861547578_signature_31.png
| 2025-05-13 | 2025-05-08 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | gsd-biomed@nkti.gov.ph |  | 30781 | MACARE-HBDC | molecularlabnkti@gmail.com | 27 |  |  | GX16D | 110005587 | Due for maintenance Disabled A1,A2,C2,D2 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | A1,A2,C2,D2 reported already for replacement A3,B3 failed xpertcheck last 9/30/2024. reported already for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229364249667788718/6229364249667788718_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6229364249667788718/6229364249667788718_signature_31.png
| 2025-05-13 | 2025-05-08 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | gsd-biomed@nkti.gov.ph |  | 30780 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 104 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 36 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 125 ASSAY: Xpert HCV VL, COUNT: 10 |  | GX4 D | 838524 | Due for maintenance B4 disabled due to temp error 1002 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenancd | B4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229363819665865572/6229363819665865572_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6229363819665865572/6229363819665865572_signature_31.png
| 2025-05-13 | 2025-05-08 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30779 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  |  | MCS + | 11H204 | As per end user they encounter fluid detect in centrifuge error Due for maintenance | Performed trouble shooting Cleaned centrifuge and test fluid sensor Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance and troubleshooting | Unit ready to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229363429666285072/6229363429666285072_signature_28.png | Josephine Montilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6229363429666285072/6229363429666285072_signature_31.png
| 2025-05-13 | 2025-05-08 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30778 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  |  | MCS + | 20168 | As per end user they encounter fluid detect in centrifuge error Due for maintenance Line sensor unable to adjust potentiometer to desire value of 3071. actual line sensor value 2921 | Performed trouble shooting Cleaned centrifuge and test fluid sensor Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed troubleshooting and maintenance | Unit ready to use During simulation run no error encounter |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229362999663115986/6229362999663115986_signature_28.png | Josephine Montilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6229362999663115986/6229362999663115986_signature_31.png
| 2025-05-13 | 2025-05-09 | Lung Center of the Philippines | Quezon Avenu QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30777 | MACARE-HBDC | nfloresca@macare-medicals.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 156 ASSAY: Xpert SARS-CoV-2, COUNT: 116 ASSAY: Xpert HIV-1 Viral Load, COUNT: 34 ASSAY: Xpert SARS-CoV-2, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 16 |  | GX4 D | 838837 | Due for maintenance B1,B4 reported already for replacement B3 for reporting encounter temp error1002,4009 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B2 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229362659663191478/6229362659663191478_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6229362659663191478/6229362659663191478_signature_31.png
| 2025-05-13 | 2025-05-13 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30776 | DOH | cauayanmholab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 687 |  | GX4-6C | 838477 | CPU power on problem; Monitor problem | . | #NAME? | - 4 modules operational - for isolation of CPU problem: possible issue with LCD consequently causing problem during power on of CPU. - defective LCD monitor- for replacement | Waiting for replacement LCD monitor |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229322292237846484/6229322292237846484_signature_28.png | Doriza Joy Tizon | https://www.jotform.com/uploads/macarewecare/222927978372067/6229322292237846484/6229322292237846484_signature_31.png
| 2025-05-13 | 2025-05-05 | Don Jose S. Monfort Medical Center | Barotac Nuevo, Iloilo | 6 | Iloilo | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30775 | UNOPS | judyboglosa35@gmail.com | 25 |  |  | Gx4 | 838501 | Defective module B2 SN 687320 | New module Installed B2 210003521 Clean Fan Filter Gather data | New module Installed B2 210003521 Clean Fan Filter Gather data | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229158614133114399/6229158614133114399_signature_28.png | Judy boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6229158614133114399/6229158614133114399_signature_31.png
| 2025-05-13 | 2025-05-05 | Don Jose S. Monfort Medical Center | Tabocan, | 6 | Iloilo | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30774 | UNOPS | judyboglosa35@gmail.com | 25 |  |  | Gx4 | 838501 | Defective module B2 SN 687320 | New module Installed B2 210003521 Clean Fan Filter Gather data | New module Installed B2 210003521 Clean Fan Filter Gather data | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229158614133114399/6229158614133114399_signature_28.png | Judy boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6229158614133114399/6229158614133114399_signature_31.png
| 2025-05-13 | 2025-05-09 | Cardinal Santos Medical Center | Wilson St. San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 30772 | COMMERCIAL | microbiology.csmc@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 123 ASSAY: Xpert HPV, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1104 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B3 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229069293617559259/6229069293617559259_signature_28.png | sheilahmae carandang | https://www.jotform.com/uploads/macarewecare/222927978372067/6229069293617559259/6229069293617559259_signature_31.png
| 2025-05-13 | 2025-05-09 | Cardinal Santos Medical Center | Wilson St San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 30773 | COMMERCIAL | csmclab@csmc.ph | 22 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6229069293618702855/6229069293618702855_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/6229069293618702855/6229069293618702855_signature_31.png
| 2025-05-11 | 2025-05-10 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | CALIBRATION | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30771 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 02F035 | Due for Calibration | Performed calibration in apheresis machine | Performed calibration in apheresis machine | Unit is operational and passed during the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6227287613444458037/6227287613444458037_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6227287613444458037/6227287613444458037_signature_31.png
| 2025-05-11 | 2025-05-10 | Chong Hua Hospital- Fuente MAB | FUNTE OSMENA, CEBU | 7 | Cebu | Molecular Laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | chh_microbiology@chonghua.com.ph |  | 30770 | COMMERCIAL | chhc_laboratory_@chonghua.com.ph | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 72 ASSAY: Xpert HBV VL, COUNT: 752 |  | GX4 | 110020777 | Due for Calibration, Preventive Maintenance, | Conduct Preventive Maintenance, and Performed calibration in all modules Clean Icore, Plunger and Filter Generate IQR, System log, Gxx Files and Assay Statistics | Conduct Preventive Maintenance, and Performed calibration in all modules Clean Icore, Plunger and Filter Generate IQR, System log, Gxx Files and Assay Statistics | All modules passed and genexpert machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6227283593444042802/6227283593444042802_signature_28.png | Chong Hua Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/6227283593444042802/6227283593444042802_signature_31.png
| 2025-05-09 | 2025-05-09 | CEBU DOCTOR'S UNIVERSITY HOSPITAL | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30769 | COMMERCIAL | cduhcovidpoct@gmail.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 88 ASSAY: Xpert CT/NG, COUNT: 264 ASSAY: Xpert HBV VL, COUNT: 189 ASSAY: Xpert HPV, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1109 ASSAY: Xpert HCV VL, COUNT: 145 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GX-IV | 110008524 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | MODULE 3 & MODULE 4 DISABLED |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225886558164922808/6225886558164922808_signature_28.png | Colleen Allison Libi | https://www.jotform.com/uploads/macarewecare/222927978372067/6225886558164922808/6225886558164922808_signature_31.png
| 2025-05-09 | 2025-05-09 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | daguilon.intp@gmail.com |  | 30768 | MACARE-HBDC | msff-tondo-labtech@paris.msf.org | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4503 | 2025-05-07 | GX16-10C | 110016343 | Unit for preventive maintenance. For module replacement of modules A2, A3, B1 and C3. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module: Defective Modules: A2 (210075770), A3 (210075498), B1 (210082298) and C3 (210075410) New Modules: A2 (210024238), A3 (210067146), B1 (210076777) and C3 (210078338) | DR#13001 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225849093113122815/6225849093113122815_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6225849093113122815/6225849093113122815_signature_31.png
| 2025-05-09 | 2025-05-08 | Batangas City Health Office | Batangas City Hall, Batangas City | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30767 | MACARE-HBDC | gjanlorenzo@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1286 |  | GXIV-10C | 110024615 | Stucked cartridge on module B1. Unit for preventive maintenance. | Repair and preventive maintenance. | Performed removal of stucked cartridge on module B1. Performed several selftest with cartridge on B1 after removing the stucked cartridge. Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225839173119494861/6225839173119494861_signature_28.png | Jan Lorenzo Godoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6225839173119494861/6225839173119494861_signature_31.png
| 2025-05-09 | 2025-05-09 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30766 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1893 |  | 6C GXIV-4D | 838561 | Due for monthly PM | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225776025673984058/6225776025673984058_signature_28.png | Hannah Jane Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/6225776025673984058/6225776025673984058_signature_31.png
| 2025-05-09 | 2025-05-09 | Mandaluyong Social Hygiene Clinic | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30765 | MACARE-HBDC | mandaluyongshc.laboratory@gmail.com | 25 |  |  | GXIV-10C | 110025111 | Unit for maintenance Printer for installation | Installed printer driver | - checked and cleaned cartridge bay, plunger rod, filter, icore slot - installed printer (Brother DCP-T720DW) - generated reports - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225729150319122122/6225729150319122122_signature_28.png | Revena Sibayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6225729150319122122/6225729150319122122_signature_31.png
| 2025-05-09 | 2025-05-09 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30764 | COMMERCIAL | dionellynretiza99@gmail.com | 23 |  |  | MCS+ | 02D167 | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225700242567482719/6225700242567482719_signature_28.png | Dionellyn Marie M. Retiza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6225700242567482719/6225700242567482719_signature_31.png
| 2025-05-09 | 2025-05-09 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30763 | COMMERCIAL | dionellynretiza99@gmail.com | 23 |  |  | MCS+ | 11H307 | #NAME? | #NAME? | #NAME? | -MCS+ EQUIPMENT IS OPERATIONAL -CONFIRMED UPS S/N:40460571107 IS DEFECTIVE AND FOR OUT OF WARRANTY REPLACEMENT |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225697227769657668/6225697227769657668_signature_28.png | Dionellyn Marie M. Retiza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6225697227769657668/6225697227769657668_signature_31.png
| 2025-05-09 | 2025-05-08 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30760 | COMMERCIAL | bloodbank@themedicalcity.com | 22 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational | Recommended to calibrate |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225663641413245348/6225663641413245348_signature_28.png | Jan danielle delos reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6225663641413245348/6225663641413245348_signature_31.png
| 2025-05-09 | 2025-05-08 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30762 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 61 ASSAY: Xpert CT/NG, COUNT: 53 ASSAY: Xpert HBV VL, COUNT: 644 ASSAY: Xpert HPV, COUNT: 397 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6378 ASSAY: Xpert MTB-XDR, COUNT: 3 ASSAY: Others, COUNT: 5 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 152 ASSAY: Xpert HIV-1 Viral Load, COUNT: 876 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1442 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225663641414335659/6225663641414335659_signature_28.png | Janica Katheryn P. Zamora | https://www.jotform.com/uploads/macarewecare/222927978372067/6225663641414335659/6225663641414335659_signature_31.png
| 2025-05-09 | 2025-05-08 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30761 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 22 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225663641417843983/6225663641417843983_signature_28.png | Febbie Lou Birang | https://www.jotform.com/uploads/macarewecare/222927978372067/6225663641417843983/6225663641417843983_signature_31.png
| 2025-05-08 | 2025-05-08 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 30759 | MACARE-HBDC | laboratory@slec.ph | 22 |  |  | Infinity 48s | 809643 | Automation failure due to door latching error | Verified functionality in INFTeach Adjusted module door tension Cleared automation failures | Performed functionality test | Unit is operational, affected modules for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6225023440228798929/6225023440228798929_signature_28.png | Ma. Sarah R.Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6225023440228798929/6225023440228798929_signature_31.png
| 2025-05-08 | 2025-05-08 | Makilala Rural Health Unit | Makilala Poblacion | 11 | Cotabato | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30758 | UNOPS | abegailweill17@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 386 |  | GX-IV 10C 4D | 110023521 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6224905722218171623/6224905722218171623_signature_28.png | Abegail Publico | https://www.jotform.com/uploads/macarewecare/222927978372067/6224905722218171623/6224905722218171623_signature_31.png
| 2025-05-08 | 2025-05-08 | Quezon City Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30757 | MACARE-HBDC | molecularlab.cityhealth@quezoncity.gov.ph | 25 |  |  | GXIV-10C | 110020696 | Unit for installation | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6224904244318693302/6224904244318693302_signature_28.png | Cyrell Arcuelo | https://www.jotform.com/uploads/macarewecare/222927978372067/6224904244318693302/6224904244318693302_signature_31.png
| 2025-05-08 | 2025-05-06 | Biopath QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30756 | COMMERCIAL | ainocentes@biopath.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 33 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 47 ASSAY: Others, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 10 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX16 D | 839306 | Due for maintenance Disabled A1,A3,A4,B2,B4,C2,D2 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | 9 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6224699913121808316/6224699913121808316_signature_28.png | Amanda inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6224699913121808316/6224699913121808316_signature_31.png
| 2025-05-07 | 2025-05-07 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 30755 | DOH | johnpaullabutap@gmail.com | 28 |  | 2025-05-07 | GXIV | 838548 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6224012638984498556/6224012638984498556_signature_28.png | Mary Grace Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/6224012638984498556/6224012638984498556_signature_31.png
| 2025-05-07 | 2025-05-07 | TB DOTS Tagoloan | Poblacion, Tagoloan | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30754 | DOH | macalisangglezeal@gmail.com | 25 |  | 2025-04-21 | GX4-6 Colors | 838574 | Module B2 found temperature errors | Module Replacement | Replacing Module B2 with SN: 862673 to SN: 894548 | All modules and units are functional |  | Certified pre-owned GeneXpert Module | 12949 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6223941143756043006/6223941143756043006_signature_28.png | Glezeal J. Macalisang | https://www.jotform.com/uploads/macarewecare/222927978372067/6223941143756043006/6223941143756043006_signature_31.png
| 2025-05-07 | 2025-04-29 | Ospital ng Makati | Sampaguita St., Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 30753 | COMMERCIAL | ctmlee.0304@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 440 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6223926003613498678/6223926003613498678_signature_28.png | Cristina Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6223926003613498678/6223926003613498678_signature_31.png
| 2025-05-06 | 2025-04-29 | Naguilian District Hospital | Naguilian, La Union | 1 | La Union | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30752 | DOH | marjelynrayray01@gmail.com | 25 |  | 2025-04-21 | GXIV-4-D-6C | 838246 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B1 861154 >>>806518 - Generated Reports | Performed Module Replacement | Unit is Fully Operational | Work with Engineer Venzon | GeneXpert Module 806518 | DR 12946 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6223264813968992426/6223264813968992426_signature_28.png | Marjelyn Rayray | https://www.jotform.com/uploads/macarewecare/222927978372067/6223264813968992426/6223264813968992426_signature_31.png
| 2025-05-06 | 2025-04-29 | Mangatarem Rural Health Unit | Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30751 | DOH | slasheiji1988new@gmail.com | 28 |  |  | GXIV-4-D-6C | 838284 | For Troubleshooting and Maintenance | #NAME? | Performed Troubleshooting and Maintenance | Unit is Operational but for BIOS battery Replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6223263653964885435/6223263653964885435_signature_28.png | Ken Lee Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/6223263653964885435/6223263653964885435_signature_31.png
| 2025-05-06 | 2025-05-06 | Diit District Health Center | Diit, Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30750 | DOH | dalepenalosa189@gmail.com | 26 |  | 2025-05-06 | GXIV | 838421 | N/a | N/a | Monthly Preventive Maintenance | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6223231259143192500/6223231259143192500_signature_28.png | Dale Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6223231259143192500/6223231259143192500_signature_31.png
| 2025-05-06 | 2025-05-06 | National Center for Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30749 | COMMERCIAL | iancristrono@gmail.com | 25 |  |  | GXIV-10C | 120003480 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6223147525315178234/6223147525315178234_signature_28.png | Ian Cris Tono | https://www.jotform.com/uploads/macarewecare/222927978372067/6223147525315178234/6223147525315178234_signature_31.png
| 2025-05-06 | 2025-04-29 | Columbio Rural Health Unit | Columbio, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30747 | DOH | rochelle.toledo19@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 709 |  | GxIV-6C | 838556 | Module 2 under observation due to temperature errors Module 3 multiple invalid results No Fan filter | Perform XPC Calibration to module 2&3 for troubleshooting Perform Genexpert Cleaning Module Cleaning Plunger rod and Optics Cleaning Fan filter Cleaning Checking of errors Errors/problem reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert module 2&3 troubleshooting | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6222951846861014086/6222951846861014086_signature_28.png | Rochelle S. Toledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6222951846861014086/6222951846861014086_signature_31.png
| 2025-05-06 | 2025-05-02 | Polomolok Rural Health Unit | Poblacion, Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30748 | DOH | jmmirasolrmt@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2170 | 2025-05-02 | GxIV-6C | 838265 | Module 2, H/W FAILED | Replace/repair module 2 SN:685448 with new module SN:613279 Perform Genexpert machine Cleaning Module cleaning Plunger rod and Optics cleaning Fan filter cleaning Checking of errors Data collection | Genexpert machine Preventive maintenance Genexpert module repair | 4 functional modules |  | Genexpert 6 color module SN:685448 | DR#: 12945 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6222951856864782110/6222951856864782110_signature_28.png | June Michele J. Mirasol | https://www.jotform.com/uploads/macarewecare/222927978372067/6222951856864782110/6222951856864782110_signature_31.png
| 2025-05-05 | 2025-04-25 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30746 | DOH | allan.vinluan2013@gmail.com | 30 |  |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6222330783964782013/6222330783964782013_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6222330783964782013/6222330783964782013_signature_31.png
| 2025-05-05 | 2025-04-25 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30745 | DOH | seansimonpaloyo@gmail.com | 27 |  |  | GXIV-4-D-6C | 838273 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6222328263966474293/6222328263966474293_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6222328263966474293/6222328263966474293_signature_31.png
| 2025-05-05 | 2025-04-24 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30744 | COMMERCIAL | annabellelcatbagan@gmail.com | 25 |  |  | GXIV-4-D-6C | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6222325533965002893/6222325533965002893_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6222325533965002893/6222325533965002893_signature_31.png
| 2025-05-05 | 2025-04-24 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30743 | DOH | serly_reyes@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 771 |  | GXIV-4-D-6C | 838516 | CPU for Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Fully Operational CPU for Observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6222324083965317917/6222324083965317917_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6222324083965317917/6222324083965317917_signature_31.png
| 2025-05-05 | 2025-04-24 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30742 | DOH | dianeruiz147@gmail.com | 30 |  |  | GXIV-4-D-6C | 838438 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Operational UPS for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6222311453962313122/6222311453962313122_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6222311453962313122/6222311453962313122_signature_31.png
| 2025-05-05 | 2025-05-02 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 30741 | MACARE-HBDC | florescanoel77@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 49 ASSAY: Xpert BCR-ABL Ultra, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1373 ASSAY: Xpert SARS-CoV-2, COUNT: 20 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1558 |  | GX16D | 110014171 | For maintenance and data gathering | Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | C2,C3,C4,D1,D2,D3,D4 enable A1,A2,A3,A4,B1,B2,B3,B4,C1 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6222117430718994462/6222117430718994462_signature_28.png | National Children's Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6222117430718994462/6222117430718994462_signature_31.png
| 2025-05-02 | 2025-05-02 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30740 | DOH | luzvimindanavidad2@gmail.com | 26 |  |  | 6C GXIV-4D | 838297 | Module 3 and 4 for XPC | -Run XpertCheck on module 3 and 4 | #NAME? | Module 3 for retry but skipped, module 4 for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6219857591539710751/6219857591539710751_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6219857591539710751/6219857591539710751_signature_31.png
| 2025-05-02 | 2025-05-02 | Jones Rural Health Unit | Brgy. 1, Jones, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30739 | DOH | oroprecy19@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 575 | 2025-05-21 | GXIV-6C | 838484 | Module 1 and 4 Hardware Failure | Performed Module Replacement B1(862050 - 638316) B4(862085 - 613605) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Modules | 12941 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6219761922529246062/6219761922529246062_signature_28.png | Preciosa Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/6219761922529246062/6219761922529246062_signature_31.png
| 2025-05-02 | 2025-05-02 | Jones Rural Health Unit | Brgy. 1, Jones, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30738 | DOH | oroprecy19@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 575 | 2025-05-21 | GXIV-6C | 838484 | Module 1 and 4 Hardware Failure | Performed Module Replacement B1(862050 - 638316) B4(862085 - 613605) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Modules | 12941 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6219761902523591277/6219761902523591277_signature_28.png | Preciosa Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/6219761902523591277/6219761902523591277_signature_31.png
| 2025-04-30 | 2025-04-30 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30737 | DOH | sajecopooh20@gmail.com | 27 |  |  | 6C GXIV-4D | 838555 | Due for PM | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6218181490195096065/6218181490195096065_signature_28.png | Sarah Jean G. Pasana | https://www.jotform.com/uploads/macarewecare/222927978372067/6218181490195096065/6218181490195096065_signature_31.png
| 2025-04-30 | 2025-04-30 | Nasipit Rural Health Unit | Brgy. Talisay, Nasipit | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30736 | DOH | peterdolotina@gmail.com | 23 |  | 2025-04-23 | GX4-6 Colors | 837611 | Module B1 found temperature error | Module Replacement | Replacing Module B1 with SN: 806517 to SN: 633808 Preventive Maintenance | All modules and units are functional. |  | Certified pre-owned Genexpert Module | 12943 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6218062938814975017/6218062938814975017_signature_28.png | Peter T. Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/6218062938814975017/6218062938814975017_signature_31.png
| 2025-04-30 | 2025-04-30 | Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30735 | DOH | miniforestlaboratory@gmail.com | 23.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 485 | 2025-04-21 | GX-IV 6C 4D | 838507 | Module 3 encounter multiple temperature related errors For Preventive Maintenance Module 1 encounter multiple errors | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replaced the Defective Module SN:861941 with new Module SN:794496 Run functionality test on Module 1 | Perform Preventive Maintenance, Module Replacement, and troubleshooting | Module 3 Successfully replaced, 4 modules are operational | Module 1 for under observation, if error persist, enduser was advised to report the said module | 900-0491rs | DR 12948 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6218053751213820436/6218053751213820436_signature_28.png | Blez Pepito | https://www.jotform.com/uploads/macarewecare/222927978372067/6218053751213820436/6218053751213820436_signature_31.png
| 2025-04-30 | 2025-04-30 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | CALIBRATION | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30734 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  | 2025-04-30 | MCS+ | 01J038 | Unit for calibration | #NAME? | Performed calibration | Unit is operational Calibration certificate to follow |  | Calibration Tools | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6218045398387553267/6218045398387553267_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6218045398387553267/6218045398387553267_signature_31.png
| 2025-04-30 | 2025-04-30 | Tigaon Municipal Health Office | Caraycayon, Tigaon | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30733 | DOH | j8ngimperial24@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1081 |  | GX4-6C | 838558 | B4/614067 - not detected, for module replacement UPS no power | Replaced module: B4 - 614067>>687278 UPS - tested/replaced fuse(working), checked power cable connections from UPS going to power source Cleaned cartridge bays, plunger rod, gx fan, icore Generate Data | Replaced module B4/614067 Performed Troubleshooting for UPS Performed Preventive Maintenance | Unit is operational UPS non operational | UPS request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6218044818182675151/6218044818182675151_signature_28.png | Maria Jesusa P. Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6218044818182675151/6218044818182675151_signature_31.png
| 2025-04-30 | 2025-04-30 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | CALIBRATION | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30732 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  | 2025-04-30 | MCS+ | 08B049 | Unit for calibration | #NAME? | Performed calibratio | Unit is operational Calibration certificate to follow |  | Calibration Tools | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6218044818383072574/6218044818383072574_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6218044818383072574/6218044818383072574_signature_31.png
| 2025-04-29 | 2025-04-29 | Estevez Memorial Hospital, Inc. | Estevez Street, Guevarra Subdivision, Legazpi City | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | reylovendino83@gmail.com |  | 30731 | COMMERCIAL | melosept22@gmail.com | 28 |  |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6217229662513831451/6217229662513831451_signature_28.png | Estevez Memorial Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6217229662513831451/6217229662513831451_signature_31.png
| 2025-04-29 | 2025-04-29 | Davao Reproductive Health and Wellness Center | Jacinto Street, Davao City | 11 | Davao del Sur | TB-HIV | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30730 | DOH | rhwclab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 405 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1628 | 2025-04-21 | GX-IV 6C 4D | 838262 | Module 1 reported due to multiple temperature related errors | Replaced the Defective Module SN: 656376 With New Module SN:623864 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Module Replacement | All Modules are operational |  | 900-0491rs | DR 12942 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6217190301113488262/6217190301113488262_signature_28.png | Ellaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6217190301113488262/6217190301113488262_signature_31.png
| 2025-04-29 | 2025-04-29 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30729 | DOH | luzvimindanavidad2@gmail.com | 26 |  |  | 6C GXIV-4D | 838297 | Due for PM | #NAME? | #NAME? | Machine is functional. To run XPC on module 3 & 4. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6217178104225529241/6217178104225529241_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6217178104225529241/6217178104225529241_signature_31.png
| 2025-04-29 | 2025-04-29 | UP-ICCHD | Central Laboratories, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30728 | MACARE-HBDC | dcsantos9@up.edu.ph | 25 |  |  | GXIV | 808015 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6217135778113152293/6217135778113152293_signature_28.png | Mary Grace Hernaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6217135778113152293/6217135778113152293_signature_31.png
| 2025-04-29 | 2025-04-29 | National Institutes of Health | Central Laboratories, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30727 | MACARE-HBDC | nihcentrallab@gmail.com | 25 | ASSAY: Xpert Carba-R, COUNT: 27 ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Xpert HBV VL, COUNT: 62 ASSAY: Xpert HCV VL, COUNT: 44 ASSAY: Xpert HIV-1 Viral Load, COUNT: 429 |  | GX16-10C | 110014186 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6217135358114643790/6217135358114643790_signature_28.png | Mary Grace Hernaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6217135358114643790/6217135358114643790_signature_31.png
| 2025-04-29 | 2025-04-29 | Philippine General Hospital | Taft Ave. Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30726 | MACARE-HBDC | pghmolecularlab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 19 ASSAY: Xpert SARS-CoV-2, COUNT: 281 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 835 |  | GX16 | 844975 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6217108699224023987/6217108699224023987_signature_28.png | Kaye bautro | https://www.jotform.com/uploads/macarewecare/222927978372067/6217108699224023987/6217108699224023987_signature_31.png
| 2025-04-29 | 2025-04-29 | Philippine General Hospital | Taft Ave. Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30725 | MACARE-HBDC | pghmolecularlab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 19 ASSAY: Xpert SARS-CoV-2, COUNT: 281 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 835 |  | GX16 | 844975 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6217107649221675444/6217107649221675444_signature_28.png | Kaye bautro | https://www.jotform.com/uploads/macarewecare/222927978372067/6217107649221675444/6217107649221675444_signature_31.png
| 2025-04-29 | 2025-04-25 | Cebu South Medical Center | San Isidro, Talisay City Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30724 | DOH | csmcgenexpert@gmail.com | 24 |  |  | GX4 | 838559 | #NAME? | generate data and clean up module B2 - Generate IQR, System Log, and Gxx files | #NAME? | Unit of module 2 required for calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6217039018652587728/6217039018652587728_signature_28.png | Kristel Sheen Miranda, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6217039018652587728/6217039018652587728_signature_31.png
| 2025-04-29 | 2025-04-25 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30723 | DOH | joydarlie72@gmail.com | 29 |  |  | Gx IV | 838527 | N/A | Monthly Preventive Maintenance | Monthly Preventive Maintenance | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6217034055105981788/6217034055105981788_signature_28.png | Joy Ma. Darlie Labrado | https://www.jotform.com/uploads/macarewecare/222927978372067/6217034055105981788/6217034055105981788_signature_31.png
| 2025-04-29 | 2025-04-28 | Cebu South Medical Center | San Isidro, Talisay City Cebu | 7 | Cebu | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30722 | DOH | csmcgenexpert@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1070 |  | GX4 | 838559 | Module B2 SN 862869 required calibration due to repeated error of 5007 in MTB-ultra and error 2097 in HIV-viral load assay as per report by the end user. | - Performed troubleshooting and calibration on Module B2 862869 failed results | Performed calibration of the Module B2 SN # 862869 failed results. | The genexpert machine is operational and module B2 failed awating for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216956368658157986/6216956368658157986_signature_28.png | Kristel Sheen Miranda, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6216956368658157986/6216956368658157986_signature_31.png
| 2025-04-29 | 2025-04-24 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30721 | COMMERCIAL | vsmmcbloodservices@gmail.com | 24 |  |  | MCS+ | 02A090 | #NAME? | #NAME? | #NAME? | MCS+ unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216934440801453806/6216934440801453806_signature_28.png | Vicente Sotto Memorial Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6216934440801453806/6216934440801453806_signature_31.png
| 2025-04-28 | 2025-04-22 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30720 | DOH | jan19.cam@gmail.com | 26.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1340 |  | GxIV-6C | 838504 | None | Perform Genexpert Cleaning and maintenance Module cleaning Plunger rod and Optics cleaning Fan and Filter cleaning Checking of errors Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216566846865761424/6216566846865761424_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/6216566846865761424/6216566846865761424_signature_31.png
| 2025-04-28 | 2025-04-25 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30719 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS+ | 08H188 | Due for maintenance and calibration Need to replace urethane rollers 6pcs UPS unable to back up defective battery Error 68 appear during simulation running | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running Performed calibration | Performed maintenance | Need to replace urethane rollers 6pcs For further checking |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216511698102558836/6216511698102558836_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/6216511698102558836/6216511698102558836_signature_31.png
| 2025-04-28 | 2025-04-25 | Lung Center of the Philippines | Queson Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbnak | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30718 | MACARE-HBDC | lcpbloodbank@gmail.com | 25 |  |  | MCS + | 02D166 | Due for maintenance and calibration For replacement of urethane rollers 6pcs slight wornout | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance and calibration | For replacement of urethane rollers 6pcs |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216511188105218838/6216511188105218838_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/6216511188105218838/6216511188105218838_signature_31.png
| 2025-04-28 | 2025-04-25 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | eamc.biomed@gmail.com |  | 30717 | MACARE-HBDC | blooddonorsection@gmail.com | 26 |  |  | MCS + | 15H044 | Due for maintenance Due for calibration last calibration june 2023 Line sensor unable to adjust the desire value of 3071. | Cleaned unit Checked rollers Performed valve test Checked and test keypad and lcd display. Checked program parameters Performed functionality test | Performed maintenance | For quotation of Calibration and line sensor assy |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216510708103128519/6216510708103128519_signature_28.png | Christopher Paul Bernaldo | https://www.jotform.com/uploads/macarewecare/222927978372067/6216510708103128519/6216510708103128519_signature_31.png
| 2025-04-28 | 2025-04-28 | Jose B. Lingad Memorial and General Hospital | San Fernando City Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30716 | MACARE-HBDC | jblmgh.is@gmail.com | 30 | ASSAY: Xpert HBV VL, COUNT: 297 ASSAY: Xpert HCV VL, COUNT: 81 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1949 |  | GXIV-D 10C | 110013748 | Quarterly Preventive Maintenance UPS troubleshooting | Troubleshooting Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational UPS for replacement | UPS not backing up |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216342421613764310/6216342421613764310_signature_28.png | Jem Suba | https://www.jotform.com/uploads/macarewecare/222927978372067/6216342421613764310/6216342421613764310_signature_31.png
| 2025-04-28 | 2025-04-28 | Jose B. Lingad Memorial and General Hospital | San Fernando City Pampanga | 3 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30715 | MACARE-HBDC | jblmghmpl@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 821 ASSAY: Xpert MTB-XDR, COUNT: 155 ASSAY: Xpert SARS-CoV-2, COUNT: 0 |  | GXIV-D 10C | 110012825 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216326921611971852/6216326921611971852_signature_28.png | Jovee De Jesus | https://www.jotform.com/uploads/macarewecare/222927978372067/6216326921611971852/6216326921611971852_signature_31.png
| 2025-04-28 | 2025-04-28 | Pio Duran Rural Health Unit | Caratagan, Pio Duran | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30714 | DOH | mitzijc@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 |  | GX4-6C | 838264 | B1/807820 for module replacement | Replaced module: B1 - 807820>>688876 Cleaned cartridge bays, plunger rod, icore, Gx fan Generate Data | Replaced module B1/807820 Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216264242518481535/6216264242518481535_signature_28.png | Mitzi D. Cedeño | https://www.jotform.com/uploads/macarewecare/222927978372067/6216264242518481535/6216264242518481535_signature_31.png
| 2025-04-28 | 2025-04-25 | Asian Hospital And Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 30713 | COMMERCIAL | bloodbank@asianhospital.com | 25 |  |  | MCS+ | 10B129 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Surface Area. Cleaned Pump Rollers. Checked Sensors and Functionality. Performed Simulation Test. | Unit is Operational. | Unit without UPS. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201743614465274/6216201743614465274_signature_28.png | Angel Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201743614465274/6216201743614465274_signature_31.png
| 2025-04-28 | 2025-04-25 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 30712 | COMMERCIAL | bloodbank@asianhospital.com | 25 |  |  | MCS+ | 01F16 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Surface Area. Cleaned Pump Rollers. Checked Sensors and Functionality. Performed Simulation Test. | Unit is Operational. | Unit without UPS. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663611186906/6216201663611186906_signature_28.png | Angel Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663611186906/6216201663611186906_signature_31.png
| 2025-04-28 | 2025-04-23 | General Trias Social Hygiene Clinic | Hospital Road, General Trias, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30709 | MACARE-HBDC | chogentri.shc@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 836 |  | GXIV-D-10C | 110016785 | Unit is due for XpertCheck. | Preventive Maintenance. XpertCheck. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Maintenance. Performed Self-test. Performed XpertCheck. Archived Test Results. Generate Reports. | All Modules Passed and Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663612853541/6216201663612853541_signature_28.png | John Joseph Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663612853541/6216201663612853541_signature_31.png
| 2025-04-28 | 2025-04-15 | San Lazaro Hospital | Quiricada St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | SACCL | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30708 | MACARE-HBDC | blessieestrelles@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 76 ASSAY: Xpert HCV VL, COUNT: 204 |  | GXIV-D-6C | 839147 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Archived Test Results. Generate Reports. | GX System is in Good Condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663615099672/6216201663615099672_signature_28.png | Blessie Estrelles | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663615099672/6216201663615099672_signature_31.png
| 2025-04-28 | 2025-04-24 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Road, Las Piñas City | NCR | NCR 4th - Las Piñas | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | uphdmc_lab@yahoo.com |  | 30711 | COMMERCIAL | phmclp.mdl@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Xpert HBV VL, COUNT: 13 ASSAY: Xpert HPV, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 119 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GXIV-D-6C | 831316 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | 4 Modules were Service Required in XpertCheck. | For Replacement of Modules. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663615130241/6216201663615130241_signature_28.png | Donn Dave Clerigo | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663615130241/6216201663615130241_signature_31.png
| 2025-04-28 | 2025-04-25 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | rpescador@asianhospital.com |  | 30710 | COMMERCIAL | agalang@asianhospital.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 3 ASSAY: Xpert CT/NG, COUNT: 31 ASSAY: Others, COUNT: 82 ASSAY: Xpert HBV VL, COUNT: 173 ASSAY: Xpert HPV, COUNT: 201 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1136 ASSAY: Xpert HIV-1 Viral Load, COUNT: 118 |  | GXIV-D-10C | 110014996 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663618708026/6216201663618708026_signature_28.png | Alyssa Galang | https://www.jotform.com/uploads/macarewecare/222927978372067/6216201663618708026/6216201663618708026_signature_31.png
| 2025-04-28 | 2025-04-24 | Metro Antipolo Hospital and Medical Center | Marcos Highway Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30707 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 18 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 283 ASSAY: Xpert HIV-1 Viral Load, COUNT: 15 |  | GX-IV-D-6C | 110005265 | Module Be and B4 for re-xpertcheck | Run XpertCheck on module B3 and B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck on module B3 and B4 Generated reports | Unit is operational. Module B3 and B4 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216132773618595915/6216132773618595915_signature_28.png | Fatima Busar | https://www.jotform.com/uploads/macarewecare/222927978372067/6216132773618595915/6216132773618595915_signature_31.png
| 2025-04-28 | 2025-04-24 | Antipolo City Health Office | M. Santos St. Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30706 | UNOPS | acho.laboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert MTB-XDR, COUNT: 40 ASSAY: Xpert HIV-1 Viral Load, COUNT: 496 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6216132773619799654/6216132773619799654_signature_28.png | Elaine Ouano | https://www.jotform.com/uploads/macarewecare/222927978372067/6216132773619799654/6216132773619799654_signature_31.png
| 2025-04-26 | 2025-04-25 | Santa Rosa City Health Office 1 | Santa Rosa Laguna | 4A | Laguna | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30704 | MACARE-HBDC | csrlcho1covidlab2@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 27 ASSAY: Xpert SARS-CoV-2, COUNT: 6 |  | GX16 | 844981 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6214462257322817347/6214462257322817347_signature_28.png | Kathleen Claire L. Yambao | https://www.jotform.com/uploads/macarewecare/222927978372067/6214462257322817347/6214462257322817347_signature_31.png
| 2025-04-26 | 2025-04-25 | Santa Rosa City Health Office 1 | Santa Rosa Laguna | 4A | Laguna | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30705 | MACARE-HBDC | csrlcho1covidlab2@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 7 ASSAY: Xpert SARS-CoV-2, COUNT: 4 |  | GXIV | 838522 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6214462257328600507/6214462257328600507_signature_28.png | Kathleen Claire L. Yambao | https://www.jotform.com/uploads/macarewecare/222927978372067/6214462257328600507/6214462257328600507_signature_31.png
| 2025-04-25 | 2025-04-25 | Pilar Rural Health Unit | Calongay, Pilar | 5 | Sorsogon | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30703 | DOH | manalangmarkramir@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1759 |  | GX4-6C | 838266 | B4/615874 temperature error - request for module replacement | Replaced module: B4 - 615874>>635384 Cleaned cartridge bays, plunger rod, icore Generate Data | Replaced module B4/615874 Performed Preventive Maintenance | Unit is Operational Monitor - grid lines started to appear below the screen - request for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6213789812378146160/6213789812378146160_signature_28.png | Mark Ramir C. Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/6213789812378146160/6213789812378146160_signature_31.png
| 2025-04-25 | 2025-04-25 | Zamboanga del sur medical center | Zamboanga del sur medical center | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30702 | DOH | limboemillie@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2047 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 | 2025-04-25 | GX 4 | 838204 | UPS unable to supply the machine properly | Preventive Maintenence, replaced 2 pcs batteries UPS with 12v and 7AH. | Preventive Maintenence and battery replacement for UPS. | All modules are ready to use | UPS must be charged 8hrs for it to work properly. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6213686915816915558/6213686915816915558_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/6213686915816915558/6213686915816915558_signature_31.png
| 2025-04-25 | 2025-04-25 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30701 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2873 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 |  | 6C GXIV-4D | 838285 | Due for monthly PM | #NAME? | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6213684014713033538/6213684014713033538_signature_28.png | Allan Villanobos Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6213684014713033538/6213684014713033538_signature_31.png
| 2025-04-25 | 2025-04-25 | Santa Rosa City Health Office 1 | Santa Rosa Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30700 | UNOPS | santarosacho1clinlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 345 |  | GXIV-10C | 110023497 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6213641779324508877/6213641779324508877_signature_28.png | ALONA NINA ALVARO | https://www.jotform.com/uploads/macarewecare/222927978372067/6213641779324508877/6213641779324508877_signature_31.png
| 2025-04-25 | 2025-04-24 | MMC Medillinx | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30699 | COMMERCIAL | janeshyrelenriquez.medilinx@outlook.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 30 ASSAY: Xpert CT/NG, COUNT: 54 ASSAY: Xpert HBV VL, COUNT: 18 ASSAY: Xpert HPV, COUNT: 37 ASSAY: Xpert MTB/RIF Ultra, COUNT: 230 ASSAY: Others, COUNT: 4 ASSAY: Xpert HCV VL, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 127 |  | GX16D 10C | 110026448 | C1 encounter error 4014,2034 optical signal, after loading of cartridge in a few second module c1 eject the cartridge error 4014 appear. | Cleaned icore Reset Gateway board Performed realibility test but error still occur | Performed troubleshooting | A1,C2,C4 reported already and confirmed for replacement C1 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6213610508916419094/6213610508916419094_signature_28.png | Jane Shyrel Enriquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6213610508916419094/6213610508916419094_signature_31.png
| 2025-04-25 | 2025-04-25 | Cebu City Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30698 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1893 | 2025-04-21 | 6C GXIV-4D | 838561 | Module A3 - H/W failed | -Replace defective module A3/862899 with module SN: 733337 | None | All modules are functional. |  | 900-0491RS Certified pre-own GeneXpert module SN. 733337 | DR no. 12944 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6213589606727728142/6213589606727728142_signature_28.png | Hannah Jane Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/6213589606727728142/6213589606727728142_signature_31.png
| 2025-04-24 | 2025-04-23 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | Tb Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 30697 | DOH | johnpaullabutap@gmail.com | 27 |  | 2025-04-23 | Genexpert | 838551 | Due for monthly PM | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212880043242599995/6212880043242599995_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6212880043242599995/6212880043242599995_signature_31.png
| 2025-04-24 | 2025-04-24 | Dumarao Rural Health Unit | Dumarao, Capiz | 6 | Capiz | TB Dots | NO | ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30696 | UNOPS | doregoshielamarie@gmail.com | 27 |  |  | 10C GXIV-4D | 110023542 | None | None | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212867725611948903/6212867725611948903_signature_28.png | Shiela Marie Dorego | https://www.jotform.com/uploads/macarewecare/222927978372067/6212867725611948903/6212867725611948903_signature_31.png
| 2025-04-24 | 2025-04-24 | Vinzons Rural Health 3 | Vinzons | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30695 | DOH | bardonamabel@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1207 |  | GX4-6C | 838576 | Gx for Preventive Maintenance Request for gx fan filter Request for keyboard replacement | Provide 1 fan filter Replaced keyboard Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212858009711758996/6212858009711758996_signature_28.png | Amabel A. Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/6212858009711758996/6212858009711758996_signature_31.png
| 2025-04-24 | 2025-04-24 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | INSTALLATION | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | dolly.banares@redcross.org.ph |  | 30694 | MACARE-HBDC | prcclinlab@redcross.org.ph | 25 |  | 2025-04-23 | MCS+ | 16K521SPG | For installation of new UPS. | Installation of UPS. | Performed installation of new UPS. Performed functionality testing of UPS. | Unit is operational. |  | New UPS (AT1K240323ST0038) Defective UPS (AT1000210708ST0032) | DR#12960 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212842805911051651/6212842805911051651_signature_28.png | Yna Cabradilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6212842805911051651/6212842805911051651_signature_31.png
| 2025-04-24 | 2025-04-24 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | INSTALLATION TROUBLESHOOTING | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | dolly.banares@redcross.org.ph |  | 30693 | MACARE-HBDC | prcclinlab@redcross.org.ph | 25 |  | 2025-04-23 | MCS+ | 12D063 | For replacement of UPS. For checking of transfer pump motor. | Installation of ups and troubleshooting. | Performed installation of new UPS. Performed functionality testing of UPS. Performed checking of transfer pump motor. | Unit is operational. For replacement of transfer pump motor due to wear and tear. Transfer pump has resistance on spinning and noisy when powered on. |  | New UPS (AT1K240323ST0039) Defective UPS (AT1000190402ST0376) | DR#12960 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212842795912070506/6212842795912070506_signature_28.png | Yna Cabradilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6212842795912070506/6212842795912070506_signature_31.png
| 2025-04-24 | 2025-04-23 | St. Luke's Medical Center - BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 30691 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 27 ASSAY: Xpert CT/NG, COUNT: 136 ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert HPV, COUNT: 190 ASSAY: Xpert MTB/RIF Ultra, COUNT: 373 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 82 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A3 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212841926282056663/6212841926282056663_signature_28.png | Jericho Jalandoni | https://www.jotform.com/uploads/macarewecare/222927978372067/6212841926282056663/6212841926282056663_signature_31.png
| 2025-04-24 | 2025-04-23 | St. Luke's Medical Center - BGC | Rizal Dr, BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 30692 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 153 ASSAY: Xpert CT/NG, COUNT: 576 ASSAY: Xpert HBV VL, COUNT: 48 ASSAY: Xpert HPV, COUNT: 789 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1982 ASSAY: Xpert MTB-XDR, COUNT: 5 ASSAY: Xpert HCV VL, COUNT: 36 ASSAY: Xpert HIV-1 Viral Load, COUNT: 226 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212841926284743171/6212841926284743171_signature_28.png | Jericho Jalandoni | https://www.jotform.com/uploads/macarewecare/222927978372067/6212841926284743171/6212841926284743171_signature_31.png
| 2025-04-24 | 2025-04-24 | Marine Environment & Resources Foundation, Inc. | Marine Science Institute, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Phycology Lab | NO | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 30690 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  | 2025-03-13 | PrimoStar HD | 3150000452 | None | None | Installed PrimoStar HD Performed functionality test Performed basic operator's training | Unit is operational | For demo | PrimoStar HD SN:3150000452 | DR 12825 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212836653617554687/6212836653617554687_signature_28.png | Leah Marie Mahomoc | https://www.jotform.com/uploads/macarewecare/222927978372067/6212836653617554687/6212836653617554687_signature_31.png
| 2025-04-24 | 2025-04-24 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30689 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 27 | ASSAY: Xpert SARS-CoV-2, COUNT: 0 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 981 |  | GXIV-D 6C | 829390 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212821540423106206/6212821540423106206_signature_28.png | Thea A. Dula | https://www.jotform.com/uploads/macarewecare/222927978372067/6212821540423106206/6212821540423106206_signature_31.png
| 2025-04-24 | 2025-04-24 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30688 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 27 | ASSAY: Xpert HBV VL, COUNT: 4 ASSAY: Xpert HPV, COUNT: 27 ASSAY: Xpert SARS-CoV-2, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1567 |  | GXIV-D 10C | 110017894 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212821070424474353/6212821070424474353_signature_28.png | Thea A. Dula | https://www.jotform.com/uploads/macarewecare/222927978372067/6212821070424474353/6212821070424474353_signature_31.png
| 2025-04-24 | 2025-04-24 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30687 | MACARE-HBDC | bghmcpcrlab@yahoo.com | 27 | ASSAY: Xpert HBV VL, COUNT: 57 ASSAY: Xpert HPV, COUNT: 52 ASSAY: Xpert HCV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 324 |  | GXIV-D 10C | 120003587 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit us operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212820650426459711/6212820650426459711_signature_28.png | Thea A. Dula | https://www.jotform.com/uploads/macarewecare/222927978372067/6212820650426459711/6212820650426459711_signature_31.png
| 2025-04-24 | 2025-04-24 | Marine Environment & Resources Foundation, Inc. | Marine Science Institute, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Phycology Lab | YES | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 30686 | COMMERCIAL | lcmahomoc@up.edu.ph | 25 |  | 2025-04-14 | Stemi 305 | 3947020979 | None | None | Installed Stemi 305 Trino Installed Axiocam 208 Color Performed functionality Test Performed basic operator's training | Unit is operational |  | Stemi 305 Trino SN:3947020979 Stand K Edu SN:3943037119 Axiocam 208 Color SN:5321007214 Spot Illuminator Router SN:2245065001299 | SI 1000135 and DR 12927 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212806654672075819/6212806654672075819_signature_28.png | Leah Marie Mahomoc | https://www.jotform.com/uploads/macarewecare/222927978372067/6212806654672075819/6212806654672075819_signature_31.png
| 2025-04-24 | 2025-04-24 | Labo District Hospital | Talobatib, Labo | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30685 | DOH | ldh.lab17@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1126 |  | GX4-6C | 838207 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data Check visibility of grid lines appearing on screen | Performed Preventive Maintenance | Unit is operational | Monitor request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212790919716151708/6212790919716151708_signature_28.png | ANALYN C. CONSUELO | https://www.jotform.com/uploads/macarewecare/222927978372067/6212790919716151708/6212790919716151708_signature_31.png
| 2025-04-24 | 2025-04-24 | Claver Rural Health Unit | Brgy. Tayaga, Claver | 13 | Surigao del Norte | TB-DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30684 | DOH | rhodaruthdalogdog@gmail.com | 27 |  |  | GX-4 6 Colors | 838514 | Module 1 found Error in Temperature after troubleshooting. Stuck Cartridge in Module 4. | Reported to Main Office. Manually removing of stuck cartridge. | Stuck Cartridge Removal Troubleshooting of Module 1 Preventive Maintenance | 3 modules functional, 1 module for reporting. Peripherals are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212751470414305155/6212751470414305155_signature_28.png | Joyce E. Andoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6212751470414305155/6212751470414305155_signature_31.png
| 2025-04-24 | 2025-04-24 | Mariveles District Hospital | Mariveles Bataan | 3 | Bataan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30683 | MACARE-HBDC | mdh.mdl2020@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 0 |  | GXIV-D | 806377 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212747302998518122/6212747302998518122_signature_28.png | Marigae Krista Monilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6212747302998518122/6212747302998518122_signature_31.png
| 2025-04-24 | 2025-04-24 | Midsalip Rural health unit | Midsalip, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30682 | DOH | genaliner@rocketmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 377 | 2025-04-24 | GX IV | 837664 | Module 3 h/w failed error 1001 | Preventive Maintenence and xpertcheck calibration. | All modules undergone preventive Maintenance. Only module 3 undergone Xpertcheck calibration | Module 1, 2 and 4 are available and ready to use. Module 3 needs more data. | None | XPERTCHECK-CE-5 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212744791219761878/6212744791219761878_signature_28.png | Genaline Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/6212744791219761878/6212744791219761878_signature_31.png
| 2025-04-24 | 2025-04-22 | Piñan Rural Health Unit | Piñan Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | MAINTENANCE Data collection | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 30681 | UNOPS | ivannagmarchan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 178 | 2025-04-22 | GX IV | 110023516 | None | Preventive Maintenence and data gathering | Preventive Maintenence and data collection | All modules are available and ready to use. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212730171214054846/6212730171214054846_signature_28.png | Ivanna Marchan | https://www.jotform.com/uploads/macarewecare/222927978372067/6212730171214054846/6212730171214054846_signature_31.png
| 2025-04-23 | 2025-04-23 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30680 | MACARE-HBDC | mongeshiela@gmail.com | 25 |  |  | MCS+ | 15H027G | Machine for PM | Cleaned unit surface Cleaned rotor pump rollers, valves, sensor | Performed Preventive Maintenance | Unit is not operational, CRC Bios defective |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212019554816728451/6212019554816728451_signature_28.png | Shiela P. Monge | https://www.jotform.com/uploads/macarewecare/222927978372067/6212019554816728451/6212019554816728451_signature_31.png
| 2025-04-23 | 2025-04-23 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30679 | MACARE-HBDC | mongeshiela@gmail.com | 25 |  |  | MCS+ | 15H031G | Leaked fluid/possible debris inside Centrifuge | Checked/Tested Centrifuge parameters Checked for any leftover debris/fluid Cleaned Centrifuge basin Cleaned unit surface Cleaned rotor pump rollers, valves, sensors | Performed Troubleshooting Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6212015834814483959/6212015834814483959_signature_28.png | Shiela P. Monge | https://www.jotform.com/uploads/macarewecare/222927978372067/6212015834814483959/6212015834814483959_signature_31.png
| 2025-04-23 | 2025-04-23 | Tagum City Health Office | Mabini Street, Tagum | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | tagumcholab.mabini@gmail.com |  | 30678 | DOH | cristiarafernandez@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2528 |  | GX-IV 6C 4D | 838553 | For Preventive Maintenance Module 2 encounter multiple temperature related errors | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance Perform Module Troubleshooting | Module B2 Disable due to frequent temperature related errors Reported to Cepheid for further Assessment Only Three Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211956110513893616/6211956110513893616_signature_28.png | Cristina Tiara A. Fernandez,. RMT. | https://www.jotform.com/uploads/macarewecare/222927978372067/6211956110513893616/6211956110513893616_signature_31.png
| 2025-04-23 | 2025-04-23 | Pontevedra Rural Health Unit | Pontevedra, Capiz | 6 | Capiz | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30677 | DOH | rtabngjr@gmail.com | 27 |  |  | 6C GXIV-4D | 838296 | -CPU defective -Module 2 frequent errors | -Replace CPU power supply -Run Xpertcheck on module 2 for checking | #NAME? | CPU is operational. Module 2 for replacement, failed XPC |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211945801601617524/6211945801601617524_signature_28.png | Mayeth Demonarca | https://www.jotform.com/uploads/macarewecare/222927978372067/6211945801601617524/6211945801601617524_signature_31.png
| 2025-04-23 | 2025-04-23 | Armed Forces Research Institute of Medical Sciences | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30676 | MACARE-HBDC | jaynueva19@gmail.com | 31 |  |  | GXIV-10C | 845462 | Unit for maintenance | NA | #NAME? | Unit is operational Room temp is advised to be monitored (Ave. temperature: 31C) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211930463549948595/6211930463549948595_signature_28.png | Jay Nueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6211930463549948595/6211930463549948595_signature_31.png
| 2025-04-23 | 2025-04-23 | Siargao Island Medical Center | Dapa, Siargao Island | 13 | Surigao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30675 | DOH | carlynvirtudazo10@gmail.com | 27 |  |  | GX-4 6 Colors | 838444 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211884846319162620/6211884846319162620_signature_28.png | Carlyn C. Virtudazo | https://www.jotform.com/uploads/macarewecare/222927978372067/6211884846319162620/6211884846319162620_signature_31.png
| 2025-04-22 | 2025-04-22 | Dasol Community Hospital | Dasol Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30674 | DOH | jeremypulido0521@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 596 |  | GXIV-D | 838287 | Due for preventive maintenance B2: 808250 invalid results error 2016 B3: 808251 Temperature error | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational B2 and B3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211236802523711246/6211236802523711246_signature_28.png | Jeremy Pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/6211236802523711246/6211236802523711246_signature_31.png
| 2025-04-22 | 2025-04-22 | Infanta Rural Health Unit | Infanta, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30673 | DOH | charissealmandres8@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 |  | GXIV-D | 110023501 | Quarterly Preventive Maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211201812911912073/6211201812911912073_signature_28.png | Charisse Alferez | https://www.jotform.com/uploads/macarewecare/222927978372067/6211201812911912073/6211201812911912073_signature_31.png
| 2025-04-22 | 2025-04-22 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 30672 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 33 |  |  | Gx4 L | 807155 | As per end user they encounter loss communication in module C1 | Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Reset gateway board Replaced lan cable of module C1 Test run unit | Performed troubleshooting | C1,C2,C3 operational C4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211163552224694718/6211163552224694718_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6211163552224694718/6211163552224694718_signature_31.png
| 2025-04-22 | 2025-04-22 | Surigao City Health Office | Brgy. San Juan, Surigao City | 13 | Surigao del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30671 | UNOPS | sanjuandistricthealthcenter23@gmail.com | 27 |  |  | GX-4 10 Colors | 110023522 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211155553983854849/6211155553983854849_signature_28.png | Anarose Ogdol | https://www.jotform.com/uploads/macarewecare/222927978372067/6211155553983854849/6211155553983854849_signature_31.png
| 2025-04-22 | 2025-04-22 | Philippine Red Cross Olongapo | 38 Magsaysay Dr. Olongapo Zambales | 3 | Zambales | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Mark Joseph Venzon |  | mark10venzon@gmail.com | olongapo.bloodservice@redcross.org.ph |  | 30670 | MACARE-HBDC | olongapo@redcross.org.ph | 25 |  |  | MCS+ | 03A043 | Due for calibration | Preventive maintenance Calibration | Cleaned surface area Checked parameters Performed calibration Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211079882919049207/6211079882919049207_signature_28.png | Philippine Red Cross - Olongapo City | https://www.jotform.com/uploads/macarewecare/222927978372067/6211079882919049207/6211079882919049207_signature_31.png
| 2025-04-22 | 2025-04-22 | President Ramon Magsaysay Memorial Hospital | Iba, Zambales | 3 | Zambales | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30669 | MACARE-HBDC | prmmhmoleclab@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 0 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 21 |  | GXIV-D | 838520 | Quarterly Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except modules B3:888978 Temperature error B4: 688212 Temperature Error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211079612914361646/6211079612914361646_signature_28.png | Bianca Yzbel Yap | https://www.jotform.com/uploads/macarewecare/222927978372067/6211079612914361646/6211079612914361646_signature_31.png
| 2025-04-22 | 2025-04-22 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30668 | DOH | rhwclab@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 407 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1576 |  | GX-IV 6C 4D | 838262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Only Three Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211068531525693900/6211068531525693900_signature_28.png | Ellaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6211068531525693900/6211068531525693900_signature_31.png
| 2025-04-22 | 2025-04-22 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30667 | MACARE-HBDC | rhwclab@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 402 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1576 |  | GX-IV 6C 4D | 839236 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Module B3 disabled due to multiple temperature related Only Three Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211066121524832230/6211066121524832230_signature_28.png | Ellaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6211066121524832230/6211066121524832230_signature_31.png
| 2025-04-22 | 2025-04-22 | Vicente Mendiola Center for Health | Naga City | 7 | Cebu | TB Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30666 | DOH | geccaprocianos26@gmail.com | 23 |  |  | Gx4 | 838564 | Due for monthly PM | Save Installation Qualification Report Save System Log file Save Assay Statistics Report. | Perform Preventive Maintenance for Genexpert Machine. | All module's are function ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211045678028543357/6211045678028543357_signature_28.png | Gene Carmel T. Procianos | https://www.jotform.com/uploads/macarewecare/222927978372067/6211045678028543357/6211045678028543357_signature_31.png
| 2025-04-22 | 2025-04-22 | Chinese General Hospital | Blumentrit, Sta. Cruz Manila | NCR | NCR 1st - Manila | Bacteriology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 30665 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1639 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 129 |  | GXIV-10C | 120002026 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211011081817336439/6211011081817336439_signature_28.png | Maria Odessa Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/6211011081817336439/6211011081817336439_signature_31.png
| 2025-04-22 | 2025-04-22 | Chinese General Hospital | Blumentrit, Sta. Cruz Manila | NCR | NCR 1st - Manila | Bacteriology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 30664 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1639 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 129 |  | GXIV-10C | 120002026 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6211010681818946622/6211010681818946622_signature_28.png | Maria Odessa Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/6211010681818946622/6211010681818946622_signature_31.png
| 2025-04-22 | 2025-04-22 | Dumarao Rural Health Unit | Dumarao, Capiz | 6 | Capiz | TB Dots | NO | REPAIR TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30663 | UNOPS | doregoshielamarie@gmail.com | 27 |  |  | 10C GXIV-4D | 110023542 | CPU to monitor connection issue | -Install CPU temporarily while waiting for CPU replacement -Backup test results -Upgrade software to 6.5 | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6210975295226954122/6210975295226954122_signature_28.png | Shiela Marie Dorego | https://www.jotform.com/uploads/macarewecare/222927978372067/6210975295226954122/6210975295226954122_signature_31.png
| 2025-04-21 | 2025-04-15 | South Cotabato Provincial Hospital / IPHO Covid lab | Koronadal City | 12 | South Cotabato | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30662 | MACARE-HBDC | dappmc.tblab@gmail.com | 27.6 | ASSAY: Xpert SARS-CoV-2, COUNT: 881 |  | Gx-XVI | 839314 | Module C2, D2,D3 and D4 Disabled due to multiple errors | Perform Genexpert Cleaning and maintenance Module Cleaning Plunger rod and Optics Cleaning Filter Cleaning Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 12 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6210569725377872610/6210569725377872610_signature_28.png | Jasmine Libona | https://www.jotform.com/uploads/macarewecare/222927978372067/6210569725377872610/6210569725377872610_signature_31.png
| 2025-04-21 | 2025-04-16 | Naguilian District Hospital | Naguilian, La Union | 1 | La Union | Laboratory | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30661 | DOH | marjelynrayray01@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 781 | 2025-04-10 | GXIV-4-D-6C | 838426 | For UPS Replacement | #NAME? | Replaced Defective UPS | #NAME? |  | UPS AT1000181128ST0151 | DR 12921 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6210211567673133383/6210211567673133383_signature_28.png | Marjelyn Rayray | https://www.jotform.com/uploads/macarewecare/222927978372067/6210211567673133383/6210211567673133383_signature_31.png
| 2025-04-21 | 2025-04-14 | Cebu South Medical Center | San Isidro, Talisay City Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | csmcgenexpert@gmail.com |  |  | 30660 | DOH | csmcgenexpert@gmail.com | 24 | ASSAY: , COUNT: 1065 |  | GX4 | 838559 | Genexpert process encountered 2097 for HIV-Assay 5007 errors for MTB-Ultra in Module 2. | Recommend user to strictly follow the sample processing of HIV-Viral Load and load only in Module 2 for HIV-Viral load only for isolation and determine repeated error of 2097 | Troubleshoot error and generate IQR, System Log and Gxx files. Clean up Icore and Airfilter cleaner. Observed sample process and review cartridge integrity of HIV-Viral Load. | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6210139373443854705/6210139373443854705_signature_28.png | Cebu South Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6210139373443854705/6210139373443854705_signature_31.png
| 2025-04-21 | 2025-04-21 | Calamba Rural Health Unit | Calamba misamis occidental | 10 | Misamis Occidental | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 30659 | DOH | palubonibn@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 382 | 2025-04-21 | GX IV | 110023518 | None | Preventive Maintenence | Preventive Maintenence. | All modules are available and ready to use | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6210106986313321287/6210106986313321287_signature_28.png | Mary-ibn E Palubon | https://www.jotform.com/uploads/macarewecare/222927978372067/6210106986313321287/6210106986313321287_signature_31.png
| 2025-04-21 | 2025-04-21 | Manila Doctors Hospital | United Nations Ave, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 30658 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 21 |  |  | GXIV | 110001648 | Unit for transfer back to original location after facilities maintenance. | #NAME? | #NAME? | Unit is operational. Original keyboard has defective keys. End-user temporarily using own keyboard. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6210100887397205189/6210100887397205189_signature_28.png | Kaye martelino | https://www.jotform.com/uploads/macarewecare/222927978372067/6210100887397205189/6210100887397205189_signature_31.png
| 2025-04-16 | 2025-04-16 | Argao Primary Health Care (RHU II) | Poblacion Argao Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30657 | DOH | victoria.jumamil@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1500 |  | GX4 | 838283 | CPU power supply problem. | Check power supply of UPS Check outlet of the 220VAC Replacement of power cable Generate IQR, System log, Assay, Archive data files, copy database | Check power supply of UPS Check outlet of the 220VAC Replacement of power cable Generate IQR, System log, Assay, Archive data files, copy database | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205985573443145705/6205985573443145705_signature_28.png | Victoria Jumamil, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6205985573443145705/6205985573443145705_signature_31.png
| 2025-04-16 | 2025-04-16 | Manila Doctors Hospital | United Nations Ave, Ermita, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 30656 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 27 | ASSAY: Xpert CT/NG, COUNT: 74 ASSAY: Xpert HBV VL, COUNT: 109 ASSAY: Xpert HPV, COUNT: 25 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1220 ASSAY: Xpert HCV VL, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 127 |  | GXIV | 110001648 | Customer requested transfer of unit to different room temporarily due to ongoing facility maintenance. | None | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205973113598172854/6205973113598172854_signature_28.png | Kaye Martelino | https://www.jotform.com/uploads/macarewecare/222927978372067/6205973113598172854/6205973113598172854_signature_31.png
| 2025-04-16 | 2025-04-14 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30655 | MACARE-HBDC | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 48 |  | GX16D | 804476 | For disassemble of A3 882239,C1 881354,C2 882120,C3 760518,C4 882137, gateway board 6001161 | Remove module and gateway board for condemnation. | Performed module disassemble and gateway board | Unit for condemn |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205731893616503504/6205731893616503504_signature_28.png | Ma,Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6205731893616503504/6205731893616503504_signature_31.png
| 2025-04-16 | 2025-04-15 | Metro Antipolo Hospital and Medical Center | Marcos Highway Antipolo Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30653 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 19 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 285 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GX-IV-D-6C | 110005265 | Module B3 and B4 for XpertCheck (retry) | Run XpertCheck on module B3 and B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck on module B3 and B4 Generated reports | Unit os operational. Module B3 and B4 for retry, insufficient XpertCheck cartridge. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205720583611771360/6205720583611771360_signature_28.png | Graziela Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/6205720583611771360/6205720583611771360_signature_31.png
| 2025-04-16 | 2025-04-14 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30654 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 61 ASSAY: Xpert CT/NG, COUNT: 49 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 658 ASSAY: Xpert HPV, COUNT: 403 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6385 ASSAY: Xpert MTB-XDR, COUNT: 3 ASSAY: Others, COUNT: 2 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 153 ASSAY: Xpert HIV-1 Viral Load, COUNT: 858 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1423 |  | GX-XVI-D-10C | 110015767 | Error 1001 on module D1/210082300 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -icore -plunger rod -filter Checked module D1 Generated reports Generated gxx file | Unit is operational. Module D1 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205720583618053400/6205720583618053400_signature_28.png | Maria Angelica Medel | https://www.jotform.com/uploads/macarewecare/222927978372067/6205720583618053400/6205720583618053400_signature_31.png
| 2025-04-15 | 2025-04-03 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30652 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 20 |  |  | GXIV-4-D-6C | 110016112 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205060617679071725/6205060617679071725_signature_28.png | Melvin Magbalita | https://www.jotform.com/uploads/macarewecare/222927978372067/6205060617679071725/6205060617679071725_signature_31.png
| 2025-04-15 | 2025-04-03 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30651 | UNOPS | benguetphortdl@gmail.com | 29 |  |  | GXIV-4-D-10C | 110023498 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205059237672322627/6205059237672322627_signature_28.png | Tillber Baniaga Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6205059237672322627/6205059237672322627_signature_31.png
| 2025-04-15 | 2025-04-03 | Baguio City Health Office | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30650 | MACARE-HBDC | sherryannemarrero@gmail.com | 19 |  |  | GXIV-4-D-6C | 838281 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205056967674216059/6205056967674216059_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6205056967674216059/6205056967674216059_signature_31.png
| 2025-04-15 | 2025-04-03 | Itogon Rural Health Unit | Itogon, Benguet | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30649 | DOH | ggdepayso@gmail.com | 30 |  |  | GXIV-4-D-6C | 838282 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational Module B4 for Reporting | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205055937675131606/6205055937675131606_signature_28.png | Gloria G Depayso | https://www.jotform.com/uploads/macarewecare/222927978372067/6205055937675131606/6205055937675131606_signature_31.png
| 2025-04-15 | 2025-04-02 | Manabo Rural Health Unit | Manabo, Abra | 1 | Abra | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30648 | DOH | lenmayclairelumiwag15@gmail.com | 26 |  | 2025-03-18 | GXIV-4-D-6C | 838274 | For Replacement of Desktop | #NAME? | Installed Original Desktop | - Unit is Partially Operational - Modules 1 and 2 for Reporting | Work with Engr. Venzon | GENEXPERT DESKTOP CPH0520F08280 | DR 12864 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205049977679286722/6205049977679286722_signature_28.png | Ellen May Claire S. Lumiwag | https://www.jotform.com/uploads/macarewecare/222927978372067/6205049977679286722/6205049977679286722_signature_31.png
| 2025-04-15 | 2025-04-15 | Mabinay Medicare Community Hospital | Mabinay, Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30647 | DOH | mmchlaboratory2017@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 430 |  | GX4 6C | 838481 | Module A1/8629088 temperature-related errors, HW Failed in IQR | . | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205046196698564727/6205046196698564727_signature_28.png | Dianne Mi-ot | https://www.jotform.com/uploads/macarewecare/222927978372067/6205046196698564727/6205046196698564727_signature_31.png
| 2025-04-15 | 2025-04-15 | Our Lady of Mt. Carmel Medical Center | San Fernando City, Pampanga | 3 | Pampanga | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30646 | COMMERCIAL | mt.carmel_laboratory@yahoo.com | 25 |  |  | MCS+ | 15G043 | Safety System Error | Preventive maintenance Troubleshooting Reset safety board Adjust value of Line sensor, SM and DPM | Performed preventive maintenance Checked parameters Cleaned and checked rollers Performed functionality test | Unit is operational, for monitoring | Due for calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205046172324036978/6205046172324036978_signature_28.png | Patricia Puyat | https://www.jotform.com/uploads/macarewecare/222927978372067/6205046172324036978/6205046172324036978_signature_31.png
| 2025-04-15 | 2025-04-11 | Virac Municipal Health Office | Virac | 5 | Catanduanes | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.comq |  |  | 30645 | DOH | limarmariano50@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1901 |  | GX4-6C | 838563 | UPS defective CPU restarts | UPS replacement performed Cleaned RAM, removed cobwebs inside CPU Cleaned cartridge bays, plunger rod, | Replaced UPS Performed Preventive Maintenance Troubleshoot CPU | Unit is operational CPU under observation if problem persists |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205030208307112434/6205030208307112434_signature_28.png | Ramil A. Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/6205030208307112434/6205030208307112434_signature_31.png
| 2025-04-15 | 2025-04-15 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30644 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  | 2025-04-11 | GXIV | 110008530 | Unit for maintenance Module B1 for troubleshooting | Performed blank cartridge test on B1 | #NAME? | Unit is operational Data to be reported |  | MTB/Rif Ultra Cartridge | 12922 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205028830412587879/6205028830412587879_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6205028830412587879/6205028830412587879_signature_31.png
| 2025-04-15 | 2025-04-08 | Rapu-Rapu Rural Health Unit | Poblacion, Rapu-Rapu | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30643 | DOH | suzetteoguan1@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 561 |  | GX4-6C | 838565 | Monitor request for replacement (defective) | Replaced monitor Cleaned cartridge bays, plunger rod, icore Generate Data | Performed replacement of Monitor Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6205026718306155572/6205026718306155572_signature_28.png | Suzette C. Oguan | https://www.jotform.com/uploads/macarewecare/222927978372067/6205026718306155572/6205026718306155572_signature_31.png
| 2025-04-14 | 2025-04-08 | Isulan Rural Health Unit | Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | nancymoscoso18@gmail.com |  | 30641 | UNOPS | nanzmoscoso18@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 877 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 |  | GxIV-10C | 110023520 | None | Perform Genexpert Cleaning and maintenance Module Cleaning Plunger rod and Optics Cleaning Fan filter Cleaning Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6204428626862253627/6204428626862253627_signature_28.png | Nancy A. Moscoso | https://www.jotform.com/uploads/macarewecare/222927978372067/6204428626862253627/6204428626862253627_signature_31.png
| 2025-04-14 | 2025-04-10 | Tampakan Rural Health Unit | Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30642 | DOH | lotisglegada@gmail.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 614 |  | GxIV-6C | 838544 | Request for Gx Fan filter | Genexpert Cleaning and maintenance Module Cleaning Plunger rod and Optics Cleaning Fan filter Cleaning Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6204428626863932382/6204428626863932382_signature_28.png | Lotis G. Legada | https://www.jotform.com/uploads/macarewecare/222927978372067/6204428626863932382/6204428626863932382_signature_31.png
| 2025-04-14 | 2025-04-14 | Governor Generoso Rural Health Unit | Poblacion, Governor Generoso | 11 | Davao Oriental | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30640 | DOH | josephmagallanes72@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 465 |  | GX-IV 6C 4D | 838490 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6204216575411545606/6204216575411545606_signature_28.png | Charlene V. Magallanes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6204216575411545606/6204216575411545606_signature_31.png
| 2025-04-14 | 2025-04-14 | Poro Rural Health Unit | Poro, Camotes, Cebu | 7 | Cebu | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30639 | DOH | avilacharmaine6@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 431 |  | 6C GXIV-4D | 838279 | Due for PM | #NAME? | #NAME? | Machine is functional. | To run XPC on module 2 and replace fan filter |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6204207338504684410/6204207338504684410_signature_28.png | Charmaine Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/6204207338504684410/6204207338504684410_signature_31.png
| 2025-04-14 | 2025-04-14 | Magsaysay Rural Health Unit | Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30638 | DOH | mclmagsaysay@gmail.com | 27 |  |  | GX-4 6 Colors | 838557 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6204173661262352554/6204173661262352554_signature_28.png | May Anisa Jedd Gerona | https://www.jotform.com/uploads/macarewecare/222927978372067/6204173661262352554/6204173661262352554_signature_31.png
| 2025-04-14 | 2025-03-20 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Road, Las Piñas | NCR | NCR 4th - Las Piñas | Molecular | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | uphdmc_lab@yahoo.com |  | 30637 | COMMERCIAL | phmclp.mdl@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 84 ASSAY: Xpert HCV VL, COUNT: 1 |  | GXIV-D-6C | 831316 | Barcode Scanning Problem | Troubleshoot | Reprogram Barcode Configuration. Upgrade Software from 6.1-6.4. Generate Reports. Performed Self-test. | GX is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6204113603618177013/6204113603618177013_signature_28.png | Donn Dave Clerigo | https://www.jotform.com/uploads/macarewecare/222927978372067/6204113603618177013/6204113603618177013_signature_31.png
| 2025-04-14 | 2025-04-11 | Metro Antipolo Hospital and Medical Center | Marcos Highway Antipolo Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30636 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 19 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 286 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GX-IV-D-6C | 110005265 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. Module B1 passed, Module B2 service required. Module B3 and B4 for retry, insufficient XpertCheck cartridge. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6204099533615379893/6204099533615379893_signature_28.png | Graziela Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/6204099533615379893/6204099533615379893_signature_31.png
| 2025-04-14 | 2025-04-11 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 30635 | UNOPS | molecular.laboratory@eamc.doh.gov.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 825 ASSAY: Xpert HPV, COUNT: 74 ASSAY: Xpert SARS-CoV-2, COUNT: 606 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 281 |  | GX4D (10C) | 110016781 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6203982603612079651/6203982603612079651_signature_28.png | East Avenue Medical center- Molecular | https://www.jotform.com/uploads/macarewecare/222927978372067/6203982603612079651/6203982603612079651_signature_31.png
| 2025-04-14 | 2025-04-11 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | TB dots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 30634 | MACARE-HBDC | nelsonmiranda791@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3096 |  | GX4D(10C) | 110014085 | Due for maintenance B4 disabled due to error 5007 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Due for xpertcheck last xpertcheck 3/22/24 | For quotation of xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6203982093617031764/6203982093617031764_signature_28.png | Nelson miranda | https://www.jotform.com/uploads/macarewecare/222927978372067/6203982093617031764/6203982093617031764_signature_31.png
| 2025-04-14 | 2025-04-11 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30633 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 16K521SPG | Unit for preventive maintenance. For checking of UPS. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. Checked UPS backing system. | Unit is operational. Rotor pumps roller are okay. UPS AT1000 (sn:AT1000210708ST0032) not backing up after disconnecting to supply. UPS for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6203964546763898891/6203964546763898891_signature_28.png | Dolla Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6203964546763898891/6203964546763898891_signature_31.png
| 2025-04-14 | 2025-04-11 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30632 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. For checking of UPS. For checking of rotor pumps. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. Checked UPS backing system. | Unit is operational. Rotor pumps roller are for replacement. UPS AT1000 (sn:AT1000190402ST0376) not backing up after disconnecting to supply. UPS for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6203964186762095283/6203964186762095283_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6203964186762095283/6203964186762095283_signature_31.png
| 2025-04-14 | 2025-04-11 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30631 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. For checking of UPS. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. Checked UPS backing system. | Unit is operational. Rotor pumps roller are okay. UPS AT1000 (sn:AT1000190402ST0376) not backing up after disconnecting to supply. UPS for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6203963796763636295/6203963796763636295_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6203963796763636295/6203963796763636295_signature_31.png
| 2025-04-14 | 2025-04-11 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30630 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 16K521SPG | Unit for preventive maintenance. For checking of UPS. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. Checked UPS backing system. | Unit is operational. Rotor pumps roller are okay. UPS AT1000 (sn:AT1000210708ST0032) not backing up after disconnecting to supply. UPS for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6203963796765697582/6203963796765697582_signature_28.png | Dolla Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6203963796765697582/6203963796765697582_signature_31.png
| 2025-04-11 | 2025-04-11 | Cebu Doctor's University Hospital | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30629 | COMMERCIAL | jdnoel.eisi@gmail.com | 22 |  |  | Gx4 | 110008542 | Due for monthly preventive maintenance. | Save installation qualification report Save Assay Statistics report Save system log file | Clean filter Clean plunger rod maintenance Clean cable wires for module's | Module B1, B2 ok to use Module B3, B4 not ok to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201614765886290478/6201614765886290478_signature_28.png | Gillian Celis | https://www.jotform.com/uploads/macarewecare/222927978372067/6201614765886290478/6201614765886290478_signature_31.png
| 2025-04-11 | 2025-04-08 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco Ave, Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30628 | MACARE-HBDC | peachmango318@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1108 |  | GX-IV-D-10C | 110024245 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201606702449064412/6201606702449064412_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6201606702449064412/6201606702449064412_signature_31.png
| 2025-04-11 | 2025-04-11 | Santolan Super Health Center | Evangelista Ave., Santolan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30626 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1420 |  | GX-IV-D-10C | 110024247 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201606692444101599/6201606692444101599_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6201606692444101599/6201606692444101599_signature_31.png
| 2025-04-11 | 2025-04-08 | Rosario Super Health Center | Rosario, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30627 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1239 |  | GX-IV-D-10C | 110024246 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201606702443525269/6201606702443525269_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6201606702443525269/6201606702443525269_signature_31.png
| 2025-04-11 | 2025-04-11 | Solano Rural Health Unit | Roxas, Solano, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30625 | DOH | rhusolanolab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 744 | 2025-03-31 | GXIV-6C | 838201 | UPS not working | Replaced UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | UPS | 12891 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201570493101834691/6201570493101834691_signature_28.png | May Vanessa Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6201570493101834691/6201570493101834691_signature_31.png
| 2025-04-11 | 2025-04-08 | Teresita L. Jalandoni Provincial Hospital | Silay City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30624 | DOH | tljpmollab_02@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2537 ASSAY: Xpert MTB-XDR, COUNT: 164 |  | GX4 6C | 838464 | E2120 in module A1 and A2 E1002 in module A2 | Preventive maintenance - cleaned optics, plunger rods, module bays, filter; archived tests. | #NAME? | -Modules A1 and A2 passed the Xpertcheck; -E5015 encountered in Module A1, for evaluation by Application Specialist; -E1002 and E4009 in module A2, for replacement by Cepheid. - Modules A3 and A4 functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201570296298915873/6201570296298915873_signature_28.png | Ma Kriska Awacay | https://www.jotform.com/uploads/macarewecare/222927978372067/6201570296298915873/6201570296298915873_signature_31.png
| 2025-04-11 | 2025-04-11 | Monkayo Rural Health Unit | Poblacion Monkayo | 11 | Davao de Oro | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30623 | DOH | kendralois730@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1438 |  | GX-IV 6C 4D | 838487 | UPS Battery Mode not working For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replace the defective UPS Battery with new Battery | Perform Preventive Maintenance and UPS Battery Replacement | All Modules are operational UPS Battery successfully replaced | High temp in GeneXpert, enduser is advised to achieve optimal room temperature before running test to avoid temperature related errors |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201550379773697635/6201550379773697635_signature_28.png | Karen S. Anuta,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6201550379773697635/6201550379773697635_signature_31.png
| 2025-04-11 | 2025-04-11 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30622 | DOH | bmccrcl2022@gmail.com | 30 |  |  | GX-4 6 Colors | 838403 | None | None | Installation of HIV Viral Load Assay and Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201519167324314756/6201519167324314756_signature_28.png | Joechita L. Doysabas | https://www.jotform.com/uploads/macarewecare/222927978372067/6201519167324314756/6201519167324314756_signature_31.png
| 2025-04-11 | 2025-04-10 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30621 | DOH | suzzettekho34@gmail.com | 30 |  |  | GX-4 6 Colors | 838554 | None | None | Installation of HIV Viral Load Assay and Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201516987324719312/6201516987324719312_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/6201516987324719312/6201516987324719312_signature_31.png
| 2025-04-11 | 2025-04-11 | Chinese General Hospital | Blumentritt, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30620 | COMMERCIAL | cghmcbloodbank2016@gmail.com | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. For checking of rubber roller. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. For replacement of rubber roller. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6201430164876303755/6201430164876303755_signature_28.png | Jacqueline J. Zatzoso | https://www.jotform.com/uploads/macarewecare/222927978372067/6201430164876303755/6201430164876303755_signature_31.png
| 2025-04-10 | 2025-04-10 | Cauayan City Health Office | Cauayan City, Isabela | 2 | Isabela | Microbiology | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30619 | UNOPS | cho1lab14@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1260 |  | GXIV-10c | 110023504 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6200786685163940520/6200786685163940520_signature_28.png | Arlyn Daggao | https://www.jotform.com/uploads/macarewecare/222927978372067/6200786685163940520/6200786685163940520_signature_31.png
| 2025-04-10 | 2025-04-10 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30618 | MACARE-HBDC | treatmenthub.qcgh@gmail.com | 25 |  |  | GXIV-10C | 110012766 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6200732237766323406/6200732237766323406_signature_28.png | Arizza Sheena Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/6200732237766323406/6200732237766323406_signature_31.png
| 2025-04-10 | 2025-04-10 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laborstory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30617 | UNOPS | jomarilitonjua@gmail.com | 25 |  |  | GXIV-10C | 110023495 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6200728537765886509/6200728537765886509_signature_28.png | Jomari Litonjua | https://www.jotform.com/uploads/macarewecare/222927978372067/6200728537765886509/6200728537765886509_signature_31.png
| 2025-04-10 | 2025-04-10 | Cebu City Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30616 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2073 |  | 6C GXIV-4D | 838561 | Due for monthly PM | -Perorm preventive maintenance; fan filter, plunger rod, optic slot, cartridge bay | #NAME? | Module A3 waiting for replacement, other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6200726999238889268/6200726999238889268_signature_28.png | Hannah Jane Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/6200726999238889268/6200726999238889268_signature_31.png
| 2025-04-10 | 2025-04-10 | Baybay City Health Office | Baybay Leyte | 8 | Leyte | Tb Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30615 | DOH | sajecopooh20@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1171 | 2025-04-10 | GXIV | 838555 | Module B4 valve positioning error (1018) | #NAME? | #NAME? | Unit is Operational |  | 1 Xpertcheck cartridge | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6200708174538323444/6200708174538323444_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/6200708174538323444/6200708174538323444_signature_31.png
| 2025-04-10 | 2025-04-10 | Santo Tomas Rural Health Unit | Bolinao, Santo Tomas Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30614 | MACARE-HBDC | gloangelier@gmail.com | 29 |  | 2025-03-12 | GXIV-10c | 110024248 | Due for Calibration | Performed Xpertcheck) (B1 = Passed) (B2 = Passed) (B3 = Passed) (B4 = Service Required) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | 12834 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6200677808379327238/6200677808379327238_signature_28.png | Elyssa Joy Tagulao | https://www.jotform.com/uploads/macarewecare/222927978372067/6200677808379327238/6200677808379327238_signature_31.png
| 2025-04-09 | 2025-04-09 | Maddela District Hospital | Maddela, Quirino | 2 | Quirino | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30613 | DOH | mdhlaboratory00@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 361 | 2025-04-02 | GXIV-6C | 838469 | Keyboard Problem UPS no power Barcode scanner not working | Replaced Keyboard Replaced UPS Replaced Barcode Scanner | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Keyboard UPS Barcode Scanner | 12898 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6199987293224289219/6199987293224289219_signature_28.png | Donna May Cañega | https://www.jotform.com/uploads/macarewecare/222927978372067/6199987293224289219/6199987293224289219_signature_31.png
| 2025-04-08 | 2025-04-08 | Balamban Rural Health Unit | Poblacion Balamban | 7 | Cebu | TB Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30612 | DOH | balambanrhu45@gmail.com | 22 |  | 2025-04-08 | GX4 | 838231 | Monthly Preventive Maintenance. | Save System Log file Save Installation qualification reports Save Assay Statistic teports | Perform preventive maintenance for all modules. | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6199037184617969128/6199037184617969128_signature_28.png | Mildred E. Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6199037184617969128/6199037184617969128_signature_31.png
| 2025-04-08 | 2025-04-07 | MMC (Medilinx) | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30611 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 20 ASSAY: Xpert CT/NG, COUNT: 13 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 82 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 10 ASSAY: Xpert HIV-1 Viral Load, COUNT: 54 |  | GX16D(10C) | 110026448 | As per end user they encounter intermittent error in c2 4014,2034 optical signal detector, C4 2034,4014 optical signal detector A2 loose door ejector | Gathered reports Cleaned ICore Fixed door ejector of A2 module | Troubleshooting | C2 and C4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6198921413613446824/6198921413613446824_signature_28.png | Aiza G. Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6198921413613446824/6198921413613446824_signature_31.png
| 2025-04-08 | 2025-04-08 | Chong Hua Hospital | Fuente Osmena Capitol Site | 7 | Cebu | Bloodbank | NO | REPAIR CALIBRATION | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30610 | COMMERCIAL | zandro_tisoy@honghua.com.ph | 23 |  |  | Mcs+ | 11J219 | Defective CPU and Pneumatic Assembly | 1.Replacement of New cpu board and pneumatic assembly 2. Calibrate mcs+ bowl optics, line sensor, weigher, dpm and spm pressure monitor | 1.Replacement of New cpu board and pneumatic assembly 2. Calibrate mcs+ bowl optics, line sensor, weigher, dpm and spm pressure monitor | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6198917825329944059/6198917825329944059_signature_28.png | Chong Hua Hospital Fuente | https://www.jotform.com/uploads/macarewecare/222927978372067/6198917825329944059/6198917825329944059_signature_31.png
| 2025-04-08 | 2025-04-08 | Bacolod City Health Office | Bacolod City, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30609 | DOH | recelenylrej@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6875 |  | GX4 6C | 837612 | None | Preventive maintenance - cleaned optics, plunger rods, module bays, filter; archived tests. | Preventive Maintenance | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6198856169749725314/6198856169749725314_signature_28.png | Recele Nylrej Geronga | https://www.jotform.com/uploads/macarewecare/222927978372067/6198856169749725314/6198856169749725314_signature_31.png
| 2025-04-07 | 2025-04-03 | Bagumbayan Rural Health Unit | Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30608 | DOH | marizalenoid@yahoo.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1295 |  | GxIV-6C | 838261 | Module 1 - H/W failed | Perform module 1 troubleshooting Genexpert Cleaning: Plunger rod and Optics Cleaning Module Cleaning Filter Cleaning Errors/problem reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert module troubleshooting | 3 functional modules | Module 1 requesting for module replacement due to multiple temperature errors. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6198394736869845842/6198394736869845842_signature_28.png | Ma . Riza R. Dionela | https://www.jotform.com/uploads/macarewecare/222927978372067/6198394736869845842/6198394736869845842_signature_31.png
| 2025-04-07 | 2025-04-07 | Binalbagan Municipal Health Office | Binalbagan, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30607 | DOH | anne_galilea@yahoo.com | 27 |  |  | GX4 6C | 838280 | CPU intermittent power problem; System date problem | Preventive maintenance - cleaned optics, module bays, filter; archived tests. | #NAME? | CPU operational; 4 modules functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6198195323525306107/6198195323525306107_signature_28.png | Xerxanne Galilea | https://www.jotform.com/uploads/macarewecare/222927978372067/6198195323525306107/6198195323525306107_signature_31.png
| 2025-04-07 | 2025-04-04 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | YES | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | nch.jeff22@gmail.com |  | 30606 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 45 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1536 ASSAY: Xpert SARS-CoV-2, COUNT: 68 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1624 | 2025-04-04 | GX16D | 110014171 | Due for xpertcheck and maintenance. A2 s/n 210068930,B3 s/n 210067748 reported already for replacement due to error 5007 | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck and maintenance | Failed in xpert check A1 s/n 887323,A3 s/n 210077159,A4 s/n 210069006,B1 s/n 210002752,B2 s/n 210069275,B4 s/n 210062093,C1 s/n 210069271 Passed in xpertcheck C3 s/n 898257,C4 s/n 210067266 For replacement of module A1,A2,A3,A4,B1,B2,B3,B4,C1 | Retry for xpertcheck due to lack of kits D1 614214,D2 210078182,D3 210025172,D4210069227,C2 210069270 | Xpertcheck kits | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6197969003611394484/6197969003611394484_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/6197969003611394484/6197969003611394484_signature_31.png
| 2025-04-07 | 2025-04-03 | Puerto Princesa City Health Office | Valencia St. Puerto Princesa Palawan | 4B | Palawan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30605 | MACARE-HBDC | jocal0324@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2251 |  | GX-XVI-D-10C | 110021518 | For maintenance | Run XpertCheck | For maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6197930623612514256/6197930623612514256_signature_28.png | Jocelyn Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/6197930623612514256/6197930623612514256_signature_31.png
| 2025-04-07 | 2025-04-03 | Puerto Princesa City Health Office | Valencia St. Puerto Princesa, Palawan | 4B | Palawan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30603 | UNOPS | jocal0324@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 196 |  | GX-IV-D-10C | 110023510 | For maintenance | Performed maintenance | For maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6197930623614800246/6197930623614800246_signature_28.png | Jocelyn Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/6197930623614800246/6197930623614800246_signature_31.png
| 2025-04-07 | 2025-04-02 | San Vicente Municipal Health Office | San Vicente, Palawan | 4B | Palawan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30604 | MACARE-HBDC | santosdollethorecel69@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 257 |  | GX-IV-L-10C | 120002441 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6197930623618451142/6197930623618451142_signature_28.png | Dollethorecel Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6197930623618451142/6197930623618451142_signature_31.png
| 2025-04-04 | 2025-04-02 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Apheresis room | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30602 | MACARE-HBDC | strobecatcher@gmail.com | 26 |  | 2025-03-28 | MCS + | 15HO29G | CRC bios failed,defective battery | Delivered and installed battery Performed calibration. Swap cuff compressor from mcs s/n 15H028G to 15H029G Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed repair | Operational |  | Ltc-7PN Keeper 3.5v 750mAH battery | 12895 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6195802280823435745/6195802280823435745_signature_28.png | Ivan Ritz P. Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6195802280823435745/6195802280823435745_signature_31.png
| 2025-04-04 | 2025-04-04 | Tuburan rural health unit | Tuburan, basilan | BARMM | Basilan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30601 | DOH | rjakilan20@gmail.com | 28 |  | 2025-04-04 | GX 4 | 838832 | Genexpert application unable to open. | Replaced hard drive of the system unit, disabled automatic updates to evade same problem occurance. Preventive Maintenance. | Replaced hard drive of the system unit, disabled automatic updates to evade same problem occurance. Preventive Maintenance. Set up new account for the end user. | All modules are available and ready to use. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6195449571671875465/6195449571671875465_signature_28.png | Raiyana Jakilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6195449571671875465/6195449571671875465_signature_31.png
| 2025-04-04 | 2025-04-04 | Southern Philippines Medical Center | JP. Laurel Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30600 | UNOPS | spmcrhivda@gmail.com | 26.8 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1277 ASSAY: Xpert SARS-CoV-2, COUNT: 6 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1108 |  | GX-IV 10C 4D | 110023505 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6195388984526784909/6195388984526784909_signature_28.png | Flordeliza Go,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6195388984526784909/6195388984526784909_signature_31.png
| 2025-04-04 | 2025-04-04 | Southern Philippines Medical Center | JP Laurel Avenue Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30599 | UNOPS | spmcrhivda@gmail.com | 26.8 | ASSAY: Xpert SARS-CoV-2, COUNT: 53 ASSAY: Xpert MTB/RIF Ultra, COUNT: 33 ASSAY: Xpert HIV-1 Viral Load, COUNT: 320 |  | GX-IV 10C 4D | 110023506 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6195388654525311425/6195388654525311425_signature_28.png | Flordeliza Go,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6195388654525311425/6195388654525311425_signature_31.png
| 2025-04-03 | 2025-04-03 | Manuel J. Santos Hospital | 554 Montilla Blvd, Butuan City | 13 | Agusan del Sur | Molecular Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30598 | COMMERCIAL | molecularlab@mjsantos.com.ph | 23 |  |  | GX-4 6 Colors | 831292 | Module A1 and A2 found temperature errors and failed in calibration last year. | None | Installation of HIV-1 Viral Load Assay Test and Preventive Maintenance | 2 modules running, 2 module for replacement Computer set and peripherals are functional. | Advised to request for module replacement of A1 and A2 to the Main Office. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6194699059719594077/6194699059719594077_signature_28.png | Daniella Andrei Plaza | https://www.jotform.com/uploads/macarewecare/222927978372067/6194699059719594077/6194699059719594077_signature_31.png
| 2025-04-03 | 2025-04-03 | PRC-Olongapo City Chapter | Olongapo City | 3 | Zambales | Blood Bank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com | jona.liza@redcross.org.ph |  | 30597 | COMMERCIAL | olongapo@redcross.org.ph | 29 |  | 2025-04-03 | MCS + | 03A043 | Centrifuge unable to close and lock/unlock. | Performed centrifuge troubleshooting and repair. | Checked and clean the centriuge locking mechanism. Adjust centrifuge cover height. Performed functionalith. Observed performance during actual UPP procedure. | Unit is operational. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6194665780026481252/6194665780026481252_signature_28.png | Vince Aurey Leal | https://www.jotform.com/uploads/macarewecare/222927978372067/6194665780026481252/6194665780026481252_signature_31.png
| 2025-04-03 | 2025-04-03 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30596 | DOH | sajecopooh20@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1327 |  | 6C GCIV-4D | 838555 | Stuck cartridge module B4 | #NAME? | #NAME? | Module B4 for retry XpertCheck. Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6194577366867441958/6194577366867441958_signature_28.png | Sarah Jean G. Pasana | https://www.jotform.com/uploads/macarewecare/222927978372067/6194577366867441958/6194577366867441958_signature_31.png
| 2025-04-02 | 2025-04-02 | Tomas Claudio Health Center | Camus St. Corner E. Quirino, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30595 | DOH | arleen_p_alconera@yahoo.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 782 |  | GX-IV 6C 4D | 838492 | For Preventive Maintenance and New Printer Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Download and Install new Printer provided by the laboratory | Perform Preventive Maintenance and Printer Installation | All Modules are operational, new printer installed | New Printer model Epson L3150 tank printer with WIFI |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6193849374118579560/6193849374118579560_signature_28.png | Arleen Alconera,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6193849374118579560/6193849374118579560_signature_31.png
| 2025-04-02 | 2025-04-02 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30594 | DOH | luzvimindanavidad2@gmail.com | 28 |  |  | 6C GXIV-4D | 838297 | Due for PM | Perform preventive maintenance: fan filter, cartridge bay, plunger rod, optic slot | Preventive Maintenance | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6193824215252717398/6193824215252717398_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6193824215252717398/6193824215252717398_signature_31.png
| 2025-04-02 | 2025-04-02 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30593 | UNOPS | bmcrhivda@gmail.com | 25 |  |  | GXIV-10C | 110023507 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6193804616121408402/6193804616121408402_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6193804616121408402/6193804616121408402_signature_31.png
| 2025-04-02 | 2025-04-02 | Diit District Health Center | Diit,Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30592 | DOH | dalepenalosa189@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1182 | 2025-04-02 | GXIV | 838421 | N/a | N/a | Monthly Preventive Maintenance | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6193799143719510830/6193799143719510830_signature_28.png | Dale D. Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6193799143719510830/6193799143719510830_signature_31.png
| 2025-04-01 | 2025-03-27 | West Dadiangas Rural Health Unit | Brgy. Dadiangas Rural Health Unit | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30590 | DOH | celinejean16@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1609 |  | GxIV-6C | 838570 | None | Perform Genexpert module cleaning Optics and plunger Rod cleaning Fan filter cleaning Checking of errors Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6193248016865760637/6193248016865760637_signature_28.png | Celine villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6193248016865760637/6193248016865760637_signature_31.png
| 2025-04-01 | 2025-03-25 | Norala District Hospital | Norala, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30591 | DOH | noralaDHlaboratory@gmail.com | 21.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1105 |  | GxIV-6C | 838543 | None | Module Cleaning Optics and Plunger rod Cleaning Filter Cleaning Checking of errors Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6193248016866786511/6193248016866786511_signature_28.png | Geraldine Magan | https://www.jotform.com/uploads/macarewecare/222927978372067/6193248016866786511/6193248016866786511_signature_31.png
| 2025-04-01 | 2025-03-26 | Cagayan de Oro City CHO | Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30589 | DOH | chocdotbdots@gmail.com | 27 |  |  | GX4-6 colors | 838441 | UPS not functioning | Reporting to Main Office | Preventive Maintenance and Troubleshooting | Units are functional except for UPS |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6193215896577238196/6193215896577238196_signature_28.png | Charissa Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/6193215896577238196/6193215896577238196_signature_31.png
| 2025-04-01 | 2025-03-26 | Philippine Red Cross Cagayan de Oro | Cagayan de Oro City | 10 | Misamis Oriental | Blood Center | NO | TROUBLESHOOTING | HAEMONETICS | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30588 | COMMERCIAL | mrbc2@redcross.com.ph | 26 |  |  | MCS+ | 02H080 | None | None | Readjustment and Recalibrating of Bowl Optics | All units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6193212296574570981/6193212296574570981_signature_28.png | Christian Gabriel Suaybaguio | https://www.jotform.com/uploads/macarewecare/222927978372067/6193212296574570981/6193212296574570981_signature_31.png
| 2025-04-01 | 2025-04-01 | Cardinal Santos Medical Center | Wilson St., GreenHills, San Juan City | NCR | NCR 2nd - San Juan | Blood Bank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 30587 | COMMERCIAL | robileanching@gmail.com | 28 |  | 2025-04-01 | MCS+ | 08M461 | Unit unable to power On. | Replaced fuse. Run functionality and sumilation. Monitor actual TPE procedure. | Found busted fuse. Replaced fuse. | Unit is operational. Recommended to use UPS for the unit to have better power source and with back up power in case of power outage. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6192944050112265683/6192944050112265683_signature_28.png | Robi-Lean P Ching | https://www.jotform.com/uploads/macarewecare/222927978372067/6192944050112265683/6192944050112265683_signature_31.png
| 2025-04-01 | 2025-03-27 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco, Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30586 | MACARE-HBDC | peachmango318@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1079 |  | GX-IV-D-10C | 110024245 | Error 2014 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B4 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6192908615093814714/6192908615093814714_signature_28.png | Arlen Delgado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6192908615093814714/6192908615093814714_signature_31.png
| 2025-03-31 | 2025-03-26 | Tuburan rural health unit | Tuburan Basilan | BARMM | Basilan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30585 | DOH | rjakilan2@gmail.com | 25 |  | 2025-03-26 | GX 4 | 838832 | Genexpert application unable to open. Error regarding SQL. | Undergone troubleshooting, uninstalled all SQL related system and uninstalled GXDX software. Pulled out System unit only for further repair. | Undergone troubleshooting, uninstalled all SQL related system and uninstalled GXDX software. | Under repair/ troubleshooting. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6192146475552439093/6192146475552439093_signature_28.png | Raiyana H. Jakilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6192146475552439093/6192146475552439093_signature_31.png
| 2025-03-31 | 2025-03-31 | Cong Semion Toribio Memorial Hospital | Carmen, Bohol | 7 | Bohol | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30584 | DOH | csgtmhlaboratory@gmail.com | 27 |  |  | Gx4 | 838394 | Defective Monitor | A video of power up monitor to automatic shutdown of monitor Copy of IQR | Troubleshooting monitor | Defective Monitor |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6192036626169425828/6192036626169425828_signature_28.png | Himaya Vale | https://www.jotform.com/uploads/macarewecare/222927978372067/6192036626169425828/6192036626169425828_signature_31.png
| 2025-03-31 | 2025-03-31 | Catigbian District Hospital | Catigbian, Bohol | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30583 | DOH | josephraytan@gmail.com | 26 |  |  | Gx4 | 838494 | Defective module B3 | New module Installed B3 862186 Clean Fan Clean Fan Filter IQR, System log and assay stats | New module Installed B3 862186 Clean Fan Clean Fan Filter IQR, System log and assay stats | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191988779426350987/6191988779426350987_signature_28.png | Joseph Ray F. Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/6191988779426350987/6191988779426350987_signature_31.png
| 2025-03-31 | 2025-03-28 | Naguilian District Hospital | Naguilian, La Union | 1 | La Union | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30582 | DOH | marjelynrayray01@gmail.com | 25 |  |  | GXIV-4-D-6C | 838426 | UPS won't turn on | #NAME? | Performed Troubleshooting | Unit is Not Operational Defective UPS |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191978985715800091/6191978985715800091_signature_28.png | Marjelyn Rayray | https://www.jotform.com/uploads/macarewecare/222927978372067/6191978985715800091/6191978985715800091_signature_31.png
| 2025-03-31 | 2025-03-28 | Ilocos Training and Regional Medical Center | San Fernando City,La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30581 | MACARE-HBDC | itrmc.mbl@gmail.com | 26 |  |  | GXIV-4-D-6C | 120003504 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191977505711620026/6191977505711620026_signature_28.png | Dominga N. PATACSIL | https://www.jotform.com/uploads/macarewecare/222927978372067/6191977505711620026/6191977505711620026_signature_31.png
| 2025-03-31 | 2025-03-26 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30580 | DOH | dianeruiz147@gmail.com | 30 |  |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191975895711967679/6191975895711967679_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6191975895711967679/6191975895711967679_signature_31.png
| 2025-03-31 | 2025-03-26 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30579 | COMMERCIAL | annabellelcatbagan@gmail.com | 28 |  |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191972835714190834/6191972835714190834_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6191972835714190834/6191972835714190834_signature_31.png
| 2025-03-31 | 2025-03-26 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30578 | DOH | serly_reyes@yahoo.com | 27 |  |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191971855712841021/6191971855712841021_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6191971855712841021/6191971855712841021_signature_31.png
| 2025-03-31 | 2025-03-21 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30577 | DOH | rhupozorrubio@yahoo.com | 30 |  |  | GXIV-4-D-6C | 838467 | For Troubleshooting | #NAME? | Performed Troubleshooting | - Unit is Partially Operational - Module 1 for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191970235714675490/6191970235714675490_signature_28.png | Jinky Grace Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/6191970235714675490/6191970235714675490_signature_31.png
| 2025-03-31 | 2025-03-19 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30576 | DOH | seansimonpaloyo@gmail.com | 25 |  | 2025-03-19 | GXIV-4-D-6C | 838473 | For Troubleshooting and XpertCheck | #NAME? | Performed Troubleshooting and XpertCheck | Unit is Fully Operational |  | XpertCheck Kits | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191968725711153242/6191968725711153242_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6191968725711153242/6191968725711153242_signature_31.png
| 2025-03-31 | 2025-03-18 | Pozorrubio Rural Health Unit | Pozorrubio | 1 | Pangasinan | Laboratory | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30575 | DOH | rhupozorrubio@yahoo.com | 30 |  |  | GXIV-4-D-6C | 838467 | Stuck Cartridge in Module 1 | #NAME? | Removed Stuck Cartridge in Module 1 | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191967145713718632/6191967145713718632_signature_28.png | Jinky Grace Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/6191967145713718632/6191967145713718632_signature_31.png
| 2025-03-31 | 2025-03-18 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30574 | DOH | allan.vinluan2013@gmail.com | 27 |  | 2025-03-14 | GXIV-4-D-6C | 838472 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 862037 >>> 888410 - Generated Reports | Performed Module Replacement | Unit is Fully Operational | Work with Engr. Venzon | GeneXpert Module 888410 | DR 12849 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191965385712547980/6191965385712547980_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6191965385712547980/6191965385712547980_signature_31.png
| 2025-03-31 | 2025-03-27 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30573 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  | 2025-03-26 | MCS+ | 16K521SPG | For installation of new display cable. Unit for preventive maintenance. | Repair. | Performed repair. Installed new display cable. Performed stress test. Performed simulation running. Verified calibration parameters. | Unit is now operational. | Work with Engr. Bagbaguen | Item Code: 112498-02 Display Cable | DR#12881 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191951053618095221/6191951053618095221_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6191951053618095221/6191951053618095221_signature_31.png
| 2025-03-31 | 2025-03-28 | MMC Medilinx | Amorosolo st Makati | NCR | NCR 4th - Makati | Molecular | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30572 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 26 |  |  | GX16D 10C | 110026448 | For installation | Installed unit Upload adf files Test run unit | Performed installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6191886183616697682/6191886183616697682_signature_28.png | Aiza G. Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6191886183616697682/6191886183616697682_signature_31.png
| 2025-03-28 | 2025-03-28 | DUMARAO RURAL HEALTH UNIT | Dumarao, Capiz | 6 | Capiz | TB Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30571 | UNOPS | doregoshielamarie@gmail.com | 23 |  |  | GX-IV 10C | 110023542 | #NAME? | #NAME? | -PERFORM TROUBLESHOOTING PROCEDURE FOR NO POWER -CONFIRMED SYSTEM UNIT SN: CPH1023F 46914 IS DEFECTIVE | FOR REPORTING TO CEPHEID |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6189589984467264262/6189589984467264262_signature_28.png | Shiela Marie Dorego | https://www.jotform.com/uploads/macarewecare/222927978372067/6189589984467264262/6189589984467264262_signature_31.png
| 2025-03-28 | 2025-03-28 | ILOILO PROVINCIAL HOSPITAL | Pototan Iloilo | 6 | Iloilo | TB DOTS | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30570 | DOH | dipragale92@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 960 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV | 838289 | Module C3 Error 2014 For Preventive Maintenance | #NAME? | #NAME? | Equipment is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6189584694466900718/6189584694466900718_signature_28.png | Rachel A. Pragale | https://www.jotform.com/uploads/macarewecare/222927978372067/6189584694466900718/6189584694466900718_signature_31.png
| 2025-03-28 | 2025-03-28 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30569 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 08B049 | DPM error | NA | #NAME? | Unit is operational No hardware problems found | With Sir Harold |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6189536876124532200/6189536876124532200_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6189536876124532200/6189536876124532200_signature_31.png
| 2025-03-28 | 2025-03-27 | Ragay Rural Health Unit | Poblacion, Ragay | 5 | Camarines Sur | Laboratory | YES | REPAIR | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30568 | DOH | virginramirez7@gmail.com | 26 |  |  | GX4-6C | 838560 | UPS no backup power | Replaced UPS - AT1000190123ST0327 Cleaned cartridge bays, plunger rod Archived 2023, 2024 results Collect Data | Replaced UPS Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6189468821385264832/6189468821385264832_signature_28.png | Virgin ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/6189468821385264832/6189468821385264832_signature_31.png
| 2025-03-28 | 2025-03-28 | Panabo City Health Office | Poblacion, Panabo | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30567 | DOH | nadineeust67@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2381 |  | GX-IV 6C 4D | 838295 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6189466234188642509/6189466234188642509_signature_28.png | Nadine Gemma Eustaquio,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6189466234188642509/6189466234188642509_signature_31.png
| 2025-03-27 | 2025-03-27 | Hinaplanon Health Center | Brgy. Hinaplanon, Iligan City | 10 | Lanao del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30566 | DOH | villacortaamy50@gmail.com | 25 |  |  | GX-4 6 Colors | 838227 | Module B1 found temperature-related error and was reported | Module Replacement | Replacing Module B1 with SN: 765568 to SN: 725419 | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6188666920323655983/6188666920323655983_signature_28.png | Amy S. Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6188666920323655983/6188666920323655983_signature_31.png
| 2025-03-27 | 2025-03-27 | Cebu TB reference lab-SRP | South Road Property-Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30565 | DOH | asmantiquilla.pbsp@gmail.com | 25 |  |  | Gx 4 | 838263 | Defective Module B4 220004971 | Install new module B4 788903 Clean Fan Filter Gather data, IQR, system log and assah stats | Install new module B4 788903 Clean Fan Filter Gather data, IQR, system log and assah stats | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6188655963205522135/6188655963205522135_signature_28.png | Allan S.Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6188655963205522135/6188655963205522135_signature_31.png
| 2025-03-27 | 2025-03-27 | WESTERN VISAYAS STATE UNIVERSITY MEDICAL CENTER | Iloilo City | 6 | Iloilo | Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30564 | DOH | mc-molecular@wvsu.edu.ph | 28 |  |  | GX-IV | 838290 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6188626172813740698/6188626172813740698_signature_28.png | Karla Danielle Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6188626172813740698/6188626172813740698_signature_31.png
| 2025-03-27 | 2025-03-27 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | UPS | 30563 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  | 2025-01-27 | ACCUPOWER | AT1000190123ST0384 | For replacement of 2pcs battery | Replaced and installed 2 pcs battery s/n 540421,540420 test run unit | Repair | operational |  |  | , |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6188617313618588133/6188617313618588133_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6188617313618588133/6188617313618588133_signature_31.png
| 2025-03-27 | 2025-03-27 | WESTERN VISAYAS STATE UNIVERSITY MEDICAL CENTER | Iloilo City | 6 | Iloilo | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30562 | UNOPS | mc-molecular@wvsu.edu.ph | 22 | ASSAY: Xpert HBV VL, COUNT: 34 ASSAY: Xpert HPV, COUNT: 17 ASSAY: Xpert MTB-XDR, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 35 |  | GX-IV | 110021716 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6188616302813063034/6188616302813063034_signature_28.png | Karla Danielle T. Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6188616302813063034/6188616302813063034_signature_31.png
| 2025-03-27 | 2025-03-27 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | UPS | 30561 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  | 2025-01-27 | ACCUPOWER | AT1000190123ST0202 | For replacement of 2pcs battery | Replaced and installed 2pcs battery s/n 540423,540422 test run unit | Repair | Operational |  | 12V 7ah battery | , |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6188614543614248906/6188614543614248906_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6188614543614248906/6188614543614248906_signature_31.png
| 2025-03-27 | 2025-03-27 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | UPS | 30560 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  | 2025-01-27 | ACCUPOWER | AT1000180910ST0259 | For replacement of 2 pcs battery | Replaced and installed 2 pcs battery s/n 540425,540424 | Repair | Operational |  | 12v 7ah battery | , |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6188612353611176139/6188612353611176139_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6188612353611176139/6188612353611176139_signature_31.png
| 2025-03-27 | 2025-03-27 | Calubian Rural Health Unit | Calubian, Leyte | 8 | Leyte | TB Dota | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30559 | DOH | mjcabardo06@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 820 |  | 6C GXIV-4D | 838399 | Due for PM | -Perform preventive maintenance; cartridge bay, fan filter, plunger rod, and optic slot | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6188534360532653533/6188534360532653533_signature_28.png | Nicholas T. Fiji | https://www.jotform.com/uploads/macarewecare/222927978372067/6188534360532653533/6188534360532653533_signature_31.png
| 2025-03-27 | 2025-03-26 | MMC Medilinx | Amorsol st Makati City | NCR | NCR 4th - Makati | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30558 | COMMERCIAL | loiszoemontemayor.medilinx@outlook.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 187 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX4- 2 D | 807128 | Monitor encounter auto shutdown | Performed monitor installation spare monitor from medilinx | Installation | Unit operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6188460293617577612/6188460293617577612_signature_28.png | Lois Zoe Montemayor | https://www.jotform.com/uploads/macarewecare/222927978372067/6188460293617577612/6188460293617577612_signature_31.png
| 2025-03-26 | 2025-03-26 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30557 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  | 2025-03-24 | GXIV | 807112 | Module C2, C3 for troubleshooting | Performed blank cartridge test on C2, C3 | #NAME? | Unit is operational Module C2: Good Module C3: Error 5007 (for follow up report) |  | MTB/RIF Ultra Cartridge | 12872 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6187766585217014901/6187766585217014901_signature_28.png | Raecel | https://www.jotform.com/uploads/macarewecare/222927978372067/6187766585217014901/6187766585217014901_signature_31.png
| 2025-03-26 | 2025-03-26 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30556 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  | 2025-03-24 | GXIV | 807957 | #NAME? | #NAME? | - Installed barcode scanner - Checked monitor (Unit shuts down automatically upon powering up: "Power Button Lockout" error) - Performed HP monitor troubleshooting - Swapped monitor with another monitor (Error still persists) - Temporarily installed another monitor to proceed - Performed blank cartridge test on D2, D3 - Generated reports - Performed self test - Checked and cleaned cartridge bay, plunger rod, filter, icore slot | Unit is operational Module D1: Awaiting for module allocation Module D2, D3: Error 5007 (for follow up report) HP P19A Monitor (SN: CNK831078T) is recommended for repair/replacement |  | MTB/RIF Ultra Cartridge GX Barcode Scanner | 12872 & 12841 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6187765625219450074/6187765625219450074_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6187765625219450074/6187765625219450074_signature_31.png
| 2025-03-25 | 2025-03-25 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30555 | MACARE-HBDC | jmoreno@ncf.edu.ph | 25 |  |  | MCS+ | 15H031G | DPM connector test failed | Checked DPM parameter(ok) Cleaned DPM surface Checked for any liquid inside DPM(None) Performed restart on machine for different protocol cards | Performed Troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6187014433414657880/6187014433414657880_signature_28.png | Jake S. Moreno | https://www.jotform.com/uploads/macarewecare/222927978372067/6187014433414657880/6187014433414657880_signature_31.png
| 2025-03-25 | 2025-03-25 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30554 | MACARE-HBDC | jmoreno@ncf.edu.ph | 25 |  |  | MCS+ | 15H027G | Bios CRC failed | Performed restart using different protocol cards Attempt to bypass error on program | Performed Troubleshooting | Unit is not operational | Battery of CRC on motherboard for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6187013103411443796/6187013103411443796_signature_28.png | Jake S. Moreno | https://www.jotform.com/uploads/macarewecare/222927978372067/6187013103411443796/6187013103411443796_signature_31.png
| 2025-03-25 | 2025-03-25 | Philippine Red Cross | Intramuros, Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30553 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 16K521SPG | Unable to push button and no button light. | Repair. | Performed troubleshooting. Reset and fixed display ribbon cable. | Buttons functionality is intermittent. For replacement of ribbon cable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6187002782126001477/6187002782126001477_signature_28.png | Lalaine ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6187002782126001477/6187002782126001477_signature_31.png
| 2025-03-25 | 2025-03-25 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30552 | MACARE-HBDC | bloodcollection.tower@gmail.com | 25 |  |  | MCS+ | 08B049 | Noise on centrifuge during procedure | Tightened centrifuge cover screws | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6186931145212555259/6186931145212555259_signature_28.png | Maybelle Monzon | https://www.jotform.com/uploads/macarewecare/222927978372067/6186931145212555259/6186931145212555259_signature_31.png
| 2025-03-25 | 2025-03-25 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30551 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2324 |  | 6C GXIV-4D | 838285 | Due for monthly PM | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6186882493687982819/6186882493687982819_signature_28.png | Leak Karen Noynay | https://www.jotform.com/uploads/macarewecare/222927978372067/6186882493687982819/6186882493687982819_signature_31.png
| 2025-03-25 | 2025-03-25 | Hi Precision Diagnostics | N. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30550 | CONSORTIUM | jose@hi-precision.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 163 ASSAY: Xpert HPV, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4832 |  | GXIV | 835601 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6186882246814852227/6186882246814852227_signature_28.png | Joselito T. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/6186882246814852227/6186882246814852227_signature_31.png
| 2025-03-25 | 2025-03-25 | Hi Precision Diagnostics | N. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30549 | CONSORTIUM | jose@hi-precision.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 124 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5782 |  | GXIV | 835602 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6186882006818126935/6186882006818126935_signature_28.png | Joselito T. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/6186882006818126935/6186882006818126935_signature_31.png
| 2025-03-25 | 2025-03-25 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TN Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 30548 | DOH | johnpaullabutap@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 | 2025-03-25 | GXIV | 838551 | Module B2 multiple temperature error | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6186831700128814860/6186831700128814860_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6186831700128814860/6186831700128814860_signature_31.png
| 2025-03-25 | 2024-12-02 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30547 | COMMERCIAL | vsmmcbloodservices@gmail.com | 24 |  |  | MCS+ | 02A090 | -Due for calibration of MCS+ | Performed Preventive in MCS+ Conduct calibration in MCS+ weigher, cuff pressure, DPM, SPM pressure, Line sensor and bowl optics sensor, Check pneumatic assembly, valve test, motor pump | Performed Preventive in MCS+ Conduct calibration in MCS+ weigher, cuff pressure, DPM, SPM pressure, Line sensor and bowl optics sensor, Check pneumatic assembly, valve test, motor pump | Machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6186772783448557825/6186772783448557825_signature_28.png | Vicente Sotto Memorial Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6186772783448557825/6186772783448557825_signature_31.png
| 2025-03-25 | 2025-03-20 | Virac Municipal Health Office | Virac | 5 | Catanduanes | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30546 | DOH | limarmariano50@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1533 |  | GX4-6C | 838563 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6186762564521822815/6186762564521822815_signature_28.png | Ramil A. Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/6186762564521822815/6186762564521822815_signature_31.png
| 2025-03-25 | 2025-03-21 | Rosario Super Health Center | Dr. Sixto Antonio Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30545 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1187 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6186715143613409476/6186715143613409476_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6186715143613409476/6186715143613409476_signature_31.png
| 2025-03-25 | 2025-03-21 | Pasig Treatment Hub - Manggahan | Manggahan, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30544 | MACARE-HBDC | pathmedtech@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GX-IV-D-6C | 821648 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -filter Generated reports | Unit is operational. Modules A2and A4 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6186715143616669764/6186715143616669764_signature_28.png | Katherine Godoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6186715143616669764/6186715143616669764_signature_31.png
| 2025-03-24 | 2025-03-21 | MMC (Medilinx) | Amorsolo st., Makati | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30543 | COMMERCIAL | janeanequirante.medilinx@outlook.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 131 |  | GX4-2 D | 807128 | As per end user they encounter flickering display A2 module encounter intermittent temp error. | Re plug in vga cable Gathered data Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed troubleshooting | During testing of monitor no flickering display encourtered Monitor for monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6185853303617645625/6185853303617645625_signature_28.png | Janeane Quirante | https://www.jotform.com/uploads/macarewecare/222927978372067/6185853303617645625/6185853303617645625_signature_31.png
| 2025-03-22 | 2025-03-22 | Philippine Red Cross Port | Manila | NCR | NCR 1st - Manila | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30542 | MACARE-HBDC | dolly.banares@redcross.org.ph | 28 |  |  | MCS+ | 16K521SPG | Display panel keys not responding | Reset ribbon | #NAME? | Unit is operational Unit is under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6184301847118756883/6184301847118756883_signature_28.png | Dolly banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6184301847118756883/6184301847118756883_signature_31.png
| 2025-03-21 | 2025-03-20 | BANTAYAN RURAL HEALTH UNIT | Bantayan, Cebu | 7 | Cebu | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30541 | DOH | mcdelle916bausin@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 932 | 2025-02-18 | GX-IV | 838509 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | LCD Monitor - 19" HP E190i IPS | 12757 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6183283965119738288/6183283965119738288_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/6183283965119738288/6183283965119738288_signature_31.png
| 2025-03-20 | 2025-03-18 | Marikina City Health Office | Shoe Ave Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30537 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 569 ASSAY: Xpert MTB-XDR, COUNT: 46 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143611885277/6182696143611885277_signature_28.png | Ronica Ana Tadtad | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143611885277/6182696143611885277_signature_31.png
| 2025-03-20 | 2025-03-18 | Marikina City Health Office | Shoe Ave Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30536 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 91 ASSAY: Xpert MTB-XDR, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 100 |  | GX-IV-D-10C | 110024073 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143615782421/6182696143615782421_signature_28.png | Ronica Ana Tadtad | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143615782421/6182696143615782421_signature_31.png
| 2025-03-20 | 2025-03-20 | Antipolo City Health Office | M. Santos St Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30539 | UNOPS | acho.laboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert MTB-XDR, COUNT: 40 ASSAY: Xpert HIV-1 Viral Load, COUNT: 363 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143617974817/6182696143617974817_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143617974817/6182696143617974817_signature_31.png
| 2025-03-20 | 2025-03-18 | Marikina City Health Office | Shoe Ave Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30540 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1097 ASSAY: Xpert MTB-XDR, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143618371967/6182696143618371967_signature_28.png | Ronica Ana Tadtad | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143618371967/6182696143618371967_signature_31.png
| 2025-03-20 | 2025-03-18 | Marikina City Health Office | Shoe Ave Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30538 | MACARE-HBDC | mcho.laboratory102018@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 106 ASSAY: Xpert HIV-1 Viral Load, COUNT: 85 |  | GX-IV-D-10C | 110024072 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143619914623/6182696143619914623_signature_28.png | Ronica Ana Tadtad | https://www.jotform.com/uploads/macarewecare/222927978372067/6182696143619914623/6182696143619914623_signature_31.png
| 2025-03-20 | 2025-03-20 | Dasol Community Hospital | Dasol Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30535 | DOH | jeremypulido0521@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 588 | 2025-03-14 | GXIV-D | 838287 | B4: 808252 for module replacement B3: 808251 temperature error | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational B4: 661031 new module |  | 1 Genexpert module | 12850 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182609080419494989/6182609080419494989_signature_28.png | Jeremy pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/6182609080419494989/6182609080419494989_signature_31.png
| 2025-03-20 | 2025-03-20 | Dasol Community Hospital | Dasol Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30534 | DOH | jeremypulido0521@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 588 | 2025-03-14 | GXIV-D | 838287 | B4: 808252 for module replacement B3: 808251 temperature error | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational B4: 661031 new module |  | 1 Genexpert module | 12850 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182608940419792537/6182608940419792537_signature_28.png | Jeremy pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/6182608940419792537/6182608940419792537_signature_31.png
| 2025-03-20 | 2025-03-20 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30533 | MACARE-HBDC | uermbloodbank@gmail.com | 28 |  | 2025-03-17 | MCS+ | 13C270 | Right IV Pole for replacement | Replaced IV Pole | #NAME? | Unit is operational |  | IV Pole | 12861 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182607233116311424/6182607233116311424_signature_28.png | Leonila Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6182607233116311424/6182607233116311424_signature_31.png
| 2025-03-20 | 2025-03-20 | Malay Primary Care Facility | Caticlan, Malay, Aklan | 6 | Aklan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30532 | DOH | mhomalayclinicallaboratory@gmail.com | 26 |  | 2025-02-27 | 6C GXIV-4D | 838457 | #NAME? | -Replace defective module B3/890553 with module SN:210112052 -Reset CPU connections | #NAME? | Machine is functional, CPU is operational. |  | 900-0491RS Certified pre-owned GeneXpert module SN. 210112052 | DR no. 12790 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182593793524513664/6182593793524513664_signature_28.png | May Soledad | https://www.jotform.com/uploads/macarewecare/222927978372067/6182593793524513664/6182593793524513664_signature_31.png
| 2025-03-20 | 2025-03-19 | Sta Rosa Community District Hospital | Sta Rosa, Olango, Lapulapu City | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30531 | DOH | ariestumz@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 300 |  | GX4 | 838534 | Due for Preventive Maintenance | Performed Preventive in all modules Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. | Performed Preventive in all modules Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182591453448387712/6182591453448387712_signature_28.png | Ariston Tumulak, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6182591453448387712/6182591453448387712_signature_31.png
| 2025-03-20 | 2025-03-20 | Labuan General hospital | Labuan, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30530 | DOH | jenne060213@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2184 | 2025-03-20 | GX IV | 838568 | Module 2 gives an error of Error 4012 multiple times. | Module 2 undergone troubleshooting and Xpertcheck calibration. | Module 2 undergone troubleshooting and Xpertcheck calibration. | Module 1, 3 and 4 are all available and ready to use. While module 2 failed the said calibration and recommended for replacement. | Module 2 for replacement. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182565689021483226/6182565689021483226_signature_28.png | Jennerose Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6182565689021483226/6182565689021483226_signature_31.png
| 2025-03-20 | 2025-03-20 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30529 | DOH | buhanginlaboratory@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1320 |  | GX-IV 6C 4D | 838232 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182557764224522724/6182557764224522724_signature_28.png | Catherine B. Deocampo ,.RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6182557764224522724/6182557764224522724_signature_31.png
| 2025-03-20 | 2025-02-20 | National Center for Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30528 | COMMERCIAL | iancristrono@gmail.com | 25 |  |  | GXIV-10C | 120003480 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182536073111227172/6182536073111227172_signature_28.png | Ian Cris Tono | https://www.jotform.com/uploads/macarewecare/222927978372067/6182536073111227172/6182536073111227172_signature_31.png
| 2025-03-20 | 2025-03-20 | Pinamalayan Municipal Health Office | Pinamalayan, Oriental Mindoro | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30527 | DOH | medtechdave22@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1067 | 2025-03-17 | GXIV | 838230 | Unit for preventive maintenance. For module replacement of module D4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | New module (635717) in replacement to defective module (807253) | DR#12853 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182436412983518607/6182436412983518607_signature_28.png | Dave Erwin M. Festin | https://www.jotform.com/uploads/macarewecare/222927978372067/6182436412983518607/6182436412983518607_signature_31.png
| 2025-03-20 | 2025-03-20 | Pinamalayan Municipal Health Office | Pinamalayan, Oriental Mindoro | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30526 | DOH | medtechdave22@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1067 | 2025-03-17 | GXIV | 838230 | Unit for preventive maintenance. For module replacement of module D4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | New module (635717) in replacement to defective module (807253) | DR#12853 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182436402982888471/6182436402982888471_signature_28.png | Dave Erwin M. Festin | https://www.jotform.com/uploads/macarewecare/222927978372067/6182436402982888471/6182436402982888471_signature_31.png
| 2025-03-20 | 2025-03-20 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | Accupower | 30525 | MACARE-HBDC | nfloresca@macare-medicals.com | 28 |  |  | 1kva UPS | 190123ST0384,190123ST0202,180910ST0259 | No power Battery low voltage 4volts | Checked unit Checked wear tear Checked and test voltage output of battery | Performed troubleshooting | For replacement of ups battery 6pcs for 3 unit of ups. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182420883616638326/6182420883616638326_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6182420883616638326/6182420883616638326_signature_31.png
| 2025-03-20 | 2025-03-19 | Best Diagnostic Corp. | Masikap ext.,QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30524 | COMMERCIAL | bestdiagnosticlaboratory@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 28 ASSAY: Xpert HBV VL, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert HCV VL, COUNT: 3 |  | GX4 D | 110008529 | As per end user they encounter 2096 and 2097 error | Performed sample running hbv vl Gathered data. | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6182383493616212596/6182383493616212596_signature_28.png | Best Diagnostic Corp. | https://www.jotform.com/uploads/macarewecare/222927978372067/6182383493616212596/6182383493616212596_signature_31.png
| 2025-03-19 | 2025-03-14 | Tanza Wellness Center | A. Soriano Highway, Tanza, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30523 | MACARE-HBDC | tanzawellnesscenter.rhu@gmail.com | 25 | ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 51 |  | GXIV-D-10C | 120003582 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6181818313615164090/6181818313615164090_signature_28.png | Jocelyn Cadaño | https://www.jotform.com/uploads/macarewecare/222927978372067/6181818313615164090/6181818313615164090_signature_31.png
| 2025-03-19 | 2025-03-19 | Likhaan Center for Women's Health | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30522 | COMMERCIAL | clatras@likhaan.org | 28 |  |  | GXII-10C | 110025174 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6181729223111125612/6181729223111125612_signature_28.png | Jocelyn Latras | https://www.jotform.com/uploads/macarewecare/222927978372067/6181729223111125612/6181729223111125612_signature_31.png
| 2025-03-19 | 2025-03-18 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30521 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1600 ASSAY: Xpert CT/NG, COUNT: 18 ASSAY: Others, COUNT: 116 ASSAY: Xpert HBV VL, COUNT: 1002 ASSAY: Xpert HPV, COUNT: 149 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2546 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 3835 ASSAY: Xpert HCV VL, COUNT: 167 | 2025-03-11 | GX16D | 110005587 | For module replacement c4 210059476 - 616251 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed module replacement | A1,A2,C2,D2 reported already for replacement A3,B3 failed xpertcheck last 9/30/2024. reported already for replacement |  | Genexpert 6 color module | 12827 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6181559723618812282/6181559723618812282_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6181559723618812282/6181559723618812282_signature_31.png
| 2025-03-18 | 2025-03-18 | Rapu-Rapu Rural Health Unit | Poblacion, Rapu-Rapu | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30520 | DOH | suzetteoguan1@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX4-6C | 838565 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data Reported B2/730958(H/W failed) Reported monitor (grid lines appear on screen) | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6180971083328492280/6180971083328492280_signature_28.png | Suzette Oguan | https://www.jotform.com/uploads/macarewecare/222927978372067/6180971083328492280/6180971083328492280_signature_31.png
| 2025-03-18 | 2025-03-18 | Nueva Vizcaya Provincial Hospital | Bambang, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30519 | DOH | dumlaojoy4@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 950 | 2025-03-04 | GXIV-6C | 838272 | Monitor Defective | Performed Monitor Replacement | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Monitor | 12806 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6180896965017435761/6180896965017435761_signature_28.png | Bernadette Dumlao | https://www.jotform.com/uploads/macarewecare/222927978372067/6180896965017435761/6180896965017435761_signature_31.png
| 2025-03-18 | 2025-03-18 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30518 | COMMERCIAL | mahmci_labcmt@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 21 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 282 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 |  | GX-IV-D-6C | 110005265 | Error on module B4. UPS no backup. | Checked UPA and connections. | Check UPS backup power. Check Connections Run reliability test Performed maintenance Generated reports Generated gxx file | Unit is operational. UPS has no backup power. | Recommended to replace UPS. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6180805123615545157/6180805123615545157_signature_28.png | Graziela Marie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/6180805123615545157/6180805123615545157_signature_31.png
| 2025-03-18 | 2025-03-18 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | daguilon.intp@gmail.com |  | 30517 | MACARE-HBDC | msff-tondo-labtech@paris.msf.org | 25 |  | 2025-03-17 | GX16-10C | 110016343 | Unit for preventive maintenance. For module replacement of module B3. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module (210113012) in replacement to defective (210081555). | DR#12859 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6180788250118392613/6180788250118392613_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6180788250118392613/6180788250118392613_signature_31.png
| 2025-03-18 | 2025-03-18 | Vicente Mendiola Center for Health | Naga City | 7 | Cebu | Tb/hiv laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 30516 | DOH | geccaprocianos@gmail.com | 23 |  | 2025-03-18 | Gx4 | 838564 | . Monthly preventive maintenance. . Error code while running all modules 2008 | Save Installation qualifications report. Save System log reports Save Assay statistics report. | Perform preventive maintenance | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6180787836225647700/6180787836225647700_signature_28.png | Gene Carmel T. Procianos | https://www.jotform.com/uploads/macarewecare/222927978372067/6180787836225647700/6180787836225647700_signature_31.png
| 2025-03-18 | 2025-03-18 | Ramon Rural Health Unit | Ramon, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30515 | DOH | rhuramonlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 520 |  | GXIV-6C | 838480 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6180725137325184190/6180725137325184190_signature_28.png | Janette Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/6180725137325184190/6180725137325184190_signature_31.png
| 2025-03-17 | 2025-03-13 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30514 | DOH | crmcmolecularlab@gmail.com | 22.6 | ASSAY: Xpert SARS-CoV-2, COUNT: 2994 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 31 ASSAY: Xpert SARS-CoV-2, COUNT: 2994 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 55 |  | GxIV-6C | 837662 | None | Perform Genexpert machine Cleaning and maintenance Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6180306306866558318/6180306306866558318_signature_28.png | Deanne Grace Cabaña | https://www.jotform.com/uploads/macarewecare/222927978372067/6180306306866558318/6180306306866558318_signature_31.png
| 2025-03-17 | 2025-03-17 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com | Accupower | 30513 | MACARE-HBDC | florescanoel77@gmail.com | 28 |  | 2025-03-12 | Accupower | AT1000190123ST0327 | For repair For replacement of 2 pcs 12v 7ah battery | Replaced and installed 2pcs battery Checked voltage output 240volts | Performed repair | Operational |  | 2pcs 12v,7ah battery s/n 533233,533234 | G7821 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6180022893611615664/6180022893611615664_signature_28.png | Macare in house | https://www.jotform.com/uploads/macarewecare/222927978372067/6180022893611615664/6180022893611615664_signature_31.png
| 2025-03-17 | 2025-03-17 | Macare In-House | Aurora Grand | NCR | NCR 1st - Manila | Engineering | NO | REPAIR | OTHERS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  | Cepheid Workstation (Lenovo) | 30512 | DOH | engineering@macare-medicals.com | 25 |  |  | Thinkcentre | CPH0520F08280 | Unable to power on. | Repair. | Performed workstation repair. Replaced mother board, CD-ROM and power supply taken from Lenovo Thinkcentre (SN: PC1132B9). Tested the unit with GeneXpert machine. | Workstation is now operational. For return to site. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6180022843615712541/6180022843615712541_signature_28.png | Macare Medicals, Inc. | https://www.jotform.com/uploads/macarewecare/222927978372067/6180022843615712541/6180022843615712541_signature_31.png
| 2025-03-17 | 2025-03-12 | Naga City General Hospital | Balatas, Naga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30511 | MACARE-HBDC | tin2_citydrug@yahoo.com | 24 |  |  | GX4-10C | 120003510 | Gx for Preventive Maintenance | Gx unit was stored for now and is not used yet due to no available kits. Checked the location where Gx is placed(no dust) | Checked Gx status | Unit is not operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179891156813482072/6179891156813482072_signature_28.png | Christine n. SANTA MARIA | https://www.jotform.com/uploads/macarewecare/222927978372067/6179891156813482072/6179891156813482072_signature_31.png
| 2025-03-17 | 2025-03-11 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | gonzagakevin4@gmail.com |  | 30510 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 | ASSAY: Others, COUNT: 3 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 247 ASSAY: Xpert SARS-CoV-2, COUNT: 1465 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-XVI-D-6C | 838627 | Modules not detected. | Transferred godo modules to bank A. | Checked network switch power supply Checked connections Checked cables Transferred good modules(701496, 667521 and 690650) to bank A | Network switch not operational. Four modules operational. (690650, 701496, 667521, 647085) | Defective internal network switch power supply. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179887573619456656/6179887573619456656_signature_28.png | Justine Lorenz V. Falzado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6179887573619456656/6179887573619456656_signature_31.png
| 2025-03-17 | 2025-03-12 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan City | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 30509 | COMMERCIAL | larozaericka25@gmail.com | 22 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179887423611634676/6179887423611634676_signature_28.png | Ericka Laroza | https://www.jotform.com/uploads/macarewecare/222927978372067/6179887423611634676/6179887423611634676_signature_31.png
| 2025-03-17 | 2025-03-12 | Cardinal Santos Medical Center | Wilson St, Greenhills, San Juan City | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 30508 | COMMERCIAL | gbuenaventurarmt@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 143 ASSAY: Xpert HPV, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1079 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179887253618058655/6179887253618058655_signature_28.png | Gabriella Buenaventura | https://www.jotform.com/uploads/macarewecare/222927978372067/6179887253618058655/6179887253618058655_signature_31.png
| 2025-03-17 | 2025-03-14 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | Demo | OTHERS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  | Limusave | 30507 | COMMERCIAL | yeye_delacruz@yahoo.com | 22 |  |  | MT-7500 | 37020508 | For demo/installation | Installed the Limusave System | Delivered and installed Limusave System: Limusave MT-7500 Thermostation Cooling station UPS Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179887103614134626/6179887103614134626_signature_28.png | Yehleen Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6179887103614134626/6179887103614134626_signature_31.png
| 2025-03-17 | 2025-03-13 | Kalinga Health Hub - Marikina | JP Rizal St. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30506 | CONSORTIUM | bryllcruz08@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1610 |  | GX-IV-D-6C | 110008418 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B2 under observation for error 5007. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179886493616703557/6179886493616703557_signature_28.png | Lorenzo Bryll Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6179886493616703557/6179886493616703557_signature_31.png
| 2025-03-17 | 2025-03-14 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30505 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 60 ASSAY: Xpert CT/NG, COUNT: 41 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 637 ASSAY: Xpert HPV, COUNT: 404 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6283 ASSAY: Xpert MTB-XDR, COUNT: 3 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 137 ASSAY: Xpert HIV-1 Viral Load, COUNT: 856 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1381 |  | GX-XVI-D-10C | 10015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger -filter Generated gxx file Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179886263611563511/6179886263611563511_signature_28.png | Kristine Parilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6179886263611563511/6179886263611563511_signature_31.png
| 2025-03-17 | 2025-03-14 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30504 | DOH | rhupozorrubio@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 885 |  | GXIV-4-D-6C | 838467 | Stuck Cartridges in Modules 1,2 and 3 | - Removed Stuck Cartridges in Modules 2 and 3 - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Generated Reports | Removed stuck cartridges | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179857065711910091/6179857065711910091_signature_28.png | Jinky Grace Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/6179857065711910091/6179857065711910091_signature_31.png
| 2025-03-17 | 2025-03-14 | MMC Medilinx | Amorsolo st., Makati | NCR | NCR 4th - Makati | Molecular | NO | INSTALLATION | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30503 | COMMERCIAL | janeanequirante.medilinx@outlook.com | 26 |  | 2025-03-13 | GX4-2D | 807128 | For standby and installation | Delivered and installed unit Upload ADF Test run unit | Performed installation | Operational |  |  | 12840 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179800843617308909/6179800843617308909_signature_28.png | Janeane Quirante | https://www.jotform.com/uploads/macarewecare/222927978372067/6179800843617308909/6179800843617308909_signature_31.png
| 2025-03-17 | 2025-03-14 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | Indiko plus | 30502 | MACARE-HBDC | janettequirino@yahoo.com | 26 |  | 2025-03-14 | 864 | 864000000213 | For replacement of timing belt and dispensing probe | Delivered and installed timing belt and dispensing probe. Performed functionality test Checked parameters Test run unit | Performed repair | Unit operational Ready to use |  | Timing belt Dispensing probe | 12842 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6179800443611107770/6179800443611107770_signature_28.png | Janette Quirino | https://www.jotform.com/uploads/macarewecare/222927978372067/6179800443611107770/6179800443611107770_signature_31.png
| 2025-03-15 | 2025-03-15 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30501 | COMMERCIAL | ireneclumicao09@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 188 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 52 | 2025-05-11 | GXII-D | 820303 | Module A2: 720228 for replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational A2 new Module: 220021213 |  | Genexpert module SN: 220021213 | 12822 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6178365403524378727/6178365403524378727_signature_28.png | Irene Lumicao | https://www.jotform.com/uploads/macarewecare/222927978372067/6178365403524378727/6178365403524378727_signature_31.png
| 2025-03-15 | 2025-03-14 | Minglanilla Rural Health Unit | Pider 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30500 | DOH | brendagonzagaparedes@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1450 |  | GX4 | 838512 | Preventive Maintenance monthly | Performed Preventive Maintenance, Clean ICore, Plunger, Filter, Cartridge Bay Check UPS and Printer, Bar Code Generate System log, and Archive database | Performed Preventive Maintenance, Clean ICore, Plunger, Filter, Cartridge Bay Check UPS and Printer, Bar Code Generate System log, and Archive database | All modules functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6178101070803377480/6178101070803377480_signature_28.png | Brenda Paredes, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6178101070803377480/6178101070803377480_signature_31.png
| 2025-03-14 | 2025-03-14 | Valenzuela Social Hygiene Clinic | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30499 | COMMERCIAL | markanthony.manaois@gmail.com | 25 |  |  | GXIV-10C | 110020699 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6177450633119759893/6177450633119759893_signature_28.png | Cristian Palteng | https://www.jotform.com/uploads/macarewecare/222927978372067/6177450633119759893/6177450633119759893_signature_31.png
| 2025-03-14 | 2025-03-14 | Bindoy District Hospital | National Hiway, Bindoy, Negros Oriental | 7 | Negros Oriental | Laboratory | NO | REPAIR | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30498 | DOH | bindoydistricthosp.laboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 301 |  | GX4 | 838475 | Defective UPS, No power backup during power interruptions. | Preventive maintenance - cleaned optics, plunger rods, module bays, filter; archived tests. | -Replaced batteries for UPS, with new set of batteries: 2 pcs 12 V 9AH. - Functional testing for UPS. - Replaced batteries for wireless keyboard and mouse. | 4 modules operational; UPS functional and with backup power during simulated power interruption. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6177422107654872663/6177422107654872663_signature_28.png | LEMUEL QUIRIT RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6177422107654872663/6177422107654872663_signature_31.png
| 2025-03-14 | 2025-03-14 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30497 | DOH | ccmclaboratory01@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1884 |  | 6C GXIV-4D | 838561 | Module B3 H/W failed | NA | #NAME? | All other modules are functional. Module B3 failed XpertCheck, for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6177391882598036932/6177391882598036932_signature_28.png | Hannah Jane Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/6177391882598036932/6177391882598036932_signature_31.png
| 2025-03-14 | 2025-03-14 | NUEVA VALENCIA RURAL HEALTH UNIT | NUEVA VALENCIA, GUIMARAS | 6 | Guimaras | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30496 | DOH | xxxrhingringxxx@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 533 | 2025-02-18 | GX-IV | 838291 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | LCD Monitor - 19" HP E190i IPS | 12756 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6177381186534756258/6177381186534756258_signature_28.png | Rhea April Q Jaleco, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6177381186534756258/6177381186534756258_signature_31.png
| 2025-03-14 | 2025-03-14 | Central Luzon Doctor's Hospital | Hospital Drive Brgy. San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30495 | COMMERCIAL | pjlsantos@cldh.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 |  | GXIV-D | 839232 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6177377833528479994/6177377833528479994_signature_28.png | Floryvy Antonio | https://www.jotform.com/uploads/macarewecare/222927978372067/6177377833528479994/6177377833528479994_signature_31.png
| 2025-03-14 | 2025-03-14 | Central Luzon Doctor's Hospital | Hospital drive Brgy San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30494 | COMMERCIAL | pjlsantos@cldh.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 |  | GXII-D | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6177377153529505354/6177377153529505354_signature_28.png | Floryvy Antonio | https://www.jotform.com/uploads/macarewecare/222927978372067/6177377153529505354/6177377153529505354_signature_31.png
| 2025-03-14 | 2025-03-13 | Medellin Rural Health Unit | Poblacion Medellin Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | rhumedellin@gmail.com |  | 30493 | DOH | rhumedellin@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 1122 |  | GX4 | 838506 | -Due for Preventive Maintenance. - UPS during power failure did not provide back up power. Needs for replacement Defective UPS SN# AT1000190402ST0465 | Performed Preventive Maintenance, Clean ICore, Plunger, Filter, Cartridge Bay | Performed Preventive Maintenance, Clean ICore, Plunger, Filter, Cartridge Bay | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6177186410804692110/6177186410804692110_signature_28.png | Raul Galano, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6177186410804692110/6177186410804692110_signature_31.png
| 2025-03-13 | 2025-03-13 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | daguilon.intp@gmail.com |  | 30492 | MACARE-HBDC | msff-tondo-labtech@paris.msf.org | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4852 | 2025-03-12 | GX16-10C | 110016343 | Unit for Preventive Maintenance. Unit for Xpertcheck. | Preventive Maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Failed Modules: 210075770/A2 210075498/A3 210082298/B1 210075410/C3 |  | Xpertcheck Kit (4 boxes). | DR#12836 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6176558838028297768/6176558838028297768_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6176558838028297768/6176558838028297768_signature_31.png
| 2025-03-13 | 2025-03-13 | University of The East Ramon Magsaysay Medical Center | Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30491 | CONSORTIUM | lpcomendador@uerm.edu.ph | 25 |  |  | GXIV | 847578 | Unit for maintenance Errors on B2, B3 | NA | - checked and cleaned cartridge bay, plunger rod, filter, icore slot - generated reports - scanned system logs for hardware errors (5007 on B2, temp. errors on B3) - performed self test | Unit is operational Unit is due for calibration (Mar. 21, 2025) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6176505193111394662/6176505193111394662_signature_28.png | Osang Comendador | https://www.jotform.com/uploads/macarewecare/222927978372067/6176505193111394662/6176505193111394662_signature_31.png
| 2025-03-13 | 2025-03-13 | University of The East Ramon Magsaysay Medical Center | Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30490 | MACARE-HBDC | bloodbank@uerm.edu.ph | 28 |  |  | MCS+ | 13C270 | Unit for maintenance | Cleaned centrifuge Plugged unit to UPS | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6176481839029377749/6176481839029377749_signature_28.png | Ruel Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6176481839029377749/6176481839029377749_signature_31.png
| 2025-03-13 | 2025-03-13 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB LABORATORY | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30489 | DOH | buraeunlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 486 | 2025-03-13 | GXIV | 838551 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6176457421516715081/6176457421516715081_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6176457421516715081/6176457421516715081_signature_31.png
| 2025-03-13 | 2025-03-13 | Philippine Red Cross - Cagayan de Oro | Cagayan de Oro City | 10 | Misamis Oriental | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30488 | COMMERCIAL | mrbc2@redcross.com.ph | 25 |  |  | MCS+ | 02H080 | There's a mixture between platelet and red blood. | None | Installation of Incident Reporting Qr and Conversion factor sticker For observation due to the problem of a mix platelet and red blood. | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6176454799536760761/6176454799536760761_signature_28.png | Karen Torniado | https://www.jotform.com/uploads/macarewecare/222927978372067/6176454799536760761/6176454799536760761_signature_31.png
| 2025-03-13 | 2025-03-13 | PONTEVEDRA PRIMARY CARE FACILITY | Pontevedra, Capiz | 6 | Capiz | LABORATORY | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30487 | DOH | rtabngjr@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: Na |  | GX-IV | 838296 | - NO POWER(SYSTEM UNIT & UPS) | #NAME? | -ON-SITE STAFF REPLACEMENT OF UPS FUSE -INSPECT SYSTEM UNIT, CONFIRMED NO POWER(SN: *CPH0320F05928*) -PERFORM SYSTEM UNIT TROUBLESHOOTING, NO POWER ISSUE PERSISTS -PERFORM GENEXPERT PREVENTIVE MAINTENANCE PROCEDURE | SYSTEM UNIT FOR REPORTING TO CEPHEID |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6176385844111126328/6176385844111126328_signature_28.png | Reynaldo Tabang Jr. , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6176385844111126328/6176385844111126328_signature_31.png
| 2025-03-13 | 2025-03-12 | Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30486 | DOH | miniforestlaboratory@gmail.com | 23.8 |  |  | GX-IV 6C 4D | 838507 | Module 1 and 3 encounter multiple temperature related errors Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XPC on Module 1 and 3 | Perform Preventive Maintenance and Module Troubleshooting | Module 3 for Service Required Module 1 passed Only three Modules are operational | Module 1 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6176071398019448153/6176071398019448153_signature_28.png | Lolijell Torralba | https://www.jotform.com/uploads/macarewecare/222927978372067/6176071398019448153/6176071398019448153_signature_31.png
| 2025-03-12 | 2025-03-12 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | daguilon.intp@gmail.com |  | 30485 | MACARE-HBDC | msff-tondo-labtech@paris.msf.org | 25 |  | 2025-03-11 | GX16-10C | 110016343 | Unit for maintenance. For module replacement of module D4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module D4 (609286) in replacement to defective D4 (633865). | DR#12828 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175785220655140081/6175785220655140081_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6175785220655140081/6175785220655140081_signature_31.png
| 2025-03-12 | 2025-03-12 | Tagoloan Rural Health Unit | Poblacion, Tagoloan | 10 | Misamis Oriental | RTDL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30484 | DOH | laboratory.mhotagoloan@gmail.com | 25 |  |  | GX-4 6 Colors | 838574 | Module B2 found temperature errors Date and time always changing | CMOS replacement Module B2 reported to Cepheid | CMOS Battery Replacement Preventive Maintenance Data Collection | 3 module functional, 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175759986453509015/6175759986453509015_signature_28.png | Venz Michael Bacoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6175759986453509015/6175759986453509015_signature_31.png
| 2025-03-12 | 2025-03-12 | Tomas Cabili Health Center | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30483 | DOH | mariacielovenezuela@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 973 |  | 6C GXIV-4D | 838294 | -Module B1 and B2 error 2005 -Due for PM | -Preventive Maintenance; plunger rod, optic slot, fan filter, cartridge bay -Run XpertCheck on module B1 and B2 -Run dummy test on module B1 and B2 | -Preventive Maintenance; plunger rod, optic slot, fan filter, cartridge bay | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175759114569354670/6175759114569354670_signature_28.png | Maria Cielo Venezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/6175759114569354670/6175759114569354670_signature_31.png
| 2025-03-12 | 2025-03-12 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30482 | UNOPS | bailynmali10@gmail.com | 29.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 671 |  | GxIV-10C | 110023524 | CPU problem | Perform CPU repair Perform Genexpert machine Cleaning and maintenance Checking of errors/problem Problem/errors reporting Data collection | Cpu troubleshooting and repair Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175744785676168730/6175744785676168730_signature_28.png | Bailyn Ali | https://www.jotform.com/uploads/macarewecare/222927978372067/6175744785676168730/6175744785676168730_signature_31.png
| 2025-03-12 | 2025-03-12 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30481 | DOH | bailynmali10@gmail.com | 29.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 672 |  | GxIV-6C | 838528 | Genexpert temperature not within acceptable range | Perform Genexpert maintenance and cleaning Checking of errors Problem/errors reporting Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175741825673002156/6175741825673002156_signature_28.png | Bailyn Ali | https://www.jotform.com/uploads/macarewecare/222927978372067/6175741825673002156/6175741825673002156_signature_31.png
| 2025-03-12 | 2025-03-10 | Polomolok Rural Health Unit | Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30480 | DOH | jmmirasolrmt@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GxIV-6C | 838265 | Module 2 SN:685448 H/W failed due to multiple temperature errors (error 1002) | Perform XPC Calibration at module 2 for troubleshooting Perform Genexpert Cleaning and maintenance Checking of errors/problem Problem/errors reporting Data collection | Genexpert module 2 troubleshooting Genexpert machine Preventive maintenance | 3 functional modules XPC troubleshooting results: Module 2 SN:68544 - Service Required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175741805675062595/6175741805675062595_signature_28.png | June Michele Mirasol | https://www.jotform.com/uploads/macarewecare/222927978372067/6175741805675062595/6175741805675062595_signature_31.png
| 2025-03-12 | 2025-03-12 | Silliman University Medical Center | Dumaguete City | 7 | Cebu | Molecular Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30479 | COMMERCIAL | sumcmoleclab@gmail.com | 20 |  | 2025-03-12 | GX4 | 821471 | Due for Calibration. | . Save System Log file. . Save Installation Qualification Report . Save assay statistics report. . Save GXC file | Perform calibration for module 1,3 and 4 only with LOT# 06701 | Module 1,3 and 4 passed for calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175690932265117854/6175690932265117854_signature_28.png | Maryjune B. Bangalando | https://www.jotform.com/uploads/macarewecare/222927978372067/6175690932265117854/6175690932265117854_signature_31.png
| 2025-03-12 | 2025-03-12 | La Castellana Rural Health Unit | La Castellana, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | REPAIR | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30478 | DOH | milouepascual@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 728 | 2025-02-18 | GX4 | 838483 | Defective LCD monitor | Preventive Maintenance - cleaned filter, plunger rods, module bays, optic slots. | Replaced LCD monitor with model HP Elite Display E19Oi, SN CN45130SCV. | 4 modules operational |  | LCD monitor | DR 12758 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175681785396559779/6175681785396559779_signature_28.png | MIRA LUNA E PASCUAL | https://www.jotform.com/uploads/macarewecare/222927978372067/6175681785396559779/6175681785396559779_signature_31.png
| 2025-03-12 | 2025-03-12 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Molecular Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30477 | MACARE-HBDC | gemmarryp@gmail.com | 25 | ASSAY: Xpert C. difficile BT, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 104 ASSAY: Xpert HPV, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 157 ASSAY: Xpert SARS-CoV-2, COUNT: 149 |  | GX4-6C | 110008458 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, Gx fan, fan filter Generate Data | Performed Preventive Maintenance | Unit is Operational | Date and Time reset(CPU) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175678734438571406/6175678734438571406_signature_28.png | Gemmarry R. Pielago | https://www.jotform.com/uploads/macarewecare/222927978372067/6175678734438571406/6175678734438571406_signature_31.png
| 2025-03-12 | 2025-03-12 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30476 | DOH | lornsabaigar@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 593 | 2025-03-12 | GXIV | 838409 | N/a | N/@ | Preventive Maintenance | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175669441817144207/6175669441817144207_signature_28.png | Lorna C.Abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/6175669441817144207/6175669441817144207_signature_31.png
| 2025-03-12 | 2025-03-12 | Cebu South Medical Center | San Isidro, Talisay City Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30475 | DOH | csmcgenexpert@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1058 |  | GX4 | 838559 | Due for Preventive Maintenance. | Performed Preventive Maintenance, Clean ICore, Plunger, Filter, Cartridge Bay | Performed Preventive Maintenance, Clean ICore, Plunger, Filter, Cartridge Bay | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175652258931324951/6175652258931324951_signature_28.png | Kristel Sheen Miranda, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6175652258931324951/6175652258931324951_signature_31.png
| 2025-03-12 | 2025-03-12 | San Fernando Rural Health Unit | Poblacion, San Fernando | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30474 | DOH | josephus_garfin@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 827 |  | GX4-6C | 838496 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175617444852729063/6175617444852729063_signature_28.png | Josephus Garfin | https://www.jotform.com/uploads/macarewecare/222927978372067/6175617444852729063/6175617444852729063_signature_31.png
| 2025-03-12 | 2025-03-12 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | Indiko plus | 30473 | MACARE-HBDC | janettequirino@yahoo.com | 28 |  |  | 864 | 864000009213 | Sample dispenser initialization failed Deform sample probe Broken sample dispenser belt vertical | Checked sample probe | Performed troubleshooting | For replacement of sample probe |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175608153618738492/6175608153618738492_signature_28.png | Janette Quirino | https://www.jotform.com/uploads/macarewecare/222927978372067/6175608153618738492/6175608153618738492_signature_31.png
| 2025-03-12 | 2025-03-11 | MMC (Medilinx) | Amorsolo st., Makati City | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30472 | COMMERCIAL | janeshyrelenriquez.medilinx@outlook.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 258 ASSAY: Xpert CT/NG, COUNT: 832 ASSAY: Xpert HBV VL, COUNT: 31 ASSAY: Xpert HPV, COUNT: 685 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3650 ASSAY: Others, COUNT: 14 ASSAY: Xpert HCV VL, COUNT: 116 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2119 | 2025-03-11 | GX4D (10C) | 110001647 | B4 encounter continous 5007 error | Cleaned plunger rod,filter,icore,cartridge bay Performed reagent running Gathered reports Test run unit | Performed troubleshooting | B4 module passed reagent running in mtb rif ultra B1 reported already B2,B3 failed xpert check 12/5/2024 reported already |  | Xpert MTB/RIF Ultra | 12820 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175607623616081308/6175607623616081308_signature_28.png | Jane Shyrel M. Enriquez | https://www.jotform.com/uploads/macarewecare/222927978372067/6175607623616081308/6175607623616081308_signature_31.png
| 2025-03-12 | 2025-03-12 | Hinaplanon Health Center | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30471 | DOH | villacortaamy50@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 820 |  | 6C GXIV-4D | 838327 | Due for PM | -Preventive Maintenance; cartridge bay, plunger rod, fan filter, optic slot | #NAME? | Machine is functional. Module B1 waiting for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175603516768480713/6175603516768480713_signature_28.png | Amy Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6175603516768480713/6175603516768480713_signature_31.png
| 2025-03-12 | 2025-03-11 | San Jose Municipal Health Office | San Jose | 5 | Camarines Sur | Laboratory | YES | REPAIR | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30470 | DOH | jhuneric2008@yahoo.com | 25 |  |  | GX4-6C | 838389 | CPU for replacement | Replaced CPU Adjust/checked Gx information Installed mtb rif ultra assay Generate Data | Replaced CPU | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175586104857171473/6175586104857171473_signature_28.png | Rosmar V. Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6175586104857171473/6175586104857171473_signature_31.png
| 2025-03-12 | 2025-03-07 | Iriga City Health Office | Santa Cruz Sur, Iriga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30469 | UNOPS | hannahalevon1@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1037 |  | GX4-10C | 110023511 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175584344854894408/6175584344854894408_signature_28.png | Hannah Kirsten Novela | https://www.jotform.com/uploads/macarewecare/222927978372067/6175584344854894408/6175584344854894408_signature_31.png
| 2025-03-12 | 2025-03-07 | Iriga City Health Office | Santa Cruz Sur | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30468 | DOH | hannahalevon1@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1q06 |  | GX4-6C | 838391 | B4/210040454 for module replacement | Replaced module: B4 - 210040454>>210008370 Cleaned cartridge bays, plunger rod Generate Data | Replaced B4/210040454 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175582964859603138/6175582964859603138_signature_28.png | Hannah Kirsten Novela | https://www.jotform.com/uploads/macarewecare/222927978372067/6175582964859603138/6175582964859603138_signature_31.png
| 2025-03-12 | 2025-03-06 | San Jose Municipal Health Office | San Jose | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30467 | DOH | jhuneric2008@yahoo.com | 26 |  |  | GX4-6C | 838389 | CPU restarts every few seconds CPU hard to power on | Installed CPU(no cables) Cleaned cartridge bays, plunger rod | Performed Troubleshooting CPU replacement Preventive Maintenance | Unit is not operational CPU not installed (will purchase cable) | Will purchase VGA to Display port cable for monitor to cpu connection |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175578404856569923/6175578404856569923_signature_28.png | Rosmar V. Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6175578404856569923/6175578404856569923_signature_31.png
| 2025-03-12 | 2025-02-24 | National Kidney and Transplant Institute | East Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Clinical Immunology | NO | INSTALLATION | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Fujifilm | 30466 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  | 2025-02-24 | MT-7500E | 37020507 | N/A | N/A | Installed Limusave System | Unit is operational |  | MT-7500E SN: 37020507 Thermostation SN: 36920310 Cooling Station SN: 23050048 | DR 12776 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6175511023613903050/6175511023613903050_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6175511023613903050/6175511023613903050_signature_31.png
| 2025-03-11 | 2025-03-11 | ILOILO PROVINCIAL HOSPITAL | POTOTAN, ILOILO | 6 | Iloilo | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30465 | DOH | djpragale92@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 869 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV | 838289 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6174921346816411446/6174921346816411446_signature_28.png | Rachel A. Pragale | https://www.jotform.com/uploads/macarewecare/222927978372067/6174921346816411446/6174921346816411446_signature_31.png
| 2025-03-11 | 2025-03-11 | Don Jose S. Monfort Medical Center | Barotac Nuevo, Iloilo | 6 | Iloilo | LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30464 | DOH | judyboglosa35@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1843 |  | GX-IV | 838501 | -MODULE B2 ERROR 1002 | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL , MODULE B2 FOR REPORTING |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6174918856813214959/6174918856813214959_signature_28.png | Judy Boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6174918856813214959/6174918856813214959_signature_31.png
| 2025-03-11 | 2025-03-11 | Chong hua hospital - mandaue | Mandaue City, Cebu | 7 | Cebu | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com | chhm_bloodbank_@chonghua.com.ph |  | 30463 | COMMERCIAL | gerardo_ochea@chonghua.com.ph | 28 |  |  | Mcs+ | 11J217 | Due date for PM and calibration | Preventive Maintenance and Diagnostics Calibration | Preventive Maintenance and Diagnostics Calibration | Systems operational No machine parts were replaced during calibration and Preventive Maintenance |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6174766427793655161/6174766427793655161_signature_28.png | Chonghua hospital mandaue | https://www.jotform.com/uploads/macarewecare/222927978372067/6174766427793655161/6174766427793655161_signature_31.png
| 2025-03-11 | 2025-03-11 | Quezon City Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30462 | COMMERCIAL | jennyquintero90@yahoo.com | 25 |  |  | GXIV-10C | 120003581 | Unit for maintenance | #NAME? | #NAME? | Unit is operational Unit has not been used yet |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6174756207812733843/6174756207812733843_signature_28.png | Jenny Quintero | https://www.jotform.com/uploads/macarewecare/222927978372067/6174756207812733843/6174756207812733843_signature_31.png
| 2025-03-11 | 2025-03-11 | Kauswagan Rural Health Unit | Kauswagan, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30461 | DOH | dextertinoy@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 393 |  | 6C GXIV-4D | 838402 | Due for PM | -Preventive Maintenance; plunger rod, fan filter, cartridge bay, optic slot | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6174739227611685416/6174739227611685416_signature_28.png | Dexter Tinoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6174739227611685416/6174739227611685416_signature_31.png
| 2025-03-10 | 2025-03-10 | Macare In house | 1005 Aurora Grand, QC | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | ACCUPOWER | 30460 | MACARE-HBDC | florescanoel77@gmail.com | 28 |  |  | 1KVA UPS | AT1000190123ST0327 | No power Drain battery | Checked unit Checked and test battery voltage output 4volts | Performed troubleshooting | For repair for replacement of 2 pcs battery |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6174025943616195067/6174025943616195067_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6174025943616195067/6174025943616195067_signature_31.png
| 2025-03-10 | 2025-03-07 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30459 | DOH | suzzettekho34@gmail.com | 29 |  |  | GX-4 6 colors | 838554 | None | None | Preventive Maintenance | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173876265953906463/6173876265953906463_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/6173876265953906463/6173876265953906463_signature_31.png
| 2025-03-10 | 2025-03-03 | Nasipit Rural Health Unit | Brgy. Talisay, Nasipit | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30458 | DOH | peterdolotina@gmail.com | 27 |  |  | GX-4 6 Colors | 837611 | B1 found temperature related errors | Reported to Cepheid | Troubleshooting and Preventive Maintenance | 3 modules functional and 1 module reported. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173874695953804928/6173874695953804928_signature_28.png | Peter Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/6173874695953804928/6173874695953804928_signature_31.png
| 2025-03-10 | 2025-03-06 | Our Lady of Lourdes Hospital | Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30457 | CONSORTIUM | mariaysabelle.visda@medilinx.com.ph | 25 |  |  | GXIV | 840974 | Unit for maintenance | NA | #NAME? | Unit is operational Unit is due for calibration (3-20-25) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173847253611177932/6173847253611177932_signature_28.png | Maria Ysabelle Visda | https://www.jotform.com/uploads/macarewecare/222927978372067/6173847253611177932/6173847253611177932_signature_31.png
| 2025-03-10 | 2025-03-04 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30456 | DOH | sherryannemarrero@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1978 | 2025-02-20 | GXIV-4-D-6C | 838281 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B1 678653 >>> 210041917 - Generated Reports | Performed Module Replacement | Unit is Fully Operational |  | GeneXpert Module 210041917 | DR 12771 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173832228293824474/6173832228293824474_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6173832228293824474/6173832228293824474_signature_31.png
| 2025-03-10 | 2025-03-06 | Rizal Medical Center | Pasig Blvd Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | gonzagakevin4@gmail.com |  | 30455 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 |  |  | GX-XVI-D-6C | 838627 | Modules not detected. | Checked connections | Checked module connections Checked cables Checked internal switch | Unit is not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173774443615278938/6173774443615278938_signature_28.png | Justine Lorenz V. Falzado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6173774443615278938/6173774443615278938_signature_31.png
| 2025-03-10 | 2025-03-07 | Pasig Treatment Hub - Champ | Champ Bldg, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | pathmedtech@gmail.com |  | 30452 | MACARE-HBDC | dansumulong.pasigcho@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 253 |  | GX-IV-D-10C | 110025127 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173774253611167570/6173774253611167570_signature_28.png | DAN NATHANIEL SUMULONG | https://www.jotform.com/uploads/macarewecare/222927978372067/6173774253611167570/6173774253611167570_signature_31.png
| 2025-03-10 | 2025-03-06 | Rizal Medical Center | Pasig Blvd Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | gonzagakevin4@gmail.com |  | 30453 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 |  |  | GX-XVI-D-6C | 838627 | Modules not detected. | Checked connections | Checked module connections Checked cables Checked internal switch | Unit is not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173774253611308029/6173774253611308029_signature_28.png | Justine Lorenz V. Falzado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6173774253611308029/6173774253611308029_signature_31.png
| 2025-03-10 | 2025-03-05 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30454 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 |  |  | GX-XVI-D-6C | 838627 | Modules not detected | Checked connections | Checked connections Checked ports Checked cables Checked IP address Swap cable | Unit is not operational. For additional troubleshooting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173774253611567941/6173774253611567941_signature_28.png | Benita Gracia N. Viloria | https://www.jotform.com/uploads/macarewecare/222927978372067/6173774253611567941/6173774253611567941_signature_31.png
| 2025-03-10 | 2025-03-07 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30451 | MACARE-HBDC | lcpmicrolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 185 ASSAY: Xpert SARS-CoV-2, COUNT: 262 ASSAY: Xpert HIV-1 Viral Load, COUNT: 69 ASSAY: Xpert SARS-CoV-2, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 9 |  | GX4D | 838837 | Due for maintenance B1,B4 reported already due to temp error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | B3 for reporting B1,B4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173756683611386864/6173756683611386864_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6173756683611386864/6173756683611386864_signature_31.png
| 2025-03-09 | 2025-03-06 | Kidapawan City Health Office | Kidapawan City, North Cotabato | 12 | Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30450 | DOH | kidcholaboratory@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2010 |  | GxIV-6C | 838545 | None | Perform Genexpert Cleaning and maintenance Checking of Gx errors/problem Data collection | Genexpert machine Preventive maintenance | 4 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173382376311116123/6173382376311116123_signature_28.png | Neil Everson Cunanan | https://www.jotform.com/uploads/macarewecare/222927978372067/6173382376311116123/6173382376311116123_signature_31.png
| 2025-03-09 | 2025-03-09 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 30449 | COMMERCIAL | cduhbloodbank2@gmail.com | 23 |  |  | MCS+ | 17H639SPG | Purple valve problem, test failure and successive valve error problem during procedure. | Replacement of the new pneumatic assembly. | Replacement of the new pneumatic assembly. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6173229852111657623/6173229852111657623_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6173229852111657623/6173229852111657623_signature_31.png
| 2025-03-08 | 2025-03-08 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cduhbloodbank2@gmail.com |  | 30448 | COMMERCIAL | cduhbloodbank2@gmail.com | 23 |  |  | MCS+ | 17H639SPG | Purple valve problem and successive valve error problem during procedure. | Pneumatic Assembly defective need to be replace. | Reset and check pneumatic assembly of the MCS+ unit. | Unit always encountered a valve error problem. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6172041690807089328/6172041690807089328_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6172041690807089328/6172041690807089328_signature_31.png
| 2025-03-07 | 2025-03-07 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | CALIBRATION | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30447 | COMMERCIAL | cduhbloodbank2@gmail.com | 23 |  |  | MCS+ | 17H639SPG | #NAME? | Performed calibration on Centrifuge, Weigher, SPM and DPM, Cuff Generate Data, calibrate and check line sensor and bowl optics sensitivity test Performed test run | Performed calibration on Centrifuge, Weigher, SPM and DPM, Cuff Generate Data, calibrate and check line sensor and bowl optics sensitivity test Performed test run | Units is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6171378238546977916/6171378238546977916_signature_28.png | Cebu Doctors Hospital University Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6171378238546977916/6171378238546977916_signature_31.png
| 2025-03-07 | 2025-03-07 | Cebu Doctors Hospital | Cebu City | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 30446 | COMMERCIAL | cduhbloodbank2@gmail.com | 23 |  |  | Mcs+ | 14F014 | Due date for PM and Calibration | Preventive Maintenance Haemonitics Calibration | Preventive Maintenance Haemonitics Calibration | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6171375167794645267/6171375167794645267_signature_28.png | CDUH Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6171375167794645267/6171375167794645267_signature_31.png
| 2025-03-07 | 2025-03-07 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30445 | COMMERCIAL | cduhcovidpoct@gmail.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 84 ASSAY: Xpert CT/NG, COUNT: 249 ASSAY: Xpert HBV VL, COUNT: 187 ASSAY: Xpert HPV, COUNT: 42 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1132 ASSAY: Xpert HCV VL, COUNT: 147 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GX-IV | 110008524 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | Module B3 and Module B4 Disabled |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6171364980904708015/6171364980904708015_signature_28.png | Jennifer Marie Lim, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6171364980904708015/6171364980904708015_signature_31.png
| 2025-03-07 | 2025-03-07 | San Pablo Colleges Medical Centee | San Pablo Laguna | 4A | Laguna | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | spcmc.molecularlab@gmail.com |  | 30444 | CONSORTIUM | spcmc.molecularlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 58 |  | GXIV-10C | 110024552 | Unit for preventive maintenance. For installation of HPV Assay Definition File. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Installed HPV ADF. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6171358824416587904/6171358824416587904_signature_28.png | Kelly Bondoc | https://www.jotform.com/uploads/macarewecare/222927978372067/6171358824416587904/6171358824416587904_signature_31.png
| 2025-03-07 | 2025-03-07 | Project 7 Social Hygiene Clinic | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30443 | UNOPS | michelenaval@gmail.com | 25 |  |  | GXIV-10C | 120000577 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6171321737811777695/6171321737811777695_signature_28.png | Michele Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6171321737811777695/6171321737811777695_signature_31.png
| 2025-03-07 | 2025-03-07 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | aivee.alcantara@medilinx.com.ph |  | 30442 | COMMERCIAL | maria.resurreccion@medilinx.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 68 ASSAY: Xpert HPV, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 787 |  | GXII | 812916 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6171282134414101524/6171282134414101524_signature_28.png | Maria Niña R. Resurreccion | https://www.jotform.com/uploads/macarewecare/222927978372067/6171282134414101524/6171282134414101524_signature_31.png
| 2025-03-07 | 2025-03-07 | Molave Rural Health unit | Moan Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30441 | DOH | aldrinacain@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 686 | 2025-03-07 | Gx 4 | 838233 | None | Preventive Maintenence and data gathering | Preventive Maintenence and data gathering | All modules are available. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6171270973215166555/6171270973215166555_signature_28.png | Aldrin Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/6171270973215166555/6171270973215166555_signature_31.png
| 2025-03-07 | 2025-03-07 | Jose Abad Santos Rural Health Unit | Caburan, Jose Abad Santos | 11 | Davao Occidental | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30440 | DOH | bevespinosa168@gmail.com | 26.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 452 | 2025-02-20 | GX-IV 6C 4D | 838460 | Module B2 for Service Required Due for Preventive Maintenance | Replaced the Defective Module SN: 861665 with new Module SN: 702471 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Module Replacement and Preventive Maintenance | All Modules are operational |  | 900-0491rs | DR 12769 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6171201685132501719/6171201685132501719_signature_28.png | Beverly Espinosa,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6171201685132501719/6171201685132501719_signature_31.png
| 2025-03-06 | 2025-03-06 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30439 | COMMERCIAL | debbiebaga@gmail.com | 22 |  |  | MCS+ | 11H307 | FOR MAINTENANCE AND CALIBRATION | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6170487550801162122/6170487550801162122_signature_28.png | Debbie Joyce Baga | https://www.jotform.com/uploads/macarewecare/222927978372067/6170487550801162122/6170487550801162122_signature_31.png
| 2025-03-06 | 2025-03-06 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30438 | COMMERCIAL | debbiebaga@gmail.com | 22 |  |  | MCS+ | 02D167 | FOR MAINTENANCE AND CALIBRATION | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6170485130809492072/6170485130809492072_signature_28.png | Debbie Joyce Baga | https://www.jotform.com/uploads/macarewecare/222927978372067/6170485130809492072/6170485130809492072_signature_31.png
| 2025-03-06 | 2025-03-06 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30437 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2256 |  | 6C GXIV-4D | 838561 | Due for PM | -Preventive Maintenance; cartridge bay, fan filter, optic slot, plunger rod | -Preventive Maintenance -Troubleshoot module 3 | Module A3 H/W fail for replacement. Other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6170468859076420470/6170468859076420470_signature_28.png | Hannah Jane Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/6170468859076420470/6170468859076420470_signature_31.png
| 2025-03-06 | 2025-03-06 | Diit District Health Center | Diit, Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30436 | DOH | dalepenalosa189@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1158 | 2025-03-06 | GXIV | 838421 | N/a | N/a | #NAME? | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6170443731886917201/6170443731886917201_signature_28.png | Dale Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6170443731886917201/6170443731886917201_signature_31.png
| 2025-03-06 | 2025-03-06 | Alicia Rural Health Unit | Alicia Rural Health Unit | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30435 | DOH | arielbaclig6@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 523 | 2025-03-06 | GX 4 | 838276 | UPS unable to work properly. | Replaced Uninterrupted power supply batteries and conducted Preventive maintenance and data collection. | Replaced Uninterrupted power supply batteries and conducted Preventive maintenance and data collection. | All modules are available and UPS is working properly. | Newly replaced UPS batteries must be charged to 6 to 8grs before use. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6170385928421787364/6170385928421787364_signature_28.png | Ariel tolentino Baclig | https://www.jotform.com/uploads/macarewecare/222927978372067/6170385928421787364/6170385928421787364_signature_31.png
| 2025-03-05 | 2025-03-05 | Bataan General Hospital and Medical Center | Balanga City, Bataan | 3 | Bataan | PCR Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30434 | MACARE-HBDC | raedionisio@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 18 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 212 |  | GXIV-D 10C | 120003587 | Machine Relocation | Preventive maintenance Relocation | Performed Preventive maintenance Relocation machine from RHIVDA to PCR Laboratory Performed functionality test | Machine is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169989061871793860/6169989061871793860_signature_28.png | Raeann Dionisio | https://www.jotform.com/uploads/macarewecare/222927978372067/6169989061871793860/6169989061871793860_signature_31.png
| 2025-03-05 | 2025-03-05 | Dr. Amado B. Diaz Provincial Foundation Hospital | Midsayap, North Cotabato | 12 | Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30433 | MACARE-HBDC | charlenemaeamatac526@gmail.com | 23.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 793 ASSAY: Xpert HIV-1 Viral Load, COUNT: 106 | 2025-03-05 | GxIV-10C | 110024249 | None | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors / Problem Data collection | Genexpert machine Preventive maintenance Genexpert machine annual Xpertcheck Calibration | 4 functional modules Xpertcheck Calibration Results: Module 1,2,3 and 4 - Passed | All modules passed Xpertcheck Calibration | Xpertcheck CE-5 cartridge Lot #:1001120421 | DR number: 5432 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169814245511166147/6169814245511166147_signature_28.png | Charlene Mae A. Amatac | https://www.jotform.com/uploads/macarewecare/222927978372067/6169814245511166147/6169814245511166147_signature_31.png
| 2025-03-05 | 2025-03-05 | Antipas Rural Health Unit | Antipas, North Cotabato | 12 | Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30432 | DOH | codillaalvin06@gmail.com | 20.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 496 |  | GxIV-6C | 838259 | None | Genexpert machine Preventive maintenance | Genexpert machine Preventive maintenance | Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169814245519079063/6169814245519079063_signature_28.png | Alvin A. Codilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6169814245519079063/6169814245519079063_signature_31.png
| 2025-03-05 | 2025-03-05 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30431 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 0 | 2025-02-26 | GXIV | 807956 | Desktop workstation for installation for 2 GX Macare units | NA | - installed desktop - switched 2 GeneXpert units (807112) - generated reports - performed self test - installed assay definition files | Unit is operational Intermittent errors for endrosement to technical team | Switched 807112 | GX Desktop SN: SGH530PD4N GX Monitor Display SN: CNK831078T | 12788 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169654083543615705/6169654083543615705_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6169654083543615705/6169654083543615705_signature_31.png
| 2025-03-05 | 2025-03-05 | Ramon magsaysay rural health unit | Ramon magsaysay Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30430 | DOH | borinagagina29@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 507 | 2025-03-05 | GX 4 | 838443 | None | Preventive Maintenence and data gathering. | Preventive Maintenence and data gathering. | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169561169219855048/6169561169219855048_signature_28.png | Gina Borinaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6169561169219855048/6169561169219855048_signature_31.png
| 2025-03-05 | 2025-03-05 | Davao Regional Medical Center - BloodBank | Apokon, Tagum City | 11 | Davao del Norte | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 30429 | COMMERCIAL | bloodbank@drmc.doh.gov.ph | 27.9 |  |  | MCS+ | 21M389SPG | For Preventive Maintenance | Clean Surface Area, check parameters of SPM, DPM, and Line Sensor Check the roller assays Perform functionality test Install QR Sticker for Reporting and Install Conversion Factor Table | Perform Preventive Maintenance | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169558120456426789/6169558120456426789_signature_28.png | Christine Porquiado,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6169558120456426789/6169558120456426789_signature_31.png
| 2025-03-05 | 2025-03-04 | St. Luke's Medical Center - Global City | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 30428 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 132 ASSAY: Xpert CT/NG, COUNT: 490 ASSAY: Xpert HBV VL, COUNT: 36 ASSAY: Xpert HPV, COUNT: 651 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1610 ASSAY: Xpert MTB-XDR, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 31 ASSAY: Xpert HIV-1 Viral Load, COUNT: 179 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169499243615710958/6169499243615710958_signature_28.png | Sapla, Ma Grace | https://www.jotform.com/uploads/macarewecare/222927978372067/6169499243615710958/6169499243615710958_signature_31.png
| 2025-03-05 | 2025-03-04 | St. Luke's Medical Center - Global City | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 30427 | COMMERCIAL | jmtentica@stlukes.com.ph | 22 |  | 2025-03-03 | MCS+ | 09C125 | Variocar Stand Replacement For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test Replaced Variocar Stand | Unit is operational. |  | Variocar Stand | DR# 12800 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169499243616194574/6169499243616194574_signature_28.png | Jen Michelle Entica | https://www.jotform.com/uploads/macarewecare/222927978372067/6169499243616194574/6169499243616194574_signature_31.png
| 2025-03-05 | 2025-03-04 | St. Luke's Medical Center - Global City | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 30426 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 22 ASSAY: Xpert CT/NG, COUNT: 122 ASSAY: Xpert HPV, COUNT: 165 ASSAY: Xpert MTB/RIF Ultra, COUNT: 297 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 64 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A3/866793 for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169499243618462678/6169499243618462678_signature_28.png | Sapla, Ma Grace | https://www.jotform.com/uploads/macarewecare/222927978372067/6169499243618462678/6169499243618462678_signature_31.png
| 2025-03-04 | 2025-03-04 | Zamboanga del sur medical center | Zamboanga del sur medical center, Pagadian city | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30425 | DOH | limboemillie@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1881 ASSAY: Xpert HIV-1 Viral Load, COUNT: 114 | 2025-03-04 | GX 4 | 838204 | None | Preventive Maintenence and data gathering. | Preventive Maintenence and data gathering. | All modules are available. | Replaced filters | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6169005722216211628/6169005722216211628_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/6169005722216211628/6169005722216211628_signature_31.png
| 2025-03-04 | 2025-03-04 | Davao Reproductive Health and Wellness Center | Jacinto Street, Poblacion District Davao | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30424 | DOH | rhwclab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 410 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1752 ASSAY: Xpert HBV VL, COUNT: 1 |  | GX-IV 6C 4D | 838262 | For Preventive Maintenance Module 1 encounter multiple errors | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XPC on Module 1 Only | Perform Preventive Maintenance, Functionality Test and XPC Calibration on Module 1 | Module 1 for Service Required Only Three Modules are operational | Module 1 disabled due to multiple errors |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6168742480456437948/6168742480456437948_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6168742480456437948/6168742480456437948_signature_31.png
| 2025-03-04 | 2025-03-04 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30423 | MACARE-HBDC | rhwclab@gmail.com | 26.5 | ASSAY: Xpert HBV VL, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 410 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1754 |  | GX-IV 6C 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Only Three Modules are operational | Module 3 disabled due to frequent temperature related errors |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6168742070454539219/6168742070454539219_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6168742070454539219/6168742070454539219_signature_31.png
| 2025-03-04 | 2025-03-03 | Candijay Community Hospital | Candijay, Bohol | 7 | Bohol | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30422 | DOH | angelinaredulla@yahoo.com.ph | 28 |  |  | Gx4 | 838396 | Occasionaly CPU lagging | Clean RAM Installing GX new version 6.5 from previous GX version 5.1 Install new RAM from cebu office | Clean RAM Installing GX new version 6.5 from previous GX version 5.1 Install new RAM from cebu office | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6168627315418260363/6168627315418260363_signature_28.png | Angelina O. Redulla | https://www.jotform.com/uploads/macarewecare/222927978372067/6168627315418260363/6168627315418260363_signature_31.png
| 2025-03-03 | 2025-02-25 | Datu Odin Sinsuat District Hospital | Datu Odin Sinsuat, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30421 | DOH | rowenasalazar1234@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 828 |  | GxIV-6C | 838205 | None | Perform Genexpert machine Cleaning and maintenance Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6168160936862719618/6168160936862719618_signature_28.png | Rowena Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/6168160936862719618/6168160936862719618_signature_31.png
| 2025-03-03 | 2025-03-03 | CTRL | Cebu City | 7 | Cebu | TB Laboratory | NO | Pull out CPU | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30420 | DOH | asmantiquilla.pbsp@gmail.com | 23 |  | 2025-03-03 | GX4 | 838263 | LCD no data on Screen | Pull out CPU | Pull out CPU with serial number CPH05020F08297 | Not ok to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167941567026678542/6167941567026678542_signature_28.png | Allan Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6167941567026678542/6167941567026678542_signature_31.png
| 2025-03-03 | 2025-02-27 | Balamban Rural Health Unit | Poblacion Balamban Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30419 | DOH | climaco.mildred@yahoo.com | 23 |  | 2025-02-27 | GX4 | 838231 | . Module A1 Error code 5006 . Due for monthly maintenance. | . Save System log file . Save Installation qualification report. . Save Assay statistics report | . Clean cartridge bay . Clean filter . Perform preventive maintenance for GX system . Clean UPS, CPU, LCD , Printer, AVR | . All modules are functional . Module A1 observation for running |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167934987027621100/6167934987027621100_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6167934987027621100/6167934987027621100_signature_31.png
| 2025-03-03 | 2025-02-28 | Manabo Rural Health Unit | Manabo, Abra | 1 | Abra | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30418 | DOH | lenmayclairelumiwag15@gmail.com | 30 |  | 2025-02-20 | GXIV-4-D-6C | 838274 | For CPU Replacement and Module Replacement | - Installed Backup CPU - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 807809 >>> 210068341 - Generated Reports | Performed CPU Replacement and Module Replacement | Unit is Fully Operational |  | GeneXpert Module 210068341 GeneXpert Computer SGH530PD9W | DR 12743, DR 12770 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167905558292416064/6167905558292416064_signature_28.png | Merceditas Emma Quilala | https://www.jotform.com/uploads/macarewecare/222927978372067/6167905558292416064/6167905558292416064_signature_31.png
| 2025-03-03 | 2025-02-27 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30417 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 57 ASSAY: Xpert CT/NG, COUNT: 42 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 41 ASSAY: Xpert HPV, COUNT: 414 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6274 ASSAY: Xpert MTB-XDR, COUNT: 2 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 129 ASSAY: Xpert HIV-1 Viral Load, COUNT: 853 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1344 |  | GX-XVI-D-10C | 110015767 | Error 5007 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports Generated gxx file | Unit is operational. Modules with error 5007 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167904483611286868/6167904483611286868_signature_28.png | Kristine Chelle Parilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6167904483611286868/6167904483611286868_signature_31.png
| 2025-03-03 | 2025-02-28 | Ospital ng Makati | Sampaguita St. Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@gmail.com |  | 30416 | COMMERCIAL | Sheirylpia021@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 332 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167904483614504680/6167904483614504680_signature_28.png | Sheiryl E. Pia | https://www.jotform.com/uploads/macarewecare/222927978372067/6167904483614504680/6167904483614504680_signature_31.png
| 2025-03-03 | 2025-02-27 | Laoag City Health Office | Laoag City, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30415 | DOH | denver.mahor@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 792 |  | GXIV-4-D-6C | 838416 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167900328294450131/6167900328294450131_signature_28.png | Rodenver Mahor | https://www.jotform.com/uploads/macarewecare/222927978372067/6167900328294450131/6167900328294450131_signature_31.png
| 2025-03-03 | 2025-02-27 | Black Nazarene Hospital | San Nicholas, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30414 | CONSORTIUM | tbnhi.lab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 34 |  | GXIV-4-D-6C | 835542 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167898788291069553/6167898788291069553_signature_28.png | WEENA LORRAINE MACUSI | https://www.jotform.com/uploads/macarewecare/222927978372067/6167898788291069553/6167898788291069553_signature_31.png
| 2025-03-03 | 2025-02-27 | Sinait District Hospital | Sinait, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30413 | DOH | isdhsinaitlaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 606 |  | GXIV-4-D-6C | 838277 | For Preventive Maintenance and Troubleshooting | #NAME? | Performed Preventive Maintenance and Troubleshooting | Unit is Fully Operational For Observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167897158292246341/6167897158292246341_signature_28.png | MARK JAYBEE TACATA | https://www.jotform.com/uploads/macarewecare/222927978372067/6167897158292246341/6167897158292246341_signature_31.png
| 2025-03-03 | 2025-02-25 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30412 | DOH | seansimonpaloyo@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 1140 |  | GXIV-4-D-6C | 838473 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is operational but for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167891768298506720/6167891768298506720_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6167891768298506720/6167891768298506720_signature_31.png
| 2025-03-03 | 2025-02-27 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30411 | DOH | piaramuj@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 801 |  | GXIV-4-D-6C | 838270 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167886188291627571/6167886188291627571_signature_28.png | Mikka Gaerlan | https://www.jotform.com/uploads/macarewecare/222927978372067/6167886188291627571/6167886188291627571_signature_31.png
| 2025-03-03 | 2025-02-25 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30410 | DOH | annabellelcatbagan@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 52 |  | GXII-2D | 820303 | For additional Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Partially Operational For Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167884388291109650/6167884388291109650_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6167884388291109650/6167884388291109650_signature_31.png
| 2025-03-03 | 2025-02-25 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30409 | DOH | seansimonpaloyo@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1140 |  | GXIV-4-D-6C | 838473 | Error 2031 Modules detected but cannot perform XpertCheck | #NAME? | Performed Troubleshooting | Unit is Operational but for reporting due to Calibration Error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167882538298287333/6167882538298287333_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6167882538298287333/6167882538298287333_signature_31.png
| 2025-03-03 | 2025-03-03 | Monkayo Rural Health Unit | Poblacion, Monkayo | 11 | Davao de Oro | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30408 | DOH | kendralois730@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1512 | 2025-02-18 | GX-IV 6C 4D | 838487 | Monitor LCD Display Problem Due for Preventive Maintenance | Replaced the Defective Monitor SN:CNK9430B0Y with Monitor SN: CN461017M Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Monitor Replacement and Preventive Maintenance | All Modules are operational Monitor successfully replaced |  | LCD Monitor HP E190 | DR 12755 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167853673425359214/6167853673425359214_signature_28.png | Karen S. Anuta | https://www.jotform.com/uploads/macarewecare/222927978372067/6167853673425359214/6167853673425359214_signature_31.png
| 2025-03-03 | 2025-02-28 | De La Salle Health and Science Institute | Congressional Road, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30406 | CONSORTIUM | cevillete@dlsmhsi.edu.ph | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2092 |  | GX16-D-6C | 842166 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filters if Clogged. Generate Reports. Performed Self-test. | No Modules were Disabled. GX System is Good. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167711753614002009/6167711753614002009_signature_28.png | Chlorish Ann E. Villete | https://www.jotform.com/uploads/macarewecare/222927978372067/6167711753614002009/6167711753614002009_signature_31.png
| 2025-03-03 | 2025-02-28 | De La Salle Health and Science Institute | Congressional Road, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30407 | CONSORTIUM | mgrioflorido@tbhivicr.org.ph | 25 |  |  | GX16-8-D-10C | 110012951 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filters if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167711753618364335/6167711753618364335_signature_28.png | Ma. Cristina Rioflorido | https://www.jotform.com/uploads/macarewecare/222927978372067/6167711753618364335/6167711753618364335_signature_31.png
| 2025-03-03 | 2025-02-28 | De La Salle Health and Science Institute | Congressional Road, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30405 | CONSORTIUM | cevillete@dlsmhsi.edu.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 584 |  | GXIV-D-6C | 812380 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | No Modules were Disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167711753619853788/6167711753619853788_signature_28.png | Chlorish Ann E. Villete | https://www.jotform.com/uploads/macarewecare/222927978372067/6167711753619853788/6167711753619853788_signature_31.png
| 2025-03-03 | 2025-02-28 | Tropical Disease Foundation | Amorsolo st., Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30404 | CONSORTIUM | krcueno@tdf.org.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2066 ASSAY: Xpert HIV-1 Viral Load, COUNT: 82 |  | GXIV D | 812560 | Due for maintenance C4 disabled due to error 5007 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | C4 disable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167698413616865542/6167698413616865542_signature_28.png | Karl Francis R. Cueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6167698413616865542/6167698413616865542_signature_31.png
| 2025-03-03 | 2025-02-26 | Southern Philippines Medical Center | JP Laurel Avenue, Bajada Davao City | 11 | Davao del Sur | RHIVDA | NO | ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30403 | UNOPS | spmcrhivda@gmail.com | 24.5 |  |  | GX-IV 10C 4D | 110023506 | Onsite Assistance for printer driver installation | Install printer driver of Epson M1120 series | Perform Printer Driver Installation | Successfully install printer drivers, unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167418322422181965/6167418322422181965_signature_28.png | Flordeliza B. Go, RMT, MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/6167418322422181965/6167418322422181965_signature_31.png
| 2025-03-02 | 2025-02-25 | Estevez Memorial Hospital, Inc. | Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City, 4500 Albay | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30402 | COMMERCIAL | melosept22@gmail.com | 27 |  |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6167043856417745241/6167043856417745241_signature_28.png | Estevez MHI | https://www.jotform.com/uploads/macarewecare/222927978372067/6167043856417745241/6167043856417745241_signature_31.png
| 2025-02-28 | 2025-02-28 | Batangas City Health Office | Batangas City, Batangas | 4A | Batangas | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | gjanlorenzo@gmail.com |  | 30401 | MACARE-HBDC | kaimoral68@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 857 |  | GXIV-10C | 110024615 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165392645115376352/6165392645115376352_signature_28.png | Catherine | https://www.jotform.com/uploads/macarewecare/222927978372067/6165392645115376352/6165392645115376352_signature_31.png
| 2025-02-28 | 2025-02-28 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30400 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2358 ASSAY: Xpert HIV-1 Viral Load, COUNT: 74 |  | 6C GCIV-4D | 838285 | Due for PM | -Preventive Maintenance; plunger rod, fan filter, cartridge bay | #NAME? | Machine is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165303687768909262/6165303687768909262_signature_28.png | Leah Karen Noynay | https://www.jotform.com/uploads/macarewecare/222927978372067/6165303687768909262/6165303687768909262_signature_31.png
| 2025-02-28 | 2025-02-28 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30399 | UNOPS | laureanjo@yahoo.com | 28 |  |  | Gx4 | 110023513 | N/A | Install new fan filter Gather data IQR, system log | Install new fan filter Gather data IQR, system log | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165294981412293141/6165294981412293141_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6165294981412293141/6165294981412293141_signature_31.png
| 2025-02-28 | 2025-02-28 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30398 | DOH | laureanjo@yahoo.com | 29 |  |  | Gx 4 | 838527 | N/A | Install new fan filter Gather data, IQR, system log | Install new fan filter Gather data, IQR, system log | Sytems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165293771417605148/6165293771417605148_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6165293771417605148/6165293771417605148_signature_31.png
| 2025-02-28 | 2025-02-28 | Labuan General hospital | Labuan, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Preventive Maintenence | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30397 | DOH | jenne06021@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2156 | 2025-02-25 | GX 4 | 838568 | Module 2 h/w failed, CPU not turning on. | Preventive Maintenence, analyse data collected and troubleshoot. | Preventive Maintenence, analyse data collected and troubleshoot. | Module 1, 2 and 3 are available. Module pending for Xpertcheck calibration. System unit working fine. | Module 2 scheduled for Xpertcheck calibration | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165256343409392940/6165256343409392940_signature_28.png | Jennerose Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6165256343409392940/6165256343409392940_signature_31.png
| 2025-02-28 | 2025-02-27 | Manila Medical Center | U.N. Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | Pull-out | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30395 | CONSORTIUM | bacezar@manilamed.com.ph | 25 |  |  | GXIV-D-6C | 835540 | CPU Problem. | Pull-out. | Checked System's Functionality. Pulled-out CPU for Repair. | CPU OS Problem. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150375341186180/6165150375341186180_signature_28.png | Boylyn A.Cezar | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150375341186180/6165150375341186180_signature_31.png
| 2025-02-28 | 2025-02-27 | Manila Medical Center | U.N. Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30396 | CONSORTIUM | kimxavierpestano@gmail.com | 25 |  |  | GXIV-D-6C | 835540 | CPU Problem. | Repair | Installed Repaired CPU. Generate Reports. Checked System's Functionality. | Recommendation to have UPS (minimum 1kVA). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150375345319704/6165150375345319704_signature_28.png | Kim Xavier Pestano | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150375345319704/6165150375345319704_signature_31.png
| 2025-02-28 | 2025-02-21 | Research Institute for Tropical Medicine | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30392 | MACARE-HBDC | hiv.laboratory.ritm@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1188 |  | GXIV-L-10C | 110017869 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Checked System's Functionality. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325345069876/6165150325345069876_signature_28.png | Nico Antonio Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325345069876/6165150325345069876_signature_31.png
| 2025-02-28 | 2025-02-25 | UST Hospital | Dapitan St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30394 | CONSORTIUM | ustimp@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1070 |  | GXIV-D-6C | 110000711 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked if Filter is Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325345117440/6165150325345117440_signature_28.png | Sherleen Calaunan | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325345117440/6165150325345117440_signature_31.png
| 2025-02-28 | 2025-02-20 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Road, Las Piñas | NCR | NCR 4th - Las Piñas | Molecular | NO | CALIBRATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | uphdmc_lab@yahoo.com |  | 30391 | COMMERCIAL | phmclp.mdl@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 45 |  | GXIV-D-6C | 831316 | Unit is due for XpertCheck. | XpertCheck. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Performed XpertCheck. Checked System's Functionality. Generate Reports. | Modules 1 and 4 Service Required 2024. Modules 2 and 3 Service Required 2025 | A2/ 784660 & A3/ 784664 - Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325345570469/6165150325345570469_signature_28.png | Donn Dave Clerigo | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325345570469/6165150325345570469_signature_31.png
| 2025-02-28 | 2025-02-21 | Dasmariñas City Health Office 2 | DBB-E, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30393 | DOH | robertborilla7@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1605 | 2025-02-19 | GXIV-D-6C | 838535 | Monitor with Output Problem. | Monitor Replacement. | Replaced Monitor. Checked System's Functionality. Generate Reports. | All Modules are Operational. |  | LCD Monitor. | DR 12767 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325346662242/6165150325346662242_signature_28.png | Roberto c. Borilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325346662242/6165150325346662242_signature_31.png
| 2025-02-28 | 2025-02-21 | Research Institute for Tropical Medicine | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30390 | MACARE-HBDC | hiv.laboratory.ritm@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1391 |  | GXIV-L-10C | 110017870 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Checked System's Functionality. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325348786419/6165150325348786419_signature_28.png | Nico Antonio Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/6165150325348786419/6165150325348786419_signature_31.png
| 2025-02-27 | 2025-02-27 | Mandaluyong Social Hygiene Clinic | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30389 | COMMERCIAL | mandaluyongshc.laboratory@gmail.com | 25 |  |  | GXIV-10C | 110025111 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6164391075319991519/6164391075319991519_signature_28.png | Geraldine Fajardo | https://www.jotform.com/uploads/macarewecare/222927978372067/6164391075319991519/6164391075319991519_signature_31.png
| 2025-02-27 | 2025-02-25 | MMC (Medilinx) | Amorsolo st.,Makati | NCR | NCR 4th - Makati | Molecular | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30388 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 30 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 257 ASSAY: Xpert CT/NG, COUNT: 851 ASSAY: Xpert HBV VL, COUNT: 19 ASSAY: Xpert HPV, COUNT: 677 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3586 ASSAY: Others, COUNT: 13 ASSAY: Xpert HCV VL, COUNT: 114 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2109 |  | GX4D(10C) | 110001647 | For relocation from molecular to microbiology | Performed relocation. Reinstalled and setup unit. Cleaned plunger rod,cartridge bay Gathered reports Test run unit | Relocation and reinstallation | Ready to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6164292193615403279/6164292193615403279_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6164292193615403279/6164292193615403279_signature_31.png
| 2025-02-26 | 2025-02-25 | Rizal Medical Center | Pasig Blvd Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | gonzagakevin4@gmail.com |  | 30387 | COMMERCIAL | rmcbloodbank@gmail.com | 24 |  |  | MCS+ | 15H045G | For maintenance | Performed maintenance | Performed maintenance Verified calibration parameters Checked rotor pumps Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6163687153611086071/6163687153611086071_signature_28.png | Ferdinand Barroga | https://www.jotform.com/uploads/macarewecare/222927978372067/6163687153611086071/6163687153611086071_signature_31.png
| 2025-02-26 | 2025-02-25 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30386 | MACARE-HBDC | peachmango318@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 923 |  | GX-IV-D-10C | 110024245 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6163687153612536673/6163687153612536673_signature_28.png | Arlen Delgado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6163687153612536673/6163687153612536673_signature_31.png
| 2025-02-26 | 2025-02-25 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | gonzagakevin4@gmail.com |  | 30385 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 | ASSAY: Others, COUNT: 2 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 242 ASSAY: Xpert SARS-CoV-2, COUNT: 1465 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-XVI-D-6C | 838627 | Loss Communication | Reseat connections | Performed maintenance Checked connections Checked IP address Reseat internal connection Run reliability test Run Xpert MTB-Rif Ultra Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6163687153617419023/6163687153617419023_signature_28.png | Justine Lorenz V. Falzado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6163687153617419023/6163687153617419023_signature_31.png
| 2025-02-26 | 2025-02-26 | Medical Center Manila | U.N. Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30384 | CONSORTIUM | microbiologylab.manilamed@gmail.com | 25 |  |  | GXIV | 835540 | CPU not proceeding to windows. | On-site troubleshooting. | Performed on-site troubleshooting. Resets internal components. Removed and re-insert RAM and SSD. Performed auto repair. Tried running on Windows Safe Mode. | CPU is not operational. CPU is always returning to Advance Repair on every option. CPU for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6163439297111879676/6163439297111879676_signature_28.png | Jericho Cabasug | https://www.jotform.com/uploads/macarewecare/222927978372067/6163439297111879676/6163439297111879676_signature_31.png
| 2025-02-25 | 2025-02-25 | Southern Philippines Medical Center | Bajada, JP Laurel Avenue, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30383 | UNOPS | spmcrhivda@gmail.com | 25.6 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1084 ASSAY: Xpert SARS-CoV-2, COUNT: 5 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 922 |  | GX-IV 10C 4D | 110023505 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6162825177184250129/6162825177184250129_signature_28.png | Flordeliza Go, RMT MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/6162825177184250129/6162825177184250129_signature_31.png
| 2025-02-25 | 2025-02-25 | Southern Philippines Medical Center | JP Laurel Avenue Bajada, Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30382 | UNOPS | spmcrhivda@gmail.com | 24.6 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 120 ASSAY: Xpert SARS-CoV-2, COUNT: 53 |  | GX-IV 10C 4D | 110023506 | Relocation Transfer of GeneXpert Preventive Maintenance | Relocate the GeneXpert System Set from Pop-Up Laboratory to Main Laboratory Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Relocation Transfer of the Machine Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6162823657181488136/6162823657181488136_signature_28.png | Flordeliza Go,. RMT MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/6162823657181488136/6162823657181488136_signature_31.png
| 2025-02-25 | 2025-02-25 | Philipine Blood Center | Quezon Ave., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 30381 | COMMERCIAL | strobecatcher@gmail.com | 26 |  | 2025-02-25 | MCS + | 15H029G | Defective CPU board. | Performed unit calibration/verification. Performed functionality test and simulation run. | Replaced CPU board. | Unit is operational. | NA | CPU board | DR#:12781 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6162783936711839246/6162783936711839246_signature_28.png | Ivan Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6162783936711839246/6162783936711839246_signature_31.png
| 2025-02-25 | 2025-02-25 | Chinese General Hospital | Blumentritt, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Bacteriology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroy.macare@gmail.com | biomed@cghmc.com.ph |  | 30380 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1466 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 103 |  | GXIV-10C | 120002026 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6162782671228499442/6162782671228499442_signature_28.png | Mary Joyce Gaudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6162782671228499442/6162782671228499442_signature_31.png
| 2025-02-25 | 2025-02-25 | Chinese General Hospital | Blumentritt St. Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular pathology laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 30379 | CONSORTIUM | oncologycgh@gmail.com | 25 |  |  | GXIV | 835604 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. Advised enduser to request Xpertcheck from consortium foe upcoming calibration for March 2025. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6162750605719212428/6162750605719212428_signature_28.png | Karen Ringler | https://www.jotform.com/uploads/macarewecare/222927978372067/6162750605719212428/6162750605719212428_signature_31.png
| 2025-02-25 | 2025-02-18 | Argao Primary Care Facility I (RHU II) | Lamacan, Argao Cebu, Poblacion | 7 | Cebu | Tb laboratory | NO | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30378 | DOH | victoria.jumamil@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX4 | 838283 | Installation of the genexpert machines from CARCAR City Health Office. Installation of printer driver HP 105 | Installation of the genexpert machines from CARCAR City Health Office. Installation of printer driver HP 105 | Installation of the genexpert machines from CARCAR City Health Office. Installation of printer driver HP 105 | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6162718835582976766/6162718835582976766_signature_28.png | Victoria Jumamil, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6162718835582976766/6162718835582976766_signature_31.png
| 2025-02-24 | 2025-02-20 | Bagumbayan Rural Health Unit | Poblacion, Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30377 | DOH | bambi.vic6@gmail.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1028 |  | GxIV-6C | 838261 | None | Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6162127926869231540/6162127926869231540_signature_28.png | Videlyne Carlet | https://www.jotform.com/uploads/macarewecare/222927978372067/6162127926869231540/6162127926869231540_signature_31.png
| 2025-02-24 | 2025-02-21 | Naguilian District Hospital | Naguilian, La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30376 | DOH | marjelynrayray01@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1012 |  | GXIV-4-D-6C | 838426 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational - Module B1 is for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6161726577041861384/6161726577041861384_signature_28.png | Marjelyn Rayray | https://www.jotform.com/uploads/macarewecare/222927978372067/6161726577041861384/6161726577041861384_signature_31.png
| 2025-02-24 | 2025-02-20 | Mangatarem Rural Health Unit | Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30375 | DOH | slasheiji1988new@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 518 |  | GXIV-4-D-6C | 838284 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6161724577048977873/6161724577048977873_signature_28.png | Ken Lee Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/6161724577048977873/6161724577048977873_signature_31.png
| 2025-02-24 | 2025-02-19 | Dagupan City Health Office | Dagupan City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30374 | DOH | thisisformyfamily@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1435 |  | GXIV-4-D-6C | 838286 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6161720587045092313/6161720587045092313_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/6161720587045092313/6161720587045092313_signature_31.png
| 2025-02-24 | 2025-02-19 | Mangaldan Rural Health Unit | Mangaldan,Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30373 | DOH | arleneabad32@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 630 |  | GXIV-4-D-6C | 838288 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6161718907045466057/6161718907045466057_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/6161718907045466057/6161718907045466057_signature_31.png
| 2025-02-24 | 2025-02-19 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30372 | DOH | rhupozorrubio@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 |  | GXIV-4-D-6C | 838467 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6161717447044696439/6161717447044696439_signature_28.png | Jinky Grace Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/6161717447044696439/6161717447044696439_signature_31.png
| 2025-02-24 | 2025-02-21 | Metro Antipolo Hospital and Medical Center | Marcos Highway Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30371 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 19 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 284 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 | 2025-02-17 | GX-IV-D-6C | 110005265 | For additional troubleshooting | Run Xpert MTB-RIF Ultra | Performed maintenance Run Xpert MTB-Rif Ultra cartridges Generated reports Generated gxx file Functionality test | Unit is operational. All modules operational. |  |  | DR# 12623, 12752 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6161703163616216421/6161703163616216421_signature_28.png | Kia Mae Kabingue | https://www.jotform.com/uploads/macarewecare/222927978372067/6161703163616216421/6161703163616216421_signature_31.png
| 2025-02-24 | 2025-02-21 | De los Santos Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed@dlsmc.ph |  | 30370 | CONSORTIUM | nspathology@dlsmc.ph | 23 | ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 35 ASSAY: Xpert HPV, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 376 |  | GX4D | 110020700 | Due for Maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6161665653613060342/6161665653613060342_signature_28.png | De Los Santos Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6161665653613060342/6161665653613060342_signature_31.png
| 2025-02-24 | 2025-02-21 | Best Diagnostic Corp. | Masikap ext. QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | bestdx@yahoo.com |  | 30369 | COMMERCIAL | bestdiagnosticlaboratory@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 37 ASSAY: Xpert HBV VL, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GX4D | 110008529 | Due for maintenance Due for xpertcheck Last xpertcheck feb 7,2024 B4 encounter 2097 error hbv | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6161665293616900826/6161665293616900826_signature_28.png | Best Diagnostic Corp. | https://www.jotform.com/uploads/macarewecare/222927978372067/6161665293616900826/6161665293616900826_signature_31.png
| 2025-02-24 | 2025-02-20 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Apheresis room | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30368 | MACARE-HBDC | michaeliza1205@gmail.com | 26 |  |  | MCS + | 17H632SPG | Loose left IV pole | Checked IV pole Retighten screw Aligned spring stopper | Troubleshooting and repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6161664713611734262/6161664713611734262_signature_28.png | Rendelle Pacis | https://www.jotform.com/uploads/macarewecare/222927978372067/6161664713611734262/6161664713611734262_signature_31.png
| 2025-02-22 | 2025-02-21 | Ragay Rural Health Unit | Poblacion-Ilawod, Ragay | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30367 | DOH | phylgomez0115@gmail.com | 26 |  |  | GX4-6C | 838560 | Gx for Preventive Maintenance UPS automatically shuts down during power outage | Cleaned cartridge bays, plunger rod, icore Check UPS cable connections(power cord) Tested the lenght of UPS backup battery when pulled off from power source( 1hr charged) Generate Data | Performed Preventive Maintenance Perfect Troubleshooting (UPS) | Unit is Operational | UPS battery defective(shuts down when pulled from power source). Request for UPS replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6159700261112172830/6159700261112172830_signature_28.png | Phyl Frances C. Gomez | https://www.jotform.com/uploads/macarewecare/222927978372067/6159700261112172830/6159700261112172830_signature_31.png
| 2025-02-22 | 2025-02-20 | San Jose Municipal Health Office | San Jose, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30366 | DOH | jhuneric2008@yahoo.com | 28 |  |  | GX4-6C | 838389 | CPU restarting | Check/swapped CPU power cord Cleaned CPU RAM tested the performance of CPU(1hr) Generate Data | Performed Troubleshooting for CPU | Unit is non operational (CPU keeps restarting) | CPU restarts in less than 5mins. Request for CPU replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6159699591114437671/6159699591114437671_signature_28.png | Rosmar V. Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6159699591114437671/6159699591114437671_signature_31.png
| 2025-02-21 | 2025-02-21 | PHILIPPINE RED CROSS - ILOILO | Iloilo City | 6 | Iloilo | BLOOD BANK | NO | MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30365 | COMMERCIAL | wvrbc@redcross.org.ph | 24 |  |  | MCS+ | 98G049 | FOR PREVENTIVE MAINTENANCE | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -REPLACE ANTICOAGULANT MOTOR RUBBER ROLLER (2 RUBBER ROLLERS USED) -PERFORM CYCLE TEST, PASSED | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6159240558024441283/6159240558024441283_signature_28.png | Josette Saltin, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6159240558024441283/6159240558024441283_signature_31.png
| 2025-02-21 | 2025-02-21 | Don Jose S. Monfort Medical Center | Barotac Nuevo, Iloilo | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30364 | DOH | judyboglosa35@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1898 |  | GX-IV | 838501 | MODULE B2 ERROR 4009 AND 1002 | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6159170113716575337/6159170113716575337_signature_28.png | Judy boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6159170113716575337/6159170113716575337_signature_31.png
| 2025-02-21 | 2025-02-18 | Pasig Treatment Hub | Dr. Garcia St. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30362 | MACARE-HBDC | pathmedtech@gmail.com | 24 |  | 2025-02-18 | GX-IV-D-10C | 110025128 | CPU for replacement | Replaced CPU | Replaced CPU Performed maintenance Generated reports Functionality test | Unit is operational. All modules operational. |  | Old CPU: CPH0424F51370 New CPU: CPH0424F51326 | DR# 12753 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6159101233611446832/6159101233611446832_signature_28.png | Jhozza Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/6159101233611446832/6159101233611446832_signature_31.png
| 2025-02-21 | 2025-02-18 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30363 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 57 ASSAY: Xpert CT/NG, COUNT: 53 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 633 ASSAY: Xpert HPV, COUNT: 417 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6244 ASSAY: Xpert MTB-XDR, COUNT: 1 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 120 ASSAY: Xpert HIV-1 Viral Load, COUNT: 854 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1308 | 2025-02-17 | GX-XVI-D-10C | 110015767 | Error 5007 | Run Xpert MTB-RIF Ultra | Performed maintenance Run Xpert MTB-Rif Ultra on modules A1, A3, A4, B2, B3, C1, C3, D1 and D2 Generated reports Generated gxx file | Unit is operational. All modules operational. | Used 9 Xpert MTB-Rif Ultra cartridges only. |  | DR# 12754 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6159101233616209308/6159101233616209308_signature_28.png | Xchosa Kay C. Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/6159101233616209308/6159101233616209308_signature_31.png
| 2025-02-21 | 2025-02-20 | Philippine Red Cross - Laguna | Sta. Cruz, Laguna | 4A | Laguna | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30361 | COMMERCIAL | lagunablood@redcross.org.ph | 24 |  |  | MCS+ | 02J046 | For calibration | Performed calibration | Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. UPS has no backup. | UPS SN: 20163711301 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6159101233616214045/6159101233616214045_signature_28.png | Josefina S. Bangay | https://www.jotform.com/uploads/macarewecare/222927978372067/6159101233616214045/6159101233616214045_signature_31.png
| 2025-02-20 | 2025-02-20 | MIAGAO RURAL HEALTH UNIT | Miagao, Iloilo | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30360 | UNOPS | joannaleesadiong@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 139 | 2025-01-27 | GX-IV | 110023531 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B1 REPLACEMENT: DEF SN:210122315 NEW SN: 210047524 | EQUIPMENT IS OPERATIONAL |  | Certified pre-owned Genexpert 10-Color Module | 12603 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6158532965605725499/6158532965605725499_signature_28.png | Joanna Lee M Sadiong | https://www.jotform.com/uploads/macarewecare/222927978372067/6158532965605725499/6158532965605725499_signature_31.png
| 2025-02-20 | 2025-02-20 | MIAGAO RURAL HEALTH UNIT | Miagao, Iloilo | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30359 | UNOPS | joannaleesadiong@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 139 | 2025-01-27 | GX-IV | 110023531 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B1 REPLACEMENT: DEF SN:210122315 NEW SN: 210047524 | EQUIPMENT IS OPERATIONAL |  | Certified pre-owned Genexpert 10-Color Module | 12603 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6158532865601118274/6158532865601118274_signature_28.png | Joanna Lee M Sadiong | https://www.jotform.com/uploads/macarewecare/222927978372067/6158532865601118274/6158532865601118274_signature_31.png
| 2025-02-20 | 2025-02-20 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30358 | DOH | bmccrcl2022@gmail.com | 29 |  |  | GX-4 6 Colors | 838403 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6158433030294918825/6158433030294918825_signature_28.png | Joechita Doysabas | https://www.jotform.com/uploads/macarewecare/222927978372067/6158433030294918825/6158433030294918825_signature_31.png
| 2025-02-20 | 2025-02-19 | Talacogon District Hospital | San Agustin, Talacogon | 13 | Agusan del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30357 | DOH | shyratortor@gmail.com | 27 |  |  | GX-4 6 Colors | 838275 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6158429330294997054/6158429330294997054_signature_28.png | Shyra Ericka T. Calamba | https://www.jotform.com/uploads/macarewecare/222927978372067/6158429330294997054/6158429330294997054_signature_31.png
| 2025-02-20 | 2025-02-20 | Tigaon Municipal Health Office | Gingaroy, Tigaon | 5 | Albay | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30356 | DOH | jingimperial24@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 839 |  | GX4-6C | 838558 | B4/614067 not detected, steady red light indicator | Used different cable from B4 to gateway board Swapped B3 and B4 cable connection going to gateway board port Cleaned module 4 port and gateway board port | Performed Troubleshooting | Unit is operational | B4/614067 not detected |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6158390917321380251/6158390917321380251_signature_28.png | Maria Jesusa Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6158390917321380251/6158390917321380251_signature_31.png
| 2025-02-20 | 2025-02-20 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30355 | DOH | lornaabaigar@gmail.com | 27 |  | 2025-02-20 | GXIV | 838409 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6158381385529009052/6158381385529009052_signature_28.png | Lorna C. Abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/6158381385529009052/6158381385529009052_signature_31.png
| 2025-02-20 | 2025-02-20 | Manila Central University | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | lab_mcuhospital@yahoo.com |  | 30354 | CONSORTIUM | purchasing@mcufdtmf.org | 25 |  |  | GXIV-10C | 110018008 | Unit for maintenance | NA | #NAME? | Unit is operational Unit calibration due date (March 6, 2025) XpertCheck Kit required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6158350506432495366/6158350506432495366_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6158350506432495366/6158350506432495366_signature_31.png
| 2025-02-20 | 2025-02-20 | Tubungan Rural Health Unit | Tubungan, Iloilo | 6 | Iloilo | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30353 | UNOPS | tubunganpcf.laboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 48 | 2025-02-11 | GX IV | 110023534 | FOR MODULE B2 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B2 REPLACEMENT: DEF SN:210121941 NEW SN:210099472 | EQUIPMENT IS OPERATIONAL |  | Certified Pre-owned Genexpert 10-Color Module | 12741 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6158253463169781603/6158253463169781603_signature_28.png | Ma. Delmar Tagabi | https://www.jotform.com/uploads/macarewecare/222927978372067/6158253463169781603/6158253463169781603_signature_31.png
| 2025-02-19 | 2025-02-18 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30352 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 57 |  | GX16D | 804476 | All module encounter loss communication error due outdated operating system.B4 encounter temp error 4012,4010,A3 882239,C1 881354,C3 760518,C4 882137 error loss communication,C2 882120 error loss communication and temp error 4012.,Gateway board 6001161 error loss communication. | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance and troubleshooting | Unable to upload other assay definition file for processing of sample bec.of workstation operating system outdated. Workstation unable to upgrade operating system windows 7 already discontinued and no support from microsoft. Unit installation date 4/16/2014 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157341743611622777/6157341743611622777_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6157341743611622777/6157341743611622777_signature_31.png
| 2025-02-19 | 2025-02-18 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30351 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 57 |  | GX16D | 804476 | All module encounter loss communication error due outdated operating system.B4 encounter temp error 4012,4010,A3 882239,C1 881354,C3 760518,C4 882137 error loss communication,C2 882120 error loss communication and temp error 4012.,Gateway board 6001161 error loss communication. | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance and troubleshooting | Unable to upload other assay definition file for processing of sample bec.of workstation operating system outdated. Workstation unable to upgrade operating system windows 7 already discontinued and no support from microsoft. Unit installation date 4/16/2014 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157341743611622777/6157341743611622777_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6157341743611622777/6157341743611622777_signature_31.png
| 2025-02-19 | 2025-02-19 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30350 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 807956 | Unit for maintenance Modules for check up | NA | #NAME? | Unit is operational Module A1(608871) for follow up report |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157552706432151529/6157552706432151529_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6157552706432151529/6157552706432151529_signature_31.png
| 2025-02-19 | 2025-02-19 | Manila Doctor's Hospital | U.N. Ave, Paco Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | xcaguas@maniladoctorscomph.onmicrosoft.com |  | 30349 | CONSORTIUM | rfbuentiempo@maniladoctors.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 73 ASSAY: Xpert HBV VL, COUNT: 111 ASSAY: Xpert HPV, COUNT: 25 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1200 ASSAY: Xpert HCV VL, COUNT: 18 ASSAY: Xpert HIV-1 Viral Load, COUNT: 110 |  | GXIV-10C | 110001648 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157551002023639711/6157551002023639711_signature_28.png | Kaye Martelino | https://www.jotform.com/uploads/macarewecare/222927978372067/6157551002023639711/6157551002023639711_signature_31.png
| 2025-02-19 | 2025-02-19 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30348 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 807112 | Unit for maintenance Modules for check up | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157549276434006203/6157549276434006203_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6157549276434006203/6157549276434006203_signature_31.png
| 2025-02-19 | 2025-02-19 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30347 | COMMERCIAL | laboratory@loveyourself.com.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 0 |  | GXIV-10C | 120002891 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157548786438547405/6157548786438547405_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6157548786438547405/6157548786438547405_signature_31.png
| 2025-02-19 | 2025-02-19 | Pilar Rural Health Unit | Calongay, Pilar | 5 | Sorsogon | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30346 | DOH | manalangmark88@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1501 |  | GX4-6C | 838266 | Monitor issue - auto-adjustment in progress, display stretching | Check VGA cable connection from monitor going to CPU use different cable on monitor Collect Data | Performed Troubleshooting | Unit is Operational | Monitor issue resolved |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157525522001248754/6157525522001248754_signature_28.png | Ma.Veronica Carrascal | https://www.jotform.com/uploads/macarewecare/222927978372067/6157525522001248754/6157525522001248754_signature_31.png
| 2025-02-19 | 2025-02-19 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30345 | DOH | buhanginlaboratory@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1414 |  | GX-IV 6C 4D | 838232 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157505704618734440/6157505704618734440_signature_28.png | Mary Grace Layson | https://www.jotform.com/uploads/macarewecare/222927978372067/6157505704618734440/6157505704618734440_signature_31.png
| 2025-02-19 | 2025-02-19 | Medical Center Manila | U.N. Avenue, Paco Manila | NCR | NCR 1st - Manila | Microbiology Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30344 | CONSORTIUM | microbiologylab.manilamed@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1254 |  | GXIV | 835540 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157492037716346758/6157492037716346758_signature_28.png | Jericho Cabasug | https://www.jotform.com/uploads/macarewecare/222927978372067/6157492037716346758/6157492037716346758_signature_31.png
| 2025-02-19 | 2025-02-07 | University of Perpetual Helps Dalta Medical Center | Alabang-Zapote Rd., Las Piñas | NCR | NCR 4th - Las Piñas | Molecular | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | uphdmc_lab@yahoo.com |  | 30342 | COMMERCIAL | phmclp.mdl@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 27 |  | GXIV-D-6C | 831316 | Barcode Scanning Error. | Troubleshoot. | Checkes Barcode's Functionality. Reinstall Assay. Import Lot. Generate Reports. | Barcode is Operational. 1 Module Disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157350743611765993/6157350743611765993_signature_28.png | Donn Dave Clerigo | https://www.jotform.com/uploads/macarewecare/222927978372067/6157350743611765993/6157350743611765993_signature_31.png
| 2025-02-19 | 2025-02-12 | San Lazaro Hospital | Quiricada St, Manila | NCR | NCR 1st - Manila | SACCL | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30343 | MACARE-HBDC | robethsonfp@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 320 ASSAY: Xpert HIV-1 Viral Load, COUNT: 232 |  | GXIV-D-6C | 821257 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | 1 Module is disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157350743612329583/6157350743612329583_signature_28.png | Robethson Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/6157350743612329583/6157350743612329583_signature_31.png
| 2025-02-19 | 2025-02-14 | General Trias Social Hygiene Clinic | Hospital Road, General Trias, Cavite | 4A | Cavite | Social Hygiene Clinic | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30341 | MACARE-HBDC | chogentri.shc@gmail.com | 26 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 802 |  | GXIV-D-10C | 110016785 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Check Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157350743614040207/6157350743614040207_signature_28.png | John Joseph Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6157350743614040207/6157350743614040207_signature_31.png
| 2025-02-19 | 2025-02-18 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30340 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1584 ASSAY: Xpert CT/NG, COUNT: 14 ASSAY: Others, COUNT: 117 ASSAY: Xpert HBV VL, COUNT: 1004 ASSAY: Xpert HPV, COUNT: 152 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2498 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 4942 ASSAY: Xpert HCV VL, COUNT: 164 |  | GX16D | 110005587 | Due for maintenance A1,A2,C2,D2 reported already for replacement C4 encounter intermittent temp error 2014,1001 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | A1,A2,C2,D2 disabled C4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157342613612808092/6157342613612808092_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6157342613612808092/6157342613612808092_signature_31.png
| 2025-02-19 | 2025-02-18 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30339 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 100 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 34 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 116 ASSAY: Xpert HCV VL, COUNT: 9 |  | GX4D | 838524 | Due for maintenance B4 disabled encounter temp error 1002 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | B1,B2,B3 operational B4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157342193615998933/6157342193615998933_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6157342193615998933/6157342193615998933_signature_31.png
| 2025-02-19 | 2025-02-18 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30338 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 57 |  | GX16D | 804476 | All module encounter loss communication error due outdated operating system B4 encounter temp error 4012,4010 C2 encounter temp error 4012 Due for maintenance C1 881354 error loss communication C2 882120 error loss communication and temp error 4012 C3 760518 error loss communication C4 882137 error loss communication. Gateway board 6001161 error loss communication | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance and troubleshooting | Unable to upload other assay definition file for processing of sample bec.of workstation operating system outdated Workstation unable to upgrade operating system windows 7 already discontinued and no support from microsoft. Unit installation date 4/16/2014 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6157341743611622777/6157341743611622777_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6157341743611622777/6157341743611622777_signature_31.png
| 2025-02-18 | 2025-02-18 | Bago City Health Office | Bago City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30337 | DOH | swabteambago@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2970 |  | GX4 | 838486 | None. | Machine cleaning. | Performed Preventive Maintenance - cleaned filter, optic slots, plunger rods. module bays; archive tests, January 2025 to Feb. 18, 2025 only. | 4 modules operational; no available cartridges as of this day. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6156674580411671524/6156674580411671524_signature_28.png | Glenly Casona | https://www.jotform.com/uploads/macarewecare/222927978372067/6156674580411671524/6156674580411671524_signature_31.png
| 2025-02-18 | 2025-02-18 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30336 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 01J038 | Unit for check up | NA | #NAME? | Unitis operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6156594101998627280/6156594101998627280_signature_28.png | Glaiza Chavez | https://www.jotform.com/uploads/macarewecare/222927978372067/6156594101998627280/6156594101998627280_signature_31.png
| 2025-02-18 | 2025-02-18 | Tuburan rural health unit | Tuburan basilan | BARMM | Basilan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING Monitor replacement and Preventive Maintenence | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30335 | DOH | rjakilan20@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 450 | 2025-01-09 | GX 4 | 838832 | Monitor LCD problem | Replaced monitor and conducted Preventive Maintenence | Replaced monitor and conducted Preventive Maintenence | All modules are available and ready to use | None | FC-E-MONITOR | 12549 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6156560820181489598/6156560820181489598_signature_28.png | Raiyana Jakilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6156560820181489598/6156560820181489598_signature_31.png
| 2025-02-18 | 2025-02-14 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30334 | DOH | suzzettekho3@gmail.com | 27 |  |  | GX4 6 Colors | 838554 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6156465902244704333/6156465902244704333_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/6156465902244704333/6156465902244704333_signature_31.png
| 2025-02-17 | 2025-02-11 | Malungon Rural Health Unit | Poblacion Malungon, Sarangani | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30333 | DOH | mjdetocil@gmail.com | 28.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 865 |  | GxIV-6C | 838546 | None | Perform Genexpert machine Cleaning and maintenance Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6156103066864765740/6156103066864765740_signature_28.png | Michael Jay detocil | https://www.jotform.com/uploads/macarewecare/222927978372067/6156103066864765740/6156103066864765740_signature_31.png
| 2025-02-17 | 2025-02-12 | Sarangani Health Care Facility | Alabel, Sarangani Province | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30332 | DOH | ronayadao@yahoo.com | 22.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 771 ASSAY: Xpert HIV-1 Viral Load, COUNT: 11 |  | GxIV-6C | 838542 | None | Perform Genexpert Cleaning and maintenance Checking of errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6156102246863983977/6156102246863983977_signature_28.png | RONORA YADAO | https://www.jotform.com/uploads/macarewecare/222927978372067/6156102246863983977/6156102246863983977_signature_31.png
| 2025-02-17 | 2025-02-12 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30331 | DOH | vsmmcbloodservices@gmail.com | 22 |  |  | MCS+ | 02A090 | Donor Pressure Monitor encountered low pressure during start up. | Replacement of motor pump blood and transfer pump | Replacement of motor pump blood and transfer pump | Units is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155966555582926221/6155966555582926221_signature_28.png | Charlotte Mae Ygoña, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6155966555582926221/6155966555582926221_signature_31.png
| 2025-02-17 | 2025-02-14 | Minglanilla Rural Health Unit | Pide 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30330 | DOH | brendagonzagaparedes@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1408 |  | GX4 | 838512 | Due for preventive maintenance | Performed preventive maintenance clean up machine, Icore, cartridge bay, ultrasonic hour, removed and clean filter Performed database management, archive results, Generate IQR, Sys Log, Assay Statistics | Performed preventive maintenance clean up machine, Icore, cartridge bay, ultrasonic hour, removed and clean filter Performed database management, archive results, Generate IQR, Sys Log, Assay Statistics | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155964355581772254/6155964355581772254_signature_28.png | Brenda Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/6155964355581772254/6155964355581772254_signature_31.png
| 2025-02-17 | 2025-02-12 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30329 | DOH | gbmmbg@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 758 |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155776897048905385/6155776897048905385_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/6155776897048905385/6155776897048905385_signature_31.png
| 2025-02-17 | 2025-02-11 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30328 | DOH | allan.vinluan2013@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 9e3 |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155773947041405325/6155773947041405325_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6155773947041405325/6155773947041405325_signature_31.png
| 2025-02-17 | 2025-02-11 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30327 | DOH | seansimonpaloyo@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1122 |  | GXIV-4-D-6C | 838473 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155772257043002430/6155772257043002430_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6155772257043002430/6155772257043002430_signature_31.png
| 2025-02-17 | 2025-02-17 | Sta Cruz Rural Health Unit | Poblacion, Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30326 | DOH | chendalagan72@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 516 |  | GX-IV 6C 4D | 838505 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155723929425954767/6155723929425954767_signature_28.png | Dorothy Dalagan RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6155723929425954767/6155723929425954767_signature_31.png
| 2025-02-17 | 2025-02-13 | Rosario Super Health Center | Dr. Sixto Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30325 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 975 |  | GX-IV-D-10C | 110024246 | For relocation | Relocated the GeneXpert System Performed maintenance Generated reports | Transferred the GeneXpert to other room | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627283611297149/6155627283611297149_signature_28.png | Eden R. Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627283611297149/6155627283611297149_signature_31.png
| 2025-02-17 | 2025-02-14 | Klinika Eastwood | CityWalk, Eastwood, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30324 | MACARE-HBDC | klinikaestwoodlaboratory@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 |  | GX-IV-D-10C | 110020698 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627283613456353/6155627283613456353_signature_28.png | Mary Claire P Baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627283613456353/6155627283613456353_signature_31.png
| 2025-02-17 | 2025-02-14 | Santolan Super Health Center | Evangelista Ave Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30323 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1119 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627283616264959/6155627283616264959_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627283616264959/6155627283616264959_signature_31.png
| 2025-02-17 | 2025-02-13 | Antipolo City Health Office | M. Santos St Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30319 | UNOPS | acho.laboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert MTB-XDR, COUNT: 34 ASSAY: Xpert HIV-1 Viral Load, COUNT: 247 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233612167802/6155627233612167802_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233612167802/6155627233612167802_signature_31.png
| 2025-02-17 | 2025-02-12 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30321 | COMMERCIAL | bloodbank@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Run simulation Verified calibration parameters Functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233613737243/6155627233613737243_signature_28.png | Janine Chelsie Uy | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233613737243/6155627233613737243_signature_31.png
| 2025-02-17 | 2025-02-12 | The Medical city | Ortigas Ave Pasig City | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30320 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 57 ASSAY: Xpert CT/NG, COUNT: 57 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 635 ASSAY: Xpert HPV, COUNT: 418 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6126 ASSAY: Xpert MTB-XDR, COUNT: 1 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 114 ASSAY: Xpert HIV-1 Viral Load, COUNT: 843 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1285 | 2025-02-11 | GX-XVI-D-10C | 110015767 | Error 5007 | Replaced module C4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports Generated.nxx file Replaced module C4 | Unit is operational. Modules with 5007 error under observation. |  | Old: C4/210082296 New: C4/210026854 | DR# 12742 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233614300468/6155627233614300468_signature_28.png | Xchosa Kay C. Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233614300468/6155627233614300468_signature_31.png
| 2025-02-17 | 2025-02-12 | The Medical City | Ortigas Ave Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30322 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233617125465/6155627233617125465_signature_28.png | Gilian dacanay | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233617125465/6155627233617125465_signature_31.png
| 2025-02-17 | 2025-02-12 | Metro Antipolo Hospital and Medical Center | Marcos Highway Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30318 | COMMERCIAL | mahmci_labcmt@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 20 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 282 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 |  | GX-IV-D-6C | 110005265 | Error 5007 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated .gxx file Generated reports | Unit is operational. Modules with error 5007 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233617398109/6155627233617398109_signature_28.png | Graziela Marie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/6155627233617398109/6155627233617398109_signature_31.png
| 2025-02-17 | 2025-02-13 | Batasan Social Hygiene Clinic | Sandigan QC | NCR | NCR 2nd - Quezon City | Lab | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30317 | MACARE-HBDC | moniqueteves99@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 28 ASSAY: Xpert HBV VL, COUNT: 39 ASSAY: Xpert HIV-1 Viral Load, COUNT: 749 |  | Gx4D 10C | 110020697 | As per end user they encounter undetected module | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Taught how to use basic operation of the machine Gathered reports Test run unit | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155606423615107466/6155606423615107466_signature_28.png | Monique Teves | https://www.jotform.com/uploads/macarewecare/222927978372067/6155606423615107466/6155606423615107466_signature_31.png
| 2025-02-16 | 2025-02-11 | Cong. Semion G. Toribio Memorial Hospital | Carmen, Bohol | 7 | Bohol | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30316 | DOH | himaya.vale@yahoo.com | 29 |  |  | Gx4 | 838394 | Occasionally monitor blinking | Check up AVG cable Check up power cable | Common hp monitor problem. Instruct End-User the solution if the error pop-up again | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6155285669614269794/6155285669614269794_signature_28.png | Himaya Vale | https://www.jotform.com/uploads/macarewecare/222927978372067/6155285669614269794/6155285669614269794_signature_31.png
| 2025-02-14 | 2025-02-14 | Lapulapu City Health Office | Lapulapu City Health Office | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30315 | DOH | tanya_pax@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2388 |  | GX-IV | 837661 | For Preventive Maintenance | #NAME? | #NAME? | Equipment is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6153203133671789078/6153203133671789078_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/6153203133671789078/6153203133671789078_signature_31.png
| 2025-02-14 | 2025-02-14 | Vinzons Rural Health Unit | Poblacion, Vinzons | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30314 | DOH | bardonamabel@gmail.com | 301 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 33 |  | GX4-6C | 838576 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6153199369192076249/6153199369192076249_signature_28.png | Amabel A. Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/6153199369192076249/6153199369192076249_signature_31.png
| 2025-02-14 | 2025-02-14 | Labo District Hospital | Talobatib, Labo | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30313 | DOH | ldh.lab17@gmail.com | 24 |  |  | GX4-6C | 838207 | GX for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Replaced Filter Generate Data | Performed Preventive Maintenance Replaced Filter | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6153196149193114446/6153196149193114446_signature_28.png | Analyn C. Consuelo | https://www.jotform.com/uploads/macarewecare/222927978372067/6153196149193114446/6153196149193114446_signature_31.png
| 2025-02-14 | 2025-02-14 | Bulacan Medical Center | Malol9s, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30312 | UNOPS | bmclab_micro@yahoo.com | 25 |  | 2025-02-11 | GXIV-10C | 110023507 | Module B3 for replacement | Replaced module B3 | #NAME? | Unit is operational |  | GeneXpert Module SN: 210031653 | 12749 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6153145437125623221/6153145437125623221_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6153145437125623221/6153145437125623221_signature_31.png
| 2025-02-13 | 2025-02-13 | Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30311 | DOH | miniforestlaboratory@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 345 |  | GX-IV 6C 4D | 838507 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bays, Plunger Rods, Fan Filter, and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6152336989027202268/6152336989027202268_signature_28.png | Lolijell Torralba | https://www.jotform.com/uploads/macarewecare/222927978372067/6152336989027202268/6152336989027202268_signature_31.png
| 2025-02-13 | 2025-02-13 | San Jose Municipal Health Unit | San Jose | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30310 | DOH | jhuneric2008@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1942 |  | GX4-6C | 838389 | CPU restarts randomly | Check CPU power cord connection Unplugged and replugged power cords going to UPS Cleaned RAM Removed dust inside CPU cleaned cartridge bays,plunger rod Generate Data | Performed Troubleshooting Performed Preventive Maintenance | Unit is operational | CPU is for observation during Gx test |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6152268556616455514/6152268556616455514_signature_28.png | Rosmar V. Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6152268556616455514/6152268556616455514_signature_31.png
| 2025-02-13 | 2025-02-12 | Tigaon Municipal Health Office | Tigaon | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30309 | DOH | jingimperial24@gmail.com | 26 |  |  | GX4-6C | 838558 | B4/614067 stuck cartridge | B4/614067 stuck cartridge resolved Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance Performed Troubleshooting for module B4 | Unit is Operational | B4/614067 not detected, request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6152238876614500855/6152238876614500855_signature_28.png | Maria Jesusa p.Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6152238876614500855/6152238876614500855_signature_31.png
| 2025-02-12 | 2025-02-12 | St Lukes Extension Clinic | Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 30308 | MACARE-HBDC | laboratory@slec.ph | 23 |  |  | Infinity 48s | 809643 | LIS connectivity problems, results not being uploaded to server. | Cleared communication buffer Redjusted host communication settings Observed test results uploading for following batches | Performed troubleshooting | Unit is operational, LIS able to upload results again |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6151605194525968051/6151605194525968051_signature_28.png | Ma. Sarah R. Rodil | https://www.jotform.com/uploads/macarewecare/222927978372067/6151605194525968051/6151605194525968051_signature_31.png
| 2025-02-12 | 2025-02-12 | Biopath QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30307 | COMMERCIAL | ainocentes@biopath.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 33 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 39 ASSAY: Xpert SARS-CoV-2, COUNT: 11 ASSAY: Xpert HCV VL, COUNT: 1 | 2025-02-12 | GX16 D | 839306 | Due for xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed xpertcheck and maintenance | A3 disabled hw failed C2 failed xpertcheck 01/2024 A1,A4,B2,B4,D2 failed xpertcheck 9 modules operational passed in xpertcheck. |  | Xpertcheck | G7772 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6151502814423785754/6151502814423785754_signature_28.png | Amanda inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6151502814423785754/6151502814423785754_signature_31.png
| 2025-02-12 | 2025-02-12 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30306 | DOH | sajecopooh20@gmail.com | 27 |  |  | 6C GXIV-4D | 838555 | Due for PM | -Preventive Maintenance; plunger rod, fan filter, optic slot, cartridge bay | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6151369239432465601/6151369239432465601_signature_28.png | Sarah Jean G. Pasana | https://www.jotform.com/uploads/macarewecare/222927978372067/6151369239432465601/6151369239432465601_signature_31.png
| 2025-02-12 | 2025-02-11 | Makati Medical Center Medilinx | Amorsolo st., Makati City | NCR | NCR 4th - Makati | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | John.Manzano@sodexo.com |  | 30305 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 250 ASSAY: Xpert CT/NG, COUNT: 863 ASSAY: Xpert HBV VL, COUNT: 9 ASSAY: Xpert HPV, COUNT: 681 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3533 ASSAY: Others, COUNT: 13 ASSAY: Xpert HCV VL, COUNT: 110 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2061 |  | GX4 D 10C | 110001647 | Due for maintenance B2,B3 failed xpertcheck reported already for module replacement B1 reported already for module replacement B4 encounter 5017,5007 for reporting | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | B1 disabled B2 ,B3 run other assay except mtb ultra |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6151342373617179978/6151342373617179978_signature_28.png | Aiza G. Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6151342373617179978/6151342373617179978_signature_31.png
| 2025-02-12 | 2025-02-12 | The Medical City - South Luzon | Sta. Rosa Laguna | 4A | Laguna | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30304 | COMMERCIAL | tmcsl-molecularlab@themedicalcity.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 194 ASSAY: Xpert HIV-1 Viral Load, COUNT: 26 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 2 |  | GXIV | 839233 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6151341057421553899/6151341057421553899_signature_28.png | Izza Mae Maroto | https://www.jotform.com/uploads/macarewecare/222927978372067/6151341057421553899/6151341057421553899_signature_31.png
| 2025-02-12 | 2025-02-12 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30303 | DOH | sarah_yangzon@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 682 | 2025-02-12 | GXIV | 838406 | N/a | N/a | #NAME? | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6151327833026723184/6151327833026723184_signature_28.png | Sarah V Yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/6151327833026723184/6151327833026723184_signature_31.png
| 2025-02-11 | 2025-02-04 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30302 | DOH | dianeruiz147@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 455 |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150687365725426319/6150687365725426319_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6150687365725426319/6150687365725426319_signature_31.png
| 2025-02-11 | 2025-02-04 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30301 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 206 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 52 |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is partially operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150684155724969149/6150684155724969149_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6150684155724969149/6150684155724969149_signature_31.png
| 2025-02-11 | 2025-02-04 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30300 | DOH | serly_reyes@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 715 |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150681455724375689/6150681455724375689_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6150681455724375689/6150681455724375689_signature_31.png
| 2025-02-11 | 2025-02-11 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30299 | DOH | mendozarey42@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3916 ASSAY: Xpert HCV VL, COUNT: 104 ASSAY: Xpert HIV-1 Viral Load, COUNT: 486 |  | GX-IV | 838510 | MODULE C1 ERROR 5017 | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | MODULE C1 FOR OBSERVATION |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150636093026737194/6150636093026737194_signature_28.png | Rey C Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6150636093026737194/6150636093026737194_signature_31.png
| 2025-02-11 | 2025-02-11 | Kalinga Health Hub Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30298 | CONSORTIUM | pearl.balderama@innovationsch.org | 25 |  |  | GXIV | 110008415 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150629657124477604/6150629657124477604_signature_28.png | Carl Babasa | https://www.jotform.com/uploads/macarewecare/222927978372067/6150629657124477604/6150629657124477604_signature_31.png
| 2025-02-11 | 2025-02-11 | Kalinga Health Hub Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30297 | CONSORTIUM | pearl.balderama@innovationsch.org | 25 |  |  | GXIV | 110008420 | Unit for maintenance | NA | #NAME? | Unit is operational B4 is under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150628987127622252/6150628987127622252_signature_28.png | Carl Babasa | https://www.jotform.com/uploads/macarewecare/222927978372067/6150628987127622252/6150628987127622252_signature_31.png
| 2025-02-11 | 2025-02-11 | Biopath Clinical Diagnostics | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30296 | COMMERCIAL | jkee@biopath.ph | 23 |  | 2025-02-11 | GX4 | 806201 | GeneXpert Machine Installation. | Save Installation Qualification Report Save System Log File | Install Back up GeneXpert Install Laptop | Module 1,3,4 ok to use. Module 2 not ok to use hardware fail. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150583115844040533/6150583115844040533_signature_28.png | Mary Joy Kee | https://www.jotform.com/uploads/macarewecare/222927978372067/6150583115844040533/6150583115844040533_signature_31.png
| 2025-02-11 | 2025-02-11 | Cabadbaran City Health Office | S.Curato St., Cabadbaran City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30295 | DOH | samstercem@gmail.com | 26 |  |  | GX4 6 colors | 838550 | None | Maintenance | Preventive Maintenance and Data Collection | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150582387314749495/6150582387314749495_signature_28.png | Sammy A. Cempron | https://www.jotform.com/uploads/macarewecare/222927978372067/6150582387314749495/6150582387314749495_signature_31.png
| 2025-02-11 | 2025-02-11 | Catigbian Distric Hospital - RTDL | Catigbian, Bohol | 7 | Bohol | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30294 | DOH | rheranes@gmail.com | 27 |  |  | Gx4 | 838494 | Error 1002 and Error 4012 | Clean Fan Filter | Run xpertcheck calibration Clean Fan Filter Gather data system log, IQR and gxc file | Systems Operational Module 3 Service Require during calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150567991914927330/6150567991914927330_signature_28.png | CHARISSE MAE T. CLAROS | https://www.jotform.com/uploads/macarewecare/222927978372067/6150567991914927330/6150567991914927330_signature_31.png
| 2025-02-11 | 2025-02-11 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma2020@gmail.com |  |  | 30293 | DOH | luzvimindanavidad2@gmail.com | 27 |  |  | 6C GXIV-4D | 838297 | Due for PM | -Preventive Maintenance; fan filter, plunger rod, cartridge bay, optic slot | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150508100895689883/6150508100895689883_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6150508100895689883/6150508100895689883_signature_31.png
| 2025-02-10 | 2025-02-10 | Lake Sebu Rural Health Unit | Purok Rosas, Brgy. Poblacion, Lake Sebu,South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30292 | DOH | jennypingkian23@yahoo.com | 28.9 |  |  | GxIV-6C | 838569 | None | Perform Genexpert Cleaning and maintenance Checking of errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150097866866088793/6150097866866088793_signature_28.png | Jenny A. Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/6150097866866088793/6150097866866088793_signature_31.png
| 2025-02-10 | 2025-02-05 | Norala District Hospital | Poblacion Norala, South Cotabato | 12 | South Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30291 | DOH | noralaDHlaboratory@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 958 |  | GxIV-6C | 838543 | None | Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150093096861282611/6150093096861282611_signature_28.png | Reziel D. Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/6150093096861282611/6150093096861282611_signature_31.png
| 2025-02-10 | 2025-02-06 | Isulan Rural Health Unit | Poblacion, Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30290 | UNOPS | isulanrhuclin.lab@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 81 |  | GxIV-10C | 110023520 | None | Perform Genexpert Cleaning and maintenance Viral load Assay installation Checking of errors Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6150090936864095685/6150090936864095685_signature_28.png | February jen Apiado | https://www.jotform.com/uploads/macarewecare/222927978372067/6150090936864095685/6150090936864095685_signature_31.png
| 2025-02-10 | 2025-02-10 | St Lukes Medical Center QC | E Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 30289 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 24 | ASSAY: Xpert HBV VL, COUNT: 234 ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert HPV, COUNT: 456 ASSAY: Xpert MTB/RIF Ultra, COUNT: 23 ASSAY: Xpert HCV VL, COUNT: 24 ASSAY: Xpert HIV-1 Viral Load, COUNT: 82 |  | GX4 D | 807155 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | Operational C1,C2,C3 Disabled C4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149912487123790517/6149912487123790517_signature_28.png | Cha Caibigan | https://www.jotform.com/uploads/macarewecare/222927978372067/6149912487123790517/6149912487123790517_signature_31.png
| 2025-02-10 | 2025-02-10 | St Lukes Medical Center | E Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbilogy | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 30288 | COMMERCIAL | megapo@stlukes.com.ph | 24 | ASSAY: Xpert CT/NG, COUNT: 206 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2937 |  | GX4 D | 839234 | Due for maintenance Last xpertcheck march 2023 As per end user they encounter error 5007 in module 1 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Reset gateway board Gathered reports Test run unit | Performed maintenance | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149912077125261489/6149912077125261489_signature_28.png | Ma Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6149912077125261489/6149912077125261489_signature_31.png
| 2025-02-10 | 2025-02-10 | CEBU DOCTORS UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30287 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 84 ASSAY: Xpert CT/NG, COUNT: 251 ASSAY: Xpert HBV VL, COUNT: 189 ASSAY: Xpert HPV, COUNT: 37 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1131 ASSAY: Xpert HCV VL, COUNT: 152 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GX-IV | 110008524 | FOR PREVENTIVE MAINTENANCE AND CALIBRATION | COLLECT IQR, SYSTEM LOG, ARCHIVE, AND CALIBRATION DATA | #NAME? | EQUIPMENT IS OPERATIONAL RECOMMENDATION: FOR MODULE B3 AND MODULE B4 REPLACEMENT | #NAME? |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149777475587769804/6149777475587769804_signature_28.png | Jennifer Marie Lim, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6149777475587769804/6149777475587769804_signature_31.png
| 2025-02-10 | 2025-02-07 | Marikina City Health Office | Shoe Ave Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 30285 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 480 ASSAY: Xpert MTB-XDR, COUNT: 37 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596303618597640/6149596303618597640_signature_28.png | Caryl I. Abrera | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596303618597640/6149596303618597640_signature_31.png
| 2025-02-10 | 2025-02-07 | Marikina City Health Office | Shoe Ave Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 30286 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1057 ASSAY: Xpert MTB-XDR, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596303619444485/6149596303619444485_signature_28.png | Caryl I. Abrera | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596303619444485/6149596303619444485_signature_31.png
| 2025-02-10 | 2025-02-07 | Marikina City Health Office | Sumulong Highway Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 30281 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 50 ASSAY: Xpert MTB-XDR, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 |  | GX-IV-D-10C | 110024073 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233611230702/6149596233611230702_signature_28.png | Estrelita S. Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233611230702/6149596233611230702_signature_31.png
| 2025-02-10 | 2025-02-07 | Metro Antipolo Hospital and Medical Center | Marcos Highway Antipolo Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30282 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 20 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 274 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233615606708/6149596233615606708_signature_28.png | Graziela Marie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233615606708/6149596233615606708_signature_31.png
| 2025-02-10 | 2025-02-06 | Cardinal Santos Medical Center | Wilson St. San Juan City | NCR | NCR 2nd - San Juan | Microbiology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 30284 | COMMERCIAL | jaywilliamstee@gmail.com | 24 | ASSAY: Xpert HBV VL, COUNT: 147 ASSAY: Xpert HPV, COUNT: 61 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1072 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B3/210066495 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233617174146/6149596233617174146_signature_28.png | Marc Maglipon | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233617174146/6149596233617174146_signature_31.png
| 2025-02-10 | 2025-02-07 | Marikina City Health Office | Sumulong Highway Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 30283 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 71 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110024072 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233618808714/6149596233618808714_signature_28.png | ESTRELITA S. ARAGON | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233618808714/6149596233618808714_signature_31.png
| 2025-02-10 | 2025-02-06 | Cardinal Santos Medical Center | Wilson St. San Juan City | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 30280 | COMMERCIAL | csmclab@csmc.ph | 24 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233619086403/6149596233619086403_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/6149596233619086403/6149596233619086403_signature_31.png
| 2025-02-07 | 2025-02-04 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30279 | MACARE-HBDC | nchmolecularlab2021@gmaip.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2001 ASSAY: Xpert SARS-CoV-2, COUNT: 559 ASSAY: Xpert SARS-CoV-2, COUNT: 1506 | 2025-02-03 | GX16 D | 110014171 | Encounter continous error 5007 Due for maintenance | Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Performed reliability test Performed reagent running Test run unit | Performed maintenance and troubleshooting | A2,A3,A4,B3,C1,C2,D2,D4 passed in reliability test A2,B3 encounter error 5007 twice failed in reagent running A3,A4,C1,C2,D2,D4 passed reagent running mtbrif ultra Disabled A2,B3 Disabled A2,B3 |  | Mtb/rif ultra | 12619 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6147175251114421788/6147175251114421788_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/6147175251114421788/6147175251114421788_signature_31.png
| 2025-02-07 | 2025-02-05 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Apheresis room | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30278 | MACARE-HBDC | strobecatcher@gmail.com | 27 |  |  | MCS + | 15H029G | Noisy vibration from centrifuge Centrifuge bearing wornout | Swap centrifuge assy from 15H028G to 15H029G. Performed calibration Checked program parameters Performed functionality test | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6147173961112251159/6147173961112251159_signature_28.png | Ivan Ritz P. Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6147173961112251159/6147173961112251159_signature_31.png
| 2025-02-07 | 2025-02-04 | National Childrens Hospital | E.R9driguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 30277 | MACARE-HBDC | zarabellehilario@yahoo.com | 28 |  |  | MCS + | 17H632SPG | As per end user they encounter noisy sound from pump rotor. Line sensor malfunction of potentiometer unable to adjust to desired value 3071 line sensor potentionmeter can adjust only to 2838 Left pole loose rod | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | For observation For quotation of Maintenance and Calibration and line sensor |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6147173561113749168/6147173561113749168_signature_28.png | National Childrens Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6147173561113749168/6147173561113749168_signature_31.png
| 2025-02-07 | 2025-02-07 | Philippine Red Cross Olongapo | Olongapo | 1 | Zambales | Blood Bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30276 | MACARE-HBDC | jona.liza@redcross.org.ph | 25 |  | 2025-02-05 | MCS+ | 03A043 | Unit for maintenance Rotor roller rubbers for replacement | Replaced rotor roller rubbers | #NAME? | Unit is operational |  | Rotor Roller Rubbers | 12630 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6147112701126333004/6147112701126333004_signature_28.png | Jona Mae Liza | https://www.jotform.com/uploads/macarewecare/222927978372067/6147112701126333004/6147112701126333004_signature_31.png
| 2025-02-07 | 2025-02-07 | Biopath-AFRIMS | CAP Building Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30275 | COMMERCIAL | jkee@biopath.ph | 23 |  | 2025-02-02 | GX16 | 838657 | Monthly Maintenance. | .Perform preventive maintenance for the GeneXert Machine . Clean plunger rod . Clean cartridge bay . Clean filter at the back of the machine. | . Save Installation Qualification Report . Save System Log file. . Save Assay Statistics Report . Save GXX file | Module A1,A3, B3, D1, D2, ok to use. Module A2, A4, B1, B2, B4, C1, C2, C3, C4, D3, D4 not ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6147079755846427811/6147079755846427811_signature_28.png | Mary Joy Kee | https://www.jotform.com/uploads/macarewecare/222927978372067/6147079755846427811/6147079755846427811_signature_31.png
| 2025-02-07 | 2025-02-06 | Burauen District Hospital | Burauen leyte | 8 | Leyte | TB Laboratory | NO | CALIBRATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30274 | DOH | buraeunlab@gmail.com | 27 |  |  | GXIV | 838551 | N/a | N/a | Monthly Preventive Maintenance | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6147054465254009937/6147054465254009937_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6147054465254009937/6147054465254009937_signature_31.png
| 2025-02-06 | 2025-02-06 | Love Your Slef - Whitehouse | Cebu City | 7 | Cebu | HIV | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30273 | COMMERCIAL | preforzado@loveyourself.ph | 22 |  | 2025-02-06 | GX4 | 814150 | . Bios Battery low battery. . Due for Preventive Maintenance | Replace new Bios Battery | . Save Installation Qualifications Report . Save Assay Statistics Report . Save System Log file | Module 1,2,3 ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6146403185584893413/6146403185584893413_signature_28.png | Paulo Crispo Reforzado | https://www.jotform.com/uploads/macarewecare/222927978372067/6146403185584893413/6146403185584893413_signature_31.png
| 2025-02-06 | 2025-01-24 | Asian Hospital and Medical Center | Civic Drive, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | molecularpathology@asianhospital.com |  | 30269 | COMMERCIAL | agalang@asianhospital.com | 21 | ASSAY: Xpert CT/NG, COUNT: 12 ASSAY: Others, COUNT: 72 ASSAY: Xpert HBV VL, COUNT: 110 ASSAY: Xpert HPV, COUNT: 182 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1089 ASSAY: Xpert HIV-1 Viral Load, COUNT: 88 |  | GXIV-D-10C | 110014996 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Surface Unit, Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality (Monitor, CPU, and UPS). Generate Reports. | GeneXpert System is Good and Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6146308908311384541/6146308908311384541_signature_28.png | Alyssa Galang | https://www.jotform.com/uploads/macarewecare/222927978372067/6146308908311384541/6146308908311384541_signature_31.png
| 2025-02-06 | 2025-01-31 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | Relocation | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30270 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 196 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 16 |  | GX16-D-10C | 110013856 | Unit is due for Relocation. | Relocation. | Relocate GX16 unit from Molecular to Old Molecular. Performed Self-test. Checked System's Functionality. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6146308908314427665/6146308908314427665_signature_28.png | Nelson Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6146308908314427665/6146308908314427665_signature_31.png
| 2025-02-06 | 2025-01-31 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | Relocation | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30272 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 816 |  | GXIV-D-10C | 110014099 | Unit for Relocation. | Relocate. | Transfer GX Unit from New Molecular to Old Molecular. Performed Self-test. Generate Reports. | 1 Module is Disabled. 3 are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6146308908318231052/6146308908318231052_signature_28.png | Nelson Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6146308908318231052/6146308908318231052_signature_31.png
| 2025-02-06 | 2025-01-28 | UST Hospital | Dapitan St, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30271 | MACARE-HBDC | bloodbank@usthospital.com.ph | 25 |  | 2024-12-12 | MCS+ | 02F069 | UPS Problem. Unit is due for Maintenance. | Repair. Preventive Maintenance. | Replaced UPS. Checked Diagnostic And Sensor Values. Checked Pump Rollers. Run Simulation Test. | Unit is Operational. |  | UPS | DR 12427 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6146308908318560993/6146308908318560993_signature_28.png | Mylene Balboa | https://www.jotform.com/uploads/macarewecare/222927978372067/6146308908318560993/6146308908318560993_signature_31.png
| 2025-02-06 | 2025-02-06 | Calbiga Rural Health Unit | Calbiga, Samar | 8 | Samar | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30268 | DOH | lornaabaigar@gmail.com | 25 |  |  | 6C GXIV-4D | 838409 | None | -Preventive Maintenance; cartridge bay, plunger rod, optic slot, fan filter | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6146265339207334238/6146265339207334238_signature_28.png | Lorna Abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/6146265339207334238/6146265339207334238_signature_31.png
| 2025-02-06 | 2025-02-06 | Panabo City Health Office | Poblacion, Panabo City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30267 | DOH | nadineeust67@gmail.com | 22.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2416 |  | GX-IV 6C 4D | 838295 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6146265110128730867/6146265110128730867_signature_28.png | Nadine Gemma S. Eustaquio | https://www.jotform.com/uploads/macarewecare/222927978372067/6146265110128730867/6146265110128730867_signature_31.png
| 2025-02-06 | 2025-02-06 | Claveria Rural Health Unit | Poblacion 1, Claveria | 5 | Masbate | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30266 | DOH | mgrmartinez757@gmail.com | 28 |  |  | GX4-6C | 838572 | Gx for Preventive Maintenance Fan filter request for replacement | Cleaned cartridge bays, plunger rod, icore Replaced fan filter Generate Data | Performed Preventive Maintenance Replaced fan filter | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6146212413127492193/6146212413127492193_signature_28.png | Mary Grace R.Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/6146212413127492193/6146212413127492193_signature_31.png
| 2025-02-06 | 2025-02-06 | Detoxicare Molecular Diagnostics Laboratory Inc | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30265 | CONSORTIUM | joy.bagsic@detoxicare.com.ph | 25 |  |  | GXIV-10C | 110018312 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6146129784121102625/6146129784121102625_signature_28.png | Mary Joy Bagsic | https://www.jotform.com/uploads/macarewecare/222927978372067/6146129784121102625/6146129784121102625_signature_31.png
| 2025-02-05 | 2025-01-28 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30264 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 208 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 52 |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145593358834376263/6145593358834376263_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6145593358834376263/6145593358834376263_signature_31.png
| 2025-02-05 | 2025-02-05 | Pio Duran Rural Health Unit | Caratagan, Pio Duran | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30263 | DOH | mitzijc@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 649 |  | GX4-6C | 838264 | Gx for Preventive Maintenance B1/807820 - blinking red, temperature error | Cleaned cartridge bays, plunger rod, icore, fan filter, Gx fan Collect data | Performed Preventive Maintenance Performed Troubleshooting | Unit is Operational | B1/807820 temperature Error: 1002, 4009. request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145497380518589591/6145497380518589591_signature_28.png | Mitzi D. Cedeño | https://www.jotform.com/uploads/macarewecare/222927978372067/6145497380518589591/6145497380518589591_signature_31.png
| 2025-02-05 | 2025-02-05 | Central Luzon Doctor's Hospital | Hospital Drive Brgy San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 30262 | COMMERCIAL | pjlsantos@cldh.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 |  | GXIV-D | 839232 | B1:742501 ( Temperature Error) B4:738203 ( Temperature Error) | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational | B1, B4 andfor reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145462482424827830/6145462482424827830_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6145462482424827830/6145462482424827830_signature_31.png
| 2025-02-05 | 2025-02-05 | Central Luzon Doctor's Hospital | Hospital drive Brgy San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 30261 | COMMERCIAL | pjlsantos@cldh.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 |  | GXII-D | 830841 | C2:620096 ( temperature error) | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports preventive maintenance | Unit is operational | C2 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145461312423260137/6145461312423260137_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6145461312423260137/6145461312423260137_signature_31.png
| 2025-02-05 | 2025-02-05 | Overseas Filipino Hospital | San Fernando City, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30260 | MACARE-HBDC | ofwhlaboratory@gmail.com | 15 | ASSAY: Xpert HPV, COUNT: 17 |  | GXIV-D | 110018556 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145452012425864660/6145452012425864660_signature_28.png | James Ferdinand S. Bacani | https://www.jotform.com/uploads/macarewecare/222927978372067/6145452012425864660/6145452012425864660_signature_31.png
| 2025-02-05 | 2025-02-05 | Diit Health Center | Diit, Tacloban City, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30259 | DOH | cholaboratory23@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1108 |  | 6C GXIV-4D | 838421 | Due for PM, Module B2 and B3 error 2037 | -Preventive Maintenance -Troubleshoot Error 2037 in module B2 and B3 | -Preventive Maintenance -Troubleshoot Error -Archive Tests from 2023-2025 | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145442656599254173/6145442656599254173_signature_28.png | Dale Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6145442656599254173/6145442656599254173_signature_31.png
| 2025-02-05 | 2025-02-05 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30258 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  |  | GXIV | 110008530 | Modules B1, B4 for additional troubleshooting | NA | #NAME? | Unit is operational Data for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145424444518283012/6145424444518283012_signature_28.png | Margaret Camua | https://www.jotform.com/uploads/macarewecare/222927978372067/6145424444518283012/6145424444518283012_signature_31.png
| 2025-02-05 | 2025-02-05 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30257 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  |  | GXIV | 110008530 | Modules B1, B4 for additional troubleshooting | NA | #NAME? | Unit is operational Data for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145424424517908326/6145424424517908326_signature_28.png | Margaret Camua | https://www.jotform.com/uploads/macarewecare/222927978372067/6145424424517908326/6145424424517908326_signature_31.png
| 2025-02-05 | 2025-02-05 | Tomas Claudio Health Center | J. Camus St, Corner Quirino Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30256 | DOH | arleen_p_alconera@yahoo.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 712 |  | GX-IV 6C 4D | 838492 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145388115312394154/6145388115312394154_signature_28.png | Arleen Alconera,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6145388115312394154/6145388115312394154_signature_31.png
| 2025-02-05 | 2025-02-05 | Bureau of Quarantine | Intramuros Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30255 | MACARE-HBDC | dklo.lab@quarantine.doh.gov.ph | 25 |  |  | GX16 | 844418 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145333346676997289/6145333346676997289_signature_28.png | DULCE KARMELA LO | https://www.jotform.com/uploads/macarewecare/222927978372067/6145333346676997289/6145333346676997289_signature_31.png
| 2025-02-05 | 2025-02-05 | Vicente Mendiola Center for Health | Naga City Cebu | 7 | Cebu | TB Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30254 | DOH | geccaprocianos26@gmail.com | 22 |  | 2025-02-05 | GX4 | 838564 | Due for Monthly Maintenance. | Save GXX file . Save System Log file . Save Assay Statistics Report . Save Installation Qualifications Report | . Clean Filter at the back of the Machine . Clean Plunger rod. . Clean Cartridge bay. . Check UPS if functional. . Check AVR if functional. . Check CPU and LCD if functional. | . All modules are functional ok to use . UPS functional ok to use . AVR functional ok to use . CPU functional ok to use . LCD monitor functional ok to use. . Canon printer functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145320288029247289/6145320288029247289_signature_28.png | Gene Carmel Procianos | https://www.jotform.com/uploads/macarewecare/222927978372067/6145320288029247289/6145320288029247289_signature_31.png
| 2025-02-05 | 2025-02-05 | Jose B. Lingad Memorial and General Hospital | San Fernando City, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30253 | MACARE-HBDC | jblmghmpl@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 68 ASSAY: Xpert MTB/RIF Ultra, COUNT: 550 ASSAY: Xpert MTB-XDR, COUNT: 139 ASSAY: Xpert SARS-CoV-2, COUNT: 692 |  | GXIV-D 10C | 110012825 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unut id operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145273734861028737/6145273734861028737_signature_28.png | Justin Carlos | https://www.jotform.com/uploads/macarewecare/222927978372067/6145273734861028737/6145273734861028737_signature_31.png
| 2025-02-05 | 2025-02-05 | Jose B. Lingad Memorial and General Hospital | San Fernando City, Pamanga | 3 | Pampanga | RHIVDA | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30252 | MACARE-HBDC | jblmgh.is@gmail.com | 31 | ASSAY: Xpert HBV VL, COUNT: 278 ASSAY: Xpert HCV VL, COUNT: 69 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1899 |  | GXIV-D 10C | 110013748 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6145255004863598379/6145255004863598379_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/6145255004863598379/6145255004863598379_signature_31.png
| 2025-02-04 | 2025-02-04 | Victor R. Potenciano Medical Center | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30251 | CONSORTIUM | oro_janet@yahoo.com | 25 |  |  | GXIV | 835600 | Unit for maintenance Module B4 for troubleshooting | NA | #NAME? | Unit is operational B4 under observation | Possible problem on catridge batch |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6144614796818014958/6144614796818014958_signature_28.png | Janet Oro | https://www.jotform.com/uploads/macarewecare/222927978372067/6144614796818014958/6144614796818014958_signature_31.png
| 2025-02-04 | 2025-02-04 | President Ramon Magsaysay Memorial Hospital | Iba, Zambales | 3 | Zambales | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30250 | MACARE-HBDC | prmmhmoleclab@gmail.com | 30 | ASSAY: Xpert SARS-CoV-2, COUNT: 402 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 6 |  | GXIV-D 6C | 838520 | B3:888979 HW Failed Error( 4010, 4012,1002,1001,2022) B4:688212 HW Failed Error(4017) | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except for modules B3 and B4 | B3 and B4 for relacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6144609164382973530/6144609164382973530_signature_28.png | Bianca Yzabel Yap | https://www.jotform.com/uploads/macarewecare/222927978372067/6144609164382973530/6144609164382973530_signature_31.png
| 2025-02-04 | 2025-02-04 | Pilar Rural Health Unit | Pilae | 5 | Sorsogon | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30249 | DOH | manalangmarkramir@gmail.com | 23 |  |  | GX4-6C | 838266 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6144511981818772640/6144511981818772640_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/6144511981818772640/6144511981818772640_signature_31.png
| 2025-02-04 | 2025-02-04 | Philippine Red Cross Olongapo | Olongapo City, Zambales | 3 | Zambales | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30248 | MACARE-HBDC | jona.liza@redcross.org.ph | 25 |  |  | MCS+ | 03A043 | Noise on pumps | Check up rollers Cleaning of rollers Preventive maintenance | Performed surface cleaning Check up and clean Rollers Performed preventive maintenance Performed functionality tes | Unit is operational | Roller for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6144461804716489801/6144461804716489801_signature_28.png | Jona Mae Liza | https://www.jotform.com/uploads/macarewecare/222927978372067/6144461804716489801/6144461804716489801_signature_31.png
| 2025-02-04 | 2025-02-04 | Maimbung District Hospital | Laum Maimbung, sulu | BARMM | Sulu | Laboratory | NO | MAINTENANCE Preventive Maintenence and Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30247 | DOH | alvinmadjid68@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 616 | 2024-11-28 | GX 4 | 838838 | None | Preventive Maintenence, module 863156 and 861972 replaced with Module 652176 and 632084 respectively. | Preventive Maintenence, module 863156 and 861972 replaced with Module 652176 and 632084 respectively. | All modules are available and ready to use | None | 900-0491RS, 900-0491RS | 12393, 12392 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6144422888023236171/6144422888023236171_signature_28.png | ALVIN MADJID | https://www.jotform.com/uploads/macarewecare/222927978372067/6144422888023236171/6144422888023236171_signature_31.png
| 2025-02-03 | 2025-01-31 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila City | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30243 | MACARE-HBDC | fabella.mollabgene@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 74 |  | GX-IV-D-6C | 834953 | For relocation | Relocated the unit. | Relocated the GeneXpert unit Performed maintenance Functionality test Generated reports | Unit is operational. Modules A1, A2 and A3 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143698453611441894/6143698453611441894_signature_28.png | Renee Lyn Robles | https://www.jotform.com/uploads/macarewecare/222927978372067/6143698453611441894/6143698453611441894_signature_31.png
| 2025-02-03 | 2025-01-30 | Pasig Treatment Hub - Manggahan | Kaginhawaan St. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30244 | MACARE-HBDC | pathmedtech@gmail.com | 24 |  |  | GX-IV-L-6C | 821648 | Modules not detected | Configure IP address | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Configure IP address Generated reports | Unit is operational. Modules A2 and A4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143698453612019373/6143698453612019373_signature_28.png | Katherine Godoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6143698453612019373/6143698453612019373_signature_31.png
| 2025-02-03 | 2025-01-30 | Pasig Treatment Hub - CHAMP | Caruncho Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | pathmedtech@gmail.com |  | 30245 | MACARE-HBDC | pasigtreatment@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 160 |  | GX-IV-D-10C | 110025127 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143698453613932160/6143698453613932160_signature_28.png | Pasig Treatment Hub | https://www.jotform.com/uploads/macarewecare/222927978372067/6143698453613932160/6143698453613932160_signature_31.png
| 2025-02-03 | 2025-01-31 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila City | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30246 | MACARE-HBDC | fabella.mollabgene@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 126 ASSAY: Xpert MTB/RIF Ultra, COUNT: 82 ASSAY: Xpert SARS-CoV-2, COUNT: 320 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 591 |  | GX-IV-D-10C | 110015482 | For relocation | Relocated the GeneXpert unit. | Relocated the GeneXpert unit Performed maintenance Functionality test Generated reports | Unit is operational. Module B3 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143698453619022940/6143698453619022940_signature_28.png | Renee Lyn Robles | https://www.jotform.com/uploads/macarewecare/222927978372067/6143698453619022940/6143698453619022940_signature_31.png
| 2025-02-03 | 2025-02-03 | Cebu City Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30242 | DOH | ccmclaboratory01@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1863 |  | 6C GXIV-4D | 838561 | Due for PM | -Preventive Maintenance; fan filter, cartridge bay, optic slot, plunger rod | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143696602422058889/6143696602422058889_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/6143696602422058889/6143696602422058889_signature_31.png
| 2025-02-03 | 2025-02-03 | Macare (In-house Unit) | J. Llorente, Cebu City | 7 | Cebu | Inhouse | NO | REPAIR | HAEMONETICS | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30241 | COMMERCIAL | engineering@macare-medicals.com | 25 |  | 2025-01-14 | MCS+ | 99J027 | Defective Parts: -Power Supply -Pneumatic Module Assembly -Compressor with filter Assembly -Display Cable -Line Sensor Assembly | . | Replaced Defective Parts: -Power Supply -Pneumatic Module Assembly -Compressor with filter Assembly -Display Cable -Line Sensor Assembly | Parts mentioned above are functional. Waiting additional parts: - Weigher Assembly - Centrifuge Cover - Lithium Battery for CPU board For repair/Adjustment/replacement: - Centrifuge Cover switch - Pump Rotors with rollers | For Calibration and Diagnostics after replacement of lacking parts mentioned. | -Power Supply (PN 48102-00) -Pneumatic Module Assembly (PN 83337-10) -Compressor with filter Assembly (PN 106789-00) -Display Cable (PN 112496-02) -Line Sensor Assembly (39436-10) | Turn Over Slip 7722 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143679095582034011/6143679095582034011_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6143679095582034011/6143679095582034011_signature_31.png
| 2025-02-03 | 2025-02-03 | Macare in-house unit | Cebu City | 7 | Cebu | Laboratory | NO | REPAIR | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 30240 | COMMERCIAL | macareengineering@gmail.com | 28 |  | 2025-01-14 | Mcs+ | 10M104 | In-house unit Defective pneumatic module assymbly | Install pneumatic module assymbly PN 83337-10 | Install pneumatic module assymbly PN 83337-10 | Unit functional |  | Pneumatic module assembly PN 83337-10 | Turn Over Slip Form 7722 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143582745582956816/6143582745582956816_signature_28.png | Macare medicals Inc. | https://www.jotform.com/uploads/macarewecare/222927978372067/6143582745582956816/6143582745582956816_signature_31.png
| 2025-02-03 | 2025-01-31 | Philippine Red Cross QC | Kalayaan Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbnak | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30239 | MACARE-HBDC | villamarevangeline103@yahoo.com | 28 |  |  | MCS+ | 02A227 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143517453614745946/6143517453614745946_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/6143517453614745946/6143517453614745946_signature_31.png
| 2025-02-03 | 2025-01-31 | Philippine Heart Center | East Avemue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30238 | MACARE-HBDC | phc_bb@yahoo.com | 28 |  |  | MCS + | 11H204 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143517123612752134/6143517123612752134_signature_28.png | philippine heart center | https://www.jotform.com/uploads/macarewecare/222927978372067/6143517123612752134/6143517123612752134_signature_31.png
| 2025-02-03 | 2025-01-30 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Blood collecting | NO | REPAIR CALIBRATION | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30237 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15H029G | CRC bios failed for replacement of cpu battery | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Replaced and installed battery Performed calibration Performed functionality test Performed simulation running | Repair and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143516733619927343/6143516733619927343_signature_28.png | Ivan Ritz P. Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6143516733619927343/6143516733619927343_signature_31.png
| 2025-02-03 | 2025-01-31 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30236 | MACARE-HBDC | phc_bb@yahoo.com | 28 |  |  | MCS+ | 20168 | Due for maintenance Encounter intermittent weigher self test error | Cleaned unit Performed self test multiple times Reset cable Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143516213617122888/6143516213617122888_signature_28.png | Philippine Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6143516213617122888/6143516213617122888_signature_31.png
| 2025-02-02 | 2025-01-28 | MEDI LINX LAB./ St. Elizabeth Hospital Inc. | General Santos City | 12 | South Cotabato | Micrology lab | NO | INSTALLATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | CL_SEHI.maintest2@medilinx.com.ph |  | 30235 | COMMERCIAL | claireramos.medilinx@outlook.com | 26.4 |  |  | GxIV-6C | 831313 | Module 1, 3, and 4 requesting for module replacement | Gx Dx Software installation MTB rif-ULTRA Assays Installation Genexpert Cleaning and maintenance Checking of errors/problem Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine software installation | 1 operational modules (module 2) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6143076836865971624/6143076836865971624_signature_28.png | Angel Claire Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/6143076836865971624/6143076836865971624_signature_31.png
| 2025-01-31 | 2025-01-31 | OTON MUNICIPAL HEALTH OFFICE | OTON,ILOILO | 6 | Iloilo | LABORATORY | NO | MAINTENANCE GX TRANSFER | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30234 | UNOPS | gracie05191980@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 183 |  | GX-IV 10C | 110023530 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6141263434832659593/6141263434832659593_signature_28.png | Mary Grace Ong-Lumasag | https://www.jotform.com/uploads/macarewecare/222927978372067/6141263434832659593/6141263434832659593_signature_31.png
| 2025-01-31 | 2025-01-31 | Macare In house | 1005 Aurora Blvd, Project 3, Quezon City, 1102 Metro Manila | NCR | NCR 2nd - Quezon City | Office | NO | REPAIR | OTHERS | Carlos Leniel Origenes |  | corigenes@macare-medicals.com |  | Aspex | 30233 | COMMERCIAL | corigenes@macare-medicals.com | 25 |  | 2025-01-24 | UPS | 533231 | No back up | Replaced battery | Replaced battery | Operational |  | Battery | Battery |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6141186503611949517/6141186503611949517_signature_28.png | Carlos Leniel Origenes | https://www.jotform.com/uploads/macarewecare/222927978372067/6141186503611949517/6141186503611949517_signature_31.png
| 2025-01-31 | 2025-01-31 | Macare In house | 1005 Aurora Blvd, Project 3, Quezon City, 1102 Metro Manila | NCR | NCR 2nd - Quezon City | Office | NO | REPAIR | OTHERS | Carlos Leniel Origenes |  | corigenes@macare-medicals.com |  | Aspex | 30232 | COMMERCIAL | corigenes@macare-medicals.com | 25 |  | 2025-01-24 | UPS | 533235 | No back up | Replaced battery | Replaced battery | Operational |  | Battery | Battery |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6141183803618896506/6141183803618896506_signature_28.png | Carlos Leniel Origenes | https://www.jotform.com/uploads/macarewecare/222927978372067/6141183803618896506/6141183803618896506_signature_31.png
| 2025-01-31 | 2025-01-31 | Calabanga Rural Health Unit | Poblacion, calabanga | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30231 | DOH | martyarc11@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 860 |  | GX4-6C | 838398 | Barcode scanner not working | Replaced barcode scanner Replaced fan filter | Replaced barcode scanner Performed Preventive Maintenance Replaced fan filter | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6141047878313655330/6141047878313655330_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/6141047878313655330/6141047878313655330_signature_31.png
| 2025-01-31 | 2025-01-30 | Bicol Medical Center - Blood Bank | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30230 | MACARE-HBDC | ma.charitydelosreyes@yahoo.com | 25 |  |  | MCS+ | 15H027G | MCS+ for Calibration Fluid in centrifuge | Cleaned Unit Check parameters Cleaned unit surface, centrifuge bowl, rotor pump, valve Calibrate/checked spm, dpm, centrifuge, line sensor, valves, cuff, rotor pumps, weighing rod Performed unit 3cycle test run | Performed Preventive Maintenance Performed Calibration | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6141047658313132798/6141047658313132798_signature_28.png | Ma. Charity P. de los Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6141047658313132798/6141047658313132798_signature_31.png
| 2025-01-31 | 2025-01-30 | Bicol Medical Center - Blood Bank | Concepcion, Naga | 5 | Camarines Sur | Laboratory | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30229 | MACARE-HBDC | ma.charitydelosreyes@yahoo.com | 25 |  |  | MCS+ | 15H031G | MCS+ for Calibration | Cleaned Unit Check parameters Cleaned unit surface, centrifuge bowl, rotor pump, valve Calibrate/checked spm, dpm, centrifuge, line sensor, valves, cuff, rotor pumps, weighing rod Performed unit 3cycle test run | Performed Calibration Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6141045878313971755/6141045878313971755_signature_28.png | Ma. Charity P. de los Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6141045878313971755/6141045878313971755_signature_31.png
| 2025-01-31 | 2025-01-31 | Cebu TB Reference Lab - SRP | SRP - Cebu | 7 | Cebu | Labortory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30228 | DOH | asmantiquilla.pbsp@gmail.com | 25 |  |  | Gx4 | 838263 | Defective Module B4 220004971 Error code 1001 and 2014 | Xpertcheck module 4 | Clean Fan Filter Plunger rod maintenance | Module 3 under observation since it has only 1 error Module 4 service require |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6141002807423884532/6141002807423884532_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6141002807423884532/6141002807423884532_signature_31.png
| 2025-01-31 | 2025-01-30 | Philippine Red Cross - Cagayan de Oro City | Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | Blood Center | NO | REPAIR | HAEMONETICS | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30227 | COMMERCIAL | mrbc2@redcross.org | 27 |  | 2025-01-23 | MCS+ | 02H080 | Defective Pneumatic Pump Assembly causing in opening/closing of the valves. | Repairing MCS+ Apheresis Machine | Replacing Pneumatic Pump Assembly into a new one. Checks if the problem still exist | MCS+ and other peripherals are now functional and working. |  | Pneumatic Pump Assembly | 12586 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6140983136615413971/6140983136615413971_signature_28.png | Coleen M. Rante | https://www.jotform.com/uploads/macarewecare/222927978372067/6140983136615413971/6140983136615413971_signature_31.png
| 2025-01-31 | 2025-01-31 | Liloy rural health unit | Liloy, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | MAINTENANCE Module Replacement and Preventive Maintenence | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30226 | DOH | nova.timusan79@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1211 | 2025-01-08 | GX 4 | 838536 | Module 4 h/w failed | Replaced module 862279 with Module 687952 and conducted Preventive maintenence | Replaced module 862279 with Module 687952 and conducted Preventive maintenence | All modules are available and ready to use. | None | 900-0491RS | 12537 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6140965494461093377/6140965494461093377_signature_28.png | Nova L. Timusan | https://www.jotform.com/uploads/macarewecare/222927978372067/6140965494461093377/6140965494461093377_signature_31.png
| 2025-01-30 | 2025-01-30 | Philippine Red Cross Davao City | Roxas Boulevard, Davao City | 11 | Davao del Sur | Blood Bank | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 30225 | COMMERCIAL | davaocitybloodservices@gmail.com | 23.8 |  |  | Haemonetics MCS+ | 02J021 | Inaccurate Weigher | Replaced the Defective Weigher SN:929980 with service Weigher SN: 904788 Check and adjusted SPM, DPM and Line Sensor values Check roller assays and perform cycle test | Weigher Replacement Calibrate the MCS+ Perform Cycle Test Check other parameters if it is in range | Successful replaced the weigher, unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6140511382224463661/6140511382224463661_signature_28.png | Joy Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/6140511382224463661/6140511382224463661_signature_31.png
| 2025-01-30 | 2025-01-30 | Don Jose S. Monfort Medical Center | Barotac Nuevo, Iloilo | 6 | Iloilo | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30224 | DOH | judyboglosa35@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2000 |  | GX-IV | 838501 | -MODULE B2 ERROR 4009 -FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL, MODULE B2 FOR OBSERVATION |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6140456986812562921/6140456986812562921_signature_28.png | Judy Boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6140456986812562921/6140456986812562921_signature_31.png
| 2025-01-30 | 2025-01-30 | Philippine Red Cross Benguet | Benguet | CAR | Benguet | Blood Bank | NO | INSTALLATION | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30223 | MACARE-HBDC | lykabagano@yahoo.com | 25 |  | 2025-01-27 | MCS+ | 17H636SPG | Unit for installation | NA | #NAME? | Unit is operational | W/ Sir Pua and Sir Salih |  | 12597 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6140274077286977287/6140274077286977287_signature_28.png | Lykaine Kate Bagano | https://www.jotform.com/uploads/macarewecare/222927978372067/6140274077286977287/6140274077286977287_signature_31.png
| 2025-01-30 | 2025-01-30 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30222 | DOH | meggysantiago@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 406 | 2025-01-30 | GXIV | 838548 | N/a | N/a | - Performed Preventive Maintenance - Performed functionality test - | Unit it Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6140163394794652963/6140163394794652963_signature_28.png | Mary Grace Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/6140163394794652963/6140163394794652963_signature_31.png
| 2025-01-30 | 2025-01-30 | Silliaman University Medical Center | Damaguete City | 7 | Cebu | Laboratory | YES | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 30221 | COMMERCIAL | bangalando23@gmail.com | 20 |  | 2025-01-28 | GX4 | 821471 | . Sometimes curent fluctuating. . Error 5017 . Aborted the cartridge while running. | . Perform preventive maintenance for the Genexpert Machine . Replace new cable wire for 4 module's inside due to lost communication. | . Save system log file . Save Installation qualification report . Save assay statistics report. . Save Gxx file | .Module A2 need to replace due to calibration fail. Module A1, 3,4 ol to use. . Ups are good condition. | Need AVR for the Genexpert machine. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6140091437313732762/6140091437313732762_signature_28.png | Maryjune B. Bangalando | https://www.jotform.com/uploads/macarewecare/222927978372067/6140091437313732762/6140091437313732762_signature_31.png
| 2025-01-29 | 2025-01-28 | Visayas Med | Cebu City | 7 | Cebu | Laboratory | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 30220 | COMMERCIAL | mirasol.trinidad@singaporediagnostics.com | 28 |  |  | Mcs+ | 05M204 | Watch Dog Error Valve test Error | Swapped parts to isolate problem. | Swapped Valve Driver PCB Assembly with service part to isolate problem with Pinch valves. | Confirmed defective - Valve Driver PCB Assembly, for request of quotation. Confirmed functional - Pneumatic Module Assembly. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6139528455428082092/6139528455428082092_signature_28.png | Maria Morasol Trinidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6139528455428082092/6139528455428082092_signature_31.png
| 2025-01-28 | 2025-01-27 | Madrid District Hospital | Madrid, Surigao del Sur | 13 | Surigao del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30219 | DOH | madrid_dh@gmail.com | 26 |  | 2025-01-08 | GX4 6 Colors | 838439 | Reported Module 4, found Temperature Errors | Module Replacement and Preventive Maintenance | Replacing Module 4 with SN: 889069 to SN: 912610 Preventive Maintenance | All modules are functional |  | Certified pre-owned GeneXpert Module | 12548 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6138328188411793801/6138328188411793801_signature_28.png | Rochelle Pueblas | https://www.jotform.com/uploads/macarewecare/222927978372067/6138328188411793801/6138328188411793801_signature_31.png
| 2025-01-27 | 2025-01-23 | Columbio Rural Health Unit | Poblacion, Columbio,Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30218 | DOH | rochelle.toledo19@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 481 | 2025-01-23 | GxIV-6C | 838556 | Module 3 for Service Required | Replace defective module 3 SN:692567 with new module SN: 634162 Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection | Genexpert module repair Genexpert Preventive maintenance | 4 operational modules |  | Genexpert 6-color module SN:634162 | DR#: 12531 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137948496869168333/6137948496869168333_signature_28.png | Rochelle S toledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6137948496869168333/6137948496869168333_signature_31.png
| 2025-01-27 | 2025-01-27 | Philippine Red Cross Davao Chapter | Roxas Boulevard, Davao City | 11 | Davao del Sur | Blood Bank | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 30217 | COMMERCIAL | davaocitybloodservices@gmail.com | 25.7 |  |  | MCS+ | 02J021 | Weigher problem | Reset Weigher Connections Reset CPU Board Recalibrate weigher paramaters | Perform Troubleshooting and reset board connections | Unit not operational, weigher displaying incorrect values for weight. Unit errors documented, sent to supervisor for evaluation | Waiting for replacement parts |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137734872218572550/6137734872218572550_signature_28.png | Joy Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/6137734872218572550/6137734872218572550_signature_31.png
| 2025-01-27 | 2025-01-24 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30216 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 02D166 | Due for maintenance Urethane roller assy slight worn out | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | For quotation of urethane roller assy |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137601323614146267/6137601323614146267_signature_28.png | Krystle Unlayao | https://www.jotform.com/uploads/macarewecare/222927978372067/6137601323614146267/6137601323614146267_signature_31.png
| 2025-01-27 | 2025-01-24 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30215 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 08H188 | Due for maintenance Urethane roller assy worn out already UPS unable to backup battery | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | For quotation of urethane roller assy need to replace |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137600373615185643/6137600373615185643_signature_28.png | Krystle Unlayao | https://www.jotform.com/uploads/macarewecare/222927978372067/6137600373615185643/6137600373615185643_signature_31.png
| 2025-01-27 | 2025-01-24 | Lung Center of the Philippines | Quezon avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30214 | COMMERCIAL | monet.laxina@yahoo.com.ph | 26 |  |  | GX4 L | 804205 | Due for maintenance D1,D2 reported already failed xpertcheck | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | D2,D3,D4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137599943616748127/6137599943616748127_signature_28.png | Simonette R. Laxina | https://www.jotform.com/uploads/macarewecare/222927978372067/6137599943616748127/6137599943616748127_signature_31.png
| 2025-01-27 | 2025-01-23 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30210 | COMMERCIAL | moleculardx@themedicalcity.com | 20 |  |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137549603612071831/6137549603612071831_signature_28.png | Stephanie Obida | https://www.jotform.com/uploads/macarewecare/222927978372067/6137549603612071831/6137549603612071831_signature_31.png
| 2025-01-27 | 2025-01-24 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30212 | COMMERCIAL | stephobida90@gmail.com | 22 |  |  | GX-XVI-D-10C | 110015767 | Cannot upload test to LIS | Reset communication buffer | Check host connection Reset communication buffer Upload pending test results to LIS | Unit is operational. All modules operational. Test uploading successful. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137549603612187491/6137549603612187491_signature_28.png | Stephanie obida | https://www.jotform.com/uploads/macarewecare/222927978372067/6137549603612187491/6137549603612187491_signature_31.png
| 2025-01-27 | 2025-01-24 | Ospital ng Makati | Sampaguita St. Makati City | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 30213 | MACARE-HBDC | osmakpatho_labs@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 253 | 2025-01-14 | GX-IV-D-10C | 110023460 | For XpertCheck | Performed XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  | DR# 12569 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137549603613142907/6137549603613142907_signature_28.png | Shairene Aubrey S. Espiritu, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6137549603613142907/6137549603613142907_signature_31.png
| 2025-01-27 | 2025-01-23 | Marikina City Health Office | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30211 | MACARE-HBDC | estyaragon2019@gmail.com | 20 |  |  | GX-IV-D-10C | 110024073 | Module B3 for replacement | Replaced module B3 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter -replaced module B3 Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137549603617184789/6137549603617184789_signature_28.png | ESTRELITA S. ARAGON | https://www.jotform.com/uploads/macarewecare/222927978372067/6137549603617184789/6137549603617184789_signature_31.png
| 2025-01-27 | 2025-01-24 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30209 | DOH | annabellelcatbagan@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 208 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 51 |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137537798834201710/6137537798834201710_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6137537798834201710/6137537798834201710_signature_31.png
| 2025-01-27 | 2025-01-24 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30208 | DOH | ganapinhazel22@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137535188835816896/6137535188835816896_signature_28.png | Hazel Ganapin | https://www.jotform.com/uploads/macarewecare/222927978372067/6137535188835816896/6137535188835816896_signature_31.png
| 2025-01-27 | 2025-01-23 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30207 | DOH | allan.vinluan2013@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 934 |  | GXIV-4-D-6C | 838472 | For Troubleshooting and Preventive Maintenance | #NAME? | Performed Preventive Maintenance and Troubleshooting | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137532338839417688/6137532338839417688_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6137532338839417688/6137532338839417688_signature_31.png
| 2025-01-27 | 2025-01-23 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30206 | DOH | seansimonpaloyo@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1103 |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137531348839263081/6137531348839263081_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6137531348839263081/6137531348839263081_signature_31.png
| 2025-01-27 | 2025-01-27 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | UPS | 30205 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  | 2025-01-27 | ACCUPOWER | AT1000170815ST0163 | For replacement of 2pcs battery | Replaced and installed 2pcs battery Performed functionality test | Repair | Operational |  | UPS Battery | , |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137503863616414483/6137503863616414483_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6137503863616414483/6137503863616414483_signature_31.png
| 2025-01-27 | 2025-01-27 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | UPS | 30204 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  | 2025-01-27 | ACCUPOWER | AT1000170815ST0163 | No power Low voltage battery ouput Battery wornout already | Checked battery voltage output Checked ups | troubleshooting | For replacement of 2pcs battery. |  |  | , |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137502183611012625/6137502183611012625_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6137502183611012625/6137502183611012625_signature_31.png
| 2025-01-27 | 2025-01-27 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | UPS | 30203 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  | 2025-01-27 | ACCUPOWER | AT1000190123ST0184 | For replacement of 2pcs Battery | Replaced and installed 2pcs battery Performed functionality test | Performed repair | Operational |  |  | , |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137491673615917901/6137491673615917901_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6137491673615917901/6137491673615917901_signature_31.png
| 2025-01-27 | 2025-01-27 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | nfloresca@macare-medicals.com | UPS | 30202 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  | 2025-01-27 | ACCUPOWER | AT1000181128ST0151 | For replacement of 2pcs battery | Replaced and installed 2pcs battery Performed functionality test | Performed repair | Operational |  | UPS Battery | , |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137489023614041539/6137489023614041539_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6137489023614041539/6137489023614041539_signature_31.png
| 2025-01-26 | 2025-01-22 | Tagum City Health Office | Poblacion, Tagum City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30201 | DOH | tagumcholab.mabini@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2107 |  | GX-IV 6C 4D | 838553 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational | Module 2 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6137008372213202323/6137008372213202323_signature_28.png | Cristina Tiara A. Fernandez,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6137008372213202323/6137008372213202323_signature_31.png
| 2025-01-24 | 2025-01-24 | Bolinao Rural Health Unit | Bolinao Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | ruralhealthunitlaboratory@gmail.com |  | 30200 | DOH | nylemae14@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 677 |  | GXIV-D | 838278 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6135149307177571217/6135149307177571217_signature_28.png | Nicole Caasi | https://www.jotform.com/uploads/macarewecare/222927978372067/6135149307177571217/6135149307177571217_signature_31.png
| 2025-01-24 | 2025-01-24 | Mabuhay Rural Health unit | Poblacion Mabuhay Zamboanga Sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | CALIBRATION MAINTENANCE Xpertcheck calibration and Preventive Maintenence | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30199 | UNOPS | tessanavarro.20.tn@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 209 | 2025-01-08 | GX 4 | 110023517 | None | Preventive Maintenence and Xpertcheck calibration for all modules | Preventive Maintenence and Xpertcheck calibration for all modules | All modules are available and passed the Xpertcheck Calibration | None | XPERTCHECK-CE-5 | 12487 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6135138138444588612/6135138138444588612_signature_28.png | Tessaphaidal A. Navarro | https://www.jotform.com/uploads/macarewecare/222927978372067/6135138138444588612/6135138138444588612_signature_31.png
| 2025-01-24 | 2025-01-24 | Puerto Princesa CHO | Puerto Princesa, Palawa | 4B | Palawan | Laboratory | NO | MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 30198 | MACARE-HBDC | jocal0324@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2018 | 2025-01-24 | GX16 | 110021518 | None. For maintebance. | Performed maintwnance. Generate reports. | Performed cleaning and naintenance. | All 16 modules operational. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6135104194012342119/6135104194012342119_signature_28.png | JOCELYN CALALIN | https://www.jotform.com/uploads/macarewecare/222927978372067/6135104194012342119/6135104194012342119_signature_31.png
| 2025-01-24 | 2025-01-24 | Puerto Princesa CHO | Puerto Princesa, Palawa | 4B | Palawan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 30197 | UNOPS | jocal0324@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 143 | 2025-01-08 | GXIV | 10023510 | None. For maintenance and xpertcheck calibration. | Performed naitenance and xpertcheck calibration. | Performed unit maintennce and xpertcheck calibration. | All 4 modules operational. All 4 modules PASSED the xpertcheck calibration. | NA | Xpertcheck kit | DR Number: 12504 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6135100224011163272/6135100224011163272_signature_28.png | Jocelyn Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/6135100224011163272/6135100224011163272_signature_31.png
| 2025-01-24 | 2025-01-24 | Dasol Community Hospital | Dasol Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30196 | DOH | jeremypulido0521@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 604 |  | GXIV-D 6C | 838287 | Due for preventive maintenance B3:808251 temperature error B4:808252 temperature error | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational B3 and B4, temperature error | B3 and B4 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6135092496393833178/6135092496393833178_signature_28.png | Jeremy Pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/6135092496393833178/6135092496393833178_signature_31.png
| 2025-01-24 | 2025-01-24 | Infanta Rural Health Unit | Poblacion Infanta, Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30195 | UNOPS | charissealmandres8@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 122 | 2025-01-08 | GXIV-D 10C | 110023501 | Due for calibration | Preventive maintenance Xertcheck | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed xpertcheck Performed functionality test Gathered reports | Unit is operational |  | Xpertcheck | 12499 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6135051356395038768/6135051356395038768_signature_28.png | Charisse A. Alferez | https://www.jotform.com/uploads/macarewecare/222927978372067/6135051356395038768/6135051356395038768_signature_31.png
| 2025-01-24 | 2025-01-24 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30194 | DOH | calambatbdots@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1824 |  | GXIV | 838225 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6135046989819552557/6135046989819552557_signature_28.png | Elizabeth Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6135046989819552557/6135046989819552557_signature_31.png
| 2025-01-24 | 2025-01-24 | Santa Rosa City Health Office 1 | City Health Office Compound | NCR | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | alonaalvaro27@gmail.com |  | 30193 | UNOPS | santarosacho1clinlab@gmail.com | 25 |  | 2025-01-08 | GXIV-10C | 110023497 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck kit | DR#12497 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6135043559811909260/6135043559811909260_signature_28.png | Alona Nina Alvaro | https://www.jotform.com/uploads/macarewecare/222927978372067/6135043559811909260/6135043559811909260_signature_31.png
| 2025-01-24 | 2025-01-24 | PGP-Molecular Diagnostics Laboratory | Puerto Princesa, Palawa | 4B | Palawan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 30192 | UNOPS | anopheles73@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 483 | 2025-01-08 | GXIV | 110023509 | None, For unit maintenance and Xpertcheck calibration. | Performed maintenance and xpertchek calibration. | Performed unit maintenance and run xpertcheck calibration. | All 4 modules operational. Xpertcheck results: All 4 modules PASSED | NA | 1 kit of xpertcheck. | DR Number: 12504 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134957514017197783/6134957514017197783_signature_28.png | Cyrus Cecilio A. Caabay RMT MPHA | https://www.jotform.com/uploads/macarewecare/222927978372067/6134957514017197783/6134957514017197783_signature_31.png
| 2025-01-23 | 2025-01-23 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | REPAIR | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30191 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  | 2025-01-22 | MCS+ | 01J038 | Rotor roller rubbers for replacement | Replaced rotor roller rubbers | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134247095337615132/6134247095337615132_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6134247095337615132/6134247095337615132_signature_31.png
| 2025-01-23 | 2025-01-23 | Philippine Red Cross-Ormoc Chapter | Ormoc City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 30190 | COMMERCIAL | ormocbloodservice@redcross.org.ph | 28 |  |  | Gx 4 | 15H043G | For Installation brand new line sensor | Install brand new line sensor Replace a worn out pump rubber | Install brand new line sensor Replace a worn out pump rubber | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134247060711036336/6134247060711036336_signature_28.png | Hannah de claro | https://www.jotform.com/uploads/macarewecare/222927978372067/6134247060711036336/6134247060711036336_signature_31.png
| 2025-01-23 | 2025-01-23 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30189 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 08B049 | Centrifuge for troubleshooting | Adjusted centrifuge bolts | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134246615337471889/6134246615337471889_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6134246615337471889/6134246615337471889_signature_31.png
| 2025-01-23 | 2025-01-23 | Tamparan District Hospital | Tamparan, Lanao del Sur | BARMM | Lanao del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30188 | UNOPS | pmamowalas@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 10 | 2025-01-08 | 10C GXIV-4D | 110023525 | Due for XPC and PM | #NAME? | #NAME? | All modules passed XpertCheck and is functional. |  | XpertCheck-CE-5 Lot No. 1001075092 Exp. June 01, 2025 | DR no. 12526 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134243894125011834/6134243894125011834_signature_28.png | Pangilamun Mamowalas Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6134243894125011834/6134243894125011834_signature_31.png
| 2025-01-23 | 2025-01-23 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | UPS | 30187 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  |  | ACCUPOWER | AT1000190123STO244 | For replacement of 2pcs battery | Replaced and installed 2pcs battery. Performed functionality test. | Performed repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134156813614352839/6134156813614352839_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6134156813614352839/6134156813614352839_signature_31.png
| 2025-01-23 | 2025-01-23 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | UPS | 30186 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  |  | ACCUPOWER | AT1000181128ST0151 | No power Low voltage output of battery | Checked UPS Checked and test voltage output of battery 3.5volts | Troubleshooting | For replacement of 2pcs battery |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134147863614981244/6134147863614981244_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6134147863614981244/6134147863614981244_signature_31.png
| 2025-01-23 | 2025-01-23 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | UPS | 30185 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  |  | ACCUPOWER | AT1000190123STO244 | No power Low voltage output of battery | Checked ups Checked and test voltage output of battery 2.0V | troubleshooting | For replacement of 2pcs battery |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134146163614352169/6134146163614352169_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6134146163614352169/6134146163614352169_signature_31.png
| 2025-01-23 | 2025-01-23 | Macare in house | 1005 Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com |  | UPS | 30184 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  |  | ACCUPOWER | AT1000190123ST0184 | No power No voltage output battery | Performed troubleshooting Checked and test voltage output of battery. 2.6volts | Troubleshooting | For replacement of battery 2pcs. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134144393618128284/6134144393618128284_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6134144393618128284/6134144393618128284_signature_31.png
| 2025-01-23 | 2025-01-23 | Philippine General Hospital | Taft Ave., Ermita Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30183 | COMMERCIAL | pghmolecularlab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 85 ASSAY: Others, COUNT: SA Nasal Complete G3 = 1 ASSAY: Xpert SARS-CoV-2, COUNT: 557 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 858 |  | GX16 | 844975 | Unit for preventive maintenance. For checking of monitor. | Preventive maintenance and troublesbooting of monitor. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed checking of monitor. Generated reports. | Unit is operational. Modules A4, B2, B3 and C3 for reporting due to temperature errors. Monitor for replacement due to half display. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134108008219217841/6134108008219217841_signature_28.png | Kleante bautro | https://www.jotform.com/uploads/macarewecare/222927978372067/6134108008219217841/6134108008219217841_signature_31.png
| 2025-01-23 | 2025-01-23 | Cawayan Rural Health Unit | Poblacion, Cawayan | 5 | Masbate | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30182 | DOH | Gellatisimo@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 917 |  | GX4-6C | 838267 | Gx for Preventive Maintenance Request for fan filter | Cleaned cartridge bays, plunger rod, icore Provide 1 extra fan filter Generate Data | Performed Preventive Maintenance Provide extra fan filter | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134089193216389574/6134089193216389574_signature_28.png | Gellie Jane Gader | https://www.jotform.com/uploads/macarewecare/222927978372067/6134089193216389574/6134089193216389574_signature_31.png
| 2025-01-23 | 2025-01-23 | The Medical City Clark | Clark Pampanga | 3 | Pampanga | Microbiology Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | tmcclarkbiomed@themedicalcity.com |  | 30181 | COMMERCIAL | tmci_laboratory@themedicalcity.com | 24 | ASSAY: Xpert CT/NG, COUNT: 12 ASSAY: Xpert HBV VL, COUNT: 64 ASSAY: Xpert HPV, COUNT: 31 ASSAY: Xpert MTB/RIF Ultra, COUNT: 147 |  | GXIV-D 6C | 839235 | Module B4: 866908 temperature error Due for preventive maintenance Due for Calibration ( last calibration June 21,2022) | Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except for module B4: 866908 Machine for Calibration ( Last calibration June 21,2022) | Module B4: 866908 for relacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134088518447040650/6134088518447040650_signature_28.png | Jessa Rae Remorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6134088518447040650/6134088518447040650_signature_31.png
| 2025-01-23 | 2025-01-23 | University of the Philippines - NIH | Pedro Gil St. Ermita Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30180 | COMMERCIAL | nihcentrallab@gmail.com | 25 | ASSAY: Xpert Carba-R, COUNT: 122 ASSAY: Xpert CT/NG, COUNT: 10 ASSAY: Xpert HBV VL, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert MTB-XDR, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 40 ASSAY: Xpert HIV-1 Viral Load, COUNT: 521 |  | GX16-10C | 110014186 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134086187122403134/6134086187122403134_signature_28.png | Mary Grace Hernaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6134086187122403134/6134086187122403134_signature_31.png
| 2025-01-23 | 2025-01-23 | Macare In House | 1005 Aurora grand,QC | NCR | NCR 2nd - Quezon City | Enginnering | NO | CALIBRATION TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 30179 | MACARE-HBDC | florescanoel77@gmail.com | 28 |  |  | MCS + | 12D166 | For calibration Urethane roller assy worn out already | Performed calibration Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed calibration and troubleshooting | For replacement of urethane roller assy |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134054513614836764/6134054513614836764_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6134054513614836764/6134054513614836764_signature_31.png
| 2025-01-23 | 2025-01-23 | Macare In House | 1005 Aurora grand,QC | NCR | NCR 2nd - Quezon City | Enginnering | NO | CALIBRATION TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 30178 | MACARE-HBDC | florescanoel77@gmail.com | 28 |  |  | MCS + | 12D166 | For calibration Urethane roller assy worn out already | Performed calibration Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed calibration and troubleshooting | For replacement of urethane roller assy |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6134054493615602182/6134054493615602182_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6134054493615602182/6134054493615602182_signature_31.png
| 2025-01-22 | 2025-01-22 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30177 | MACARE-HBDC | honeyeichel@gmail.com | 24 | ASSAY: Xpert HBV VL, COUNT: 0 ASSAY: Xpert SARS-CoV-2, COUNT: 3293 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1580 |  | GXIV-D 10C | 110017894 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133627884118679978/6133627884118679978_signature_28.png | Hobey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/6133627884118679978/6133627884118679978_signature_31.png
| 2025-01-22 | 2025-01-22 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | PCR Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 30176 | MACARE-HBDC | honeyeichel@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 3574 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1000 |  | GXIV-D 6C | 829390 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133627674112616515/6133627674112616515_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/6133627674112616515/6133627674112616515_signature_31.png
| 2025-01-22 | 2025-01-22 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30175 | DOH | jan19.cam@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1219 |  | GxIV-6C | 838504 | None | Perform Genexpert Cleaning and maintenance Checking of errors Data collection | Genexpert preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133575686867551152/6133575686867551152_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/6133575686867551152/6133575686867551152_signature_31.png
| 2025-01-22 | 2025-01-21 | Malapatan Rural Health Unit | Malapatan, Sarangani Province | 12 | Sarangani | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30174 | DOH | kamilahnichol@yahoo.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 662 | 2025-01-21 | GxIV-6C | 838419 | Module 4 for Service Required | Replace defective module 4 SN:635997 with new module SN:761549 Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection | Genexpert module repair Genexpert Preventive maintenance | 4 operational modules |  | Genexpert 6-color module SN:761549 | DR#:12541 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133575686869694643/6133575686869694643_signature_28.png | Kamilah Nichol Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/6133575686869694643/6133575686869694643_signature_31.png
| 2025-01-22 | 2025-01-21 | Marawe City Health Office | Marawe City, Lanao del Norte | 10 | Lanao del Norte | Tb laboratory | NO | INSTALLATION REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30173 | DOH | faridahmamao@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 |  | GX4 | 838293 | Display monitor numerous horizontal lines display. | Replacement of display monitor. | Replacement of display monitor HP SN: CN461017D3 | The machine and the display monitor is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133412645582072545/6133412645582072545_signature_28.png | Faridah BUsara,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6133412645582072545/6133412645582072545_signature_31.png
| 2025-01-22 | 2025-01-20 | Marawe City Health Office | Marawe City, Lanao del Norte | 10 | Lanao del Norte | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30172 | DOH | faridahmamao@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 950 |  | GX4 | 838260 | Module 4 SN: 808095 failed during calibration. | Replacement of defective module and replacement of new module 4 SN: 885935 Generate IQR, System Log, GXX files and Assay Stats Clean filter and Icore, plunger. | Replacement of defective module and replacement of new module 4 SN: 885935 Generate IQR, System Log, GXX files and Assay Stats Clean filter and Icore, plunger. | All modules are operational. Need an additional filter. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133410395586839767/6133410395586839767_signature_28.png | Faridah BUsara,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6133410395586839767/6133410395586839767_signature_31.png
| 2025-01-22 | 2025-01-13 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank Service Unit (Sub-National) | NO | INSTALLATION ON-SITE ASSISTANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 30171 | COMMERCIAL | vsmmcbloodservices@gmail.com | 24 |  |  | MCS+ | 02A090 | Installation in the Sub-national laboratory for permanent transfer. | Dismantle UPS and AVR, Stand and Mcs+ Assemble and install in the Subnational Bloodbank | Dismantle UPS and AVR, Stand and Mcs+ Assemble and install in the Subnational Bloodbank | The machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133403575583692848/6133403575583692848_signature_28.png | Charlotte Mae Ygoña, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6133403575583692848/6133403575583692848_signature_31.png
| 2025-01-22 | 2025-01-22 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | RHIVDA | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | bataan.bghmc.rhivda@gmail.com |  | 30170 | MACARE-HBDC | mariaana.hizon@gmail.com | 24 | ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 15 ASSAY: Xpert HIV-1 Viral Load, COUNT: 108 |  | GXIV-D 10C | 120003587 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133389448365286894/6133389448365286894_signature_28.png | Maria Ana Hizon | https://www.jotform.com/uploads/macarewecare/222927978372067/6133389448365286894/6133389448365286894_signature_31.png
| 2025-01-22 | 2025-01-22 | Esperanza Rural Health Unit | Poblacion, Esperanza | 5 | Masbate | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30169 | DOH | esperanzarhu@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 671 |  | GX4-6C | 838268 | Gx for Preventive Maintenance Request fan filter | Cleaned cartridge bays, plunger rod Provide extra fan filter Generate Data | Performed Preventive Maintenance Provide fan filter | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133386330935267098/6133386330935267098_signature_28.png | Airyn L. Glory | https://www.jotform.com/uploads/macarewecare/222927978372067/6133386330935267098/6133386330935267098_signature_31.png
| 2025-01-22 | 2025-01-22 | Binalbagan Municipal Health Office | Binalbagan, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30168 | DOH | anne_galilea@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 768 |  | GX4 | 838280 | High room ambient temperature; more than 32 C temp around the Genexpert machine. | . | Preventive Maintenance - cleaned optics, plunger rods, module bays, filter; provide additional filter. | 4 modules operational | Genexpert operating temperature should be 15-30C to avoid module hardware failures. Recommendations: Clean ACU; temporarily install electric fan for the Genexpert machine; install additinal ACU. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133381313424511351/6133381313424511351_signature_28.png | Anne Galilea | https://www.jotform.com/uploads/macarewecare/222927978372067/6133381313424511351/6133381313424511351_signature_31.png
| 2025-01-22 | 2025-01-22 | Philippine Red Cross - Cagayan de Oro City | Cagayan de Oro, Misamis Oriental | 10 | Misamis Oriental | Red Cross Blood Center | NO | TROUBLESHOOTING | HAEMONETICS | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30167 | COMMERCIAL | mrbc2@redcross.org | 27 |  |  | Apheresis MCS+ | 02H080 | Problem with Pneumatic Assembly causing problem with opening and closing of Valves | Reported to Main Office | Checking status of the equipment Troubleshooting | MCS+ not functional | For replacement of pneumatic assembly, waiting for parts from Main Office |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133356767146838589/6133356767146838589_signature_28.png | Phoenix James Dahili | https://www.jotform.com/uploads/macarewecare/222927978372067/6133356767146838589/6133356767146838589_signature_31.png
| 2025-01-22 | 2025-01-22 | Hi Precision Diagnostics | N. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30166 | CONSORTIUM | jose@hi-precision.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 121 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6681 |  | GXIV | 835602 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133327466015203529/6133327466015203529_signature_28.png | Andrei Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/6133327466015203529/6133327466015203529_signature_31.png
| 2025-01-22 | 2025-01-22 | Hi Precision Diagnostics | N. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30165 | CONSORTIUM | jose@hi-precision.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 149 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5862 |  | GXIV | 835601 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133326636018501250/6133326636018501250_signature_28.png | Andrei Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/6133326636018501250/6133326636018501250_signature_31.png
| 2025-01-22 | 2025-01-22 | Piñan Rural Health Unit | Piñan, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | CALIBRATION MAINTENANCE Xpertcheck calibration and Preventive Maintenence | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30164 | UNOPS | ivannagmarchan@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 77 | 2025-01-08 | GX 4 | 110023516 | None | All modules undergone Xpertcheck calibration and Preventive Maintenence | All modules undergone Xpertcheck calibration and Preventive Maintenence | All modules passed the calibration and are available | None | XPERTCHECK-CE-5 | 12486 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133301714426446557/6133301714426446557_signature_28.png | Ivanna G. Marchan | https://www.jotform.com/uploads/macarewecare/222927978372067/6133301714426446557/6133301714426446557_signature_31.png
| 2025-01-22 | 2025-01-09 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30163 | DOH | dianeruiz147@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 465 |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133297010946817597/6133297010946817597_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6133297010946817597/6133297010946817597_signature_31.png
| 2025-01-22 | 2025-01-09 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30162 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 206 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 45 |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133295870941961455/6133295870941961455_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6133295870941961455/6133295870941961455_signature_31.png
| 2025-01-22 | 2025-01-22 | Cagayan De Oro City Health Office | CHO TB-Dots, Cagayan de Oro | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30161 | DOH | chocdotbdots@gmail.com | 24 |  | 2025-01-08 | GX4 6 Colors | 838435 | Module C2 found Service Required thru XPC Calibration | Module Replacement and PM | Replacing module C2 with SN: 791261 to SN: 806222 Preventive Maintenance | All modules and units are functional |  | Certified pre-owned GeneXpert Module | 12543 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133278585651846525/6133278585651846525_signature_28.png | Charisse Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/6133278585651846525/6133278585651846525_signature_31.png
| 2025-01-22 | 2025-01-22 | Baybay City Health Office | Baybay Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30160 | DOH | sajecopooh20@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1146 | 2025-01-22 | GXIV | 838555 | n/a | N/a | #NAME? | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133236774539937529/6133236774539937529_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/6133236774539937529/6133236774539937529_signature_31.png
| 2025-01-22 | 2025-01-22 | Carmen Health Center | Brgy. Carmen, Cagayan de Oro City | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30159 | DOH | murielmelchor11@gmail.com | 26 |  | 2025-01-08 | GX4 6 Colors | 838442 | Reported Module B2 found Temperature Errors | Module Replacement and Maintenance | Replacing Module B2 with SN: 210021428 to SN: 885917 Preventive Maintenance | All modules and units are now functional |  | Certified pre-owned GeneXpert Module | 12545 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133218015656309194/6133218015656309194_signature_28.png | Muriel G. Melchor | https://www.jotform.com/uploads/macarewecare/222927978372067/6133218015656309194/6133218015656309194_signature_31.png
| 2025-01-22 | 2025-01-22 | Masbate City Health Office | Kalipay, Masbate City | 5 | Masbate | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30158 | UNOPS | Juedevan.ancajas14@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 148 ASSAY: Xpert MTB-XDR, COUNT: 2 |  | GX4-10C | 110023512 | Gx for Calibration | XpertCheck Result: A1/210121669 - passed A2/210122064 - passed A3/210117513 - passed A4/210120091 - passed Cleaned cartridge bays, plunger rod, icore Generate Data | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6133196701349502929/6133196701349502929_signature_28.png | JUEDE VAN ANCAJAS | https://www.jotform.com/uploads/macarewecare/222927978372067/6133196701349502929/6133196701349502929_signature_31.png
| 2025-01-21 | 2025-01-21 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@protonmail.com |  |  | 30157 | DOH | cauayanmholab@gmail.com | 26 |  |  | GX4 | 838477 | Due for Preventive maintenance | . | Performed Preventive Maintenance - cleaned optics, plunger rods, module bays, filter; archive tests. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6132581873429248383/6132581873429248383_signature_28.png | Doriza Joy Tizon | https://www.jotform.com/uploads/macarewecare/222927978372067/6132581873429248383/6132581873429248383_signature_31.png
| 2025-01-21 | 2025-01-21 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30156 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2361 ASSAY: Xpert HIV-1 Viral Load, COUNT: 205 |  | 6C GXIV-4D | 838285 | Due for monthly PM | -Preventive Maintenance; cartridge bay, fan filter, plunger rod, optic slot -Archive test results -Backup database | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6132482789945546221/6132482789945546221_signature_28.png | Leah Karen Noynay | https://www.jotform.com/uploads/macarewecare/222927978372067/6132482789945546221/6132482789945546221_signature_31.png
| 2025-01-21 | 2025-01-21 | Calamba Rural Health Unit | Calamba misamis occidental | 10 | Misamis Occidental | Laboratory | NO | CALIBRATION MAINTENANCE Xpertcheck calibration and Preventive Maintenence | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 30155 | UNOPS | palubonibn@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 125 | 2025-01-08 | GX 4 | 110023518 | None | Preventive Maintenence and Xpertcheck calibration for all four modules | Preventive Maintenence and Xpertcheck calibration for all four modules | All modules are available. All modules passed Xpertcheck calibration | All module are available | XPERTCHECK-CE-5 | 12488 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6132447769518117679/6132447769518117679_signature_28.png | Mary-ibn E. Palubon | https://www.jotform.com/uploads/macarewecare/222927978372067/6132447769518117679/6132447769518117679_signature_31.png
| 2025-01-21 | 2025-01-21 | National Center for Mental Health | Mandaluyong City | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30154 | MACARE-HBDC | iancristrono@gmail.com | 28 |  |  | GXIV 10C | 120003480 | Unit for additional troubleshooting by Cepheid | NA | - checked if B3's module door is completely closed during procedure - checked if there are external light falling directly onto the icore of the module - generated reports | Unit is operational Previous tests done in module B3 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6132423335336482771/6132423335336482771_signature_28.png | Ian Cris Tono | https://www.jotform.com/uploads/macarewecare/222927978372067/6132423335336482771/6132423335336482771_signature_31.png
| 2025-01-21 | 2025-01-21 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Mabini, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30153 | UNOPS | bryanmichaelbalunes@gmail.com | 28 |  | 2025-01-21 | GXIV-10C | 110023508 | Due for Calibration | Performed Xpertcheck (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | 12493 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6132320681128989145/6132320681128989145_signature_28.png | Bryan Michael Balnues | https://www.jotform.com/uploads/macarewecare/222927978372067/6132320681128989145/6132320681128989145_signature_31.png
| 2025-01-20 | 2025-01-20 | Tayasan Rural Health Unit | Tayasan, Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30152 | DOH | porshcabz151@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 244 | 2025-01-08 | GX4 | 110023502 | Due for Xpertcheck | Preventive maintenance - cleaned optics, plunger rods, module bays, filter. | Performed Xpertcheck in all 4 modules | all 4 modules passed Xpertcheck; all 4 modules are operational. |  | Xpertcheck Kit | 12525 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131880583425494201/6131880583425494201_signature_28.png | Portia Cabanting, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6131880583425494201/6131880583425494201_signature_31.png
| 2025-01-20 | 2025-01-15 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30151 | MACARE-HBDC | bloodcollection.tower@gmail.com | 25 |  |  | MCS+ | 01J038 | Unit for maintenance ACDA pump error | Cleaned ACDA rotor roller pump | #NAME? | Unit is operational Unit is under observation If ACDA error occurs, rubbers are subject for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131628513618740097/6131628513618740097_signature_28.png | Ardel Brecinio | https://www.jotform.com/uploads/macarewecare/222927978372067/6131628513618740097/6131628513618740097_signature_31.png
| 2025-01-20 | 2025-01-20 | Makilala Rural Health Unit | P. 16 Poblacion Makilala, North Cotabato | 12 | Cotabato | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30150 | UNOPS | abegailweill17@gmail.com | 22.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 211 | 2025-01-08 | GX-IV 6C 4D | 110023521 | For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install XPC and run XPC on four modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration,Unit is Operational |  | XPERTCHECK CE 5 | DR 12483 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131628446942760505/6131628446942760505_signature_28.png | Shayne S. Belandres | https://www.jotform.com/uploads/macarewecare/222927978372067/6131628446942760505/6131628446942760505_signature_31.png
| 2025-01-20 | 2025-01-17 | Candon City Health Office | Candon City | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30149 | DOH | candoncholaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 637 |  | GXIV-4-D-6C 27 | 838273 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131617990943818294/6131617990943818294_signature_28.png | Michelle Bucalen | https://www.jotform.com/uploads/macarewecare/222927978372067/6131617990943818294/6131617990943818294_signature_31.png
| 2025-01-20 | 2025-01-17 | Sta. Maria Rural Health Unit | Sta. Maria, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30148 | UNOPS | jomarcaberto24@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 156 | 2025-01-08 | GXIV-4-D-6C | 110023500 | For Preventive Maintenance and Calibration | #NAME? | Performed Preventive Maintenance and Calibration | Unit is Fully Operational All Modules Passed XpertCheck |  | XpertCheck Kits | DR 12489 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131616970948586233/6131616970948586233_signature_28.png | Jomar De Guzman Caberto | https://www.jotform.com/uploads/macarewecare/222927978372067/6131616970948586233/6131616970948586233_signature_31.png
| 2025-01-20 | 2025-01-17 | Manabo Rural Health Unit | Manabo, Abra | CAR | Abra | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30147 | DOH | lenmayclairelumiwag15@gmail.com | 28 |  |  | GXIV-4-D-6C | 838274 | For Preventive Maintenance and Troubleshooting | #NAME? | Performed Preventive Maintenance and Troubleshooting | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131614960941981132/6131614960941981132_signature_28.png | Ellen May Claire S. Lumiwag | https://www.jotform.com/uploads/macarewecare/222927978372067/6131614960941981132/6131614960941981132_signature_31.png
| 2025-01-20 | 2025-01-17 | Abra Provincial Hospital | Banguer, Abra | CAR | Abra | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30146 | UNOPS | aphpathologydepartment@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 577 | 2025-01-08 | GXIV-4-D-10C | 110023499 | For Preventive Maintenance and Calibration | #NAME? | Performed Preventive Maintenance and Calibration | Unit is Fully Operational All Modules Passed XpertCheck | Work with Engineer Venzon | XpertCheck Kits | DR 12500 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131612330946618910/6131612330946618910_signature_28.png | Riverlie Valera | https://www.jotform.com/uploads/macarewecare/222927978372067/6131612330946618910/6131612330946618910_signature_31.png
| 2025-01-20 | 2025-01-15 | San Jose Municipal Health Office | Zone 1, Buklod, San Jose | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30145 | DOH | katrinamartirez36@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2111 |  | GX4-6C | 838389 | UPS defective Damaged fan filter | Replaced UPS (AT1000) Replaced fan filter Cleaned cartridge bays, plunger rod Generate Data | Replaced UPS Performed Preventive Maintenance Replaced fan filter | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131610709124279589/6131610709124279589_signature_28.png | Maria Katrina P. Martirez | https://www.jotform.com/uploads/macarewecare/222927978372067/6131610709124279589/6131610709124279589_signature_31.png
| 2025-01-20 | 2025-01-16 | Mariano Marcos State University | Batac Ilocos Norte | 1 | Ilocos Norte | Microbiology Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30144 | MACARE-HBDC | mpolivar@mmsu.edu.ph | 26 |  | 2025-01-10 | GXIV-4-D-10C | 110022474 | For Preventive Maintenance and XpertCheck | #NAME? | Performed Preventive Maintenance and XpertCheck | Unit is Fully Operational All Modules passed XpertCheck | Work with Engineer Venzon | XpertCheck Kits | G 7716 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131607770947278631/6131607770947278631_signature_28.png | Marvielyn Olivar | https://www.jotform.com/uploads/macarewecare/222927978372067/6131607770947278631/6131607770947278631_signature_31.png
| 2025-01-20 | 2025-01-15 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30143 | DOH | sherryannemarrero@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2163 | 2025-01-08 | GXIV-4-D-6C | 838281 | For Module Replacement and Preventive Maintenance | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B2 651808 >>> 776635 - Generated Reports | Performed Preventive Maintenance | - Unit is Partially Operational - Module 1 has already been Reported | Work with Engineer Venzon | GeneXpert Module 776635 | DR 12440 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131606850947689692/6131606850947689692_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6131606850947689692/6131606850947689692_signature_31.png
| 2025-01-20 | 2025-01-15 | Benguet Provincial Health | La Trinidad, Benguet | CAR | Benguet | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30142 | UNOPS | benguetphortdl@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 837 | 2025-01-08 | GXIV-4-D-10C | 110023498 | For Preventive Maintenance and Calibration | #NAME? | Performed Preventive Maintenance and Calibration | #NAME? | Work with Engr. Venzon | XpertCheck Kits | DR 12495 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131603080949495331/6131603080949495331_signature_28.png | Frannie Calixto | https://www.jotform.com/uploads/macarewecare/222927978372067/6131603080949495331/6131603080949495331_signature_31.png
| 2025-01-20 | 2025-01-14 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 30141 | DOH | thisisformyfamily@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1422 | 2025-01-08 | GXIV-4-D-6C | 838286 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 808255 >>>646204 - Generated Reports | Performed Module Replacement | Unit is Fully Operational | Work with Engineer Venzon | GeneXpert Module 646204 | DR 12534 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131601980942114176/6131601980942114176_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/6131601980942114176/6131601980942114176_signature_31.png
| 2025-01-20 | 2025-01-20 | Magsaysay Rural Health Unit | Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30140 | DOH | jkakilgan12@gmail.com | 27 |  | 2025-01-08 | GX-4 6 Colors | 838557 | Module B2 found Service Required thru XPC Calibration | Module Replacement | Replacing B2 with SN: 862892 to SN: 634361 Preventive Maintenance | All modules and units are functional |  | Certified pre-owned GeneXpert Module | 12536 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131571830031011214/6131571830031011214_signature_28.png | Jesusa N. Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/6131571830031011214/6131571830031011214_signature_31.png
| 2025-01-20 | 2025-01-16 | Ospital ng Muntinlupa | Civic Drive, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30136 | UNOPS | osmunmicro@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 303 | 2025-01-08 | GXIV-D-10C | 110023494 | Unit is due for XpertCheck. | Xpertcheck. Preventive Maontenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Performed XpertCheck. Checked System's Functionality. Genrate Reports. | All Modules are Operational. | All Modules Passed in XpertCheck. | XpertCheck Kit. | DR 12496 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131519583612658022/6131519583612658022_signature_28.png | Jesse Jasper Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/6131519583612658022/6131519583612658022_signature_31.png
| 2025-01-20 | 2025-01-17 | Dasmariñas City Health Office 2 | DBB-E, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30139 | DOH | hazelmilan0777@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1264 |  | GXIV-D-6C | 838535 | Monitor Problem. Unit is due for Maintenance. | Troubleshoot | Cleaned Cartridge Bay, Plunger Rod and Filter. Checked/Reset Monitor Connections. Performed Self-test. Checked System's Functionality. Generate Reports. | Monitor is Defective. All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131519583615265004/6131519583615265004_signature_28.png | Hazel Tabucol | https://www.jotform.com/uploads/macarewecare/222927978372067/6131519583615265004/6131519583615265004_signature_31.png
| 2025-01-20 | 2025-01-18 | Dr. Jose Fabella Memorial Hospital | Tayuman St, Manila | NCR | NCR 1st - Manila | Molecular | NO | Relocate | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30137 | MACARE-HBDC | fabella.mollabpcr@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 130 ASSAY: Xpert MTB/RIF Ultra, COUNT: 69 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 591 |  | GXIV-D-10C | 110015482 | GX System for Relocation. | Relocation. | Checked GeneXpert Functionality. Performed Self-test. Relocate GeneXpert From Molecular Lab to Main Lab. Generate Reports. | GX System is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131519583618452188/6131519583618452188_signature_28.png | Renee Lyn Robles | https://www.jotform.com/uploads/macarewecare/222927978372067/6131519583618452188/6131519583618452188_signature_31.png
| 2025-01-20 | 2025-01-16 | University of Perpetual Helps Dalta Medical Center | Alabang-Zapote Road, Las Piñas | NCR | NCR 4th - Las Piñas | Molecular | NO | MAINTENANCE Software Upgrade | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30138 | COMMERCIAL | phmclp.mdl@gmail.com | 27 |  |  | GXIV-D-6C | 831316 | Unit si due for Maintenance. Software Compatibility with the Assay. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Upgrade GX Software from 4.8 to 6.1. Generate Reports. | 1 Modules is Disabled. 3 are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131519583619125809/6131519583619125809_signature_28.png | Donn Dave B. Clerigo | https://www.jotform.com/uploads/macarewecare/222927978372067/6131519583619125809/6131519583619125809_signature_31.png
| 2025-01-20 | 2025-01-17 | St. Luke's Medical Center - BGC | Rizal Dr. BGC Taguig City | NCR | NCR 4th - Taguig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@st.lukes.com.ph |  | 30134 | COMMERCIAL | jkcidiesca@stlukes.com.ph | 22 |  |  | MCS+ | 09C125 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Run simulation Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153611414635/6131422153611414635_signature_28.png | Kenneth Idiesca | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153611414635/6131422153611414635_signature_31.png
| 2025-01-20 | 2025-01-17 | St. Luke's Medical Center - BGC | Rizal Dr. BGC Taguig City | NCR | NCR 4th - Taguig | CI | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@st.lukes.com.ph |  | 30135 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 15 ASSAY: Xpert CT/NG, COUNT: 100 ASSAY: Xpert HPV, COUNT: 123 ASSAY: Xpert MTB/RIF Ultra, COUNT: 224 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 62 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A3/866793 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153611520228/6131422153611520228_signature_28.png | Vincent Raquiza-Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153611520228/6131422153611520228_signature_31.png
| 2025-01-20 | 2025-01-16 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30133 | UNOPS | acho.laboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 ASSAY: Xpert MTB-XDR, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 166 | 2025-01-08 | GX-IV-D-10C | 110023496 | For XpertCheck No black ink when printing | Performed maintenance Run XpertCheck Cleaned ink cartridge | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports Checked printer Cleaned ink cartridge Print test page | Unit is operational. All modules operational. All modules passed. Printer ok. |  |  | DR# 12494 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153613466167/6131422153613466167_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153613466167/6131422153613466167_signature_31.png
| 2025-01-20 | 2025-01-17 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | CI | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@st.lukes.com.ph |  | 30131 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 107 ASSAY: Xpert CT/NG, COUNT: 405 ASSAY: Xpert HBV VL, COUNT: 22 ASSAY: Xpert HPV, COUNT: 527 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1245 ASSAY: Xpert MTB-XDR, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 26 ASSAY: Xpert HIV-1 Viral Load, COUNT: 128 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153617800612/6131422153617800612_signature_28.png | Vincent Raquiza-Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153617800612/6131422153617800612_signature_31.png
| 2025-01-20 | 2025-01-16 | Kalinga Health Hub | JP Rizal St. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30132 | CONSORTIUM | krisna.morante@innovationsch.org | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1616 |  | GX-IV-D-6C | 110008418 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153613235852/6131422153613235852_signature_28.png | Krisna Morante | https://www.jotform.com/uploads/macarewecare/222927978372067/6131422153613235852/6131422153613235852_signature_31.png
| 2025-01-19 | 2025-01-17 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30130 | DOH | renlizros11@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1354 | 2025-01-17 | GxIV-6C | 838570 | Monitor problem | Replace defective monitor SN:CNK9430B0W with monitor SN:CN4801057K | Genexpert monitor repair | 4 operational modules |  | hp monitor SN: CN4801057K | DR#: 12551 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6130976496865509112/6130976496865509112_signature_28.png | Renee Liz Gomez | https://www.jotform.com/uploads/macarewecare/222927978372067/6130976496865509112/6130976496865509112_signature_31.png
| 2025-01-19 | 2025-01-15 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | gmpelaez144@gmail.com |  | 30129 | DOH | onehope834@gmail.com | 26.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 | 2025-01-15 | GxIV-10C | 110023524 | None | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert Cleaning and maintenance MTB-XDR assays installation Checking of errors Errors/problem reporting Data collection | Genexpert Xpertcheck Calibration Genexpert Preventive maintenance | 4 operational modules Calibration Results: All modules passed |  | Xpertcheck CE-5 cartridge Lot #: 1001075092 | DR #: 12484 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6130976496868463908/6130976496868463908_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6130976496868463908/6130976496868463908_signature_31.png
| 2025-01-19 | 2025-01-19 | PRC Port | Port Area Manila | NCR | NCR 1st - Manila | Blood Bank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 30128 | COMMERCIAL | dolly.banares@redcross.org.ph | 28 |  | 2025-01-19 | MCS + | 16K521SPG | Loose centrifuge cover lid screw and was came off during platelet pricedure. | Reinstalled and tighten the centrifuge cover lid screw. Run funtionality test. | Performed troubleshooting on centrifuge assembly. Reinstalled and tighten the cebtrifuge cover lid screw.. Performed functionlity test and simulation run. | Unit is operational.. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6130601572128884353/6130601572128884353_signature_28.png | Dolly banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6130601572128884353/6130601572128884353_signature_31.png
| 2025-01-18 | 2025-01-17 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30127 | MACARE-HBDC | lcpmicrolab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 246 ASSAY: Xpert SARS-CoV-2, COUNT: 451 ASSAY: Xpert HIV-1 Viral Load, COUNT: 83 ASSAY: Xpert SARS-CoV-2, COUNT: 2 |  | GX4 D | 838837 | Due for maintenance B1 879215,B4 780860 reported already failed xpertcheck B3 602135 for reporting | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | B1,B4 disabled B2 operational B3 enable to gathered additional data |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6130176465533126459/6130176465533126459_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6130176465533126459/6130176465533126459_signature_31.png
| 2025-01-18 | 2025-01-17 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | avabichuela3216@gmail.com |  | 30126 | MACARE-HBDC | molecularlab@phc.gov.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 653 ASSAY: Others, COUNT: 6 ASSAY: Xpert SARS-CoV-2, COUNT: 861 |  | GX16 | 838655 | Due for maintenance A2 666159,B2 625297,C2 639257 failed xpertcheck. Reported already | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Performed selftest Gathered reports Test run unit | Performed maintenance | A2 disabled due to ultrasonic error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6130175995531157028/6130175995531157028_signature_28.png | Philippine Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6130175995531157028/6130175995531157028_signature_31.png
| 2025-01-17 | 2024-12-10 | Bindoy District Hospital | Bindoy, Negros Oriental | 7 | Negros Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30125 | DOH | bindoydistricthosp.laboratory@gmail.com | 25 |  | 2024-11-28 | GX4 6C | 838475 | Module B2/784239 failed latest Xpertcheck. | #NAME? | - Replaced module B2 with new module SN 676631. | 4 modules operational |  | Genexpert Module 6c | 12383 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6129110033424594402/6129110033424594402_signature_28.png | Lemuel T.Quirit | https://www.jotform.com/uploads/macarewecare/222927978372067/6129110033424594402/6129110033424594402_signature_31.png
| 2025-01-17 | 2025-01-17 | Pontevedra Rural Health Unit | Pontevedra, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30124 | DOH | pontevedra.laboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 706 | 2025-01-08 | GX4 | 838491 | Defective module B2 with SN 862258 | Preventive maintenance - cleaned optics, filter, module bay areas, plunger rods, archive tests. | - Replaced module B2 with new module SN 891156 - Provide 2 new filters | 4 modules operational |  | 900-0491RS- Genexpert module 6c | 12547 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6129032786219878549/6129032786219878549_signature_28.png | Love Joy Sulleza | https://www.jotform.com/uploads/macarewecare/222927978372067/6129032786219878549/6129032786219878549_signature_31.png
| 2025-01-17 | 2025-01-17 | Padre Burgos Rural Health Unit | Pob. Padre Burgos | 8 | Southern Leyte | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30123 | UNOPS | keishaelezze14@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 223 | 2025-01-17 | GXIV 10C | 110023514 | Due for Calibration | #NAME? | #NAME? | Unit it Operational |  | 1 box Xpertcheck kit | 12482 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6129023437892744899/6129023437892744899_signature_28.png | Keisha Elezze Batalon | https://www.jotform.com/uploads/macarewecare/222927978372067/6129023437892744899/6129023437892744899_signature_31.png
| 2025-01-17 | 2025-01-17 | Padre Burgos Rural Health Unit | Pob. Padre Burgos | 8 | Southern Leyte | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30122 | UNOPS | keishaelezze14@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 223 | 2025-01-17 | GXIV 10C | 110023514 | Due for Calibration | #NAME? | #NAME? | Unit it Operational |  | 1 box Xpertcheck kit | 12482 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6129022657899725793/6129022657899725793_signature_28.png | Keisha Elezze Batalon | https://www.jotform.com/uploads/macarewecare/222927978372067/6129022657899725793/6129022657899725793_signature_31.png
| 2025-01-17 | 2025-01-17 | Zamboanga City Medical Centre | Zamboanga City, Zamboanga del sur | 9 | Zamboanga del Sur | Molecular Laboratory | YES | INSTALLATION | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 30121 | COMMERCIAL | cabayacruzdela@gmail.com | 26 |  | 2025-01-10 | GX 4 | 110025742 | None | Installation of GX 4(10c) and conducted genexpert self test. | Installation of GX 4(10c) and conducted genexpert self test. | All modules are available and ready to use | None | GXIV-4-D-10c, GX-DESKTOP, GX-MONITOR, FC-IT-PRINTER, FC-IT-UPS-ACCU1K, SVC-1KVA | 12555 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128975325213227534/6128975325213227534_signature_28.png | Dela P. Cabayacruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6128975325213227534/6128975325213227534_signature_31.png
| 2025-01-17 | 2025-01-17 | Malay Primary Care Facility | Malay, Aklan | 6 | Aklan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30120 | DOH | mhomalayclinicallaboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1127 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 |  | 6C GXIV-4D | 838457 | #NAME? | -Perform preventive maintenance; filter, catridge bay, plunger rod, optic slot | #NAME? | Module B3 for reporting. Other modules is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128927770361529049/6128927770361529049_signature_28.png | May Soledad | https://www.jotform.com/uploads/macarewecare/222927978372067/6128927770361529049/6128927770361529049_signature_31.png
| 2025-01-16 | 2025-01-14 | Southern Philippines Medical Center | J.P Laurel Avenue Bajada, Davao City | 11 | Davao del Sur | IM | NO | For return /turnover | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 30119 | COMMERCIAL | spmcimdept@gmail.com | 24.5 |  | 2024-11-22 | MCS+ | 15H034 | For return of Centrifuge Assembly Assy | Return and turnover the Defective Centrifuge Assembly Assy with SN:5014 | Return and turnover the Defective Centrifuge Assembly Assy with SN:5014 | Successfully return the defective Centrifuge Assembly Assy |  | 39627-00 Centrifuge Assembly Assy | 12348 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128423877818904893/6128423877818904893_signature_28.png | DONNA MAE TONG-AN | https://www.jotform.com/uploads/macarewecare/222927978372067/6128423877818904893/6128423877818904893_signature_31.png
| 2025-01-16 | 2025-01-16 | Manila Med - Medical Center Manila | United Nations Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 30118 | CONSORTIUM | microbiologylab.manilamed@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1255 |  | GX IV 6C-D | 835540 | Unit due for maintenance | Performed preventive maintenance All modulea passed reliability test Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test Archived results | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128239835016167167/6128239835016167167_signature_28.png | Frances Velarde | https://www.jotform.com/uploads/macarewecare/222927978372067/6128239835016167167/6128239835016167167_signature_31.png
| 2025-01-16 | 2025-01-16 | DR. CATALINO GALLEGO NAVA PROVINCIAL HOSPITAL | Jordan , Guimaras | 6 | Guimaras | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30117 | UNOPS | dcgnphlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2025-01-08 | GX-IV | 110023540 | FOR CALIBRATION | COLLECT IQR AND CALIBRATION DATA | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK CE-IVD 5 Cartridges/kit | 12517 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128212428026619821/6128212428026619821_signature_28.png | Glaed Joy Joplo , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6128212428026619821/6128212428026619821_signature_31.png
| 2025-01-16 | 2025-01-16 | Mambajao Municipal Health Office PCF | Poblacion, Mambajao | 10 | Camiguin | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30116 | UNOPS | mirasolflores76@gmail.com | 28 |  | 2025-01-08 | GX-4 10 Colors | 110023519 | None | None | Preventive Maintenance XpertCheck Calibration | All modules and units are functional | All modules passed in XPC | XpertCheck, CE IVD 5 Cartridges | 12475 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128191810919126907/6128191810919126907_signature_28.png | Mirasol B. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6128191810919126907/6128191810919126907_signature_31.png
| 2025-01-16 | 2025-01-16 | FEU - Dr. Nicanor Reyes Medical Foundation | Regalado Avenue, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 30115 | CONSORTIUM | kmtejoso@feu-nrmf.edu.ph | 26 | ASSAY: Xpert HPV, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 404 |  | GX IV 6C-D | 840975 | Unit due for maintenance | Performed preventive maintenance All modules passed reliability test Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test Archived results | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128155225014647487/6128155225014647487_signature_28.png | Kenneth Anthony M. Tejoso | https://www.jotform.com/uploads/macarewecare/222927978372067/6128155225014647487/6128155225014647487_signature_31.png
| 2025-01-16 | 2025-01-16 | Maco Rural Health Unit | Poblacion, Maco, Davao de Oro | 11 | Davao de Oro | TB-DOTS | NO | REPAIR CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30114 | DOH | edlienkai26@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1122 | 2025-01-08 | GX-IV 6C 4D | 838500 | Module B3 encounter multiple temperature related errors Due for Preventive Maintenance | Replaced the Defective Module SN:861093 with New Module SN:866293 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | GeneXper Module 6C | 12544 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128149044219884534/6128149044219884534_signature_28.png | Edlien Kai Prado,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6128149044219884534/6128149044219884534_signature_31.png
| 2025-01-16 | 2025-01-16 | Quezon City General Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30113 | UNOPS | jomarilitonjua@gmail.com | 25 |  | 2025-01-08 | GXIV-10C | 110023495 | Unit for maintenance Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: passed B2: passed B3: passed B4: passed |  | XpertCheck Kit | 12492 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128123864915173954/6128123864915173954_signature_28.png | Jomari Bryan Litonjua | https://www.jotform.com/uploads/macarewecare/222927978372067/6128123864915173954/6128123864915173954_signature_31.png
| 2025-01-16 | 2025-01-16 | BUENAVISTA EMERGENCY HOSPITAL | Mclain, Buenavista Guimaras | 6 | Guimaras | LABORATORY | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30112 | UNOPS | campo_janice@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2025-01-08 | GX-IV | 110023539 | FOR XPERTCHECK CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK CE-IVD 5 Catridges/kit | 12521 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128093866912063969/6128093866912063969_signature_28.png | Janice de la Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6128093866912063969/6128093866912063969_signature_31.png
| 2025-01-16 | 2025-01-16 | LEON RURAL HEALTH UNIT | Leon, Iloilo | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30111 | UNOPS | camiquejolyn10@gmail.com | 27 |  |  | Gx4 | 110023533 | Due date for xpertcheck calibration | Xpertcheck calibration Clean Fan filter Gather data IQR, system log | Xpertcheck calibration Clean Fan filter Gather data IQR, system log | Systems operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128079364911058620/6128079364911058620_signature_28.png | Jolyn C. Maloco | https://www.jotform.com/uploads/macarewecare/222927978372067/6128079364911058620/6128079364911058620_signature_31.png
| 2025-01-16 | 2025-01-16 | Tubungan Rural Health Unit | Tubungan, Iloilo | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30110 | UNOPS | sheryltacadao@gmail.com | 28 |  |  | Gx4 | 110023534 | Due date for xpertcheck calibration | Xpertcheck calibration Clean Fan Filter Gather data IQR, System log | Xpertcheck calibration Clean Fan Filter Gather data IQR, System log | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128038028019184689/6128038028019184689_signature_28.png | SHERYL TACADAO | https://www.jotform.com/uploads/macarewecare/222927978372067/6128038028019184689/6128038028019184689_signature_31.png
| 2025-01-16 | 2025-01-15 | Jamindan Primary Care Facility | Jamindan, Capiz | 6 | Capiz | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30109 | DOH | louiseannaries.laa@gmail.com | 24 |  | 2025-01-08 | GX4 | 110023543 | Due for Xpertcheck | Preventive maintenance - cleaned optics, plunger rods, module bays, filter. | Perform Xpertcheck in all 4 modules | 4 modules passed Xpertcheck. 4 modules operational |  | Xpertcheck Kit | 12511 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128023302492855504/6128023302492855504_signature_28.png | Louise Ann B. Aries | https://www.jotform.com/uploads/macarewecare/222927978372067/6128023302492855504/6128023302492855504_signature_31.png
| 2025-01-16 | 2025-01-16 | Carles Primary Care Facility | Carles, Iloilo | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30108 | DOH | vangieyuro@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 131 | 2025-01-08 | GX4 | 110023526 | Due for Xpertcheck | Preventive maintenance - cleaned optics, plunger rods, module bays, filter. | Performed Xpertcheck in all 4 modules | 4 modules passed Xpertcheck 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6128023002496797097/6128023002496797097_signature_28.png | VANGIE C. YURO | https://www.jotform.com/uploads/macarewecare/222927978372067/6128023002496797097/6128023002496797097_signature_31.png
| 2025-01-15 | 2025-01-15 | San Lorenzo Rural Health Unit | San Lorenzo, Guimaras | 6 | Guimaras | Laboratory | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30107 | UNOPS | rhusanlorenzo@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-IV 10C | 110023537 | FOR INSTALLATION(GX UNIT TRANSFERRED FROM NUEVA VALENCIA RURAL HEALTH UNIT) | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127529046911419993/6127529046911419993_signature_28.png | Ellen Pearl Fernandez Talaban | https://www.jotform.com/uploads/macarewecare/222927978372067/6127529046911419993/6127529046911419993_signature_31.png
| 2025-01-15 | 2025-01-15 | NUEVA VALENCIA RURAL HEALTH UNIT | NUEVA VALENCIA, GUIMARAS | 6 | Guimaras | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30106 | UNOPS | xxxrhingringxxx@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 123 | 2025-01-08 | GX-IV | 110023537 | #NAME? | COLLECT IQR, SYSTEM LOG, ARCHIVE, AND CALIBRATION DATA | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK CE-IVD 5 Cartridges/kit | 12519 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127527636914283574/6127527636914283574_signature_28.png | Rhea April Q Jaleco | https://www.jotform.com/uploads/macarewecare/222927978372067/6127527636914283574/6127527636914283574_signature_31.png
| 2025-01-15 | 2025-01-15 | Cuartero Primary Care Facility | Cuartero, Capiz | 6 | Capiz | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30105 | DOH | cuarteroruralhealthunit@gmail.com.ph | 24 |  | 2025-01-08 | GX4 | 110022580 | Due for Xpertcheck | Preventive maintenance - cleaned optics, filter, plunger rods, module bays. | Performed Xpertcheck in all 4 modules | 4 modules passed Xpertcheck 4 modulea operational |  | Xpertcheck Kit | 12509 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127374438974574683/6127374438974574683_signature_28.png | Marilyn M Cornepillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6127374438974574683/6127374438974574683_signature_31.png
| 2025-01-15 | 2025-01-15 | Quezon City General Hospital (1) | Seminary Rd. Project 8, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 30104 | UNOPS | treatmenthub.qcgh@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 196 |  | GXIV 10C D | 110012766 | Unit due for maintenance | Performed preventive maintenance All modules passed reliability test Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test Archived results | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127353733718848465/6127353733718848465_signature_28.png | Arizza Sheena B. Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6127353733718848465/6127353733718848465_signature_31.png
| 2025-01-15 | 2025-01-15 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | gmpelaez144@gmail.com |  | 30103 | DOH | onehope834@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 |  | GxIV-6C | 838528 | None | Perform Genexpert cleaning and maintenance Checking of errors/problem Errors/problem reporting Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127331942315629482/6127331942315629482_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6127331942315629482/6127331942315629482_signature_31.png
| 2025-01-15 | 2025-01-15 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | gmpelaez144@gmail.com |  | 30102 | DOH | onehope834@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 |  | GxIV-6C | 838528 | None | Perform Genexpert cleaning and maintenance Checking of errors/problem Errors/problem reporting Data collection | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127331492315846297/6127331492315846297_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6127331492315846297/6127331492315846297_signature_31.png
| 2025-01-15 | 2025-01-15 | SIBUNAG RURAL HEALTH UNIT | Sibunag Guimaras | 6 | Guimaras | Laboratory | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30101 | UNOPS | rhodatrance1994@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2025-01-08 | GX-IV 10C | 110023538 | FOR XPERTCHECK CALIBRATION | COLLECT IQR AND CALIBRATION DATA | #NAME? | EQUIPMENT IS OPERATIONAL |  | Xpertcheck CE-IVD 5 Cartridges/kit | 12520 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127330282766521041/6127330282766521041_signature_28.png | Ma. Rhoda Lyn Villarta | https://www.jotform.com/uploads/macarewecare/222927978372067/6127330282766521041/6127330282766521041_signature_31.png
| 2025-01-15 | 2025-01-15 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | dy_rosemarie71@gmail.com |  | 30100 | DOH | dyrosemarie71@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1248 | 2025-01-09 | GX-IV 6C 4D | 838401 | Monitor Problem Due for Preventive Maintenance | Replaced the Defective Module SN:CNK94301WM with Monitor SN: CN461016VJ Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Monitor Replacement and Preventive Maintenance | Monitor replaced successfully All Modules are operational |  | FC-E-MONITOR | 12553 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127302653217198962/6127302653217198962_signature_28.png | Rhodora Mae Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/6127302653217198962/6127302653217198962_signature_31.png
| 2025-01-15 | 2025-01-15 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30099 | DOH | sanjoseantiquegxsite@gmail.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1100 ASSAY: Xpert HIV-1 Viral Load, COUNT: 90 |  | 6C GXIV-4D | 838412 | Due for PM | -Perform preventive maintenance of machine; plunger rod, optic slot, module bay, filter -Archived test results | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127285753526476901/6127285753526476901_signature_28.png | Oscar Delano Salazar, IV | https://www.jotform.com/uploads/macarewecare/222927978372067/6127285753526476901/6127285753526476901_signature_31.png
| 2025-01-15 | 2025-01-15 | AFPMC - AFRIMS 6C | V. Luna, Quezon City | NCR | NCR 2nd - Quezon City | Mo | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 30098 | MACARE-HBDC | mtvalderama@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 105 ASSAY: Xpert SARS-CoV-2, COUNT: 32 |  | GX IV 6C-D | 805294 | Unit due for maintenance | Performed preventive maintenance Modules 1,3,4 apssed reliability test Archives and purged 2014-2020 results | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test Archived results | Unit is operational Module 2 not working but reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127272813714238422/6127272813714238422_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/6127272813714238422/6127272813714238422_signature_31.png
| 2025-01-15 | 2025-01-15 | AFPMC - AFRIMS 10C | V. Luna, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 30097 | MACARE-HBDC | mtvalderama@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 553 ASSAY: Xpert SARS-CoV-2, COUNT: 829 ASSAY: Xpert HCV VL, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 ASSAY: Xpert HIV-1 Viral Load, COUNT: 112 |  | GXIV 10C D | 845462 | Unit due for maintenance | Performed preventive maintenance All modules passed reliability test Archived results | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test Archived results | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127265973711125581/6127265973711125581_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/6127265973711125581/6127265973711125581_signature_31.png
| 2025-01-15 | 2025-01-15 | Kairos Laboratories and Diagnostics, Inc. | A. Mabini St. Ermita Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30096 | CONSORTIUM | kairos.jamartinez@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 110 | 2025-01-10 | GXIV (2 modules) | 110022622 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck Kit | DR#12561 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127262049208260464/6127262049208260464_signature_28.png | Jude martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/6127262049208260464/6127262049208260464_signature_31.png
| 2025-01-15 | 2025-01-15 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30095 | MACARE-HBDC | bloodcollection.tower@gmail.com | 25 |  |  | MCS+ | 01J038 | Unit for maintenance ACDA pump error | Cleaned ACDA rotor roller pump | #NAME? | Unit is operational Unit is under observation If ACDA error occurs, rubbers are subject for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127253114915263194/6127253114915263194_signature_28.png | Ardel Brecinio | https://www.jotform.com/uploads/macarewecare/222927978372067/6127253114915263194/6127253114915263194_signature_31.png
| 2025-01-15 | 2025-01-15 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30094 | MACARE-HBDC | bloodcollection.tower@gmail.com | 25 |  |  | MCS+ | 08B049 | Unit for maintenance Unit for check up | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127252254914010606/6127252254914010606_signature_28.png | Ardel Brecinio | https://www.jotform.com/uploads/macarewecare/222927978372067/6127252254914010606/6127252254914010606_signature_31.png
| 2025-01-15 | 2025-01-15 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 30093 | UNOPS | laureanjo@yahoo.com | 26 |  |  | Gx4 | 110023513 | Due date for xpertcheck calibration | Xpertcheck Calibration Clean Fan Filter Clean Module Bay | Xpertcheck Calibration Clean Fan Filter Clean Module Bay | Systems operational All modules passed in xpertheck calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127247778328086900/6127247778328086900_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6127247778328086900/6127247778328086900_signature_31.png
| 2025-01-15 | 2025-01-15 | Pamplona Rural Health Unit | Centro, Pamplona, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30092 | UNOPS | wickizena@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 295 | 2025-01-08 | GXIV-10C | 110023503 | Due for Calibration | Performed Xpertcheck (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Xpertcheck Kit | 12501 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127230209227308158/6127230209227308158_signature_28.png | Wicki Zena Gerona | https://www.jotform.com/uploads/macarewecare/222927978372067/6127230209227308158/6127230209227308158_signature_31.png
| 2025-01-15 | 2025-01-15 | San Rafael Rural Health Unit | San Rafael, Iloilo | 6 | Iloilo | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30091 | DOH | apilmayfairh@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 89 | 2025-01-08 | GX4 | 110023536 | Due for Xpertcheck. | PM- cleaned optics, filter, module bays plunger rods. | Perform Xpertcheck in all 4 modules | 4 modules operational, 4 modules passed Xpertcheck. |  | Xpertcheck Kit | 12514 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127223987518805388/6127223987518805388_signature_28.png | May Faith B Apil | https://www.jotform.com/uploads/macarewecare/222927978372067/6127223987518805388/6127223987518805388_signature_31.png
| 2025-01-15 | 2025-01-15 | Dumarao Rural Health Unit | Dumarao, Iloilo | 6 | Iloilo | TB Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30090 | DOH | doregoshielamarie@gmail.com | 24 |  | 2025-01-08 | GX4 | 110023542 | #NAME? | . | #NAME? | - all 4 modules passed Xpertcheck using laptop. - CPU problem confirmed, to be reported to Cepheid for replacement. |  | Xpertcheck kit | 12510 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127215647518471569/6127215647518471569_signature_28.png | Shiela Marie Dorego | https://www.jotform.com/uploads/macarewecare/222927978372067/6127215647518471569/6127215647518471569_signature_31.png
| 2025-01-15 | 2025-01-15 | San Isidro Rural Health Unit | San Isidro Northern Samar | 8 | Northern Samar | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30089 | UNOPS | hsdelamitas@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 144 | 2025-01-15 | GXIV 10C | 110023515 | Due for Calibration | #NAME? | #NAME? | Unit is operational |  | 1 box Xpertcheck Kit | 12481 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127191855311062103/6127191855311062103_signature_28.png | Hazel G. Salinas | https://www.jotform.com/uploads/macarewecare/222927978372067/6127191855311062103/6127191855311062103_signature_31.png
| 2025-01-15 | 2025-01-15 | Socorro Rural Health Unit | Brgy. Navarro, Socorro | 13 | Surigao del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30088 | UNOPS | dynamite.tux@gmail.com | 27 |  | 2025-01-08 | GX-4 10 Colors | 110023523 | None | None | Preventive Maintenance XpertCheck Calibration | All modules and units are functional | All modules passed in XPC Calibration | XpertCheck, CE-IVD 5 Cartridges/Kit | 12478 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127134478432748397/6127134478432748397_signature_28.png | James Anastacio | https://www.jotform.com/uploads/macarewecare/222927978372067/6127134478432748397/6127134478432748397_signature_31.png
| 2025-01-15 | 2025-01-15 | Jordan Primary Care Facility | Jordan, Guimaras | 6 | Guimaras | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30087 | UNOPS | badette1815@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2025-01-08 | GX-IV 10C | 110023528 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK, CE-IVD 5 Cartridges/kit | 12518 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127128185428219372/6127128185428219372_signature_28.png | Bernadette I. Galila, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6127128185428219372/6127128185428219372_signature_31.png
| 2025-01-15 | 2025-01-14 | Rizal Medical Center | Pasig Blvd, Paaig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | Stephenbryan.1621@gmail.com |  | 30086 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 |  |  | GX-XVI-D-6C | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules A1/860841, B1/701496, C2/668748, C3/662448, D1/860832, D2/620837, D3/860742 for replacement. Modules A2/860940, B3/702534 and D4/661946 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127125713613644148/6127125713613644148_signature_28.png | Justine Lorenz V. Falzado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6127125713613644148/6127125713613644148_signature_31.png
| 2025-01-15 | 2025-01-14 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | Stephenbryan.1621@gmail.com |  | 30085 | COMMERCIAL | rmcbloodbank@gmail.com | 22 |  |  | MCS+ | 15H045G | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Run simulation Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6127125713617473627/6127125713617473627_signature_28.png | Ferdinand Barroga, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6127125713617473627/6127125713617473627_signature_31.png
| 2025-01-14 | 2025-01-14 | Tampakan Rural Health Unit | Poblacion, Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30084 | DOH | jmedtech_lab@yahoo.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 523 | 2025-01-14 | GxIV-6C | 838544 | Defective Monitor Module 2 and 3 for Service Required | -Replace defective module 2 SN:860765 with new module SN:210025707 -Replace defective module 3 SN:807344 with new module SN:619178 -Replace defective monitor SN:CNK9430BT8 with hp monitor SN:CN4610177J -Genexpert Cleaning and maintenance -Checking of errors -Data collection | Perform Genexpert Module repair Perform Genexpert Monitor repair Preventive maintenance | 4 operational modules |  | Genexpert 6-color module SN:210025707 SN: 619178 Hp monitor SN: CN4610177j | DR #'s: 12533, 12550, and 12532 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126676996869142227/6126676996869142227_signature_28.png | Jovi Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/6126676996869142227/6126676996869142227_signature_31.png
| 2025-01-14 | 2025-01-14 | Iriga City Health Office | Sta. Cruz Sur, Iriga | 5 | Camarines Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30083 | UNOPS | hannahalevon1@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 763 |  | GX4-10C | 110023511 | GX4-10C for Calibration | XpertCheck Result: B1/210117798 - passed B2/210121970 - passed B3/210119501 - passed B4/210120184 - passed Cleaned cartridge bays, plunger rod, icore Generate Data | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126559416118252664/6126559416118252664_signature_28.png | Hannah Kirsten Novela | https://www.jotform.com/uploads/macarewecare/222927978372067/6126559416118252664/6126559416118252664_signature_31.png
| 2025-01-14 | 2025-01-14 | Iriga City Health Office | Sta. Cruz Sur, Iriga | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30082 | DOH | hannahalevon1@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1219 |  | GX4-6C | 838391 | Defective keyboard | Replaced defective keyboard Cleaned cartridge bays, plunger rod, icore Generate Data | Replaced keyboard Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126558876113151907/6126558876113151907_signature_28.png | Hannah Kirsten Novela | https://www.jotform.com/uploads/macarewecare/222927978372067/6126558876113151907/6126558876113151907_signature_31.png
| 2025-01-14 | 2025-01-14 | Kalinga Provincial Hospital | Tabuk City, Kalinga | CAR | Kalinga | roxasmatthew659@yahoo.com.ph | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30081 | DOH | leombelgica139@gmail.com | 26 |  |  | GXIV-6C | 838466 | GX Software cannot open | Reinstalled GX Software (5.1) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126505262376839174/6126505262376839174_signature_28.png | Franchet L. Campilis | https://www.jotform.com/uploads/macarewecare/222927978372067/6126505262376839174/6126505262376839174_signature_31.png
| 2025-01-14 | 2025-01-14 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | TB dots | YES | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | eamc.biomed@gmail.com |  | 30080 | MACARE-HBDC | roseannparnes@yahoo.com | 28 |  |  | GX 4D 10C | 110014085 | Due xpertcheck March 22,2025. Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational | For quotation of xpertcheck kits ,need to perform xpertcheck march 22,2025 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126472702312247375/6126472702312247375_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/6126472702312247375/6126472702312247375_signature_31.png
| 2025-01-14 | 2025-01-14 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta.Cruz, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30079 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 229 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 16 |  | GX16-D-10C | 110013856 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, And Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126458846325805072/6126458846325805072_signature_28.png | Ervin Emmett P Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6126458846325805072/6126458846325805072_signature_31.png
| 2025-01-14 | 2025-01-14 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta.Cruz, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30078 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 818 |  | GXIV-D-10C | 110014099 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | 1 Module is Disabled. 3 Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126458186327692712/6126458186327692712_signature_28.png | Ervin Emmett P Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6126458186327692712/6126458186327692712_signature_31.png
| 2025-01-14 | 2025-01-14 | Calinog Rural Health Unit | Calinog, Iloilo | 6 | Iloilo | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30077 | DOH | icjssca@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 60 | 2025-01-08 | GX4 | 110023535 | Due for Xpertcheck. | Preventive maintenance - cleaned optics, filter, plunger rods, module bays. | Performed Xpertcheck in all 4 modules. | 4 modules operational |  | Xpertcheck Kit | 12506 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126457053729458738/6126457053729458738_signature_28.png | IC Jessica Coronado | https://www.jotform.com/uploads/macarewecare/222927978372067/6126457053729458738/6126457053729458738_signature_31.png
| 2025-01-14 | 2025-01-14 | Surigao City Health Office | Brgy. San Juan, Surigao City | 13 | Surigao del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30076 | UNOPS | sanjuandistrictcenter@gmail.com | 29 |  | 2025-01-08 | GX-4 10 colors | 110023522 | None | None | Preventive Maintenance XpertCheck Calibration | All modules and units are functional |  | XpertCheck, CE-IVD 5 Cartridges/ Kit | 12476 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126438648971800879/6126438648971800879_signature_28.png | Anrose Ogdol | https://www.jotform.com/uploads/macarewecare/222927978372067/6126438648971800879/6126438648971800879_signature_31.png
| 2025-01-14 | 2025-01-14 | Southern Philippines Medical Center | J.P Laurel Avenue, Bajada Davao City | 11 | Davao del Sur | RHIVDA | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30075 | UNOPS | spmcrhivda@gmail.com | 24.5 |  | 2025-01-08 | GX-IV 10C 4D | 110023506 | Due For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install XPC version 1.5a Run XpertCheck on four modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules passed and machine is operational |  | XPERTCHECK CE 5 | 12480 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126428583308063576/6126428583308063576_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6126428583308063576/6126428583308063576_signature_31.png
| 2025-01-14 | 2025-01-14 | Southern Philippines Medical Center | J.P Laurel Avenue, Davao City | 11 | Davao del Sur | RHIVDA | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30074 | UNOPS | spmcrhivda@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 832 ASSAY: Xpert HIV-1 Viral Load, COUNT: 754 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert SARS-CoV-2, COUNT: 4 ASSAY: Xpert HCV VL, COUNT: 2 | 2025-01-08 | GX-IV 10C 4D | 110023505 | For Preventive Maintenance and due for XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install XPC Version 1.5a Run XPC at four Modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration |  | XPERTCHECK CE 5 | 12479 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126428223309263398/6126428223309263398_signature_28.png | Christelle Clarisse Sienes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6126428223309263398/6126428223309263398_signature_31.png
| 2025-01-14 | 2025-01-14 | CEBU DOCTORS UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30073 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 79 ASSAY: Xpert CT/NG, COUNT: 241 ASSAY: Xpert HBV VL, COUNT: 188 ASSAY: Xpert HPV, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1127 ASSAY: Xpert HCV VL, COUNT: 152 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GX-IV | 110008524 | MONTHLY PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126412105581138000/6126412105581138000_signature_28.png | Gieril Respecia,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6126412105581138000/6126412105581138000_signature_31.png
| 2025-01-14 | 2025-01-14 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30072 | UNOPS | bmcrhivda@gmail.com | 28 |  | 2025-01-08 | GXIV-10C | 110023507 | Unit for maintenance Unit for calibration | NA | #NAME? | Unit is operational A1: passed A2: passed A3: service required A4: passed |  | XpertCheck Kit | 12491 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126409024911018935/6126409024911018935_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6126409024911018935/6126409024911018935_signature_31.png
| 2025-01-14 | 2025-01-14 | Cauayan City Health Office | Cauayan City, Isabela | 2 | Isabela | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30071 | UNOPS | cho1lab14@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 634 ASSAY: Xpert SARS-CoV-2, COUNT: 2 | 2025-01-08 | GXIV-10C | 110023504 | Due for Calibration | Performed Xpertcheck (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | 12490 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126400385441304579/6126400385441304579_signature_28.png | Arlyn Daggao | https://www.jotform.com/uploads/macarewecare/222927978372067/6126400385441304579/6126400385441304579_signature_31.png
| 2025-01-14 | 2025-01-14 | Chinese General Hospital | Blumentritt, Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com | jmescolano@cghmc.com.ph |  | 30070 | MACARE-HBDC | cghbloodbank2016@gmail.com | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Performed valve diagnostics running. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126388435111957820/6126388435111957820_signature_28.png | Joan escolano | https://www.jotform.com/uploads/macarewecare/222927978372067/6126388435111957820/6126388435111957820_signature_31.png
| 2025-01-14 | 2025-01-14 | Miagao Rural Health Unit | Miagao, Iloilo | 6 | Iloilo | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30069 | UNOPS | joannaleesadiong@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 93 | 2025-01-08 | 10C GXIV-4D | 110023531 | Due for XPC | #NAME? | #NAME? | Module B1 for replacement; service required. Other modules passed and is functional. |  | XpertCheck-CE-5 Lot no. 1001075092 Exp. June 01, 2025 | DR no. 12522 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126340521858193103/6126340521858193103_signature_28.png | Joanna Lee Sadiong | https://www.jotform.com/uploads/macarewecare/222927978372067/6126340521858193103/6126340521858193103_signature_31.png
| 2025-01-14 | 2025-01-14 | Concepcion Rural Health Unit | Concepcion, Iloilo | 6 | Iloilo | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30068 | DOH | jescelmaymesias@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 42 | 2025-01-08 | GX4 | 110023527 | Due for Xpertcheck | Preventive Maintenance - clean optics, filter, plunger rods, modules bays. | Perform Xpertcheck on all 4 modules. | 4 modules operational |  | Xpertcheck Kit | 12523 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126337148922858061/6126337148922858061_signature_28.png | Jescel May Melgar | https://www.jotform.com/uploads/macarewecare/222927978372067/6126337148922858061/6126337148922858061_signature_31.png
| 2025-01-14 | 2025-01-14 | Ramon Rural Health Unit | Bugallon Proper, Ramon, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 30067 | DOH | rhuramonlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 402 | 2025-01-08 | GXIV-6C | 838480 | Module B4 Temperature Error | Performed Module Replacement B4(715915 - 641322) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Module | 12535 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126301320026237879/6126301320026237879_signature_28.png | Janette Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/6126301320026237879/6126301320026237879_signature_31.png
| 2025-01-14 | 2025-01-14 | Oton Municipal Health Office Diagnostic Center | Oton, Iloilo | 6 | Iloilo | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30066 | UNOPS | gracie05191980@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 144 | 2025-01-08 | 10C GXIV-4D | 110023530 | Due for XPC | #NAME? | #NAME? | All modules passed XpertCheck and is functional. |  | XpertCheck-CE-5 Lot no. 1001075092 Exp. June 01, 2025 | DR no. 12513 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126262088317454412/6126262088317454412_signature_28.png | Mary Grace Ong-Lumasag | https://www.jotform.com/uploads/macarewecare/222927978372067/6126262088317454412/6126262088317454412_signature_31.png
| 2025-01-14 | 2025-01-14 | Dumangas Primary Health Care Facility | Dumangas, Iloilo | 6 | Iloilo | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30065 | DOH | marilyn_dusaran@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 81 | 2025-01-08 | GX4 | 110023532 | Due for Xpertcheck | . | Perform Xpertcheck on all 4 modules | all 4 modules Passed the Xpertcheck; 4 modules operational |  |  | 12516 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6126251091418996085/6126251091418996085_signature_28.png | marilyn p. dusaran | https://www.jotform.com/uploads/macarewecare/222927978372067/6126251091418996085/6126251091418996085_signature_31.png
| 2025-01-13 | 2025-01-13 | Isulan Rural Health Unit | Poblacion Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30064 | DOH | nancymoscoso18@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 651 | 2025-01-13 | GxIV-6C | 110023520 | None | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors/problem Errors/problem reporting Data collection | Genexpert Xpertcheck Calibration Genexpert Preventive maintenance | 4 operational modules Calibration Results: All modules passed |  | Xpertcheck CE-5 cartridge Lot#1001075092 | DR#: 12485 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125824116861845435/6125824116861845435_signature_28.png | Nancy A. Moscoso | https://www.jotform.com/uploads/macarewecare/222927978372067/6125824116861845435/6125824116861845435_signature_31.png
| 2025-01-13 | 2025-01-13 | Mandaue City Health Office | Mandaue City Cebu | 7 | Cebu | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | mendozarey42@gmail.com |  | 30063 | DOH | mendozarey42@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4036 ASSAY: Xpert HCV VL, COUNT: 86 ASSAY: Xpert HIV-1 Viral Load, COUNT: 467 | 2025-01-08 | GX-IV | 838508 | #NAME? | #NAME? | -PERFORM PM PROCEDURE -PERFORM MODULE REPLACEMENT: B1 DEF SN: 629957 ; NEW SN: 698586 B3 DEF SN: 695932 ; NEW SN: 708784 | EQUIPMENT IS OPERATIONAL |  | Certified Pre-owned GeneXpert module | 12538, 12539 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125666806715521463/6125666806715521463_signature_28.png | Rey C Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6125666806715521463/6125666806715521463_signature_31.png
| 2025-01-13 | 2025-01-13 | Atimonan Rural Health Unit | G. Boulevard, Brgy. Zone 1, Atimonan, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 30062 | DOH | AMCL2008@yqhoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 662 | 2025-01-08 | GX IV 6C-D | 838206 | Module B2 for replacement | Performed Preventive maintenance Performed module replacement on B2 | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced module B2 807294 -> 794896 | Unit is operational All modules working |  | GX module 6C | DR 12542 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125609948603360470/6125609948603360470_signature_28.png | marife Villamiel | https://www.jotform.com/uploads/macarewecare/222927978372067/6125609948603360470/6125609948603360470_signature_31.png
| 2025-01-13 | 2025-01-13 | Kabankalan City Hospital | Tabugon, Kabankalan City, Negros Occidental | 6 | Negros Occidental | TB DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 30061 | DOH | kchclinicallaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX4 | 838499 | For Installation | . | Install unit - transferred from Cadiz CHO | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125596489735939746/6125596489735939746_signature_28.png | Renly A. Layson | https://www.jotform.com/uploads/macarewecare/222927978372067/6125596489735939746/6125596489735939746_signature_31.png
| 2025-01-13 | 2025-01-10 | Jordan Primary Care Facility | Jordan, Guimaras | 6 | Guimaras | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30060 | UNOPS | jordanrhu20@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-IV | 110023528 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125522015585697548/6125522015585697548_signature_28.png | Badette Galila, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6125522015585697548/6125522015585697548_signature_31.png
| 2025-01-13 | 2025-01-10 | Libon Rural Health Unit | Zone-1, Poblacion, Libon | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30059 | DOH | marilynrazal@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1007 |  | GX4-6C | 838397 | UPS for replacement | Replaced UPS: AT1000 Cleaned cartridge bays, plunger rod Generate Data Replaced fan filter | Performed Preventive Maintenance Replaced UPS Fan filter replaced | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125464815514125997/6125464815514125997_signature_28.png | Marilyn S. Razal | https://www.jotform.com/uploads/macarewecare/222927978372067/6125464815514125997/6125464815514125997_signature_31.png
| 2025-01-13 | 2025-01-09 | Pili Rural Health Unit | Poblacion, Pili | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 30058 | DOH | lfsanao724@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1117 |  | GX4-6C | 838408 | Defective UPS | Replaced UPS (AT1000) Cleaned cartridge bays, plunger rod Replaced Fan filter | Replaced UPS Performed Preventive Maintenance Replaced Filter | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125463955512063845/6125463955512063845_signature_28.png | Liezl F. Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/6125463955512063845/6125463955512063845_signature_31.png
| 2025-01-13 | 2025-01-10 | Asian Hospital and Medical Center | Civic Drive, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | YES | CALIBRATION | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 30057 | COMMERCIAL | agimang@asianhospital.com | 25 |  | 2024-12-17 | MCS+ | 10B129 | Unit is due for Calibration. | Calibration. | Checked A/D Values. Checked Transducers. Adjust Values. Checked Functionality. Replaced Roller. Performed Simulation Test. | Unit is Operational. |  | Urethane Roller Assembly | SI 99320 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125397373615943419/6125397373615943419_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6125397373615943419/6125397373615943419_signature_31.png
| 2025-01-13 | 2025-01-10 | Asian Hospital and Medical Center | Civic Drive, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | YES | CALIBRATION | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 30056 | COMMERCIAL | agimang@asianhospital.com | 25 |  | 2024-12-17 | MCS+ | 01F116 | Unit is due for Calibration. | Calibration. | Checked A/D Values. Checked Transducers. Adjust Values. Checked Functionality. Checked UPS. Performed Simulation Test. Replaced Rollers. | Unit is Operational. UPS is Not Functional. | UPS S/N: 40405361107 | Urethane Roller Assembly. | SI 99320 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125397243616300367/6125397243616300367_signature_28.png | Angel Danica Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6125397243616300367/6125397243616300367_signature_31.png
| 2025-01-13 | 2025-01-10 | Asian Hospital and Medical Center | Civic Drive, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | YES | CALIBRATION | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 30055 | COMMERCIAL | agimang@asianhospital.com | 25 |  | 2024-12-17 | MCS+ | 10B129 | Unit is due for Calibration. | Calibration. | Checked A/D Values. Checked Transducers. Adjust Values. Checked Functionality. Replaced Roller. Performed Simulation Test. | Unit is Operational. |  | Urethane Roller Assembly | SI 99320 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125396993613743297/6125396993613743297_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6125396993613743297/6125396993613743297_signature_31.png
| 2025-01-13 | 2025-01-10 | Asian Hospital and Medical Center | Civic Drive, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | YES | CALIBRATION | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 30054 | COMMERCIAL | agimang@asianhospital.com | 25 |  | 2024-12-17 | MCS+ | 01F116 | Unit is due for Calibration. | Calibration. | Checked A/D Values. Checked Transducers. Adjust Values. Checked Functionality. Checked UPS. Performed Simulation Test. Replaced Rollers. | Unit is Operational. UPS is Not Functional. | UPS S/N: 40405361107 | Urethane Roller Assembly. | SI 99320 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125396993613059995/6125396993613059995_signature_28.png | Angel Danica Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/6125396993613059995/6125396993613059995_signature_31.png
| 2025-01-13 | 2025-01-09 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30053 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2241 ASSAY: Xpert SARS-CoV-2, COUNT: 769 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1464 |  | GX16 D 10C | 110014171 | Due for maintenance Encounter continous error 5007 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125387063613816921/6125387063613816921_signature_28.png | National Children's Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6125387063613816921/6125387063613816921_signature_31.png
| 2025-01-13 | 2025-01-09 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30052 | COMMERCIAL | mahmci_laboratory@yahoo.com | 20 |  |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125392513619476884/6125392513619476884_signature_28.png | Sarah Almeria | https://www.jotform.com/uploads/macarewecare/222927978372067/6125392513619476884/6125392513619476884_signature_31.png
| 2025-01-13 | 2025-01-09 | Santolan Super Health Center | Evangelista Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30051 | MACARE-HBDC | samandlet17@gmail.com | 20 |  |  | GX-XVI-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125392403615696736/6125392403615696736_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6125392403615696736/6125392403615696736_signature_31.png
| 2025-01-13 | 2025-01-08 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30050 | COMMERCIAL | mscaina@themedicalcity.com | 20 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125391943617239672/6125391943617239672_signature_28.png | Maria Zoraida Caina | https://www.jotform.com/uploads/macarewecare/222927978372067/6125391943617239672/6125391943617239672_signature_31.png
| 2025-01-13 | 2025-01-08 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30049 | COMMERCIAL | moleculardx@themedicalcity.com | 20 |  |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125391603611972568/6125391603611972568_signature_28.png | Stephanie Obida | https://www.jotform.com/uploads/macarewecare/222927978372067/6125391603611972568/6125391603611972568_signature_31.png
| 2025-01-13 | 2025-01-08 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30048 | COMMERCIAL | bloodbank@themedicalcity.com | 22 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125391253618697609/6125391253618697609_signature_28.png | Armin Francis Betan | https://www.jotform.com/uploads/macarewecare/222927978372067/6125391253618697609/6125391253618697609_signature_31.png
| 2025-01-13 | 2025-01-10 | Cardinal Santos Medical Center | Wilson St. San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 30047 | COMMERCIAL | jovincelim5@gmail.com | 24 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390973614875925/6125390973614875925_signature_28.png | Josh Vincent Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390973614875925/6125390973614875925_signature_31.png
| 2025-01-13 | 2025-01-09 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30044 | COMMERCIAL | mahmci_laboratory@yahoo.com | 20 |  |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913613827752/6125390913613827752_signature_28.png | Sarah Almeria | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913613827752/6125390913613827752_signature_31.png
| 2025-01-13 | 2025-01-09 | Santolan Super Health Center | Evangelista Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30046 | MACARE-HBDC | samandlet17@gmail.com | 20 |  |  | GX-XVI-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913613862094/6125390913613862094_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913613862094/6125390913613862094_signature_31.png
| 2025-01-13 | 2025-01-08 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30045 | COMMERCIAL | mscaina@themedicalcity.com | 20 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913615637278/6125390913615637278_signature_28.png | Maria Zoraida Caina | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913615637278/6125390913615637278_signature_31.png
| 2025-01-13 | 2025-01-08 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30043 | COMMERCIAL | moleculardx@themedicalcity.com | 20 |  |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913617441561/6125390913617441561_signature_28.png | Stephanie Obida | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913617441561/6125390913617441561_signature_31.png
| 2025-01-13 | 2025-01-08 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30042 | COMMERCIAL | bloodbank@themedicalcity.com | 22 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913617992784/6125390913617992784_signature_28.png | Armin Francis Betan | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390913617992784/6125390913617992784_signature_31.png
| 2025-01-13 | 2025-01-07 | Pinagbuhatan TB Microscopy Laboratory | 182 Urbano Velasco Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30041 | MACARE-HBDC | peachmango318@gmail.com | 22 |  |  | GX-IV-D-10C | 110024245 | For maintenance | Performed maintenance | Performed maintenance Cleaned -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853611174956/6125390853611174956_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853611174956/6125390853611174956_signature_31.png
| 2025-01-13 | 2025-01-07 | Marikina City Health Office | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30037 | MACARE-HBDC | amymitra22@gmail.com | 20 |  |  | GX-IV-D-10C | 110024072 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853611835961/6125390853611835961_signature_28.png | Amelia Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853611835961/6125390853611835961_signature_31.png
| 2025-01-13 | 2025-01-07 | Marikina City Health Office | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30039 | MACARE-HBDC | amymitra22@gmail.com | 18 |  |  | GX-IV-D-10C | 110024073 | Module B3 has error 4008 | Checked module B3 | Performed maintenance Checked module B3 Checked connection Checked F1 reading Checked fan functionality | Unit is operational. Module B3/210123737 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853613701157/6125390853613701157_signature_28.png | Amelia Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853613701157/6125390853613701157_signature_31.png
| 2025-01-13 | 2025-01-07 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 30038 | MACARE-HBDC | estyaragon2019@gmail.com | 22 |  |  | GX-IV-D-10C | 120000612 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853617938583/6125390853617938583_signature_28.png | Estrelita S. Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853617938583/6125390853617938583_signature_31.png
| 2025-01-13 | 2025-01-07 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 30040 | MACARE-HBDC | estyaragon2019@gmail.com | 22 |  |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853618225691/6125390853618225691_signature_28.png | Estrelita s. Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6125390853618225691/6125390853618225691_signature_31.png
| 2025-01-13 | 2025-01-10 | East Avenue Medical Center | East ave QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30036 | MACARE-HBDC | roseannparnes@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3510 |  | GX-4D 10C | 110014085 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125387853611051071/6125387853611051071_signature_28.png | Rose Ann Parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/6125387853611051071/6125387853611051071_signature_31.png
| 2025-01-13 | 2025-01-10 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 30035 | UNOPS | eamcmolecularlab@gmail.com | 20 | ASSAY: Xpert CT/NG, COUNT: 521 ASSAY: Xpert HPV, COUNT: 248 ASSAY: Xpert SARS-CoV-2, COUNT: 1105 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 197 | 2025-01-08 | GX4D 10C | 110016781 | Due for maintenance and module replacement. B1 220007358 - 210048951 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and module replacement | Operational |  | Genexpert 10 color module | 12530 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125387483611445454/6125387483611445454_signature_28.png | ANN GYLENE I. SISANTE | https://www.jotform.com/uploads/macarewecare/222927978372067/6125387483611445454/6125387483611445454_signature_31.png
| 2025-01-13 | 2025-01-09 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30034 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2241 ASSAY: Xpert SARS-CoV-2, COUNT: 769 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1464 |  | GX16 D 10C | 110014171 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6125387063613816921/6125387063613816921_signature_28.png | National Children's Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6125387063613816921/6125387063613816921_signature_31.png
| 2025-01-12 | 2025-01-09 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30033 | DOH | celinejean16@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1315 |  | GxIV-6C | 838570 | None | Perform Genexpert machine Cleaning and maintenance Checking of errors Problem/errors reporting Data collection | Genexpert preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6124960636861801156/6124960636861801156_signature_28.png | Celine villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6124960636861801156/6124960636861801156_signature_31.png
| 2025-01-11 | 2025-01-11 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Miceobiology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | aivee.alcantara@medilinx.com.ph |  | 30032 | COMMERCIAL | rosandhreijose.medilinx@outlook.com | 25 |  |  | GXII | 812916 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6123883743126374974/6123883743126374974_signature_28.png | Rosandhrei Katherine D. Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/6123883743126374974/6123883743126374974_signature_31.png
| 2025-01-10 | 2025-01-10 | Bureau of Plant Industry - Lucena | Dalahican Road, Lucena, Quezon City | 4A | Quezon | Laboratory | NO | INSTALLATION | CARL ZEISS | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 30031 | MACARE-HBDC | pqsquezon4a@gmail.com | 25 |  | 2025-01-08 | Microscope Stemi 305 Trino | 3947010009 | Demo unit for installation | Installed stemi 305 microscope for demo | Unpacked demo unit Setup demo unit Performed demonstration | Unit is operational | Installation with Sir Vince | Microscope Body Stemi 305 trino Stand K EDU Microscopy Camera Axiocam 208 Color | DR 12529 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6123004798797764792/6123004798797764792_signature_28.png | Gilbert Romano | https://www.jotform.com/uploads/macarewecare/222927978372067/6123004798797764792/6123004798797764792_signature_31.png
| 2025-01-10 | 2025-01-10 | Bureau of Plant Industry - Lucena | Dalahican Road, Lucena, Quezon City | 4A | Quezon | Labor | NO | INSTALLATION | CARL ZEISS | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 30030 | MACARE-HBDC | pqsquezon4a@gmail.com | 25 |  | 2025-01-10 | Stemi 305 EDU | 3946033510 | Unit due for installation | Installed carl zeiss microscope | Unpacked microscope Setup microscope Performed demonstration | Unit is operational | Installation with Sir Vince | CARL Zeiss microscope | DR 12471 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6123001658798662619/6123001658798662619_signature_28.png | Gilbert Romano | https://www.jotform.com/uploads/macarewecare/222927978372067/6123001658798662619/6123001658798662619_signature_31.png
| 2025-01-10 | 2025-01-10 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 30029 | DOH | suzzettekho34@gmail.com | 28 |  |  | GX-4 6 Colors | 838554 | No reported problem/s. | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122956863507519866/6122956863507519866_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/6122956863507519866/6122956863507519866_signature_31.png
| 2025-01-10 | 2025-01-10 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | synch28@yahoo.com |  | 30028 | MACARE-HBDC | msff-tondo-labtech@paris.msf.org | 25 |  | 2025-01-08 | GX16-10C | 110016343 | Unit for preventive maintenance. For module replacement of module C4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module C4 (685089) in replacement to defective C4 (210117673). | DR#12477 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122913702322893916/6122913702322893916_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6122913702322893916/6122913702322893916_signature_31.png
| 2025-01-10 | 2025-01-10 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30027 | COMMERCIAL | gramanmolecularlab@gmail.com | 21 |  |  | GXIV | 110008530 | Unit for additional troubleshooting by Cepheid Unit for maintenance | NA | #NAME? | Unit is operational Report to be forwarded to technical specialist |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122901179315224066/6122901179315224066_signature_28.png | Margareth Camua | https://www.jotform.com/uploads/macarewecare/222927978372067/6122901179315224066/6122901179315224066_signature_31.png
| 2025-01-10 | 2025-01-10 | WESTERN VISAYAS MEDICAL CENTER | Mandurriao, Iloilo | 6 | Iloilo | rHIVda | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30026 | UNOPS | wvmcrhivdacrcl@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 300 |  | GX-IV | 110023529 | FOR MAINTENANCE AND CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122890016021708951/6122890016021708951_signature_28.png | Jessa Krista Riina Jalandoni | https://www.jotform.com/uploads/macarewecare/222927978372067/6122890016021708951/6122890016021708951_signature_31.png
| 2025-01-10 | 2025-01-10 | Muntinlupa Reproductive Health and Wellness Center | Centennial Avenue, Tunasan, Muntinlupa | NCR | NCR 4th - Muntinlupa | Social Hygiene Clinic | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30025 | MACARE-HBDC | petelerica@gmail.com | 25 |  |  | GXIV-D-10C | 110025706 | Unit is due for Installation. | Installation. | Checked Outlet Voltage. Room with A/C Unit. No Direct Sunlight Exposure. Performed Self-test. Checked System's Functionality. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122884766554088092/6122884766554088092_signature_28.png | Erica Monique Petel | https://www.jotform.com/uploads/macarewecare/222927978372067/6122884766554088092/6122884766554088092_signature_31.png
| 2025-01-10 | 2025-01-03 | Dasmariñas City Health Office 2 | DBB-E, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30023 | DOH | hazelmilan0777@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1118 |  | GXIV-D-6C | 838535 | UPS Problem. | Troubleshoot. | Checked Voltage Output. Checked Backup Voltage Output. Checked UPS Functionality. | UPS is Functional. For Further Observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122823106557272067/6122823106557272067_signature_28.png | Hazel Tabucol | https://www.jotform.com/uploads/macarewecare/222927978372067/6122823106557272067/6122823106557272067_signature_31.png
| 2025-01-10 | 2025-01-09 | Rosario Rural Health Unit | Municipal Compound, Rosario, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 30024 | DOH | arlenepenaarmijo@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1195 |  | GXIV-D-6C | 838537 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Replaced Filter. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122823106559544167/6122823106559544167_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/6122823106559544167/6122823106559544167_signature_31.png
| 2025-01-09 | 2025-01-09 | SIBUNAG RURAL HEALTH UNIT | SIBUNAG, GUIMARAS | 6 | Guimaras | LABORATORY | NO | INSTALLATION Site-transfer from Buenavista RHU | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30022 | UNOPS | rhodatrance1994@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-IV | 110023538 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122197424623566530/6122197424623566530_signature_28.png | Ma. Rhoda Lyn Villarta | https://www.jotform.com/uploads/macarewecare/222927978372067/6122197424623566530/6122197424623566530_signature_31.png
| 2025-01-09 | 2025-01-09 | Philippine Red Cross Zamboanga | Pettit baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 30021 | COMMERCIAL | zamboangacity@redcross.org.ph | 22 |  | 2025-01-09 | MCS + | 01M256 | Display monitor is flickering, machine don't turn on. | Troubleshoot and determined what is the cause of the problem, Run cycle test. | Troubleshoot and determined what is the cause of the problem, Run cycle test. | Unit condition is operational, no flickering problem and machine turns on properly | Request for display cable due to recent problem. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122135569414450798/6122135569414450798_signature_28.png | Philippine Red Cross Zamboanga City chapter | https://www.jotform.com/uploads/macarewecare/222927978372067/6122135569414450798/6122135569414450798_signature_31.png
| 2025-01-09 | 2025-01-09 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | INSTALLATION REPAIR | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30020 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  | 2025-01-08 | MCS+ | 16K521SPG | For installation of MCS+. | Installation. | Performed installation of MCS+. Performed simulation running. Verified calibration parameters. | Unit is operational. |  | Apheresis System | DR#12474 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122125222125059213/6122125222125059213_signature_28.png | Adelle Florendo | https://www.jotform.com/uploads/macarewecare/222927978372067/6122125222125059213/6122125222125059213_signature_31.png
| 2025-01-09 | 2025-01-09 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Bloodbank | NO | REPAIR | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 30019 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. For installation of rotor pumps. | Preventive maintenance. | Performed preventive maintenance. Performed simulation running. Verfied calibration parameters. Installed rotor pumps. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122125222126991629/6122125222126991629_signature_28.png | Adelle Florendo | https://www.jotform.com/uploads/macarewecare/222927978372067/6122125222126991629/6122125222126991629_signature_31.png
| 2025-01-09 | 2025-01-09 | National Center For Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30018 | COMMERCIAL | iancristrono@gmail.com | 29 |  |  | GXIV 10C | 120003480 | Unit for maintenance Module B3 optics error | Cleaned module B3's optics | #NAME? | Unit is operational Module B3 is under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122121089312735234/6122121089312735234_signature_28.png | Ian Cris Tono | https://www.jotform.com/uploads/macarewecare/222927978372067/6122121089312735234/6122121089312735234_signature_31.png
| 2025-01-09 | 2025-01-09 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30017 | DOH | rhwclab@gmail.com | 25.6 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1556 ASSAY: Xpert MTB/RIF Ultra, COUNT: 384 ASSAY: Xpert HBV VL, COUNT: 9 |  | GX-IV 6C 4D | 838262 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational | Module B1 foe under observation, error sent to cepheid for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122108664524623898/6122108664524623898_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6122108664524623898/6122108664524623898_signature_31.png
| 2025-01-09 | 2025-01-09 | Davao Reproductive Health and Wellness Center | Jacinto Street, Poblacion District Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30016 | MACARE-HBDC | rhwclab@gmail.com | 25.6 | ASSAY: Xpert HBV VL, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 384 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1556 |  | GX-IV 6C 4D | 839236 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Only Three Modules are operational | Module B3 reported to Cepheid for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122107624522200658/6122107624522200658_signature_28.png | Elaine Malano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6122107624522200658/6122107624522200658_signature_31.png
| 2025-01-09 | 2025-01-09 | Poro Rural Health Unit | Poro, Camotes, Cebu | 7 | Cebu | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30015 | DOH | avilacharmaine6@gmail.com | 38 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 531 |  | 6C GXIV-4D | 838279 | Due for PM | #NAME? | #NAME? | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6122100506619902295/6122100506619902295_signature_28.png | Charmaine Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/6122100506619902295/6122100506619902295_signature_31.png
| 2025-01-09 | 2025-01-09 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30014 | DOH | burauenlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 469 | 2025-01-09 | GXIV | 838551 | n/a | n/a | Performed preventive maintenance | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6121990290318595238/6121990290318595238_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6121990290318595238/6121990290318595238_signature_31.png
| 2025-01-09 | 2025-01-08 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Donor room | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 30013 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15H029G | CRC bios failed, Drain output voltage of CPU battery | Performed main reset but error still occur Checked output voltage of battery 0.101 volts Reset CPU board but error still occur. | Performed troubleshooting | For replacement of CPU board | Work with Ronnel Belizar |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6121907803617451686/6121907803617451686_signature_28.png | Ivan Ritz Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6121907803617451686/6121907803617451686_signature_31.png
| 2025-01-08 | 2025-01-08 | Polomolok Rural Health Unit | Poblacion Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | CALIBRATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 30012 | DOH | jmmirasolrmt@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1694 |  | GxIV-6C | 838265 | Module 2 and 4 waiting for replacement module | Perform Genexpert Cleaning and maintenance Checking of errors Problem/errors reporting Data collection | Genexpert machine Preventive maintenance | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6121428826864029297/6121428826864029297_signature_28.png | JUNE MICHELE J. MIRASOL | https://www.jotform.com/uploads/macarewecare/222927978372067/6121428826864029297/6121428826864029297_signature_31.png
| 2025-01-08 | 2025-01-08 | Visayas Med | Cebu City | 7 | Cebu | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 30011 | COMMERCIAL | SGDPH1123@pathologyasia.com | 28 |  |  | Mcs+ | 05M204 | Watchdog error Defective valve pump | Preventive Maintenance & Troubleshooting | Preventive Maintenance & Troubleshooting | Defective valve pump need for quotation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6121291339615943249/6121291339615943249_signature_28.png | JUSTIN MASECAMPO | https://www.jotform.com/uploads/macarewecare/222927978372067/6121291339615943249/6121291339615943249_signature_31.png
| 2025-01-08 | 2025-01-08 | Klinika Eastwood | City Walk 2, Eastwood, Bagumabayn, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | giandesilvamagnaye@gmail.com |  |  | 30010 | MACARE-HBDC | klinikaeastwoodlaboratory@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 |  | GXIV 10C D | 110020698 | Unit duebfor maintenance | Performed preventive maintenance All modules passed reliability test | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Archived results Performed reliability test | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6121257770712712840/6121257770712712840_signature_28.png | Mary Claire P Baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/6121257770712712840/6121257770712712840_signature_31.png
| 2025-01-08 | 2025-01-08 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 30009 | DOH | buhanginlaboratory@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1524 |  | GX-IV 6C 4D | 838232 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6121257733408609030/6121257733408609030_signature_28.png | Catherine B. Deocampo,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6121257733408609030/6121257733408609030_signature_31.png
| 2025-01-08 | 2025-01-08 | Maayon Rural Health Unit | Maayon , Capiz | 6 | Capiz | Laboratory | NO | INSTALLATION CALIBRATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 30008 | UNOPS | pamskii_38@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-IV | 110023541 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6121185107916853000/6121185107916853000_signature_28.png | Pamela Cerujano | https://www.jotform.com/uploads/macarewecare/222927978372067/6121185107916853000/6121185107916853000_signature_31.png
| 2025-01-08 | 2025-01-07 | Makati Medical Center | Amorsolo St.,Makati | NCR | NCR 4th - Makati | CRM | NO | MAINTENANCE | OTHERS | Noel Floresca |  | florescanoel77@gmail.com | biomedsupervisor@makatimed.net.ph | Miltenyi | 30007 | COMMERCIAL | stemcell@makatimed.net.ph | 28 |  |  | Clinimacs | 797 | Due for maintenance | Cleaned unit Checked rollers Checked and test tube sensor Checked and test keypad and lcd display. Checked program parameters Performed functionality | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6120993958816196909/6120993958816196909_signature_28.png | Princess Jerica Alondra T. Mararac | https://www.jotform.com/uploads/macarewecare/222927978372067/6120993958816196909/6120993958816196909_signature_31.png
| 2025-01-08 | 2025-01-07 | Makati Medical Center | Amorsolo St,.Makati | NCR | NCR 4th - Makati | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | biomedsupervisor@makatimed.net.ph |  | 30006 | COMMERCIAL | bloodbank@makatimed.net.ph | 27 |  |  | MCS + | 97H330 | Due for maintenance Noisy vibration from centrifuge cover Power cord wornout insulation rubber | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Tighten screw of cebtrifuge cover Performed functionality test Performed simulation running | Performed maintenance | Operational | Need to replace powercord |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6120991738819538421/6120991738819538421_signature_28.png | Mark Dalisay | https://www.jotform.com/uploads/macarewecare/222927978372067/6120991738819538421/6120991738819538421_signature_31.png
| 2025-01-07 | 2025-01-07 | Batasan Social Hygiene Clinic | Batasan Road, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 30005 | MACARE-HBDC | prestozatricia@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 23 ASSAY: Xpert HBV VL, COUNT: 30 ASSAY: Xpert HIV-1 Viral Load, COUNT: 712 |  | GXIV 10C D | 110020697 | Unit due for maintenance | Peeformed preventive maintenance Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Archived results | Unit is operational All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6120432366174050938/6120432366174050938_signature_28.png | Patricia Anne Prestoza | https://www.jotform.com/uploads/macarewecare/222927978372067/6120432366174050938/6120432366174050938_signature_31.png
| 2025-01-07 | 2025-01-07 | Diit District Health Center | DIIT , Tacloban City | 8 | Leyte | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 30004 | DOH | cholaboratory23@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1149 | 2025-01-07 | GXIV | 838421 | #NAME? | n/a | #NAME? | Unit is operational | Unit transferred from Tacloban CHO |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6120415860962017340/6120415860962017340_signature_28.png | Dale D. Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6120415860962017340/6120415860962017340_signature_31.png
| 2025-01-07 | 2025-01-07 | Philippine Red Cross - Davao City Chapter | Roxas Boulevard, Davao City | 11 | Davao del Sur | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 30003 | COMMERCIAL | davaocitybloodservices@gmail.com | 24.5 |  | 2024-10-25 | MCS+ | 02J021 | Procedure Termination due to Return Flow Problem and High Pressure on DPM UPS Troubleshooting | Clean Surface Area Change Rollers Check the parameters for DPM, SPM and Line Sensor Run Cycle Test Check for Line Voltages of the UPS and test its Functionality | Perform Troubleshooting on MCS+. Machine and at UPS | Machine completed the cycle test. Unit is Operational, UPS is operational | Bought an adapter for MCS+ Machine | MCS+ Urethane Roller Assay | DR 12282 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6120328557361590367/6120328557361590367_signature_28.png | Joy Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/6120328557361590367/6120328557361590367_signature_31.png
| 2025-01-07 | 2025-01-07 | Philippine Red Cross Tower | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 30002 | MACARE-HBDC | tchristianpaolo84@gmail.com | 25 |  |  | MCS+ | 08B049 | Centifuge noise during procedure | NA | #NAME? | Unit is operational Unit is under observation For PM next visit |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6120296158103758556/6120296158103758556_signature_28.png | Christian Paolo Taguiam | https://www.jotform.com/uploads/macarewecare/222927978372067/6120296158103758556/6120296158103758556_signature_31.png
| 2025-01-07 | 2025-01-06 | Rosario Super Health Center | Dr. Antonio Sixto Avenue, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | Relocation/Packing | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 30001 | MACARE-HBDC | consalaneljaine@gmail.com | 24 |  |  | GX-IV-D-10C | 110024246 | GeneXpert System for relocation/packing | Transferred to other room | Packed the GeneXpert system -GeneXpert machine -Workstation -UPS -Barcode scanner and stand Transferred to other room | Unit is operational. | Not yet installed. No acu in the new room. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6120255414198950134/6120255414198950134_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6120255414198950134/6120255414198950134_signature_31.png
| 2025-01-06 | 2025-01-06 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 30000 | DOH | ccmclaboratory01@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2174 |  | 6C GXIV-4D | 838561 | Due for PM | #NAME? | #NAME? | Unit is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6119511348124137272/6119511348124137272_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/6119511348124137272/6119511348124137272_signature_31.png
| 2025-01-03 | 2025-01-03 | Central Luzon Doctor's Hospital | Hospital Drive Brgy San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 29999 | COMMERCIAL | pjlsantos@cldh.com.ph | 24 | ASSAY: Xpert MTB/RIF, COUNT: 193 |  | GXII-D | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6116973338215837562/6116973338215837562_signature_28.png | Maria Rowena Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6116973338215837562/6116973338215837562_signature_31.png
| 2025-01-03 | 2025-01-03 | Central Luzon Doctor's Hospital | Hospital Drive Brgy San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 29998 | COMMERCIAL | pjlsantos@cldh.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 |  | GXIV-D | 839232 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6116972078216971834/6116972078216971834_signature_28.png | Maria Rowena Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6116972078216971834/6116972078216971834_signature_31.png
| 2025-01-02 | 2025-01-02 | Manila Doctors Hospital | United Nations Ave., Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29997 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 21 | ASSAY: Xpert CT/NG, COUNT: 75 ASSAY: Xpert HBV VL, COUNT: 97 ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1213 ASSAY: Xpert HCV VL, COUNT: 18 ASSAY: Xpert HIV-1 Viral Load, COUNT: 106 |  | GXIV 10C D | 110001648 | Unit to be returned to Molecular | Disassembled machine Setup the machine in Molecular Checkes and cleaned cartridge bays, plunger rod, icore slot and filter | Transferred machine from Bacteriology to Molecular Performed Preventive maintenance | Unit is operational All modulea working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6116116639916651683/6116116639916651683_signature_28.png | Alexa Beatrice Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/6116116639916651683/6116116639916651683_signature_31.png
| 2025-01-02 | 2024-12-05 | Rosario Super Health Center | Dr. Sixto Antonio Avenue. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29996 | MACARE-HBDC | consalaneljaine@gmail.com | 22 |  |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6116094933619635569/6116094933619635569_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6116094933619635569/6116094933619635569_signature_31.png
| 2025-01-02 | 2024-12-05 | Santolan Super Health Center | Evangelista Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29995 | MACARE-HBDC | samandlet17@gmail.com | 22 |  |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6116094793613837720/6116094793613837720_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6116094793613837720/6116094793613837720_signature_31.png
| 2024-12-27 | 2024-12-27 | Estevez Memorial Hospital, Inc. | Don Juan Estevez Sr. st., Legazpi City | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | emhi@gmail.com |  | 29994 | COMMERCIAL | reylovendino83@gmail.com | 24 |  |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Disabled A2/220009328 - temperature Error Generate Data | Performed Preventive Maintenance | Unit is operational | A2/220009328 - temperature Error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6111215788905133459/6111215788905133459_signature_28.png | Lovendino, rey a | https://www.jotform.com/uploads/macarewecare/222927978372067/6111215788905133459/6111215788905133459_signature_31.png
| 2024-12-27 | 2024-12-09 | Tigaon Municipal Health Office | Tigaon | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29993 | DOH | jingimperial24@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 839 |  | GX4-6C | 838558 | Relocation of Gx from temporary to permanent facility | Relocation of GX4-6C from vinagre brgy health center(temporary facility) to Tigaon Municipal Health Office Cleaned cartridge bays, plunger rod Generate data | Relocation of Gx from temporary to permanent facility Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6111211538903793665/6111211538903793665_signature_28.png | Maria K esusa P. Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6111211538903793665/6111211538903793665_signature_31.png
| 2024-12-27 | 2024-12-27 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 29992 | MACARE-HBDC | laboratory@slec.ph | 22 |  | 2024-12-26 | GeneXpert Infinity | 809643 | Modules that failed XpertCheck are for replacement | Replaced modules A4 (210050828 > 616697), C4 (720535 > 804782), D2 (729869 > 769035), D4 (618607 > 734837), E3 (896445 > 731600), H1 (608345 > 710090), H2 (696244 > 651942), H3 (654533 > 635029) Performed preventive maintenance cleaning Performed functionality test | Replaced modules that failed XpertCheck | Unit is operational |  | GeneXpert Modules (8pcs), modified ejector for GeneXpert Infinity module (8pcs), Assy Door GeneXpert module (8pcs) | DR12462 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6110985508421628322/6110985508421628322_signature_28.png | Ma. Sarah Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6110985508421628322/6110985508421628322_signature_31.png
| 2024-12-23 | 2024-12-23 | MDH | United Nations Avenue, Manila | NCR | NCR 1st - Manila | Laboratory | NO | Relocation | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 29991 | COMMERCIAL | molecular.lab@maniladoctors.com.ph | 22 |  |  | GXIV | 110001648 | Unit for temporary relocation | Relocated unit Performed functionality test Generated reports | Relocated GeneXpert unit from 6th floor molecular lab to 2nd floor laboratory | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6107492394619857144/6107492394619857144_signature_28.png | Alyssa Mae Jolo | https://www.jotform.com/uploads/macarewecare/222927978372067/6107492394619857144/6107492394619857144_signature_31.png
| 2024-12-23 | 2024-12-23 | St Lukes Medical Center QC | E.Rodriguez | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 29990 | COMMERCIAL | megapo@stlukes.com.ph | 24 | ASSAY: Xpert CT/NG, COUNT: 214 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3007 |  | GX4 D | 839234 | B1 Module for repair door | Repaire door module Returned and installed module B1 s/n 210072911 | Performed repair | B1,B2 operational B3,B4 disabled failed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6107446675429984765/6107446675429984765_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6107446675429984765/6107446675429984765_signature_31.png
| 2024-12-23 | 2024-12-19 | Marikina City Health Office | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 29989 | MACARE-HBDC | estyaragon2019@gmail.com | 22 |  |  | GX-IV-D-10C | 110024073 | Error 4008/2006 on module B3 | Performed troubleshooting | Performed troubleshooting Checked module B3 Checked internal connections Removed stuck cartridge Generated reports | Unit is operational. Module B3/210123737 not operational, for reporting. | Disabled module B3 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6107425379655678547/6107425379655678547_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6107425379655678547/6107425379655678547_signature_31.png
| 2024-12-23 | 2024-12-23 | Mandaluyong Social Hygiene Clinic | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29988 | MACARE-HBDC | mandaluyongshc.laboratory@gmail.com | 27 |  |  | GXIV 10C | 110025111 | Unit for installation | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6107309885214565873/6107309885214565873_signature_28.png | Revena Sibayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6107309885214565873/6107309885214565873_signature_31.png
| 2024-12-21 | 2024-12-21 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29987 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 03J030 | For checking and troubleshooting of MCS+ due to reported centrifuge. | Troubleshooting. | Performed checking and troubleshooting. Verified centrifuge calibration parameters. Performed simulation running with latham bowl from the reported kit. Performed simulation running with latham bowl from another kit. | Unit is operational. Endorsed to technical for kit replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6105691792122277326/6105691792122277326_signature_28.png | Lalaine ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6105691792122277326/6105691792122277326_signature_31.png
| 2024-12-17 | 2024-12-17 | Marikina City Health Office | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 29986 | MACARE-HBDC | estyaragon2019@gmail.com | 22 |  |  | GX-IV-D-10C | 110024073 | Relocation of GeneXpert system | Installed the GeneXpert System | Transferred GeneXpert system to other room Installed: -GeneXpert unit -Workstation -Printer -Barcode scanner -UPS Performed maintenance | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6102092290022294016/6102092290022294016_signature_28.png | Estrelita S.Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6102092290022294016/6102092290022294016_signature_31.png
| 2024-12-17 | 2024-12-17 | Marikina City Health Office | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 29985 | MACARE-HBDC | estyaragon2019@gmail.com | 22 |  |  | GX-IV-D-10C | 110024072 | For relocation | Installed the GeneXpert System | Transferred GeneXpert system to other room Installed: -GeneXpert unit -Workstation -Printer -Barcode scanner -UPS Performed maintenance | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6102091450021700914/6102091450021700914_signature_28.png | Estrelita S. Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6102091450021700914/6102091450021700914_signature_31.png
| 2024-12-16 | 2024-12-16 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | Pull out | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 29984 | COMMERCIAL | megapo@stlukes.com.ph | 24 | ASSAY: Xpert CT/NG, COUNT: 251 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3029 |  | GX4 D | 839234 | B1 module broken uniframe | Pull out B1 module s/n 210072911 for repair broken uniframe | Pull out | B2 operational B3,B4 failed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6101415253112061941/6101415253112061941_signature_28.png | Ma. Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6101415253112061941/6101415253112061941_signature_31.png
| 2024-12-16 | 2024-12-11 | Lugait Rural Health Unit | Lugait, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com | rodelynlimocon@gmail.com |  | 29983 | DOH | joerizmikeflorida@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 687 | 2024-11-28 | 6C GXIV-4D | 838562 | Module B2 defective | -Replace defective module B2/690358 with new module SN: 607941 | #NAME? | All modules are functional. |  | 900-0491RS Certified pre-owned GeneXpert module SN. 607941 | DR no. 12394 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6101228273612309782/6101228273612309782_signature_28.png | Joeriz Michael Florida | https://www.jotform.com/uploads/macarewecare/222927978372067/6101228273612309782/6101228273612309782_signature_31.png
| 2024-12-16 | 2024-12-05 | Rosario Super Health Center | Dr. Sixto Antonio Avenue. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29982 | MACARE-HBDC | consalaneljaine@gmail.com | 22 |  |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6101174253619840375/6101174253619840375_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6101174253619840375/6101174253619840375_signature_31.png
| 2024-12-16 | 2024-12-05 | Santolan Super Health Center | Evangelista Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29981 | MACARE-HBDC | samandlet17@gmail.com | 22 |  |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6101174153614617188/6101174153614617188_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6101174153614617188/6101174153614617188_signature_31.png
| 2024-12-16 | 2024-12-11 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29980 | MACARE-HBDC | mahmci_labcmt@yahoo.com | 22 |  |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports Generated .gxx file | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6101174013612301672/6101174013612301672_signature_28.png | Graziela Marie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/6101174013612301672/6101174013612301672_signature_31.png
| 2024-12-16 | 2024-12-11 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29979 | DOH | lornsabaiga@gmail.com | 27 |  | 2024-12-11 | GXIV | 838409 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6101059165918803112/6101059165918803112_signature_28.png | lorna abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/6101059165918803112/6101059165918803112_signature_31.png
| 2024-12-12 | 2024-12-11 | St Lukes Medical Center QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 29978 | COMMERCIAL | megapo@stlukes.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 258 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3059 |  | GX4D | 839234 | As per end user they unable to close the door module properly Broken Uniframe | Checked uniframe Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed trpubleshooting | B1 broken uniframe B3,B4 failed xpertcheck B2 operational | B1 module for replacement uniframe module unrepairable |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097731763618475959/6097731763618475959_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6097731763618475959/6097731763618475959_signature_31.png
| 2024-12-12 | 2024-12-11 | National Kidney and Transplant Institute | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 29977 | MACARE-HBDC | molecularlabnkti@gmail.com | 27 |  | 2024-12-10 | GX16D | 110005587 | For module replacement A4 210071092 - 722336 C4 210071450 - 210059476 A1,A2,C2,D2 reported already for replacement | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | 12 modules operational 4 modules disabled |  | Genexpert 6 color module | 12424, 12419 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097728193614752449/6097728193614752449_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6097728193614752449/6097728193614752449_signature_31.png
| 2024-12-11 | 2024-12-10 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29976 | DOH | rhupozorrubio@yahoo.com | 27 |  |  | GXIV-4-D-6C | 838467 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097147369852552916/6097147369852552916_signature_28.png | Caroline Nacis | https://www.jotform.com/uploads/macarewecare/222927978372067/6097147369852552916/6097147369852552916_signature_31.png
| 2024-12-11 | 2024-12-10 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29975 | DOH | arleneabad32@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 657 |  | GXIV-4-D-6C | 838288 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097146349859222884/6097146349859222884_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/6097146349859222884/6097146349859222884_signature_31.png
| 2024-12-11 | 2024-12-05 | Asian Hospital and Medical Center | Corporate Ave., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | agimang@asianhospital.com |  | 29974 | COMMERCIAL | mvflores@asianhospital.com | 25 |  |  | MCS+ | 10B129 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Surface Area. Cleaned Pumps and Sensors. Checked A/D Values. Performed Simulation Test | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065643614505739/6097065643614505739_signature_28.png | Maria Vanessa Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065643614505739/6097065643614505739_signature_31.png
| 2024-12-11 | 2024-11-19 | San Lazaro Hospital | Quiricada St., Manila | NCR | NCR 1st - Manila | SACCL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | piaciudadano1331@gmail.com |  | 29971 | MACARE-HBDC | ciudadanosophiapatrice@gmail.com | 26.5 | ASSAY: Xpert HBV VL, COUNT: 412 ASSAY: Xpert HIV-1 Viral Load, COUNT: 219 ASSAY: Xpert HCV VL, COUNT: 28 |  | GXIV-L-6C | 821247 | Unit has 5007 Error. | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Reliability Test. Performed Self-test. Generate Reports. | 3 Modules are Operational. 1 is Disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065593611638978/6097065593611638978_signature_28.png | Sophia Patrice Ciudadano | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065593611638978/6097065593611638978_signature_31.png
| 2024-12-11 | 2024-11-26 | Jose Reyes Memorial Medical Center | Rizal Ave.,Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29972 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 ASSAY: Xpert SARS-CoV-2, COUNT: 134 |  | GXIV-D-10C | 110014099 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | 1 Module Disabled. 3 Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065593613785335/6097065593613785335_signature_28.png | Ervin Emmett P. Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065593613785335/6097065593613785335_signature_31.png
| 2024-12-11 | 2024-12-04 | UST Hospital | Dapitan St., Manila | NCR | NCR 1st - Manila | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | leriza_rabanal@yahoo.com |  | 29973 | MACARE-HBDC | usth_bloodbank@yahoo.com | 25 |  |  | MCS+ | 02F069 | UPS Problem. | Troubleshoot. | Checked UPS Output Voltage. Checked UPS Functionality (with and w/o Load). | UPS is Defective ( with No Backup Power) | UPS S/N: AT1000180910ST0259 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065593617704412/6097065593617704412_signature_28.png | Leriza Rabanal | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065593617704412/6097065593617704412_signature_31.png
| 2024-12-11 | 2024-11-26 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29969 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 21 | ASSAY: Xpert HPV, COUNT: 253 ASSAY: Xpert SARS-CoV-2, COUNT: 868 |  | GX16-D-10C | 110013856 | Unit is dur for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065583616389527/6097065583616389527_signature_28.png | Ervin Emmett P. Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065583616389527/6097065583616389527_signature_31.png
| 2024-12-11 | 2024-12-05 | Asian Hospital and Medical Center | Corporate Ave., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | agimang@asianhospital.com |  | 29970 | COMMERCIAL | mvflores@asianhospital.com | 25 |  |  | MCS+ | 01F116 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Surface Area. Cleaned Pumps and Sensors. Checked A/D Values. Performed Simulation Test. | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065583616594263/6097065583616594263_signature_28.png | Maria Vanessa Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6097065583616594263/6097065583616594263_signature_31.png
| 2024-12-11 | 2024-12-11 | LAPULAPU CITY HEALTH OFFICE | LAPULAPU CEBU | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.acebedo@gmail.com |  |  | 29968 | DOH | tanya_pax@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2572 |  | GX-IV | 837661 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097013001368193923/6097013001368193923_signature_28.png | Maria Lucille Arong, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6097013001368193923/6097013001368193923_signature_31.png
| 2024-12-11 | 2024-12-11 | University of The East Ramon Magsaysay Medical Center | Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29967 | MACARE-HBDC | uermbloodbank@gmail.com | 28 |  | 2024-12-03 | MCS+ | 13C270 | UPS for replacement | Replaced UPS | #NAME? | Unit is operational |  | UPS 1KVA SN: AT1000190123ST0327 | 12413 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6097008488555009803/6097008488555009803_signature_28.png | Leonila | https://www.jotform.com/uploads/macarewecare/222927978372067/6097008488555009803/6097008488555009803_signature_31.png
| 2024-12-10 | 2024-12-10 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29966 | DOH | gbmmbg@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 749 |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6096252523705305965/6096252523705305965_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/6096252523705305965/6096252523705305965_signature_31.png
| 2024-12-10 | 2024-12-10 | Lung Center of the Philippines | Quezon Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | HAEMONETICS | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29965 | MACARE-HBDC | lcpbloodbank@gmail.coma | 25 |  |  | MCS+ | 08H188 | DPM Pressure error | Checked program parameters Adjust full scale of bowl optics Adjust line sensor value to 3071 Adjust DPM and SPm pressue to 300mmHg Adjust weigher value to 1000 Performed functionality test | Performed troubleshooting | Unit is oeprational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6096252389117760931/6096252389117760931_signature_28.png | Rhonaq Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/6096252389117760931/6096252389117760931_signature_31.png
| 2024-12-10 | 2024-12-10 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29964 | DOH | thisisformyfamily@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1469 |  | GXIV-4-D-6C | 838286 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | - Unit is Partially Operational - Module 4(808255) Has already been Reported |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6096160215612364988/6096160215612364988_signature_28.png | Jezra Jong | https://www.jotform.com/uploads/macarewecare/222927978372067/6096160215612364988/6096160215612364988_signature_31.png
| 2024-12-10 | 2024-12-10 | Pasig Treatment Hub | Dr. Garcia St. | NCR | NCR 2nd - Pasig | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29963 | MACARE-HBDC | marcabapo@gmaul.com | 22 |  |  | GX-IV-D-10C | 110025128 | SQL Problem | Performed troubleshooting | Performed troubleshooting of SQL Tried to reinstall GeneXpert software Manual installation of SQL Change settings to original | Unit is not operational. For additional troubleshooting of Workstation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6096139586269103256/6096139586269103256_signature_28.png | Marc francis abapo | https://www.jotform.com/uploads/macarewecare/222927978372067/6096139586269103256/6096139586269103256_signature_31.png
| 2024-12-10 | 2024-12-10 | San Jacinto Rural Health Unit | San Jacinto, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29962 | DOH | allen.indong@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 328 |  | GXIV-4-D-6C | 838518 | For Troubleshooting and Preventive Maintenance | #NAME? | Performed Troubleshooting and Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6096038905423239223/6096038905423239223_signature_28.png | Alexander Indong | https://www.jotform.com/uploads/macarewecare/222927978372067/6096038905423239223/6096038905423239223_signature_31.png
| 2024-12-09 | 2024-12-02 | Bagumbayan Rural Health Unit | Poblacion, Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29960 | DOH | larrozarey27@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 919 | 2024-12-02 | GxIV-6C | 838261 | Module 1 for Service Required | Perform Genexpert machine repair Perform Genexpert CPU repair | Replace defective module SN:807762 with new module SN: 882094 Replace CPU CMOS battery Genexpert Cleaning and checking of errors Data collection | 4 operational modules |  | Genexpert 6-color module SN:882094 | DR#:12362 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095599459494323063/6095599459494323063_signature_28.png | Rey D. LARROZA | https://www.jotform.com/uploads/macarewecare/222927978372067/6095599459494323063/6095599459494323063_signature_31.png
| 2024-12-09 | 2024-12-04 | Sultan Mastura Rural Health Unit | Sultan Mastura, Maguindanao Del norte | BARMM | Maguindanao | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29961 | DOH | salehshahinaz@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 341 | 2024-12-04 | GxIV-6C | 837613 | Module 1 and 2 for Service Required | Genexpert machine module repair Preventive maintenance | Replace defective module SN:806639 with new module SN:804307 Replace defective module SN:797793 with new module SN:837613 Checking of Errors/problem Data collection | 4 operational modules |  | Genexpert 6-color modules Sn:804307 Sn:769600 | DR#:12376 and 12377 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095599459497004899/6095599459497004899_signature_28.png | Shahinaz Saleh | https://www.jotform.com/uploads/macarewecare/222927978372067/6095599459497004899/6095599459497004899_signature_31.png
| 2024-12-09 | 2024-12-06 | Norala District Hospital | Poblacion Norala, South Cotabato | 12 | South Cotabato | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29959 | DOH | ndhlab2023@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 955 | 2024-12-06 | GxIV-6C | 838543 | Module 2 for Service Required | Perform Genexpert repair Perform Genexpert preventive maintenance | Replace defective module B2 SN:210068700 with new module SN:630592 Genexpert machine Cleaning and maintenance Checking of errors Data collection | 4 operational modules |  | Genexpert 6-color module SN:630592 | DR#:12395 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095599459498752392/6095599459498752392_signature_28.png | Geraldine N. Magan | https://www.jotform.com/uploads/macarewecare/222927978372067/6095599459498752392/6095599459498752392_signature_31.png
| 2024-12-09 | 2024-12-09 | Tomas Cabili Health Center | Brgy. Tominobo, Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29958 | DOH | mariacielovenezuela@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1130 | 2024-11-28 | 6C GXIV-4D | 838294 | Module B2 defective | -Replace defective module B2/808192 with module SN: 838294 | #NAME? | All modules are functional. |  | 900-0491RS Certified pre-owned GeneXpert module SN. 652201 | DR no. 12391 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095411122322837664/6095411122322837664_signature_28.png | Maria Cielo Venezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/6095411122322837664/6095411122322837664_signature_31.png
| 2024-12-09 | 2024-12-06 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29957 | MACARE-HBDC | molecularlabnkti@gmail.com | 29 |  |  | GX16D | 804476 | Additional troubleshooting for C2 module loss communication | Swap module position c1 to c2 no loss communication encountered Replaced lan cable Checked gateway board | Performed troubleshooting | A3,A4,B2 in use Disabled A1,A2,B1,B3,,C1,C2,C3,C4 B4 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095386803617411603/6095386803617411603_signature_28.png | Ma. Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6095386803617411603/6095386803617411603_signature_31.png
| 2024-12-09 | 2024-12-09 | Hinaplanon Health Center | Brgy. Hinaplanon, Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29956 | DOH | villacortaamy50@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 899 | 2024-11-28 | 6C GXIV-4D | 838227 | Module B3 defective | -Replace defective module B3/641982 with module SN: 630963 | #NAME? | All modules are functional. |  | 900-0491RS Certified pre-owned GeneXpert module SN. 630963 | DR no. 12389 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095360659018857538/6095360659018857538_signature_28.png | Amy Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6095360659018857538/6095360659018857538_signature_31.png
| 2024-12-09 | 2024-12-09 | CEBU DOCTORS UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29955 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 77 ASSAY: Xpert CT/NG, COUNT: 222 ASSAY: Xpert HBV VL, COUNT: 191 ASSAY: Xpert HPV, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1077 ASSAY: Xpert HCV VL, COUNT: 157 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-IV | 110008524 | FOR MONTHLY PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095345655582905516/6095345655582905516_signature_28.png | Katrina Marie Balcera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6095345655582905516/6095345655582905516_signature_31.png
| 2024-12-09 | 2024-12-09 | Claver Rural Health Unit | Brgy. Tayaga, Claver | 13 | Surigao del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29954 | DOH | rhodaruthdalogog@gmail.com | 27 |  | 2024-11-27 | GX-4 | 838514 | Module B4 found service required thru XPC | Module Replacement | Replacing module B4 with SN: 694289 to SN: 718376 | All modules and units are functional |  | Certified pre-owned GeneXpert Module | 12378 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095325575426949694/6095325575426949694_signature_28.png | Joyce E. Andoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6095325575426949694/6095325575426949694_signature_31.png
| 2024-12-09 | 2024-12-09 | Dinas Rural Health Unit | Dinas, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29953 | DOH | bernantereynacres@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 679 | 2024-12-09 | GX 4 | 838533 | System unit unable to turn on | Replaced CMOS battery and re installed RAM, Preventive maintenence. | Replaced CMOS battery and re installed RAM, Preventive maintenence. | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095293624229072751/6095293624229072751_signature_28.png | Reyna Cres B. Lazaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6095293624229072751/6095293624229072751_signature_31.png
| 2024-12-09 | 2024-12-06 | Samboan Rural Health Unit | Samboan Poblacion, Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29952 | DOH | sheneannesenagan@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 274 |  | GX4 | 838513 | Module 1 862517 failed during calibration. | Replaced new module SN # 892506. Generate IQR , System log, Assay Statistics and Gxx files | Replaced new module SN # 892506. Generate IQR , System log, Assay Statistics and Gxx files | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095257585585497427/6095257585585497427_signature_28.png | Shene Anne A. Senagan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6095257585585497427/6095257585585497427_signature_31.png
| 2024-12-09 | 2024-12-09 | Pili Rural Health Unit | Poblacion, San Isidro, Pili | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29951 | DOH | lfsanao724@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1085 |  | GX4-6C | 838408 | Request for fan filter UPS no power | Replaced Fan filter Checked UPS power dependability Charged UPS for 4hrs | Performed Troubleshooting for UPS replaced fan filter | Unit is Operational | UPS lasts 5seconds after pulling the plug. Request for UPS battery replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095252458412116718/6095252458412116718_signature_28.png | Liezl F. Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/6095252458412116718/6095252458412116718_signature_31.png
| 2024-12-09 | 2024-12-09 | Zamboanga del sur | Pagadian city | 9 | Zamboanga del Sur | Laboratory | NO | Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29950 | DOH | limboemillie@gmail.com | 25 |  | 2024-11-27 | GX 4 | 838204 | None | Replaced module 210009828 with Module 890132 | Replaced module 210009828 with Module 890132 | All modules are available | None | 900-0491RS | 12370 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6095207722514033707/6095207722514033707_signature_28.png | Emillie limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/6095207722514033707/6095207722514033707_signature_31.png
| 2024-12-08 | 2024-12-07 | Marikina CHO | Marikina city | NCR | NCR 2nd - Marikina | Lab | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29949 | MACARE-HBDC | estyaragon2019@gmail.com | 28 |  |  | GX4 D 10C | 110024072 | As per end user they encounter water pipe leakage in laboratory | Checked unit Checked workstation,barcode scanner,printer,monitor Test run unit ok | Performed checked up | UPS AT1K240323ST0047 for observation because of some water residue,need to dry up |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6094695976399428949/6094695976399428949_signature_28.png | Estrelita s. Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6094695976399428949/6094695976399428949_signature_31.png
| 2024-12-08 | 2024-12-07 | Marikina CHO | Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29948 | MACARE-HBDC | estyaragon2019@gmail.com | 28 |  |  | GX4 D 10C | 110024073 | As per enduser they encounter water leakage inside laboratory | Checked unit Checked workstation,barcode scanner,printer,monitor Test run unit ok | Performed checked up | UPS AT1K240323ST0045 for observation because of some water residue,need to dry up |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6094695576394380504/6094695576394380504_signature_28.png | Estrelita s. Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/6094695576394380504/6094695576394380504_signature_31.png
| 2024-12-08 | 2024-12-06 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29947 | UNOPS | eamcmolecularlab@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 481 ASSAY: Xpert HPV, COUNT: 250 ASSAY: Xpert SARS-CoV-2, COUNT: 1251 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 197 | 2024-11-28 | GX4 D 10C | 110016781 | For module replacement B3 210084566 - 220008197 B1 220007358 encounter temp error for reporting. | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed module replacement | B1 disabled due to temp error B2,B3,B4 operational |  | Genexpert module | 12399 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6094695226396807384/6094695226396807384_signature_28.png | Aldwin Kristofer M. Orpilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6094695226396807384/6094695226396807384_signature_31.png
| 2024-12-07 | 2024-12-07 | Labuan General hospital | Labuan, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29946 | DOH | shairaanmaulod@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2301 | 2024-11-27 | GX IV | 838568 | Module 3 and 4 h/w failed | Replaced module 768535 and 674481 with new Modules 688833 and 701396 respectively. | Replaced module 768535 and 674481 with new Modules 688833 and 701396 respectively. | All modules are available | None | 900-0491RS | 12375 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6093564364419200391/6093564364419200391_signature_28.png | Shaira-An Asma F. Maulod | https://www.jotform.com/uploads/macarewecare/222927978372067/6093564364419200391/6093564364419200391_signature_31.png
| 2024-12-06 | 2024-12-06 | Chinese General Hospital | Blumentritt, Sta. Cruz Manila | NCR | NCR 1st - Manila | Bacteriology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 29945 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1091 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 57 |  | GXIV-10C | 120002026 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092768276429759274/6092768276429759274_signature_28.png | Mary Joyce Gaudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6092768276429759274/6092768276429759274_signature_31.png
| 2024-12-06 | 2024-12-06 | Bantayan Rural Health Unit | Bantayan Island | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29944 | DOH | mcdelle916bausin@gmail.com | 28 |  |  | Gx4 | 838509 | Defective module B3 712987 | New module installed b3 210044938 Clean Fan Filter data IQR, system log and Assay stats | New module installed b3 210044938 Clean Fan Filter data IQR, system log and Assay stats | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092767870126734360/6092767870126734360_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/6092767870126734360/6092767870126734360_signature_31.png
| 2024-12-06 | 2024-12-05 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29943 | DOH | danicaescorpion@gmail.com | 24 |  |  | Gx4 | 837612 | Defective module B4 678429 | New installed module B4 803293 Clean Fan Filter IQR, Sytem log and assay stats | New installed module B4 803293 Clean Fan Filter IQR, Sytem log and assay stats | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092767350123843572/6092767350123843572_signature_28.png | Danica Escorpuion | https://www.jotform.com/uploads/macarewecare/222927978372067/6092767350123843572/6092767350123843572_signature_31.png
| 2024-12-06 | 2024-12-06 | Pozorubio Rural Health Unit | Pozorubio Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29942 | DOH | rhupozorrubio@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 | 2024-11-29 | GXIV-D | 838467 | Module B1:638011 for module replacement | Preventive maintenance Repair | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unitis operational |  | 1 Genexpert module | 12402 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092766175517250277/6092766175517250277_signature_28.png | Jinky Grace C. Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/6092766175517250277/6092766175517250277_signature_31.png
| 2024-12-06 | 2024-12-06 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | annabellelcatbagan@gmail.com |  | 29941 | COMMERCIAL | labmedsci@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 206 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 34 |  | GXII-D | 820303 | Error 5007 | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational | Under observation for error 5007 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092765875514276142/6092765875514276142_signature_28.png | Annabelle L. Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6092765875514276142/6092765875514276142_signature_31.png
| 2024-12-06 | 2024-12-06 | Chinese General Hospital | Blumentritt, Sta. Cruz Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 29940 | CONSORTIUM | oncologycgh@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 14 ASSAY: Xpert HBV VL, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 596 ASSAY: Xpert SARS-CoV-2, COUNT: 378 ASSAY: Xpert HCV VL, COUNT: 29 ASSAY: Xpert HIV-1 Viral Load, COUNT: 11 |  | GXIV | 835604 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092756926424286666/6092756926424286666_signature_28.png | Jing Jing Hong | https://www.jotform.com/uploads/macarewecare/222927978372067/6092756926424286666/6092756926424286666_signature_31.png
| 2024-12-06 | 2024-12-06 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29939 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  | 2024-12-06 | GXIV 10C | 120002891 | Unit for installation | NA | #NAME? | Unit is operational |  |  | 99241 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092742358337283609/6092742358337283609_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6092742358337283609/6092742358337283609_signature_31.png
| 2024-12-06 | 2024-12-06 | WESTERN VISAYAS STATE UNIVERSITY HOSPITAL | Iloilo City | 6 | Iloilo | MOLECULAR LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29938 | UNOPS | mc-molecular@wvsu.edu.ph | 23 | ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert HPV, COUNT: 13 ASSAY: Xpert MTB-XDR, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV | 110021716 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092740475532146981/6092740475532146981_signature_28.png | R. MANALO JR. | https://www.jotform.com/uploads/macarewecare/222927978372067/6092740475532146981/6092740475532146981_signature_31.png
| 2024-12-06 | 2024-12-06 | Atimonan Rural Health Unit | G. Boulevard, Brgy. Zone 1, Atimonan, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29937 | DOH | sandycecille.perlado26@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 679 | 2024-11-20 | GXIV 6C D | 838206 | Module B3 for replacement | Performed PM Replacrd Module B3 Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced B3 625372 -> 210003782 | Unit is operational |  | Gx Module 6C | DR 12400 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092730883617430894/6092730883617430894_signature_28.png | Sandy Cecille Velasco | https://www.jotform.com/uploads/macarewecare/222927978372067/6092730883617430894/6092730883617430894_signature_31.png
| 2024-12-06 | 2024-12-06 | GUIMARAS PROVINCIAL HEALTH OFFICE | GUIMARAS, ILOILO | 6 | Guimaras | LABORATORY | NO | PACKING FOR TRANSFER | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29936 | UNOPS | mylene210@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1250 |  | GX-IV | 110023528 | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092642614629121156/6092642614629121156_signature_28.png | Mylene Publico | https://www.jotform.com/uploads/macarewecare/222927978372067/6092642614629121156/6092642614629121156_signature_31.png
| 2024-12-06 | 2024-12-06 | Abra Provincial Hospital | Bangued, Abra | CAR | Abra | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29935 | UNOPS | aphpathologydepartment@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 503 |  | GXIV-4-D-10C | 110023499 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092628802618483701/6092628802618483701_signature_28.png | Riverlie Valera | https://www.jotform.com/uploads/macarewecare/222927978372067/6092628802618483701/6092628802618483701_signature_31.png
| 2024-12-06 | 2024-12-05 | MMC Medilinx | Amorsolo St., Makati | NCR | NCR 4th - Makati | Molecular | NO | CALIBRATION | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29934 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 223 ASSAY: Xpert CT/NG, COUNT: 824 ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert HPV, COUNT: 729 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3346 ASSAY: Others, COUNT: 12 ASSAY: Xpert HCV VL, COUNT: 96 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1979 | 2024-12-05 | GX4D 10C | 110001647 | Due for xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck | B1 reported already for replacement B2,B3 failed xpertcheck B4 passed xpertcheck |  | Xpertcheck | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092568849708304709/6092568849708304709_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6092568849708304709/6092568849708304709_signature_31.png
| 2024-12-06 | 2024-12-06 | Manabo Rural Health Unit | Manabo, Abra | CAR | Abra | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29933 | DOH | lenmayclairelumiwag15@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1311 | 2024-11-22 | GXIV-4-D-6C | 838274 | For Troubleshooting, Preventive Maintenance and Calibration | #NAME? | Performed Troubleshooting, Preventive Maintenance and Calibration | - Unit is Partially Operational - Module A3 (807809) Failed XpertCheck, For Reporting |  | XpertCheck Kits | DR 12346 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092564582417943436/6092564582417943436_signature_28.png | Rachel A.Caddan | https://www.jotform.com/uploads/macarewecare/222927978372067/6092564582417943436/6092564582417943436_signature_31.png
| 2024-12-05 | 2024-12-05 | Negros Oriental Provincial Hospital | Dumaguete City, Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29932 | DOH | rezelmarielagos@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 ASSAY: Xpert MTB-XDR, COUNT: 38 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1803 | 2024-11-28 | GX-4 6c | 838525 | Module A4/690680 failed latest Xpertcheck. | #NAME? | - Replaced module A4 with new module SN 717194. | - A1, A2, A4 operational; - A3 for observation, encountering E4016. |  | Genexpert Module 6c | DR 12382 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6092074674478203327/6092074674478203327_signature_28.png | Rezel Marie Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/6092074674478203327/6092074674478203327_signature_31.png
| 2024-12-05 | 2024-12-05 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29931 | DOH | chaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1780 |  | GX4 | 838283 | Module 1 SN 808271 failed during calibration. | Replacement of New Module SN# 210051956 | Replacement of New Module SN# 210051956 | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091902808584185442/6091902808584185442_signature_28.png | Lucille Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6091902808584185442/6091902808584185442_signature_31.png
| 2024-12-05 | 2024-12-05 | Kalinga Provincial Hospital | Bolinao, Tabuk City, Kalinga | CAR | Kalinga | TB Dots Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | roxasmatthew659@yahoo.com.ph |  | 29930 | DOH | leombelgica139@gmail.com | 29 |  |  | GXIV-6C | 838466 | Due for Calibration | Performed Xpertcheck (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091895116612484561/6091895116612484561_signature_28.png | Leo Mark Belgica | https://www.jotform.com/uploads/macarewecare/222927978372067/6091895116612484561/6091895116612484561_signature_31.png
| 2024-12-05 | 2024-12-05 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29929 | DOH | luzvimindanavidad2@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1045 |  | 6C GXIV-4D | 838297 | None | Preventive Maintenance | Preventive Maintenance | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091880271418899391/6091880271418899391_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6091880271418899391/6091880271418899391_signature_31.png
| 2024-12-05 | 2024-12-05 | Nazareth Health Center | Brgy. Nazareth, Cagayan de Oro City | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29928 | DOH | nanettebtolentino@gmail.com | 25 |  |  | GX-4 | 838269 | Stuck Cartridge at Module B1 | Stuck Cartridge Removal | Manually removed Stuck Cartridge | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091872397473161391/6091872397473161391_signature_28.png | Nanette B. Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/6091872397473161391/6091872397473161391_signature_31.png
| 2024-12-05 | 2024-12-05 | Sta Cruz Rural Health Unit | Poblacion, Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29927 | DOH | chendalagan72@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 519 | 2024-11-28 | GX-IV 6C 4D | 838505 | Module B2 and B3 for Service Required | Replaced the Defective Module SN: 861943 with new Module SN: 662813 Replaced the Defective Module SN: 86144 with new Modules SN: 668209 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | 900-0491rs | DR 12388 and DR 12387 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091854566517373410/6091854566517373410_signature_28.png | Ilyssa Jann J Ibañez | https://www.jotform.com/uploads/macarewecare/222927978372067/6091854566517373410/6091854566517373410_signature_31.png
| 2024-12-05 | 2024-12-05 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29926 | UNOPS | benguetphortdl@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 656 |  | GXIV-4-D-10C | 110023498 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091841659128649989/6091841659128649989_signature_28.png | Frannie Calixto | https://www.jotform.com/uploads/macarewecare/222927978372067/6091841659128649989/6091841659128649989_signature_31.png
| 2024-12-05 | 2024-12-05 | Quezon City Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29925 | MACARE-HBDC | molecularlab.cityhealth@quezoncity.gov.ph | 25 |  |  | GXIV 10C | 120003581 | Unit for installation | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091796246917345366/6091796246917345366_signature_28.png | Cyrell Arcueno | https://www.jotform.com/uploads/macarewecare/222927978372067/6091796246917345366/6091796246917345366_signature_31.png
| 2024-12-05 | 2024-12-05 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29924 | MACARE-HBDC | ruellamorena@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 340 ASSAY: Xpert SARS-CoV-2, COUNT: 1909 ASSAY: Xpert HCV VL, COUNT: 20 ASSAY: Xpert HIV-1 Viral Load, COUNT: 252 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 886 | 2024-11-29 | GXIV-6C | 110008457 | Module 4 Error | Performed Module Replacement B4(210034759 - 611093) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Module | 12401 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091765128618803470/6091765128618803470_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/6091765128618803470/6091765128618803470_signature_31.png
| 2024-12-05 | 2024-12-05 | Baguio Health Services Office | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29923 | DOH | sherryannemarrero@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2286 |  | GXIV-4-D-6C | 838281 | For Preventive Maintenance and Troubleshooting | #NAME? | Performed Preventive Maintenance and Troubleshooting | - Unit is Partially Operational - Module 1 (678653) For Reporting - Module 2 (651808) Already Reported |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091760247518857555/6091760247518857555_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/6091760247518857555/6091760247518857555_signature_31.png
| 2024-12-05 | 2024-12-05 | ILOILO PROVINCIAL HOSPITAL | Pototan, Iloilo | 6 | Iloilo | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29922 | DOH | djpragale92@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 650 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 | 2024-11-25 | GX-IV | 838289 | FOR MODULE C4 REPLACEMENT | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | -PERFORM PM PROCEDURE -PERFORM MODULE C4 REPLACEMENT DEF SN:808247 NEW SN: 890802 | EQUIPMENT IS OPERATIONAL |  | GeneXpert Module | 12357 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091747887421479717/6091747887421479717_signature_28.png | Rachel A. Pragale | https://www.jotform.com/uploads/macarewecare/222927978372067/6091747887421479717/6091747887421479717_signature_31.png
| 2024-12-05 | 2024-12-05 | Garcia Memorial Hospital | Talibon Bohol | 7 | Bohol | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29921 | DOH | ybanezrosalina19@gmail.com | 23 |  | 2024-12-05 | GX4 | 838203 | Defective Module B4 with serial number 807026. | . Save Installation Qualification Report. . Save System Log file. . Save Assay Statistics Report | Replace Module B4 with new serial number 807026. | All module's are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091719943321941080/6091719943321941080_signature_28.png | Rosalina A. Ybañez | https://www.jotform.com/uploads/macarewecare/222927978372067/6091719943321941080/6091719943321941080_signature_31.png
| 2024-12-05 | 2024-12-05 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29920 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 22 | ASSAY: Xpert HBV VL, COUNT: 212 ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert SARS-CoV-2, COUNT: 198 ASSAY: Xpert HCV VL, COUNT: 46 ASSAY: Xpert HIV-1 Viral Load, COUNT: 38 |  | GXIV-4-D-10C | 110016112 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091713017515950086/6091713017515950086_signature_28.png | Rhea Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6091713017515950086/6091713017515950086_signature_31.png
| 2024-12-05 | 2024-12-05 | DON JOSE S. MONFORT MEDICAL CENTER | Barotac Nuevo, Iloilo | 6 | Iloilo | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29919 | DOH | judyboglosa35@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2170 | 2024-11-25 | GX-IV | 838501 | FOR MODULE B3 REPLACEMENT | COLLECT IQR,SYSTEM LOG, AND ARCHIVE | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B3 DEF SN: 862361 REPLACEMENT WITH NEW SN:210030546 | EQUIPMENT IS OPERATIONAL | MODULE B2 FOR REPORTING | Genexpert Module | 12359 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091697335417284331/6091697335417284331_signature_28.png | Judy boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6091697335417284331/6091697335417284331_signature_31.png
| 2024-12-04 | 2024-12-04 | Teresita L. Jalandoni Provincial Hospital | Silay City | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29918 | DOH | tljphmollab02@gmail.com | 28 |  |  | Gx 4 | 838464 | Defective module A4 sn 805849 | New module installed A4 SN 675288 Clean Fan Filter Lastest IQR, system log and assay stats | New module installed A4 SN 675288 Clean Fan Filter Lastest IQR, system log and assay stats | Systems operational Module A2 Error 1002 from 10/4/2024 to 10/11/2024 Module A1 Error 5015, 5011, 5016 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091107618427237667/6091107618427237667_signature_28.png | Richelle Anne A. APUYAN | https://www.jotform.com/uploads/macarewecare/222927978372067/6091107618427237667/6091107618427237667_signature_31.png
| 2024-12-04 | 2024-12-04 | Monkayo Rural Health Unit | Poblacion, Monkayo Davao de Oro | 11 | Davao de Oro | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29917 | DOH | kendralois730@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1799 | 2024-12-27 | GX-IV 6C 4D | 838487 | Module B1 and B3 for Service Required Due for Preventive Maintenance UPS problem Monitor Display Problem | Replaced the Defective Module SN: 643605 with new Module SN: 730899 Replaced the Defective Module SN: 862220 with new Module SN: 761214 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Check line voltages for UPS | Perform Module Replacement and Preventive Maintenance | All Modules are operational Documented the error at UPS, no battery during battery mode Documented LCD Display Error Errors will be sent for further assessment |  | 900-0491rs | DR 12373 and DR 12374 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091045964444670498/6091045964444670498_signature_28.png | Karen S. Anuta,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6091045964444670498/6091045964444670498_signature_31.png
| 2024-12-04 | 2024-12-04 | Baggao Rural Health Unit | San Jose, Baggao, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29916 | DOH | baggaorhulaboratory@gmail.com | 29 |  |  | GXIV-6C | 838531 | CPU Hard to power on | Troubleshoot CPU Replaced CMOS Battery | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | CPU for Observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091043592918453459/6091043592918453459_signature_28.png | Jeany Anne Galera | https://www.jotform.com/uploads/macarewecare/222927978372067/6091043592918453459/6091043592918453459_signature_31.png
| 2024-12-04 | 2024-12-04 | Valencia City Health Office | New Terminal, Hagkol, Valencia | 10 | Bukidnon | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29915 | DOH | caluyajopay@gmail.com | 24 |  | 2024-11-26 | GX-4 | 838271 | None | Module Replacement | Replacing Module B4 with SN: 808072 to SN: 879826 | All modules and units are functional |  | Certified pre-owned GeneXpert Module | 12364 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091032914732465819/6091032914732465819_signature_28.png | Jonalyn Dablio | https://www.jotform.com/uploads/macarewecare/222927978372067/6091032914732465819/6091032914732465819_signature_31.png
| 2024-12-04 | 2024-12-04 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | daguilon.intp@gmail.com |  | 29914 | MACARE-HBDC | msff-tondo-labtech@paris.msf.org | 25 |  |  | GX16-10C | 110016343 | For troubleshooting c/o Cepheid instructions. | Troubleshooting. | Performed troubleshooting. Swapped module C3 to C4. Performed blank test running. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6091013033961525057/6091013033961525057_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6091013033961525057/6091013033961525057_signature_31.png
| 2024-12-04 | 2024-12-04 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29913 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 03J030 | For checking of centrifuge. | Troubleshooting. | Performed troubleshooting of MCS+. Verified calibration parameters. Performed simulation running. Performed multiple cycles. | Unit is operational. Centrifuge calibration parameters are okay. For replacement of worn-out rubber roller. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090974402126768108/6090974402126768108_signature_28.png | ROSELYN JRA C ALUA | https://www.jotform.com/uploads/macarewecare/222927978372067/6090974402126768108/6090974402126768108_signature_31.png
| 2024-12-04 | 2024-12-04 | Baybay City Health Office | Baybay City | 8 | Leyte | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29912 | DOH | sajecopooh20@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1152 | 2024-12-04 | GXIV | 838555 | Module A4/862905 defective | #NAME? | #NAME? | Unit is Operational |  | 1 Genexpert Module 6C | 12379 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090965134533268691/6090965134533268691_signature_28.png | Sarah Jean G. Pasana | https://www.jotform.com/uploads/macarewecare/222927978372067/6090965134533268691/6090965134533268691_signature_31.png
| 2024-12-04 | 2024-12-04 | Alfonso Ponce Enrile Memorial District Hospital | Brgy Smart, Gonzaga, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29911 | DOH | apemdh_laboratory@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1480 | 2024-11-25 | GXIV-6C | 838474 | Module 1 Error UPS problem | Performed Module Replacement B1(698770 - 210004167) Replaced UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | UPS and Genexpert Module | 12356 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090957818801698018/6090957818801698018_signature_28.png | Regina Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6090957818801698018/6090957818801698018_signature_31.png
| 2024-12-04 | 2024-12-04 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mcho.laboratory102018@gmail.com |  | 29910 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1162 ASSAY: Xpert MTB-XDR, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110016107 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules operational. | All modules passed. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090949387025308118/6090949387025308118_signature_28.png | ESTRELITA S. ARAGON | https://www.jotform.com/uploads/macarewecare/222927978372067/6090949387025308118/6090949387025308118_signature_31.png
| 2024-12-04 | 2024-12-04 | Calubian Rural Health Unit | Calubian, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29909 | DOH | mjcabardo06@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 939 |  | 6C GXIV-4D | 838399 | Due for PM | Preventive Maintenance | Preventive Maintenance | Machine is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090919723677206562/6090919723677206562_signature_28.png | Sherlita Esmeralda | https://www.jotform.com/uploads/macarewecare/222927978372067/6090919723677206562/6090919723677206562_signature_31.png
| 2024-12-04 | 2024-12-04 | Quezon Rural Health Unit | Poblacion, Quezon | 10 | Bukidnon | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29908 | DOH | quenxlopez@gmail.com | 28 |  | 2024-11-27 | GX-4 6 Colors | 838440 | None | Module Replacement | Replacing Module B2 with SN: 629512 to SN: 882700 Preventive Maintenance | All modules and units are functional |  | Certified pre-owned GeneXpert Module | 12372 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090888310526770284/6090888310526770284_signature_28.png | Quenie Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6090888310526770284/6090888310526770284_signature_31.png
| 2024-12-04 | 2024-12-04 | TAPAZ RURAL HEALTH UNIT | TAPAZ, CAPIZ | 6 | Capiz | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29907 | DOH | rhulabtapaz@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 638 ASSAY: Xpert HIV-1 Viral Load, COUNT: 68 | 2024-11-27 | GX-IV | 837615 | FOR MODULE B2 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B2 REPLACEMENT DEF SN: 806570 NEW SN: 890453 | EQUIPMENT IS OPERATIONAL |  | Certified Pre-Owned GeneXpert Module | 12397 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090887877357980907/6090887877357980907_signature_28.png | Ma. Jorlyn Oliva | https://www.jotform.com/uploads/macarewecare/222927978372067/6090887877357980907/6090887877357980907_signature_31.png
| 2024-12-04 | 2024-12-04 | Lasam District Hospital | IBJ, Lasam, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29906 | DOH | lasamdh.laboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 334 | 2024-09-26 | GXIV-6C | 838530 | UPS Defective | Replaced UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | UPS | 12172 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090880577092850152/6090880577092850152_signature_28.png | Kimberly Nicolas | https://www.jotform.com/uploads/macarewecare/222927978372067/6090880577092850152/6090880577092850152_signature_31.png
| 2024-12-04 | 2024-12-04 | Lasam District Hospital | IBJ, Lasam, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29905 | DOH | lasamdh.laboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 334 | 2024-09-26 | GXIV-6C | 838530 | UPS Defective | Replaced UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | UPS | 12172 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090880567092981559/6090880567092981559_signature_28.png | Kimberly Nicolas | https://www.jotform.com/uploads/macarewecare/222927978372067/6090880567092981559/6090880567092981559_signature_31.png
| 2024-12-03 | 2024-12-03 | Mabinay Medicare Community Hospital | Mabinay, Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29904 | DOH | diannecm8@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 602 | 2024-11-25 | GX4 6C | 838481 | Module B4/703735 failed latest Xpertcheck; status HW Fail | #NAME? | Replaced module B4 with new module SN 210074162. | 4 modules operational. |  | Genexpert module 6c | DR 12352 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090362125425780860/6090362125425780860_signature_28.png | Dianne Mi-ot, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6090362125425780860/6090362125425780860_signature_31.png
| 2024-12-03 | 2024-12-03 | Rapu-Rapu Rural Health Unit | Poblacion, Rapu-Rapu | 5 | Albay | RTDL Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29903 | DOH | suzetteoguan1@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 607 |  | GX4-6C | 838565 | B1/862627 for Module Replacement | Module Replaced: B1 - 862627>>601367 Cleaned cartridge bays, plunger rod Generate | B1/862627 module replaced Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090278504319110083/6090278504319110083_signature_28.png | Suzette C. Oguan | https://www.jotform.com/uploads/macarewecare/222927978372067/6090278504319110083/6090278504319110083_signature_31.png
| 2024-12-03 | 2024-12-03 | Likhaan Center for Women's Health | Bahay Toro, Project 8, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29902 | MACARE-HBDC | office@likhaan.org | 25 |  |  | GXIV 10C L | 110025174 | Unit due for maintenance | Performed preventive maintenance Performed reliability test | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Installed TV assay | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090204157898253125/6090204157898253125_signature_28.png | Anne Kristine Oquila | https://www.jotform.com/uploads/macarewecare/222927978372067/6090204157898253125/6090204157898253125_signature_31.png
| 2024-12-03 | 2024-12-03 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | dy_rosemarie71@gmail.com |  | 29901 | DOH | dyrosemarie71@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1325 | 2024-11-25 | GX-IV 6C 4D | 838401 | Module B3 for Service Required Due for Preventive Maintenance | Replaced the Defective Module SN: 860973 with new Module SN: 912643 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Performed Preventive Maintenance and Module Replacement | All Modules are operational |  | 900-0491rs | DR 12361 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090168626773396126/6090168626773396126_signature_28.png | Rosemarie B. Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/6090168626773396126/6090168626773396126_signature_31.png
| 2024-12-03 | 2024-12-03 | Project 7 Social Hygiene Clinic | Project 7, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29900 | MACARE-HBDC | michelenaval@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 21 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HCV VL, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 465 |  | GXIV 10C D | 120000577 | Unit due for maintenance | Performed preventive maintenance Performed reliability test Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Archived results | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090150388917314125/6090150388917314125_signature_28.png | Michele Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6090150388917314125/6090150388917314125_signature_31.png
| 2024-12-03 | 2024-12-03 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29899 | DOH | cholaboratory23@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1159 | 2024-12-03 | GXIV | 838421 | n/a | n/a | Monthly preventive maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6090134245423849426/6090134245423849426_signature_28.png | Ma. Gennette Dolly D. Garapan | https://www.jotform.com/uploads/macarewecare/222927978372067/6090134245423849426/6090134245423849426_signature_31.png
| 2024-12-03 | 2024-12-03 | BIOPATH CAP | Cebu City | 7 | Cebu | Molecular Lab | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29898 | COMMERCIAL | jkee@biopath.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HPV, COUNT: 27 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 18 ASSAY: Xpert HCV VL, COUNT: 4 |  | GX-XVI | 838657 | #NAME? | #NAME? | -PERFORM DIAGNOSIS AND TROUBLESHOOTING, CONFIRM UPS DEFECTIVE -UPS MODEL PRO902-WS DEFECTIVE SN: 565701183400184 | GX EQUIPMENT IS OPERATIONAL, UPS/UPS BATTER FOR REPLACEMENT |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089984935847250855/6089984935847250855_signature_28.png | Mary Joy C. kee | https://www.jotform.com/uploads/macarewecare/222927978372067/6089984935847250855/6089984935847250855_signature_31.png
| 2024-12-02 | 2024-11-29 | Jose Abad Santos Rural Health Unit | Poblacion, Jose Abad Santos | 11 | Davao Occidental | TB-DOTS | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29897 | DOH | bevespinosa168@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 345 | 2024-09-04 | GX-IV 6C 4D | 838460 | Due for Annual XpertCheck Calibration For Preventive Maintenance UPS no battery load | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Run XpertCheck on four modules Check Line Voltages and troubleshoot the UPS Install GX Sticker | Perform Preventive Maintenance, XPC Calibration and UPS Troubleshooting | Module B2 for Service Required Defective UPS, no load during Battery Mode | UPS error Documented for further assessment | XPERTCHECK CE 5 | DR 12002 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089338879123115815/6089338879123115815_signature_28.png | Beverly Espinosa,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6089338879123115815/6089338879123115815_signature_31.png
| 2024-12-02 | 2024-11-29 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29895 | MACARE-HBDC | madriaga860@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 391 ASSAY: Xpert MTB-XDR, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089290013617524725/6089290013617524725_signature_28.png | John Michael Madriaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6089290013617524725/6089290013617524725_signature_31.png
| 2024-12-02 | 2024-11-29 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29896 | MACARE-HBDC | madriaga860@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1174 ASSAY: Xpert MTB-XDR, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089290013619157983/6089290013619157983_signature_28.png | John Michael Madriaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6089290013619157983/6089290013619157983_signature_31.png
| 2024-12-02 | 2024-12-02 | Mini Forest Health Center (District C) | Quezon Boulevard, Brgy Poblacion, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29894 | DOH | miniforestlaboratory@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 392 | 2024-11-25 | GX-IV 6C 4D | 838507 | Module B2 for service required , and due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replace the defective Module SN: 860769 with new Module SN: 210061057 | Perform module replacement and Preventive Maintenance | All Modules are operational |  | 900-0491rs | DR 12353 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089283606771045494/6089283606771045494_signature_28.png | Joy L. Nio,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6089283606771045494/6089283606771045494_signature_31.png
| 2024-12-02 | 2024-11-28 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29893 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 205 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 29 |  | GXII-2D | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089277868448658999/6089277868448658999_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6089277868448658999/6089277868448658999_signature_31.png
| 2024-12-02 | 2024-11-27 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29892 | DOH | allan.vinluan2013@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 938 |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089276688443707878/6089276688443707878_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6089276688443707878/6089276688443707878_signature_31.png
| 2024-12-02 | 2024-11-27 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29891 | DOH | seansimonpaloyo@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1224 | 2024-04-04 | GXIV-4-D-6C | 838473 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  | XpertCheck Kits | DR 11361 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089272608446055868/6089272608446055868_signature_28.png | Manolo Paloyo | https://www.jotform.com/uploads/macarewecare/222927978372067/6089272608446055868/6089272608446055868_signature_31.png
| 2024-12-02 | 2024-11-14 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29890 | DOH | dianeruiz147@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 468 |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089271038449392826/6089271038449392826_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6089271038449392826/6089271038449392826_signature_31.png
| 2024-12-02 | 2024-11-14 | Urdaneta Sacred Heart Hospital | Urdaneta, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29889 | COMMERCIAL | annabellelcatbagan@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 205 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 24 |  | GXIV-4-D-6C | 820303 | For Preventive Maintenance due to Error 5007 | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089268978442096741/6089268978442096741_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6089268978442096741/6089268978442096741_signature_31.png
| 2024-12-02 | 2024-11-14 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29888 | DOH | serly_reyes@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089268028441185032/6089268028441185032_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6089268028441185032/6089268028441185032_signature_31.png
| 2024-12-02 | 2024-12-02 | Detoxicare Molecular Diagnostics Laboratory | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | fortuna.mandapat@detoxicare.com.ph |  | 29887 | CONSORTIUM | moses.pineda@detoxicare.com.ph | 25 |  |  | GXIV 10C | 110018312 | Unit for installation Unit from Detoxicare Shaw | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089221673371224619/6089221673371224619_signature_28.png | John Moses Pineda | https://www.jotform.com/uploads/macarewecare/222927978372067/6089221673371224619/6089221673371224619_signature_31.png
| 2024-12-02 | 2024-12-02 | Detoxicare Molecular Diagnostics Laboratory | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29886 | CONSORTIUM | molecular.onco@detoxicare.com.ph | 25 |  |  | GXIV 10C | 110018312 | Unit for relocation | NA | Pulled out unit | Unit for transfer |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089216893372797183/6089216893372797183_signature_28.png | Mary Joy Bagsic | https://www.jotform.com/uploads/macarewecare/222927978372067/6089216893372797183/6089216893372797183_signature_31.png
| 2024-12-02 | 2024-11-26 | National Center for Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29885 | MACARE-HBDC | iancristrono@gmail.com | 25 |  |  | GXIV 10C | 120003480 | Unit for installation | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089170473377898560/6089170473377898560_signature_28.png | Ian Cris Tono | https://www.jotform.com/uploads/macarewecare/222927978372067/6089170473377898560/6089170473377898560_signature_31.png
| 2024-12-02 | 2024-11-28 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29884 | COMMERCIAL | cduhbloodbank2@gmail.com | 25 |  |  | MCS+ | 17H639SPG | Line sensor no minimum sensitivity limit problem and UPS battery problem. | Check line sensor and UPS battery supply voltage. Replacement of pump in blood draw line. | Check line sensor and UPS battery supply voltage. | The unit is standby for UPS battery replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089161320524741040/6089161320524741040_signature_28.png | CDUH BLOODBANK | https://www.jotform.com/uploads/macarewecare/222927978372067/6089161320524741040/6089161320524741040_signature_31.png
| 2024-12-02 | 2024-11-29 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29883 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2306 ASSAY: Xpert SARS-CoV-2, COUNT: 1071 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1425 |  | GX16D | 110014171 | Encounter continous error 5007 in xpress cov2/flu/rsv plus Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational in other assay |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089093543617431161/6089093543617431161_signature_28.png | Jessica Faye Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6089093543617431161/6089093543617431161_signature_31.png
| 2024-12-02 | 2024-11-29 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29882 | MACARE-HBDC | lcpmicrolab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 309 ASSAY: Xpert SARS-CoV-2, COUNT: 600 ASSAY: Xpert HIV-1 Viral Load, COUNT: 107 |  | GX4 D | 838837 | Due for maintenance B1,B4 reported already failed in xpertcheck B3 encounter intermittent temp error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B1,B3,B4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089093193619956446/6089093193619956446_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6089093193619956446/6089093193619956446_signature_31.png
| 2024-12-02 | 2024-11-28 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29881 | MACARE-HBDC | molecularlabnkti@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1545 ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Others, COUNT: 121 ASSAY: Xpert HBV VL, COUNT: 1065 ASSAY: Xpert HPV, COUNT: 145 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2369 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 8441 ASSAY: Xpert HCV VL, COUNT: 151 |  | GX16D | 110005587 | Due for maintenance A4 for reporting encounter temp error 1002,4009 A1,A2,C2,,D2 reported already for replacement C4 reported already warranty | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | 10 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089092503619703960/6089092503619703960_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6089092503619703960/6089092503619703960_signature_31.png
| 2024-12-02 | 2024-11-28 | National Kidney and Transplant Institue | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29880 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 64 ASSAY: Others, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 24 ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 100 ASSAY: Xpert HCV VL, COUNT: 8 |  | GX4 D | 838524 | Due for maintenance B4 encounter intermmitent temp error 4012,1002 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B4 for observation | Work with Gian Magnaye |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089091863611453771/6089091863611453771_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6089091863611453771/6089091863611453771_signature_31.png
| 2024-12-02 | 2024-11-28 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29879 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 115 |  | GX16D | 804476 | Due for maintenance A1,A2,B1,B3,C1,,C3,C4 encounter loss comm. B4 encounter continous temp error 4010,4009,4012 C2 intermittent encounter 4012 Malfunction of lan cable rj45 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | A3,A4,B2 operational | For replacement of lan cable rj45 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6089090833614954389/6089090833614954389_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6089090833614954389/6089090833614954389_signature_31.png
| 2024-12-01 | 2024-11-27 | Isulan Rural Health Unit | Poblacion Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29878 | UNOPS | nancymoscoso18@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 467 |  | GxIV-10C | 110023520 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6088641416865779062/6088641416865779062_signature_28.png | Nancy A. Moscoso | https://www.jotform.com/uploads/macarewecare/222927978372067/6088641416865779062/6088641416865779062_signature_31.png
| 2024-12-01 | 2024-11-28 | Lake Sebu Rural Health Unit | Prk. Rosas, Lake Sebu, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | HAEMONETICS | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29877 | DOH | jennypingkian23@yahoo.com | 28.7 |  |  | GxIV-6C | 838569 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6088641416869865261/6088641416869865261_signature_28.png | Jenny A. Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/6088641416869865261/6088641416869865261_signature_31.png
| 2024-11-29 | 2024-11-29 | Talavera General Hospital | Maestrang Kikay, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29876 | MACARE-HBDC | pjgtehlab@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 94 ASSAY: Xpert MTB/RIF Ultra, COUNT: 692 ASSAY: Xpert CT/NG, COUNT: 53 ASSAY: Xpert HCV VL, COUNT: 8 ASSAY: Xpert HIV-1 Viral Load, COUNT: 205 |  | GXIV-6C | 110013773 | Due for Maintenance | Performed Maintenance Cleaned Plunger Rod Cleaned Cartridge Bay Cleaned Filter Cleaned Icore | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6086736102229795477/6086736102229795477_signature_28.png | Sharey R. Capati | https://www.jotform.com/uploads/macarewecare/222927978372067/6086736102229795477/6086736102229795477_signature_31.png
| 2024-11-29 | 2024-11-29 | Love Your Self | Cebu City | 7 | Cebu | HIV Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 29875 | COMMERCIAL | paulo14reforzado@gmail.com | 23 |  | 2024-11-29 | GX4 | 814150 | Due for Preventive maintenance. | Save System Log Save Installation Qualification Save Assay statistics | . Perform preventive maintenance fo GeneXert Machine. . Plunger rod maintenance . Clean filter | Module 1,2,3 ok to use 4 not ok use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6086734610214681655/6086734610214681655_signature_28.png | Paulo Crispo Reforzado Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/6086734610214681655/6086734610214681655_signature_31.png
| 2024-11-29 | 2024-11-29 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29874 | MACARE-HBDC | madriaga860@gmail.com | 22 |  |  | GX-IV-D-10C | 110024073 | For installation | Installed the GeneXpert system | Installed the GeneXpert system: GeneXpert unit Workstation Barcode scanner Printer UPS Performed maintenance Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6086690722558918770/6086690722558918770_signature_28.png | John Michael Madriaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6086690722558918770/6086690722558918770_signature_31.png
| 2024-11-29 | 2024-11-29 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29873 | MACARE-HBDC | madriaga860@gmail.com | 22 |  |  | GX-IV-D-10C | 110024072 | For installation | Installed the GeneXpert system | Installated the GeneXpert system: GeneXpert unit Workstation Barcode scanner Printer UPS Performed maintenance Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6086687172559310214/6086687172559310214_signature_28.png | John Michael Madriaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6086687172559310214/6086687172559310214_signature_31.png
| 2024-11-29 | 2024-11-29 | Labo District Hospital | Talobatib, Labo | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29872 | DOH | analyn221977@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1202 |  | GX4-6C | 838207 | GX4-6C for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6086666981329133018/6086666981329133018_signature_28.png | Analyn Consuelo | https://www.jotform.com/uploads/macarewecare/222927978372067/6086666981329133018/6086666981329133018_signature_31.png
| 2024-11-29 | 2024-11-29 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29871 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2377 ASSAY: Xpert HIV-1 Viral Load, COUNT: 306 |  | 6C GXIV-4D | 838285 | Due for monthly PM | #NAME? | #NAME? | Unit is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6086645474867193926/6086645474867193926_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/6086645474867193926/6086645474867193926_signature_31.png
| 2024-11-29 | 2024-11-29 | Vinzons Rural Health Unit | Poblacion, Vinzons | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29870 | DOH | bardonamabel@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1122 |  | GX4-6C | 838576 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6086570371324000443/6086570371324000443_signature_28.png | Amabel A. Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/6086570371324000443/6086570371324000443_signature_31.png
| 2024-11-29 | 2024-11-28 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29869 | UNOPS | acho.laboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 94 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6086567267517846473/6086567267517846473_signature_28.png | Elaine Ouano | https://www.jotform.com/uploads/macarewecare/222927978372067/6086567267517846473/6086567267517846473_signature_31.png
| 2024-11-28 | 2024-11-28 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29868 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  | 2024-11-28 | MCS+ | 12D063 | For installation of CPU board. | Repair and maintenance. | Performed installation of CPU board. Verified calibration parameters. Performed simulation running. | Unit is now operational. For configuration of UPP parameters c/o Marvin. |  | MCS+ CPU PCB Board Assy. | DR#12380 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6085908942122638859/6085908942122638859_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6085908942122638859/6085908942122638859_signature_31.png
| 2024-11-28 | 2024-11-28 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29867 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  | 2024-11-28 | MCS+ | 03J030 | For replacement of centrifuge. | Repair and maintenance. | Performed installation of centrifuge. Verified calibration parameters. Performed simulation running. | Unit is now operational. |  | MCS+ Centrifuge Assy. | DR#12381 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6085907682127121034/6085907682127121034_signature_28.png | Dolla Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6085907682127121034/6085907682127121034_signature_31.png
| 2024-11-28 | 2024-11-28 | Balamban Rural Health Unit | Poblacion Balamban | 7 | Cebu | TB Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29866 | DOH | climaco.mildred@yahoo.com | 23 |  | 2024-11-28 | GX 4 | 838231 | . Always error code 2008 . Due for Preventive Maintenance | Save System Log file Save Installation Qualifications Save Assay statistics report | Perform Preventive Maintenance gor all module's. | All module's are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6085882854119859186/6085882854119859186_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6085882854119859186/6085882854119859186_signature_31.png
| 2024-11-28 | 2024-11-28 | Ragay Rural Health Unit | Poblacion, Ragay | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29865 | DOH | phylgomez0115@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 686 |  | GX4-6C | 838560 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Checked fan filter Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6085881196639576298/6085881196639576298_signature_28.png | Dr. Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/6085881196639576298/6085881196639576298_signature_31.png
| 2024-11-28 | 2024-11-27 | Philippine Blood Center | Quezon Avemue QC | NCR | NCR 2nd - Quezon City | Donor room | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | strobecatcher@gmail.com |  | 29864 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15H029G | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6085628553615935835/6085628553615935835_signature_28.png | Ivan Ritz Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/6085628553615935835/6085628553615935835_signature_31.png
| 2024-11-28 | 2024-11-27 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Donor Room | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | strobecatcher@gmail.com |  | 29863 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS+ | 15H030G | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenancs | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6085628023615818947/6085628023615818947_signature_28.png | IVAN RITZ ROMERO | https://www.jotform.com/uploads/macarewecare/222927978372067/6085628023615818947/6085628023615818947_signature_31.png
| 2024-11-27 | 2024-11-27 | Jose B. Lingad Memorial and General Hospital | San Fernando City Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29862 | MACARE-HBDC | jblmghmpl@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 84 ASSAY: Xpert MTB/RIF Ultra, COUNT: 279 ASSAY: Xpert SARS-CoV-2, COUNT: 848 ASSAY: Others, COUNT: 65 |  | GXIV-D 10C | 110012825 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6085069875515127331/6085069875515127331_signature_28.png | Vince Daryll Navarro | https://www.jotform.com/uploads/macarewecare/222927978372067/6085069875515127331/6085069875515127331_signature_31.png
| 2024-11-27 | 2024-11-27 | Jose B. Lingad Memorial and General Hospital | San Fernando City, Pampanga | 3 | Pampanga | RHIVDA Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29861 | MACARE-HBDC | jblmgh.is@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 242 ASSAY: Xpert HCV VL, COUNT: 54 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1683 |  | GXIV-D 10C | 110013748 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6085069605519354285/6085069605519354285_signature_28.png | Criselle C. Medina | https://www.jotform.com/uploads/macarewecare/222927978372067/6085069605519354285/6085069605519354285_signature_31.png
| 2024-11-27 | 2024-11-27 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29860 | DOH | laureanjo@yahoo.com | 26 |  |  | Gx4 | 838527 | Monthly Preventive Maintenance | Preventive Maintenance Clean Fan Filter IQR, SYSTEM LOG & ASSAY STATS | Preventive Maintenance Clean Fan Filter IQR, SYSTEM LOG & ASSAY STATS | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6085021267414777430/6085021267414777430_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6085021267414777430/6085021267414777430_signature_31.png
| 2024-11-27 | 2024-11-26 | Ospital ng Makati | Sampaguita St. Makati | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 29859 | MACARE-HBDC | Kaylareynaldo77@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 198 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6084870079633597903/6084870079633597903_signature_28.png | Cristina Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6084870079633597903/6084870079633597903_signature_31.png
| 2024-11-27 | 2024-11-26 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | avabichuela3216@gmail.com |  | 29858 | MACARE-HBDC | molecularlab@phc.gov.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 622 ASSAY: Others, COUNT: 6 ASSAY: Xpert SARS-CoV-2, COUNT: 835 | 2024-11-21 | GX16 D | 838655 | Fue for xpertcheck A1,A2,B1,B2,C2,D1,D3,D4 | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck | A1,D4,B1 passed xpertcheck D1,D3 retry lack of xpertcheck kits A2,B2,C2 failed in xpertcheck |  | Xpertcheck kits | 12341 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6084867172596110679/6084867172596110679_signature_28.png | Philippine Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/6084867172596110679/6084867172596110679_signature_31.png
| 2024-11-26 | 2024-11-26 | Quezon City General Hospital (2) | Seminary Rd. Proj 8, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | gpdmagnaye30@gmail.com |  | 29857 | MACARE-HBDC | jomarilitonjua@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 52 |  | GXIV 10C D | 110023495 | Unit due for maintenance | Performed preventive maintenance Performed reliability test Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6084107071419452058/6084107071419452058_signature_28.png | Jomari Bryan D. Litonjua | https://www.jotform.com/uploads/macarewecare/222927978372067/6084107071419452058/6084107071419452058_signature_31.png
| 2024-11-26 | 2024-11-26 | Magsaysay Rural Health Unit | Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29856 | DOH | jkakilgan12@gmail.com | 25.6 |  |  | GX-4 | 838557 | Time not sync. Unable to boot cpu. | Troubleshooting | CMOS Battery Replacement, Cleaning and Reseating RAM Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6084089372411923097/6084089372411923097_signature_28.png | Jesusa Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/6084089372411923097/6084089372411923097_signature_31.png
| 2024-11-26 | 2024-11-26 | Hi-Precision Diagnostics | N.S. Amoranto Sr. St., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29855 | CONSORTIUM | pcr@hi-precision.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 118 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7582 |  | GXIV 6C D | 835602 | A2 Error 5007 | Performed preverntive maintenance Archived results Generated report | View results Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability Test | Unut is operational A2 for monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6084060571114436975/6084060571114436975_signature_28.png | Andrei Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/6084060571114436975/6084060571114436975_signature_31.png
| 2024-11-26 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | RBC BIOSCIENCE | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29854 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MagCore Plus II | P26R021720135 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Checked and verified alignment values Performed water run Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083936123617211626/6083936123617211626_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083936123617211626/6083936123617211626_signature_31.png
| 2024-11-25 | 2024-11-25 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29853 | DOH | burauenlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 416 | 2024-11-25 | GXIV | 838551 | UPS defective SN/ AT1000190123ST0195 | Unit is operational | #NAME? | Unit is operational |  | 1 unit UPS | 5429 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083355124694118920/6083355124694118920_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6083355124694118920/6083355124694118920_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | PERKIN ELMER | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29851 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | Pre Nat II | 3310-2141 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Checked and verified alignment values Performed water run Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083287896314838981/6083287896314838981_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083287896314838981/6083287896314838981_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | RBC BIOSCIENCE | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29852 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MagCore Plus II | P26R022020135 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Checked and verified alignment values Performed water run Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083287896319853796/6083287896319853796_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083287896319853796/6083287896319853796_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29850 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | 88400D | 160414501161220 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified temperatures at -80°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230533611724718/6083230533611724718_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230533611724718/6083230533611724718_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph | MedX | 29847 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | LX-B50L | F5020078 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified temperatures at 129°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230533618421339/6083230533618421339_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230533618421339/6083230533618421339_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29848 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | TSX2330FZ | 1161322301200520 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified temperatures at -20°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230533619577714/6083230533619577714_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230533619577714/6083230533619577714_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph | Biobase | 29849 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | BKQ-B50 II | BKQ50II2007541 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified temperatures at 134°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230533619830276/6083230533619830276_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230533619830276/6083230533619830276_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph | Condura | 29846 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | CTD3113MNI | 18956716056 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified temperatures at 2°C, 8°C, and -20°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473612840706/6083230473612840706_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473612840706/6083230473612840706_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph | Esco | 29842 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | IFA-32-8 | 2020-T05163 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified temperatures at 56°C and 60°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473615204543/6083230473615204543_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473615204543/6083230473615204543_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph | Esco | 29845 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | IFA-32-8 | 2020-T05006 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified temperatures at 56°C and 60°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473615936615/6083230473615936615_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473615936615/6083230473615936615_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph | Condura | 29843 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | CBC283-NL | 688103101W | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified temperatures at 2°C and 8°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473619248883/6083230473619248883_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473619248883/6083230473619248883_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph | Condura | 29844 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | CBC283-NL | 688103125W | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified temperatures at 2°C and 8°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473619944648/6083230473619944648_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230473619944648/6083230473619944648_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29838 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | Medifuge | 42358139 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 1000rpm and 5000rpm Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413612143504/6083230413612143504_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413612143504/6083230413612143504_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29839 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MySpin 12 | HSH12116 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 500rpm and 12500rpm Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413614715880/6083230413614715880_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413614715880/6083230413614715880_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29837 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MySpin 12 | HSH12119 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 500rpm and 12500rpm Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413614955320/6083230413614955320_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413614955320/6083230413614955320_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29841 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MySpin 12 | HSH12117 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 500rpm and 12500rpm Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413617248010/6083230413617248010_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413617248010/6083230413617248010_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph | LW Scientific | 29840 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | EZ Swing | 1101273 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 1000rpm and 5000rpm Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413617307879/6083230413617307879_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230413617307879/6083230413617307879_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph | Iris | 29834 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | StatSpin Express 2 | 731M50106146 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 1000rpm and 3000rpm Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353611937211/6083230353611937211_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353611937211/6083230353611937211_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29836 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | Vortex Mixer | K3CT82010445 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 200rpm and 3000rpm Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353613412316/6083230353613412316_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353613412316/6083230353613412316_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29835 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | Thermal Mixer | 1014319110746 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 500rpm and 2000rpm Verified temperatures at 55°C and 60°C Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353613923563/6083230353613923563_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353613923563/6083230353613923563_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29832 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | Vortex Mixer | K3CT82010446 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 200rpm and 3000rpm Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353617580995/6083230353617580995_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353617580995/6083230353617580995_signature_31.png
| 2024-11-25 | 2024-11-25 | Biopath Clinical Diagnostics -ERod | Garden Heights Condominium, 269 E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rmlao@biopath.ph |  | 29833 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | Vortex Mixer | K3CT82010021 | None | None | Performed surface cleaning Performed preventive maintenance and calibration Verified speeds at 200rpm and 3000rpm Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353619734678/6083230353619734678_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6083230353619734678/6083230353619734678_signature_31.png
| 2024-11-25 | 2024-11-25 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY, CEBU | 7 | Cebu | LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29831 | DOH | mendozarey42@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4449 ASSAY: Xpert HCV VL, COUNT: 54 ASSAY: Xpert HIV-1 Viral Load, COUNT: 454 |  | GX-IV | 838510 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083217180529637608/6083217180529637608_signature_28.png | Rey C. Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6083217180529637608/6083217180529637608_signature_31.png
| 2024-11-25 | 2024-11-25 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY, CEBU | 7 | Cebu | LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29830 | DOH | mendozarey42@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4449 ASSAY: Xpert HCV VL, COUNT: 54 ASSAY: Xpert HIV-1 Viral Load, COUNT: 454 |  | GX-IV | 838508 | #NAME? | #NAME? | #NAME? | #NAME? | MODULE B1 AND B3 PENDING |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083214520521566884/6083214520521566884_signature_28.png | Rey C. Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6083214520521566884/6083214520521566884_signature_31.png
| 2024-11-25 | 2024-11-25 | Ilocos Training and Regional Medical Center | San Fernando City, La Union | 1 | Pangasinan | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29829 | MACARE-HBDC | itrmc.mbl@gmail.com | 25 |  | 2024-11-25 | GXIV-4-D-10C | 120003504 | For Installation | #NAME? | Installed GeneXpert 10C System | Unit is Fully Operational |  | GeneXpert System | DR 12331 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6083158770289407449/6083158770289407449_signature_28.png | Dominga Patacsil | https://www.jotform.com/uploads/macarewecare/222927978372067/6083158770289407449/6083158770289407449_signature_31.png
| 2024-11-22 | 2024-07-05 | St. Luke's Medical Center -Global City | Rizal Dr., BGC, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | Pull out | CEPHEID | Noel Floresca |  | jmdoctor@macare-medicals.com | biomedreports@st.lukes.com.ph |  | 29828 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 |  |  | GX-IV-D-6C | 808411 | For pull out | Pull out GeneXpert system | Pull out the GeneXpert system -GeneXpert unit -UPS -HP Laptop | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6080715216269296482/6080715216269296482_signature_28.png | Ange Sapla | https://www.jotform.com/uploads/macarewecare/222927978372067/6080715216269296482/6080715216269296482_signature_31.png
| 2024-11-22 | 2024-11-22 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29827 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 19 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 277 ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 |  | GX-IV-D-6C | 110005265 | No test result generated in module B2 and B3. | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Checked modules Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6080703115056055396/6080703115056055396_signature_28.png | Sarah Almeria | https://www.jotform.com/uploads/macarewecare/222927978372067/6080703115056055396/6080703115056055396_signature_31.png
| 2024-11-22 | 2024-11-22 | CEBU DOCTORS UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29826 | COMMERCIAL | montezakars2199@gmail.com | 22 |  |  | MCS+ | 14F014 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6080612670523716059/6080612670523716059_signature_28.png | Karen Kay Monteza | https://www.jotform.com/uploads/macarewecare/222927978372067/6080612670523716059/6080612670523716059_signature_31.png
| 2024-11-22 | 2024-11-22 | San Fernando Rural Health Unit | Zone 1, San Fernando | 5 | Camarines Sur | Laboratory Gx | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29825 | DOH | rhusanfernando.2018@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 838 |  | GX4-6C | 838496 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6080580760114983865/6080580760114983865_signature_28.png | Wilda t. Bona | https://www.jotform.com/uploads/macarewecare/222927978372067/6080580760114983865/6080580760114983865_signature_31.png
| 2024-11-21 | 2024-11-21 | Polomolok Rural Health Unit | Poblacion, Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29824 | DOH | jmmirasolrmt@gmail.com | 26.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1651 |  | GxIV-6C | 838265 | Module 2 and 4 Waiting for replacement module | Preventive Maintenance | Perform Genexpert cleaning and maintenance Checking of errors Data collection | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079865949495482465/6079865949495482465_signature_28.png | June Michele Mirasol | https://www.jotform.com/uploads/macarewecare/222927978372067/6079865949495482465/6079865949495482465_signature_31.png
| 2024-11-21 | 2024-11-21 | Biopath QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29823 | COMMERCIAL | ainocentes@biopath.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 19 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 25 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 4 |  | GX16 D | 839306 | Due for maintenance A3 disabled due to temp error. C2 disabled failed xpertcheck | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | 14 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079847398624275475/6079847398624275475_signature_28.png | Amanda inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6079847398624275475/6079847398624275475_signature_31.png
| 2024-11-21 | 2024-11-21 | Calabanga Rural Health Unit | San Pablo St. Calabanga, Cam.Sur | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29822 | DOH | martyarc11@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 787 |  | GX4-6C | 838398 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079835291132130043/6079835291132130043_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/6079835291132130043/6079835291132130043_signature_31.png
| 2024-11-21 | 2024-11-21 | Baganga Rural Health Unit | Poblacion Baganga | 11 | Davao Oriental | TB-DOTS | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29821 | DOH | rhubaganga@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 505 | 2024-09-04 | GX-IV 6C 4D | 838395 | For Preventive Maintenance No power at System Unit For XpertCheck Calibration | Reset Connections and change CMOS battery Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck on four modules | Perform System Unit Troubleshooting Perform Preventive Maintenance and XpertCheck Calibration | System unit working All Modules passed the Calibration All Modules are operational |  | XPERTCHECK CE - 5 | DR 11997 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079772558515704230/6079772558515704230_signature_28.png | Liza N. Narciso,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6079772558515704230/6079772558515704230_signature_31.png
| 2024-11-21 | 2024-11-21 | Pasig Treatment Hub | Dr. Garcia St. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29820 | MACARE-HBDC | pathmedtech@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 799 | 2024-11-14 | GX-XVI-D-10C | 821648 | Module A1 for replacement | Replaced module A1 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module A1 | Unit is operational. Modules A1 and A3 operational. |  | Old: A1/733228 New: A1/210067862 | SI# 99098 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079748305468358483/6079748305468358483_signature_28.png | Jhozza Rei Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/6079748305468358483/6079748305468358483_signature_31.png
| 2024-11-21 | 2024-11-21 | PaTH CHAMP | Caruncho Complex Pasig City Mega Market, 1895, 16 Caruncho Ave, San Nicholas, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | pathmedtech@gmail.com |  | 29819 | MACARE-HBDC | dansumulong.pasigcho@gmail.com | 25 |  |  | GXIV 10C D | 110025127 | Unit due for installation | Geberated report | Setup and installed GX and accessories | Unut ia operational. All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079747995124546377/6079747995124546377_signature_28.png | DAN NATHANIEL SUMULONG | https://www.jotform.com/uploads/macarewecare/222927978372067/6079747995124546377/6079747995124546377_signature_31.png
| 2024-11-21 | 2024-11-21 | Pasig Treatment Hub | Dr. Garcia St. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29818 | MACARE-HBDC | pathmedtech@gmail.com | 22 |  |  | GX-IV-D-10C | 110025128 | For installation | Installed the GeneXpert system | Installed the GeneXpert system: -GeneXpert machine -workstation -printer -barcode scanner -UPS | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079747815469922122/6079747815469922122_signature_28.png | Jhozza Rei Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/6079747815469922122/6079747815469922122_signature_31.png
| 2024-11-21 | 2024-11-21 | Pili Rural Health Unit | San Isidro, Pili | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29817 | DOH | mariechristineorcine@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1117 |  | GX4-6C | 838408 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079708554521463316/6079708554521463316_signature_28.png | Marie Christine Nobleza | https://www.jotform.com/uploads/macarewecare/222927978372067/6079708554521463316/6079708554521463316_signature_31.png
| 2024-11-21 | 2024-11-21 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 29816 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 31 |  |  | GX4 L | 807155 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational | To follow EST |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079675428326397823/6079675428326397823_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/6079675428326397823/6079675428326397823_signature_31.png
| 2024-11-21 | 2024-11-21 | St lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiolgy | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 29815 | COMMERCIAL | megapo@stlukes.com.ph | 27 | ASSAY: Xpert CT/NG, COUNT: 259 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3103 |  | GX4 D | 839234 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Back up gxx files Test run unit | Performed maintenance | B1,B2 operational B3,B4 disabled | To follow EST |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079637721714135556/6079637721714135556_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6079637721714135556/6079637721714135556_signature_31.png
| 2024-11-21 | 2024-11-20 | Phililpine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29814 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | CRC bios. | Troubleshooting | Performed onsite troubleshooting. Resets internal connections. Checked CPU board. | Unit is not operational. For replacement of CPU board. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6079560770118389563/6079560770118389563_signature_28.png | Lalaine ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6079560770118389563/6079560770118389563_signature_31.png
| 2024-11-20 | 2024-11-20 | Tigaon Municipal Health Office | Vinagre | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29813 | DOH | jingimperial@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 855 |  | GX4-6C | 838558 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6078987659049859217/6078987659049859217_signature_28.png | Maria Jesusa P. Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6078987659049859217/6078987659049859217_signature_31.png
| 2024-11-20 | 2024-11-20 | San Jose Municipal Health Office | Poblacion, San Jose | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29812 | DOH | jhuneric2008@yahoo.com | 22 |  |  | GX4-6C | 838389 | Gx for Preventive Maintenance B3/739734 - invalid UPS no power | Run cartridge with no specimen and reagent for B3739734 - 2026 error, H/W failed Collected GXX file Tested/Check the battery dependability of UPS Replaced fuse Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance Performed Troubleshooting for module B3 and UPS | Unit is Operational UPS non-operational | UPS request for battery replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6078926699045486216/6078926699045486216_signature_28.png | Rosmar V. Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6078926699045486216/6078926699045486216_signature_31.png
| 2024-11-20 | 2024-11-19 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29811 | COMMERCIAL | jenmichelleentica@gmail.com | 22 |  |  | MCS+ | 09C125 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Run simulation Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6078778737451215953/6078778737451215953_signature_28.png | Jen Michelle T. Entica | https://www.jotform.com/uploads/macarewecare/222927978372067/6078778737451215953/6078778737451215953_signature_31.png
| 2024-11-20 | 2024-11-19 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29810 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 72 ASSAY: Xpert CT/NG, COUNT: 314 ASSAY: Xpert HBV VL, COUNT: 11 ASSAY: Xpert HPV, COUNT: 380 ASSAY: Xpert MTB/RIF Ultra, COUNT: 908 ASSAY: Xpert HCV VL, COUNT: 19 ASSAY: Xpert HIV-1 Viral Load, COUNT: 72 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6078778737454803303/6078778737454803303_signature_28.png | Jordan T. Purpura | https://www.jotform.com/uploads/macarewecare/222927978372067/6078778737454803303/6078778737454803303_signature_31.png
| 2024-11-20 | 2024-11-19 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29809 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 13 ASSAY: Xpert CT/NG, COUNT: 69 ASSAY: Xpert HPV, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 153 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 35 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A3/866793 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6078778727459821481/6078778727459821481_signature_28.png | Jordan T. Purpura | https://www.jotform.com/uploads/macarewecare/222927978372067/6078778727459821481/6078778727459821481_signature_31.png
| 2024-11-19 | 2024-11-19 | Kalinga Health Hub Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29807 | CONSORTIUM | pearl.balderama@innovationsch.org | 25 |  |  | GXIV | 110008420 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6078043579322834864/6078043579322834864_signature_28.png | Pearl Balderama | https://www.jotform.com/uploads/macarewecare/222927978372067/6078043579322834864/6078043579322834864_signature_31.png
| 2024-11-19 | 2024-11-19 | Kalinga Health Hub Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29808 | CONSORTIUM | pearl.balderama@innovationsch.org | 25 |  |  | GXIV | 110008415 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6078043579323848262/6078043579323848262_signature_28.png | Pearl Balderama | https://www.jotform.com/uploads/macarewecare/222927978372067/6078043579323848262/6078043579323848262_signature_31.png
| 2024-11-19 | 2024-11-19 | Cardinal Santos Medical Center | Wilson St., Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | TROUBLESHOOTING | HAEMONETICS | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29806 | COMMERCIAL | ftejada@csmc.ph | 25 |  |  | MCS+ | 08M461 | No power Fuse blown out | Replacrd and installed new fuse Test run unit | Perform troubleshooting | Unit is operational | With Sir Noel Floresca |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6078008661517254840/6078008661517254840_signature_28.png | Fidel Tejada | https://www.jotform.com/uploads/macarewecare/222927978372067/6078008661517254840/6078008661517254840_signature_31.png
| 2024-11-19 | 2024-11-19 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | synch28@yahoo.com |  | 29805 | MACARE-HBDC | msff-tondo-labtech@paris.msf.org | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5515 | 2024-11-13 | GX16-10C | 110016343 | Unit for preventive maintenance. For module replacement of module D4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module D4 (633865) in replacement to defective D4 (210075616). | DR#12316 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077962348116387523/6077962348116387523_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6077962348116387523/6077962348116387523_signature_31.png
| 2024-11-19 | 2024-11-18 | Makati Medical Center | Amorsolo st , Makati | NCR | NCR 4th - Makati | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | bryan.pitero@makatimed.net.ph |  | 29804 | COMMERCIAL | annark3ene@yahoo.com | 26 |  |  | MCS + | 97H330 | As per end user they encounter bowl breakage | Cleaned centrifuge Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077892593616092218/6077892593616092218_signature_28.png | Mark Fil | https://www.jotform.com/uploads/macarewecare/222927978372067/6077892593616092218/6077892593616092218_signature_31.png
| 2024-11-18 | 2024-11-15 | St. Elizabeth Hospital Inc/Meli Linx | General Santos City | 12 | South Cotabato | Tb-Lab | NO | INSTALLATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | sofia.maclang@medilinx.com.ph |  | 29802 | COMMERCIAL | therese.cariaga@medilinx.com.ph | 24.3 |  |  | GxIV-6C | 831313 | CPU problem | Software Installation Preventive maintenance | GxDx software installation Assays installation | 1 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077455889492196186/6077455889492196186_signature_28.png | Therese Christia Diana Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6077455889492196186/6077455889492196186_signature_31.png
| 2024-11-18 | 2024-11-15 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29803 | DOH | renlizros11@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1450 |  | GxIV-6C | 838570 | Monitor problem | Troubleshooting Preventive maintenance | Monitor problem Troubleshooting Preventive maintenance Problem reporting Data collection | 4 operational modules Monitor problem request for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077455889493337614/6077455889493337614_signature_28.png | Renee Liz Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6077455889493337614/6077455889493337614_signature_31.png
| 2024-11-18 | 2024-11-18 | Chong Hua Hospital | Cebu City | 7 | Cebu | Bloodbank | YES | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29801 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ haemonitics | 11J219 | BIOS CRC Failed and Pneumatic assembly motor did not stop or cut off. | For replacement of CPU and pneumatic assembly. | #NAME? | DEFECTIVE UNIT. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077256554918124428/6077256554918124428_signature_28.png | CHONG HUA BLOODBANK | https://www.jotform.com/uploads/macarewecare/222927978372067/6077256554918124428/6077256554918124428_signature_31.png
| 2024-11-18 | 2024-11-15 | Cebu South Medical Center | Talisay City | 7 | Cebu | TB Laboratory | NO | TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29800 | DOH | mkrytelshean@gmail.com | 23 |  | 2024-11-15 | GX 4 | 838559 | Error for module B3 during selected for running. | Save System Logfile Save Installation Qualification | Inform end user to disconnect the Internet. | All module's are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077126015138409113/6077126015138409113_signature_28.png | Krystel Shean Miranda | https://www.jotform.com/uploads/macarewecare/222927978372067/6077126015138409113/6077126015138409113_signature_31.png
| 2024-11-18 | 2024-11-14 | San Lazaro Hospital | Quiricada St, Sta. Cruz, Manila | NCR | NCR 1st - Manila | SACCL | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | piaciudadano1331@gmail.com |  | 29798 | MACARE-HBDC | ciudadanosophiapatrice@gmail.com | 28 | ASSAY: Xpert HBV VL, COUNT: 397 ASSAY: Xpert HCV VL, COUNT: 28 ASSAY: Xpert HIV-1 Viral Load, COUNT: 219 |  | GXIV-L-6C | 821257 | Module 3 with H/W Error. | Reliabilty Test. | Cleaned Icore Optics. Performed Reliability Test in Module 3. Performed Self-test. Generate Reports. | 4 Modules are Available. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077086193613841155/6077086193613841155_signature_28.png | Sophia Patrice Ciudadano | https://www.jotform.com/uploads/macarewecare/222927978372067/6077086193613841155/6077086193613841155_signature_31.png
| 2024-11-18 | 2024-11-15 | UST Hospital | Dapitan St, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29799 | CONSORTIUM | ustimp@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1021 |  | GXIV-D-6C | 110000711 | Module 1 with Errors. | Preventive Maintenance. Self-test. | Cleaned Cartridge Bay, Plunger Rod, Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | Module 1 with 5007 Error. (For Observation after Cleaning of Optics.) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077086193618536242/6077086193618536242_signature_28.png | Sherleen Calaunan | https://www.jotform.com/uploads/macarewecare/222927978372067/6077086193618536242/6077086193618536242_signature_31.png
| 2024-11-18 | 2024-11-15 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | Stephenbryan.1621@gmail.com |  | 29796 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 24 | ASSAY: Others, COUNT: 2 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 173 ASSAY: Xpert SARS-CoV-2, COUNT: 1456 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-XVI-D-6C | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules A3, A4, B1, B2 C1 and C4 are operational. Modules A1/860941, A2/860940, B3/702534, B4/860923, C2/668748, C3/662448, D1/860832, D2/620837, D3/860742, D4/661946 not operational, for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077059813617251838/6077059813617251838_signature_28.png | Justine Lorenz Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/6077059813617251838/6077059813617251838_signature_31.png
| 2024-11-18 | 2024-11-15 | Rizal Medical Center | Pasig Blvd. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | stephenbryan.1621@gmail.com |  | 29797 | MACARE-HBDC | rmcbloodbank@gmail.com | 24 |  |  | MCS+ | 15H045G | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077059813617729626/6077059813617729626_signature_28.png | Marilyn Ang | https://www.jotform.com/uploads/macarewecare/222927978372067/6077059813617729626/6077059813617729626_signature_31.png
| 2024-11-18 | 2024-11-15 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29795 | COMMERCIAL | lcpmicrolab@gmail.com | 29 |  | 2024-11-15 | GX4 L | 804205 | For xpertcheck For relocation microbiology to molecular lab | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck and relocation | D1,D2 failed xpertcheck D3,D4 retry lack of xpertcheck kits |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077014193616674358/6077014193616674358_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6077014193616674358/6077014193616674358_signature_31.png
| 2024-11-18 | 2024-11-15 | Best Diagnostic Corporation | 94 Matalino st., Diliman QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | bestdx@yahoo.com |  | 29794 | COMMERCIAL | laboratory@bestdiagnostic.com.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 75 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 12 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GX4 D | 110008529 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Back up gxx files Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6077013643611824774/6077013643611824774_signature_28.png | Best Diagnostic Corporation | https://www.jotform.com/uploads/macarewecare/222927978372067/6077013643611824774/6077013643611824774_signature_31.png
| 2024-11-17 | 2024-11-11 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29793 | DOH | chaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1775 |  | GX4 | 838283 | Due for monthly maintenance For transfer to Argao Rural Health Unit | Conduct preventive maintenance for Icore, Plunger, ultrasonic horn Clean up cartridge bay and filter | Conduct preventive maintenance for Icore, Plunger, ultrasonic horn Clean up cartridge bay and filter | All units of module are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6076031879314540310/6076031879314540310_signature_28.png | Lucille Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6076031879314540310/6076031879314540310_signature_31.png
| 2024-11-15 | 2024-11-15 | St Lukes Extension Clinic | J Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 29792 | MACARE-HBDC | laboratory@slec.com | 25 |  | 2024-11-15 | Infinity 48s | 809643 | Unit for XpertCheck | None | Performed maintenance cleaning Performed XpertCheck procedure Replaced IAI batteries | Unit is operational, failed modules for reporting: 210050828/A4 720535/C4 729869/D2 618607/D4 896445/E3 608345/H1 696244/H2 654533/H3 |  | XpertCheck kit - 8 boxes IAI Battery - 5 pcs | 13232 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074729900021553566/6074729900021553566_signature_28.png | MA. SARAH R. SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/6074729900021553566/6074729900021553566_signature_31.png
| 2024-11-15 | 2024-11-15 | Saguiaran Rural Health Unit | Saguiaran, Marawe Lanao Del Sur | 10 | Lanao del Sur | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29791 | DOH | aiylacalandada01@gmail.com | 24 |  |  | GX4 | 838298 | Due for Annual Calibration of the machine. | Performed calibration of the genexpert machine in all modules Conduct preventive maintenance for Icore, Plunger, ultrasonic horn Clean up cartridge bay and filter | Performed calibration of the genexpert machine in all modules Conduct preventive maintenance for Icore, Plunger, ultrasonic horn Clean up cartridge bay and filter | All modules passed during calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074673799318316817/6074673799318316817_signature_28.png | Sahara Muti, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6074673799318316817/6074673799318316817_signature_31.png
| 2024-11-15 | 2024-11-15 | Marawe City Health Office | Marawe City, Lanao del Sur | 10 | Lanao del Sur | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29790 | DOH | faridahmamao@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 306 |  | GX4 | 838293 | Due for Annual Calibration of the machine. | Performed calibration of the genexpert machine in all modules Conduct preventive maintenance for Icore, Plunger, ultrasonic horn Clean up cartridge bay and filter | Performed calibration of the genexpert machine in all modules Conduct preventive maintenance for Icore, Plunger, ultrasonic horn Clean up cartridge bay and filter | All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074669082109742472/6074669082109742472_signature_28.png | Faridah BUsara,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6074669082109742472/6074669082109742472_signature_31.png
| 2024-11-15 | 2024-11-15 | Marawe City Health Office | Marawe City, Lanao del Sur | 10 | Lanao del Sur | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29789 | DOH | faridahmamao@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1098 |  | GX4 | 838260 | Due for Annual Calibration of the machine | Performed calibration of the genexpert machine in all modules Conduct preventive maintenance for Icore, Plunger, ultrasonic horn Clean up cartridge bay and filter | Performed calibration of the genexpert machine in all modules Conduct preventive maintenance for Icore, Plunger, ultrasonic horn Clean up cartridge bay and filter | All modules passed except B4. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074666882108294649/6074666882108294649_signature_28.png | Faridah BUsara,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6074666882108294649/6074666882108294649_signature_31.png
| 2024-11-15 | 2024-11-15 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29788 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 75 ASSAY: Xpert CT/NG, COUNT: 212 ASSAY: Xpert HBV VL, COUNT: 203 ASSAY: Xpert HPV, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1072 ASSAY: Xpert HCV VL, COUNT: 157 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GXIV | 110008524 | MONTHLY PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074623334939012339/6074623334939012339_signature_28.png | Jennifer Marie Lim, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6074623334939012339/6074623334939012339_signature_31.png
| 2024-11-15 | 2024-11-15 | Manila Doctors Hospital | United Nations Ave., Ermita, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | gpdmagnaye30@gmail.com |  | 29787 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 75 ASSAY: Xpert HBV VL, COUNT: 88 ASSAY: Xpert HPV, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1236 ASSAY: Xpert HCV VL, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GXIV 10C D | 110001648 | Unit due for maintenance | Performed preventive maintenance Performed rewlibility test Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Run reliability test | Unit is operational all modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074592359216368443/6074592359216368443_signature_28.png | Alyssa Mae Jolo | https://www.jotform.com/uploads/macarewecare/222927978372067/6074592359216368443/6074592359216368443_signature_31.png
| 2024-11-15 | 2024-11-15 | Cebu City Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29786 | DOH | ccmclaboratory01@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1865 |  | 6C GXIV-4D | 838561 | Desktop unit resets date/time every power on. | #NAME? | None | Unit is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074583532131136552/6074583532131136552_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/6074583532131136552/6074583532131136552_signature_31.png
| 2024-11-15 | 2024-11-15 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29785 | CONSORTIUM | lpcomendador@uerm.edu.ph | 25 |  |  | GXIV | 847578 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074538889296941726/6074538889296941726_signature_28.png | Osang Comendador | https://www.jotform.com/uploads/macarewecare/222927978372067/6074538889296941726/6074538889296941726_signature_31.png
| 2024-11-15 | 2024-11-15 | Naga City General Hospital | Leon SA Aureus St. Balatas | 5 | Camarines Sur | Laboratory | YES | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | ncgh.clinicallaboratory@gmail.com |  | 29784 | MACARE-HBDC | caizelleinfeliz97@gmail.com | 25 |  |  | GX4-10C | 120003510 | GX4-10C for Installation | Installed Gx4-10C 120003510 Installed UPS AT1000 Installed CPU CPH0424F50953 Installed Monitor Installed Printer HP Deskjet 2337 Performed Preventive Maintenance Training | GX Installation completed | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074522388614047881/6074522388614047881_signature_28.png | Infeliz Caizelle | https://www.jotform.com/uploads/macarewecare/222927978372067/6074522388614047881/6074522388614047881_signature_31.png
| 2024-11-15 | 2024-11-15 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | dyrosemarie71@gmail.com |  | 29783 | DOH | dy_rosemarie71@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1373 |  | GX-IV 6C 4D | 838401 | Monitor Display Problem | Reset cable connections, Change power chord, change vga cables | Perform Monitor troubleshooting | Monitor display still broken, error reported to Cepheid for further Assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074509664615781810/6074509664615781810_signature_28.png | Rosemarie Dy,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6074509664615781810/6074509664615781810_signature_31.png
| 2024-11-15 | 2024-11-15 | Midsalip Rural health unit | Midsalip Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29782 | DOH | genaliner@rocketmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 410 | 2024-11-15 | GX 4(6c) | 837664 | None | Preventive Maintenence, checked error logs and room/ module temperatures. | Preventive Maintenence, checked error logs and room/ module temperatures. | All modules are available and ready to use | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074502723011172557/6074502723011172557_signature_28.png | Genaline Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/6074502723011172557/6074502723011172557_signature_31.png
| 2024-11-15 | 2024-11-15 | Central Luzon Doctor's Hospital | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29781 | COMMERCIAL | pjlsantos@cldh.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 ASSAY: Xpert SARS-CoV-2, COUNT: 16 |  | GXII-D 6C | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074471438613508265/6074471438613508265_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6074471438613508265/6074471438613508265_signature_31.png
| 2024-11-15 | 2024-11-15 | Central Luzon Doctor's Hospital | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29780 | COMMERCIAL | pjlsantos@cldh.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 ASSAY: Xpert SARS-CoV-2, COUNT: 16 |  | GXIV-D 6C | 839232 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074470858619757553/6074470858619757553_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6074470858619757553/6074470858619757553_signature_31.png
| 2024-11-15 | 2024-11-14 | The Medical City | Ortigas Ave., Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29778 | COMMERCIAL | moleculardx@themedicalcity.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 56 ASSAY: Xpert CT/NG, COUNT: 67 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 599 ASSAY: Xpert HPV, COUNT: 426 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5657 ASSAY: Xpert MTB-XDR, COUNT: 1 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 93 ASSAY: Xpert HIV-1 Viral Load, COUNT: 753 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 13220 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit os operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074403127591085354/6074403127591085354_signature_28.png | Stephanie Obida | https://www.jotform.com/uploads/macarewecare/222927978372067/6074403127591085354/6074403127591085354_signature_31.png
| 2024-11-15 | 2024-11-14 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29777 | COMMERCIAL | mscaina@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | For calibration | Calibrated the unit | Performed maintenance Calibrated the unit Checked rotor pumps Adjusted calibration parameters Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074403127596961194/6074403127596961194_signature_28.png | Maria Zoraida Caina | https://www.jotform.com/uploads/macarewecare/222927978372067/6074403127596961194/6074403127596961194_signature_31.png
| 2024-11-15 | 2024-11-14 | The Medical City | Ortigas Ave, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29779 | COMMERCIAL | bloodbank@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6074403127598842689/6074403127598842689_signature_28.png | Armin Francis Betan | https://www.jotform.com/uploads/macarewecare/222927978372067/6074403127598842689/6074403127598842689_signature_31.png
| 2024-11-14 | 2024-11-14 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | Check up | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29776 | MACARE-HBDC | bloodcollection.tower@gmail.com | 25 |  |  | MCS+ | 01J038 | Blood spillage on centrifuge | - Unit has ongoing procedure before I performed the check up (procedure success) - cleaned centrifuge's sensors | #NAME? | Unit is operational No issues found |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6073636580729549860/6073636580729549860_signature_28.png | Ardel brecinio | https://www.jotform.com/uploads/macarewecare/222927978372067/6073636580729549860/6073636580729549860_signature_31.png
| 2024-11-14 | 2024-11-14 | Ramon magsaysay rural health unit | Ramon magsaysay, Zamboanga del 2 | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 29775 | DOH | ramonmagsaysay@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 | 2024-11-14 | GX 4 (6c) | 838443 | None | Preventive Maintenance. Checked module errors and room temperature. | Preventive Maintenance. Checked module errors and room temperature. | All modules are available and ready to use. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6073598335182018990/6073598335182018990_signature_28.png | Gina Borinaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6073598335182018990/6073598335182018990_signature_31.png
| 2024-11-13 | 2024-11-13 | Pio Duran Rural Health Unit | Caratagan, Pio Duran | 5 | Albay | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29774 | DOH | mitzijc@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 592 |  | GX4-6C | 838264 | Gx for Preventive Maintenance Date and Time resets when restarting PC | Replace CMOS/BIOS battery of CPU for Date and Time Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance Performed CPU troubleshooting | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072907923619999150/6072907923619999150_signature_28.png | Mitzi D. Cedeño | https://www.jotform.com/uploads/macarewecare/222927978372067/6072907923619999150/6072907923619999150_signature_31.png
| 2024-11-13 | 2024-11-12 | Pontevedra Rural Health Unit | Pontevedra, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29773 | DOH | pontevedra.laboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: 744 | 2024-10-30 | GX-4 6C | 838491 | Defective UPS | Preventive Maintenance | #NAME? | 4 modules operational |  | UPS | DR 12288 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072860031915338348/6072860031915338348_signature_28.png | Love Joy Sulleza | https://www.jotform.com/uploads/macarewecare/222927978372067/6072860031915338348/6072860031915338348_signature_31.png
| 2024-11-13 | 2024-11-12 | East Avenue Medical Center | East Ave. QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29772 | MACARE-HBDC | eamcmolecularlab@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 365 ASSAY: Xpert HPV, COUNT: 245 ASSAY: Xpert SARS-CoV-2, COUNT: 1315 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 145 |  | GX4 10C | 110016781 | B3 disabled due to temp error 2014 Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B3 disabled B1,B2,B4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072855293613690473/6072855293613690473_signature_28.png | EAMC Molecular Lab | https://www.jotform.com/uploads/macarewecare/222927978372067/6072855293613690473/6072855293613690473_signature_31.png
| 2024-11-13 | 2024-11-12 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Tbdots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29771 | MACARE-HBDC | roseannparnes@yahoo.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3067 |  | GX4 10C | 110014085 | Due for maintenance and relocation | Relocate unit Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and relocation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072854483612993390/6072854483612993390_signature_28.png | Rose Ann Parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/6072854483612993390/6072854483612993390_signature_31.png
| 2024-11-13 | 2024-11-12 | Philippine Red Cross QC | Kalayaan Ave. QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29770 | MACARE-HBDC | villamarevangeline103@gmail.com | 28 |  |  | MCS + | 02A227 | As per end user they encounter red valve error | Performed valve test Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running Fixed red valve | Performed troubleshooting | Operational | Workwith Gian Magnaye |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072854013613425498/6072854013613425498_signature_28.png | Evangeline P. Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/6072854013613425498/6072854013613425498_signature_31.png
| 2024-11-13 | 2024-11-13 | Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29769 | DOH | miniforestlaboratory@gmail.com | 23.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 387 |  | GX-IV 6C 4D | 838507 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat, and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072849361619062159/6072849361619062159_signature_28.png | Joy L. Nio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6072849361619062159/6072849361619062159_signature_31.png
| 2024-11-13 | 2024-11-13 | Baybay City Health Office | Baybay City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29768 | DOH | sajecopooh20@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1145 | 2024-11-13 | GXIV | 83855 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072801954533272347/6072801954533272347_signature_28.png | Sarah Jean Pasana | https://www.jotform.com/uploads/macarewecare/222927978372067/6072801954533272347/6072801954533272347_signature_31.png
| 2024-11-13 | 2024-11-12 | AFPMC-AFRIMS 6C | V. Luna, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | gpdmagnaye30@gmail.com |  | 29767 | MACARE-HBDC | mtvalderama@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 174 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GXIV 6C -D | 805294 | Unit due for maintenance | Performed Preventive maintenance Performed Reliability Test Generated Report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Installed MTB Rif - Ultra ADF | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072794193612358651/6072794193612358651_signature_28.png | maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/6072794193612358651/6072794193612358651_signature_31.png
| 2024-11-13 | 2024-11-13 | Dinas Rural Health Unit | Dinas Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 29766 | DOH | bernantereynacres@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 722 | 2024-10-21 | GX 4 (6c) | 838533 | Module 4 h/w failed | Replaced module 862260 with Module 789674. Conducted Preventive maintenence. | Replaced module 862260 with Module 789674. Conducted Preventive maintenence. | All modules are available | None | 900-0491RS | 12260 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072791622778660584/6072791622778660584_signature_28.png | Reyna Cres Bernante | https://www.jotform.com/uploads/macarewecare/222927978372067/6072791622778660584/6072791622778660584_signature_31.png
| 2024-11-13 | 2024-11-12 | Rosario Super Health Center | Dr. Sixto Antonio Avenue, Rosario, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29765 | MACARE-HBDC | consalaneljaine@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 836 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072725733618864823/6072725733618864823_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6072725733618864823/6072725733618864823_signature_31.png
| 2024-11-13 | 2024-11-13 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | Cebu City | 7 | Cebu | BLOOD BANK | NO | MAINTENANCE ON-SITE ASSISTANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29764 | COMMERCIAL | montezakars2199@gmail.com | 23.2 |  |  | MCS+ | 14F014 | REPORTED CENTRIFUGE ISSUE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072696552913243106/6072696552913243106_signature_28.png | Karen Kay Monteza | https://www.jotform.com/uploads/macarewecare/222927978372067/6072696552913243106/6072696552913243106_signature_31.png
| 2024-11-13 | 2024-10-30 | Dasmariñas City Health Office II | DBB-E, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29762 | DOH | hazelmilan0777@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1038 |  | GXIV-D-6C | 838535 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. Checked System's Functionality. | All Modules are Good. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072681953618938377/6072681953618938377_signature_28.png | Hazel Tabucol | https://www.jotform.com/uploads/macarewecare/222927978372067/6072681953618938377/6072681953618938377_signature_31.png
| 2024-11-13 | 2024-11-07 | General Trias Social Hygiene Clinic | Hospital Road, General Trias, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29763 | MACARE-HBDC | chogentri.shc@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 817 |  | GXIV-D-10C | 110016785 | Unit is due for Maintenance. | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System's Functionality. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072681953619076735/6072681953619076735_signature_28.png | John Joseph Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6072681953619076735/6072681953619076735_signature_31.png
| 2024-11-12 | 2024-11-12 | AFPMC - AFRIMS | V. Luna, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29761 | MACARE-HBDC | mtvalderama@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 45 ASSAY: Xpert MTB/RIF Ultra, COUNT: 648 ASSAY: Xpert SARS-CoV-2, COUNT: 961 ASSAY: Xpert HCV VL, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GXIV 10C - D | 845462 | Unit due for maintenance | Performed Preventive maintenance Performed Reliability Test Generated report | checked and cleaned cartridge bays, plunger rod, icore slot and filter | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072006965303798476/6072006965303798476_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/6072006965303798476/6072006965303798476_signature_31.png
| 2024-11-12 | 2024-11-12 | Philippine General Hospital | Taft Ave Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29760 | MACARE-HBDC | eomercado@up.edu.ph | 25 |  |  | GX16 | 844975 | Unit for preventive maintenance. For generating of data for technical team. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6072006910613302808/6072006910613302808_signature_28.png | Erma Grace Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/6072006910613302808/6072006910613302808_signature_31.png
| 2024-11-12 | 2024-11-12 | Pascual Laboratories, Inc. | Km 31, Mac Arthur Highway, Balagtas, Bulacan, 3016 Philippines | 3 | Bulacan | QC Laboratory | YES | CALIBRATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29759 | COMMERCIAL | jnvalerio@pascuallab.com | 24 |  |  | AXIO Scope.A1 | 3321001650 | Due for Preventive Maintenance and Calibration | Aligned Kohler | Performed Preventive Maintence and Calibration Cleaned Surfaces, Eyepieces, Objectives, Condenser, Diaphragm, and Camera adapter Aligned Kohler Performed Functionality Test | Unit is operational | Recommended to add blue filter to have white light and off-site servicing of the camera and its adapter to access some dust. Services done under Service PO # 7310004178 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071997784314057544/6071997784314057544_signature_28.png | John Wilson N. Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/6071997784314057544/6071997784314057544_signature_31.png
| 2024-11-12 | 2024-11-12 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29758 | DOH | ccmclaboratory01@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1864 |  | 6C GXIV-4D | 8383561 | Due for monthly PM | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071993282023210198/6071993282023210198_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/6071993282023210198/6071993282023210198_signature_31.png
| 2024-11-12 | 2024-11-12 | Pinagbuhatan TB Microscopy Laboratory | 182 Urbano Velasco Ave, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29757 | MACARE-HBDC | peachmango318@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 418 |  | GX-IV-D-10C | 110024245 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071983933362156210/6071983933362156210_signature_28.png | Arlen Delgado, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6071983933362156210/6071983933362156210_signature_31.png
| 2024-11-12 | 2024-11-12 | Alicia Rural Health Unit | Alicia, Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia@gmail.com |  | 29756 | DOH | artoba33@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 767 | 2024-10-21 | GX 4 (6c) | 838276 | None | Replaced module 685309 with module 731666. Conducted Preventive maintenence for all modules. | Replaced module 685309 with module 731666. Conducted Preventive maintenence for all modules. | All modules are available | None | 900-0491RS | 12259 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071977429013049373/6071977429013049373_signature_28.png | Ariel T. Baclig | https://www.jotform.com/uploads/macarewecare/222927978372067/6071977429013049373/6071977429013049373_signature_31.png
| 2024-11-12 | 2024-11-12 | Pilar Rural Health Unit | Calongay, Pilar | 5 | Sorsogon | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29755 | DOH | manalangmarkramir@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1460 |  | GX4-6C | 838266 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Checked fan filter Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071971449221668109/6071971449221668109_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/6071971449221668109/6071971449221668109_signature_31.png
| 2024-11-12 | 2024-11-12 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29754 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  |  | GXIV | 110008530 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071966290723686327/6071966290723686327_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6071966290723686327/6071966290723686327_signature_31.png
| 2024-11-12 | 2024-11-12 | Sta. Cruz Rural Health Unit | Poblacion, Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29753 | DOH | chendalagan72@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 550 | 2024-10-18 | GX-IV 6C 4D | 838505 | Module B1 encounter multiple temperature related errors | Replaced the Defective Module SN: 861939 with new Module SN: 647666 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log | Perform Preventive Maintenance and Module Replacement | Only two Modules are Operational Module B2 and B3 reported to Cepheid |  | 900-0491rs | DR 12246 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071965583955378431/6071965583955378431_signature_28.png | Dorothy Dalagan,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6071965583955378431/6071965583955378431_signature_31.png
| 2024-11-12 | 2024-11-12 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29752 | MACARE-HBDC | roselyn.alua@redcross.org.ph | 25 |  |  | MCS+ | O3J030 | For checking of centrifuge lock mechanism. | Troubleshooting and repair. | Performed checking and troubleshooting of centrifuge. Performed functionality testing. Secured all lock. | Unit is operational. Centrifuge lock plate is loosed and recommended for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071934092127436477/6071934092127436477_signature_28.png | ROSELYN JRA C ALUA | https://www.jotform.com/uploads/macarewecare/222927978372067/6071934092127436477/6071934092127436477_signature_31.png
| 2024-11-12 | 2024-11-08 | Tapaz Rural Health Unit | Tapaz, Capiz | 6 | Capiz | TB Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29751 | DOH | rhutapaz@gmail.com | 35 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 733 ASSAY: Xpert HIV-1 Viral Load, COUNT: 68 |  | GX-4 6c | 837615 | None | Perform Preventive Maintenance | #NAME? | 3 modules operational, 1 module waiting for replacement due to failed Xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071875235426483758/6071875235426483758_signature_28.png | Sherlym G. Fancubila | https://www.jotform.com/uploads/macarewecare/222927978372067/6071875235426483758/6071875235426483758_signature_31.png
| 2024-11-12 | 2024-11-08 | Tapaz Rural Health Unit | Tapaz, Capiz | 6 | Capiz | TB Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29750 | DOH | rhutapaz@gmail.com | 35 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 733 ASSAY: Xpert HIV-1 Viral Load, COUNT: 68 |  | GX-4 6c | 837615 | None | Perform Preventive Maintenance | #NAME? | 3 modules operational, 1 module waiting for replacement due to failed Xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071875235426483758/6071875235426483758_signature_28.png | Sherlym G. Fancubila | https://www.jotform.com/uploads/macarewecare/222927978372067/6071875235426483758/6071875235426483758_signature_31.png
| 2024-11-11 | 2024-11-11 | Bicol Regional Blood Center | Rizal st., Legazpi City | 5 | Albay | Blood Center | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | versonolicia27@gmail.com |  | 29749 | MACARE-HBDC | bsl-snrl@bicol.doh.gov.ph | 25 |  |  | MCS+ | 15H032G | MCS+ for Calibration | Performed Calibration Cleaned Unit Checked Program Parameters Checked and adjust line sensor value Checked rollers, DPM, SPM, valve & centrifuge Performed functionality test Performed simulation running | Performed Calibration Performed Preventive Maintenance | Unit is operational | Cuff is for observation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071211227941361341/6071211227941361341_signature_28.png | Verson Olicia | https://www.jotform.com/uploads/macarewecare/222927978372067/6071211227941361341/6071211227941361341_signature_31.png
| 2024-11-11 | 2024-11-11 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29748 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 25 |  |  | MCS+ | 11J219 | Noise of the centrifuge during platelet apheresis. | Check and repair chuck of the centrifuge of the machines. | Check and repair chuck of the centrifuge of the machines. | Unit machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071175059312843098/6071175059312843098_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6071175059312843098/6071175059312843098_signature_31.png
| 2024-11-11 | 2024-11-11 | Sta Rosa Community District Hospital | Sta Rosa, Olango, Lapulapu City | 7 | Cebu | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29747 | DOH | ariestumz@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 227 |  | GX4 | 838534 | Due for Annual Calibration, Preventive Maintenance | Performed annual calibration in all modules Generate and back-up database, IQR, System Log, and Archive data results, Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. | Performed annual calibration in all modules Generate and back-up database, IQR, System Log, and Archive data results, Performed preventive maintenance and clean filer, Icore, Ultrasonic Heater and plunger cleaning. | All units are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071172219318377837/6071172219318377837_signature_28.png | Ariston Tumulak, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6071172219318377837/6071172219318377837_signature_31.png
| 2024-11-11 | 2024-11-04 | Project 7 Social Hygiene Clinic | Veterans Village, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29746 | MACARE-HBDC | gpdmagnaye30@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Xpert HBV VL, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 490 |  | GX IV 10C D | 120000577 | Unit due for maintenance | Performed Preventive maintenance | Checked and cleaned cartridge bays, plunger rod, icore slot and filter | unit is operational all modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071165333613306566/6071165333613306566_signature_28.png | Michele Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/6071165333613306566/6071165333613306566_signature_31.png
| 2024-11-11 | 2024-11-07 | Kalilangan Rural Health Unit | West Poblacion, Kalilangan | 10 | Bukidnon | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29745 | DOH | myrnasuyao17@gmail.com | 27 |  |  | GX-4 6 Colors | 838571 | None | None | Xpertcheck Calibration and Preventive Maintenance | All modules passed in calibration and functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071049429971684354/6071049429971684354_signature_28.png | Myrna Suyao | https://www.jotform.com/uploads/macarewecare/222927978372067/6071049429971684354/6071049429971684354_signature_31.png
| 2024-11-11 | 2024-11-05 | Sarangani Health Care Facility | Alabel, Sarangani | 12 | Sarangani | Tb-Lab | NO | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29744 | DOH | ronayadao@yahoo.com | 22.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 764 |  | GxIV-6C | 838542 | None | Preventive maintenance | Installed Hiv viral load Assay | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071016686862043856/6071016686862043856_signature_28.png | Ronora Yadao | https://www.jotform.com/uploads/macarewecare/222927978372067/6071016686862043856/6071016686862043856_signature_31.png
| 2024-11-11 | 2024-11-06 | Teresita L Jalandoni Provincial Hospital | Silay, Negros Occendintal | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29743 | DOH | aklit26@yahoo.com | 30 |  |  | Gx 4 | 838464 | Defective Module 3 sn 210028917 Clean Fan Filter Clean Fan IQR, System log & assay stats | New module A3 699622 installed Clean Fan Filter Clean Fan IQR, System log & assay stats | New module A3 699622 installed Clean Fan Filter Clean Fan IQR, System log & assay stats | Defective Module A2 and A4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071014689196481755/6071014689196481755_signature_28.png | Ma Kriska Awacay | https://www.jotform.com/uploads/macarewecare/222927978372067/6071014689196481755/6071014689196481755_signature_31.png
| 2024-11-11 | 2024-11-06 | Tampakan Rural Health Unit | Poblacion, Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29742 | DOH | jmedtech_lab@yahoo.com | 27.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 612 |  | GxIV-6C | 838544 | Monitor problem | Check up and Troubleshooting Preventive maintenance | Troubleshoot monitor problem Genexpert Cleaning Checking of errors Problem and errors reporting Data collection | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071014026866452412/6071014026866452412_signature_28.png | JOVI G GELIO | https://www.jotform.com/uploads/macarewecare/222927978372067/6071014026866452412/6071014026866452412_signature_31.png
| 2024-11-11 | 2024-11-05 | Malapatan Rural Health Unit | Poblacion, Malapatan Rural Health Unit | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | kamilahnichol@yahoo.com |  | 29741 | DOH | kamilahnichol@yahoo.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 673 |  | GxIV-6C | 838419 | Module 4 (H/W Failed) multiple error:2014 - Temperature errors | Genexpert troubleshooting Genexpert Preventive maintenance | Perform module 4 problem troubleshooting Perform Genexpert Cleaning and maintenance Checking of errors Problem/errors reporting Data collection | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6071014026867170861/6071014026867170861_signature_28.png | Kamilah Nichol Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/6071014026867170861/6071014026867170861_signature_31.png
| 2024-11-10 | 2024-11-06 | Rapu-Rapu Rural Health Unit | Poblacion, Rapu-Rapu | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29740 | DOH | suzetteoguan1@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 605 |  | GX4-6C | 838565 | B4/673571 - service required (calibration) | Replaced module: B4 - 673571 >> 683709 Cleaned cartridge bays, plunger rod, icore Generate Data | Replaced module B4/673571 Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6070549127916442230/6070549127916442230_signature_28.png | Suzette C.Oguan | https://www.jotform.com/uploads/macarewecare/222927978372067/6070549127916442230/6070549127916442230_signature_31.png
| 2024-11-10 | 2024-11-05 | Iriga City Health Office | Iriga City Hall Compound, Sta. Cruz Sur, Iriga City | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29739 | UNOPS | mongealyzza@yahoo.com.ph | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 |  | GX4-10C | 110023511 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6070548317911782131/6070548317911782131_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/6070548317911782131/6070548317911782131_signature_31.png
| 2024-11-10 | 2024-11-05 | Iriga City Health Office | Iriga City Hall Compound, Brgy. Sta. Cruz Sur, Iriga City | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29738 | DOH | mongealyzza@yahoo.com.ph | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1447 |  | GX4-6C | 838391 | B2/210044677 - service required | Replaced module: B2 - 210044677 >> 718383 Cleaned cartridge bays, plunger rod, Icore Generate Data | Replaced B2/210044677 Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6070547927915393394/6070547927915393394_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/6070547927915393394/6070547927915393394_signature_31.png
| 2024-11-10 | 2024-10-29 | Estevez Memorial Hospital, Inc. | Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City, 4500 Albay | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29737 | COMMERCIAL | reylovendino83@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 193 ASSAY: Xpert MTB/RIF Ultra, COUNT: 148 ASSAY: Xpert MTB-XDR, COUNT: 11 ASSAY: Xpert SARS-CoV-2, COUNT: 12 |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore, fan Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6070547367918066663/6070547367918066663_signature_28.png | Rey a. Lovendino | https://www.jotform.com/uploads/macarewecare/222927978372067/6070547367918066663/6070547367918066663_signature_31.png
| 2024-11-10 | 2024-11-07 | Naujan Municipal Health Office | Naujan, Oriental Mindoro | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29736 | DOH | ellainehernandez27@yahoo.com | 25 |  | 2024-10-18 | GXIV | 838226 | Unit for preventive maintenance. For module replacement of module B1. | Preventive maintenance and module replacement | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B1 (634939) in replacement to defective B1 (632561). | DR#12247 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6070174978503612559/6070174978503612559_signature_28.png | Ellaine Villas | https://www.jotform.com/uploads/macarewecare/222927978372067/6070174978503612559/6070174978503612559_signature_31.png
| 2024-11-10 | 2024-11-07 | Naujan Municipal Health Office | Naujan, Oriental Mindoro | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29735 | DOH | ellainehernandez27@yahoo.com | 25 |  | 2024-10-18 | GXIV | 838226 | Unit for preventive maintenance. For module replacement of module B1. | Preventive maintenance and module replacement | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B1 (634939) in replacement to defective B1 (632561). | DR#12247 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6070174728504227911/6070174728504227911_signature_28.png | Ellaine Villas | https://www.jotform.com/uploads/macarewecare/222927978372067/6070174728504227911/6070174728504227911_signature_31.png
| 2024-11-08 | 2024-11-08 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29734 | DOH | burauenlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 483 | 2024-11-08 | GXIV | 838551 | UPS defective (Not charging) SN/ AT1000190123ST0195 | Preventive Maintenance Troubleshoot | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6068520073118172982/6068520073118172982_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6068520073118172982/6068520073118172982_signature_31.png
| 2024-11-08 | 2024-11-08 | Maria Reyna University - Xavier Hospital | Hayes Street, Pinikitan, Cagayan de Oro City | 10 | Misamis Oriental | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29733 | COMMERCIAL | Jmrabe84@gmail.com | 27 |  |  | MCS+ | 05M131 | Unable to read protocol card. CRC failure. Self test problem. BIOS CRC failure also occur last inspection. 6 roller rubbers are thin and need to be replace. | Replacement of CPU PCB/ Motherboard is needed. Replacement of 6 rollers is also needed. | Inspection and troubleshooting of MCS+ with SN: 05M131 | MCS+ unable to start. Equipment not functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6068369137023067438/6068369137023067438_signature_28.png | John Michael E. Rabe | https://www.jotform.com/uploads/macarewecare/222927978372067/6068369137023067438/6068369137023067438_signature_31.png
| 2024-11-08 | 2024-11-07 | Philippine Red Cross | Sta. Cruz, Laguna | 4A | Laguna | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29732 | COMMERCIAL | lagunablood@redcross.org.ph | 22 |  |  | MCS+ | 02J046 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Cleaned centrifuge Cleaned line sensor Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6068340493617999968/6068340493617999968_signature_28.png | Hiltor Jonhn M. Abustan | https://www.jotform.com/uploads/macarewecare/222927978372067/6068340493617999968/6068340493617999968_signature_31.png
| 2024-11-08 | 2024-11-06 | Cardinal Santos Medical Center | Wilson St. San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 29731 | COMMERCIAL | microbiology.csmc@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 144 ASSAY: Xpert HPV, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1075 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules B1, B2 and B4 operational. Module B3 disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6068340493619936085/6068340493619936085_signature_28.png | Ellaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/6068340493619936085/6068340493619936085_signature_31.png
| 2024-11-07 | 2024-11-07 | Medilinx MMC | Amorsolo st., Makati | NCR | NCR 4th - Makati | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | johnjayresco.medilinx@outlook.com |  | 29730 | COMMERCIAL | johnjayresco@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 227 ASSAY: Xpert CT/NG, COUNT: 837 ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert HPV, COUNT: 744 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3282 ASSAY: Others, COUNT: 13 ASSAY: Xpert HCV VL, COUNT: 92 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1977 |  | GX 4 D (10C) | 110001647 | Due for maintenance and for installation of printer Xpertcheck due 11/15/2024 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Back up files Test run unit | Performed maintenance and installation of printer | B2,B3,B4 B1 encounter 5007 error | For xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067864163223705335/6067864163223705335_signature_28.png | John Jay Dawang Resco | https://www.jotform.com/uploads/macarewecare/222927978372067/6067864163223705335/6067864163223705335_signature_31.png
| 2024-11-07 | 2024-11-06 | Philippine Red Cross | Kalayaan Ave. QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29729 | MACARE-HBDC | villamarevangeline103@gmail.com | 26 |  |  | MCS + | 02A227 | As per end user they encounter reddish platelet | Checked program parameters Checked and adjust line sensor value to 3071 Checked rollers,dpm,spm,centrifuge,weigher Adjust bowl optics to 2800 and 25 without bowl Performed functionality test Performed simulation running | Performed troubleshooting | Operational for observation | Workwith Gian Magnaye |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067863863223644327/6067863863223644327_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/6067863863223644327/6067863863223644327_signature_31.png
| 2024-11-07 | 2024-11-04 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29728 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2404 ASSAY: Xpert SARS-CoV-2, COUNT: 1284 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1310 |  | GXXVI D | 110014171 | Encounter continous error 5007 in different module location. Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed troubleshooting. Performed maintenance | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067863393225041749/6067863393225041749_signature_28.png | National Children's Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6067863393225041749/6067863393225041749_signature_31.png
| 2024-11-07 | 2024-11-07 | Maco Rural Health Unit | Poblacion, Maco | 11 | Davao del Sur | TB-DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29727 | DOH | edlienkai26@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1137 | 2024-10-21 | GX-IV 6C 4D | 838500 | Module B4 encounter multiple temperature related errors For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replaced the Defective Module SN: 862350 with new Module SN: 688471 Run XpertCheck on three Modules | Perform Module Replacement Perform Xpertcheck Perform Preventive Maintenance | Module B3 for Service Required due to multiple temperature related errors Only 3 Modules are operational |  | Xpertcheck CE-5 900-0491rs | DR 12254 & DR 12003 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067747832317947420/6067747832317947420_signature_28.png | Edlien Kai Prado,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6067747832317947420/6067747832317947420_signature_31.png
| 2024-11-07 | 2024-11-07 | National Crop Protection Center | Los Baños, Laguna | 4A | Laguna | Plant Nematology | NO | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | jincarino@gmail.com |  | 29726 | COMMERCIAL | jncarino1@up.edu.ph | 25 |  |  | Axiolab 5 | 3165000305 | N/A | N/A | Installation of Axiolab 5 Kohler alignment Functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067677193917628650/6067677193917628650_signature_28.png | Jaybee Ian N. Cariño | https://www.jotform.com/uploads/macarewecare/222927978372067/6067677193917628650/6067677193917628650_signature_31.png
| 2024-11-07 | 2024-11-07 | National Crop Protection Center | Los Baños, Laguna | 4A | Laguna | Plant Nematology | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | jincarino@gmail.com |  | 29725 | COMMERCIAL | jncarino1@up.edu.ph | 25 |  |  | STEMI 305 | 3947012945 | Connection to Labscope | Screwed the head properly | Checked connection Functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067673483917097059/6067673483917097059_signature_28.png | Jaybee Ian N. Cariño | https://www.jotform.com/uploads/macarewecare/222927978372067/6067673483917097059/6067673483917097059_signature_31.png
| 2024-11-07 | 2024-11-07 | University of the Philippines - Los Baños | UPLB, Los Baños, Laguna | 4A | Laguna | Phycology Lab | NO | REPAIR | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29724 | COMMERCIAL | mmartinezgoss@gmail.com | 25 |  |  | Axio Scope.A1 | 3331000862 | HAL 100 not working | Repaired HAL 100 Re-attached Condenser Aligned Kohler | Repaired HAL 100 Re-attached Condenser Aligned Kohler Functionality Test | Unit is operational | Due for preventive maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067673463916802873/6067673463916802873_signature_28.png | Milagrosa R. Martinez-Goss | https://www.jotform.com/uploads/macarewecare/222927978372067/6067673463916802873/6067673463916802873_signature_31.png
| 2024-11-07 | 2024-11-06 | Libacao Rural Health Unit | Libacao, Aklan | 6 | Aklan | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29723 | DOH | haladeeb18@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 391 | 2024-10-21 | GX-4 6C | 838540 | Module B1/616665 failed the latest Xpertcheck. | Cleaned optics, plunger rods, module bay areas, filter; Archive tests | Replaced module B1 with new module SN 863340. | 4 modules operational |  |  | DR 12257 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067670364731225386/6067670364731225386_signature_28.png | MARIA HALA DEEB ZAULDA | https://www.jotform.com/uploads/macarewecare/222927978372067/6067670364731225386/6067670364731225386_signature_31.png
| 2024-11-07 | 2024-11-07 | Labuan General hospital | Labuan, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 29722 | DOH | jenne060213@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2329 | 2024-11-07 | GX 4 | 838568 | System unit unable to turn on. | Replaced system unit CMOS batteries and re udjust RAM. | Replaced system unit CMOS batteries and re udjust RAM | Module 1 and 3 are available while modules 3 and 4 are for replacement | Module 3 and 4 are for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067667002313945874/6067667002313945874_signature_28.png | Jennerose Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6067667002313945874/6067667002313945874_signature_31.png
| 2024-11-07 | 2024-11-06 | Altavas Rural Health Unit | Altavas, Aklan | 6 | Aklan | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29721 | DOH | nycolebenedicto@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 601 | 2024-10-21 | GX-4 6C | 838458 | Modules B2/861749 and B3/635368 failed the latest Xpertcheck. | Cleaned module bay areas, optics, filter, plunger rods; Archive tests. | - Replaced modules B2 and B3 with new modules with SN 801411 and 793171 respectively. | 4 modules operational |  |  | DR 12256 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067593025284783253/6067593025284783253_signature_28.png | MA. NYCOLE S. BENEDICTO | https://www.jotform.com/uploads/macarewecare/222927978372067/6067593025284783253/6067593025284783253_signature_31.png
| 2024-11-07 | 2024-11-07 | NAZARETH HEALTH CENTER | Nazareth , Cagayan de Oro City | 10 | Misamis Oriental | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29720 | DOH | nanettebtolentino@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1360 | 2024-10-23 | GX-IV | 838269 | FOR MODULE B3 AND MODULE B3 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE REPLACEMENT: B3 DEF SN:688652 WITH NEW SN:874395 B4 DEF SN:808125 WITH NEW SN:882093 | EQUIPMENT IS OPERATIONAL |  | Certified pre-owned GeneXpert module | 12272 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067582081122744298/6067582081122744298_signature_28.png | Nanette Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/6067582081122744298/6067582081122744298_signature_31.png
| 2024-11-07 | 2024-11-07 | Cong. Simeon G. Toribio Memerial Hospital | Carmen Bohol | 7 | Bohol | TB laboratory | NO | INSTALLATION REPAIR | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29719 | DOH | cristalroy@yahoo.com | 23 |  | 2424-11-07 | GX4 | 838394 | Defective LCD with serial number CNK9430BTS | .Pull out defective LCD .Save Installation Qualification Report .Save System Log file | Install new LCD with serial number CNK83224Q2 | LCD are functional ok to use. All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067569228711714150/6067569228711714150_signature_28.png | Roy Cristal | https://www.jotform.com/uploads/macarewecare/222927978372067/6067569228711714150/6067569228711714150_signature_31.png
| 2024-11-07 | 2024-11-07 | CARMEN HEALTH CENTER | Carmen, Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29718 | DOH | juvymadarang@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 657 | 2024-09-03 | GX-IV | 838442 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT OPERATIONAL | MODULE B2 DISABLED | XPERTCHECK, 5Cartridges/kit | 11982 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6067511806827381787/6067511806827381787_signature_28.png | Juvy | https://www.jotform.com/uploads/macarewecare/222927978372067/6067511806827381787/6067511806827381787_signature_31.png
| 2024-11-06 | 2024-11-06 | Philippine Red Cross | Cagayan de Oro City | 10 | Misamis Oriental | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29717 | COMMERCIAL | mrbc2@redcross.org.ph | 24 |  |  | MCS+ | 02H080 | None | None | Preventive Maintenance and Calibration | Equipment is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6066957201125393970/6066957201125393970_signature_28.png | Karen Torniado | https://www.jotform.com/uploads/macarewecare/222927978372067/6066957201125393970/6066957201125393970_signature_31.png
| 2024-11-06 | 2024-11-06 | Talakag Rural Health Unit | Poblacion, Talakag | 10 | Bukidnon | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29716 | DOH | mhotalakaglaboratory@gmail.com | 29 |  |  | GX-4 | 838566 | None | Calibration | XpertCheck Calibration Preventive Maintenance | All modules and units are operational All modules passed in XPC |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6066871781121638594/6066871781121638594_signature_28.png | Khiza Gessel Tinoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6066871781121638594/6066871781121638594_signature_31.png
| 2024-11-06 | 2024-11-06 | CAGAYAN DE ORO CITY HEALTH OFFICE | Cagayan De Oro, Misamis Oriental | 10 | Misamis Oriental | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | chocdotbdots@gmail.com |  | 29715 | DOH | chocdodiagnosticlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2907 | 2024-09-03 | GX-IV | 848435 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK, 5 Cartridges/Kit | 11980 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6066851941128420867/6066851941128420867_signature_28.png | Charissa B Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/6066851941128420867/6066851941128420867_signature_31.png
| 2024-11-06 | 2024-11-06 | CAGAYAN DE ORO CITY HEALTH OFFICE | Cagayan De Oro, Misamis Oriental | 10 | Misamis Oriental | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | chocdotbdots@gmail.com |  | 29714 | DOH | chocdodiagnosticlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2907 | 2024-09-03 | GX-IV | 838441 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK, 5Cartridges/Kit | 11981 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6066823711104471529/6066823711104471529_signature_28.png | Charissa B Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/6066823711104471529/6066823711104471529_signature_31.png
| 2024-11-06 | 2024-11-06 | Matanao Rural Health Unit | Poblacion, Matanao, Davao del Sur | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29713 | DOH | pamsustento@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 552 | 2024-10-23 | +GX-IV 6C 4D | 838456 | Module 4 for Service Required during last Calibration For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Replaced the Defective Module SN: 861233 with new Module SN: 210057751 Generate IQR, Assay Stat and System Log | Perform Module Replacement and Preventive Maintenance | All Modules are operational |  | GeneXpert Module 6c | DR 12273 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6066820758029843587/6066820758029843587_signature_28.png | Pamela Joyce P. Sustento,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6066820758029843587/6066820758029843587_signature_31.png
| 2024-11-06 | 2024-11-06 | LAGUINDINGAN RURAL HEALTH UNIT | LAGUINDINGAN, MISAMIS ORIENTAL | 10 | Misamis Oriental | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29712 | DOH | chellesambaan@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1200 |  | GX-IV | 838575 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6066735771121997354/6066735771121997354_signature_28.png | Rochelle M. Samban | https://www.jotform.com/uploads/macarewecare/222927978372067/6066735771121997354/6066735771121997354_signature_31.png
| 2024-11-06 | 2024-11-05 | Mabinay Medicare Community Hospital | Mabinay, Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29711 | DOH | zyrusbohol@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 587 |  | GX-4 6C | 838481 | #NAME? | . | #NAME? | Modules 1,2,3 operational. Module 4 waiting for new module replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6066693338022349234/6066693338022349234_signature_28.png | Zyrus Bohol | https://www.jotform.com/uploads/macarewecare/222927978372067/6066693338022349234/6066693338022349234_signature_31.png
| 2024-11-05 | 2024-11-05 | Tomas Claudio Health Center (District A) | Camus Street, Corner E. Quirino Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29710 | DOH | arleen_p_alconera@yahoo.com | 27.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 575 | 2024-10-23 | GX-IV 6C 4D | 838492 | Module B3, failed during calibration For Preventive Maintenance | Replaced the Defective Module SN: 862220 with new Module SN:870134 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | Item Code: 900-0491rs GeneXpert 6c Module | DR 12271 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6066033704725903827/6066033704725903827_signature_28.png | Arleen P. Alconera,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6066033704725903827/6066033704725903827_signature_31.png
| 2024-11-05 | 2024-11-05 | Kalinga Health Hub - Marikina | JP Rizal St. Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29709 | CONSORTIUM | bryllcruz08@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1597 | 2024-10-30 | GX-IV-D-6C | 11008418 | Module B1 for replacement | Replaced module B1 | Performed maintenance Replaced module B1 Generated reports | Unit is operational. All modules operational. |  | Old: B1/678461 Old: B2/210016600 | DR# 12287 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6066009420028755144/6066009420028755144_signature_28.png | Lorenzo Bryll Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6066009420028755144/6066009420028755144_signature_31.png
| 2024-11-05 | 2024-11-04 | Santolan Super Health Center | Evangelista Ave. Paaig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29707 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 834 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065947133611684645/6065947133611684645_signature_28.png | Marilyn Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6065947133611684645/6065947133611684645_signature_31.png
| 2024-11-05 | 2024-10-30 | Ospital ng Makati | Sampaguita St. Makati | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 29708 | MACARE-HBDC | jinx01gs@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 169 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065947133617922559/6065947133617922559_signature_28.png | Cristina Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6065947133617922559/6065947133617922559_signature_31.png
| 2024-11-05 | 2024-10-30 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29706 | DOH | gbmmbg@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 691 |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065943720341968897/6065943720341968897_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/6065943720341968897/6065943720341968897_signature_31.png
| 2024-11-05 | 2024-10-30 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29705 | DOH | seansimonpaloyo@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1251 |  | GXIV-4-D-6C | 838473 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Fully Operational - Unable to perform XpertCheck to be reported |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065942070343484265/6065942070343484265_signature_28.png | Manolo C. Paloyo | https://www.jotform.com/uploads/macarewecare/222927978372067/6065942070343484265/6065942070343484265_signature_31.png
| 2024-11-05 | 2024-11-05 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29704 | DOH | ammegzedd@gmail.com | 21 |  | 2024-11-05 | GXIV | 838421 | n/a | n/a | Performed monthly preventive maintenance | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065915298692925697/6065915298692925697_signature_28.png | gemma malate | https://www.jotform.com/uploads/macarewecare/222927978372067/6065915298692925697/6065915298692925697_signature_31.png
| 2024-11-05 | 2024-11-05 | Padre Burgos Rural Health Unit | Padre Burgos, Southern Leyte | 8 | Southern Leyte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29703 | UNOPS | keishaelezze14@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 186 |  | 10C GXIV-4D | 110023514 | Stuck cartridge at module B4 | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065874913491960870/6065874913491960870_signature_28.png | Keisha Elezze Batalon | https://www.jotform.com/uploads/macarewecare/222927978372067/6065874913491960870/6065874913491960870_signature_31.png
| 2024-11-04 | 2024-10-30 | Datu Odin Sinsuat District Hospital | Datu Odin Sinsuat, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29702 | DOH | dosdhlaboratory@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 820 | 2024-10-30 | GxIV-6C | 838205 | None | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors Data collection and Reporting | 4 operational modules Calibration Results: Module 1 - Passed Module 2 - Skipped (new module) Module 3 - Passed Module 4 - Passed |  | Xpertcheck cartridge/kits Lot#: 1001075092 | DR#: 12029 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065345436861217096/6065345436861217096_signature_28.png | ROWENA B. SALAZAR | https://www.jotform.com/uploads/macarewecare/222927978372067/6065345436861217096/6065345436861217096_signature_31.png
| 2024-11-04 | 2024-11-04 | Samboan Rural Health Unit | Samboan Rural Health Unit | 7 | Cebu | Tb laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29701 | DOH | sheneannesenagan@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1550 |  | GX4 | 838513 | Defective Air Filter Generate IQR, Gxx files, System Log Performed calibration on genexpert machine Conduct Preventive Maintenance | Installation of New Filter Generate IQR, Gxx files, System Log Performed calibration on genexpert machine Conduct Preventive Maintenance | Installation of New Filter Generate IQR, Gxx files, System Log Performed calibration on genexpert machine Conduct Preventive Maintenance | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065163220655374351/6065163220655374351_signature_28.png | Shene Anne A. Senagan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6065163220655374351/6065163220655374351_signature_31.png
| 2024-11-04 | 2024-11-04 | Cebu South Medical Center | San Isidro, Talisay City Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29700 | DOH | csmcgenexpert@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1151 |  | GX4 | 838559 | Defective Module 4 SN # 707812 during calibration. | Replaced module 4 with new module SN # 730678 Conduct Preventive maintenance Clean Icore, ultrasonic horn and filter Generate GXX files, IQR and System Log, | Replaced module 4 with new module SN # 730678 Conduct Preventive maintenance Clean Icore, ultrasonic horn and filter Generate GXX files, IQR and System Log, | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065156163517051817/6065156163517051817_signature_28.png | Kristel Sheen Miranda, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6065156163517051817/6065156163517051817_signature_31.png
| 2024-11-04 | 2024-11-04 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | Pull-Out UPS | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29699 | DOH | luzvimindanavidad2@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 965 |  | 6C GXIV-4D | 838297 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6065087782406244986/6065087782406244986_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6065087782406244986/6065087782406244986_signature_31.png
| 2024-11-04 | 2024-10-30 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29698 | DOH | dorothygundaya@gmail.com | 28 |  | 2024-10-18 | GX-4 6 Colors | 838403 | Module B2 found Service Required thru XPC | Module Replacement | Replacing Module B2 with SN: 672225 with SN: 630596 | All modules and units are functional |  | Certified pre-owned GeneXpert Module | 12248 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6064997254284008568/6064997254284008568_signature_28.png | Dorothy Gundaya | https://www.jotform.com/uploads/macarewecare/222927978372067/6064997254284008568/6064997254284008568_signature_31.png
| 2024-11-04 | 2024-11-04 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | biomed@nch.doh.gov.ph |  | 29697 | MACARE-HBDC | nchmbdteam@gmail.com | 27 |  |  | MCS + | 17H632SPG | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Performed visual checkinf Checked rollers,dpm,spm,centrifuge,weigher,air detector Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6064995096701208055/6064995096701208055_signature_28.png | NATIONAL CHILDREN'S HOSPITAL | https://www.jotform.com/uploads/macarewecare/222927978372067/6064995096701208055/6064995096701208055_signature_31.png
| 2024-11-04 | 2024-10-30 | Palompon Rural Health Unit | Palompon Leyte | 8 | Leyte | TB Laboratory | NO | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29696 | DOH | luzvimindanavidad2@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 965 | 2024-10-30 | GXIV | 838297 | UPS defective | UPS replacement | #NAME? | Unit is Operational |  | 1-UPS (ACCUPOWER) | 12281 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6064919680121369810/6064919680121369810_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6064919680121369810/6064919680121369810_signature_31.png
| 2024-11-02 | 2024-11-02 | Cardinal Santos Medical Center | Wilson st,Greenhills SanJuan | NCR | NCR 2nd - San Juan | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29695 | COMMERCIAL | csmclab@csmc.ph | 26 |  |  | MCS + | 08M461 | As per enduser they encounter pressure problem during process. | Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6063426293802913918/6063426293802913918_signature_28.png | Andrea Sumilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6063426293802913918/6063426293802913918_signature_31.png
| 2024-11-02 | 2024-10-31 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29694 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1510 ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Others, COUNT: 116 ASSAY: Xpert HBV VL, COUNT: 1065 ASSAY: Xpert HPV, COUNT: 138 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2338 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 9625 ASSAY: Xpert HCV VL, COUNT: 147 |  | GX 16 D | 110005587 | A1 encounter error 4009,4017 temp error A2 encounter error 4010,4017 temp error C2 encounter error 4010, temp error D2 encounter error 4017 temp error C4 newly replaced encounter error 4004 valve motion | Gathered data Performed troubleshooting | Performed troubleshooting | A1,A2,C2,C4,D2 unoperational reported already. 11 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6063425903802941100/6063425903802941100_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6063425903802941100/6063425903802941100_signature_31.png
| 2024-11-02 | 2024-10-31 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29693 | MACARE-HBDC | phcmolecularlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 508 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 764 |  | GX16 D | 838655 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6063425563802510348/6063425563802510348_signature_28.png | ISON M. SAN JOSE | https://www.jotform.com/uploads/macarewecare/222927978372067/6063425563802510348/6063425563802510348_signature_31.png
| 2024-11-02 | 2024-11-02 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29692 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  |  | MCS + | 02D168 | As per end user they encounter reddish color of platelet. Line sensor potentiometer unable to adjust value of desire value Line sensor value 2987 | Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Troubleshooting | Unit operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6063423033808763927/6063423033808763927_signature_28.png | Cathlyn I. Pua | https://www.jotform.com/uploads/macarewecare/222927978372067/6063423033808763927/6063423033808763927_signature_31.png
| 2024-10-31 | 2024-10-31 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29691 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2416 ASSAY: Xpert HIV-1 Viral Load, COUNT: 339 | 2024-10-21 | 6C GXIV-4D | 838285 | Module A1 defective | -Replace defective module A1/698265 with module SN: 669847 | #NAME? | All modules are functional. |  | 900-0491RS Certified pre-owned GeneXpert module SN. 669847 | DR no. 12252 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6061547350197550645/6061547350197550645_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/6061547350197550645/6061547350197550645_signature_31.png
| 2024-10-31 | 2024-10-31 | Valenzuela Social Hygiene Clinic | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29690 | COMMERCIAL | markanthony.manaois@gmail.com | 25 |  |  | GXIV-10C | 110020699 | Modules not detected | Set GX network settings | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6061527070842195451/6061527070842195451_signature_28.png | Mark Manaois | https://www.jotform.com/uploads/macarewecare/222927978372067/6061527070842195451/6061527070842195451_signature_31.png
| 2024-10-30 | 2024-10-29 | Tropical Disease Foundation | Amorsolo st , Makati City | NCR | NCR 4th - Makati | Laboaratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29689 | CONSORTIUM | gsmondejar@tdf.org.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2079 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 |  | Gx4 D | 812560 | Due for maintenance and troubleshooting | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and troubleshooting | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6060853833612176360/6060853833612176360_signature_28.png | Gabriel Mondejar | https://www.jotform.com/uploads/macarewecare/222927978372067/6060853833612176360/6060853833612176360_signature_31.png
| 2024-10-30 | 2024-10-30 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29688 | DOH | buhanginlaboratory@gmail.com | 25.6 |  | 2024-10-25 | GX-IV 6C 4D | 838232 | CPU System Unit not working For Monthly Preventive Maintenance | Replace the defective CPU System Unit SN: CPH0420F06673 with new CPU System Unit SN: CPH0420F08168 Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log | Perform CPU System Unit replacement and Monthly Preventive Maintenance | Successfully Installed the replacement CPU All Modules are Operational | No recorded assay due to new CPU System Unit | Item Code: GX-DM720S GeneXpert Desktop (Lenovo Thinkcentre m720) | DR 12279 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6060755327212971072/6060755327212971072_signature_28.png | Catherine B. Deocampo,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6060755327212971072/6060755327212971072_signature_31.png
| 2024-10-30 | 2024-10-15 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29687 | MACARE-HBDC | honeyeichel@gmail.com | 23 | ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 3643 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 ASSAY: Others, COUNT: 1554 |  | GXIV-D 10C | 110017894 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6060682625514212998/6060682625514212998_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/6060682625514212998/6060682625514212998_signature_31.png
| 2024-10-30 | 2024-10-15 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29686 | MACARE-HBDC | honeyeichel@gmail.com | 23 | ASSAY: Xpert SARS-CoV-2, COUNT: 6357 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Others, COUNT: 940 |  | GXIV-D 6C | 829390 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6060682325518670608/6060682325518670608_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/6060682325518670608/6060682325518670608_signature_31.png
| 2024-10-30 | 2024-10-30 | Datu Halun Sakilan Memorial Hospital | Bongo, Tawi-Tawi | BARMM | Tawi-Tawi | Laboratory | YES | INSTALLATION CALIBRATION MAINTENANCE GX Barcode Sticker Installation | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | abdulmuidabdilla@gmail.com |  | 29685 | DOH | joel_kamlian2003@yahoo.com | 23 |  | 2024-09-05 | GX 4 | 838834 | None | Xpertcheck calibration for modules 2, 3, 4 all modules passed. Preventive Maintenance, Barcode sticker installation. | Xpertcheck calibration for modules 2, 3, 4 all modules passed. Preventive Maintenance, Barcode sticker installation. | All modules are available and ready to use | None | XPERTCHECK-CE-5 | 12040 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6060630105218780978/6060630105218780978_signature_28.png | Almu-eed M. Abdilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6060630105218780978/6060630105218780978_signature_31.png
| 2024-10-29 | 2024-10-29 | Macare In-House | Quezon City | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Ronnel Belizar |  | rcb.macare@gmail.com |  | Uninterruptible Power Supply | 29684 | MACARE-HBDC | htoco@macare-medicals.com | 25 |  | 2024-10-29 | AT1000 | AT1000181128ST0059, AT1000181128ST0106, AT1000181128ST0188, AT1000180910ST0568, AT1000190123ST0375, AT1000190402ST0439, AT1000180910ST0078, AT1000181128ST0100, AT1000181128ST0032 | Battery for replacement | Replaced battery | #NAME? | Unit is operational SN: AT1000181128ST0059, AT1000181128ST0106, AT1000181128ST0188, AT1000180910ST0568, AT1000190123ST0375, AT1000190402ST0439, AT1000180910ST0078, AT1000181128ST0100, AT1000181128ST0032 |  | Battery | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6059941403612552488/6059941403612552488_signature_28.png | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/6059941403612552488/6059941403612552488_signature_31.png
| 2024-10-29 | 2024-10-29 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29683 | DOH | claritomirasol@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 448 | 2024-10-29 | GXIV | 838548 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6059817210014124246/6059817210014124246_signature_28.png | Mirasol Clarito | https://www.jotform.com/uploads/macarewecare/222927978372067/6059817210014124246/6059817210014124246_signature_31.png
| 2024-10-29 | 2024-10-29 | Vicente Mendiola Center for Health | Poblacion Naga City, Cebu | 7 | Cebu | TB Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29682 | DOH | geccaprocianos26@gmail.com | 23 |  | 2024-10-29 | GX 4 | 838564 | GeneXpert Machine Installation Due for Xpertcheck | Save Installation Qualification Report Save System Log file Save gxc file. | . Install GeneXpert Machine . Install CPU . Install UPS . Install AVR . Install Printer | All modules are functional ok to use. | Unit Pullout from Balamban Rural Health Unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6059801960912396224/6059801960912396224_signature_28.png | Gecca Carmel T. Procianos | https://www.jotform.com/uploads/macarewecare/222927978372067/6059801960912396224/6059801960912396224_signature_31.png
| 2024-10-29 | 2024-10-29 | Cebu City Health Office | Carreta, Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29681 | DOH | joydarlie72@gmail.com | 27 |  |  | Gx4 | 838527 | Monthly PM | Clean Fan Filter Plunger rod maintenance Clean Module Bay | Clean Fan Filter Plunger rod maintenance Clean Module Bay | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6059776557325060426/6059776557325060426_signature_28.png | Joy Ma. Darlie Labrado | https://www.jotform.com/uploads/macarewecare/222927978372067/6059776557325060426/6059776557325060426_signature_31.png
| 2024-10-28 | 2024-10-23 | Malungon Rural Health Unit | Poblacion, Malungon, Sarangani | 12 | Sarangani | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29680 | DOH | madameerrmie@gmail.com | 28.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 777 | 2024-10-23 | GxIV-6C | 838546 | None | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert Preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert machine Preventive maintenance Checking of errors Data collection | 4 operational modules Calibration Results: All modules Passed Module 1 - Passed Module 2 - Passed Module 3 - Passed Module 4 - Passed |  | Xpertcheck cartridge/kits Lot#:1001075092 | DR#: 12023 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6059294206864380324/6059294206864380324_signature_28.png | Ermie D Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/6059294206864380324/6059294206864380324_signature_31.png
| 2024-10-28 | 2024-10-24 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City b | 12 | South Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29679 | DOH | celinejean16@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-10-24 | GxIV-6C | 838570 | None | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: Module 1 - Skipped (new module) Module 2 - Passed Module 3 - Passed Module 4 - Passed | 4 functional modules | Xpertcheck cartridge/kits Lot#:1001075092 | DR#: 12027 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6059294206864712119/6059294206864712119_signature_28.png | Celine Jean Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6059294206864712119/6059294206864712119_signature_31.png
| 2024-10-28 | 2024-10-28 | Laoag City Health Office | Laoag City, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29678 | DOH | denver.mahor@gmail.com | 28.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 723 | 2024-10-18 | GXIV-D 6C | 838416 | Module B2:86118 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational B2:615885 new module |  | Genexpert module ( 615885) | 12245 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6059137347308638363/6059137347308638363_signature_28.png | Rodenver Mahor | https://www.jotform.com/uploads/macarewecare/222927978372067/6059137347308638363/6059137347308638363_signature_31.png
| 2024-10-28 | 2024-10-09 | Libon Rural Health Unit | Poblacion, Libon | 5 | Albay | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29677 | DOH | marilyn@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1119 |  | GX4-6C | 838397 | UPS no power | Checked power cable Checked/replaced Fuse Tested UPS battery dependability(defective) | Performed Troubleshooting | Unit is operational UPS non functional | UPS battery reported for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6058938175028547830/6058938175028547830_signature_28.png | Marilyn S. Razal | https://www.jotform.com/uploads/macarewecare/222927978372067/6058938175028547830/6058938175028547830_signature_31.png
| 2024-10-28 | 2024-10-22 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29676 | DOH | sherryannemarrero@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2324 |  | GXIV-4-D-6C | 838281 | For Preventive Maintenance and Troubleshooting | #NAME? | Performed Preventive Maintenance and Troubleshooting | - Unit is partially operational - Module B2 651808 for reporting - Module B1 678653 for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6058902369323394587/6058902369323394587_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/6058902369323394587/6058902369323394587_signature_31.png
| 2024-10-28 | 2024-10-22 | Baguio General Hospital and Medical Center | Baguio City | 1 | Benguet | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29675 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 19 | ASSAY: Xpert HBV VL, COUNT: 195 ASSAY: Xpert SARS-CoV-2, COUNT: 198 ASSAY: Xpert HCV VL, COUNT: 48 ASSAY: Xpert HIV-1 Viral Load, COUNT: 34 | 2024-10-22 | GXIV-4-D-10C | 110016112 | For Preventive Maintenance and Calibration | #NAME? | Performed Preventive Maintenance and Calibration | Unit is Fully Operational |  | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6058901009321938793/6058901009321938793_signature_28.png | Elpidio Cadut III | https://www.jotform.com/uploads/macarewecare/222927978372067/6058901009321938793/6058901009321938793_signature_31.png
| 2024-10-28 | 2024-10-28 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | synch28@yahoo.com |  | 29674 | MACARE-HBDC | msff-tondo-labtech@pris.msf.org | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5504 | 2024-10-17 | GX16-10C | 110016343 | Unit for preventive maintenance. For module replacement of module C3. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module C3 (210117673) in replacement to defective C3 (914567). | DR#12236 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6058860863017599717/6058860863017599717_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6058860863017599717/6058860863017599717_signature_31.png
| 2024-10-27 | 2024-10-27 | Baybay City Health Office | Baybay Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29673 | DOH | sajecopooh20@gmail.com | 27 |  | 2024-10-23 | GXIV | 838555 | n/a | PM | #NAME? | Unit is operational (Module A4 disabled due to multiple temperature error) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6058053427383683185/6058053427383683185_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/6058053427383683185/6058053427383683185_signature_31.png
| 2024-10-26 | 2024-10-25 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29672 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 02D166 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6057041907286042898/6057041907286042898_signature_28.png | Chrestien Cule | https://www.jotform.com/uploads/macarewecare/222927978372067/6057041907286042898/6057041907286042898_signature_31.png
| 2024-10-26 | 2024-10-25 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29671 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 08H188 | Due for maintenance Urethane rollers yellowish color | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | For replacement of roller Unit Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6057041497285066037/6057041497285066037_signature_28.png | Chrestien Cule | https://www.jotform.com/uploads/macarewecare/222927978372067/6057041497285066037/6057041497285066037_signature_31.png
| 2024-10-26 | 2024-10-25 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29670 | COMMERCIAL | lcpmicrolab@gmail.com | 28 |  |  | GX4 L | 804205 | Unit not in use due to loss communication error. Last xpertcheck 2021 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | D2,D3,D4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6057041037287715457/6057041037287715457_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6057041037287715457/6057041037287715457_signature_31.png
| 2024-10-26 | 2024-10-25 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29669 | MACARE-HBDC | lcpmicrolab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 335 ASSAY: Xpert SARS-CoV-2, COUNT: 718 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX4 D | 838837 | Due for maintenance B1 and B4 disabled failed xpertcheck reported already B3 encounter temp error for reporting | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B2 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6057039557286395566/6057039557286395566_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6057039557286395566/6057039557286395566_signature_31.png
| 2024-10-25 | 2024-10-25 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29668 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 |  |  | 6C GXIV-4D | 838285 | None | None | -Preventive Maintenance -Archive and purge tests earlier than 2024 | All modules are functional. | Waiting for module A1 replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6056432523818926558/6056432523818926558_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/6056432523818926558/6056432523818926558_signature_31.png
| 2024-10-25 | 2024-10-25 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29667 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2940 ASSAY: Xpert HIV-1 Viral Load, COUNT: 405 |  | 6C GXIV-4D | 838285 | Non | None | -Preventive Maintenance -Archive and purges tests later than 2024 | All modules are functional. Waiting for module A1 replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6056430123815756364/6056430123815756364_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/6056430123815756364/6056430123815756364_signature_31.png
| 2024-10-25 | 2024-10-24 | Madrid District Hospital | Poblacion, Madrid | 13 | Surigao del Sur | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29666 | DOH | madrid_dh@gmail.com | 27 |  |  | GX-4 | 838439 | None | None | XpertCheck Calibration Preventive Maintenance | All modules passed in XpertCheck. No errors found. All units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6056318158153732492/6056318158153732492_signature_28.png | Rochelle Pueblas | https://www.jotform.com/uploads/macarewecare/222927978372067/6056318158153732492/6056318158153732492_signature_31.png
| 2024-10-25 | 2024-10-22 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29665 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 13 ASSAY: Xpert CT/NG, COUNT: 52 ASSAY: Xpert HPV, COUNT: 59 ASSAY: Xpert MTB/RIF Ultra, COUNT: 109 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 21 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A1/210124694 has error 4016, modules A3/866793 has error 1001/2014, for reporting. Modules A3 and A4 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6056253683616860222/6056253683616860222_signature_28.png | Morillo, Gerald | https://www.jotform.com/uploads/macarewecare/222927978372067/6056253683616860222/6056253683616860222_signature_31.png
| 2024-10-25 | 2024-10-24 | The Medical City | Ortigas Avenue. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | YES | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29663 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 56 ASSAY: Xpert CT/NG, COUNT: 68 ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 590 ASSAY: Xpert HPV, COUNT: 423 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5577 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 87 ASSAY: Xpert HIV-1 Viral Load, COUNT: 714 ASSAY: Xpress SARS-CoV-2_Flu_RSV plus, COUNT: 1235 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6056253683617393648/6056253683617393648_signature_28.png | Janica Katheryn Zamora | https://www.jotform.com/uploads/macarewecare/222927978372067/6056253683617393648/6056253683617393648_signature_31.png
| 2024-10-25 | 2024-10-22 | St. Luke's Medical Center - BGC | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29664 | COMMERCIAL | ci.bgc@stlukes.com.ph | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 65 ASSAY: Xpert CT/NG, COUNT: 270 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 327 ASSAY: Xpert MTB/RIF Ultra, COUNT: 712 ASSAY: Xpert HCV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 51 |  | GX-IV-D-10C | 120002383 | For maintenance | performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6056253683618554289/6056253683618554289_signature_28.png | Morillo, Gerald P | https://www.jotform.com/uploads/macarewecare/222927978372067/6056253683618554289/6056253683618554289_signature_31.png
| 2024-10-24 | 2024-10-24 | St Lukes Extension Clinic | J Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 29662 | MACARE-HBDC | laboratory@slec.ph | 22 |  |  | Infinity 48s | 809643 | Connection to LIS server not working | Cepheid Tech Support troubleshooted GeneXpert software settings | Coordinated remote troubleshooting by Cepheid Tech Support | Unit is operational, LIS functionality is working again |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6055664763219843838/6055664763219843838_signature_28.png | MA. SARAH R. SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/6055664763219843838/6055664763219843838_signature_31.png
| 2024-10-24 | 2024-10-24 | Philippine Heart Center | East Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29661 | MACARE-HBDC | phc_bb@yahoo.com | 27 |  | 2024-10-23 | MCS+ | 11H204 | For replacement of rollers | Replaced and installed | Performed repair | Unit operational |  | Urethane roller assembly | 12270 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6055617450017061932/6055617450017061932_signature_28.png | CLAIRE JADSAC | https://www.jotform.com/uploads/macarewecare/222927978372067/6055617450017061932/6055617450017061932_signature_31.png
| 2024-10-24 | 2024-10-24 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | dyrosemarie71@gmail.com |  | 29660 | DOH | dy_rosemarie71@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1400 |  | GX-IV 6C 4D | 838401 | UPS problem and For Preventive Maintenance | Check Line Voltages of UPS, change the fuse and replace with new fuse Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform UPS Troubleshooting and Preventive Maintenance | Machine is operational, UPS is working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6055536000751026567/6055536000751026567_signature_28.png | Rosemarie Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/6055536000751026567/6055536000751026567_signature_31.png
| 2024-10-24 | 2024-10-24 | Liloy rural health unit | Liloy, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | YES | CALIBRATION MAINTENANCE GX barcode sticker installation | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29659 | DOH | gisellegipulan75@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1421 | 2024-09-05 | GX 4 (6c) | 838536 | None | Module 1, 2 and 4 undergone Xpertcheck calibration while module 3 skipped due recently replaced module. Genexpert Machine undergone preventive Maintenence and installed GX barcode sticker. | Module 1, 2 and 4 undergone Xpertcheck calibration while module 3 skipped due recently replaced module. Genexpert Machine undergone preventive Maintenence and installed GX barcode sticker. | Module 1, 2 and 3 are available while module 4 is for replacement | None | XPERTCHECK-CE-5 | 12035 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6055454190711798360/6055454190711798360_signature_28.png | Giselle M.Gipulan | https://www.jotform.com/uploads/macarewecare/222927978372067/6055454190711798360/6055454190711798360_signature_31.png
| 2024-10-24 | 2024-10-24 | LAPU LAPU CITY HEALTH OFFICE | LAPULAPU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29658 | DOH | tanya_pax@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2590 |  | GX-IV | 837661 | PREVENTIVE MAINTENANCE | COLLECT IQR,SYSTEM LOG, AND ARCHIVE | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6055435033837122285/6055435033837122285_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/6055435033837122285/6055435033837122285_signature_31.png
| 2024-10-23 | 2024-10-23 | Philippine Red Cross - Davao Chapter | Roxas Boulevard, Davao City | 11 | Davao del Sur | BloodBank | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 29657 | COMMERCIAL | davaocitybloodservices@gmail.com | 25.8 |  |  | MCS+ | 02J021 | Error during last procedure regarding on return flow | Clean Surface Area, Check rollers, check sensor parameters, Perform Cycle Test | Perform troubleshooting and Maintenance | Successfully finished the cycletest, Unit is operational Visble threading on the rollers | Request roller replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6054870925217702649/6054870925217702649_signature_28.png | Arlene dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6054870925217702649/6054870925217702649_signature_31.png
| 2024-10-23 | 2024-10-23 | Piñan Rural health unit | Piñan, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29656 | UNOPS | ivannagmarchan@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5 | 2024-10-23 | GX 4 (10) | 110023516 | Printer black cartridge gives unstable output. | Site inspection follow up, checked room and Module temperatures, Preventive maintenence and Gx barcode sticker installation | Site inspection follow up, checked room and Module temperatures, Preventive maintenence and Gx barcode sticker installation | All modules are available and ready to use. | Printer black cartridge under observation. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6054587571114255805/6054587571114255805_signature_28.png | Ivanna G. Marchan | https://www.jotform.com/uploads/macarewecare/222927978372067/6054587571114255805/6054587571114255805_signature_31.png
| 2024-10-23 | 2024-10-23 | Piñan Rural health unit | Piñan, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION MAINTENANCE GX barcode sticker installation | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29655 | UNOPS | ivannagmarchan@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5 | 2024-10-23 | GX 4 (6c) | 110023516 | Printer black cartridge gives unstable output. | Site inspection follow up, checked room and Module temperatures, Preventive maintenence and Gx barcode sticker installation | Site inspection follow up, checked room and Module temperatures, Preventive maintenence and Gx barcode sticker installation | All modules are available and ready to use. | Black cartridge under observation | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6054587571114255805/6054587571114255805_signature_28.png | Ivanna G. Marchan | https://www.jotform.com/uploads/macarewecare/222927978372067/6054587571114255805/6054587571114255805_signature_31.png
| 2024-10-23 | 2024-10-22 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29654 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 51 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2462 ASSAY: Xpert SARS-CoV-2, COUNT: 1406 ASSAY: Xpert SARS-CoV-2, COUNT: 1209 |  | GX16 | 110014171 | Encounter continous error 5007 in different module. | Check unit Restart system Cleaned plunger | Troubleshooting and on site assistance | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6054535663618438533/6054535663618438533_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/6054535663618438533/6054535663618438533_signature_31.png
| 2024-10-22 | 2024-10-22 | Bureau of Plant Industry | Malate Manila, Metro Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29653 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 25 |  |  | Axio Lab.A1 | 3135010036 | N/A | N/A | Preventive Maintenance and Calibration | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949269101656534/6053949269101656534_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949269101656534/6053949269101656534_signature_31.png
| 2024-10-22 | 2024-10-22 | Bureau of Plant Industry | Malate Manila, Metro Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29652 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 25 |  |  | Stemi 305 Trino | 3947010025 | N/A | N/A | Preventive Maintenance and Calibration | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949269101979254/6053949269101979254_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949269101979254/6053949269101979254_signature_31.png
| 2024-10-22 | 2024-10-22 | Bureau of Plant Industry | Malate Manila, Metro Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29650 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 25 |  |  | Primo Star | 3116003729 | 40x Objective needs heavy maintenance | N/A | Preventive Maintenance and Calibratipn | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949269104386879/6053949269104386879_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949269104386879/6053949269104386879_signature_31.png
| 2024-10-22 | 2024-10-22 | Bureau of Plant Industry | Malate Manila, Metro Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29651 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 25 |  |  | Primo Star | 3116006880 | Wobbly stand, needs rubber stopper or alike | N/A | Preventive Maintenance and Calibration | Unit is operational | Reccomendation: Blue Filter |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949269104683106/6053949269104683106_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949269104683106/6053949269104683106_signature_31.png
| 2024-10-22 | 2024-10-22 | Bureau of Plant Industry | Malate Manila, Metro Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29647 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 25 |  |  | Stemi 305 Trino | 3947010021 | Gooseneck Illumination: loose handle of Left Side | N/A | Preventive Maintenance and Calibration | Unit is operational | Recommendation: Replacement of gooseneck illumination |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209101098980/6053949209101098980_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209101098980/6053949209101098980_signature_31.png
| 2024-10-22 | 2024-10-22 | Bureau of Plant Industry | Malate, Manila, Metro Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29645 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 25 |  | 2024-10-17 | Stemi 305 Trino | 3947020968 | N/A | N/A | Installation | Unit is operational |  | N/A | DR No. 12232 & SI No. 98947 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209104261242/6053949209104261242_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209104261242/6053949209104261242_signature_31.png
| 2024-10-22 | 2024-10-22 | Bureau of Plant Industry | Malate Manila, Metro Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.orignes@gmail.com |  |  | 29648 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 25 |  |  | Primo Star iLED | 3134006922 | 100x objective needs heavy maintenance | None | Preventive Maintenance and Calibration | Unit is operational | Recommendation for Blue Filter |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209109141845/6053949209109141845_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209109141845/6053949209109141845_signature_31.png
| 2024-10-22 | 2024-10-22 | Bureau of Plant Industry | Malte, Manila, Metro Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29646 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 25 |  | 2024-10-17 | Stemi 305 Trino | 3947020967 | N/A | N/A | Installation | Unit is operational |  | N/A | DR No. 12232 & SI No. 98947 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209109373993/6053949209109373993_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209109373993/6053949209109373993_signature_31.png
| 2024-10-22 | 2024-10-22 | Bureau of Plant Industry | Malate Manila, Metro Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29649 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 25 |  |  | Primo Star iLED | 3134006954 | Flickering light bulb, recommendation to observe | N/A | Preventive Maintenance and Calibration | Unit is operational | Recommendation: Blue Filter |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209109642805/6053949209109642805_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053949209109642805/6053949209109642805_signature_31.png
| 2024-10-22 | 2024-10-22 | Manila Central University Hospital | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29644 | CONSORTIUM | purchasing@mcufdtmf.org | 19 |  |  | GXIV-10C | 110018008 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053896906813588662/6053896906813588662_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6053896906813588662/6053896906813588662_signature_31.png
| 2024-10-22 | 2024-10-22 | Molave Rural Health unit | Molave, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29643 | DOH | rhumolavelaboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 612 | 2024-09-05 | GX 4 (6c) | 838233 | None | Modules 1, 2 and 4 undergone Xpertcheck calibration while module 3 skipped due to recent module Replacement. | Modules 1, 2 and 4 undergone Xpertcheck calibration while module 3 skipped due to recent module Replacement. | All modules are available and ready to use. Modules 1,2 and 4 passed while module 3 skipped the xpertcheck calibration. | None | XPERTCHECK-CE-5 | 12037 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053841962315521143/6053841962315521143_signature_28.png | Mariah Christine Jane S.Jungco | https://www.jotform.com/uploads/macarewecare/222927978372067/6053841962315521143/6053841962315521143_signature_31.png
| 2024-10-22 | 2024-10-17 | Columbio Rural Health Unit | Poblacion, Columbio, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29642 | DOH | rochelle.toledo19@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-10-17 | GxIV-6C | 838556 | Module 3 for Service Required | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert machine Preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: Module 1 - Passed Module 2 - Passed Module 3 - Service required Module 4 - Passed |  | Xpertcheck cartridge/kit Lot#: 1001075092 | DR#: 12017 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053572436864585085/6053572436864585085_signature_28.png | Rochelle S. Toledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6053572436864585085/6053572436864585085_signature_31.png
| 2024-10-22 | 2024-10-18 | Tampakan Rural Health Unit | Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29641 | DOH | lotisglegada@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 617 | 2024-10-18 | GxIV-6C | 838544 | Module 2 for Service Required Module 3 for Service Required | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert cleaning and maintenance Checking of errors Errors reporting Data collection | 3 operational modules Calibration Results: Module 1 - Passed Module 2 - Service required Module 3 - Service required Module 4 - Skipped (new module) |  | Xpertcheck cartridge/kits Lot#:1001075092 | DR#:12026 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053571736863841119/6053571736863841119_signature_28.png | Lots G. Legada | https://www.jotform.com/uploads/macarewecare/222927978372067/6053571736863841119/6053571736863841119_signature_31.png
| 2024-10-21 | 2024-10-18 | Chong Hua Hospital- Mandaue | Subangdaku , Mandaue City Cebu | 7 | Cebu | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29640 | COMMERCIAL | irl.maderazo@gmail.com | 24 |  |  | MCS+ | 11J217 | Safety System error # 68 state fault. | Check Blood pump return and draw pump, ACD pump. Check supply voltage of UPS. Replaced blood pump. | Check Blood pump return and draw pump, ACD pump. Check supply voltage of UPS. | Unit apheresis machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053108547736220857/6053108547736220857_signature_28.png | Inna Maderazo,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6053108547736220857/6053108547736220857_signature_31.png
| 2024-10-21 | 2024-10-17 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29639 | COMMERCIAL | vsmmcbloodservices@gmail.com | 24 |  |  | MCS+ | 02A090 | For Maintenance and check MCS+ | Check for blood pump and ACD pump. Check rubber roller of the pump, check pneumatic assembly of apheresis machine. Check supply voltage of the UPS. | Check for blood pump and ACD pump. Check rubber roller of the pump, check pneumatic assembly of apheresis machine. Check supply voltage of the UPS. | Unit MCS+ is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053103747735615239/6053103747735615239_signature_28.png | Charlotte Mae Ygoña, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6053103747735615239/6053103747735615239_signature_31.png
| 2024-10-21 | 2024-10-16 | Chong Hua Hospital- Fuente | Fuente Osmena, Capitol Site, Cebu City | 7 | Cebu | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29638 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  | 2024-10-16 | MCS+ | 02F035 | Check for Maintenance of blood pump and ACD pump. | Check for blood pump and ACD pump. Check rubber roller of the pump, check pneumatic assembly of apheresis machine. Check supply voltage of the UPS. | Check for blood pump and ACD pump. Check rubber roller of the pump, check pneumatic assembly of apheresis machine. Check supply voltage of the UPS. | The unit machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053100307735069828/6053100307735069828_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6053100307735069828/6053100307735069828_signature_31.png
| 2024-10-21 | 2024-10-14 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Tb laboratory | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29637 | COMMERCIAL | cduhlabresults2@yahoo.com | 24 |  | 2024-10-14 | MCS+ | 14F014 | Safety error #68 during procedure. | Replace pump of the ACD and Blood pump of the motor. Check power voltage of the UPS. | Troubleshoot problem check blood pump and ACD pump. Restart procedure and orient end-user during the next occurrence of error. | Unit of MCS+ are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6053095027735877451/6053095027735877451_signature_28.png | Eula Quijada, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6053095027735877451/6053095027735877451_signature_31.png
| 2024-10-21 | 2024-10-16 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | rpescador@asianhospital.com |  | 29636 | COMMERCIAL | agalang@asianhospital.com | 22 | ASSAY: Xpert CT/NG, COUNT: 9 ASSAY: Others, COUNT: 85 ASSAY: Xpert HBV VL, COUNT: 77 ASSAY: Xpert HPV, COUNT: 177 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1089 ASSAY: Xpert HIV-1 Viral Load, COUNT: 56 |  | GXIV-D-10C | 110014996 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Cleaned Filter. Performed Self-test. Checked System Functionality. Generate Reports. | GeneXpert System is in Good Condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913613158239/6052995913613158239_signature_28.png | Alyssa Galang | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913613158239/6052995913613158239_signature_31.png
| 2024-10-21 | 2024-10-09 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29632 | MACARE-HBDC | fabella.mollabgene@gmail.com | 22 |  |  | GXIV-L-6C | 834953 | Module 2 with Test Errors. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Cleaned Filter. Performes Self-test. Checked System Functionality. Generate Reports. | Module 2 for further observation after PMs. 2 Modules are Defective. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913613862008/6052995913613862008_signature_28.png | Rallion Gio Paguia, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913613862008/6052995913613862008_signature_31.png
| 2024-10-21 | 2024-10-09 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29634 | MACARE-HBDC | fabella.mollabgene@gmail.com | 22 |  |  | GXIV-D-10C | 110015482 | Module 3 with 5007 Error. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Cleaned Filter. Performes Self-test. Checked System Functionality. Generate Reports. | Module 3 for further observation after PMs. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913615350003/6052995913615350003_signature_28.png | Rallion gio paguia, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913615350003/6052995913615350003_signature_31.png
| 2024-10-21 | 2024-10-17 | Seamen's Hospital | Intramuros, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29635 | MACARE-HBDC | shmolecularlab@gmail.com | 23 | ASSAY: Xpert SARS-CoV-2, COUNT: 85 |  | GX16-D-6C | 838777 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System Functionality. Generate Reports. | 14 Modules are Operational. 2 Modules were Disabled (with Errors) | Unit is due for XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913617778220/6052995913617778220_signature_28.png | Queenie Lea Serquina | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913617778220/6052995913617778220_signature_31.png
| 2024-10-21 | 2024-10-17 | Seamen's Hospital | Intramuros, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29633 | MACARE-HBDC | shmolecularlab@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 5 |  | GXIV-D-6C | 838836 | Unit is due for XpertCheck | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Checked System Functionality. Generate Reports. | All Modules are Operational. | Unit is due for XpertCheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913618666983/6052995913618666983_signature_28.png | Queenie Lea Serquina | https://www.jotform.com/uploads/macarewecare/222927978372067/6052995913618666983/6052995913618666983_signature_31.png
| 2024-10-21 | 2024-10-18 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29631 | DOH | allan.vinluan2013@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 974 |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052966081038261485/6052966081038261485_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6052966081038261485/6052966081038261485_signature_31.png
| 2024-10-21 | 2024-10-18 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29630 | DOH | seansimonpaloyo@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1268 |  | GXIV-4-D-6C | 838473 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is Fully Operational For additional Troubleshooting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052965261031961646/6052965261031961646_signature_28.png | Manolo C. Paloyo | https://www.jotform.com/uploads/macarewecare/222927978372067/6052965261031961646/6052965261031961646_signature_31.png
| 2024-10-21 | 2024-10-18 | Antipolo City Health Office | M. Santos St. Antipolo City | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29628 | UNOPS | acho.laboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 82 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052804707021739582/6052804707021739582_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6052804707021739582/6052804707021739582_signature_31.png
| 2024-10-21 | 2024-10-17 | Kalinga Health Hub - Marikina | JP Rizal St. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29627 | CONSORTIUM | bryllcruz08@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1576 | 2024-10-16 | GX-IV-D-6C | 110008418 | Module B3 for replacement | Replaced module B3 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B3 Generated reports | Unit is operational. Module B1/678461 has error 1002, for reporting. Modules B2, B3, and B4 operational. |  | Old: B3/616087 New: B3/210012781 | DR# 12223 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052804707023921485/6052804707023921485_signature_28.png | Lorenzo Bryll Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6052804707023921485/6052804707023921485_signature_31.png
| 2024-10-21 | 2024-10-18 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29629 | COMMERCIAL | mahmci_laboratory@yahoo.com | 22 | ASSAY: Xpert HBV VL, COUNT: 18 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 270 ASSAY: Xpert HIV-1 Viral Load, COUNT: 11 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052804707024313249/6052804707024313249_signature_28.png | Sarah Lee Almeria | https://www.jotform.com/uploads/macarewecare/222927978372067/6052804707024313249/6052804707024313249_signature_31.png
| 2024-10-21 | 2024-10-18 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29626 | MACARE-HBDC | phc_bb@yahoo.com | 27 |  |  | MCS + | 11H204 | Due for maintenance Roller for replacement yellowish color | Performed maintenance Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | For replacement of rollers |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052794853613068158/6052794853613068158_signature_28.png | Cayle Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/6052794853613068158/6052794853613068158_signature_31.png
| 2024-10-21 | 2024-10-18 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29625 | MACARE-HBDC | phc_bb@yahoo.com | 28 |  |  | MCS + | 20168 | Due for maintenance Roller for replacement yellowish color | Performed maintenance Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | For replacement of rollers |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052794443611375952/6052794443611375952_signature_28.png | Cayle Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/6052794443611375952/6052794443611375952_signature_31.png
| 2024-10-21 | 2024-10-14 | Medilinx (MMC) | Amorsolo st, Makati City | NCR | NCR 4th - Makati | Molecular | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | John.Manzano@sodexo.com |  | 29624 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 221 ASSAY: Xpert CT/NG, COUNT: 824 ASSAY: Xpert HBV VL, COUNT: 1 ASSAY: Xpert HPV, COUNT: 758 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3191 ASSAY: Others, COUNT: 13 ASSAY: Xpert HCV VL, COUNT: 88 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1956 |  | GX4 D (10C) | 110001647 | For relocation | Relocate and re install unit Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed relocation and maintenance | B1 disabled B2,B3,B4 in use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052794083619635966/6052794083619635966_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6052794083619635966/6052794083619635966_signature_31.png
| 2024-10-21 | 2024-10-15 | Saragani Islands Rural Health Unit | Poblacion Mabila, Saragani Islands Davao Occidental | 11 | Davao Occidental | TB-DOTS | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29623 | DOH | lovitosquennie26@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-09-04 | GX-IV 6C 4D | 837614 | Cpu system unit not working For Preventive Maintenance and XpertCheck Calibration For GX Sticker Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Change Cmos Battery Clean ram slots Reset network and gateway board connection Run XPC of Four Modules Install GX Sticker | Perform CPU System Unit Troubleshooting Perform Preventive Maintenance, and XpertCheck Calibration Install GX Sticker | All Modules are Passed the Calibration CPU System Unit successfully repair All Modules are Operational |  | Xpertcheck CE 5 | DR 12010 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052521499822134024/6052521499822134024_signature_28.png | Queenie Lovitos,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6052521499822134024/6052521499822134024_signature_31.png
| 2024-10-20 | 2024-10-18 | Bataan General Hospital and Medical Center | Balanga City, Bataan | 3 | Bataan | RHIVDA | NO | INSTALLATION | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29622 | MACARE-HBDC | mariaana.hizon@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 ASSAY: Xpert HBV VL, COUNT: 0 ASSAY: Xpert HCV VL, COUNT: 0 ASSAY: Xpert HPV, COUNT: O0 | 2024-10-16 | GXIV-D 10C | 120003587 | For installation | Installation | Installation ad setup of Genexpert machine | Unit is operational |  | Genexpert machine | 98942 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6052023686183982820/6052023686183982820_signature_28.png | Maria Ana Hizon | https://www.jotform.com/uploads/macarewecare/222927978372067/6052023686183982820/6052023686183982820_signature_31.png
| 2024-10-18 | 2024-10-18 | Hi-Precision Diagnostics | NS Amoranto Sr. St., Queozn City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29621 | CONSORTIUM | jose@hi-precision.com.ph | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 139 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6410 |  | GX IV 6C-D | 835601 | Module B2 and B3 disabled due to error | Performed Preventive maintenance Performed reagent run No error occurred Generated report | View tests Checked and cleaned cartridge bays, plunger rod, icore slot and filter Archived results Performed reagent run twice on 2 modules | unit is operational all modules working B2 and B3 made available again |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6050513498258561003/6050513498258561003_signature_28.png | Andrei N. Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/6050513498258561003/6050513498258561003_signature_31.png
| 2024-10-18 | 2024-10-18 | Valenzuela Social Hygiene Clinic | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29620 | COMMERCIAL | markanthony.manaois@gmail.com | 25 |  |  | GXIV-10C | 110020699 | Unit for maintenance | Cleaned slightly clogged filter | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6050377923152661228/6050377923152661228_signature_28.png | Mark Manaois | https://www.jotform.com/uploads/macarewecare/222927978372067/6050377923152661228/6050377923152661228_signature_31.png
| 2024-10-18 | 2024-10-18 | University of the Philippines - ICCHD | Pedro Gil St., Ermita Manila | NCR | NCR 1st - Manila | Central lab | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29619 | COMMERCIAL | dcsantos9@up.edu.ph | 25 |  |  | GXIV | 808015 | Unit for Preventive Maintenance. | Preventive Maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6050377078915971160/6050377078915971160_signature_28.png | Dana Kamylle Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/6050377078915971160/6050377078915971160_signature_31.png
| 2024-10-18 | 2024-10-18 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29618 | DOH | buhanginlaboratory@gmail.com | 25.6 |  |  | GX-IV 6C 4D | 838232 | System Unit not working | Change cmos battery, change power chords, reset connections, clean ram and slots, and check for line voltages | Perform System Unit Troubleshooting | CPU System Unit not working, documented the error and reported to Cepheid for further assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6050372864768050820/6050372864768050820_signature_28.png | Catherine B. Deocampo,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6050372864768050820/6050372864768050820_signature_31.png
| 2024-10-18 | 2024-10-18 | University of the Philippines - National Institutes of Health | Pedro Gil St., Ermita Manila | NCR | NCR 1st - Manila | Central lab | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | vvluzentales@up.edu.ph |  | 29617 | COMMERCIAL | nihcentrallab@gmail.com | 28 | ASSAY: Others, COUNT: Xpert Carba-R = 122 ASSAY: Xpert CT/NG, COUNT: 194 ASSAY: Xpert HBV VL, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Others, COUNT: Xpert MTB-XDR = 1 ASSAY: Xpert HCV VL, COUNT: 53 ASSAY: Xpert HIV-1 Viral Load, COUNT: 430 |  | GX16-10C | 110014186 | Unit for Preventive Maintenance. | Preventive Maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6050328118915213392/6050328118915213392_signature_28.png | Mary Grace Hernaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6050328118915213392/6050328118915213392_signature_31.png
| 2024-10-18 | 2024-10-18 | Philippine General Hospital | Taft. Ave., Ermita Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | adcanseco@up.edu.ph |  | 29616 | COMMERCIAL | eomercado@up.edu.ph | 25 | ASSAY: Xpert HPV, COUNT: 99 ASSAY: Others, COUNT: Xpert SA Nasal Complete G3 = 1 ASSAY: Xpert SARS-CoV-2, COUNT: 795 ASSAY: Others, COUNT: Xpress SARS-CoV-2_Flu_RSV plus = 275 |  | GX16 | 844975 | Unit for Preventive Maintenance. | Preventive Maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. For schedule of Calibration upon purchase of Xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6050299578916512246/6050299578916512246_signature_28.png | Erma Grace Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/6050299578916512246/6050299578916512246_signature_31.png
| 2024-10-18 | 2024-10-18 | GUIMARAS PROVINCIAL HEALTH OFFICE | GUIMARAS, ILOILO | 6 | Guimaras | LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29615 | UNOPS | mylene210@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 300 ASSAY: Others, COUNT: XDR -1 |  | GX-IV | 110023528 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6050239103599847772/6050239103599847772_signature_28.png | Mylene Publico | https://www.jotform.com/uploads/macarewecare/222927978372067/6050239103599847772/6050239103599847772_signature_31.png
| 2024-10-17 | 2024-10-17 | NUEVA VALENCIA RURAL HEALTH UNIT | NUEVA VALENCIA, GUIMARAS | 6 | Guimaras | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29614 | DOH | xxxrhingringxxx@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 373 |  | GX-IV | 838291 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049627433545766150/6049627433545766150_signature_28.png | Rhea April Q Jaleco | https://www.jotform.com/uploads/macarewecare/222927978372067/6049627433545766150/6049627433545766150_signature_31.png
| 2024-10-17 | 2024-10-17 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29613 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 807112 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049533233155825628/6049533233155825628_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6049533233155825628/6049533233155825628_signature_31.png
| 2024-10-17 | 2024-10-17 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29612 | MACARE-HBDC | laboratory@loveyourself.com.ph | 27 |  |  | GXIV | 807956 | Unit for maintenance Module A2 for check up | NA | #NAME? | Unit is operational A2 under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049532293157873871/6049532293157873871_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/6049532293157873871/6049532293157873871_signature_31.png
| 2024-10-17 | 2024-10-17 | Magsaysay Rural Health Unit | Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29611 | DOH | jkakilgan12@gmail.com | 26 |  |  | GX-4 6 Colors | 838557 | Module B2 found Service Required | XPC Calibration and Preventive Maintenance | XPC Calibration and Preventive Maintenance | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049531733426479466/6049531733426479466_signature_28.png | Jesusa Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/6049531733426479466/6049531733426479466_signature_31.png
| 2024-10-17 | 2024-10-17 | Riverside Medical Center, Inc. | BS Aquino Drive, Bacolod City | 6 | Negros Occidental | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29610 | COMMERCIAL | rjmanito@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 22 ASSAY: Xpert HBV VL, COUNT: 169 ASSAY: Xpert HPV, COUNT: 37 ASSAY: Xpert MTB/RIF Ultra, COUNT: 493 ASSAY: Xpert HCV VL, COUNT: 49 |  | GX-4 | 110005268 | - B3 - operational; - B1, B2, B4 - temperature -related errors during specimen runs | #NAME? | #NAME? | B1- encountering error E2014, hardware-related, thermistor problem, for report to Cepheid Tech Support. B2 - encoutering errors E1002 and E4009, hardware-related, thermistor problem, for report to Cepheid Tech Support. B3 - operational B4 - operational, but encounting intermittent temperature-related errors. recommendation: perform Xpertcheck on this module. | Unit due for Xpertcheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049521860439510727/6049521860439510727_signature_28.png | Rumary Joan Manito | https://www.jotform.com/uploads/macarewecare/222927978372067/6049521860439510727/6049521860439510727_signature_31.png
| 2024-10-17 | 2024-10-15 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29609 | DOH | dorothygundaya@gmail.com | 30 |  |  | GX-4 6 Colors | 838403 | Module B2 found Service Required thru XPC Calibration | Calibration | XPC Calibration and Preventive Maintenance | 3 modules working, 1 module for replacement (Error in XPC) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049513443429472876/6049513443429472876_signature_28.png | Dorothy Gundaya | https://www.jotform.com/uploads/macarewecare/222927978372067/6049513443429472876/6049513443429472876_signature_31.png
| 2024-10-17 | 2024-10-17 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29607 | MACARE-HBDC | amymitra22@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 323 ASSAY: Others, COUNT: 23 ASSAY: Xpert HIV-1 Viral Load, COUNT: 98 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049507853613049975/6049507853613049975_signature_28.png | Amelia Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/6049507853613049975/6049507853613049975_signature_31.png
| 2024-10-17 | 2024-10-17 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29608 | MACARE-HBDC | amymitra22@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1186 ASSAY: Others, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049507853613096910/6049507853613096910_signature_28.png | Amelia Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/6049507853613096910/6049507853613096910_signature_31.png
| 2024-10-17 | 2024-10-16 | Makati Medical Center | Amorsolo St., Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29606 | COMMERCIAL | reinaluz.bactol@medilinx.com.ph | 24 |  |  | GX-IV-D-10C | 110001647 | For relocation | Performed maintenance Transferred the GeneXpert system | Transferred the unit to other room Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049507853619152649/6049507853619152649_signature_28.png | Lois Zoe Montemayor | https://www.jotform.com/uploads/macarewecare/222927978372067/6049507853619152649/6049507853619152649_signature_31.png
| 2024-10-17 | 2024-10-17 | PHILIPPINE RED CROSS - ILOILO CITY | Iloilo City | 6 | Iloilo | BLOOD BANK | NO | MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29605 | COMMERCIAL | aprilys014@gmail.com | 22 |  |  | MCS+ | 98G049 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6049437027912626604/6049437027912626604_signature_28.png | April Rose Yso | https://www.jotform.com/uploads/macarewecare/222927978372067/6049437027912626604/6049437027912626604_signature_31.png
| 2024-10-16 | 2024-10-16 | San Roque Rural Health Unit | San Roque Northern Samar | 8 | Northern Samar | TB Laboratory | NO | REPAIR | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29604 | DOH | lynnjarito@yahoo.com | 29 |  | 2024-10-16 | GXIV | 838498 | Module C3/862236 Defective | #NAME? | -Performed preventive maintenance -Install new module on C3 w/ SN:620877 - Generate IQR and System log - Pullout defective module | Unit is Operational |  | 1 Genexpert Module 6C | 12076 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6048569888168226631/6048569888168226631_signature_28.png | LYNN N JARITO | https://www.jotform.com/uploads/macarewecare/222927978372067/6048569888168226631/6048569888168226631_signature_31.png
| 2024-10-16 | 2024-10-16 | Maluso Rural Health Unit | Maluso, Basilan | BARMM | Basilan | Laboratory | NO | CALIBRATION | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29603 | DOH | sharminahussin@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 576 | 2024-09-05 | GX 4 | 838835 | None | Module 1, 2, 3 and 4 undergone Xpertcheck calibration. | Module 1, 2, 3 and 4 undergone Xpertcheck calibration. All modules passed the Calibration. | All modules are available | None | XPERTCHECK-CE-5 | 12043 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6048545762266129476/6048545762266129476_signature_28.png | Sharmina A. Hussin | https://www.jotform.com/uploads/macarewecare/222927978372067/6048545762266129476/6048545762266129476_signature_31.png
| 2024-10-16 | 2024-10-15 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | lleerrssoonn@gmail.com |  | 29602 | MACARE-HBDC | Cabacasmelvin@gmail.com | 22 |  |  | MCS+ | 15H045G | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6048488043618425792/6048488043618425792_signature_28.png | Cabacas, Melvin D. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6048488043618425792/6048488043618425792_signature_31.png
| 2024-10-15 | 2024-10-15 | Saragani Islands Rural Health Unit | Poblacion Mabila, Saragani Islands Davao Occidental | 11 | Davao Occidental | TB-DOTS | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29601 | DOH | lovitosquennie26@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-09-04 | GX-IV 6C 4D | 837614 | Cpu system unit not working For Preventive Maintenance and XpertCheck Calibration For GX Sticker Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Change Cmos Battery Clean ram slots Reset network and gateway board connection Run XPC of Four Modules Install GX Sticker | Perform CPU System Unit Troubleshooting Perform Preventive Maintenance, and XpertCheck Calibration Install GX Sticker | All Modules passed the Calibration and are operational |  | Xpertcheck CE 5 | DR 12010 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6047869444703629128/6047869444703629128_signature_28.png | Quennie Lovitos,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6047869444703629128/6047869444703629128_signature_31.png
| 2024-10-15 | 2024-10-15 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Microbiology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | marijo.siat@medilinx.com.ph |  | 29600 | COMMERCIAL | aivee.alcantara@medilinx.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 64 ASSAY: Xpert HPV, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 710 |  | GXII | 812916 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. Next preventive maintenance is on January 2025. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6047840052285830258/6047840052285830258_signature_28.png | Aivee Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/6047840052285830258/6047840052285830258_signature_31.png
| 2024-10-15 | 2024-10-15 | Tanza Wellness Center | A. Soriano Highway, Tanza, Cavite | 4A | Cavite | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29599 | MACARE-HBDC | tanzawellnesscenter.rhu@gmail.com | 27 |  |  | GXIV-D-10C | 120003582 | Unit is for Installation. | Installation. | Checked Facility Setup. Install GeneXpert System. Generate Reports. | GX System is in Good Condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6047817786091554898/6047817786091554898_signature_28.png | Sheradee Sorino | https://www.jotform.com/uploads/macarewecare/222927978372067/6047817786091554898/6047817786091554898_signature_31.png
| 2024-10-15 | 2024-10-15 | OTON MUNICIPAL HEALTH OFFICE | OTON, ILOILO | 6 | Iloilo | LABORATORY | NO | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29598 | UNOPS | gracie05191980@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-IV | 110023530 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6047749887013927720/6047749887013927720_signature_28.png | Mary Grace Ong-Lumasag | https://www.jotform.com/uploads/macarewecare/222927978372067/6047749887013927720/6047749887013927720_signature_31.png
| 2024-10-15 | 2024-10-15 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29597 | DOH | calambatbdots@gmail.com | 25 |  |  | GXIV | 838225 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6047730865892138780/6047730865892138780_signature_28.png | Lily Banalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6047730865892138780/6047730865892138780_signature_31.png
| 2024-10-15 | 2024-10-11 | Ospital ng Muntinlupa | Centennial Rd., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29596 | UNOPS | jjpsanchez21@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 110 |  | GXIV-D-10C | 110023494 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | GX System is in Good Condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6047709476092289519/6047709476092289519_signature_28.png | Jesse Jasper Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/6047709476092289519/6047709476092289519_signature_31.png
| 2024-10-15 | 2024-10-15 | ST. PAULS HOSPITAL OF ILOILO | ILOILO CITY | 6 | Iloilo | BLOOD BANK | NO | MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29595 | COMMERCIAL | sphi.bloodbank@yahoo.com | 23 |  |  | MCS+ | 98B020 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6047654183333338738/6047654183333338738_signature_28.png | Mariejol G. Dureza | https://www.jotform.com/uploads/macarewecare/222927978372067/6047654183333338738/6047654183333338738_signature_31.png
| 2024-10-14 | 2024-10-14 | Mabinay Medicare Community Hospital | Mabinay,Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29594 | DOH | diannecm8@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 611 |  | GX-4 6C | 838481 | E5015 and E5011 in module B2/769404 | #NAME? | - Reported to Cepheid under case no. 1663597. - Cepheid TS requested to run blank cartridge using SR only and report result as troubleshooting to isolate specimen and processing-related issues. Result of blank cartridge - No Error encountered. - Run actual specimens on B2 to determine module performance. - Copy archive (Gxx) file of dummy and specimen runs in module B2/769404. - Generate IQR and System Log. | Module B1 and B3 operational. Module B4 failed Xpertcheck and awaiting new module from Cepheid. No error encountered in module B2 during blank cartridge run. Continue using module B2 and report new errors encountered during specimen runs. Module B2 for followup report to Cepheid Tech Support. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046962608975114707/6046962608975114707_signature_28.png | DIANNE C. MI-OT, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6046962608975114707/6046962608975114707_signature_31.png
| 2024-10-14 | 2024-10-08 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29593 | DOH | thisisformyfamily@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1496 |  | GXIV-4-D-6C | 838286 | For Preventive Maintenance and Troubleshooting | #NAME? | Performed Preventive Maintenance and Troubleshooting | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046953611033851947/6046953611033851947_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/6046953611033851947/6046953611033851947_signature_31.png
| 2024-10-14 | 2024-10-08 | San Jacinto Rural Health Unit | San Jacinto, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29592 | DOH | allen.indong@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 294 |  | GXIV-4-D-6C | 838518 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046952621039583025/6046952621039583025_signature_28.png | Alexander Indong | https://www.jotform.com/uploads/macarewecare/222927978372067/6046952621039583025/6046952621039583025_signature_31.png
| 2024-10-14 | 2024-10-08 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29591 | DOH | arleneabad32@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 693 |  | GXIV-4-D-6C | 838288 | For Preventive Maintenance and Monitor Replacement | #NAME? | Performed Preventive Maintenance and Monitor Replacement | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046951521034868592/6046951521034868592_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/6046951521034868592/6046951521034868592_signature_31.png
| 2024-10-14 | 2024-10-08 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29590 | DOH | dianeruiz147@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 460 |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046949671039443805/6046949671039443805_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6046949671039443805/6046949671039443805_signature_31.png
| 2024-10-14 | 2024-10-14 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29589 | DOH | burauenlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 476 | 2024-10-14 | GXIV | 838551 | n/a | n/a | Performed monthly preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046922822817713521/6046922822817713521_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6046922822817713521/6046922822817713521_signature_31.png
| 2024-10-14 | 2024-10-11 | East Avenue Medical Inc. | East Avenue QC | NCR | NCR 2nd - Quezon City | TB dots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | eamc.biomed@gmail.com |  | 29588 | MACARE-HBDC | noriel_ruiz7@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2895 |  | GX4 (10C) | 110014085 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046767493616890424/6046767493616890424_signature_28.png | Noriel ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6046767493616890424/6046767493616890424_signature_31.png
| 2024-10-14 | 2024-10-11 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | eamc.biomed@gmail.com |  | 29587 | MACARE-HBDC | blooddonorsection@gmail.com | 29 |  |  | MCS + | 15H044 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046766913615923113/6046766913615923113_signature_28.png | Christopher Paul Bernaldo | https://www.jotform.com/uploads/macarewecare/222927978372067/6046766913615923113/6046766913615923113_signature_31.png
| 2024-10-13 | 2024-10-11 | Norala District Hospital | Poblacion, Norala, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29586 | DOH | noralaDHlaboratory@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 988 | 2024-10-11 | GxIV-6C | 838543 | Module 2 for Service Required | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: Module 1 - Passed Module 2 - Service required Module 3 - Passed Module 4 - Passed |  | Xpertcheck cartridge/kits Lot#: 1001075092 | DR#: 12022 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046348546861498846/6046348546861498846_signature_28.png | REZIEL D. SABIDO | https://www.jotform.com/uploads/macarewecare/222927978372067/6046348546861498846/6046348546861498846_signature_31.png
| 2024-10-13 | 2024-10-09 | Kidapawan City Health Office | Kidapawan City, North Cotabato | 12 | Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29584 | DOH | kidcholaboratory@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2009 | 2024-10-09 | GxIV-6C | 838545 | None | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: Module 1 - Passed Module 2 - Passed Module 3 - Passed Module 4 - Passed |  | Xpertcheck cartridge/kits Lot#:1001075092 | DR#: 12019 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046348486864678291/6046348486864678291_signature_28.png | Neil Everson P. Cunanan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6046348486864678291/6046348486864678291_signature_31.png
| 2024-10-13 | 2024-10-08 | Antipas Rural Health Unit | Poblacion, Antipas, Cotabato | 12 | Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29585 | DOH | antipasclinicallaboratory@gmail.com | 22.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-10-08 | GxIV-6C | 838259 | None | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: Module 1 - Passed Module 2 - Passed Module 3 - Passed Module 4 - Passed | All modules passed Xpertcheck Calibration | Xpertcheck cartridge/kit Lot#:1001075092 | DR#: 12014 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046348486865143307/6046348486865143307_signature_28.png | ALVIN A. CODILLA, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6046348486865143307/6046348486865143307_signature_31.png
| 2024-10-13 | 2024-10-11 | Tigan Municipal Health Office | Tigaon, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | ON-SITE ASSISTANCE Gx relocation to temporary facility | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29583 | DOH | jingimperial24@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 848 |  | GX4-6C | 838558 | Gx relocation to temporary facility due to Tigaon MHO building renovation. | GX installed/relocated @Vinagre Barangay Health Center | Moving/Reinstalling Gx to another facility | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6046166610911294931/6046166610911294931_signature_28.png | Maria Jesusa P.Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6046166610911294931/6046166610911294931_signature_31.png
| 2024-10-11 | 2024-10-11 | Lugait Rural Health Unit | Brgy. Masilakon 1, Lugait, Misamis Occidental | 10 | Misamis Occidental | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29582 | DOH | lugaitmaternity@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 709 |  | GX4 | 838562 | -Due for annual calibration. - Module 2 encountered 2014 hardware failure during running. | #NAME? | #NAME? | All modules are operational except Module 2 SN 690358 failed during calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044581079848137960/6044581079848137960_signature_28.png | Rodelyn B. Limocon, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6044581079848137960/6044581079848137960_signature_31.png
| 2024-10-11 | 2024-10-11 | Tomas Cabili Health Center | Tominobo,Tomas Cabili, Iligan City, Lanao del Norte | 10 | Misamis Occidental | Tb laboratory | NO | REPAIR CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29581 | DOH | mariacielovenezuela@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1067 |  | GX4 | 838294 | - Defective module in Module 1 SN # 807847 - Due for calibration in genexpert machine - Due for Preventive Maintenance clean Icore, Plunger, filter - Stucked cartridge in Module 2 and Module 3 | -Replacement of new module in Module 1 SN # 766189 -Performed calibration in genexpert machine -Conduct Preventive Maintenance clean Icore, Plunger, filter -Removal of stucked cartridge in Module 2 and Module 3 | - Replacement of new module in Module 1 SN # 766189 - Performed calibration in genexpert machine - Conduct Preventive Maintenance clean Icore, Plunger, filter - Removal of stucked cartridge in Module 2 and Module 3 | All modules are operational. -B2 failed during calibration. Usable while waiting for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044569419844854681/6044569419844854681_signature_28.png | Maria Cielo C. Venezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/6044569419844854681/6044569419844854681_signature_31.png
| 2024-10-11 | 2024-10-11 | Hinaplanon RTDL | Hinaplanon, Iligan City | 10 | Misamis Occidental | Tb laboratory | NO | REPAIR CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29580 | DOH | Villacortaamy50@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 922 |  | GX4 | 838227 | Replacement of defective module 1 SN # 792036 | Replacement of new module in Module 1 SN # 765568 Performed calibration in genexpert machine Conduct Preventive Maintenance clean Icore, Plunger, filter | Replacement of new module in Module 1 SN # 765568 Performed calibration in genexpert machine Conduct Preventive Maintenance clean Icore, Plunger, filter | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044562749846472666/6044562749846472666_signature_28.png | Amy S. Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6044562749846472666/6044562749846472666_signature_31.png
| 2024-10-11 | 2024-10-11 | Allied Care Experts Medical | Tagbilaran, Bohol | 7 | Bohol | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29579 | CONSORTIUM | laboratory@acemc.com | 26 |  |  | Gx4 | 840969 | N/A | Site visit Fan Filter Cleaning Module bay cleaning Plunger rod maintenance | Site visit Fan Filter Cleaning Module bay cleaning Plunger rod maintenance | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044401296518814123/6044401296518814123_signature_28.png | Dianne Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/6044401296518814123/6044401296518814123_signature_31.png
| 2024-10-11 | 2024-10-11 | Allied Care Experts Medical | Tagbilaran, Bohol | 7 | Bohol | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29578 | COMMERCIAL | laboratory@acemc.com | 26 |  |  | Gx4 | 840969 | N/A | Site visit Fan Filter Cleaning Module bay cleaning Plunger rod maintenance | Site visit Fan Filter Cleaning Module bay cleaning Plunger rod maintenance | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044401296518814123/6044401296518814123_signature_28.png | Dianne Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/6044401296518814123/6044401296518814123_signature_31.png
| 2024-10-11 | 2024-10-11 | The Medical City Clark | Clark City, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | tmcclarkbiomed@themedicalcity.com |  | 29577 | COMMERCIAL | tmci_laboratory@themedicalcity.com | 26 | ASSAY: Xpert CT/NG, COUNT: 14 ASSAY: Xpert HBV VL, COUNT: 64 ASSAY: Xpert HPV, COUNT: 34 ASSAY: Xpert MTB/RIF Ultra, COUNT: 161 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GXIV-D | 839235 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044376654293935845/6044376654293935845_signature_28.png | Jessa Rae Remorta | https://www.jotform.com/uploads/macarewecare/222927978372067/6044376654293935845/6044376654293935845_signature_31.png
| 2024-10-11 | 2024-10-11 | CEBU DOCTORS UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29576 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 75 ASSAY: Xpert CT/NG, COUNT: 211 ASSAY: Xpert HBV VL, COUNT: 200 ASSAY: Xpert HPV, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1096 ASSAY: Xpert HCV VL, COUNT: 167 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-IV | 110008524 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044342893746754866/6044342893746754866_signature_28.png | Alyssa C. Sayson, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6044342893746754866/6044342893746754866_signature_31.png
| 2024-10-11 | 2024-10-11 | Batasan Social Hygiene Clinic | Batasan Rd, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29575 | MACARE-HBDC | prestozatricia@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 17 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HIV-1 Viral Load, COUNT: 534 |  | GXIV 10C D | 110020697 | Unit due for maintenance | Performed Preventive maintenance Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter archived Results | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044342191325194606/6044342191325194606_signature_28.png | Patricia Anne Prestoza | https://www.jotform.com/uploads/macarewecare/222927978372067/6044342191325194606/6044342191325194606_signature_31.png
| 2024-10-11 | 2024-10-11 | Cebu City Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29574 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1857 |  | 6C GXIV-4D | 838561 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044333508518701975/6044333508518701975_signature_28.png | Lariosa, Sheila Mae | https://www.jotform.com/uploads/macarewecare/222927978372067/6044333508518701975/6044333508518701975_signature_31.png
| 2024-10-11 | 2024-10-11 | Jose B. Lingad Memorial General Hospital | San Fernando City, Pampanga | 3 | Pampanga | RHIVDA | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29573 | MACARE-HBDC | jblmgh.is@gmail.com | 30 | ASSAY: Xpert HBV VL, COUNT: 222 ASSAY: Xpert HCV VL, COUNT: 47 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1408 |  | GXIV-D 10C | 110013748 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044299274298758803/6044299274298758803_signature_28.png | Michelle C. Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/6044299274298758803/6044299274298758803_signature_31.png
| 2024-10-11 | 2024-10-11 | Jose B. Lingad Memoria General Hospital | San Fernando City, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29572 | MACARE-HBDC | jblmghmpl@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 84 ASSAY: Xpert MTB/RIF Ultra, COUNT: 279 ASSAY: Others, COUNT: 65 ASSAY: Xpert SARS-CoV-2, COUNT: 873 |  | GXIV-D 10C | 110012825 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044278785811492812/6044278785811492812_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/6044278785811492812/6044278785811492812_signature_31.png
| 2024-10-11 | 2024-10-10 | Rosario Super Health Center | Dr. Sixto Avenue, Rosario Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29571 | MACARE-HBDC | eden.cawayan@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 649 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6044193153619202284/6044193153619202284_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6044193153619202284/6044193153619202284_signature_31.png
| 2024-10-10 | 2024-10-10 | Klinika Eastwood | 3rd Floor Citywalk 2, Eastwood City, Bagumbayan, QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | clairebaptista32786@gmail.com |  | 29570 | MACARE-HBDC | klinikaeastwoodlaboratory@gmail.com | 28 |  |  | GXIV 10C D | 110020698 | ADF for installation | Performed PM Installed HIV adf Restart GX and performed self-test Still H/w failed on module 1 Gathered report | Checked and cleaned cartridge bays, plunger rod icore slot and filter Installed HIV assay definition file Checked module 1 error 4006 | Unit is operational Module 1 for monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6043548124217543734/6043548124217543734_signature_28.png | MARY CLAIRE BAPTISTA | https://www.jotform.com/uploads/macarewecare/222927978372067/6043548124217543734/6043548124217543734_signature_31.png
| 2024-10-10 | 2024-10-10 | Governor Generoso Municipal Health Office | Poblacion, Governor Generoso | 11 | Davao Oriental | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29569 | DOH | josephmagallanes72@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 536 | 2024-09-04 | GX-IV 6C 4D | 838490 | For XpertCheck Calibration and Preventive Maintenance | Upgrade the Version of XPC into Version 1.5B Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck on four modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration All Modules are operational |  | Xpertcheck CE-5 | DR 12001 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6043526388802018401/6043526388802018401_signature_28.png | Charlene V. Magallanes,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6043526388802018401/6043526388802018401_signature_31.png
| 2024-10-10 | 2024-10-10 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29568 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 24 | ASSAY: Others, COUNT: 2 ASSAY: Xpert HPV, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 170 ASSAY: Xpert SARS-CoV-2, COUNT: 1452 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-XVI-D-6C | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules A1, B4, C2, C3, D1, D3 and D4 not operational, reported already. Modules B3 and D2 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6043511341979610345/6043511341979610345_signature_28.png | Benita Gracia N. Viloria, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6043511341979610345/6043511341979610345_signature_31.png
| 2024-10-10 | 2024-10-10 | Central Luzon Doctor's Hospital | Hospital drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 29567 | COMMERCIAL | rohmhardameg@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 206 ASSAY: Xpert SARS-CoV-2, COUNT: 59 |  | GXII-D | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6043460149713471768/6043460149713471768_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/6043460149713471768/6043460149713471768_signature_31.png
| 2024-10-10 | 2024-10-10 | Central Luzon Doctor's Hospital | Hospital Drive Brgy. San Vicente, Tarla City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 29566 | COMMERCIAL | rohmhardameg@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 206 ASSAY: Xpert SARS-CoV-2, COUNT: 59 |  | GXIV-D | 839232 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6043458049714104245/6043458049714104245_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/6043458049714104245/6043458049714104245_signature_31.png
| 2024-10-10 | 2024-10-10 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB Laboratory | NO | REPAIR | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29565 | DOH | sarah_yangzon@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 801 | 2024-10-10 | GXIV | 838406 | Module B4/861251 Defective | Preventive Maintenance Module replacement | #NAME? | Unit is Operational |  | 1 Genexpert Module 6C | 12074 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6043391017718533874/6043391017718533874_signature_28.png | Sarah V. Yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/6043391017718533874/6043391017718533874_signature_31.png
| 2024-10-10 | 2024-10-10 | Maimbung District Hospital | Laum Maimbung, sulu | BARMM | Sulu | Laboratory | YES | CALIBRATION Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29564 | DOH | alvinmadjid68@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1255 | 2024-09-10 | GX 4 (6C) | 838838 | Module 1 h/w failed. | Replaced defective module 861603 with new Module 610818. Conducted xpertcheck calibration for modules 2, 3 and 4. | Replaced defective module 861603 with new Module 610818. Conducted xpertcheck calibration for modules 2, 3 and 4. | Module 1 newly replaced, module 2 passed, module 3 and 4 to be replace. | Genexpert Machine 838838 transferred from Jolo sulu IPHO to Maimbung District Hospital | 900-0491RS, XPERTCHECK-CE-5 | 12044, 12077 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6043387361713732377/6043387361713732377_signature_28.png | Alvin B. Madjid | https://www.jotform.com/uploads/macarewecare/222927978372067/6043387361713732377/6043387361713732377_signature_31.png
| 2024-10-10 | 2024-10-10 | Garcia Memorial Provincial Hospital | Talibon, Bohol | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29563 | DOH | ybanezrosalina19@gmail.com | 27 |  |  | Gx4 | 838203 | Defected Module B4 | Run xpertcheck calibration module 4 Clean Fan Filter Gxc data, IQR and system log | Run xpertcheck calibration module 4 Clean Fan Filter Gxc data, IQR and system log | Module B4 Service Require during xpertcheck calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6043330597655814326/6043330597655814326_signature_28.png | ROSALINA A. YBAÑEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/6043330597655814326/6043330597655814326_signature_31.png
| 2024-10-10 | 2024-10-08 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29562 | MACARE-HBDC | jmoreno2020@gmail.com | 25 |  |  | MCS+ | 15H027G | MCS+ for Preventive Maintenance | Cleaned Unit Checked Program Parameters Cleaned Centrifuge bowl, rollers, pumps, valve | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041860697528244402/6041860697528244402_signature_28.png | Jake S. Moreno | https://www.jotform.com/uploads/macarewecare/222927978372067/6041860697528244402/6041860697528244402_signature_31.png
| 2024-10-09 | 2024-10-09 | Santa Cruz Rural Health Unit | Poblacion, Santa Cruz | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29561 | DOH | chendalagan72@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 584 | 2024-09-04 | GX-IV 6C 4D | 838505 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck on four modules | Perform Preventive Maintenance, and XpertCheck Calibration | Modules 1,2,3 for Service Required Module 4 Passed the Calibration | Modules for under observation, if error exist frequently modules are advised to disabled | XPERTCHECK CE 5 | DR 12011 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6042627577566012128/6042627577566012128_signature_28.png | Dorothy A. Dalagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6042627577566012128/6042627577566012128_signature_31.png
| 2024-10-09 | 2024-10-08 | University of the Philippines Los Baños - Institute of Biology | Institute of Biology Bldg, UP Los Baños | 4A | Laguna | Phycology Laboratory | NO | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | cvvelasquez@macare-medicals.com |  | 29560 | COMMERCIAL | mmartinezgoss@gmail.com | 25 |  | 2024-10-03 | Primo Star | 3150000452 | None | None | Installed Primo Star Adjusted Kohler Performed PM Performed basic operation | Unit is operational |  | Primo Star | DR 12198 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6042614743616867273/6042614743616867273_signature_28.png | Milagrosa R. Martinez-Goss | https://www.jotform.com/uploads/macarewecare/222927978372067/6042614743616867273/6042614743616867273_signature_31.png
| 2024-10-09 | 2024-10-08 | University of the Philippines Los Baños - Institute of Biology | Institute of Biology Bldg., UP Los Baños | 4A | Laguna | Phycology Laboratory | NO | TROUBLESHOOTING | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | cvvelasquez@macare-medicals.com |  | 29559 | COMMERCIAL | mmartinezgoss@gmail.com | 25 |  | 2024-10-08 | Axio Scope.A1 | 3331000862 | HAL 100 Light Source does not power on | None | Check all connections and voltages Scraped off oxidized areas for better contact Traced connectivity of supply | Unit is not operational | Pulled Out HAL 100 for further troubleshooting in the office |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6042607823617167160/6042607823617167160_signature_28.png | Milagrosa R. Martinez-Goss | https://www.jotform.com/uploads/macarewecare/222927978372067/6042607823617167160/6042607823617167160_signature_31.png
| 2024-10-09 | 2024-10-09 | Talacogon District Hospital | San Isidro, Talacogon | 13 | Agusan del Sur | RTDL | NO | CALIBRATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29558 | DOH | shyratortor@gmail.com | 18 |  |  | GX-4 6 Colors | 838275 | None | Calibration | XpertCheck Calibration and Preventive Maintenance of all modules | All module are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6042592641731727359/6042592641731727359_signature_28.png | Shyra Ericka T. Calamba | https://www.jotform.com/uploads/macarewecare/222927978372067/6042592641731727359/6042592641731727359_signature_31.png
| 2024-10-09 | 2024-10-09 | Ramon Rural Health Unit | Bugallon Proper, Ramon, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29557 | DOH | rhuramonlab@gmail.com | 28 |  |  | GXIV-6C | 838480 | PC wont Turn on Module 4 Hardware Failure | Troubleshoot PC Replaced Cmos Battery Performed Xpertcheck on module 4 (SERVICE REQUIRED) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | For observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6042496426328935926/6042496426328935926_signature_28.png | Janelle Janette Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/6042496426328935926/6042496426328935926_signature_31.png
| 2024-10-08 | 2024-10-08 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29556 | MACARE-HBDC | jmoreno2020@gmail.com | 25 |  |  | MCS+ | 15H027G | MCS+ Display Keypad Malfunction MCS+ for Preventive Maintenance | Cleaned Unit Checked Program Parameters Cleaned Centrifuge bowl, rollers, pump, valve Display ribbon(defective) | Performed Preventive Maintenance | Unit is non operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041860697528244402/6041860697528244402_signature_28.png | Jake S. Moreno | https://www.jotform.com/uploads/macarewecare/222927978372067/6041860697528244402/6041860697528244402_signature_31.png
| 2024-10-08 | 2024-10-08 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29555 | MACARE-HBDC | jmoreno2020@gmail.com | 25 |  |  | MCS+ | 15H031G | MCS+ for Preventive Maintenance | Cleaned unit Cleaned centrifuge bowl, pump, rollers, valve Checked program parameters | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041859747521752572/6041859747521752572_signature_28.png | JAKE S. Moreno | https://www.jotform.com/uploads/macarewecare/222927978372067/6041859747521752572/6041859747521752572_signature_31.png
| 2024-10-08 | 2024-10-08 | Solano Rural Health Unit | Roxas, Solano, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29554 | DOH | rhusolanolab@gmail.com | 26 |  |  | GXIV-6C | 838201 | Sometimes PC wont turn on | Troubleshoot PC Replaced CMOS Battery | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | For Observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041801065514868018/6041801065514868018_signature_28.png | Roderick Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6041801065514868018/6041801065514868018_signature_31.png
| 2024-10-08 | 2024-10-08 | St. Luke's Medical Center - Quezon City | E. Rodriguez Sr. Ave, Quezon City | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE Gathered Report | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29553 | COMMERCIAL | biomed.qc@stlukes.com.ph | 29 | ASSAY: Xpert CT/NG, COUNT: 262 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3181 |  | GXIV 6C D | 839234 | Error on Module B1 | Performed PM Archived results Generated report | Checked and cleaned cartridge bays, plunger rod and icore slot Archived results from September to October | Unit is operational B1 under monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041774620411589405/6041774620411589405_signature_28.png | Michael Fuentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6041774620411589405/6041774620411589405_signature_31.png
| 2024-10-08 | 2024-10-08 | Teresita L Jalandoni Provincial Hospital | Silay, Negros occidental | 6 | Negros Occidental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29552 | DOH | fjazzrmt13@gmail.com | 26 |  |  | Gx4 | 838464 | Due date for xpertcheck calibration | Xpertcheck module 2, 3 and 4 Clean Fan Filter Gather data IQR, system log, assay stats GXC file | Xpertcheck module 2, 3 and 4 Clean Fan Filter Gather data IQR, system log, assay stats GXC file | Module 3 and 4 was Service require during xpertcheck calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041755900321533578/6041755900321533578_signature_28.png | Fely Jasmin D. Ledesma | https://www.jotform.com/uploads/macarewecare/222927978372067/6041755900321533578/6041755900321533578_signature_31.png
| 2024-10-08 | 2024-10-08 | Calubian Rural Health Unit | Calubian, Leyte | 8 | Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29551 | DOH | mjcabardo06@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 898 |  | 6C GXIV-4D | 838399 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041659627265737355/6041659627265737355_signature_28.png | Mighty Joy V. Cabardo | https://www.jotform.com/uploads/macarewecare/222927978372067/6041659627265737355/6041659627265737355_signature_31.png
| 2024-10-08 | 2024-10-08 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB Laboratory | NO | REPAIR | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29550 | DOH | lornsabaigar@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 713 | 2024-10-08 | GXIV | 838409 | Module B2/861163 and B3/663048 Defective | #NAME? | -Performed preventive maintenance - install new module on B2/806382 and B3/807030 - Generate IQR and System log | Unit is Operational |  | 2 Genexpert Module | 12063/12064 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041639780825242296/6041639780825242296_signature_28.png | lornac.abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/6041639780825242296/6041639780825242296_signature_31.png
| 2024-10-08 | 2024-10-07 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29549 | COMMERCIAL | annabellelcatbagan@gmail.com1 | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 199 ASSAY: Others, COUNT: 17 |  | GXII-2D | 820303 | For Preventive Maintenance and Troubleshooting | #NAME? | Performed Preventive Maintenance and Troubleshooting | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041571795107016174/6041571795107016174_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6041571795107016174/6041571795107016174_signature_31.png
| 2024-10-07 | 2024-10-03 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29548 | DOH | crmcmolecularlab@gmail.com | 21.4 | ASSAY: Xpert SARS-CoV-2, COUNT: ASSAY: Xpert SARS-CoV-2, COUNT: 8990 | 2024-10-03 | GxIV-6C | 837662 | None | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: Module 1 - Passed Module 2 - skipped (new module) Module 3 - Passed Module 4 - skipped (new module) |  | Xpertcheck cartridge/kits Lot#:1001075092 | DR#: |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041158816862477044/6041158816862477044_signature_28.png | Chareasse love yu UTTO | https://www.jotform.com/uploads/macarewecare/222927978372067/6041158816862477044/6041158816862477044_signature_31.png
| 2024-10-07 | 2024-10-04 | Libertad Main Health Center | Brgy. Libertad, Butuan, Agusan del Norte | 13 | Agusan del Norte | RTDL | NO | REPAIR CALIBRATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29547 | DOH | suzzettekho34@gmail.com | 30 |  |  | GX-4 6 Colors | 838554 | B4 found service required thru XPC Error Checking | Module Replacement XpertCheck Calibration | Replacing Module B4 with SN: 862880 to SN: 619731 XpertCheck Calibration of 3 modules | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6041072241118466284/6041072241118466284_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/6041072241118466284/6041072241118466284_signature_31.png
| 2024-10-07 | 2024-10-07 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29546 | DOH | recelenylrej@gmail.com | 28 |  |  | Gx4 | 837612 | Due for xpertcheck module 4 | Xpertcheck module 4 Clean Fan filter IQR, system log and assay statistics | Xpertcheck module 4 Clean Fan filter IQR, system log and assay statistics | Module 4 service require Module 1,2 and 3 not available for xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040961680893967226/6040961680893967226_signature_28.png | Recele Nylrej Geronga | https://www.jotform.com/uploads/macarewecare/222927978372067/6040961680893967226/6040961680893967226_signature_31.png
| 2024-10-07 | 2024-10-07 | Riverside Medical Center | Bacolod City | 6 | Negros Occidental | Molecular lab | NO | ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29545 | COMMERCIAL | jehzeel.palacios@medilinx.com.ph | 23 |  |  | Gx4 | 110005268 | Module 1 to be transfered | Removing and packing of module 1 Clean Fan filter IQR before and after removing module 1 | Removing and packing of module 1 Clean Fan filter IQR before and after removing module 1 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040953480899714110/6040953480899714110_signature_28.png | Jehzeel Palacios | https://www.jotform.com/uploads/macarewecare/222927978372067/6040953480899714110/6040953480899714110_signature_31.png
| 2024-10-07 | 2024-06-21 11:21 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29544 | PBSP | bmcgenexpertlab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 1859 |  | GXIV | 834623 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040901543619453956/6040901543619453956_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6040901543619453956/6040901543619453956_signature_31.png
| 2024-10-07 | 2024-09-10 14:20 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29543 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 |  |  | GXIV | 110008530 | Unit for maintenance | NA | #NAME? | Unit is operational No issues found |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040891323617862788/6040891323617862788_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/6040891323617862788/6040891323617862788_signature_31.png
| 2024-10-07 | 2024-09-13 13:35 | University of The East Ramon Magsaysay Medical Center | Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29542 | CONSORTIUM | osangcomedador19@gmail.com | 25 |  |  | GXIV | 847578 | Unit for maintenance | NA | #NAME? | Unit is operational No issues found |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040891063611068493/6040891063611068493_signature_28.png | Lawrence Dave M. Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/6040891063611068493/6040891063611068493_signature_31.png
| 2024-10-07 | 2024-10-03 | Bulacan Medical Center | Malolos, Bulacan | NCR | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29541 | UNOPS | bmclab_micro@yahoo.com | 26 |  |  | GXIV | 1100023507 | Unit for maintenance | NA | #NAME? | Unit is operational Room air-conditioning unit is functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040890743611754725/6040890743611754725_signature_28.png | Alyssa Mae Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/6040890743611754725/6040890743611754725_signature_31.png
| 2024-10-07 | 2024-10-07 | Corazon Locsin Montelibano Regional Hospital | Bacolod City | 6 | Negros Occidental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29540 | DOH | mollab@clmmrh.doh.gov.ph | 26 |  |  | Gx 4 10c | 110020743 | Due date for xpertcheck calibration | Xpertcheck Calibration Clean Fan Filter Gather data IQR, system log, assay and GXC file | Xpertcheck Calibration Clean Fan Filter Gather data IQR, system log, assay and GXC file | Systems operational All modules passed0 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040876542932181949/6040876542932181949_signature_28.png | Santiago, dennis aaron | https://www.jotform.com/uploads/macarewecare/222927978372067/6040876542932181949/6040876542932181949_signature_31.png
| 2024-10-07 | 2024-10-07 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Rd., Las Piñas | NCR | NCR 4th - Las Piñas | Molecular | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29539 | COMMERCIAL | uphdmc_lab@yahoo.com | 25 | ASSAY: Xpert CT/NG, COUNT: 20 ASSAY: Others, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 38 ASSAY: Xpert HPV, COUNT: 37 ASSAY: Xpert MTB/RIF Ultra, COUNT: 403 ASSAY: Xpert HCV VL, COUNT: 6 |  | GXIV-D-6C | 831316 | Software Problem. | Troubleshooting. | Cleaned CPU. Cleaned and Check Functionality of RAM. Archived Tests Results. Checked Computer Settings. Generate Reports. | 2 Modules are Disabled. GX System for Further Observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040823652616286730/6040823652616286730_signature_28.png | Inger Lia Jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/6040823652616286730/6040823652616286730_signature_31.png
| 2024-10-07 | 2024-10-03 | Jose Reyes Memorial Medical Center | Rizal Ave, Sta.Curz, Manila | NCR | NCR 1st - Manila | Molecular | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29538 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 21 | ASSAY: Xpert HPV, COUNT: 303 ASSAY: Xpert SARS-CoV-2, COUNT: 1179 |  | GX16-D-10C | 110013856 | Port Problem of CPU. | Troubleshoot | Install External Port in CPU. Configure IP Address. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040810002615198454/6040810002615198454_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6040810002615198454/6040810002615198454_signature_31.png
| 2024-10-07 | 2024-10-01 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Molecular | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29536 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 25 |  |  | GX16-D-10C | 110013856 | GeneXpert Unit is Not Detectable. | Troubleshoot. | Checked Connections. Check IP Address. Checked Port Connection. Update Driver. | GX Unit is Not Detectable. Recommendation for Port Driver Reinstallation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040809712615786630/6040809712615786630_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/6040809712615786630/6040809712615786630_signature_31.png
| 2024-10-07 | 2024-09-26 | De La Salle Health Sciences Institute | Congressional Ave., Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29537 | CONSORTIUM | acpabruada@dlsmhsi.edu.ph | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2703 | 2024-09-18 | GX16-D-6C | 842166 | Modules A3 and C3 with Temperature Related Errors. | Repair. | Replaced Defective Modules. Performed Self-test. Checked Functuonality. Generate Reports. | 13 Modules are Good. 3 Modules with 2008 Errors. | 880284 -> 696374 , 688481 -> 801775 | GX Module | 12146 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040809722616209404/6040809722616209404_signature_28.png | Angie Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/6040809722616209404/6040809722616209404_signature_31.png
| 2024-10-07 | 2024-09-26 | De La Salle Health Sciences Institute | Congressional Ave., Dasmariñas City, Cavite | 4A | Cavite | THRU | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29534 | CONSORTIUM | jvfernandez@my.dlshsi.edu.ph | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 66 |  | GX16-8-D-10C | 110012951 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040809702619868815/6040809702619868815_signature_28.png | Juvy V.Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/6040809702619868815/6040809702619868815_signature_31.png
| 2024-10-07 | 2024-10-01 | San Lazaro Hospital | Quiricada St, Manila | NCR | NCR 1st - Manila | SACCL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29535 | MACARE-HBDC | ciudadanosophiapatrice@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 299 ASSAY: Xpert HCV VL, COUNT: 21 ASSAY: Xpert HIV-1 Viral Load, COUNT: 161 |  | GXIV-L-6C | 821257 | Module with 4014 Error. | Troubleshoot. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Performed Self-test. Generate Reports. | 3 Modules are Operational, 1 Module Disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040809712612461702/6040809712612461702_signature_28.png | Sophia Patrice Ciudadano | https://www.jotform.com/uploads/macarewecare/222927978372067/6040809712612461702/6040809712612461702_signature_31.png
| 2024-10-07 | 2024-10-07 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29533 | DOH | serly_reyes@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 772 |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040737656614661865/6040737656614661865_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/6040737656614661865/6040737656614661865_signature_31.png
| 2024-10-07 | 2024-10-07 | Minglanilla Rural Health Unit | Pide 1, Minglanilla ,Cebu | 7 | Cebu | Bloodbank | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29532 | DOH | brendagonzagaparedes@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1650 |  | Gx4 | 838512 | Due for calibration annual. | Performed calibration in Module 2. Clean up Icore and filter and plunger. | Performed calibration in Module 2. Clean up Icore and filter and plunger. | All units are operational. Module 2 passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040730439847788486/6040730439847788486_signature_28.png | Brenda Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/6040730439847788486/6040730439847788486_signature_31.png
| 2024-10-07 | 2024-09-30 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | HACT | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com | hact@bghmc.doh.gov.ph |  | 29531 | MACARE-HBDC | balbahazel@gmail.com | 20 |  |  | GXIV-4-D-6C | 821632 | CPU won't Boot | #NAME? | Performed Troubleshooting Instructions of Cepheid | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040714896611766451/6040714896611766451_signature_28.png | Hazel Gwyneth Balba | https://www.jotform.com/uploads/macarewecare/222927978372067/6040714896611766451/6040714896611766451_signature_31.png
| 2024-10-07 | 2024-09-30 | Baguio General Hospital and Medical Center | Baguio City | 1 | Benguet | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29530 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 18 | ASSAY: Xpert HBV VL, COUNT: 195 ASSAY: Xpert SARS-CoV-2, COUNT: 198 ASSAY: Xpert HCV VL, COUNT: 48 ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 |  | GXIV-4-D-6C | 110016112 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040714896612246785/6040714896612246785_signature_28.png | Rhe AnnetonetyeAcosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6040714896612246785/6040714896612246785_signature_31.png
| 2024-10-07 | 2024-09-30 | Baguio General Hospital and Medical Center | Baguio City | 1 | Benguet | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29529 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 18 | ASSAY: Xpert HBV VL, COUNT: 195 ASSAY: Xpert SARS-CoV-2, COUNT: 198 ASSAY: Xpert HCV VL, COUNT: 48 ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 |  | GXIV-4-D-6C | 110016112 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6040714776619567605/6040714776619567605_signature_28.png | Rhe AnnetonetyeAcosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6040714776619567605/6040714776619567605_signature_31.png
| 2024-10-04 | 2024-10-04 | BINDOY DISTRICT HOSPITAL | BINDOY, NEGROS ORIENTAL | 6 | Negros Oriental | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29528 | DOH | bindoydistricthosp.laboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 561 | 2024-09-03 | GX-IV | 838475 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK, CE-IVD 5 Cartridges/kit | 11959 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038358775415758883/6038358775415758883_signature_28.png | Lemuel T. Quirit , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6038358775415758883/6038358775415758883_signature_31.png
| 2024-10-04 | 2024-10-04 | Our Lady of Lourdes Hospital | Sta. Mesa, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29527 | CONSORTIUM | mariaysabelle.visda@medilinx.com.ph | 25 |  |  | GXIV | 840974 | Desktop freezes during sample running | Reset RAM | #NAME? | Unit is operational Desktop is under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038338372722374712/6038338372722374712_signature_28.png | Maria Ysabelle Visda | https://www.jotform.com/uploads/macarewecare/222927978372067/6038338372722374712/6038338372722374712_signature_31.png
| 2024-10-04 | 2024-10-04 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29526 | DOH | rhwclab@gmail.com | 24.6 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1405 ASSAY: Xpert MTB/RIF Ultra, COUNT: 385 ASSAY: Xpert HBV VL, COUNT: 9 | 2024-09-04 | GX-IV 6C 4D | 838262 | Stuck Cartidge Due for Preventive Maintenance and XpertCheck Calibration | Extract the stuck cartidge at Module 1, Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run Calibration on Modules 1 and 3 only | Perform Preventive Maintenance and XpertCheck Calibration Perform Stuck Cartidge extraction | All Modules are operational |  | XPERTCHECK CE 5 | 12009 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038289100126352824/6038289100126352824_signature_28.png | Kent Ferolin,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6038289100126352824/6038289100126352824_signature_31.png
| 2024-10-04 | 2024-10-04 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29525 | MACARE-HBDC | rhwclab@gmail.com | 25.6 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1405 ASSAY: Xpert MTB/RIF Ultra, COUNT: 385 ASSAY: Xpert HBV VL, COUNT: 9 |  | GX-IV 6C 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Only Three Modules are operational | Module B3 for disabled due to multiple temperature related errors |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038288690124046365/6038288690124046365_signature_28.png | Kent Ferolin,. rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/6038288690124046365/6038288690124046365_signature_31.png
| 2024-10-04 | 2024-10-04 | NEGROS ORIENTAL PROVINCIAL HOSPITAL | Dumaguete City | 7 | Negros Oriental | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29524 | DOH | rezelmarielagos@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1254 ASSAY: Xpert HCV VL, COUNT: 11 ASSAY: Xpert HIV-1 Viral Load, COUNT: 113 | 2024-09-03 | GX-IV | 838525 | FOR CALIBRATION | COLLECT IQR, SYSTEM LOG,ARCHIVE, AND CALIBRATION DATA | #NAME? | EQUIPMEMT IS OPERATIONAL, MODULE A4 DISABLED |  | XPERCHECK, CE-IVD 5 Cartridges/kit | 11972 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038211156708595002/6038211156708595002_signature_28.png | Rezel Marie S. Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/6038211156708595002/6038211156708595002_signature_31.png
| 2024-10-04 | 2024-10-04 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29523 | DOH | dalepenalosa189@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1178 | 2024-10-04 | GXIV | 838421 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038195334699186994/6038195334699186994_signature_28.png | Dale D. Peñalosa,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6038195334699186994/6038195334699186994_signature_31.png
| 2024-10-04 | 2024-10-03 | Philippine Res Cross Port Area | Intramuros Manila | NCR | NCR 1st - Manila | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29522 | MACARE-HBDC | dolly.banares@redcross.org.ph | 27 |  |  | MCS + | 03J030 | Centrifuge latch to tight when releasing to unlock position | Checked centrifuge latch Adjust relay position Performed functionality test Test unit | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038160977023787931/6038160977023787931_signature_28.png | Dolly banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6038160977023787931/6038160977023787931_signature_31.png
| 2024-10-04 | 2024-10-02 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed@nch.doh.gov.ph |  | 29521 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2542 ASSAY: Xpert SARS-CoV-2, COUNT: 1575 ASSAY: Xpert SARS-CoV-2, COUNT: 1073 |  | GX16 D | 110014171 | As per end user they encounter intermittent error 5007. In different module | Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed troubleshooting | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038160667026355050/6038160667026355050_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/6038160667026355050/6038160667026355050_signature_31.png
| 2024-10-04 | 2024-10-03 | Pasig Treatment Hub | Dr. Garcia St., Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | dansumulong.pasigcho@gmail.com |  | 29520 | MACARE-HBDC | pathmedtech@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 802 ASSAY: Xpert HIV-1 Viral Load, COUNT: 802 |  | GX-IV-L-6C | 821648 | For maintenance | Performed maintenance. | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A3 operational. Modules A1/733228, A2/734493, A4/804598 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038138027024462531/6038138027024462531_signature_28.png | Dan Nathaniel Sumulong | https://www.jotform.com/uploads/macarewecare/222927978372067/6038138027024462531/6038138027024462531_signature_31.png
| 2024-10-04 | 2024-10-01 | Manila Doctors Hospital | United Nations Ave., Ermita, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29519 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 77 ASSAY: Xpert HBV VL, COUNT: 90 ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1277 ASSAY: Xpert HCV VL, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 116 |  | GXIV 10C D | 110001648 | GX for transfer | Performed Preventive maintenance Transferred GX to Molec | Disassembled machine Transfer machine from Micro Lab to Molec Lab Assembled machine Checked and cleaned cartridge bays, plunger rod, icore slot and filter | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038107987024278891/6038107987024278891_signature_28.png | KAYE MARTELINO | https://www.jotform.com/uploads/macarewecare/222927978372067/6038107987024278891/6038107987024278891_signature_31.png
| 2024-10-04 | 2024-10-03 | Pasig Treatment Hub | Dr. Garcia St., Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | dansumulong.pasigcho@gmail.com |  | 29518 | MACARE-HBDC | pathmedtech@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 802 ASSAY: Xpert HIV-1 Viral Load, COUNT: 802 |  | GX-IV-L-6C | 821648 | For maintenance | Performed maintenance. | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A3 operational. Modules A1/733228, A2/734493, A4/804598 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6038105377029137803/6038105377029137803_signature_28.png | Dan Nathaniel Sumulong | https://www.jotform.com/uploads/macarewecare/222927978372067/6038105377029137803/6038105377029137803_signature_31.png
| 2024-10-03 | 2024-10-03 | Philippine Red Cross Port Area | Intramuros Manila | NCR | NCR 1st - Manila | Bloodbnak | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29517 | MACARE-HBDC | dolly.banares@redcross.org.ph | 26 |  |  | MCS + | 12D063 | No power, during switch on Vario stand caster wheel worn out | Checked and measure voltage input in power entry module. Test line continuity Checked and test fuse continuity After checking and testing the unit still unoperational need to replace power entry module Defective power entry module Nuetral line | Performed troubleshooting | Unit unoperational | For replacement of power entry module and vario stand |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037450801027250460/6037450801027250460_signature_28.png | Dolly banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6037450801027250460/6037450801027250460_signature_31.png
| 2024-10-03 | 2024-10-03 | Southern Philippines Medical Center | JP Laurel Avenue Bajada, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29516 | UNOPS | spmcrhivda@gmail.com | 27.6 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 585 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 371 |  | GX-IV 10C 4D | 110023505 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037441490127085755/6037441490127085755_signature_28.png | Flordeliza Go,. RMT MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/6037441490127085755/6037441490127085755_signature_31.png
| 2024-10-03 | 2024-10-03 | Southern Philippines Medical Center | JP Laurel Avenue, Bajada Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29515 | UNOPS | spmcrhivda@gmail.com | 25.6 |  |  | GX-IV 10C 4D | 110023506 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037440620121245853/6037440620121245853_signature_28.png | Flordeliza B. Go, RMT, MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/6037440620121245853/6037440620121245853_signature_31.png
| 2024-10-03 | 2024-10-03 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29514 | DOH | luzvimindanavidad2@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1022 | 2024-09-10 | 6C GXIV-4D | 838297 | Module B1 failed XPC | -Replace defective module B1/806280 with module SN: 617227 | #NAME? | All modules are functional. |  | 1 Certified Pre-owned GeneXpert Module SN. 617227 | DR no. 12079 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037434157217790314/6037434157217790314_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/6037434157217790314/6037434157217790314_signature_31.png
| 2024-10-03 | 2024-10-03 | Tuburan Rural Health Unit | Tuburan, Basilan | BARMM | Basilan | Laboratory | YES | CALIBRATION | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29513 | DOH | rjakilan20@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 481 | 2024-09-05 | GX 4 (6c) | 838832 | None | Xpertcheck calibration for modules 1, 2, 3 and 4. | Xpertcheck calibration for modules 1, 2, 3 and 4. | All modules are available and ready to use | None | XPERTCHECK-CE-5 | 12045 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037431112822803693/6037431112822803693_signature_28.png | Raiyana Hassan jakilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6037431112822803693/6037431112822803693_signature_31.png
| 2024-10-03 | 2024-10-03 | Philippine Red Cross - Davao Chapter | Roxas Boulevard, Davao City | 11 | Davao del Sur | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 29512 | COMMERCIAL | davaocitybloodservices@gmail.com | 25.8 |  |  | Haemonetics MCS+ | 02J021 | Error on Haemonetics, spillage and low to no flow | Adjust the Optic Gain parameters Check the Line Sensor parameters Clean Surface area Perform Cycle Test | Troubleshoot the MCS+ | Machine is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037376430126882077/6037376430126882077_signature_28.png | Joy Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/6037376430126882077/6037376430126882077_signature_31.png
| 2024-10-03 | 2024-10-02 | Claver Rural Health Unit | Brgy. Tayaga, Claver | 13 | Surigao del Norte | RTDL | NO | CALIBRATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29511 | DOH | rhodaruthdalugdog@gmail.com | 30 |  |  | GX-4 | 838514 | B4 found service required thru XPC | Calibration (For reporting) | XpertCheck Calibration | 3 modules functional, 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037373515695591572/6037373515695591572_signature_28.png | Rhoda Ruth Dalugdog | https://www.jotform.com/uploads/macarewecare/222927978372067/6037373515695591572/6037373515695591572_signature_31.png
| 2024-10-03 | 2024-10-03 | Chong Hua Hospital- Mandaue | Subangdaku , Mandaue City Cebu | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29510 | COMMERCIAL | irl.maderazo@gmail.com | 24 |  |  | MCS+ | 11J217 | Error in ACD pump during flow of ACD. | Replaced ACD pump in MCS+ | Replaced ACD pump in MCS+ | Unit apheresis machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037350079843280978/6037350079843280978_signature_28.png | Inna Maderazo,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6037350079843280978/6037350079843280978_signature_31.png
| 2024-10-03 | 2024-10-03 | Chinese General Hospital | Blumentritt, Sta. Cruz Manila | NCR | NCR 1st - Manila | Blood bank | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29509 | COMMERCIAL | cghmcbloodbank2016@gmail.xom | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. Unit for calibration. For troubleshooting of valves. Compressor slow to cutoff. | Preventive maintenance, calibration and troubleshooting. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Performed calibration. Verified calibration parameters. Performed simulation running. Performed valve diagnostics. | Unit is operational. For observation of compressor for valves. For request of compressor for isolation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037346137019218248/6037346137019218248_signature_28.png | Jacqueline Zarzoso | https://www.jotform.com/uploads/macarewecare/222927978372067/6037346137019218248/6037346137019218248_signature_31.png
| 2024-10-03 | 2024-10-03 | Estevez Memorial Hospital, Inc. | Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City, 4500 Albay | 5 | Albay | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29508 | COMMERCIAL | reylovendino83@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 176 ASSAY: Others, COUNT: 10 ASSAY: Xpert SARS-CoV-2, COUNT: 11 |  | GX4-10C | 110017942 | Room temperature reading | Place temperature data logger near Gx machine. Room temp to be recorded and observed for 1week Collect gxx file Generate reports | Performed Troubleshooting Collect Data | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037336994041971470/6037336994041971470_signature_28.png | Rey a. Lovendino | https://www.jotform.com/uploads/macarewecare/222927978372067/6037336994041971470/6037336994041971470_signature_31.png
| 2024-10-03 | 2024-10-03 | CONG. LAMBERTO L. MACIAS MEMORIAL HOSPITAL | Siaton, Negros Oriental | 7 | Negros Oriental | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29507 | DOH | honiemacaday619@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 416 | 2024-09-10 | GX-IV | 838482 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE REPLACEMENT: B1 DEF SN:862069 WITH NEW SN:210001403 B3 DEF SN:862073 WITH NEW SN:888597 | EQUIPMENT IS OPERATIONAL |  | CERTIFIED PRE-OWNED GENEXPERT MODULE | 1206112062 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037296102928407559/6037296102928407559_signature_28.png | Honie Dy Macaday | https://www.jotform.com/uploads/macarewecare/222927978372067/6037296102928407559/6037296102928407559_signature_31.png
| 2024-10-03 | 2024-10-02 | Sultan Mastura Rural Health Unit | Sultan Mastura, Maguindanao | 12 | Maguindanao | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29506 | DOH | salehshahinaz@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 411 | 2024-10-02 | GxIV-6C | 837613 | Module 1 and 2 service required | Perform Genexpert annual Xpertcheck Calibration Perform Genexpert preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors Problem reporting Data collection | 2 operational modules Calibration Results: Module 1 - service required Module 2 - service required Module 3 - Passed Module 4 - Passed |  | Xpertcheck cartridge/kits Lot#: 1001075092 | DR#: 12031 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037212335312830417/6037212335312830417_signature_28.png | Shahinaz Saleh | https://www.jotform.com/uploads/macarewecare/222927978372067/6037212335312830417/6037212335312830417_signature_31.png
| 2024-10-03 | 2024-10-02 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29505 | DOH | gmpelaez144@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 188 | 2024-10-02 | GxIV-6C | 838528 | Module 1 H/W failed | Perform Genexpert module repair Perform Genexpert annual Xpertcheck Calibration Perform Genexpert preventive maintenance | Replace defective module 1 SN:804507 with new module SN:210010056 Perform Genexpert machine Xpertcheck calibration Perform Genexpert machine Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: Module 1 - Skipped (new module) Module 2 - Skipped (new module) Module 3 - Passed Module 4 - Skipped (new module) |  | -Genexpert 6-color module SN:210010056 -Xpertcheck cartridge/kits lot#:1001075092 | DR#: 12060 and 12028 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6037212335313790282/6037212335313790282_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6037212335313790282/6037212335313790282_signature_31.png
| 2024-10-02 | 2024-09-27 | Kalinga Health Hub - Marikina | JP Rizal St. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29502 | CONSORTIUM | bryllcruz08@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1549 |  | GX-IV-D-6C | 110008418 | Error 1002 on module B3 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Checked module B3 Generated reports | Unit is operational. Module B3/616087 disabled, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036587667022786657/6036587667022786657_signature_28.png | Lorenzo Bryll Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6036587667022786657/6036587667022786657_signature_31.png
| 2024-10-02 | 2024-10-01 | The Medical City | Ortigas Ave. Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29504 | COMMERCIAL | bloodbank@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036587667027324277/6036587667027324277_signature_28.png | Mary Pauline Samson | https://www.jotform.com/uploads/macarewecare/222927978372067/6036587667027324277/6036587667027324277_signature_31.png
| 2024-10-02 | 2024-10-01 | The Medical City | Ortigas ave. Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29503 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036587667027450698/6036587667027450698_signature_28.png | Ana Ponferrada | https://www.jotform.com/uploads/macarewecare/222927978372067/6036587667027450698/6036587667027450698_signature_31.png
| 2024-10-02 | 2024-10-01 | The Medical City | Ortigas Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29501 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 52 ASSAY: Xpert CT/NG, COUNT: 73 ASSAY: Others, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 592 ASSAY: Xpert HPV, COUNT: 429 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5525 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 85 ASSAY: Xpert HIV-1 Viral Load, COUNT: 684 ASSAY: Xpert SARS-CoV-2, COUNT: 1151 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance: Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036587667028458163/6036587667028458163_signature_28.png | Stephanie obida | https://www.jotform.com/uploads/macarewecare/222927978372067/6036587667028458163/6036587667028458163_signature_31.png
| 2024-10-02 | 2024-10-02 | Monkayo Rural Health Unit | Poblacion, Monkayo | 11 | Davao de Oro | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29500 | DOH | kendralois730@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1959 | 2024-09-04 | GX-IV 6C 4D | 838487 | Due for Preventive Maintenance and XpertCheck Calibration UPS no power on battery mode | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Install new version of the XPC Version 1.5b and run XpertCheck on four modules | Perform Preventive Maintenance, and XpertCheck Calibration UPs Troubleshooting | Module B1 and B3 for Service Required Module B2 and B4 Passed the Calibration Modules B1 and B3 for under observation, if module encounter multiple temperature related error, enduser was advised to disable the said modules UPS wont charge up and no power at battery mode, documented the error to be reported |  | XPERTCHECK CE - 5 | DR 12005 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036585541717834792/6036585541717834792_signature_28.png | Karen S. Anuta,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6036585541717834792/6036585541717834792_signature_31.png
| 2024-10-02 | 2024-09-27 | BUENAVISTA EMERGENCY HOSPITAL | Mclain Buenavista, Guimaras | 6 | Guimaras | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29499 | UNOPS | campo_janice@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-12 | GX-IV | 110023539 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  | 12134 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036517322673938119/6036517322673938119_signature_28.png | Janice de la cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6036517322673938119/6036517322673938119_signature_31.png
| 2024-10-02 | 2024-10-02 | Labuan General hospital | Labuan, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION MAINTENANCE GX Barcode Sticker Installation | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29498 | DOH | jenne060213@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2274 | 2024-09-05 | GX 4 (6c) | 838568 | Module 4 error error 4009, 1002 | Xpertcheck calibration, Preventive Maintenence and GX Barcode sticker installation | Xpertcheck calibration, Preventive Maintenence and GX Barcode sticker installation. | Module 1,2,3 and 4 are available. Module 4 failed xpertcheck calibration while module 1 and 2 passed. Module 3 skipped due to new replaced module. | Module 4 is for replacement | XPERTCHECK-CE-5 | 12034 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036516025415283213/6036516025415283213_signature_28.png | Jennerose Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6036516025415283213/6036516025415283213_signature_31.png
| 2024-10-02 | 2024-09-30 | Dr. Rafael S. Tumbokon Memorial Hospital | Kalibo, Aklan | 6 | Aklan | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29497 | MACARE-HBDC | gonzagaevarmt@gmail.com | 29 |  | 2024-09-30 | GXIV 10C | 110020745 | Due for Calibration | #NAME? | #NAME? | Unit is Operational |  | Xpertcheck kit | 12047 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036512185026783811/6036512185026783811_signature_28.png | Eva Raquel D. Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6036512185026783811/6036512185026783811_signature_31.png
| 2024-10-02 | 2024-10-02 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29496 | DOH | sajecpooh20@gmail.com | 18 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1250 | 2024-09-10 | 6C GXIV-4D | 838555 | Module A2 failed XPC | -Replace defective module A2/878222 with module SN: 615450 | #NAME? | Module A4 for XPC retry. Machine is functional. |  | -Certified pre-own GeneXpert module SN. 615450 | DR no. 12080 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036472764129409873/6036472764129409873_signature_28.png | Letecia Bagohin | https://www.jotform.com/uploads/macarewecare/222927978372067/6036472764129409873/6036472764129409873_signature_31.png
| 2024-10-02 | 2024-10-01 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29495 | MACARE-HBDC | ljf7403@yahoo.com | 25 |  |  | MCS+ | 13C270 | Unit for maintenance Check up of UPS | NA | #NAME? | Unit is operational UPS for replacement (SN: AT1000190123ST0327) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036454507025080786/6036454507025080786_signature_28.png | Vonn De Jesus | https://www.jotform.com/uploads/macarewecare/222927978372067/6036454507025080786/6036454507025080786_signature_31.png
| 2024-10-02 | 2024-10-01 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29494 | COMMERCIAL | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HPV, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 134 ASSAY: Xpert HCV VL, COUNT: 2 |  | GX16 D | 804476 | Due for maintenance A1,A2,A3,B1,B3,C1,C2,C3,C4 encounter loss comm. B4 temp error 4010,4009,4012 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | A4,B2 operational A1,A2,A3,B1,B3,B4,C1,C2,C3,C4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036430097022490617/6036430097022490617_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6036430097022490617/6036430097022490617_signature_31.png
| 2024-10-02 | 2024-10-01 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29493 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 57 ASSAY: Others, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 20 ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 79 ASSAY: Xpert HCV VL, COUNT: 8 |  | GX4 D | 838524 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6036429727023934910/6036429727023934910_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6036429727023934910/6036429727023934910_signature_31.png
| 2024-10-01 | 2024-10-01 | Virac Municipal Health Office | Virac, Catanduanes | 5 | Catanduanes | Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29492 | DOH | limarmariano50@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 475 |  | GX4-6C | 838563 | A3/736617 - service required Monitor - blue on half of the screen Date and time resets after power on CPU slow response/lagging | Replaced module: A3 - 736617>>916398 Replaced monitor Change CPU CMOS battery for date and time Cleaned and resets RAM Cleaned cartridge bays, plunger rod, GX fan, replaced fan filter Generate Data | Replaced module A3/736617 Replaced monitor Performed Preventive Maintenance Troubleshooting CPU | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6035748256311960621/6035748256311960621_signature_28.png | Ramil A. Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/6035748256311960621/6035748256311960621_signature_31.png
| 2024-10-01 | 2024-10-01 | Panabo City Health Office | Poblacion | 11 | Davao del Norte | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29491 | DOH | nadineeust67@gmail.com | 23.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-09-04 | GX-IV 6C 4D | 838295 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install new version of XPC 1.5B Run XPC at Four Modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules passed the Calibration All Modules are operational |  | XPERTCHECK CE-5 | DR 12008 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6035711705316518978/6035711705316518978_signature_28.png | Nadine Gemma S. Eustaquio,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6035711705316518978/6035711705316518978_signature_31.png
| 2024-10-01 | 2024-10-01 | Quezon City General Hospital | Seminary Rd, Proj 8, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29490 | MACARE-HBDC | elenaventura076@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 92 |  | GXIV 6C D | 110012766 | GX due for maintenance | Performed pm Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Archived results | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6035690351023787002/6035690351023787002_signature_28.png | Elena G. Ventura RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6035690351023787002/6035690351023787002_signature_31.png
| 2024-10-01 | 2024-09-30 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29489 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1485 ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Others, COUNT: 114 ASSAY: Xpert HBV VL, COUNT: 1068 ASSAY: Xpert HPV, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2326 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 10813 ASSAY: Xpert HCV VL, COUNT: 152 | 2024-09-27 | GX16 D | 110005587 | Due for xpert check and maintenance A4 encountered 4004 underwarranty last module replacement 8.20.24 | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck and maintenance | A1,A2,C2,D2 disabled reported already C4 for reporting A4 new replaced A3 B3 failed xpertcheck C1,D1,D4 passed xpertcheck |  | Xpertcheck | 98839 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6035512488712105693/6035512488712105693_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6035512488712105693/6035512488712105693_signature_31.png
| 2024-09-30 | 2024-09-27 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29487 | DOH | jan19.cam@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1298 | 2024-09-27 | GxIV-6C | 838504 | None | Perform Genexpert machine annual Xpertcheck Calibration Perform Genexpert machine Preventive maintenance | Perform Genexpert Xpertcheck Calibration Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: All modules passed |  | Xpertcheck cartridge/kit Lot#:1001075092 | DR#:12016 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6035112646866534917/6035112646866534917_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/6035112646866534917/6035112646866534917_signature_31.png
| 2024-09-30 | 2024-09-26 | Malapatan Rural Health Unit | Malapatan, Sarangani Province | 12 | Sarangani | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29488 | DOH | kamilahnichol@yahoo.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 689 | 2024-09-26 | GxIV-6C | 838419 | None | Perform Genexpert machine annual Calibration Perform Genexpert machine Preventive maintenance | Perform Genexpert Calibration/Xpertcheck Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: All modules passed |  | Xpertcheck cartridge/kit Lot#:1001075092 | DR#: 12021 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6035112646869256727/6035112646869256727_signature_28.png | Kamilah Nichol Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/6035112646869256727/6035112646869256727_signature_31.png
| 2024-09-30 | 2024-09-24 | Philippine Red Cross | Intramuros, Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29486 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. For checking of rotor pumps. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Preventive maintenance and troubleshooting. | Unit is operational. Rubber rollers are okay. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6034841308425975976/6034841308425975976_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6034841308425975976/6034841308425975976_signature_31.png
| 2024-09-30 | 2024-09-24 | Philippine Red Cross | Intramuros, Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29485 | MACARE-HBDC | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 03J030 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6034841308428429967/6034841308428429967_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/6034841308428429967/6034841308428429967_signature_31.png
| 2024-09-30 | 2024-09-30 | Concepcion Rural Health Unit | Confessor St. Poblacion | 6 | Iloilo | TB Laboratory | NO | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29484 | UNOPS | jescelmaymesias@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-30 | GXIV 10C | 110023527 | New Installation | N/a | Install Genexpert machine SN: 110023527 System Unit SN: CPH1023F469955 | Unit is Operational |  | Genexpert Machine CPU Printer UPS Barcode Sanner Monitor | 12122 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6034803675386353990/6034803675386353990_signature_28.png | jescel may m. melgar | https://www.jotform.com/uploads/macarewecare/222927978372067/6034803675386353990/6034803675386353990_signature_31.png
| 2024-09-30 | 2024-09-24 | Our Lady of Lourdes Hospital | Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | jcmasaganda@ollh.ph |  | 29483 | CONSORTIUM | leslie.cabanilla@medilinx.com.ph | 25 |  |  | GXIV | 840974 | Unit for maintenance | NA | - checked and cleaned cartridge bay, plunger rod, filter - generated reports - performed archive (2021 - 2023) - performed self test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6034743117404213254/6034743117404213254_signature_28.png | Leslie Jhon Cabanilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6034743117404213254/6034743117404213254_signature_31.png
| 2024-09-30 | 2024-09-25 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | REPAIR | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29482 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  | 2024-09-25 | MCS+ | 08B049 | Display Assembly and Variocar for replacement | Replaced Display Assembly and Variocar | #NAME? | Unit is operational | With John Monsour Doctor | Display Assembly MCS+ VARIOCAR | 12167 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6034743117407329023/6034743117407329023_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6034743117407329023/6034743117407329023_signature_31.png
| 2024-09-29 | 2024-09-27 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29481 | MACARE-HBDC | eamcmolecularlab@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 188 ASSAY: Xpert HPV, COUNT: 269 ASSAY: Xpert SARS-CoV-2, COUNT: 1381 ASSAY: Xpert SARS-CoV-2, COUNT: 6 |  | GX4 D 10C | 110016781 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6034227775107275878/6034227775107275878_signature_28.png | EAST AVENUE MEDICAL CENTER | https://www.jotform.com/uploads/macarewecare/222927978372067/6034227775107275878/6034227775107275878_signature_31.png
| 2024-09-29 | 2024-09-27 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29480 | MACARE-HBDC | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 362 ASSAY: Xpert SARS-CoV-2, COUNT: 805 ASSAY: Xpert HIV-1 Viral Load, COUNT: 116 |  | GX4 D | 838837 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B2,B3 operational B1,B4 disabled failed xpertcheck reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6034227285105994343/6034227285105994343_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6034227285105994343/6034227285105994343_signature_31.png
| 2024-09-29 | 2024-09-27 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29479 | COMMERCIAL | lcpmicrolab@gmail.com | 26 |  |  | GX4 L | 804205 | Unit not in use due to loss communication error. Last xpertcheck 2021 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | D2,D3,D4 disable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6034226685108409514/6034226685108409514_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/6034226685108409514/6034226685108409514_signature_31.png
| 2024-09-29 | 2024-09-26 | Biopath QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29478 | COMMERCIAL | ainocentes@biopath.ph | 26 | ASSAY: Xpert CT/NG, COUNT: 48 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 20 ASSAY: Xpert SARS-CoV-2, COUNT: 6 ASSAY: Xpert SARS-CoV-2, COUNT: 4 |  | GX16 D | 839306 | Loss communication Low voltage output network switch power supply | Replaced and installed power adaptor. Installed all cables Checked and test power supply voltage Performed functionality test | Performed repair | Unit operational | Work with, Sir toco, gian paolo |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6034226185106412354/6034226185106412354_signature_28.png | Amanda inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/6034226185106412354/6034226185106412354_signature_31.png
| 2024-09-28 | 2024-09-26 | Tagum City Health Office | Mabini St, Tagum City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29477 | DOH | tagumcholab.mabini@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2311 | 2024-09-18 | GX-IV 6C 4D | 838553 | Module B1 encounter error Due for Calibration | Clean Surface Area, Cartridge Bay Replaced the Defective Module SN: 651048 with new Module SN:808354 Run XpertCheck on Module 2 only, Since Newly Replaced Modules 1,3 and 4 | Perform Module Replacement and XpertCheck Calibration | Module 2 passed the Calibration All Modules Are Operational |  | 900-0491rs XPERTCHECK CE-5 | DR 12147 and DR 12012 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6033448248621530688/6033448248621530688_signature_28.png | Cristina Tiara A. Fernandez,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6033448248621530688/6033448248621530688_signature_31.png
| 2024-09-28 | 2024-09-26 | Dr. Catalino Gallego Nava Provincial Hospital | San Miguel,Jordan, Guimaras | 6 | Guimaras | Laboratory | NO | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29476 | UNOPS | casluzuriaga1983@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-26 | GXIV 10C | 110023540 | New Installation | n/a | Install New Genexpert machine CPU SN: CPH1023F46915 Monitor SN: CN0WMYC5CC037KCLTXA05 UPS SN: AT1K231111ST0045 Printer SN: CN3B2431CG | Unit is operational |  |  | 12135 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6033147507219503344/6033147507219503344_signature_28.png | Christopher Alan Luzuriaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6033147507219503344/6033147507219503344_signature_31.png
| 2024-09-28 | 2024-09-27 | Buenavista Rural Health Unit | New Poblacion, Buenavista | 6 | Guimaras | TB-DOTS | NO | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29475 | UNOPS | clys0910@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-27 | GXIV 10C | 110023538 | New Installation | n/a | Install Genexpert machine SN: 110023538 System Unit SN: CPH1023F46913 UPS SN: AT1K231111ST0009 Printer SN: CN3V2431QP | Unit is operational |  | Genexpert machine Printer System Unit Barcode Sanner UPS Monitor | 12133 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6033146627218492083/6033146627218492083_signature_28.png | Raymund Salapio | https://www.jotform.com/uploads/macarewecare/222927978372067/6033146627218492083/6033146627218492083_signature_31.png
| 2024-09-27 | 2024-09-27 | Midsalip Rural Health unit | Poblacion Midsalip, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29474 | DOH | genaliner@rocketmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 393 | 2024-09-05 | GX 4 (6c) | 837664 | None | Module 1, 2 and 3 undergone Xpertcheck calibration. All modules passed | Module 1, 2 and 3 undergone Xpertcheck calibration. All modules passed | All modules are available and ready to use | None. Module 4 unable to undergone Xpertcheck calibration due to recently replaced module. | XPERTCHECK-CE-5 | 12036 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6032223554021740254/6032223554021740254_signature_28.png | Genaline Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/6032223554021740254/6032223554021740254_signature_31.png
| 2024-09-27 | 2024-09-27 | BUENAVISTA EMERGENCY HOSPITAL | Mclain Buenavista, Guimaras | 6 | Guimaras | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29473 | UNOPS | campo_janice@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-12 | GX-IV | 110023539 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  | 12134 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6032195414086375266/6032195414086375266_signature_28.png | Janice de la cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6032195414086375266/6032195414086375266_signature_31.png
| 2024-09-27 | 2024-09-24 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29472 | MACARE-HBDC | peachmango318@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 278 |  | GX-IV-D-10C | 110024245 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit os operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6032063687022137944/6032063687022137944_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6032063687022137944/6032063687022137944_signature_31.png
| 2024-09-26 | 2024-09-26 | Lala Rural Health Unit | Lala, Lanipao Lanao del Norte | 10 | Lanao del Norte | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29471 | DOH | villagonzalojoyb@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 688 |  | GX4 | 838229 | Due for annual calibration B1,B2 and B3 | Conduct Calibration in Module 1,2 and 3 Performed preventive maintenance on genexpert machine Generate GXX files, IQR, Sys Log, and Assay Statistics Performed preventive maintenance | Conduct Calibration in Module 1,2 and 3 Performed preventive maintenance on genexpert machine Generate GXX files, IQR, Sys Log, and Assay Statistics Performed preventive maintenance | B1/B2/B3 - passed All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031501129845319551/6031501129845319551_signature_28.png | Maria Glenda Villagonzalo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6031501129845319551/6031501129845319551_signature_31.png
| 2024-09-26 | 2024-09-26 | Kapatagan Rural Health Unit | Kapatagan , Lanao del Norte | 10 | Lanao del Norte | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29470 | DOH | aganub2005@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 |  | GX4 | 838224 | Module 4 SN # 807289 defective failed during calibration. | Replacement of new module SN # 752826 Sticker placement Calibrate module 1,2,3 Clean up ICore, plunger, valve, cartridge bay , and filter. Generate IQR, System log, GXX files, and Assay statistics | Replacement of new module SN # 752826 Sticker placement Calibrate module 1,2,3 Clean up ICore, plunger, valve, cartridge bay , and filter. Generate IQR, System log, GXX files, and Assay statistics | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031497899841889484/6031497899841889484_signature_28.png | Angelita Paradero, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6031497899841889484/6031497899841889484_signature_31.png
| 2024-09-26 | 2024-09-26 | Kauswagan Rural Health Unit -Tb Dots | Kauswagan Lanao del Norte | 10 | Lanao del Norte | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29469 | DOH | dextertinoy@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 325 |  | GX4 | 838402 | Defective module 3 SN # 749578 failed during calibration. | Replacement of Module 3 with new module SN # 638856 Conduct Calibration in Module 1,2 and 4 Performed preventive maintenance on genexpert machine Generate GXX files, IQR, Sys Log, and Assay Statistics | Replacement of Module 3 with new module SN # 638856 Conduct Calibration in Module 1,2 and 4 Performed preventive maintenance on genexpert machine | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031486169841785883/6031486169841785883_signature_28.png | Dexter Tinoy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6031486169841785883/6031486169841785883_signature_31.png
| 2024-09-26 | 2024-09-26 | Claveria Rural Health Unit | Claveria, Masbate | 5 | Masbate | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29468 | DOH | narsclav@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 722 |  | GX4-6C | 838572 | UPS no power when cut from AC source. | Replaced UPS batteries (2x) Tested UPS performance (passed) Cleaned cartridge bays, plunger rod Generate Data | Replaced UPS batteries Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031463315315812672/6031463315315812672_signature_28.png | Ivy Garcia Penano | https://www.jotform.com/uploads/macarewecare/222927978372067/6031463315315812672/6031463315315812672_signature_31.png
| 2024-09-26 | 2024-09-26 | Sarangani Health Care Facility | Alabel, Sarangani | 12 | Sarangani | Tb-Lab | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29467 | DOH | ronayadao@yahoo.com | 22.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 749 | 2024-09-26 | GxIV-6C | 838542 | Module 2 & 3 H/W failed | Perform Genexpert machine repair Perform Genexpert annual Xpertcheck/Calibration Perform Genexpert preventive maintenance | Replace/repair defective module 2 SN: 862568 with new module SN: 705626 Replace/repair defective module 3 SN: 714753 with new module SN: 671228 Perform Genexpert Calibration/Xpertcheck Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: Module 2 & 3 - Skipped (new module) Module 4 - Passed Module 1 - Service required |  | -Genexpert 6-color module SN:705626 SN:671228 -Xpertcheck cartridge/kit Lot#:1001075092 | DR#: 12025, 12068, and 12072 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031425013653652719/6031425013653652719_signature_28.png | Ronora Yadao | https://www.jotform.com/uploads/macarewecare/222927978372067/6031425013653652719/6031425013653652719_signature_31.png
| 2024-09-26 | 2024-09-26 | Quezon Rural Health Unit | Poblacion, Quezon | 10 | Bukidnon | RTDL | NO | CALIBRATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29466 | DOH | quenxlopez@gmail.com | 27 |  |  | GX-4 6 Colors | 838440 | B2 found Service Required thru XPC | For reporting | XpertCheck Calibration | All modules and units and functional 3 modules passed, 1 module for service required in XPC |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031409316215287492/6031409316215287492_signature_28.png | Quenie Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6031409316215287492/6031409316215287492_signature_31.png
| 2024-09-26 | 2024-09-26 | Western Visayas Sanitarium | Sta. Barbara, Iloilo | 6 | Iloilo | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29465 | DOH | ruthiedequilla@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1806 | 2024-09-03 | 6C GCIV-4D | 838393 | Due for XPC | #NAME? | #NAME? | All modules passed XpertCheck. Machine is functional. |  | XpertCheck, CE-IVD 5 Cartridges/kit 101120421 June 01, 2025 | DR no. 11956 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031397012339617684/6031397012339617684_signature_28.png | Ruth Dequilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6031397012339617684/6031397012339617684_signature_31.png
| 2024-09-26 | 2024-09-26 | Western Visayas Sanitarium | Sta. Barbara, Iloilo | 6 | Iloilo | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29464 | DOH | ruthiedequilla@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1806 | 2024-09-03 | 6C GCIV-4D | 838393 | Due for XPC | #NAME? | #NAME? | All modules passed XpertCheck. Machine is functional. |  | XpertCheck, CE-IVD 5 Cartridges/kit 101120421 June 01, 2025 | DR no. 11956 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031396942338344742/6031396942338344742_signature_28.png | Ruth Dequilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6031396942338344742/6031396942338344742_signature_31.png
| 2024-09-26 | 2024-09-26 | Ramon magsaysay rural health unit | Ramon magsaysay, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29463 | DOH | niagnozia54@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 561 | 2024-09-05 | GX 4 (6c) | 838443 | None | Module 1,2 and 4 undergone Xpertcheck calibration. Data gathering. | Module 1,2 and 4 undergone Xpertcheck calibration. | All modules are available and ready to use. All modules passed the Xpertcheck Calibration. | All modules passed the Xpertcheck Calibration | XPERTCHECK-CE-5 | 12038 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031381163313429372/6031381163313429372_signature_28.png | Gina Borinaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6031381163313429372/6031381163313429372_signature_31.png
| 2024-09-26 | 2024-09-26 | Nueva Valencia Rural Health Unit | Nueva Valencia | 6 | Guimaras | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29462 | DOH | xxxrhingringxxx@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 369 | 2024-09-26 | GXIV | 838291 | Due for Calibration | #NAME? | #NAME? | Unit is Operational |  | Xpertcheck kit | 11953 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031379745013563160/6031379745013563160_signature_28.png | RHEA APRIL JALECO | https://www.jotform.com/uploads/macarewecare/222927978372067/6031379745013563160/6031379745013563160_signature_31.png
| 2024-09-26 | 2024-09-26 | NUEVA VALENCIA RURAL HEALTH UNIT | Nueva Valencia, Guimaras | 6 | Iloilo | TB DOTS | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29461 | UNOPS | xxxrhingringxxx@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-12 | GX-IV | 110023537 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  | 12132 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031374867697257087/6031374867697257087_signature_28.png | Rhea April Q Jaleco | https://www.jotform.com/uploads/macarewecare/222927978372067/6031374867697257087/6031374867697257087_signature_31.png
| 2024-09-26 | 2024-09-26 | GUIMARAS PROVINCIAL HEALTH OFFICE | Guimaras , Iloilo | 6 | Iloilo | Laboratory | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29460 | UNOPS | mylene210@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-12 | GX-IV | 110023528 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  | 12123 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031301580958049752/6031301580958049752_signature_28.png | Mylene Publico | https://www.jotform.com/uploads/macarewecare/222927978372067/6031301580958049752/6031301580958049752_signature_31.png
| 2024-09-26 | 2024-09-25 | WESTERN VISAYAS MEDICAL CENTER | Iloilo City | 6 | Iloilo | Laboratory | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | jkrjalandoni@gmail.com |  | 29459 | UNOPS | wvmcrhivdacrcl@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-12 | GX-IV | 110023529 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  | 12124 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030576046751004957/6030576046751004957_signature_28.png | Jessa Krista Riina Jalandoni | https://www.jotform.com/uploads/macarewecare/222927978372067/6030576046751004957/6030576046751004957_signature_31.png
| 2024-09-26 | 2024-09-26 | Escalante City Health Office | Poblacion Escalante | 6 | Negros Occidental | TB Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | sheila.negrido@yahoo.com |  | 29458 | DOH | boy928_sink13@yahoo.com | 31.6 |  | 2024-09-26 | GX4 6C | 838523 | GeneXpert Machine due for Xpertcheck. | .Save file GXC .Save file System Log .Save file Installation qualification reports .Save file assay statistics report. | Perform Xpertcheck for all module's | Aircon defective | Room temperature is higher need to repair aircon. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031279665383830103/6031279665383830103_signature_28.png | Shiela Negrido / Jireh Cabus | https://www.jotform.com/uploads/macarewecare/222927978372067/6031279665383830103/6031279665383830103_signature_31.png
| 2024-09-26 | 2024-09-26 | Bukidnon Provincial Hospital-Kibawe | CM Recto, Palma, Kibawe | 10 | Bukidnon | RTDL | NO | CALIBRATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29457 | DOH | juliealte10@gmail.com | 28 |  |  | GX-4 6 Colors | 838547 | None | None | XpertCheck Calibration | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6031275392884359142/6031275392884359142_signature_28.png | Julie Alte | https://www.jotform.com/uploads/macarewecare/222927978372067/6031275392884359142/6031275392884359142_signature_31.png
| 2024-09-25 | 2024-09-25 | Valencia City Health Office | New Terminal Complex, Hagkol, Valencia | 10 | Bukidnon | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29456 | DOH | caluyajopay@gmail.com | 27 |  |  | GX-4 6 Colors | 838271 | Module B3 found Error in Temperature and was error checked thru XPC Module B4 found Service Required | Module Replacement Calibration | Replacing module B3 with SN: 808071 to SN: 800646 XPC Calibration of 3 modules | All modules and units are functional. 1 module for replacement from XPC Result |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030737511204544067/6030737511204544067_signature_28.png | Joefelyn Caluya | https://www.jotform.com/uploads/macarewecare/222927978372067/6030737511204544067/6030737511204544067_signature_31.png
| 2024-09-25 | 2024-09-25 | Polomolok Rural Health Unit | Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29454 | DOH | jmmirasolrmt@gmail.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1679 | 2024-09-25 | GxIV-6C | 838265 | Module 1 and 2 H/W failed | Perform Genexpert module repair Perform Genexpert Calibration Perform Genexpert Preventive maintenance | Replace defective module SN: 662331 with new module SN:710656 Replace defective module SN: 711533 with new module SN:633525 Perform Genexpert machine Calibration/Xpertcheck Perform Genexpert Cleaning Checking of errors Data collection | 2 operational modules Calibration results: Module 1&3 - skipped new module Module 2&4 - Service Required |  | Genexpert 6-color module SN:633525 and SN:710656 Xpertcheck cartridge/kit Lot#: 1001075092 | DR#:12024,12069, and 12070 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030657276867382205/6030657276867382205_signature_28.png | June Miichele Mirasol | https://www.jotform.com/uploads/macarewecare/222927978372067/6030657276867382205/6030657276867382205_signature_31.png
| 2024-09-25 | 2024-09-24 | Lake Sebu Rural Health Unit | Poblacion, Lake Sebu, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29455 | DOH | rhulablakesebu@gmail.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 542 | 2024-09-24 | GxIV-6C | 838569 | None | Perform Genexpert machine Calibration/Xpertcheck Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Calibration/Xpertcheck Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 operational modules Calibration Results: Module 4 - skipped (New module) Module 1,2,and 3 - Passed |  | Xpertcheck cartridge/kit Lot#: 1001075092 | DR#:12020 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030657286869003140/6030657286869003140_signature_28.png | Reynellen Eborde | https://www.jotform.com/uploads/macarewecare/222927978372067/6030657286869003140/6030657286869003140_signature_31.png
| 2024-09-25 | 2024-09-25 | Tapaz Rural Health Unit | Tapaz, Capiz | 6 | Capiz | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29453 | DOH | chadmayrs@gmail.com | 26 |  | 2024-09-02 | GX-4 | 837615 | Due for Xpertcheck/ PM | PM/maintenance | Performed Xpertcheck on all 4 modules | Modules A1, A3, and A4 passed Xpertcheck. Module A2 for replacement |  |  | 5424 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030584363676464956/6030584363676464956_signature_28.png | Chad May Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6030584363676464956/6030584363676464956_signature_31.png
| 2024-09-25 | 2024-09-25 | WESTERN VISAYAS MEDICAL CENTER | Iloilo City | 6 | Iloilo | Laboratory | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | jkrjalandoni@gmail.com |  | 29452 | DOH | wvmcrhivdacrcl@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-12 | GX-IV | 110023529 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  | 12124 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030576046751004957/6030576046751004957_signature_28.png | Jessa Krista Riina Jalandoni | https://www.jotform.com/uploads/macarewecare/222927978372067/6030576046751004957/6030576046751004957_signature_31.png
| 2024-09-25 | 2024-09-25 | Macare In House | 1005 Aurora grand,Aurore blvd. QC | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | OTHERS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com | UPS Accupower | 29451 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  | 2024-09-20 | 1kva | AT1000180910ST0289, AT1000181128ST0216, AT10001901123ST0040,AT1000181128ST0144,AT1000180910ST0672,AT1000180910ST0672,AT1000170815ST0035,AT1000180910ST0185 | For replacement of ups battery,defective battery | Replacement and installed ups battery Pefrormed functionality test Test run unit | Performed battery replacement | Operational |  | 12v,7AH Battery ups 14pcs for 7units of UPS | G7386 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030561107025651831/6030561107025651831_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6030561107025651831/6030561107025651831_signature_31.png
| 2024-09-25 | 2024-09-25 | Dumangas Rural Health Unit | Dumangas, Iloilo | 6 | Iloilo | TB Dots | NO | INSTALLATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29450 | UNOPS | municipalhealthoffice_dumangas@yahoo.com | 24 |  | 2024-09-12 | 10C GXIV-4D | 110023532 | None | None | -GeneXpert Unit Installation -Setup GeneXpert Workstation -Setup GeneXpert software -Install XpertCheck 1.5b | Machine is functional. |  | -GeneXpert IV R2 4-Modules, 10-Color SN. 110023532 -GeneXpert Desktop SN. CPH1023F46952 -GeneXpert Monitor Display SN. CN0WMYC5FCC0037KCKAXA05 -GeneXpert Barcode Scanner SN. 22355010550941 -Accupower 1000VA Line Interactive UPS (AT1000) SN. AT1K231111ST0013 -Printer SN. CN3B2431Q9 | DR no. 12127 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030549783744510030/6030549783744510030_signature_28.png | Marilyn P. Dusaran | https://www.jotform.com/uploads/macarewecare/222927978372067/6030549783744510030/6030549783744510030_signature_31.png
| 2024-09-25 | 2024-09-25 | District C - Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29449 | DOH | miniforestlaboratory@gmail.com | 23.5 |  |  | GX-IV 6C 4D | 838507 | Stuck Cartidge at Module 4 | Remove the stuck cartidge at Module 4, clean cartridge bay Generate IQR, Assay Stat, and System Log | Perform Stuck Cartidge extraction | Successfully removed the stuck cartidge All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030473118718202619/6030473118718202619_signature_28.png | Lolijell Torralba | https://www.jotform.com/uploads/macarewecare/222927978372067/6030473118718202619/6030473118718202619_signature_31.png
| 2024-09-25 | 2024-09-25 | Miagao Rural Health Unit | Brgy. Guibongan, Miagao | 6 | Iloilo | TB-DOTS | NO | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29448 | UNOPS | joannaleesadiong@gmail.com | 28 |  | 2024-09-25 | GXIV 10C | 110023531 | n/a | n/a | Install Genexpert machine SN:110023531 CPU SN:CPH1023F46957 Monitor SN: CNOWMYC5FCC0037KCK9XA05 UPS SN:AT1K231111ST0032 Printer SN: CN3B2431HH | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030430385803093475/6030430385803093475_signature_28.png | Joanna lee M. Sadiong | https://www.jotform.com/uploads/macarewecare/222927978372067/6030430385803093475/6030430385803093475_signature_31.png
| 2024-09-25 | 2024-09-25 | San Rafael Rural Health Unit | San Rafael, Iloilo | 6 | Iloilo | Laboratory | NO | INSTALLATION | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29447 | DOH | apilmayfaith@gmail.com | 26 |  | 2024-09-12 | GX-4 10C | 110023536 | None | Copy old patient data from old Genexpert unit. | #NAME? | 4 modules operational |  |  | DR 12131 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030429632374507172/6030429632374507172_signature_28.png | May Faith Apil, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6030429632374507172/6030429632374507172_signature_31.png
| 2024-09-25 | 2024-09-25 | Dinas Rural Health Unit | Dinas, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29446 | DOH | Bernantereynacres@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 646 | 2024-09-05 | GX 4 (6C) | 838533 | Module 4 h/w failed. | Module 2 and 3 undergone Xpertcheck calibration. | Module 2 and 3 undergone Xpertcheck calibration. | Modules 1 2 and 3 are all available ang ready to use | Module 2 and 3 passed Xpertcheck Calibration | XPERTCHECK-CE-5 | 12033 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030424484219275736/6030424484219275736_signature_28.png | Reyna Cres Bernante-Lazaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6030424484219275736/6030424484219275736_signature_31.png
| 2024-09-25 | 2024-09-25 | Cawayan Rural Health Unit | Poblacion, Cawayan. Masbate | 5 | Masbate | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29445 | DOH | clenettedollison@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 888 |  | GX4-6C | 838267 | UPS no power when plugged from AC source | Replaced defective UPS batteries Tested UPS after battery replacement (charged for 1hr) Cleaned cartridge bays, plunger rod, Generate data | Replaced UPS batteries Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030422845319444534/6030422845319444534_signature_28.png | Clenette Nyza A. Dollison | https://www.jotform.com/uploads/macarewecare/222927978372067/6030422845319444534/6030422845319444534_signature_31.png
| 2024-09-25 | 2024-09-25 | OTON MUNICIPAL HEALTH OFFICE | OTON, ILOILO | 6 | Iloilo | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29444 | UNOPS | gracie05191980@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-12 | GX-IV | 110023530 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  | 12125 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030384467567044900/6030384467567044900_signature_28.png | Mary Grace Ong-Lumasag | https://www.jotform.com/uploads/macarewecare/222927978372067/6030384467567044900/6030384467567044900_signature_31.png
| 2024-09-25 | 2024-09-25 | Tagoloan RHU | Poblacion, Tagoloan | 10 | Misamis Oriental | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29443 | DOH | macalisangglezeal@gmail.com | 27 |  |  | GX-4 6 Module | 838574 | Module B3 found Error in Temperature | Module Replacement Calibration | Replacing Module B3 with SN: 807571 to SN: 684890 XPC Calibration of 3 modules since 1 module was newly installed | All module and units are functional All modules passed in calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6030381957178838611/6030381957178838611_signature_28.png | Venz Michael Bacoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6030381957178838611/6030381957178838611_signature_31.png
| 2024-09-24 | 2024-09-24 | Southern Philippines Medical Center | JP. Laurel Avenue, Bajada Davao City | 11 | Davao del Sur | ICU 2 | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 29442 | COMMERCIAL | icu2estuya@gmail.com | 22.6 |  | 2024-09-17 | Haemonetics MCS+ | 15H034 | Cmos BIOS Failure at CPU Board | Remove the defective Battery and replaced it with new Battery. Set-up the acceptable parameters for operating procedure and perform Cycle test to assure its functionality. | Perform Battery Replacement Perform Calibration Perform Functionality Test | Passed the Cycle Test, Unit is Operational |  | Lithium Battery LTC-7PN | DR 12142 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029778258719647742/6029778258719647742_signature_28.png | Dianne Karla de leon | https://www.jotform.com/uploads/macarewecare/222927978372067/6029778258719647742/6029778258719647742_signature_31.png
| 2024-09-24 | 2024-09-24 | Esperanza Rural Health Unit | Esperanza, Masbate | 5 | Masbate | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29441 | DOH | floramarmoncada@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1177 |  | GX4-6C | 838268 | B1/210068697, B3/210068696 - service required | Replaced module: B1 - 210068697 >> 220004891 B3 - 210068696 >> 887325 Cleaned cartridge bays, plunger, Gx fan Generate Data | Replaced module B1/210068697, B3/210068696 Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029705357984226276/6029705357984226276_signature_28.png | RHU-ESPERANZA | https://www.jotform.com/uploads/macarewecare/222927978372067/6029705357984226276/6029705357984226276_signature_31.png
| 2024-09-24 | 2024-09-24 | Jamindan Rural Health Unit | Jamindan, Capiz | 6 | Capiz | TB Laboratory | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29440 | DOH | louiseannaries.laa@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-4 | 110023543 | New Installation | New Installation | - Installed new 10c unit. - User orientation on basic machine operation, maintenance, cleaning, plunger rod maintenance. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029703324702661679/6029703324702661679_signature_28.png | Louise Ann Aries, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6029703324702661679/6029703324702661679_signature_31.png
| 2024-09-24 | 2024-09-24 | Maayon Rural Health Unit | Pob. Ilaya, Maayon | 6 | Capiz | TB-DOTS | NO | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29439 | UNOPS | pamskii_28@yahoo.com | 29 |  | 2024-09-24 | GXIV | 110023541 | n/a | n/a | Install Genexpert machine SN:110023541 CPU SN: CPH1023F7090 Monitor SN:CN0WMYC5CC00398EUPXA05 Barcode Sanner SN: 22355010550931 UPS SN: AT1K231111ST0049 Printer SN:CN3B241CF Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029701689154807734/6029701689154807734_signature_28.png | pamela cerujano | https://www.jotform.com/uploads/macarewecare/222927978372067/6029701689154807734/6029701689154807734_signature_31.png
| 2024-09-24 | 2024-09-24 | Carles Rural Health Unit | Carles, Iloilo | 6 | Iloilo | Laboratory | NO | INSTALLATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29438 | UNOPS | vangieyuro3@gmail.com | 25 |  | 2024-09-12 | 10C GXIV-4D | 110023526 | None | None | -New GeneXpert Unit and Peripherals Installation -Setup GeneXpert Workstation -Setup GeneXpert Software -Install Printer -Install XpertCheck 1.5b | Machine is functional. |  | -GeneXpert IV R2 10-Color SN. 110023526 -GeneXpert Desktop SN. CPH1023F46951 -GeneXpert Monitor Display SN. CN0WMYC5FCC0037KCK8XA05 -GeneXpert Barcode Scanner SN. 23103010550426 -Accupower 1000VA Line Interactive UPS AT1000 SN. AT1K231111ST0029 -Printer SN. CN3B2431V8 | DR no. 12121 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029660831024957156/6029660831024957156_signature_28.png | Vangie Yuro | https://www.jotform.com/uploads/macarewecare/222927978372067/6029660831024957156/6029660831024957156_signature_31.png
| 2024-09-24 | 2024-09-24 | St lukes'Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 29437 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 31 | ASSAY: Xpert HBV VL, COUNT: 109 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert HPV, COUNT: 358 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 27 |  | GX4 L | 807155 | For Maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | C4 disabled C1,C2,C3 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029593991014997664/6029593991014997664_signature_28.png | Jeinz Quinnie Alcala | https://www.jotform.com/uploads/macarewecare/222927978372067/6029593991014997664/6029593991014997664_signature_31.png
| 2024-09-24 | 2024-09-24 | St luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed.qc@stlukes.com.ph |  | 29436 | COMMERCIAL | megapo@stlukes.com.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 266 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3179 |  | GX4 D | 839234 | As per end user they encounter 5007 error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Troubleshooting | B1 for observation. B3 and B4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029593631018898957/6029593631018898957_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/6029593631018898957/6029593631018898957_signature_31.png
| 2024-09-24 | 2024-09-24 | Zamboanga del sur medical center | Zamboanga del sur medical center | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29435 | DOH | zdsmc.clinlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1686 ASSAY: Xpert HIV-1 Viral Load, COUNT: 116 | 2024-08-29 | GX 4 (6c) | 838204 | Module 2 and 4 hardware failed | Replaced defective modules 690804 and 752962 with new module 706303 and 710393 respectively. Module 1 Undergone Xpertcheck calibration. | Replaced defective modules 690804 and 752962 with new module 706303 and 710393 respectively. Module 1 Undergone Xpertcheck calibration. | Module 2, 3 and 4 are available and ready to use while module 1 Undergone Xpertcheck calibration and the result is service required. | Module 1 failed xpertcheck calibration, service required and needs to be replaced. | 900-0491RS , XPERTCHECK-CE-5 | 11857, 11852, 12039 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029571392611438269/6029571392611438269_signature_28.png | Jane L.Pactol | https://www.jotform.com/uploads/macarewecare/222927978372067/6029571392611438269/6029571392611438269_signature_31.png
| 2024-09-24 | 2024-09-23 | Bagumbayan Rural Health Unit | Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29434 | DOH | marizalenoid@yahoo.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 973 | 2024-09-23 | GxIV-6C | 838261 | Module 1 for Service Required | Perform Genexpert machine Calibration Perform Genexpert machine Preventive maintenance | Perform Genexpert Xpertcheck/Calibration Perform Genexpert Cleaning and maintenance Gx ID sticker installation Checking of errors Data collection | 3 operational modules Calibration Results: Module 2 skipped - new module Module 1 service Required Module 3 and 4 Passed |  | Xpertcheck cartridge/kit Lot: 1001075092 | DR#: 12015 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029499390519557326/6029499390519557326_signature_28.png | Ma. Riza R. Dionela | https://www.jotform.com/uploads/macarewecare/222927978372067/6029499390519557326/6029499390519557326_signature_31.png
| 2024-09-24 | 2024-09-24 | Pontevedra Rural Health Unit | Pontevedra, Capiz | 6 | Capiz | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29433 | DOH | rtabngjr@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1018 | 2024-09-02 | 6C GXIV-4D | 838296 | Module B1 defective | -Replace defective module B1/694490 with new module SN: 210046361 | #NAME? | All modules are functional. |  | Certified pre-owned GeneXpert module SN. 210046361 | DR no. 11916 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6029475783382043888/6029475783382043888_signature_28.png | Reynaldo Tabang, Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6029475783382043888/6029475783382043888_signature_31.png
| 2024-09-23 | 2024-09-23 | Dumarao Rural Health Unit | Dumarao, Capiz | 6 | Capiz | TB Laboratory | NO | INSTALLATION | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29432 | DOH | doregoshielamarie@gmail.com | 25 |  |  | GX-4 10c | 110023542 | None | Checked plunger rod operation, cleaned module bays, filter | - Installed new Genexpert, 10c unit - Orient end user on basic Operation, maintenance, archiving. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028885748169559094/6028885748169559094_signature_28.png | Shiela Marie Dorego | https://www.jotform.com/uploads/macarewecare/222927978372067/6028885748169559094/6028885748169559094_signature_31.png
| 2024-09-23 | 2024-09-23 | CUARTERO RURAL HEALTH UNIT | Cuartero, Capiz | 6 | Capiz | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29431 | UNOPS | cornepillomars@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-09-12 | GX-IV | 110022580 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  | 12139 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028880668127428773/6028880668127428773_signature_28.png | Marilyn M. Cornepillo | https://www.jotform.com/uploads/macarewecare/222927978372067/6028880668127428773/6028880668127428773_signature_31.png
| 2024-09-23 | 2024-09-23 | Calinog Rural Health Unit | Brgy. Libot, Calinog | 6 | Iloilo | TB-DOTS | NO | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29430 | UNOPS | icjssca@gmail.com | 28 |  | 2024-09-23 | GXIV 10C | 110023535 | n/a | #NAME? | - Install Genexpert machine SN: 110023535 - CPU SN: CPH1023F47091 - LCD Monitor SN: CN0WMYC5FCC0038EV8XA05 - UPS SN: AT1K231111ST0044 - Printer SN :CN3B431SH | Unit is Operational |  | Genexpert machine CPU Printer Monitor UPS Barcode | UNOPS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028817267124594925/6028817267124594925_signature_28.png | IC Jessica Coronado | https://www.jotform.com/uploads/macarewecare/222927978372067/6028817267124594925/6028817267124594925_signature_31.png
| 2024-09-23 | 2024-09-23 | Malay Primary Care Facility | Caticlan, Malay, Aklan | 6 | Aklan | Laboratory | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29429 | DOH | mhomalayclinicallaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1044 | 2024-09-02 | 6C GXIV-4D | 838457 | Module B2 defective | -Replace defective module B2/860115 with module SN: 685412 | #NAME? | All modules are functional. |  |  | DR no. 11926 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028779273476632649/6028779273476632649_signature_28.png | REYO KRISTOF C. TAYCO | https://www.jotform.com/uploads/macarewecare/222927978372067/6028779273476632649/6028779273476632649_signature_31.png
| 2024-09-23 | 2024-09-23 | Malay Primary Care Facility | Caticlan, Malay, Aklan | 6 | Aklan | Laboratory | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29428 | DOH | mhomalayclinicallaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1044 | 2024-09-02 | 6C GXIV-4D | 838457 | Module B2 defective | -Replace defective module B2/860115 with module SN: 685412 | #NAME? | All modules are functional. |  |  | DR no. 11926 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028778843479643870/6028778843479643870_signature_28.png | REYO KRISTOF C. TAYCO | https://www.jotform.com/uploads/macarewecare/222927978372067/6028778843479643870/6028778843479643870_signature_31.png
| 2024-09-23 | 2024-09-23 | Padada Rural Health Unit | Rizal Street, Poblacion Padada | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29427 | DOH | zhakie.marvz@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 579 | 2024-09-04 | GX-IV 6C 4D | 838511 | For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Upgrade the XPC Software to Version 1.5B Run Calibration at Four Modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration and Operational |  | XPERTCHECK CE 5 | 12006 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028705018717471232/6028705018717471232_signature_28.png | Erica Jane O. Gamboa, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6028705018717471232/6028705018717471232_signature_31.png
| 2024-09-23 | 2024-09-20 09:43 | UST Hospital | España Blvd., Sampaloc, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29424 | MACARE-HBDC | usth_bloodbank@yahoo.com | 25 |  | 2024-09-17 | MCS+ | 02F069 | Blood Pump Roller Worn-out. MCS Variocar Problem. Unit is due for Calibration. Unit is due for Maintenance. | Preventive Maintenance. Parts Replacement. Calibration. | Cleaned Surface Unit. Replaced Pumps Roller and Variocar. Run Simulation Test. Checked A/D and Calibration Values. Replaced Filter. | All Parameters Pass. Unit is Operational and in Good Condition. |  | MCS+ Variocar | 12140 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997021456803/6028675997021456803_signature_28.png | Mylene Balboa | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997021456803/6028675997021456803_signature_31.png
| 2024-09-23 | 2024-09-18 11:40 | Research Institute for Tropical Medicine | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29426 | MACARE-HBDC | hiv.laboratory.ritm@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1138 |  | GXIV-L-10C | 110017869 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997021481279/6028675997021481279_signature_28.png | Kriztianne Giron | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997021481279/6028675997021481279_signature_31.png
| 2024-09-23 | 2024-09-19 11:04 | Universitt of Perpetual Help Dalta Medical Center | Alabang-Zapote Road, Las Piñas | NCR | NCR 4th - Las Piñas | Molecular | NO | MAINTENANCE Relocation | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29425 | COMMERCIAL | phmclp.mdl@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 18 ASSAY: Others, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 36 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 402 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GXIV-D-6C | 831316 | GX Systen for Relocation. Unit is due for Maintenance. | Preventive Maintenance. | Relocate GX System from Micro to Molecular. Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Cleaned and Replace Filter. Performed Self-test. Generate Reports. | 1 Modules is Disabled with 5007,2008 Error. 3 Modules are Operational. | 2 Modules for Quotation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997023613637/6028675997023613637_signature_28.png | Donn Dave Clerigo | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997023613637/6028675997023613637_signature_31.png
| 2024-09-23 | 2024-09-20 12:55 | Rosario Rural Health Unit | Scout Torillo, Rosario, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29423 | DOH | rhu.rosario@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1143 | 2024-09-17 | GXIV-D-6C | 838537 | Monitor with Screen Problem. | Repair. | Replaced Defective Monitor. Performed Self-test. Adjust Monitor Settings. Generate Reports. | All Modules are Operational | CNK94301WX -> CN4610170L | HP Monitor | 12097 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997027719766/6028675997027719766_signature_28.png | Angelina T.Toledo | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997027719766/6028675997027719766_signature_31.png
| 2024-09-23 | 2024-09-18 11:42 | Research Institute for Tropical Medicine | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29422 | MACARE-HBDC | hiv.laboratory.ritm@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1259 |  | GXIV-L-10C | 110017870 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997028149830/6028675997028149830_signature_28.png | Kriztianne Giron | https://www.jotform.com/uploads/macarewecare/222927978372067/6028675997028149830/6028675997028149830_signature_31.png
| 2024-09-23 | 2024-09-19 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29421 | MACARE-HBDC | amymitra22@gmail.com | 22 |  |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -filter -icore Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028645127026576566/6028645127026576566_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/6028645127026576566/6028645127026576566_signature_31.png
| 2024-09-23 | 2024-09-19 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29420 | MACARE-HBDC | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 309 ASSAY: Others, COUNT: 21 ASSAY: Xpert HIV-1 Viral Load, COUNT: 51 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance. | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028644867024915137/6028644867024915137_signature_28.png | Marikina Public Health Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/6028644867024915137/6028644867024915137_signature_31.png
| 2024-09-23 | 2024-09-19 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29419 | MACARE-HBDC | amymitra22@gmail.com | 22 |  |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -filter -icore Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642977021888150/6028642977021888150_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642977021888150/6028642977021888150_signature_31.png
| 2024-09-23 | 2024-09-20 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29418 | UNOPS | acho.laboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 60 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642977023101200/6028642977023101200_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642977023101200/6028642977023101200_signature_31.png
| 2024-09-23 | 2024-09-19 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29417 | MACARE-HBDC | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 309 ASSAY: Others, COUNT: 21 ASSAY: Xpert HIV-1 Viral Load, COUNT: 51 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance. | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642977028185994/6028642977028185994_signature_28.png | Marikina Public Health Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642977028185994/6028642977028185994_signature_31.png
| 2024-09-23 | 2024-09-18 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29414 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 46 ASSAY: Xpert CT/NG, COUNT: 186 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 510 ASSAY: Xpert HCV VL, COUNT: 14 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937022332302/6028642937022332302_signature_28.png | Vincent Raquiza-Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937022332302/6028642937022332302_signature_31.png
| 2024-09-23 | 2024-09-18 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | YES | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29412 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 8 ASSAY: Xpert CT/NG, COUNT: 29 ASSAY: Xpert HPV, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 70 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937024570024/6028642937024570024_signature_28.png | Vincent Raquiza-Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937024570024/6028642937024570024_signature_31.png
| 2024-09-23 | 2024-09-18 | Ospital ng Makati | Sampaguita St. Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 29415 | MACARE-HBDC | jinx01gs@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 91 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937027966426/6028642937027966426_signature_28.png | Cristina M. Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937027966426/6028642937027966426_signature_31.png
| 2024-09-23 | 2024-09-19 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29413 | COMMERCIAL | mahmci_laboratory@gmail.com | 24 | ASSAY: Xpert HBV VL, COUNT: 17 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 252 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937029244518/6028642937029244518_signature_28.png | Sarah Lee Almeria | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937029244518/6028642937029244518_signature_31.png
| 2024-09-23 | 2024-09-18 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29416 | COMMERCIAL | ajmbalsomo@stlukes.com.ph | 24 |  |  | MCS+ | 09C125 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937029527522/6028642937029527522_signature_28.png | Aldwin Balsom9 | https://www.jotform.com/uploads/macarewecare/222927978372067/6028642937029527522/6028642937029527522_signature_31.png
| 2024-09-23 | 2024-09-23 | Manabo Rural Health Unit | Manabo, Abra | CAR | Abra | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29411 | DOH | lenmayclairelumiwag15@gmail.com | 25 |  | 2024-09-16 | GXIV-4-D-6C | 838274 | For CPU Replacement | - Replaced Defective CPU 2UA3520VNV with Repaired Original CPU CPH0520F0820 - Replaced Defective Mouse - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Generated Reports | Replaced Defective CPU | Unit is Fully Operational | Work with Engr. Venzon | GeneXpert CPUCPH0520F0820 | DR 12095 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028637778809429956/6028637778809429956_signature_28.png | Ellen May Claire S. Lumiwag | https://www.jotform.com/uploads/macarewecare/222927978372067/6028637778809429956/6028637778809429956_signature_31.png
| 2024-09-23 | 2024-09-23 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | La Union | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29410 | DOH | piaramuj@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 837 | 2024-09-10 | GXIV-4-D-6C | 838270 | For Module Replacement and Preventive Maintenance | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B2 210039876 >>> 737128 - Generated Reports | Performed Module Replacement and Preventive Maintenance | Unit is Fully Operational | Work with Engr. Venzon | GeneXpert Module 737128 | DR 12067 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028636758804041171/6028636758804041171_signature_28.png | Pia Jumar T. Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/6028636758804041171/6028636758804041171_signature_31.png
| 2024-09-23 | 2024-09-20 12:48 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29409 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 08B049 | Display board for troubleshooting Damaged stand wheel set | Adjusted display contrast | #NAME? | Unit is operational Display board for replacement request Stand for request | With Sir Harold Toco |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028607577025866946/6028607577025866946_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6028607577025866946/6028607577025866946_signature_31.png
| 2024-09-22 | 2024-09-22 | Samboan Rural Health Unit | Poblacion, Samboan Cebu | 7 | Cebu | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29408 | DOH | sheneannesenagan@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1470 |  | GX4 | 838513 | Due for Calibration | Generate IQR, Gxx files, System Log Performed calibration on genexpert machine Conduct Preventive Maintenance | Generate IQR, Gxx files, System Log Performed calibration on genexpert machine Conduct Preventive Maintenance | B1- Failed B2/B3/B4 - Passed All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028240933225593753/6028240933225593753_signature_28.png | Shene Anne A. Senagan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6028240933225593753/6028240933225593753_signature_31.png
| 2024-09-22 | 2024-09-17 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29407 | DOH | crmcmolecularlab@gmail.com | 19.8 | ASSAY: Xpert SARS-CoV-2, COUNT: 9590 | 2024-09-17 | GxIV-6C | 837662 | Module 2 and 4 for Service Required | Perform Genexpert machine module repair Perform Genexpert machine Preventive maintenance | Replace defective module 2 SN:210005670 with new module SN:618228 Replace defective module 4 SN:758411 with new module SN:623202 Genexpert machine Cleaning and maintenance Checking of errors Data collection | 4 functional modules |  | Genexpert 6-color module SN:618228 SN:623202 | DR #: 11885 and 11886 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6028196676862357223/6028196676862357223_signature_28.png | Darianne Mae Tubilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6028196676862357223/6028196676862357223_signature_31.png
| 2024-09-21 | 2024-09-21 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | avabichuela3216@gmail.com |  | 29406 | MACARE-HBDC | molecularlab@phc.gov.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 386 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 664 |  | GX16 D | 838655 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6027169154712556297/6027169154712556297_signature_28.png | Dianne Nicole V. Jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/6027169154712556297/6027169154712556297_signature_31.png
| 2024-09-21 | 2024-09-21 | Philippine Red Cross QC | Kalayaan Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29405 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 25 |  |  | MCS + | 02A227 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Tigthen line sensor screw cover Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6027168684713403917/6027168684713403917_signature_28.png | Lambert John Cordero | https://www.jotform.com/uploads/macarewecare/222927978372067/6027168684713403917/6027168684713403917_signature_31.png
| 2024-09-20 | 2024-09-20 | Bantayan Rural Health Unit | Bantayan, Cebu | 7 | Cebu | TB Dots | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29403 | DOH | mcdelle916bausin@gmail.com | 28 |  |  | 6C GXIV-4D | 838509 | CPU Date/Time resets every use | #NAME? | #NAME? | Unit is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026304087839252504/6026304087839252504_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/6026304087839252504/6026304087839252504_signature_31.png
| 2024-09-20 | 2024-09-20 | Manuel J. Santos Hospital | 554 Montilla Blvd. Butuan City | 13 | Agusan del Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29402 | COMMERCIAL | molecularlab@mjsantos.com.ph | 20 |  |  | GX-4 6 Colors | 831292 | 2 modules found Service Required thru XPC | For reporting | Preventive Maintenance XPC Calibration | 2 modules for replacement, 2 modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026260648834840762/6026260648834840762_signature_28.png | Raulyn Queen Tabañag | https://www.jotform.com/uploads/macarewecare/222927978372067/6026260648834840762/6026260648834840762_signature_31.png
| 2024-09-20 | 2024-09-20 | Estevez Memorial Hospital, Inc. | Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City, 4500 Albay | 5 | Albay | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29401 | COMMERCIAL | melosept22@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 163 ASSAY: Xpert MTB/RIF Ultra, COUNT: 99 ASSAY: Others, COUNT: 7 ASSAY: Xpert SARS-CoV-2, COUNT: 11 |  | GX4-10C | 110017942 | A2/220009328 - H/W Failed Gx for Preventive Maintenance | Checked IQR and system log for A2 Cleaned cartridge bays, plunger rod, icore Generate Data | Performed Preventive Maintenance Checked temperature and errors of A2 | Unit is operational | A2/220009328 for module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026255863615527851/6026255863615527851_signature_28.png | Estevez Memorial Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/6026255863615527851/6026255863615527851_signature_31.png
| 2024-09-20 | 2024-09-20 | FEU - Dr. Nicanor Reyes Medical Foundation | Dhalia St., West Fairview, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29400 | CONSORTIUM | kmtejoso@feu-nrmf.edu.ph | 25 | ASSAY: Xpert HPV, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 397 | 2024-09-18 | GX IV 6C-D | 840975 | Module A3 for replacement | Performed Preventive maintenance Replaced module A3 Generated report | Checked and cleaned cartridge bays plunger rod, icore slot and filter Replqced module A3 876314 -> 737004 | Unit is operational All modules working |  | GX Module 6C | DR 12145 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026243522024558465/6026243522024558465_signature_28.png | Kahlil Djunar Djames Silvestre | https://www.jotform.com/uploads/macarewecare/222927978372067/6026243522024558465/6026243522024558465_signature_31.png
| 2024-09-20 | 2024-09-20 | Hi-Precision Diagnostics | NS Amoranto Sr. St., Queozn City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29399 | CONSORTIUM | jose@hi-precision.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 135 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6601 |  | GX IV 6C-D | 835601 | Cepheid recommended maintenance due to error 5007 | Performed Preventive maintenance Generated report | Checked adn cleaned cartridge bays, plunger rod, icore slot and filter | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026239052025165662/6026239052025165662_signature_28.png | Andrei N. Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/6026239052025165662/6026239052025165662_signature_31.png
| 2024-09-20 | 2024-09-20 | Talomo South Health Center | Puan Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | dy_rosemarie71@gmail.com |  | 29398 | DOH | dyrosemarie71@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1601 | 2024-09-10 | GX-IV 6C 4D | 838401 | Module B4 encounter temperature related error For XPC Calibration | Replaced the Defective Module SN: 860974 with new Module SN: 724756 Run XPC on Module B3 only | Perform Module Replacement and XpertCheck Calibration | Successfully Replaced the Module B4 Module B3 Xpc result : Retry and for service required Module 3 for under observation if error persist, Enduser was advised to disable the said module |  | 900-0491RS XpertCheck CE-5 | DR 12071 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026230435574506482/6026230435574506482_signature_28.png | Rosemarie Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/6026230435574506482/6026230435574506482_signature_31.png
| 2024-09-20 | 2024-09-20 | Nasipit Rural Health Unit | Brgy. Talisay, Nasipit, Agusan del Norte | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29397 | DOH | peterdolotina@gmail.com | 26 |  |  | GX-4 6 Colors | 837611 | None | None | Preventive Maintenance XpertCheck Calibration | All modules and units are functional | All modules passed in XPC Calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026175469768187806/6026175469768187806_signature_28.png | Peter C. Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/6026175469768187806/6026175469768187806_signature_31.png
| 2024-09-20 | 2024-09-20 | Tubungan Rural Health Unit | San Antonio St. Tubungan, Iloilo | 7 | Iloilo | Laboratory | NO | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29396 | UNOPS | sheryltacadao@gmail.com | 22 |  |  | Gx4 | 110023534 | New Genexpert Installation End-User Basic training in genexpert operation | New Genexpert Installation End-User Basic training in genexpert operation | New Genexpert Installation End-User Basic training in genexpert operation | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026121502435623944/6026121502435623944_signature_28.png | SHERYL TACADAO | https://www.jotform.com/uploads/macarewecare/222927978372067/6026121502435623944/6026121502435623944_signature_31.png
| 2024-09-20 | 2024-09-20 | Hi-Precision Diagnostics | N.S. Amoranto Sr. St. Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29395 | CONSORTIUM | jose@hi-precision.com.ph | 24 | ASSAY: Xpert CT/NG, COUNT: 114 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7835 |  | GX IV 6C-D | 835602 | Cepheid recommended maintenance of machine | Performed Preventive maintenance Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026088216361169449/6026088216361169449_signature_28.png | Andrei N. Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/6026088216361169449/6026088216361169449_signature_31.png
| 2024-09-20 | 2024-09-20 | Western Visayas State Univerity Medical Center | Iloilo City | 6 | Iloilo | Molecular lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29394 | MACARE-HBDC | mc-molecular@wvsu.edu.ph | 26 |  |  | Gx4 10c | 110021716 | None | Expertcheck calibration Clean Fan Filter | Expertcheck calibration Clean Fan Filter | Systems Operational | All modules passed in xpertcheck calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026023006119762907/6026023006119762907_signature_28.png | Karla Danielle T. Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6026023006119762907/6026023006119762907_signature_31.png
| 2024-09-20 | 2024-09-20 | West Visayas State University Medical Center | Poblacion | 6 | Iloilo | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29393 | DOH | mc-molecular@wvsu.edu.ph | 24 |  | 2024-09-20 | GX4 | 838290 | Due for Xpertcheck | .Save Installation Qualification Reports .Save system Log files .Save GXC file Save Assay statistics Report | .Perform Xpertcheck for all module's. .Perform preventive maintenance. | All module's passed for Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6026022078135583749/6026022078135583749_signature_28.png | Karla Danielle T. Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6026022078135583749/6026022078135583749_signature_31.png
| 2024-09-19 | 2024-09-19 | Don Jose S. Monfort Medical Center Extension Hospital | Barotac Nuevo, Iloilo | 6 | Iloilo | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29392 | DOH | judyboglosa35@gmail.com | 29 |  |  | Gx4 | 838501 | Defective Module B1 760020 New defective module B3 862361 | New module installed B1 210018929 Clean Fan Filter IQR,system log and assay stats Clean Fan | New module installed B1 210018929 Clean Fan Filter IQR,system log and assay stats Clean Fan | Systems operational Defective module B3 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6025448296121604332/6025448296121604332_signature_28.png | Judy boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6025448296121604332/6025448296121604332_signature_31.png
| 2024-09-19 | 2024-09-19 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY, CEBU | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29391 | DOH | mendozarey42@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4604 ASSAY: Xpert HCV VL, COUNT: 43 ASSAY: Xpert HIV-1 Viral Load, COUNT: 313 | 2024-09-03 | GX-IV | 838510 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK, CE-IVD 5 Cartridges/kit | 11971 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6025369076831854912/6025369076831854912_signature_28.png | Rey C. Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6025369076831854912/6025369076831854912_signature_31.png
| 2024-09-19 | 2024-09-19 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY, CEBU | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29390 | DOH | mendozarey42@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4604 ASSAY: Xpert HCV VL, COUNT: 43 ASSAY: Xpert HIV-1 Viral Load, COUNT: 313 | 2024-09-03 | GX-IV | 838508 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERCHECK, CE-IVD 5 Catridges/kit | 11970 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6025365476836443412/6025365476836443412_signature_28.png | Rey C. Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6025365476836443412/6025365476836443412_signature_31.png
| 2024-09-19 | 2024-09-19 | Leon Rural Health Unit | Poblacion | 6 | Iloilo | Laboratory | NO | INSTALLATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29389 | DOH | camiquejolyn10@gmail.com | 23 |  | 2024-09-19 | GX4 D 10C | 110023533 | New GeneXpert Machine Instalation | Save file Installation Qualificatin Save System Log file | .Instal GeneXpert Machine Sr#. 110023533 .Desktop Sr#. CPH1023F47088 .Monitor Display SR. CN0WMYC5FCC0037KCLYA05 .Barcode Scanner Sr. 23103010550606 .UPS sr. AT1K231111ST0004 .Printer Sr. CN3B2431VH | All are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6025359813221656856/6025359813221656856_signature_28.png | Jolyn c. Maloco | https://www.jotform.com/uploads/macarewecare/222927978372067/6025359813221656856/6025359813221656856_signature_31.png
| 2024-09-19 | 2024-09-19 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29388 | DOH | buhanginlaboratory@gmail.com | 25.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 878 | 2024-09-04 | GX-IV 6C 4D | 838232 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Upgrade the version of XpertCheck to 1.5B Run XPC on four Modules | Perform Preventive Maintenance and XpertCheck Calibration | Module B1 B3 : Retry and Passed Module B2 B4 :Passed All Modules Passed the Calibration |  | XPERTCHECK CE-5 | DR 11998 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6025337985573440510/6025337985573440510_signature_28.png | Catherine B. Deocampo,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6025337985573440510/6025337985573440510_signature_31.png
| 2024-09-19 | 2024-09-19 | Cebu South Medical Center | San Isidro, Talisay City Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | cscgenexpert@gmail.com |  | 29387 | DOH | csmcmicro@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 1147 ASSAY: Xpert HIV-1 Viral Load, COUNT: 62 |  | GX4 | 838559 | Module 1 SN # 806586 defective | Performed module replacement of defective module. New module installation SN # : 636355 Performed Maintenance of genexpert machine. Clean Air Filter Generate gxx files, IQR, Sys Log Back-up database. | Performed module replacement of defective module. New module installation SN # : 636355 Performed Maintenance of genexpert machine. Clean Air Filter Generate gxx files, IQR, Sys Log Back-up database. | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6025258203223875841/6025258203223875841_signature_28.png | Joseph Lorenzo , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6025258203223875841/6025258203223875841_signature_31.png
| 2024-09-19 | 2024-09-19 | Iloilo Provincial Hospital | Pototan | 6 | Iloilo | TB Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29386 | DOH | djpragale92@gmail.com | 24 |  | 2024-09-18 | GX4 | 838289 | . Module C2,C3 defective . Module C4 Hardware fail | . Save file system Log . Save file Installation Qualification . Save file assay statistics report | Replace module C2,C3 with new module's. | Module 1,2,3 ok to use. Module 4 not ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6025220543312170599/6025220543312170599_signature_28.png | Rachael Pragale | https://www.jotform.com/uploads/macarewecare/222927978372067/6025220543312170599/6025220543312170599_signature_31.png
| 2024-09-18 | 2024-09-18 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29385 | DOH | chaska.lp14@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 1563 |  | GX4 | 838283 | Schedule for Calibration for all module. | Conduct Calibration B1- Failed, B2/B3/B4- PASSED Performed Preventive Maintenance Clean Air Filter | Conduct Calibration B1- Failed, B2/B3/B4- PASSED Performed Preventive Maintenance Clean Air Filter | B1- Failed B2/B3/B4 - Passed All modules are still operational while waiting for replacement for B1. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6024724503228031345/6024724503228031345_signature_28.png | Lucille Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6024724503228031345/6024724503228031345_signature_31.png
| 2024-09-18 | 2024-09-18 | Minglanilla Rural Health Unit | Pider 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29384 | DOH | brendagonzagaparedes@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 1436 ASSAY: Xpert HIV-1 Viral Load, COUNT: 316 ASSAY: Xpert HPV, COUNT: 316 |  | GX4 | 838512 | Module 1 SN# 862340 defective due to Error 1002 | Module Replacement with New Module SN: 210058064 Generate Gxx files, Sys Log, IQR Conduct Preventive Maintenance Clean filter | Module Replacement with New Module SN: 210058064 Generate Gxx files, Sys Log, IQR Conduct Preventive Maintenance Clean filter | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6024719403224985498/6024719403224985498_signature_28.png | Brenda Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/6024719403224985498/6024719403224985498_signature_31.png
| 2024-09-18 | 2024-09-19 | Bantayan Rural Health Unit | Bantayan, Cebu | 7 | Cebu | Tb Dots | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29383 | DOH | mcdelle916bausin@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1000 | 2024-08-29 | 6C GXIV-4D | 838509 | #NAME? | -Replace defective module B2/629373 with module SN: 638306 -Run XpertCheck on module B3 | #NAME? | Module B3 for replacement. Module B1 and B4 are functional. |  | 1 Certified Pre-Owned GeneXpert Module SN. 638306 | DR no. 11892 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6024597829459119873/6024597829459119873_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/6024597829459119873/6024597829459119873_signature_31.png
| 2024-09-18 | 2024-09-18 | CEBU DOCTORS UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29382 | COMMERCIAL | cduhcovidpoct@gmail.com | 21 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 72 ASSAY: Xpert CT/NG, COUNT: 212 ASSAY: Xpert HBV VL, COUNT: 194 ASSAY: Xpert HPV, COUNT: 25 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1106 ASSAY: Xpert HCV VL, COUNT: 158 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 |  | GX-IV | 110008524 | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6024500523229760583/6024500523229760583_signature_28.png | Alyssa C. Sayson, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6024500523229760583/6024500523229760583_signature_31.png
| 2024-09-18 | 2024-09-18 | Mini Forest Health Center | Quezon Boulevard, Brgy 23-C Davao City | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29381 | DOH | miniforestlaboratory@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 400 | 2024-09-04 | GX-IV 6C 4D | 838507 | For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Upgrade the XPC to Version 1.5b Run XPC to four Modules | Perform Preventive Maintenance, and XpertCheck Calibration | Module B2 for Service Required Modules B1, B3 and B4 Passed Module B2 for under observation, if error persists enduser was advised to disable the said Module |  | XPERTCHECK CE-5 | DR 12000 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6024471570517472929/6024471570517472929_signature_28.png | Joy L. Nio,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6024471570517472929/6024471570517472929_signature_31.png
| 2024-09-18 | 2024-09-18 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29380 | DOH | burauenlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 459 | 2024-09-18 | GXIV | 838551 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6024375895112954083/6024375895112954083_signature_28.png | mary Ann C. Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/6024375895112954083/6024375895112954083_signature_31.png
| 2024-09-18 | 2024-09-18 | Medellin Rural Health Unit | Medellin, Cebu | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29379 | DOH | medellindtc@yahoo.com | 26 |  | 2024-09-03 | 6C GXIV-4D | 838506 | Due for XPC | #NAME? | #NAME? | All modules passed XpertCheck and is functional. |  | XpertCheck-CE-IVD | DR no. 11986 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6024299374148520996/6024299374148520996_signature_28.png | Raul M. Gallano | https://www.jotform.com/uploads/macarewecare/222927978372067/6024299374148520996/6024299374148520996_signature_31.png
| 2024-09-17 | 2024-09-13 13:51 | West Dadiangas Rural Health Unit | Brgy. West, Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29376 | DOH | renlizros11@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1269 | 2024-09-13 | GxIV-6C | 838570 | Module 1 for Service Required | Perform Genexpert repair Perform Genexpert machine Preventive maintenance | Repair/replace defective module 1 SN:644744 with new module SN:677306 Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 functional modules |  | Genexpert module 6-color SN:677306 | DR. Number: 11905 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6023694531222782946/6023694531222782946_signature_28.png | Renee Liz Rosa Gomez | https://www.jotform.com/uploads/macarewecare/222927978372067/6023694531222782946/6023694531222782946_signature_31.png
| 2024-09-17 | 2024-09-10 10:59 | Tampakan Rural Health Unit | Tampakan,South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29377 | DOH | jmedtech_lab@yahoo.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 641 | 2024-09-10 | GxIV-6C | 838544 | Module 4 for Service Required Module 2 under observation due to temperature errors | Perform Genexpert machine Preventive maintenance Genexpert machine repair | Replace/repair defective module 4 SN:862222 with new module SN:806160 Perform Genexpert Cleaning and maintenance Checking of errors/problem Problem/errors reporting Data collection | 4 operational modules |  | Genexpert module 6-color SN:806160 | Cepheid case: 1481068 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6023694531223306638/6023694531223306638_signature_28.png | Jovi Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/6023694531223306638/6023694531223306638_signature_31.png
| 2024-09-17 | 2024-09-11 13:02 | Datu Odin Sinsuat District Hospital | Poblacion, Datu Odin Sinsuat, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29378 | DOH | rowenasalazar1234@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 780 | 2024-09-11 | GxIV-6C | 838205 | Module 2 for Service Required | Perform Genexpert Module repair Perform Genexpert preventive maintenance | Repair/replace defective module 2 SN:807327 with new module SN:643611 Gx ID sticker installation Perform Genexpert Cleaning and maintenance Checking of errors Data collection | 4 functional modules |  | Genexpert module 6-color SN:643611 | Cepheid case:1527730 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6023694531227688942/6023694531227688942_signature_28.png | ROWENA B. SALAZAR | https://www.jotform.com/uploads/macarewecare/222927978372067/6023694531227688942/6023694531227688942_signature_31.png
| 2024-09-17 | 2024-09-17 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com | gfmselga@gmail.com |  | 29375 | DOH | ajbtrryes@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1069 |  | GX-4 | 838477 | Due for Xpertcheck | PM, cleaned optics, module bay, filter, pkunger rods. | Performed Xpertcheck in all modules. | All modules Passed Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6023629764415836149/6023629764415836149_signature_28.png | jose ali c. treyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6023629764415836149/6023629764415836149_signature_31.png
| 2024-09-16 | 2024-09-16 | University of the Philippines Los Baños - Institute of Biology | Institute of Biology Bldg., UP Los Baños | 4A | Laguna | Phycology Laboratory | NO | REPAIR TROUBLESHOOTING | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | rppadernal@up.edu.ph |  | 29374 | COMMERCIAL | mmartinezgoss@gmail.com | 23 |  | 2024-09-13 | Axio Scope.A1 | 3331000862 | HAL 100 light source does not power on | Replaced halogen bulb | Performed additional troubleshooting Checked fuses Checked output voltages of power supply | Unit is not operational. Isolated the power supply as the problem. Output is only at 4.5V when the output should be 12V | Recommended for Power Supply replacement | 12V 100W Halogen Bulb | 98762 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6022757402427421723/6022757402427421723_signature_28.png | Milagrosa R. Martinez-Goss | https://www.jotform.com/uploads/macarewecare/222927978372067/6022757402427421723/6022757402427421723_signature_31.png
| 2024-09-16 | 2024-09-16 10:50 | Mabinay Medicare Community Hospital | Mabinay, Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29373 | DOH | diannecm8@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 619 | 2024-09-03 | GX-4 | 838481 | Due for Xpertcheck | Cleaned filter, optics, plunger rod, module bays. | Paerformed PM and Xpertcheck. | Module 4 for replacement Modules 1, 2, 3 Passed Xpertcheck |  | Xpertcheck kit | 11969 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6022648198691228953/6022648198691228953_signature_28.png | Dianne C. Mi-ot, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6022648198691228953/6022648198691228953_signature_31.png
| 2024-09-16 | 2024-09-16 09:40 | Cardinal Santos Medical Center | Wilson st,San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 29372 | COMMERCIAL | florescanoel77@gmail.com | 26 |  | 2024-09-11 | MCS + | 08M461 | Encounter system error. Unusual rotation AC pump motor. | Performed repair Replaced and installed AC pump motor. Checked calibration parameters. Performed pump test Performed functionality test | Performed repair. Lent spare parts AC pump motor | Operational | For quotation of AC pump motor | AC pump motor assy | 12088 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6022603467023174484/6022603467023174484_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/6022603467023174484/6022603467023174484_signature_31.png
| 2024-09-16 | 2024-09-13 15:39 | Rosario Rural Health Unit | Scout Torillo, Rosario ,Cavite | 4A | Cavite | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29370 | DOH | arlenepenaarmijo@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1149 |  | GXIV-D-6C | 838537 | Hard Startup CPU. | Troubleshooting. | Checked Software System. Replaced and Clean RAM for Troubleshooting. Checked and Restart GX System. Run Diagnostic Tests for Memory, Storage, and Motherboard. Replaced Bios Battery. Generate Reports. | All Modules are Operational. CPU for further Observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6022565397023659541/6022565397023659541_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/6022565397023659541/6022565397023659541_signature_31.png
| 2024-09-16 | 2024-09-12 14:55 | UST Hospital | Lacson Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29371 | CONSORTIUM | ustimp@yahoo.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 962 |  | GXIV-D-6C | 110000711 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Unit's Surface Area. Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked if Filter is Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6022565397024137495/6022565397024137495_signature_28.png | Maria Liova Sherleen D. Calaunan | https://www.jotform.com/uploads/macarewecare/222927978372067/6022565397024137495/6022565397024137495_signature_31.png
| 2024-09-15 | 2024-09-11 08:53 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29369 | DOH | dianeruiz147@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 442 |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6022146161852802649/6022146161852802649_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6022146161852802649/6022146161852802649_signature_31.png
| 2024-09-15 | 2024-09-11 12:17 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29368 | DOH | arleneabad32@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 740 | 2024-09-02 | GXIV-4-D-6C | 838288 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Modules B1 808218 >>> 685424 B4 698768 >>> 210091816 | Performed Module Replacement | Unit is Fully Operational | Work with Engineer Venzon | GeneXpert Modules 685424 & 210091816 | DR 11915, DR 11925 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6022144661857632352/6022144661857632352_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/6022144661857632352/6022144661857632352_signature_31.png
| 2024-09-15 | 2024-09-11 08:45 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29367 | DOH | serly_reyes@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 724 | 2024-09-02 | GXIV-4-D-6C | 838516 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B4 862532 >>> 676927 - Generated Reports | Performed Preventive Maintenance | Unit is Fully Operational |  | GeneXpert Module 676927 | DR 11929 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6022143221855702111/6022143221855702111_signature_28.png | Marione Klaudette J. Farro | https://www.jotform.com/uploads/macarewecare/222927978372067/6022143221855702111/6022143221855702111_signature_31.png
| 2024-09-15 | 2024-09-11 08:41 | Umingan Community Hospital | Umingan Community Hospital | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29366 | DOH | jeffreyvrumbaoa@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1363 |  | GXIV-4-D-6C | 838473 | For Calibration and Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational Modules unavailable for Calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6022140731855459227/6022140731855459227_signature_28.png | Jeffrey V. Rumbaoa | https://www.jotform.com/uploads/macarewecare/222927978372067/6022140731855459227/6022140731855459227_signature_31.png
| 2024-09-14 | 2024-09-13 16:52 | Pilar Rural Health Unit | Brgy. Calongay, Pilar | 5 | Sorsogon | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29365 | DOH | manalangmark88@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1234 |  | GX4-6C | 838266 | B1/807683 for module replacement | Replaced module: B1 - 807683 >> 210028907 Cleaned cartridge bays, plunger rod, fan filter Collect data | Replaced module B1/807683 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6021013929945068966/6021013929945068966_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/6021013929945068966/6021013929945068966_signature_31.png
| 2024-09-13 | 2024-09-13 17:41 | Catigbian District Hospital | CATIGBIAN BOHOL | 7 | Bohol | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29364 | DOH | rheranes@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 628 | 2024-09-03 | GX-IV | 838494 | #NAME? | #NAME? | #NAME? | #NAME? |  | XPERTCHECK, CE-IVD 5 Cartridges/kit | 11962 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020307409311413657/6020307409311413657_signature_28.png | Rheazel G. Ranes | https://www.jotform.com/uploads/macarewecare/222927978372067/6020307409311413657/6020307409311413657_signature_31.png
| 2024-09-13 | 2024-09-12 15:09 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | biomed@nch.doh.gov.ph |  | 29363 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 48 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2574 ASSAY: Xpert SARS-CoV-2, COUNT: 1751 ASSAY: Xpert SARS-CoV-2, COUNT: 918 |  | GX16 (10C) | 110014171 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020201947027616523/6020201947027616523_signature_28.png | Jessica Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6020201947027616523/6020201947027616523_signature_31.png
| 2024-09-13 | 2024-09-13 13:34 | Kalinga Health Hub | JP Rizal St. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29362 | CONSORTIUM | pearl.balderama@innovationsch.org | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1527 |  | GX-IV-D-6C | 110008418 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020194667023798117/6020194667023798117_signature_28.png | Krisna Morante | https://www.jotform.com/uploads/macarewecare/222927978372067/6020194667023798117/6020194667023798117_signature_31.png
| 2024-09-13 | 2024-09-12 15:23 | Rosario Super Health Center | Dr. Sixto Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29361 | MACARE-HBDC | eden.cawayan@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 437 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020194477029900931/6020194477029900931_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6020194477029900931/6020194477029900931_signature_31.png
| 2024-09-13 | 2024-09-12 13:31 | Santolan Super Health Center | Evangelista Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29360 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 558 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020194337028135172/6020194337028135172_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6020194337028135172/6020194337028135172_signature_31.png
| 2024-09-13 | 2024-09-12 15:23 | Rosario Super Health Center | Dr. Sixto Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29357 | MACARE-HBDC | eden.cawayan@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 437 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020193927021383414/6020193927021383414_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/6020193927021383414/6020193927021383414_signature_31.png
| 2024-09-13 | 2024-09-13 13:34 | Kalinga Health Hub | JP Rizal St. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29359 | CONSORTIUM | pearl.balderama@innovationsch.org | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1527 |  | GX-IV-D-6C | 110008418 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020193927024354147/6020193927024354147_signature_28.png | Krisna Morante | https://www.jotform.com/uploads/macarewecare/222927978372067/6020193927024354147/6020193927024354147_signature_31.png
| 2024-09-13 | 2024-09-12 13:31 | Santolan Super Health Center | Evangelista Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29358 | MACARE-HBDC | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 558 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020193927029320579/6020193927029320579_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6020193927029320579/6020193927029320579_signature_31.png
| 2024-09-13 | 2024-09-13 14:11 | Macare In House | 1005 Aurora Grand,Aurora blvd | NCR | NCR 2nd - Quezon City | In House | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 29356 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | Lenovo M720S | CPH0520F08280 | CPU booting problem | Performed troubleshooting and repair. Reset all cable. Cleaned RAM Cleaned all data. Checked hard drive. Performed functionality test | Troubleshooting and repair | Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020173367024801520/6020173367024801520_signature_28.png | Macare In house | https://www.jotform.com/uploads/macarewecare/222927978372067/6020173367024801520/6020173367024801520_signature_31.png
| 2024-09-13 | 2024-09-13 13:51 | Atimonan Rural Health Unit | G. Boulevard, Brgy. Zone 1, Atimonan, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29355 | DOH | AMCL2008@yqhoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 783 | 2024-09-10 | GX IV 6C-D | 838206 | ModuleB3 for replacement | Performed Preventive maintenance Replaced module B3 Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced module B3 807295 -> 625372 | Unit is operational All modules working |  | GX Module 6C | DR 12075 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020162119561206154/6020162119561206154_signature_28.png | Marife Villamiel | https://www.jotform.com/uploads/macarewecare/222927978372067/6020162119561206154/6020162119561206154_signature_31.png
| 2024-09-13 | 2024-09-13 13:07 | Cebu City Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | CALIBRATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29354 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2079 | 2024-09-03 | 6C GXIV-4D | 838561 | Due for XPC | #NAME? | #NAME? | All modules passed XPC and is functional. |  | XpertCheck, CE-IVD 5 Cartridges/kit 1001120421 June 01, 2025 | DR no. 11964 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020152785488152223/6020152785488152223_signature_28.png | Jedieva Lumbab | https://www.jotform.com/uploads/macarewecare/222927978372067/6020152785488152223/6020152785488152223_signature_31.png
| 2024-09-13 | 2024-09-13 10:25 | TEODORO B. GALAGAR DISTRICT HOSPITAL | JAGNA BOHOL | 7 | Bohol | LABORATORY | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29353 | DOH | applecpalma09@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 770 | 2024-08-29 | GX-IV | 838292 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE B2 DEF SN:786200 REPLACEMENT WITH NEW SN:660740 -RUN XPERTCHECK CALIBRATION: ALL MODULES PASSED | EQUIPMENT IS OPERATIONAL |  | -GeneXpert Module -Xpertcheck, CE-IVD 5 Cartridges/kit | 11903, 11979 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6020039416173124494/6020039416173124494_signature_28.png | Therese Arjelin Palma, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6020039416173124494/6020039416173124494_signature_31.png
| 2024-09-12 | 2024-09-12 17:09 | Nueva Vizcaya Provincial Hospital | Almaguer, Bambang, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29352 | DOH | dumlaojoy4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1115 | 2024-08-29 | GXIV-6C | 838272 | Module 1 and 2 errors | Performed Module Replacement B1(650488 - 793131) B2(674508 - 794253) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Modules | 11868 & 11869 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019421640916233252/6019421640916233252_signature_28.png | Bernadette Dumlao | https://www.jotform.com/uploads/macarewecare/222927978372067/6019421640916233252/6019421640916233252_signature_31.png
| 2024-09-12 | 2024-09-12 16:52 | CANDIJAY COMMUNITY HOSPITAL | CANDIJAY BOHOL | 7 | Bohol | LABORATORY | NO | REPAIR MAINTENANCE GX ID Sticker Installation | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29351 | DOH | angelinaredulla@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 | 2024-08-29 | GX-IV | 838396 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE REPLACEMENT: B2 DEF SN: 860993 WITH NEW SN: 662120 B3 DEF SN: 860994 WITH NEW SN: 893616 | EQUIPMENT IS OPERATIONAL |  | GeneXpert module | 11877/11902 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019407414129274024/6019407414129274024_signature_28.png | ANGELINA O. REDULLA | https://www.jotform.com/uploads/macarewecare/222927978372067/6019407414129274024/6019407414129274024_signature_31.png
| 2024-09-12 | 2024-09-12 | Libon Rural Health Unit | Velzco St., Zone 1 Poblacion, Libon | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29350 | DOH | marilyn@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1094 |  | GX4-6C | 838397 | B2/860937 for module replacement | Replaced module: B2 - 860937 >> 645793 Cleaned cartridge bays, plunger rod, Gx fan Generate Data | Replaced module B2/860937 Performed Preventive Maintenance | Unit is Operational | UPS is battery problem, was already stored. Advised end user to collect data for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019369760815933249/6019369760815933249_signature_28.png | Marilyn S. Razal | https://www.jotform.com/uploads/macarewecare/222927978372067/6019369760815933249/6019369760815933249_signature_31.png
| 2024-09-12 | 2024-09-12 15:09 | Ciudad Medical Zamboanga | Mayor, Vitaliano Agan Ave, Zamboanga, 7000 Zamboanga del Sur | 9 | Davao del Sur | Laboratory | YES | INSTALLATION Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | supplychain@ciudadmedicalzamboanga.com.ph |  | 29349 | COMMERCIAL | thairaphoebegrey13@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 51 ASSAY: Xpert MTB/RIF Ultra, COUNT: 348 ASSAY: Xpert MTB/RIF Ultra, COUNT: 28 | 2021-09-06 | GX 4 (6c) | 210118423 | Module 2 gives an error of 1002, 2005 and 4009 all related to temperature problem. | Replaced module 781102 with new module 210118423. | Replaced module 781102 with new module 210118423. | All modules 1 and 2 are working properly. | None | 900-0508 | 98725 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019351868299726171/6019351868299726171_signature_28.png | Shereemer A.Cheong-Ynawat | https://www.jotform.com/uploads/macarewecare/222927978372067/6019351868299726171/6019351868299726171_signature_31.png
| 2024-09-12 | 2024-09-12 14:24 | Matanao Rural Health Unit | Poblacion Matanao | 11 | Davao del Sur | TB-DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29348 | DOH | pamsustento@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 533 | 2024-08-29 | GX-IV 6C 4D | 838456 | Module B2 for repair For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replaced the Defective Module SN:806397 with new Module SN: 685453 Run XpertCheck on Modules 1, 3 and 4 | Perform Preventive Maintenance, Module Repair and Calibration | Module 4 for Service Required Module 1 and 3 Module 2 skipped newly replaced module Only Three Modules arw operational Data sent to Cepheid for Further Assessment |  | 900-0491rs XpertCheck CE 5 | DR 11909 and DR 12004 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019331831487438906/6019331831487438906_signature_28.png | Pamela Joyce Sustento,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6019331831487438906/6019331831487438906_signature_31.png
| 2024-09-12 | 2024-09-12 11:26 | Cebu TB Reference Lab | South Road Property, Cebu | 7 | Cebu | Laboratoy | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29347 | DOH | asmantiquilla.pbsp@gmail.com | 26 |  |  | Gx4 | 838263 | Due date for xpertcheck calibration | Clean Fan Filter Xpertcheck all modules IQR, System log, GXC file | Clean Fan Filter Xpertcheck all modules IQR, System log, GXC file | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019233927675390381/6019233927675390381_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/6019233927675390381/6019233927675390381_signature_31.png
| 2024-09-12 | 2024-09-12 11:23 | Nazareth Health Center | Poblacion Cagayan City | 10 | Cagayan | TB Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29346 | DOH | nanettebtolentino@gmail.com | 23 |  | 2024-09-12 | GX4 | 838269 | . Module B1 serial number 808130 defective . Due for Xpertcheck | . Save Installation Qualification Report . System Log . Save GXC file . Save Assay stat | .Replace module B1 with new module. . Perform Xpertcheck for module B2, B3, B4 only. | Module B3,B4 fail for calibration Module B2 passed for calibration Module B1 new module functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019233184324360022/6019233184324360022_signature_28.png | Nanette Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/6019233184324360022/6019233184324360022_signature_31.png
| 2024-09-12 | 2024-09-12 10:06 | Likhaan Center for Women's Health | 27 Ofelia St., Ofelia Village, Brgy. Project 8, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29345 | MACARE-HBDC | finance@likhaan.org | 25 |  |  | GXIV - II - 10C L | 110025174 | Unit for installation | Setup and installed GX and accessories Performed PM Generated report | Unbox GX and accessories Setup GX and accessories Setup user admin | U it is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019197044515472087/6019197044515472087_signature_28.png | Ronith Susi | https://www.jotform.com/uploads/macarewecare/222927978372067/6019197044515472087/6019197044515472087_signature_31.png
| 2024-09-12 | 2024-09-12 09:37 | Baggao Rural Health Unit | San Jose, Baggao, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29344 | DOH | eslabrajan02@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 726 | 2024-09-10 | GXIV-6C | 838531 | Module 3 errors | Performed Module Replacement B3(862506 - 210091171) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Module | 12058 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019145494825087969/6019145494825087969_signature_28.png | Jan Anthony Eslabra | https://www.jotform.com/uploads/macarewecare/222927978372067/6019145494825087969/6019145494825087969_signature_31.png
| 2024-09-12 | 2024-09-12 08:44 | Laguindingan Rural Health Unit | Poblacion Laguindingan | 6 | Misamis Oriental | TB Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29343 | DOH | chellesambaan@gmail.com | 23 |  | 2024-09-12 | GX4 | 838575 | . Module B1 stuckup cartridge . Due for Xpertcheck . Module B3 defective serial number 863072 | . Save Installation Qualifications Report before and after. . Save Assay stat . Save GXC file . Save system Log | . Pullout cartridge inside the module B1 . Perform Xpertch for module B2 & B3 . Replace Module B3 with new module | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6019144884323148728/6019144884323148728_signature_28.png | CLINT RHEEVEN MAGRAMO | https://www.jotform.com/uploads/macarewecare/222927978372067/6019144884323148728/6019144884323148728_signature_31.png
| 2024-09-11 | 2024-09-11 21:27 | CONG. SIMEON G. TORIBIO MEMORIAL HOSPITAL | CARMEN BOHOL | 7 | Bohol | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29342 | DOH | himaya.vale@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 769 | 2024-09-03 | GX-IV | 838394 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK, CE-IVD 5 Cartridges/kit | 11967 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6018707696211130590/6018707696211130590_signature_28.png | HIMAYA VALE,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6018707696211130590/6018707696211130590_signature_31.png
| 2024-09-11 | 2024-09-11 21:13 | GARCIA MEMORIAL PROVINCIAL HOSPITAL | TALIBON, BOHOL | 7 | Bohol | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29341 | DOH | ybanezrosalina19@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1404 | 2024-09-03 | GX-IV | 838203 | #NAME? | #NAME? | #NAME? | #NAME? |  | XPERTCHECK, CE-IVD 5 Cartridges/kit | 11968 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6018701546216237422/6018701546216237422_signature_28.png | ROSALINA A. YBANEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/6018701546216237422/6018701546216237422_signature_31.png
| 2024-09-11 | 2024-09-11 16:02 | Tomas Claudio Health Center | Corner E. Quirino, Camus Street Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29340 | DOH | arleen_p_alconera@yahoo.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 516 | 2024-08-29 | GX-IV 6C 4D | 838492 | Due for Preventive Maintenance Due for Annual Calibration Module 2 encounter temperature related error | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replaced the Defective Module SN: 791145 with new Module SN:684209 Upgrade the XpertCheck to newer version 1.5B Run XpertCheck Calibration on Three Modules | Perform Preventive Maintenance, Module Repair, and XpertCheck Calibration | Module B1 and B4 Passed the Calibration Module 2 Skipped since newly replaced Module 3 Retry and For Service Required Module 3 for under observation | Errors will be sent to Cepheid for further Assessment | 900-0491rs XpertCheck CE - 5 | DR 11906 AND DR 11999 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6018527035222119631/6018527035222119631_signature_28.png | Arlene P. Alconera,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6018527035222119631/6018527035222119631_signature_31.png
| 2024-09-11 | 2024-09-11 14:44 | Lasam District Hospital | IBJ, Lasam, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29339 | DOH | lasamdh.laboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 413 | 2024-08-29 | GXIV-6C | 838530 | UPS Problem Module 2 errors | Performed UPS Troubleshooting Performed Module Replacement B2(862459 - 669254) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | Genexpert Module | 11931 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6018468034824512043/6018468034824512043_signature_28.png | Verbell Manuela | https://www.jotform.com/uploads/macarewecare/222927978372067/6018468034824512043/6018468034824512043_signature_31.png
| 2024-09-11 | 2024-09-11 13:51 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29338 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1854 |  | 6C GXIV-4D | 838561 | Due for monthly PM | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6018433323857062891/6018433323857062891_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/6018433323857062891/6018433323857062891_signature_31.png
| 2024-09-11 | 2024-09-11 12:24 | Alfonso Ponce Enrile Memorial District Hospital | Smart, Gonzaga, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29337 | DOH | apemdh_laboratory@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2224 | 2024-08-29 | GXIV-6C | 838474 | UPS PROBLEM | Troubleshoot UPS | Performed Maintenance Performed Module Replacement B3(861932 - 667514) Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | Genexpert Module | 11930 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6018381528418085464/6018381528418085464_signature_28.png | Regina Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/6018381528418085464/6018381528418085464_signature_31.png
| 2024-09-10 | 2024-09-10 15:31 | Cardinal Santos Medical Center | Wilson st., San juan | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomed@csmc.ph |  | 29336 | COMMERCIAL | ftejada@csmc.ph | 27 |  |  | MCS + | 08M461 | As per end user they encounter intermittent system error. AC pump motor sudden unusual rotation during pump test and simulation test | Checked unit, Checked calibration parameters Checked and cleaned pump rotor Performed and obeserved simulation testing. During simulation testing safety system error encountered. Possible defective AC pump motor. | Performed troubleshooting | Unit unoperational For further repair and problem isolation | AC pump motor assy for possible replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017636187015031206/6017636187015031206_signature_28.png | FidelTejada | https://www.jotform.com/uploads/macarewecare/222927978372067/6017636187015031206/6017636187015031206_signature_31.png
| 2024-09-10 | 2024-09-10 15:39 | Kalinga Provincial Hospital | Bulanao Centro, Tabuk City, Kalinga | CAR | Kalinga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29335 | DOH | leombelgica139@gmail.com | 28 |  | 2024-08-29 | GXIV-6C | 838466 | Module 1 errors | Performed Module Replacement B1(686593 - 693927) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Module | 11855 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017635492375503101/6017635492375503101_signature_28.png | Leo Mark Belgica | https://www.jotform.com/uploads/macarewecare/222927978372067/6017635492375503101/6017635492375503101_signature_31.png
| 2024-09-10 | 2024-09-10 14:57 | Lubang Rural Health Unit | Lubang, Occidental Mindoro | 4B | Occidental Mindoro | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29334 | DOH | lalaine.batalla@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 419 | 2024-08-29 | GX-IV-D-6C | 838228 | For XpertCheck Module B2 for replacement | Replaced module B2 Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore Replaced filter Replaced module B2 Run XpertCheck Generated reports | Unit is operational. All modules operational. Modules B1, B3 and B4 passed. Module B2 newly replaced. |  | Old: B2/807273 New: B2/797006 | DR# 11870, DR# 11793 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017620929922371768/6017620929922371768_signature_28.png | LALAINE V. BATALLA,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6017620929922371768/6017620929922371768_signature_31.png
| 2024-09-10 | 2024-09-10 14:35 | Calabanga Rural Health Unit | San Pablo, Calabanga | 5 | Camarines Sur | Lang | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29333 | DOH | martyarc11@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 717 |  | GX4-6C | 838398 | B1/807565 for module replacement | Replaced Module: B1 - 807565 >> 693641 Cleaned cartridge bays, plunger rod, icore Generate Data | Replaced module B1/807565 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017599338324809799/6017599338324809799_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/6017599338324809799/6017599338324809799_signature_31.png
| 2024-09-10 | 2024-09-10 13:19 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29332 | DOH | dyrosemarie71@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1651 | 2024-09-02 | GX-IV 6C 4D | 838401 | Module 1 and Module 2 encounter multiple temperature related errors Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replaced the Defective Module SN: 861015 with new Module SN: 872183 Replaced the Defective Module SN: 692438 with new Module SN:685132 | Perform Module Replacement Perform Preventive Maintenance | Only Three Modules are operational Module 4 encounter temperature related errors Errors sent to Cepheid for further Assessment |  | 900-0491rs | DR 11876 and DR 11927 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017592617407675667/6017592617407675667_signature_28.png | Rosemarie Dy RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6017592617407675667/6017592617407675667_signature_31.png
| 2024-09-10 | 2024-09-10 12:16 | Tondo Foreshore Super Health Centee | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29331 | MACARE-HBDC | synch28@yahoo.com | 25 |  |  | GX16-10C | 110016343 | Unit for preventive maintenance. For checking of monitor. | Preventive maintenance, troubleshooting and repair. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed troubleshooting and repair of monitor. Replaced HDMI cable. Generated reports. | Unit is operational. Monitor is now operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017524724013937245/6017524724013937245_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/6017524724013937245/6017524724013937245_signature_31.png
| 2024-09-10 | 2024-09-10 12:01 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29330 | DOH | lornsabaigar@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 814 | 2024-09-10 | GXIV | 838409 | Module B4/861165 Defective | #NAME? | - Performed preventive maintenance - Module replacement B4/861165 into 640214 - Performed functionality test - Generate IQR and System log | Unit is operational |  | 1 Genexpert Module | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017508406653760022/6017508406653760022_signature_28.png | lorna C.abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/6017508406653760022/6017508406653760022_signature_31.png
| 2024-09-10 | 2024-09-10 | Alicia Rural Health Unit | Magsaysay, Alicia, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29329 | DOH | labalicia.rhu@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 828 | 2024-08-29 | GXIV-6C | 838521 | Module 3 errors | Performed Module Replacement B3(807773 - 860270) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Module | 11874 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017453036518593724/6017453036518593724_signature_28.png | Jesile Reymunda Valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/6017453036518593724/6017453036518593724_signature_31.png
| 2024-09-10 | 2024-09-05 13:13 | Poro Rural Health Unit | Poro, Camotes Island | 7 | Cebu | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29328 | DOH | avilacharmaine6@gmail.com | 27 |  |  | Gx4 | 838279 | Defective Module 3 Auto adjustment eroor in monitor | New module 3 installed Troubleshoot monitor Clean Fan filter IQR and System log data | New module 3 installed Troubleshoot monitor Clean Fan filter IQR and System log data | Genexpert Systems operational Need CMOS battery for CPU |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017450946743264060/6017450946743264060_signature_28.png | Charmaine A. Apat | https://www.jotform.com/uploads/macarewecare/222927978372067/6017450946743264060/6017450946743264060_signature_31.png
| 2024-09-10 | 2024-09-03 09:55 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29327 | DOH | laureanjo@yahoo.com | 30 |  |  | Gx 4 | 838527 | Module D4 | New Module D4 installed 674714 Clean Fan Filter Gx Data IQR and system log | New Module D4 installed 674714 Clean Fan Filter Gx Data IQR and system log | Systems operational High Ambient temperature up to 32C |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017450346747677049/6017450346747677049_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/6017450346747677049/6017450346747677049_signature_31.png
| 2024-09-10 | 2024-09-10 09:28 | Balaman rural health unit | Poblacion Balamban Cebu | 7 | Cebu | TB Laboratory | NO | Pullout GeneXpert Machine System | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29326 | DOH | climaco.mildred@yahoo.com | 23 |  | 2024-09-10 | GX4 6c | 838564 | Pullout GeneXpert Machine Transfer to Naga City. | Pullout Printer Serial number KMMK20508 . Pullout Keyboard CT: BEXHP0BCP51CYP . Pullout Voltage regulator Model number SVC . Pullout LCD serial number CNK9430322 . Pullout CPU serial number *CPH0520F08544* . UPS serial number AT1000190402ST0391 | . Pullout Printer Serial number KMMK20508 . Pullout Keyboard CT: BEXHP0BCP51CYP . Pullout Voltage regulator Model number SVC . Pullout LCD serial number CNK9430322 . Pullout CPU serial number *CPH0520F08544* . UPS serial number AT1000190402ST0391 | All units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6017445883418198788/6017445883418198788_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6017445883418198788/6017445883418198788_signature_31.png
| 2024-09-09 | 2024-09-03 11:34 | Columbio Rural Health Unit | Poblacion, Columbio, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29325 | DOH | covivido211@gmail.com | 28.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 573 |  | GxIV-6C | 838556 | N/A | Perform Genexpert machine Preventive maintenance | Gx ID sticker installation Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection | 4 functional module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016996836868785137/6016996836868785137_signature_28.png | ROLAN E. COLOMA | https://www.jotform.com/uploads/macarewecare/222927978372067/6016996836868785137/6016996836868785137_signature_31.png
| 2024-09-09 | 2024-09-09 10:22 | La Castellana Rural Health Unit | La Castellana, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29324 | DOH | milouepascual@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX-IV | 838483 | Defective module B3/862087. | Precentive Maintenance. Cleaned optics, filter, module bay areas, plunger rods. | Replaced Module B3 with new module SN 625627. | 4 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016796041527945313/6016796041527945313_signature_28.png | Mira Luna Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/6016796041527945313/6016796041527945313_signature_31.png
| 2024-09-09 | 2024-09-04 12:57 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29323 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  | 2024-09-03 | MCS+ | 08B049 | Rotor roller rubbers for replacement | Replaced rotor roller rubbers | #NAME? | Unit is operational |  | Rotor Roller Rubbers | 11976 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016784783618296414/6016784783618296414_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6016784783618296414/6016784783618296414_signature_31.png
| 2024-09-09 | 2024-09-06 12:58 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29322 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 807112 | Module C4 (692507) booting problem for further troubleshooting | Swapped C4 (692507) with C1 (882201) | #NAME? | Unit is operational Module 692507 under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016784493617122991/6016784493617122991_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6016784493617122991/6016784493617122991_signature_31.png
| 2024-09-09 | 2024-09-06 12:58 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29321 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 807112 | Module C4 (692507) booting problem for further troubleshooting | Swapped C4 (692507) with C1 (882201) | #NAME? | Unit is operational Module 692507 under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016784173619879963/6016784173619879963_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6016784173619879963/6016784173619879963_signature_31.png
| 2024-09-09 | 2024-09-04 13:00 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | REPAIR MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29320 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  | 2024-09-03 | MCS+ | 01J038 | UPS for replacement | Replaced UPS | #NAME? | Unit is operational Defective UPS for pull out |  | Accupower 1000VA UPS SN: AT1K240323ST0015 | 11977 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016783923615114871/6016783923615114871_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6016783923615114871/6016783923615114871_signature_31.png
| 2024-09-09 | 2024-09-09 14:43 | Maddela District Hospital | Diduyon, Maddela, Quirino | 2 | Quirino | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29319 | DOH | mdhlaboratory00@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 316 | 2024-08-29 | GXIV-6C | 838469 | Module 1 Errors | Performed Module Replacement B1(861013 - 655523) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Module | 11901 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016739503522450715/6016739503522450715_signature_28.png | DONNA MAY CAN̈EGA | https://www.jotform.com/uploads/macarewecare/222927978372067/6016739503522450715/6016739503522450715_signature_31.png
| 2024-09-09 | 2024-09-09 13:53 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29318 | DOH | rhwclab@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 295 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1306 ASSAY: Xpert HBV VL, COUNT: 9 | 2024-08-29 | GX-IV 6C 4D | 838262 | Module 2 and Module 4 for Repair For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Replaced the Defective Module SN: 807733 with new Module SN:210042937 Replaced the Defective Module SN 807952 with new Module SN : 210043526 | Perform Module Replacement and Preventive Maintenance | All Modules are operational |  | 900-0491RS | DR 11878 and DR 11879 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016710377402783053/6016710377402783053_signature_28.png | Kent Ferolin | https://www.jotform.com/uploads/macarewecare/222927978372067/6016710377402783053/6016710377402783053_signature_31.png
| 2024-09-09 | 2024-09-09 14:00 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29317 | MACARE-HBDC | rhwclab@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 295 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1306 ASSAY: Xpert HBV VL, COUNT: 9 |  | GX-IV 6C 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Three Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016709697407439973/6016709697407439973_signature_28.png | Kent G. Ferolin | https://www.jotform.com/uploads/macarewecare/222927978372067/6016709697407439973/6016709697407439973_signature_31.png
| 2024-09-09 | 2024-09-03 21:04 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | La Union | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29316 | DOH | piaramuj@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 832 | 2024-08-29 | GXIV-4-D-6C | 838270 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B4 761595 >>> 675665 - Generated Reports | Performed Module Replacement | Unit is partially operational | Work with Engineer Venzon | GeneXpert Module 675665 | DR 11898 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016708631858509517/6016708631858509517_signature_28.png | Pia Jumar T. Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/6016708631858509517/6016708631858509517_signature_31.png
| 2024-09-09 | 2024-09-03 21:09 | Sinait District Hospital | Sinait, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29315 | DOH | isdhsinaitlaboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 691 | 2024-08-29 | GXIV-4-D-6C | 838277 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B4 210045938 >>>644959 - Generated Reports | Performed Module Replacement | Unit is Operational | Work with Engr. Venzon | GeneXpert Module 644959 | DR 11897 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016707481855989627/6016707481855989627_signature_28.png | Charmi faith p. Ines | https://www.jotform.com/uploads/macarewecare/222927978372067/6016707481855989627/6016707481855989627_signature_31.png
| 2024-09-09 | 2024-09-03 20:52 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29314 | DOH | dianeruiz147@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 445 | 2024-08-29 | GXIV-4-D-6C | 838438 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B3 806441 >>>624493 - Generated Reports | Replaced Defective Module | Unit is Fully Operational |  | GeneXpert Module 624493 | DR 11888 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016706511852308997/6016706511852308997_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/6016706511852308997/6016706511852308997_signature_31.png
| 2024-09-09 | 2024-09-06 09:12 | Rosario Rural Health Unit | Scollt Torillo St., Rosario, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29313 | DOH | arlenepenaarmijo@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1129 | 2024-09-02 | GXIV-D-6C | 838537 | Modules 2 and 3 with Problem/Disabled. | Module Replacement. | Cleaned Cartridge Bay, Plunger Rod and Replaced Filter. Replaced Modules 2 and 3. Performed Functionality Test. Performed Self-test. Generate Reports. | All Modules are Operational. | Def: 683305&732613 , New: 622319&782590 | GX Module Kit 6-C | 11920 & 11914 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016703413619986059/6016703413619986059_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/6016703413619986059/6016703413619986059_signature_31.png
| 2024-09-09 | 2024-09-03 20:58 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29312 | DOH | rhupozorrubio@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 738 | 2024-08-29 | GXIV-4-D-6C | 838467 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B1 861502 >>> 638011 - Performed Troubleshooting on newly replaced Module - Generated Reports | Performed Module Replacement | #NAME? | Work with Engr. Venzon | GeneXpert Module 638011 | DR 11891 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016703361855117571/6016703361855117571_signature_28.png | Jinky Grace Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/6016703361855117571/6016703361855117571_signature_31.png
| 2024-09-09 | 2024-09-03 20:43 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29311 | DOH | jeffreyvrumbaoa@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1369 | 2024-08-29 | GXIV-4-D-6C | 838473 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B2 862041 >>> 615896 - Generated Reports | Performed Module Replacement | Unit is Fully Operational | Work with Engr. Venzon | GeneXpert Module 862041 | DR 11887 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016700711857330196/6016700711857330196_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/6016700711857330196/6016700711857330196_signature_31.png
| 2024-09-09 | 2024-09-03 20:48 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29310 | DOH | allan.vinluan2013@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 985 | 2024-08-29 | GXIV-4-D-6C | 838472 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B1 862034 >>> 652028 - Generated Reports | Replaced Defective Module | Unit is Fully Operational | Work with Engr. Venzon | GeneXpert Module 652028 | DR 11896 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016699731857885149/6016699731857885149_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/6016699731857885149/6016699731857885149_signature_31.png
| 2024-09-09 | 2024-09-03 08:32 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29309 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 ASSAY: Others, COUNT: 8 |  | GXII-2D | 820303 | -Frequent Error 5007 -For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is partially operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016698021852603290/6016698021852603290_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6016698021852603290/6016698021852603290_signature_31.png
| 2024-09-09 | 2024-09-09 11:37 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29308 | DOH | dalepenalosa189@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1176 | 2024-09-09 | GXIV | 838421 | Module B2/617811 Defective | #NAME? | - Module replacement B2/617811 into 794426 - Performed preventive maintenance - Performed functionality test | Unit is operational |  | 1 Genexpert Module | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016625420616651062/6016625420616651062_signature_28.png | Dale D. Peñalosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6016625420616651062/6016625420616651062_signature_31.png
| 2024-09-09 | 2024-09-09 09:10 | Bacolod City Health Office | Bacolod City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29307 | DOH | danicaescorpion@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 7171 | 2024-08-29 | GX-4 | 837612 | Defective module B3/806656. | Precentive Maintenance - clean optics, filter, module bay area, pkunger. | Replaced Module B3 with new module SN 210074294. | 4 modules operational |  | Genexpert module 6c, PN 900-0491RS. | DR 11883 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016554155984176231/6016554155984176231_signature_28.png | Danica Escorpion | https://www.jotform.com/uploads/macarewecare/222927978372067/6016554155984176231/6016554155984176231_signature_31.png
| 2024-09-09 | 2024-09-06 14:55 | Rizal Medical Center | Pasig Blvd. Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | princeszvillaran13@gmail.com |  | 29306 | COMMERCIAL | rmcbloodbank@gmail.com | 24 |  |  | MCS+ | 15H045G | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016522423614687658/6016522423614687658_signature_28.png | Carla Nikka L. Sisno | https://www.jotform.com/uploads/macarewecare/222927978372067/6016522423614687658/6016522423614687658_signature_31.png
| 2024-09-09 | 2024-09-06 14:15 | Rizal Medical Center | Pasig Blvd. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | princeszvillaran13@gmail.com |  | 29305 | MACARE-HBDC | dynrhernandez@gmail.com | 24 | ASSAY: Others, COUNT: 2 ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 167 ASSAY: Xpert SARS-CoV-2, COUNT: 1446 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-XVI-D-6C | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules A1/860941, A2/860940, B3/702534, B4/860923, C2/668748, C3/662448, D1/860832, D3/860742 not operational, for replacement. Module D4/661946 has error 2014, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016522423616760969/6016522423616760969_signature_28.png | Dayanara C Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/6016522423616760969/6016522423616760969_signature_31.png
| 2024-09-08 | 2024-09-05 14:25 | Kidapawan City Health Office | Kidapawan City | 12 | Cotabato | Tb-lab | NO | REPAIR | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29304 | DOH | malufranz5@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2086 |  | Gx4-6C | 838545 | CPU Power Supply Problem | Perform Genexpert CPU repair | Repair/replace Genexpert Cpu power supply | 4 functional module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6016140816864632595/6016140816864632595_signature_28.png | Ma. Luisa Francisco | https://www.jotform.com/uploads/macarewecare/222927978372067/6016140816864632595/6016140816864632595_signature_31.png
| 2024-09-07 | 2024-09-05 13:55 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Tbdots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29303 | MACARE-HBDC | roseannparnes@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2684 |  | GX4 10C | 110014085 | B2 disabled due to 5007 error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and troubleshooting | B2 disabled B1,B3,B4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6014696366521679342/6014696366521679342_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/6014696366521679342/6014696366521679342_signature_31.png
| 2024-09-07 | 2024-09-05 15:08 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29302 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 08H188 | As per end user they encounter bowl breakage during process | Cleaned unit Cleaned centrifuge Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Troubleshooting and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6014695906521338225/6014695906521338225_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/6014695906521338225/6014695906521338225_signature_31.png
| 2024-09-06 | 2024-09-06 15:59 | San Lazaro Hospital (SACCL) | Quericada,Manila | NCR | NCR 2nd - Quezon City | Saccl | NO | Relocation | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29301 | MACARE-HBDC | estrellesb@yahoo.com | 26 | ASSAY: Xpert HBV VL, COUNT: 273 ASSAY: Xpert HCV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 65 |  | GX4 L | 821257 | For relocation | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed relocation | Operational | Work with Gian Paolo |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6014239289745620369/6014239289745620369_signature_28.png | Blessie estrelles | https://www.jotform.com/uploads/macarewecare/222927978372067/6014239289745620369/6014239289745620369_signature_31.png
| 2024-09-06 | 2024-09-06 15:32 | Tagum City Health Office | Mabini Street, Tagum City | 11 | Davao del Norte | TB DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29300 | DOH | tagumcholab.mabini@gmail.com | 23.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2384 | 2024-09-02 | GX-IV 6C 4D | 838553 | Module B1, B3 and B4 for repair For Preventive Maintenance UPS Troubleshooting and Gx sticker installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log Replaced the Defective Module SN 694881 with new Module 659601 Replaced the Defective Module SN 618174 with new Module SN 696367 Replaced the Defective Module SN 612606 with new Module SN 651048 Reset network connection and gateway board Reset UPS connections and test line battery voltages | Perform Preventive Maintenance, Module Repair, Module Troubleshooting, UPS Troubleshooting and GX Sticker Installation | New Module B1 encounter hardware failure, not detected by the GeneXpert Software even it is swap to other module bay. Documented the issue to be sent at Cepheid for further assessment Module B3 and B4 are successfully replaced. Only 3 Modules are operational UPS still not having response of any charging or led status changes, faulty UPS with SN:AT1000190402ST0483 |  | Item code 900-0491rs GX-IV 6C module | DR 11932 DR 11993 DR 11856 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6014187076912641341/6014187076912641341_signature_28.png | Cristina Tiara Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/6014187076912641341/6014187076912641341_signature_31.png
| 2024-09-06 | 2024-09-05 15:51 | San Jose Rural Health Unit | San Jose, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29299 | DOH | jhuneric2008@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2311 |  | GX4-6C | 838389 | Module B2/607968, B4/718686 for module replacement | Replaced module: B2 - 607968 >> 732943 B4 - 718686 >> 909230 Cleaned cartridge bays, plunger rod, gx fan Generate Data | Replaced module B2/607968, B4/718686 Performed Preventive Maintenance | Uni is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6014174650614612509/6014174650614612509_signature_28.png | Rosmar V. Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/6014174650614612509/6014174650614612509_signature_31.png
| 2024-09-06 | 2024-09-06 12:42 | Midsalip Rural health unit | Poblacion Midsalip, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | REPAIR MAINTENANCE Installation of GX ID sticker | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 29298 | DOH | genaliner@rocketmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 384 | 2024-09-06 | GX 4(6c) | 837664 | Keyboard and mouse not working properly | Replaced keyboard and mouse batteries, Preventive maintenence, Installation of GX ID sticker | Replaced keyboard and mouse batteries, Preventive maintenence, Installation of GX ID sticker. | All modules are available, mouse and keyboard are working properly. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6014073864213080809/6014073864213080809_signature_28.png | Genaline Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/6014073864213080809/6014073864213080809_signature_31.png
| 2024-09-06 | 2024-09-06 12:02 | Pasig Treatment Hub | DR. Garcia St. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | dansumulong.pasigcho@gmail.com |  | 29297 | MACARE-HBDC | pathmedtech@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 791 |  | GX-IV-L-6C | 821648 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules A1, A2 and A4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6014048107274725183/6014048107274725183_signature_28.png | DAN SUMULONG | https://www.jotform.com/uploads/macarewecare/222927978372067/6014048107274725183/6014048107274725183_signature_31.png
| 2024-09-05 | 2024-09-03 14:42 | The Medical City | Ortigas ave | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29296 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 44 ASSAY: Xpert CT/NG, COUNT: 71 ASSAY: Others, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 589 ASSAY: Xpert HPV, COUNT: 445 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5434 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 87 ASSAY: Xpert HIV-1 Viral Load, COUNT: 700 ASSAY: Xpert SARS-CoV-2, COUNT: 1105 | 2024-08-30 | GX-XVI-D-10C | 110015767 | Module B4 for replacement | Replaced module B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B4 Generated reports | Unit is operational. All modules operational. |  | Old: B4/210082262 New: B4/872864 | DR# 11913 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013280323292151965/6013280323292151965_signature_28.png | Ashley espiel | https://www.jotform.com/uploads/macarewecare/222927978372067/6013280323292151965/6013280323292151965_signature_31.png
| 2024-09-05 | 2024-09-05 14:32 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan City | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 29295 | COMMERCIAL | microbiology.csmc@gmail.com | 24 | ASSAY: Xpert HBV VL, COUNT: 140 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1130 | 2024-09-04 | GX-IV-D-10C | 110003714 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. Module B3 not operational, disabled. For possible replacement. | Module B3/210066495 service required. Modules B1 and B4 newly replaced. | XpertCheck kit | SI# 098703 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013280323294993912/6013280323294993912_signature_28.png | Ellaine Mare Gatchalia | https://www.jotform.com/uploads/macarewecare/222927978372067/6013280323294993912/6013280323294993912_signature_31.png
| 2024-09-05 | 2024-09-03 13:46 | The Medical City | Ortigas Avenue, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29294 | COMMERCIAL | bloodbank@themedicalcity.com | 20 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013280323292680936/6013280323292680936_signature_28.png | Maricka Viloria | https://www.jotform.com/uploads/macarewecare/222927978372067/6013280323292680936/6013280323292680936_signature_31.png
| 2024-09-05 | 2024-09-03 15:37 | The Medical City | Ortigas ave, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29293 | COMMERCIAL | mscaina@themedicalcity.com | 20 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013280323298670101/6013280323298670101_signature_28.png | Maria Zoraida Caina | https://www.jotform.com/uploads/macarewecare/222927978372067/6013280323298670101/6013280323298670101_signature_31.png
| 2024-09-05 | 2024-09-05 14:29 | PONTEVEDRA RURAL HEALTH UNIT | PONTEVEDRA, CAPIZ | 6 | Capiz | LABORATORY | NO | CALIBRATION MAINTENANCE GX ID Sticker Installation | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29292 | DOH | rtabngjr@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 998 | 2024-09-02 | GX-IV | 838296 | #NAME? | COLLECT IQR, SYSTEM LOG, ARCHIVE, AND CALIBRATION DATA | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK-CE-5 | 5423 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013272612716632772/6013272612716632772_signature_28.png | Reynaldo Tabang, Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/6013272612716632772/6013272612716632772_signature_31.png
| 2024-09-05 | 2024-09-05 13:17 | Dinas Rural Health Unit | Poblacion Dinas, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | REPAIR MAINTENANCE Installation of GX ID sticker | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29291 | DOH | bernantereynacres@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 660 | 2024-09-05 | GX 4 (6c) | 838533 | Module 4 h/w failed and monitor discolouration. | Replaced VGA cord, Preventive maintenance, data collection for module 4 h/w failed, GX ID Barcode sticker installation | Replaced VGA cord, Preventive maintenance, data collection for module 4 h/w failed, GX ID Barcode sticker installation. | Modules 1, 2 and 2 are available while module 4 is h/w failed. | Module 4 must be replaced. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013230484119959986/6013230484119959986_signature_28.png | Reyna Cres Bernante-Lazaga | https://www.jotform.com/uploads/macarewecare/222927978372067/6013230484119959986/6013230484119959986_signature_31.png
| 2024-09-05 | 2024-09-05 12:25 | Tigaon Municipal Health Office | Caraycayon, Tigaon, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29290 | DOH | jingimperial24@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 880 |  | GX4-6C | 838558 | B1/862620, B3/862622, B4/695719 request for module replacement (temperature error) Date and time reset CPU not turning | Replaced Module: B1 - 862620 >> 734388 B3 - 862622 >> 733402 B4 - 695719 >> 614067 Replaced coin-cell battery for Date & time in CPU CPU: cleaned ram and placed in other port Cleaned cartridge bays, plunger rod, fan, icore Generate Data | Replaced module B1/862620, B3/862622, B4/695719 Replaced coin-cell battery for date and time in cpu Performed troubleshootingon CPU Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013222016321401157/6013222016321401157_signature_28.png | Maria Jesusa P. Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/6013222016321401157/6013222016321401157_signature_31.png
| 2024-09-05 | 2024-09-05 12:37 | Garcia Memorial Provincial Hospital | Talibon, Bohol | 7 | Bohol | TB Dots | NO | MAINTENANCE ["[\"GX ID Sticker Installation \",\"GX ID Sticker Installation \"]","[\"GX ID Sticker Installation \",\"GX ID Sticker Installation \"]"] | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29289 | DOH | ybanezrosalina19@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1404 |  | 6C GXIV-4D | 838203 | None | None | -Preventive Maintenance -GeneXpert ID Sticker Installation -Upgrade XpertCheck software version to 1.5b | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013220206159852402/6013220206159852402_signature_28.png | Rosalina A. Ybañez | https://www.jotform.com/uploads/macarewecare/222927978372067/6013220206159852402/6013220206159852402_signature_31.png
| 2024-09-05 | 2024-09-05 12:29 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29288 | DOH | Rhucalbig@gmail.com | 29 |  | 2024-09-05 | GXIV | 838409 | n/a | n/a | #NAME? | Unit is operational (Fan Filter for replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013199741141595277/6013199741141595277_signature_28.png | vanessa mabag | https://www.jotform.com/uploads/macarewecare/222927978372067/6013199741141595277/6013199741141595277_signature_31.png
| 2024-09-05 | 2024-09-05 10:16 | Balamban rural health unit | Poblacion | 7 | Cebu | Tb Laboratory | NO | INSTALLATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29287 | DOH | climaco.mildred@yahoo.comm | 23 |  | 2024-09-05 | GX4 | 838231 | Defective module B3 serial number 807124, B4 serial number 806389 | . Save Installation Qualification. . Save System Log . Save Assay Stat | Replace module B3,B4 with new serial number 688963,685825 | All modules are finctional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013133312512513571/6013133312512513571_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/6013133312512513571/6013133312512513571_signature_31.png
| 2024-09-05 | 2024-09-05 09:34 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29286 | DOH | calambatbdots@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1923 | 2024-09-02 | GXIV | 838225 | Unit for preventive maintenance. For module replacement of B2 and B3. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B2 (692810) and B3 (623515) in replacement to defective B2 (750024) and B3 (687052). | DR# 11928 & 11924 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6013125417714703451/6013125417714703451_signature_28.png | Lily P.Banalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6013125417714703451/6013125417714703451_signature_31.png
| 2024-09-04 | 2024-09-04 16:17 | ALTAVAS RURAL HEALTH UNIT | ALTAVAS, AKLAN | 6 | Aklan | LABORATORY | NO | CALIBRATION MAINTENANCE GX ID Sticker Installation | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29285 | DOH | nycolebenedicto@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 602 | 2024-09-02 | GX-IV | 838458 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | -XPERTCHECK, CE-IVD 5 CARTRIDGES/kit -GX Fan Filter | 5422 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6012476842713380482/6012476842713380482_signature_28.png | MA. NYCOLE S. BENEDICTO, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6012476842713380482/6012476842713380482_signature_31.png
| 2024-09-04 | 2024-09-04 16:10 | LIBACAO RURAL HEALTH UNIT | LIBACAO, AKLAN | 6 | Aklan | LABORATORY | NO | CALIBRATION MAINTENANCE GX ID Sticker Installation | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29284 | DOH | haladeeb18@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 362 | 2024-09-02 | GX-IV | 838540 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | -XPERTCHECK, CE-IVD 5 CARTRIDGES/kit -GX Fan Filter | 5421 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6012470182714609476/6012470182714609476_signature_28.png | Ma. Hala Deed Estanislao | https://www.jotform.com/uploads/macarewecare/222927978372067/6012470182714609476/6012470182714609476_signature_31.png
| 2024-09-04 | 2024-09-04 15:27 | Talisay District Hospital | Talisay City | 7 | Cebu | Tb laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29283 | DOH | mkrystelshean@gmail.com | 23 |  | 2024-09-04 | Gx4 | 838559 | Module B4 Hardware fail | Module B4 passed for Xpertchec. .Save Installation Qualification Reports . Save System Log . Save Assay stat | Perform Xpertcheck for module B4 | . Module B1 not ok to use . Module B2, B3, B4 ok to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6012446794922659620/6012446794922659620_signature_28.png | Krystel Shean Miranda | https://www.jotform.com/uploads/macarewecare/222927978372067/6012446794922659620/6012446794922659620_signature_31.png
| 2024-09-04 | 2024-09-04 14:43 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 29282 | MACARE-HBDC | laboratory@slec.ph | 22 |  |  | Infinity 48s | 809643 | Unit turned off during operation | None | Performed basic troubleshooting of GeneXpert Checked UPS functionality | Unit is operational using newer UPS from customer. Original UPS has functionality issue with constant alarm, probably due to age of UPS. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6012418841932386311/6012418841932386311_signature_28.png | MA. SARAH R. SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/6012418841932386311/6012418841932386311_signature_31.png
| 2024-09-04 | 2024-09-04 13:32 | Manila Doctors Hospital | 667 UN Ave. Ermita, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29281 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 19 | ASSAY: Xpert CT/NG, COUNT: 86 ASSAY: Xpert HBV VL, COUNT: 91 ASSAY: Xpert HPV, COUNT: 28 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1297 ASSAY: Xpert HCV VL, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 |  | GXIV 10C D | 110001648 | End user asked assistance for printer replacement | Performed PM Replaced printer | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced printer HP deskjet 2337 -> HP 2135 | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6012409422316677442/6012409422316677442_signature_28.png | Dyan valois | https://www.jotform.com/uploads/macarewecare/222927978372067/6012409422316677442/6012409422316677442_signature_31.png
| 2024-09-04 | 2024-09-04 13:21 | Bureau of Plant Industry | PhilExport Compound, Pasay City | NCR | NCR 4th - Pasay | BPI - One Stop Export Documentation Center | YES | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29280 | COMMERCIAL | jamoragennielyn14@gmail.com | 25 |  |  | Stemi 305 | 3943037091 | N/A | N/A | Installed Stemi 305 Installed 2x attachment | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6012384941618173166/6012384941618173166_signature_28.png | GENNIELYN R JAMORA | https://www.jotform.com/uploads/macarewecare/222927978372067/6012384941618173166/6012384941618173166_signature_31.png
| 2024-09-04 | 2024-09-04 13:50 | Macare In House | 1005 Aurora Grand,Aurora blvd QC | NCR | NCR 2nd - Quezon City | In House | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 29279 | MACARE-HBDC | florescanoel77@gmail.com | 27 |  | 2024-09-04 | MCS + | 02A087 | CRC bios failed | Replaced and installed CPU board Performed calibration Performed functionality test | Performed repair | Operational |  | CPU board s/n MM021214853 | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6012384353613949185/6012384353613949185_signature_28.png | Macare In house | https://www.jotform.com/uploads/macarewecare/222927978372067/6012384353613949185/6012384353613949185_signature_31.png
| 2024-09-04 | 2024-09-04 13:25 | Alicia Rural Health Unit | Poblacion Alicia | 9 | Zamboanga Sibugay | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE Installation of GX ID barcode sticker | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29278 | DOH | artubag33@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 809 | 2023-08-01 | GX 4 (6c) | 838276 | CPU unable to turn on properly. | Replaced CMOS battery of CPU, Preventive maintenence, Replaced batteries of keyboard and mouse and undergone Xpertcheck calibration. Installation of GX ID sticker | Replaced CMOS battery of CPU, Preventive maintenence, Replaced batteries of keyboard and mouse and undergone Xpertcheck calibration. Installation of GX ID sticker | Genexpert filter must be replaced. Module 1 and 3 passed while module 2 service required and Module 4 suggested for retry. | Genexpert filter must be replaced. | XPERTCHECK-CE-5 | 10393 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6012375990718564335/6012375990718564335_signature_28.png | Ariel T. Baclig | https://www.jotform.com/uploads/macarewecare/222927978372067/6012375990718564335/6012375990718564335_signature_31.png
| 2024-09-04 | 2024-09-04 10:27 | Catigbian District Hospital | Catigbian, Bohol | 7 | Bohol | TB Dots | NO | REPAIR MAINTENANCE GX ID Sticker Installation | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29277 | DOH | rheranes@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 642 | 2024-08-29 | 6C GXIV-4D | 838494 | Module B4 defective | -Replace defective module B4/862188 with module SN: 637284 | #NAME? | All modules are functional. |  |  | DR no. 11890 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6012262268916330834/6012262268916330834_signature_28.png | Rheazel G. Ranes | https://www.jotform.com/uploads/macarewecare/222927978372067/6012262268916330834/6012262268916330834_signature_31.png
| 2024-09-03 | 2024-09-03 14:31 | Governor Generoso Municipal Health Office | Poblacion, Governor Generoso | 11 | Davao Oriental | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29276 | DOH | josephmagallanes72@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 |  | GX-IV 6C 4D | 838490 | For Preventive Maintenance and GX Sticker Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Install GX Sticker | Perform Preventive Maintenance and GX Sticker Installation | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011621006318329772/6011621006318329772_signature_28.png | Charlene V. Magallanes, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6011621006318329772/6011621006318329772_signature_31.png
| 2024-09-03 | 2024-09-03 15:54 | DR. RAFAEL S. TUMBOKON MEMORIAL HOSPITAL | Kalibo, Aklan | 6 | Aklan | HACT | NO | MAINTENANCE GX STICKER ID | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29275 | COMMERCIAL | gonzagaevarmt@gmail.com | 30 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 547 ASSAY: Xpert HIV-1 Viral Load, COUNT: 419 |  | GX-IV | 110020745 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011597652327557646/6011597652327557646_signature_28.png | Eva Raquel Gonzaga, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6011597652327557646/6011597652327557646_signature_31.png
| 2024-09-03 | 2024-09-03 13:25 | Chinese General Hospital | Blumentritt, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 29274 | CONSORTIUM | oncologycgh@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 11 ASSAY: Xpert HBV VL, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1086 ASSAY: Xpert SARS-CoV-2, COUNT: 378 ASSAY: Xpert HCV VL, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GXIV | 835604 | Unit for preventive maintenance. | Preventive maintenance | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011507161222999032/6011507161222999032_signature_28.png | Karen Ringler RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6011507161222999032/6011507161222999032_signature_31.png
| 2024-09-03 | 2024-09-03 12:53 | Chinese General Hospital | Blumentritt, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Bacteriology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 29273 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 633 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 |  | GXIV-10C | 120002026 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011497071228901552/6011497071228901552_signature_28.png | Maria Odessa Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/6011497071228901552/6011497071228901552_signature_31.png
| 2024-09-03 | 2024-09-03 11:31 | Molave Rural Health unit | Molave Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | REPAIR MAINTENANCE GX ID Barcode Installation | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 29272 | DOH | aldrinacain@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 705 | 2024-09-03 | GX 4(6c) | 838233 | Keyboard unable to work properly | Preventive Maintenence, GX ID Barcode Installation and replacement of keyboard batteries. Data collection. | Preventive Maintenence, GX ID Barcode Installation and replacement of keyboard batteries. Data collection. | All modules are available and working properly | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011439505329113024/6011439505329113024_signature_28.png | Aldrin L.Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/6011439505329113024/6011439505329113024_signature_31.png
| 2024-09-03 | 2024-09-03 10:51 | Cong. Simeon G. Toribio Memorial Hospital | Carmen, Bohol | 7 | Bohol | TB Dots | NO | MAINTENANCE GX ID Sticker Installation | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29271 | DOH | himaya.vale@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 780 |  | 6C GXIV-4D | 838394 | None | None | -Preventive Maintenance -GX ID Sticker Installation -Upgrade XpertCheck Software to version 1.5b | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011414176793643213/6011414176793643213_signature_28.png | Himaya Vale | https://www.jotform.com/uploads/macarewecare/222927978372067/6011414176793643213/6011414176793643213_signature_31.png
| 2024-09-03 | 2024-09-03 09:22 | Macare In House | 1005 Aurora Grand, Aurora blvd QC | NCR | NCR 2nd - Quezon City | In House | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 29270 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  | 2024-08-30 | MCS + | 13C269 | CRC bios failed | Replaced and installtion CPU board and safety board. Performed calibration . Test run unit | Performed repair | Operational |  | CPU board s/n C121304219 Safety board s/n C031363275 | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011359433611213175/6011359433611213175_signature_28.png | Macare Medicals Inc. | https://www.jotform.com/uploads/macarewecare/222927978372067/6011359433611213175/6011359433611213175_signature_31.png
| 2024-09-03 | 2024-08-29 15:13 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29269 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 01J038 | UPS problem | NA | #NAME? | Unit is operational Unit is temporarily plugged to other unit's UPS UPS for replacement (SN: AT1000190123ST0184) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011335433619357540/6011335433619357540_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6011335433619357540/6011335433619357540_signature_31.png
| 2024-09-03 | 2024-08-30 14:35 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29268 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 807112 | Unit for maintenance Module C4 for checkup | NA | #NAME? | Unit is operational Unable to generate IQR due to error on C4 C4 waiting for module allocation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011335203612994555/6011335203612994555_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6011335203612994555/6011335203612994555_signature_31.png
| 2024-09-03 | 2024-08-30 14:39 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29267 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 |  |  | GXIV | 807956 | Unit for maintenance A1 for checkup | NA | #NAME? | Unit is operational Waiting for module allocation for A1 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011334993612000518/6011334993612000518_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/6011334993612000518/6011334993612000518_signature_31.png
| 2024-09-03 | 2024-08-29 15:10 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29266 | MACARE-HBDC | marnelli.manuel@gmail.com | 28 |  |  | MCS+ | 08B049 | Blood and AC pump error | Rotor roller pumps for AC, WBC, RBC for replacement | #NAME? | Unit is temporarily not operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6011306793612267569/6011306793612267569_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6011306793612267569/6011306793612267569_signature_31.png
| 2024-09-02 | 2024-08-29 14:25 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29265 | DOH | crmcmolecularlab@gmail.com | 20.8 | ASSAY: Xpert SARS-CoV-2, COUNT: 10583 |  | GxIV-6C | 837662 | Module 2 and 4 H/W failed | Perform Genexpert machine Preventive maintenance | Gx ID sticker installation Perform Genexpert Cleaning and maintenance Checking of errors/problem Problem/errors reporting Data collection | 2 functional modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010948699497259121/6010948699497259121_signature_28.png | Joseph Romy Pascasio | https://www.jotform.com/uploads/macarewecare/222927978372067/6010948699497259121/6010948699497259121_signature_31.png
| 2024-09-02 | 2024-09-02 15:28 | Minglanilla rhu1 | Poblacion | 7 | Cebu | Tb laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29264 | DOH | brendagonzagaparedes@gmail.com | 23 |  | 2024-09-02 | Gx4 | 838512 | Defective Module B4 Sr. Number 862341 | Replace module | Replace module B4 with new serial number 684155 | B1 not ok to use B2, B3, B4 ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010716886118347376/6010716886118347376_signature_28.png | Brenda G. Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/6010716886118347376/6010716886118347376_signature_31.png
| 2024-09-02 | 2024-08-28 11:09 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29263 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 3 ASSAY: Xpert CT/NG, COUNT: 20 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 49 ASSAY: Xpert HCV VL, COUNT: 1 | 2024-08-21 | GX-IV-D-6C | 839231 | Module A4 for replacement | Replaced module A4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module A4 Generated reports | Unit is operational. All modules operational. |  | Old: A4/210096427 New: A4/699644 | DR# 11833 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010698203616288711/6010698203616288711_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6010698203616288711/6010698203616288711_signature_31.png
| 2024-09-02 | 2024-08-30 13:48 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29262 | COMMERCIAL | microbiology.csmc@gmail.com | 20 | ASSAY: Xpert HBV VL, COUNT: 141 ASSAY: Xpert HPV, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1122 | 2024-08-21 | GX-IV-D-10C | 110003714 | Module B4 for replacement | Replaced module B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B4 Generated reports | Unit is operational. All modules operational. | Module B3/210066495 is operational but has confirmed module replacement. | Old: B4/210113679 New: B4/210014273 | DR# 11832 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010698013618487290/6010698013618487290_signature_28.png | Sheilahmae Carandang | https://www.jotform.com/uploads/macarewecare/222927978372067/6010698013618487290/6010698013618487290_signature_31.png
| 2024-09-02 | 2024-09-02 14:24 | Davao Doctors Hospital | Quirino Avenue, Davao City | 11 | Davao del Sur | Molecular | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | alvindaniellegarcia@gmail.com |  | 29261 | COMMERCIAL | labmolecular@ddh.com.ph | 24.5 | ASSAY: Xpert HBV VL, COUNT: 10 |  | GX-IV 6C 4D | 839151 | For Preventive Maintenance and GX Sticker Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Install GX Sticker | Perform Preventive Maintenance and GX Sticker Installation | Module 1 and Module 2 Disabled for multiple temperature related errors Error sent to cepheid for further assessment Only Modules 3 and 4 are operational | Modules 1 and 2 are disabled and not operational |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010683273348389186/6010683273348389186_signature_28.png | Kien Larry T. Ylagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6010683273348389186/6010683273348389186_signature_31.png
| 2024-09-02 | 2024-09-02 14:26 | Davao Doctors Hospital | Quirino Avenue, Davao City | 11 | Davao del Sur | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | alvindaniellegarcia@gmail.com |  | 29260 | COMMERCIAL | labmolecular@ddh.com.ph | 25.4 |  |  | GX-II 6C 2D | 830842 | For data gathering and GX Sticker Installation 9 | Extract needed reports: Installation Qualification Report , Assay Statistics Report, anf System Log Install GX Sticker | Extract data needed | All modules encounter multiple temperature related errors Error sent to cepheid for further assessment | Modules are disabled and not operational |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010682573347963579/6010682573347963579_signature_28.png | Kien Larry T. Ylagan | https://www.jotform.com/uploads/macarewecare/222927978372067/6010682573347963579/6010682573347963579_signature_31.png
| 2024-09-02 | 2024-09-02 13:35 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29259 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2452 ASSAY: Xpert HIV-1 Viral Load, COUNT: 402 | 2024-08-29 | 6C GXIV-4D | 838285 | Module A1 defective | -Replace defective module A1/626532 with module SN: 686648 | #NAME? | All modules are functional. Module A4 pending for replacement. |  | 900-0491RS GeneXpert Module SN. 686648 | 11853 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010649199108991308/6010649199108991308_signature_28.png | Marc Rodney Manaban | https://www.jotform.com/uploads/macarewecare/222927978372067/6010649199108991308/6010649199108991308_signature_31.png
| 2024-09-02 | 2024-08-28 11:09 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29258 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 3 ASSAY: Xpert CT/NG, COUNT: 20 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 49 ASSAY: Xpert HCV VL, COUNT: 1 | 2024-08-21 | GX-IV-D-6C | 839231 | Module A4 for replacement | Replaced module A4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module A4 Generated reports | Unit is operational. All modules operational. |  | Old: A4/210096427 New: A4/699644 | DR# 11833 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010565463619296303/6010565463619296303_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6010565463619296303/6010565463619296303_signature_31.png
| 2024-09-02 | 2024-08-28 11:13 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29257 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 30 ASSAY: Xpert CT/NG, COUNT: 144 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 178 ASSAY: Xpert MTB/RIF Ultra, COUNT: 370 ASSAY: Xpert HCV VL, COUNT: 9 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2024-08-07 | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance Installed barcode scanner | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Installed barcode scanner Generated reports | Unit is operational. All modules operational. |  | Barcode scanner SN: 22355010550961 | DR# 11792 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010565313615600713/6010565313615600713_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6010565313615600713/6010565313615600713_signature_31.png
| 2024-09-02 | 2024-08-30 11:48 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Blood bank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 29256 | COMMERCIAL | pau.robles12@gmail.com | 28 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Run simulation Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010565153612818419/6010565153612818419_signature_28.png | Pauline Robles | https://www.jotform.com/uploads/macarewecare/222927978372067/6010565153612818419/6010565153612818419_signature_31.png
| 2024-09-02 | 2024-08-28 11:09 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29255 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 3 ASSAY: Xpert CT/NG, COUNT: 20 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 49 ASSAY: Xpert HCV VL, COUNT: 1 | 2024-08-21 | GX-IV-D-6C | 839231 | Module A4 for replacement | Replaced module A4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module A4 Generated reports | Unit is operational. All modules operational. |  | Old: A4/210096427 New: A4/699644 | DR# 11833 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010565013619465184/6010565013619465184_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6010565013619465184/6010565013619465184_signature_31.png
| 2024-09-02 | 2024-08-30 13:48 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29254 | COMMERCIAL | microbiology.csmc@gmail.com | 20 | ASSAY: Xpert HBV VL, COUNT: 141 ASSAY: Xpert HPV, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1122 | 2024-08-21 | GX-IV-D-10C | 110003714 | Module B4 for replacement | Replaced module B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B4 Generated reports | Unit is operational. All modules operational. | Module B3/210066495 is operational but has confirmed module replacement. | Old: B4/210113679 New: B4/210014273 | DR# 11832 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564873611675256/6010564873611675256_signature_28.png | Sheilahmae Carandang | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564873611675256/6010564873611675256_signature_31.png
| 2024-09-02 | 2024-08-28 11:13 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29250 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 30 ASSAY: Xpert CT/NG, COUNT: 144 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 178 ASSAY: Xpert MTB/RIF Ultra, COUNT: 370 ASSAY: Xpert HCV VL, COUNT: 9 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2024-08-07 | GX-IV-D-10C | 120002383 | For maintenance | Performed maintenance Installed barcode scanner | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Installed barcode scanner Generated reports | Unit is operational. All modules operational. |  | Barcode scanner SN: 22355010550961 | DR# 11792 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564623612885056/6010564623612885056_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564623612885056/6010564623612885056_signature_31.png
| 2024-09-02 | 2024-08-30 11:48 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Blood bank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 29252 | COMMERCIAL | pau.robles12@gmail.com | 28 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Run simulation Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564623615760453/6010564623615760453_signature_28.png | Pauline Robles | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564623615760453/6010564623615760453_signature_31.png
| 2024-09-02 | 2024-08-28 11:09 | St. Luke's Medical Center - BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29251 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 3 ASSAY: Xpert CT/NG, COUNT: 20 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 49 ASSAY: Xpert HCV VL, COUNT: 1 | 2024-08-21 | GX-IV-D-6C | 839231 | Module A4 for replacement | Replaced module A4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module A4 Generated reports | Unit is operational. All modules operational. |  | Old: A4/210096427 New: A4/699644 | DR# 11833 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564623619460923/6010564623619460923_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564623619460923/6010564623619460923_signature_31.png
| 2024-09-02 | 2024-08-30 13:48 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29253 | COMMERCIAL | microbiology.csmc@gmail.com | 20 | ASSAY: Xpert HBV VL, COUNT: 141 ASSAY: Xpert HPV, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1122 | 2024-08-21 | GX-IV-D-10C | 110003714 | Module B4 for replacement | Replaced module B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B4 Generated reports | Unit is operational. All modules operational. | Module B3/210066495 is operational but has confirmed module replacement. | Old: B4/210113679 New: B4/210014273 | DR# 11832 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564633617027311/6010564633617027311_signature_28.png | Sheilahmae Carandang | https://www.jotform.com/uploads/macarewecare/222927978372067/6010564633617027311/6010564633617027311_signature_31.png
| 2024-09-02 | 2024-08-29 | Riverside Medical Center | Negross Occidental, Bacolod | 6 | Misamis Occidental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29249 | COMMERCIAL | rjmanito@gmail.com | 24 |  |  | Gx 4 | 110005268 | Defective Module B2 and B4 | Xpertcheck Calibration B2 and B4 Clean Fan Filter | Xpertcheck Calibration B2 and B4 Clean Fan Filter | Module 1 passed, module 2 and 4 retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010536713292122271/6010536713292122271_signature_28.png | Rumary Joan Manito | https://www.jotform.com/uploads/macarewecare/222927978372067/6010536713292122271/6010536713292122271_signature_31.png
| 2024-09-02 | 2024-09-02 08:35 | CEBU DOCTORS UNIVERSITY HOSPITAL | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29248 | COMMERCIAL | cduhlabresults2@yahoo.com | 24 |  |  | MCS+ | 17H639SPG | Fuse busted in the supply. | Replacement of fuse 2A and test run for reliability. | Replacement of fuse 2A and test run for reliability. Check and clean bowl chuck | Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010467592118886022/6010467592118886022_signature_28.png | Cebu Doctors Hospital University | https://www.jotform.com/uploads/macarewecare/222927978372067/6010467592118886022/6010467592118886022_signature_31.png
| 2024-09-02 | 2024-08-29 15:13 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29247 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 01J038 | UPS problem | NA | #NAME? | Unit is operational Unit is temporarily plugged to other unit's UPS UPS for replacement (SN: AT1000190123ST0184) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010420997511621803/6010420997511621803_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6010420997511621803/6010420997511621803_signature_31.png
| 2024-09-02 | 2024-08-29 15:10 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29246 | MACARE-HBDC | marnelli.manuel@gmail.com | 28 |  |  | MCS+ | 08B049 | Blood and AC pump error | Rotor roller pumps for AC, WBC, RBC for replacement | #NAME? | Unit is temporarily not operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6010420327512238874/6010420327512238874_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6010420327512238874/6010420327512238874_signature_31.png
| 2024-08-30 | 2024-08-30 16:14 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 29245 | MACARE-HBDC | laboratory@slec.ph | 22 |  |  | GeneXpert IV | 815686 | Unit for XpertCheck | None | Performed maintenance cleaning Performed XpertCheck | Unit is operational, all modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6008161745154163115/6008161745154163115_signature_28.png | Khariz Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/6008161745154163115/6008161745154163115_signature_31.png
| 2024-08-30 | 2024-08-30 16:18 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 29244 | MACARE-HBDC | laboratory@slec.ph | 22 |  |  | GeneXpert IV | 808092 | Unit for XpertCheck | None | Performed maintenance cleaning Performed XpertCheck procedure | Unit ia operational, module B4/657202 requires servicing |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6008160465153459656/6008160465153459656_signature_28.png | Khariz Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/6008160465153459656/6008160465153459656_signature_31.png
| 2024-08-30 | 2024-08-30 14:31 | Don Jose S. Monfort Medical Center | Poblacion Barotac | 6 | Iloilo | Laboratory | YES | INSTALLATION TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | shirleymae@gmail.com |  | 29243 | DOH | judyboglosa35@gmail.com | 23 |  | 2024-08-30 | GX4 | 838501 | 3 modules are functional only due to module 1 fail for XpertCheck last March. | Save File's | . Installation of GX ID sticker . Perform preventive maintenance . Save file System Log . Save file Assay Stat . Save Installation Qualification | . Modules B1 subject for replacement fail for Xpert Check last march 2024 . Module B3 under observation due to error H/W sometimes . Module B2 and B4 ok to use. | Operator of GeneXpert Machine request for training on how to Preventive maintenance and also how to handle error's of cartridges during running. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6008095541661831302/6008095541661831302_signature_28.png | Juny B. Boglosa | https://www.jotform.com/uploads/macarewecare/222927978372067/6008095541661831302/6008095541661831302_signature_31.png
| 2024-08-30 | 2024-08-28 13:46 | Malay Primary Care Facility | Caticlan, Malay, Aklan | 6 | Aklan | Laboratory | NO | ["[\"GeneXpert ID Sticker Installation\",\"GeneXpert ID Sticker Installation\"]","[\"GeneXpert ID Sticker Installation\",\"GeneXpert ID Sticker Installation\"]"] | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29242 | DOH | malaymhoclinicallaboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1042 |  | 6C GXIV-4D | 838457 | Due for XPC | #NAME? | #NAME? | Module B2 for replacement. Other modules passed XPC and functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6008043908555529921/6008043908555529921_signature_28.png | May Soledad | https://www.jotform.com/uploads/macarewecare/222927978372067/6008043908555529921/6008043908555529921_signature_31.png
| 2024-08-30 | 2024-08-30 11:52 | Bureau of Plant Industry | 692 San Andres St., 706 Zone 77, Malate Manila | NCR | NCR 1st - Manila | BPI-CPEQS Plant Pathology Lab | YES | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29241 | COMMERCIAL | banasihan.arnel28@gmail.com | 25 |  | 2024-08-29 | Axiocam 208 Color | 5321007166 | N/A | N/A | Installed Axiocam 208 Color on the Stemi 305 at Plant Pathology Laboratory Set-up router for wireless access of Axiocam Set-up Laptop | Unit is operational | Installed at BPI-CPEQS Los Baños | Axiocam 208 Color Laptop Router | SI 98674 & DR 11844 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6008036856432057626/6008036856432057626_signature_28.png | Arnel C Banasihan | https://www.jotform.com/uploads/macarewecare/222927978372067/6008036856432057626/6008036856432057626_signature_31.png
| 2024-08-30 | 2024-08-30 12:00 | Bureau of Plant Industry | 692 San Andres St., 706 Zone 77, Malate Manila | NCR | NCR 1st - Manila | BPI-CPEQS Entomology Laboratory | YES | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 29240 | COMMERCIAL | banasihan.arnel28@gmail.com | 25 |  | 2024-08-29 | Axiocam 208 Color | 5321007325 | N/A | N/A | Installed Axiocam 208 Color on the Stemi 305 at Entomology Laboratory Set-up router for wireless access of Axiocam Set-up Laptop | Unit is operational | Installed at BPI-CPEQS Los Baños | Axiocam 208 Color | SI 98674 & DR 11844 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6008036436436977478/6008036436436977478_signature_28.png | Arnel C Banasihan | https://www.jotform.com/uploads/macarewecare/222927978372067/6008036436436977478/6008036436436977478_signature_31.png
| 2024-08-30 | 2024-08-30 12:31 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29239 | DOH | ammegzed@gmail.com | 25 |  | 2024-08-30 | GXIV | 838421 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6008016421018541966/6008016421018541966_signature_28.png | gemma malate | https://www.jotform.com/uploads/macarewecare/222927978372067/6008016421018541966/6008016421018541966_signature_31.png
| 2024-08-30 | 2024-08-30 12:04 | Kalilangan Rural Health Unit | West Poblacion, Kalilangan | 10 | Bukidnon | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29238 | DOH | myrnasuyao17@gmail.com | 30 |  |  | GX-4 6 Colors | 838571 | None | None | Preventive Maintenance GX ID Sticker Installation | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6008002690754346426/6008002690754346426_signature_28.png | Myrna Suyao | https://www.jotform.com/uploads/macarewecare/222927978372067/6008002690754346426/6008002690754346426_signature_31.png
| 2024-08-30 | 2024-08-30 11:43 | Vuena Valencia Rural Health Unit | Poblacion | 6 | Guimaras | Laboratory | YES | INSTALLATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29237 | DOH | xxxrhingringxxx@gmail.com | 24 |  | 2024-08-29 | Gx4 | 838291 | N/A | Prerntive Maintenance | . Installation of GX ID Sticker . Perform Preventive maintenance . Save file's Sytem Log, Installation Qualification . Change filter. | All modules are functional. | User of GeneXpert Machine Request for training on how to Preventive Maintenance and Error's encounter of Cartridge's during running. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007991371663046747/6007991371663046747_signature_28.png | Rhea April Q. Saleco | https://www.jotform.com/uploads/macarewecare/222927978372067/6007991371663046747/6007991371663046747_signature_31.png
| 2024-08-30 | 2024-08-29 16:58 | Iriga City Health Office | Sta Cruz, Iriga | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29236 | DOH | jahseh004200@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1930 |  | GX4-6C | 838391 | UPS battery drained/defective GX for Calibration | Replaced UPS batteries(2x) XpertCheck Result: B1/758212 - passed B2/210044677 - service required B3/669045 - passed B4/210040454 - passed Cleaned cartridge bays, plunger rod, icore Generate Data | Replaced UPS batteries Performed XpertCheck | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007963099115483953/6007963099115483953_signature_28.png | Hannah Kirsten O. Novela | https://www.jotform.com/uploads/macarewecare/222927978372067/6007963099115483953/6007963099115483953_signature_31.png
| 2024-08-30 | 2024-08-30 10:46 | Phililpine General Hospital | Taft. Ave. Ermita, Manil | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29235 | MACARE-HBDC | pghmolecularlab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 97 ASSAY: Others, COUNT: SA Nasal Complete G3 (1) ASSAY: Xpert SARS-CoV-2, COUNT: 1598 ASSAY: Xpert SARS-CoV-2, COUNT: 56 |  | GX16 | 844975 | Unit for preventive maintenance. For generating of reports for technical troubleshooting. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated GXX file for technical troubleshooting. Generated reports. | Unit is operational. Generated reports for technical troubleshooting of technical specialist for the reported errors. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007959332948621572/6007959332948621572_signature_28.png | Angela Ricamela Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/6007959332948621572/6007959332948621572_signature_31.png
| 2024-08-29 | 2024-08-29 15:21 | Maco Rural Health Unit | Poblacion, Maco | 11 | Davao de Oro | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29234 | DOH | edlienkai26@gmail.com | 29.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1034 |  | GX-IV 6C 4D | 838500 | For Preventive Maintenance and GX Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Install Sticker | Perform Preventive Maintenance and GX Sticker Installation | Module B4 H/W Failed, reported to Jotform for further assessment | Advised the Enduser to achieve the optimal GeneXpert Temperature for running |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007260779812701040/6007260779812701040_signature_28.png | Edlien Kai Prado | https://www.jotform.com/uploads/macarewecare/222927978372067/6007260779812701040/6007260779812701040_signature_31.png
| 2024-08-29 | 2024-08-29 09:18 | Valencia City Health Office | New Terminal, Hagkol, Valencia City | 10 | Bukidnon | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29233 | DOH | caluyajopay@gmail.com | 27 |  |  | GX-4 6 Colors | 838271 | Module B3 found temperature error(reported) | None | Preventive Maintenance GX ID Sticker Installation | 3 modules functional, 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007249472681199454/6007249472681199454_signature_28.png | Joefelyn Caluya | https://www.jotform.com/uploads/macarewecare/222927978372067/6007249472681199454/6007249472681199454_signature_31.png
| 2024-08-29 | 2024-08-29 14:55 | Quezon Rural Health Unit | Poblacion, Quezon | 10 | Bukidnon | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29232 | DOH | quenxlopez@gmail.com | 30 |  |  | GX-4 6 Colors | 838440 | None | None | Preventive Maintenance GX ID Sticker Installation | All module and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007239612685330695/6007239612685330695_signature_28.png | Quenie Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/6007239612685330695/6007239612685330695_signature_31.png
| 2024-08-29 | 2024-08-29 10:28 | Sultan Mastura Rural Health Unit | Sultan Mastura, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29231 | DOH | salehshahinaz@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 437 |  | GxIV-6C | 837613 | N/A | Perform Genexpert machine Preventive maintenance | Gx ID sticker installation Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007231763612652447/6007231763612652447_signature_28.png | Shahinaz S. Saleh | https://www.jotform.com/uploads/macarewecare/222927978372067/6007231763612652447/6007231763612652447_signature_31.png
| 2024-08-29 | 2024-08-29 10:28 | Sultan Mastura Rural Health Unit | Sultan Mastura, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29230 | DOH | salehshahinaz@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 437 |  | GxIV-6C | 837613 | N/A | Perform Genexpert machine Preventive maintenance | Gx ID sticker installation Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007231763612898876/6007231763612898876_signature_28.png | Shahinaz S. Saleh | https://www.jotform.com/uploads/macarewecare/222927978372067/6007231763612898876/6007231763612898876_signature_31.png
| 2024-08-29 | 2024-08-28 13:06 | Teresita Jalandoni Provincial Hospital | Negros Occidental, Silay | 6 | Negros Occidental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29229 | DOH | mrvillan.pbsp@gmail.com | 29 |  |  | Gx 4 | 838464 | Due date for xpertcheck calibration A1 and A2 | Xpertcheck calibration Module A3 Clean Fan Filter | Xpertcheck calibration Module A3 Clean Fan Filter | Systems operational Xpertcheck A3 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007220604528241828/6007220604528241828_signature_28.png | Ma. Arna R. Villan | https://www.jotform.com/uploads/macarewecare/222927978372067/6007220604528241828/6007220604528241828_signature_31.png
| 2024-08-29 | 2024-08-29 12:13 | Bukidnon Provincial Hospital | CM Recto St., Palma, Kibawe | 10 | Bukidnon | Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29228 | DOH | juliealte10@gmail.com | 30 |  |  | GX-4 6 Colors | 838547 | Module B4 found temperature error (for observation) | None | Preventive Maintenance GX ID Sticker Installation | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007153055113827651/6007153055113827651_signature_28.png | Julie Alte | https://www.jotform.com/uploads/macarewecare/222927978372067/6007153055113827651/6007153055113827651_signature_31.png
| 2024-08-29 | 2024-08-29 10:35 | LUGAIT HEALTH CENTER | Lugait, Misamis Oriental | 10 | Misamis Oriental | LABORATORY | NO | MAINTENANCE GX Sticker ID Installation | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29227 | DOH | lugaitmaternity@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 697 |  | GX-IV | 838562 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007085061449116531/6007085061449116531_signature_28.png | Rodelyn B. Limocon, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6007085061449116531/6007085061449116531_signature_31.png
| 2024-08-29 | 2024-08-29 10:25 | Baybay City Health Office | Baybay Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 29226 | DOH | sajecopooh20@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 937 | 2024-08-29 | GXIV | 838555 | N/A | N/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6007077814539775949/6007077814539775949_signature_28.png | SARAH JEAN GABIETA | https://www.jotform.com/uploads/macarewecare/222927978372067/6007077814539775949/6007077814539775949_signature_31.png
| 2024-08-29 | 2024-08-28 15:48 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29225 | DOH | gmpelaez144@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1483 |  | GxIV-6C | 838528 | Module A1 h/w failed | Perform Genexpert machine Preventive maintenance | Set up/Network 2 Gx machine in 1 computer Gx ID sticker installation Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6006979145718448491/6006979145718448491_signature_28.png | Gladiola Marie E. Pelaez | https://www.jotform.com/uploads/macarewecare/222927978372067/6006979145718448491/6006979145718448491_signature_31.png
| 2024-08-29 | 2024-08-28 15:54 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29224 | UNOPS | onehope834@gmail.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1483 |  | GxIV-10C | 110023524 | N/A | Perform Genexpert machine Preventive | Set up/Network 2 Gx machine in 1 computer Gx ID sticker installation Perform Genexpert Cleaning and maintenance Checking of errors/problem Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6006976195718580268/6006976195718580268_signature_28.png | EMELYN A,CADIZ | https://www.jotform.com/uploads/macarewecare/222927978372067/6006976195718580268/6006976195718580268_signature_31.png
| 2024-08-28 | 2024-08-28 15:30 | Tagoloan Rural Health Unit | Balagtas St, Tagoloan | 13 | Misamis Oriental | TB-DOTS | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29223 | DOH | macalisangglezeal@gmail.com | 29 |  |  | GX-4 6 Colors | 838574 | B3 was reported to Cepheid(Upcoming Shipment) | None | Preventive Maintenance GX ID Sticker Installation | 3 modules working, 1 module for replacement All GX accessories ate functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6006396511122592575/6006396511122592575_signature_28.png | Venz Michael Bacoy | https://www.jotform.com/uploads/macarewecare/222927978372067/6006396511122592575/6006396511122592575_signature_31.png
| 2024-08-28 | 2024-08-28 15:22 | LAGUINDINGAN RURAL HEALTH UNIT | Laguindingan, Misamis Oriental | 10 | Misamis Oriental | Laboratory | NO | MAINTENANCE GX Sticker ID Installation | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29222 | DOH | chellesambaan@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 |  | GX-IV | 838575 | -FOR PREVENTIVE MAINTENANCE -MODULE B3 ERROR 1001 -GX STICKER ID INSTALLATION | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6006394292137570903/6006394292137570903_signature_28.png | Rochelle M. Sambaan , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6006394292137570903/6006394292137570903_signature_31.png
| 2024-08-28 | 2024-08-28 12:42 | Monkayo Rural Health Unit | Poblacion, Monkayo | 11 | Davao de Oro | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29221 | DOH | kendralois730@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1865 |  | GX-IV 6C 4D | 838487 | Stuck Cartidge at Module B2 | Remove the stuck cartidge at Module 2 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install GX Sticker | Perform Stuck Cartidge Removal | Successfully removed the stuck cartidge |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6006324789262121615/6006324789262121615_signature_28.png | Karen Anuta RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6006324789262121615/6006324789262121615_signature_31.png
| 2024-08-28 | 2024-08-28 09:02 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | TB Dots | NO | ["GX ID Sticker Installation","GX ID Sticker Installation"] | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29220 | DOH | sanjoseantiquegxsite@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1124 ASSAY: Xpert HIV-1 Viral Load, COUNT: 87 |  | 6C GXIV-4D | 838412 | Due for XPC | #NAME? | #NAME? | All modules passed XpertCheck, functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6006175888931563146/6006175888931563146_signature_28.png | Oscar Delano Salazar, IV | https://www.jotform.com/uploads/macarewecare/222927978372067/6006175888931563146/6006175888931563146_signature_31.png
| 2024-08-27 | 2024-08-27 15:29 | Macare In House | 1005 Aurora Grand, Aurora blvd QC | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 29219 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | CPU board | MM021214853 | CPU bios battery error | Replaced and installed CPU battery Performed calibration Performed configuration Test run unit ok | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005530553613276084/6005530553613276084_signature_28.png | Macare in House | https://www.jotform.com/uploads/macarewecare/222927978372067/6005530553613276084/6005530553613276084_signature_31.png
| 2024-08-27 | 2024-08-27 15:27 | Macare In House | 1005 Aurora Grand,Aurora blvd. QC | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 29218 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | CPU board | MM061004515 | CPU bios battery error | Replaced and installed battery. Performed calibration Performed configuration Test run unit ok | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005528943618351829/6005528943618351829_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/6005528943618351829/6005528943618351829_signature_31.png
| 2024-08-27 | 2024-08-27 15:24 | Macare In House | 1005 Aurora grand Aurora blvd QC | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 29217 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | CPU board | C101429672G | CPU Bios battery error | Replaced and installed CPU battery Performed calibration Performed configuration Test run unit ok | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005527273612763899/6005527273612763899_signature_28.png | Macare In house | https://www.jotform.com/uploads/macarewecare/222927978372067/6005527273612763899/6005527273612763899_signature_31.png
| 2024-08-27 | 2024-08-27 15:20 | Macare In House | 1005 Aurora Grand Aurora blvd.QC | NCR | NCR 2nd - Quezon City | In House | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 29216 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | CPU board | C101429661G | CPU Bios battery error | Replaced and installed CPU battery. Performed calibration Performed configuration Test run unit ok | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005525533614080700/6005525533614080700_signature_28.png | Macare In house | https://www.jotform.com/uploads/macarewecare/222927978372067/6005525533614080700/6005525533614080700_signature_31.png
| 2024-08-27 | 2024-08-22 09:49 | Abra Provincial Hospital | Bangued, Abra | CAR | Abra | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29215 | UNOPS | aphpathologydepartment@gmail.com | 30 |  |  | GXIV-4-D-10C | 110023499 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005501834941069745/6005501834941069745_signature_28.png | Riverlie A. Valera | https://www.jotform.com/uploads/macarewecare/222927978372067/6005501834941069745/6005501834941069745_signature_31.png
| 2024-08-27 | 2024-08-22 08:10 | Manabo Rural Health Unit | Manabo, Abra | CAR | Abra | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29214 | DOH | lenmayclairelumiwag15@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 315 |  | GXIV-4-D-6C | 838274 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005500784946034080/6005500784946034080_signature_28.png | Ellen May Claire S. Lumiwag | https://www.jotform.com/uploads/macarewecare/222927978372067/6005500784946034080/6005500784946034080_signature_31.png
| 2024-08-27 | 2024-08-21 13:52 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29213 | UNOPS | benguetphortdl@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 113 |  | GXIV-4-D-10C | 110023498 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005499494945273072/6005499494945273072_signature_28.png | Frannie G. Calixto | https://www.jotform.com/uploads/macarewecare/222927978372067/6005499494945273072/6005499494945273072_signature_31.png
| 2024-08-27 | 2024-08-21 11:32 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29212 | DOH | sherryannemarrero@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2503 |  | GXIV-4-D-6C | 838281 | For Preventive Maintenance and Temporary Disassembly of GeneXpert System | #NAME? | Performed Preventive Maintenance and Disassembly of GeneXpert System | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498544945675189/6005498544945675189_signature_28.png | Emilee Cuison | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498544945675189/6005498544945675189_signature_31.png
| 2024-08-27 | 2024-08-13 06:28 | Romblon District Hospital | Hospital Road, Romblon, Romblon | 4B | Romblon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29209 | DOH | rdhlab5500@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 643 | 2024-05-02 | GXIV-D-6C | 838573 | Unit is due for XpertCheck. | Preventive Maintenance. XpertCheck. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Cleaned Filter. Performed Self-test. Performed XpertCheck. Generate Reports. | All Modules Passed. |  | XpertCheck Kit | DR 11457 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498253613784141/6005498253613784141_signature_28.png | SHEILAH MAE R. PUNAY | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498253613784141/6005498253613784141_signature_31.png
| 2024-08-27 | 2024-08-09 12:02 | Dasmariñas City Health Office II | DBB-E, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29210 | DOH | hazelmilan0777@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 102 | 2024-08-07 | GXIV-D-6C | 838535 | UPS Low Backup Voltage. | Repair. | Replaced UPS Battery. Performed Functionality Test. Generate Reports. | GX System Is Good. |  | UPS Battery | DR 11790 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498253614786275/6005498253614786275_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498253614786275/6005498253614786275_signature_31.png
| 2024-08-27 | 2024-08-08 11:45 | Rosario Rural Health Unit | Rosario, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29207 | DOH | arlenepenaarmijo@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1123 |  | GXIV-D-6C | 838537 | Unit is due for Maintenance | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Cartridge if Clogged. Generate Reports. Performed Self-test. | 2 Modules are Dissabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498253614820504/6005498253614820504_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498253614820504/6005498253614820504_signature_31.png
| 2024-08-27 | 2024-08-16 10:58 | General Trias Social Hygiene Clinic | Hospital Road, General Trias, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29208 | MACARE-HBDC | iveerose02@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 650 |  | GXIV-D-10C | 110016785 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Generate Reports. Performed Self-test | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498253618670028/6005498253618670028_signature_28.png | Ivee Rose R. Milay | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498253618670028/6005498253618670028_signature_31.png
| 2024-08-27 | 2024-08-09 12:02 | Dasmariñas City Health Office II | DBB-E, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 29211 | DOH | hazelmilan0777@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 102 | 2024-08-07 | GXIV-D-6C | 838535 | UPS Low Backup Voltage. | Repair. | Replaced UPS Battery. Performed Functionality Test. Generate Reports. | GX System Is Good. |  | UPS Battery | DR 11790 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498263611320090/6005498263611320090_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/6005498263611320090/6005498263611320090_signature_31.png
| 2024-08-27 | 2024-08-21 10:44 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29206 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 21 | ASSAY: Xpert HBV VL, COUNT: 177 ASSAY: Xpert SARS-CoV-2, COUNT: 203 ASSAY: Xpert HCV VL, COUNT: 41 ASSAY: Xpert HIV-1 Viral Load, COUNT: 28 |  | GXIV-4-D-6C | 110016112 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005496254948506956/6005496254948506956_signature_28.png | Helley Mae Lo | https://www.jotform.com/uploads/macarewecare/222927978372067/6005496254948506956/6005496254948506956_signature_31.png
| 2024-08-27 | 2024-07-12 13:20 | Bulacan Medical Center | Malolos, Bulacan | NCR | Bulacan | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29205 | UNOPS | bmclab_micro@yahoo.com | 27 |  |  | GXIV 10C | 110023507 | Unit for rearrangement | Rearranged unit position | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005495373617533709/6005495373617533709_signature_28.png | Maria Angelica P. Guintu | https://www.jotform.com/uploads/macarewecare/222927978372067/6005495373617533709/6005495373617533709_signature_31.png
| 2024-08-27 | 2024-08-21 09:06 | Itogon Rural Health Unit | Itogon, Benguet | CAR | Benguet | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29204 | DOH | ggdepayso@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 336 | 2024-07-19 | GXIV-4-D-6C | 838282 | For Calibration and Preventive Maintenance | #NAME? | Performed Calibration and Preventive Maintenance | - Unit is Operational - Did not perform XpertCheck on newly Replaced Module B3 (863306) - All Remaining Modules passed XpertCheck |  | XpertCheck Kits | DR 11730 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005495214941645459/6005495214941645459_signature_28.png | Gloria G. Depayso | https://www.jotform.com/uploads/macarewecare/222927978372067/6005495214941645459/6005495214941645459_signature_31.png
| 2024-08-27 | 2024-07-31 13:59 | Bagong Barrio Health Center | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29203 | DOH | cruzernitamay@gmail.com | 27 |  |  | GXIV | 831906 | Unit for relocation | #NAME? | Relocated unit | Unit is not operational Pending desktop replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005495133618189933/6005495133618189933_signature_28.png | Ernita Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/6005495133618189933/6005495133618189933_signature_31.png
| 2024-08-27 | 2024-08-27 14:26 | Carmen Health Center | Carmen , Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | MAINTENANCE GX ID Sticker | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29202 | DOH | murielmelchor11@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 644 |  | GX-IV | 838442 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005494944129779909/6005494944129779909_signature_28.png | Muriel G. Melchor | https://www.jotform.com/uploads/macarewecare/222927978372067/6005494944129779909/6005494944129779909_signature_31.png
| 2024-08-27 | 2024-08-27 14:19 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29201 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 11J219 | #NAME? | For replacement of the new pneumatic assembly as soon as possible. The pneumatic assembly is defective. | Troubleshoot, and test Mcs+ valve test. | Unit is not operational due to defective pneumatic assembly. (Service Unit) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005490423345986825/6005490423345986825_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/6005490423345986825/6005490423345986825_signature_31.png
| 2024-08-27 | 2024-08-27 13:12 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29200 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 41 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2826 ASSAY: Xpert HIV-1 Viral Load, COUNT: 431 |  | 6C GXIV-4D | 838285 | Due for XPC | -Run XpertCheck on module A1, A2, & A3 | #NAME? | Module A1 failed XPC, for replacement. Module A2 and A3, passed XPC. Module A4 waiting for allocation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005465224311236122/6005465224311236122_signature_28.png | Leah Karen Noynay, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6005465224311236122/6005465224311236122_signature_31.png
| 2024-08-27 | 2024-08-22 14:52 | The Medical City - South Luzon | Sta. Rosa, Laguna | 4A | Laguna | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29199 | COMMERCIAL | tmcsl-molecularlab@themedicalcity.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 188 ASSAY: Xpert HIV-1 Viral Load, COUNT: 29 ASSAY: Xpert SARS-CoV-2, COUNT: 2 |  | GXIV | 839233 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005462683612662781/6005462683612662781_signature_28.png | Ma Veronica Agnas | https://www.jotform.com/uploads/macarewecare/222927978372067/6005462683612662781/6005462683612662781_signature_31.png
| 2024-08-27 | 2024-08-22 15:53 | Pilar Rural Health Unit | Brgy. Calongay, Pilar | 5 | Sorsogon | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29198 | DOH | manalangmarkramir@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1133 |  | GX4-6C | 838266 | Gx for Preventive Maintenance UPS batteries for replacement | Replaced 2 UPS battery Cleaned cartridge bays, plunger rod Generate Data | Replaced UPS batteries Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005458113229960402/6005458113229960402_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/6005458113229960402/6005458113229960402_signature_31.png
| 2024-08-27 | 2024-08-15 14:21 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29197 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 08M461 | Unit for maintenance | NA | #NAME? | Unit is operational Red and white pump rubber rollers for replacement next PM |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005457483616221618/6005457483616221618_signature_28.png | Marnelli Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/6005457483616221618/6005457483616221618_signature_31.png
| 2024-08-27 | 2024-08-21 13:49 | Kalinga Health Hub Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29196 | CONSORTIUM | pearl.balderama@innovationsch.org | 25 |  |  | GXIV | 110008420 | UPS for troubleshooting (beeping sound) | Plugged GX system to power outlet | #NAME? | Unit is operational UPS is recommended for repair/replacement from supplier |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005457153618947999/6005457153618947999_signature_28.png | Pearl Balderama | https://www.jotform.com/uploads/macarewecare/222927978372067/6005457153618947999/6005457153618947999_signature_31.png
| 2024-08-27 | 2024-08-22 11:31 | Pinagbuhatan Health Center | Urbano Velasco Ave Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29195 | MACARE-HBDC | peachmango318@gmail.cim | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 | 2024-08-21 | GX-IV-D-10C | 110024245 | Module B1 for replacement | Replaced module B1 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filer Replaced module B1 Generated reports | Unit is operational. All modules operational. |  | Old: B1/210127592 New:B1/210021453 | DR# 11834 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005456853613044442/6005456853613044442_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6005456853613044442/6005456853613044442_signature_31.png
| 2024-08-27 | 2024-08-22 11:31 | Pinagbuhatan Health Center | Urbano Velasco Ave Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29194 | MACARE-HBDC | peachmango318@gmail.cim | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 202 | 2024-08-21 | GX-IV-D-10C | 110024245 | Module B1 for replacement | Replaced module B1 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filer Replaced module B1 Generated reports | Unit is operational. All modules operational. |  | Old: B1/210127592 New:B1/210021453 | DR# 11834 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6005456563613577611/6005456563613577611_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/6005456563613577611/6005456563613577611_signature_31.png
| 2024-08-26 | 2024-08-22 11:30 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29193 | DOH | celinejean16@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1326 |  | GxIV-6C | 838570 | Module 1 for Service Required - H/W failed | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Cleaning/ maintenance Gx ID sticker installation Checking of errors/problem Data collection | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6004883830967104919/6004883830967104919_signature_28.png | Celine Jean villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6004883830967104919/6004883830967104919_signature_31.png
| 2024-08-22 | 2024-08-22 14:57 | Detoxicare Molecular Diagnostics Laboratory Inc. | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29192 | CONSORTIUM | joy.bagsic@detoxicare.com.ph | 25 |  | 2024-08-21 | GXIV 10C | 110118312 | Module B1 for replacement | Replaced B1 | #NAME? | Unit is operational |  | GeneXpert Module SN: 210069174 | 11835 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6001196302026716353/6001196302026716353_signature_28.png | Mary Joy Bagsic | https://www.jotform.com/uploads/macarewecare/222927978372067/6001196302026716353/6001196302026716353_signature_31.png
| 2024-08-22 | 2024-08-22 15:01 | Madrid District Hospital | National Road, Madrid | 13 | Surigao del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29191 | DOH | madrid_dh@gmail.com | 30 |  |  | GX-4 6 Colors | 838439 | None | None | Preventive Maintenance GX ID STICKER INSTALLATION | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6001194546284431392/6001194546284431392_signature_28.png | Rochelle J. Pueblas | https://www.jotform.com/uploads/macarewecare/222927978372067/6001194546284431392/6001194546284431392_signature_31.png
| 2024-08-22 | 2024-08-22 13:52 | Our Lady Of Mt. Carmel Medical Center | San Fernando, Pampanga | 3 | Pampanga | Blood Bank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com | laboratorymt.carmel@gmail.com |  | 29190 | COMMERCIAL | gcvillanueva20@gmail.com | 25 |  | 2024-08-08 | MCS+ | 15G043 | None; Unit for calibration. | Verify and checked calibration parameters. | Performed unit calibration. Performed unit maintenance/cleaning. Replaced pump rotor rubber roller. Run functionality and simulation test. | Unit is calibrated and fully operational. | N/A | Rotor rubber roller | DR #: 11797 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6001178664616521735/6001178664616521735_signature_28.png | Gian Carlo S. Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/6001178664616521735/6001178664616521735_signature_31.png
| 2024-08-22 | 2024-08-22 10:47 | Tubural Rural Health Unit | Tuburan Basilan | BARMM | Basilan | Laboratory | YES | ["GX ID Sticker installation ","GX ID Sticker installation "] | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 29189 | DOH | rjakilan20@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 460 | 2024-08-22 | GX 4 (6c) | 838832 | None | Preventive Maintenence, GX ID sticker installation and data gathering | Preventive Maintenence, GX ID sticker installation and data gathering. | All modules are available | Filter needs to be replaced | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6001044858964979074/6001044858964979074_signature_28.png | Raiyana Hassan Jakilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6001044858964979074/6001044858964979074_signature_31.png
| 2024-08-22 | 2024-08-22 10:47 | Tubural Rural Health Unit | Tuburan Basilan | BARMM | Basilan | Laboratory | YES | REPAIR MAINTENANCE GX ID Sticker installation | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 29188 | DOH | rjakilan20@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 460 | 2024-08-22 | GX 4 (6c) | 838832 | None | Preventive Maintenence, GX ID sticker installation and data gathering | Preventive Maintenence, GX ID sticker installation and data gathering. | All modules are available | Filter needs to be replaced | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6001044858964979074/6001044858964979074_signature_28.png | Raiyana Hassan Jakilan | https://www.jotform.com/uploads/macarewecare/222927978372067/6001044858964979074/6001044858964979074_signature_31.png
| 2024-08-22 | 2024-08-22 10:01 | Teodor B. Galagar District Hospital | Jagna, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29187 | DOH | vcorre84@gmail.com | 29 |  |  | Gx 4 | 838292 | Defective Module 2 | Clean Fan Filter Gx ID Sticker Installation | Clean Fan Filter Gx ID Sticker Installation | H/W MODULE 2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6001019026241858343/6001019026241858343_signature_28.png | VENUS C. ATON,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6001019026241858343/6001019026241858343_signature_31.png
| 2024-08-22 | 2024-08-21 14:33 | Malungon Rural Health Unit | Poblacion, Malungon, Sarangani Province | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29185 | DOH | ulbenarioe@yahoo.com | 28.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 766 |  | GxIV-6C | 838546 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert Cleaning/maintenance Gx ID sticker installation Checking of errors Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000997272517288290/6000997272517288290_signature_28.png | Ermie D Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/6000997272517288290/6000997272517288290_signature_31.png
| 2024-08-22 | 2024-08-21 09:44 | Isulan Rural Health Unit | Poblacion Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29186 | UNOPS | isulanrhuclin.lab@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 52 |  | GxIV-10C | 110023520 | Module 1 stuck cartridge | Perform Genexpert machine repair | Repair module 1 stuck cartridge | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000997272519325519/6000997272519325519_signature_28.png | Glenndale Latrez | https://www.jotform.com/uploads/macarewecare/222927978372067/6000997272519325519/6000997272519325519_signature_31.png
| 2024-08-21 | 2024-08-20 15:18 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29184 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1468 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Others, COUNT: 98 ASSAY: Xpert HBV VL, COUNT: 1054 ASSAY: Xpert HPV, COUNT: 129 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2330 ASSAY: Others, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 12259 ASSAY: Xpert HCV VL, COUNT: 140 | 2024-08-15 | GX16 D | 110005587 | For module replacement A4 882420 - 210071029, C4 890075 - 210071450 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | A4 ,C4 operational newly replaced A1,A2,C2,D2 reported already for replacement. 12 modules operational. | Work with Ronnel Belizar | Genexpert module | 98621 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000338733613866679/6000338733613866679_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6000338733613866679/6000338733613866679_signature_31.png
| 2024-08-21 | 2024-08-20 14:51 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29183 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 43 ASSAY: Others, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 21 ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 49 ASSAY: Xpert HCV VL, COUNT: 5 | 2024-08-15 | GX4 D | 838524 | For module replacement B1 862543 -210062293, B3 862548 -210070806 | Performed module replacement. Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational | Work with Ronnel Belizar | Genexpert module | 98621 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000338483618515342/6000338483618515342_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/6000338483618515342/6000338483618515342_signature_31.png
| 2024-08-21 | 2024-08-21 15:01 | Macare In House | 1005 Aurora Grand,Aurora blvd | NCR | NCR 2nd - Quezon City | In house | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com | florescanoel77@gmail.com |  | 29182 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | Centrifuge Assy | 5014 | Centrifuge for troubleshooting Previously reported sudden shutdown during centrifugation. Stuck up centrifuge rotor. | Test centrifuge assy in another machine after press the run button for centrifugation sudden shutdown encounter. Checked motor and rotor | Performed troubleshooting | For replacement and quotaion of centrifuge assy. | Defective centrifuge assy |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000338023612912179/6000338023612912179_signature_28.png | Macare Medical Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/6000338023612912179/6000338023612912179_signature_31.png
| 2024-08-21 | 2024-08-21 15:08 | Cabadbaran City Health Office | Cabadbaran City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29181 | DOH | samstercem@gmail.com | 26 |  |  | GX-4 6 Colors | 838550 | None | None | Preventive Maintenance GX ID Sticker Installation | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000335992339415345/6000335992339415345_signature_28.png | Sammy Cempron | https://www.jotform.com/uploads/macarewecare/222927978372067/6000335992339415345/6000335992339415345_signature_31.png
| 2024-08-21 | 2024-08-21 13:50 | Lasam District Hospital | IBJ, Lasam, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29180 | DOH | lasamdh.laboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 432 |  | GXIV-6C | 838530 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000291221516307945/6000291221516307945_signature_28.png | Verbell Manuela | https://www.jotform.com/uploads/macarewecare/222927978372067/6000291221516307945/6000291221516307945_signature_31.png
| 2024-08-21 | 2024-08-21 12:42 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29179 | DOH | bmccrcl2022@gmail.com | 30 |  |  | GX-4 6 Colors | 838403 | B2 found error in temperature | Error checking using XpertCheck Calibration in Module B2 | Preventive Maintenance XpertCheck Calibration GX ID Sticker Installation | 3 modules working, 1 module found service required All GX accessories are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000255012337145892/6000255012337145892_signature_28.png | Kristine Pelletero | https://www.jotform.com/uploads/macarewecare/222927978372067/6000255012337145892/6000255012337145892_signature_31.png
| 2024-08-21 | 2024-08-21 12:01 | Alfonso Ponce Enrile Memorial District Hospital | Gonzaga, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29178 | DOH | apemdh_laboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2028 |  | GXIV-6C | 838474 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000253849413030993/6000253849413030993_signature_28.png | Shelly Vee Maona C. Naval | https://www.jotform.com/uploads/macarewecare/222927978372067/6000253849413030993/6000253849413030993_signature_31.png
| 2024-08-21 | 2024-08-21 11:58 | Matanao Rural Health Unit | Poblacion Matanao | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29177 | DOH | pamsustento@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 523 |  | GX-IV 4D 6C | 838456 | For XPC Error Verification at Module 2 | Install new Calibration XPC Version 1.5B Run XPC at Module 2 | Run XPC at Module 2 | Module 2, retry and for service required Module 2 for under observation, if error persist, suggested to disable the said module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000231395523534877/6000231395523534877_signature_28.png | Pamela Joyce Sustento RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/6000231395523534877/6000231395523534877_signature_31.png
| 2024-08-21 | 2024-08-21 12:07 | ACE Hospital | Bohol | 7 | Bohol | Laboratory | NO | ID sticker Installation | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29176 | DOH | laboratory@acemcbohol.com | 25 |  |  | Gx 4 | 840969 | None | ID sticker installation Clean Fan Filter Plunger rod maintenance | ID sticker installation Clean Fan Filter Plunger rod maintenance | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000227586313914991/6000227586313914991_signature_28.png | ACEMC-BOHOL | https://www.jotform.com/uploads/macarewecare/222927978372067/6000227586313914991/6000227586313914991_signature_31.png
| 2024-08-21 | 2024-08-20 15:44 | Rapu-Rapu Rural Health Unit | Poblacion, Rapu-Rapu | 5 | Albay | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29175 | DOH | suzetteoguan1@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 |  | GX4-6C | 838565 | Gx for Calibration Gx for Preventive Maintenance | XpertCheck Result: B1/862627 - passed B2/730958 - newly replaced module B3/757493 - newly replaced module B4/673571 - service required Cleaned cartridge bays, plunger rod, icore Attached QR sticker on Gx Generate Data | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000217962204603130/6000217962204603130_signature_28.png | Suzette C.Oguan | https://www.jotform.com/uploads/macarewecare/222927978372067/6000217962204603130/6000217962204603130_signature_31.png
| 2024-08-21 | 2024-08-19 13:26 | Estevez Memorial Hospital, Inc. | Guevarra Subdivision, Legazpi City | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29174 | COMMERCIAL | melosept22@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 126 ASSAY: Xpert MTB/RIF Ultra, COUNT: 76 ASSAY: Others, COUNT: 6 ASSAY: Xpert SARS-CoV-2, COUNT: 11 |  | GX4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Attached QR sticker on Gx Generate Data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000217672206865488/6000217672206865488_signature_28.png | Emh | https://www.jotform.com/uploads/macarewecare/222927978372067/6000217672206865488/6000217672206865488_signature_31.png
| 2024-08-21 | 2024-08-21 11:02 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | PCR LABORATORY | YES | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29173 | MACARE-HBDC | honeyeichel@gmail.com | 26 | ASSAY: Xpert HBV VL, COUNT: 17 ASSAY: Xpert SARS-CoV-2, COUNT: 7594 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Others, COUNT: 940 | 2024-08-20 | GXIV-D | 829390 | B1:662752 B2:623847 B3:792278 For module replacement | Module replacement Preventive maintenance | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational |  | 3 Genexpert module | 11825 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000204452176745313/6000204452176745313_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/6000204452176745313/6000204452176745313_signature_31.png
| 2024-08-21 | 2024-08-21 10:45 | Manuel J. Santos Hospital | 554 Montilla Blvd. Butuan City | 13 | Agusan del Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29172 | COMMERCIAL | molecularlab@mjsantos.com.ph | 27 |  |  | GX 4 6 Colors | 831292 | B2 error in temperature | None | Preventive Maintenance GX ID Sticker Installation | 3 modules working, 1 module disabled (error) All GX accessories are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000185852338767973/6000185852338767973_signature_28.png | Daniella Andrei Plaza | https://www.jotform.com/uploads/macarewecare/222927978372067/6000185852338767973/6000185852338767973_signature_31.png
| 2024-08-21 | 2024-08-21 09:49 | Maluso Rural Health Unit | Maluso Basilan | BARMM | Basilan | Laboratory | NO | MAINTENANCE Installation of GX Barcode Sticker | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 29171 | DOH | wadznaibrahom0706@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 569 | 2024-08-21 | GX 4 (6c) | 838835 | None | Preventive Maintenence, Data Collection and gx barcode sticker | Preventive Maintenence, Data Collection and gx barcode sticker | All modules are available | System unit battery needs to be replaced | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000157300422100698/6000157300422100698_signature_28.png | Wanna Ibrahim | https://www.jotform.com/uploads/macarewecare/222927978372067/6000157300422100698/6000157300422100698_signature_31.png
| 2024-08-21 | 2024-08-21 09:52 | Baggao Rural Health Unit | San Jose, Baggao, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29170 | DOH | baggaorhulaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 663 |  | GXIV-6C | 838531 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/6000147977914092840/6000147977914092840_signature_28.png | Jan Anthony Eslabra | https://www.jotform.com/uploads/macarewecare/222927978372067/6000147977914092840/6000147977914092840_signature_31.png
| 2024-08-20 | 2024-08-15 11:36 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29168 | DOH | jan19.cam@gmail.com | 25.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1284 |  | GxIV-6C | 838504 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert Cleaning /preventive maintenance Gx ID sticker installation Checking of errors/problem Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999662846862652633/5999662846862652633_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5999662846862652633/5999662846862652633_signature_31.png
| 2024-08-20 | 2024-08-13 14:30 | Antipas Rural Health Unit | Antipas, North Cotabato | 12 | Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29169 | DOH | antipasclinicallaboratory@gmail.com | 23.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 746 |  | GxIV-6C | 838259 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert Preventive maintenance/Cleaning Gx ID sticker installation Checking of errors/problem Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999662846865476883/5999662846865476883_signature_28.png | PRINCESS JERRUMAE PORRAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5999662846865476883/5999662846865476883_signature_31.png
| 2024-08-20 | 2024-08-20 10:52 | Isulan Rural Health Unit | Poblacion Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29167 | UNOPS | nancymoscoso18@gmail.com | 24.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 51 |  | GxIV-10C | 110023520 | Module 4 stuck cartridge | Perform Genexpert machine Preventive maintenance | Repair module 4 stuck cartridge Perform Genexpert Cleaning and maintenance Checking of error/problem Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999662756866462598/5999662756866462598_signature_28.png | Nancy A. Moscoso | https://www.jotform.com/uploads/macarewecare/222927978372067/5999662756866462598/5999662756866462598_signature_31.png
| 2024-08-20 | 2024-08-06 12:39 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29166 | COMMERCIAL | gabbyg.carag@gmail.com | 24 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Simulation run Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999507403614036640/5999507403614036640_signature_28.png | Ma Gabriela Grace Frances Carag | https://www.jotform.com/uploads/macarewecare/222927978372067/5999507403614036640/5999507403614036640_signature_31.png
| 2024-08-20 | 2024-08-06 12:05 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 29165 | COMMERCIAL | microbiology.csmc@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 145 ASSAY: Xpert HPV, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1119 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999507243617589821/5999507243617589821_signature_28.png | Ellaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5999507243617589821/5999507243617589821_signature_31.png
| 2024-08-20 | 2024-08-02 16:01 | St. Luke's Medical Center BGC | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29164 | COMMERCIAL | ci.bgc@stlukes.com.ph | 20 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 18 ASSAY: Xpert CT/NG, COUNT: 90 ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert HPV, COUNT: 109 ASSAY: Xpert MTB/RIF Ultra, COUNT: 223 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV-D-10C | 120002383 | Printer wasted pad end of service life | Cleaned printer waste pad | Checked printer Cleaned printer waste pad Performed maintenance: Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. Printer is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999507063616845473/5999507063616845473_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5999507063616845473/5999507063616845473_signature_31.png
| 2024-08-20 | 2024-08-02 15:44 | St. Luke's Medical Center BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29162 | COMMERCIAL | ci.bgc@stlukes.com.ph | 20 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999506843618983529/5999506843618983529_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5999506843618983529/5999506843618983529_signature_31.png
| 2024-08-20 | 2024-08-02 15:44 | St. Luke's Medical Center BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29163 | COMMERCIAL | ci.bgc@stlukes.com.ph | 20 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999506853612212455/5999506853612212455_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5999506853612212455/5999506853612212455_signature_31.png
| 2024-08-20 | 2024-08-02 11:13 | Pinagbuhatan Health Center | Urbano Velasco Ave., Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29161 | MACARE-HBDC | peachmango318@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 152 |  | GX-IV-D-10C | 110024245 | Module B1 has error 4012 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Checked module B1 Generated reports Disabled module B1 | Unit is operational. Module B1/210127592 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999506573611459784/5999506573611459784_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/5999506573611459784/5999506573611459784_signature_31.png
| 2024-08-20 | 2024-08-16 11:14 | Antipolo City Health Office | M. Santos St. Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29160 | UNOPS | acho.laboratory@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-IV-D-10C | 110023496 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999506363618843041/5999506363618843041_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5999506363618843041/5999506363618843041_signature_31.png
| 2024-08-20 | 2024-08-20 15:45 | Piat Municipal Health Office | Poblacion 2, Piat, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | sentrongsigla1999@yahoo.com.ph |  | 29159 | DOH | marieatalipsalas@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 95 |  | GXIV-6C | 838532 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999496055615398559/5999496055615398559_signature_28.png | Marie Atalip Salas | https://www.jotform.com/uploads/macarewecare/222927978372067/5999496055615398559/5999496055615398559_signature_31.png
| 2024-08-20 | 2024-08-20 14:49 | Tuao Rural Health Unit | Centro 2, Tuao, Cagayan | 2 | Cagayan | Tb Dots | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29158 | DOH | annrollaine@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 382 |  | GXIV-6C | 838465 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999460995615969077/5999460995615969077_signature_28.png | Rollaine Ann Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5999460995615969077/5999460995615969077_signature_31.png
| 2024-08-20 | 2024-08-20 13:54 | Conner District Hospital | Ripang, Conner, Apayao | CAR | Apayao | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29157 | MACARE-HBDC | cdhlaboratory2019@gmail.com | 28 | ASSAY: Xpert SARS-CoV-2, COUNT: 17 ASSAY: Others, COUNT: 67 |  | GXIV-6C | 847727 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999429739326118610/5999429739326118610_signature_28.png | Mark Talay | https://www.jotform.com/uploads/macarewecare/222927978372067/5999429739326118610/5999429739326118610_signature_31.png
| 2024-08-20 | 2024-08-20 12:27 | Sta Cruz Rural Health Unit | Poblacion, Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29156 | DOH | chendalagan72@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 577 |  | GX-IV 4D 6C | 838505 | Module Error at Module B1 For Preventive Maintenance and GX Sticker Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XPC at Module B1 Install GX Sticker | Perform Preventive Maintenance, Calibration and Sticker Installation | Module B1 passed the Calibration All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999376013212043185/5999376013212043185_signature_28.png | Dorothy A. Dalagan RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5999376013212043185/5999376013212043185_signature_31.png
| 2024-08-20 | 2024-08-20 11:08 | Calubian Rural Health Unit | Calubian, Leyte | 8 | Leyte | TB Dots | NO | CALIBRATION MAINTENANCE GeneXpert ID Sticker Installation | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29155 | DOH | mjcabardo06@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 861 |  | 6C GXIV-4D | 838399 | Due for XPC | #NAME? | #NAME? | All modules passed XPC. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999326435196016567/5999326435196016567_signature_28.png | Mighty Joy V. Cabardo | https://www.jotform.com/uploads/macarewecare/222927978372067/5999326435196016567/5999326435196016567_signature_31.png
| 2024-08-20 | 2024-08-20 10:45 | Delfin Albano Rural Health Unit | Ragan Sur, Delfin Albano, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29154 | DOH | valorozos@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 938 |  | GXIV-6C | 838208 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999313604524667292/5999313604524667292_signature_28.png | Stephen Valorozo | https://www.jotform.com/uploads/macarewecare/222927978372067/5999313604524667292/5999313604524667292_signature_31.png
| 2024-08-20 | 2024-08-16 16:08 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29153 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 29 |  |  | MCS + | 15H028G | Exhaust fan motor unoperational, no voltage output | Checked and test voltage output connector fan from personality board Swap power supply. Swap cable. After troubleshooting is performed system status still the same fan motor unoperational | Performed troubleshooting | Unit unoperational | Defective personality board for replacement and quotation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999238603612279249/5999238603612279249_signature_28.png | Philippine Blood Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5999238603612279249/5999238603612279249_signature_31.png
| 2024-08-20 | 2024-08-16 16:13 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Blood bank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29152 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 29 |  |  | MCS + | 15H029G | Defective CPU and display cable | Borrowed temporary CPU Board and display cable frm 15H028G to temporary use the machine | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999238253611135573/5999238253611135573_signature_28.png | Philippine Blood Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5999238253611135573/5999238253611135573_signature_31.png
| 2024-08-20 | 2024-08-20 08:30 | Cauayan City Health Office | Cabaruan, Cauayan City, Isabela | 2 | Isabela | Microbiology | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE Relocation | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29151 | UNOPS | kj102412@gmail.com | 29 |  |  | GXIV-10C | 110023504 | For Relocation | Relocate Genexpert | Performed Maintenance Gathered Reports Performed Functionality Test Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5999234274524275806/5999234274524275806_signature_28.png | Stephen P. Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5999234274524275806/5999234274524275806_signature_31.png
| 2024-08-19 | 2024-08-16 08:39 | Mangatarem Rural Health Unit | Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29150 | DOH | slasheiji1988new@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 642 |  | GXIV-4-D-6C | 838284 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998728077495121715/5998728077495121715_signature_28.png | Ken Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/5998728077495121715/5998728077495121715_signature_31.png
| 2024-08-19 | 2024-08-15 14:05 | Infanta Rural Health Unit | Infanta, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29149 | UNOPS | charissealmandres8@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 17 |  | GXIV-4-D-10C | 110023501 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998727237492007126/5998727237492007126_signature_28.png | Angelie Paz Olivar | https://www.jotform.com/uploads/macarewecare/222927978372067/5998727237492007126/5998727237492007126_signature_31.png
| 2024-08-19 | 2024-08-15 12:09 | Dasol Community Hospital | Dasol, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29148 | DOH | jeremypulido0521@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 664 |  | GXIV-4-D-6C | 838287 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998726037494862190/5998726037494862190_signature_28.png | Jeremy Pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/5998726037494862190/5998726037494862190_signature_31.png
| 2024-08-19 | 2024-08-14 11:22 | Bolinao Rural Health Unit | Bolinao, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29147 | DOH | nylemae14@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 626 |  | GXIV-4-D-6C | 838278 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998716477496073878/5998716477496073878_signature_28.png | Emylyn C Celzo | https://www.jotform.com/uploads/macarewecare/222927978372067/5998716477496073878/5998716477496073878_signature_31.png
| 2024-08-19 | 2024-08-19 15:27 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 29146 | MACARE-HBDC | laboratory@slec.ph | 21 |  |  | Infinity 48s | 809643 | - Error 6001: infinity server offline - UPS constantly beeping | #NAME? | #NAME? | Unit is operational, embedded PC did not restart properly, UPS was overheating due to stuck up coolong fans. Lubricated cooling fans for temporary fix but recommend replacing UPS as it is almost 7 years old. (Powervar Sinergy III 2KVA Model or similar) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998622621819264147/5998622621819264147_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5998622621819264147/5998622621819264147_signature_31.png
| 2024-08-19 | 2024-08-14 15:32 | Esperanza Rural Health Unit | Poblacion Iligan Rd., Brgy. Poblacion, Masbate | 5 | Masbate | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29145 | DOH | floramarmoncada@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1238 |  | GX4-6C | 838268 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Attached QR sticker on Gx Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998426709511321007/5998426709511321007_signature_28.png | Floramar M. Antipolo | https://www.jotform.com/uploads/macarewecare/222927978372067/5998426709511321007/5998426709511321007_signature_31.png
| 2024-08-19 | 2024-08-13 14:59 | Cawayan Rural Health Unit | Cawayan | 5 | Masbate | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29144 | DOH | gellatisimo@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 877 |  | GX4-6C | 838267 | Gx for Calibration | XpertCheck Result: C1/807602 - passed C2/807592 - passed C3/784120 - passed C4/621794 - newly replaced module Cleaned cartridge bays, plunger rod, icore Generate data Attached QR sticker on Gx Checked the battery dependability of UPS wheb unplugged drom AC source, checked power connection of UPS | Performed XpertCheck Performed Preventive Maintenance Performed Troubleshooting on UPS | Unit is Operational UPS battery drained | UPS last a few seconds when unplugged from power source. Drained UPS battery, request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998426369517342144/5998426369517342144_signature_28.png | Gellie Jane Gader | https://www.jotform.com/uploads/macarewecare/222927978372067/5998426369517342144/5998426369517342144_signature_31.png
| 2024-08-19 | 2024-08-13 10:09 | Masbate City Health Office | City Hall Compound, Brgy. Kalipay, Masbate City | 5 | Masbate | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29143 | UNOPS | Juedevan.ancajas14@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 33 |  | GX4-10C | 110023512 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate Data Attachment of QR sticker on Gx | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998424109514958086/5998424109514958086_signature_28.png | JUEDE VAN A. ANCAJAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5998424109514958086/5998424109514958086_signature_31.png
| 2024-08-19 | 2024-08-12 16:15 | Claveria Rural Health Unit | Poblacion, Claveria, Burias Island | 5 | Masbate | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29142 | DOH | mgrmartinez757@gmail.com | 29 |  |  | GX4-6C | 838572 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Attached QR sticker on Gx Generate Data Checked the power output of UPS, checked the power supply of UPS after unplugging from AC power source, checked the power connection from/to UPS | Performed Preventive Maintenance Performed Troubleshooting for UPS | Unit is Operational | UPS last 3mins at most when cutting connection from AC or if there is a power outage(frequent occurence within a day). UPS battery drained, request for replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998423289519863526/5998423289519863526_signature_28.png | Mary Grace R.Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5998423289519863526/5998423289519863526_signature_31.png
| 2024-08-19 | 2024-08-12 10:12 | Pioduran Rural Health Unit | Caratagan, Pioduran, Albay | 5 | Albay | Laboratory GX4-6C | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29141 | DOH | mitzijc@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 649 |  | GX4-6C | 838264 | GX for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Attached QR sticker on Gx Generate data | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5998418749515092420/5998418749515092420_signature_28.png | Mitzi D. Cedeño | https://www.jotform.com/uploads/macarewecare/222927978372067/5998418749515092420/5998418749515092420_signature_31.png
| 2024-08-16 | 2024-08-15 11:47 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29140 | MACARE-HBDC | lcpmicrolab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 396 ASSAY: Xpert SARS-CoV-2, COUNT: 924 ASSAY: Xpert HIV-1 Viral Load, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 114 |  | GX 4 D | 838837 | Due for xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck Performed maintenance | B1,B4 failed in xpertcheck B2,B3 retry. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5996210288987169672/5996210288987169672_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5996210288987169672/5996210288987169672_signature_31.png
| 2024-08-16 | 2024-08-15 10:49 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29139 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 26 |  |  | MCS + | 15H030G | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5996209838982762216/5996209838982762216_signature_28.png | Philippine Blood Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5996209838982762216/5996209838982762216_signature_31.png
| 2024-08-16 | 2024-08-15 10:27 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Blood bank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29138 | MACARE-HBDC | pbcapheresis@doh.gov.ph | 26 |  |  | MCS + | 15H028G | Due for maintenance Exhaust fan not working | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Replaced cable of exhaust fan Performed functionality test Performed simulation running | Performed maintenance and troubleshooting | For further troubleshooting exhaust fan |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5996209448983493226/5996209448983493226_signature_28.png | Philippine Blood Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5996209448983493226/5996209448983493226_signature_31.png
| 2024-08-16 | 2024-08-16 14:01 | FEU Nicanor Reyes Medical Foundation | Regalado Ave., West Fairview, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29137 | CONSORTIUM | kmtejoso@feu-nrmf.edu.ph | 24 | ASSAY: Xpert HPV, COUNT: 51 ASSAY: Xpert MTB/RIF Ultra, COUNT: 413 |  | GXIV 6C D | 840975 | Unit due for maintenance | Performed preventive maintenance Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Archived results | Unit is operational 3 modules working Module A3 reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5995978001124151552/5995978001124151552_signature_28.png | Kenneth Anthony M. Tejoso | https://www.jotform.com/uploads/macarewecare/222927978372067/5995978001124151552/5995978001124151552_signature_31.png
| 2024-08-16 | 2024-08-16 10:45 | Medical Center Manila | 850 UN Ave, Paco, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29136 | CONSORTIUM | microbiologylab.manilamed@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1095 |  | GXIV 6C D | 835540 | Unit due for maintenance | Performed preventive maintenance Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Archived test results | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5995863781124260114/5995863781124260114_signature_28.png | Jericho Cabasug | https://www.jotform.com/uploads/macarewecare/222927978372067/5995863781124260114/5995863781124260114_signature_31.png
| 2024-08-16 | 2024-08-16 10:07 | AFPMC AFRIMS 10C | V. Luna Road, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29135 | MACARE-HBDC | mtvalderama@gmail.com | 28 | ASSAY: Xpert HBV VL, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 692 ASSAY: Xpert SARS-CoV-2, COUNT: 1046 ASSAY: Xpert HCV VL, COUNT: 31 ASSAY: Xpert HIV-1 Viral Load, COUNT: 39 ASSAY: Xpert HIV-1 Viral Load, COUNT: 105 | 2024-08-09 | GXIV 10C L | 845462 | Unit due for Calibration | Performed preventive maintenance Performed Xpertcheck Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck on B1,2,4 B3 already reported All modules passed | Unit is operational 3 modules working | Service done on 8/15/24. jotform error | Xpertcheck kit | DR |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5995837651126761163/5995837651126761163_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/5995837651126761163/5995837651126761163_signature_31.png
| 2024-08-16 | 2024-08-16 09:45 | AFPMC AFRIMS 6C | V. Luna Road, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29134 | MACARE-HBDC | mtvalderama@gmail.com | 23 | ASSAY: Xpert HBV VL, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 633 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2024-08-09 | GXIV 6C L | 805294 | Modules B1 and B3!for replacement Unit due for Xpertcheck | Performed preventive maintenance Replaced modules B1 and B3 Modules B2 and B4 for reporting Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced modules B1: 35793 -> 210071525 B3: 715913 -> 210063075 Performed Xpertcheck Module B2 and B4 for reporting | Unit is operational | Service done on 8/15/24. jotform error | 2 GX Modules 6C 1 Xpertcheck kit | Dr |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5995827711123951957/5995827711123951957_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/5995827711123951957/5995827711123951957_signature_31.png
| 2024-08-15 | 2024-08-15 15:50 | Talavera General Hospital | Maestrang Kikay, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29133 | MACARE-HBDC | pjgtehlab@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 708 ASSAY: Xpert HBV VL, COUNT: 72 ASSAY: Xpert CT/NG, COUNT: 34 ASSAY: Xpert HCV VL, COUNT: 10 ASSAY: Xpert HIV-1 Viral Load, COUNT: 116 | 2024-07-18 | GXIV-10C | 110013773 | Module B1 and B2 error | Performed Module Replacement B1(210066389 - 779453) B2(210065641 - 786610) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  | 11728 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5995178351125694525/5995178351125694525_signature_28.png | Sharey Capati | https://www.jotform.com/uploads/macarewecare/222927978372067/5995178351125694525/5995178351125694525_signature_31.png
| 2024-08-15 | 2024-08-15 14:27 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY, CEBU | 7 | Cebu | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29132 | DOH | mendozarey42@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4784 ASSAY: Xpert HCV VL, COUNT: 39 ASSAY: Xpert HIV-1 Viral Load, COUNT: 243 |  | GX-IV | 838508 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5995141458336001600/5995141458336001600_signature_28.png | Rey C. Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5995141458336001600/5995141458336001600_signature_31.png
| 2024-08-15 | 2024-08-15 13:52 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29131 | DOH | luzvimindanavidad2@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1037 |  | 6C GXIV-4D | 838297 | Due for XPC | #NAME? | #NAME? | Module B1 failed XPC, for replacement. Other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5995113235003070663/5995113235003070663_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5995113235003070663/5995113235003070663_signature_31.png
| 2024-08-15 | 2024-08-15 12:03 | Batangas City Health Office | Poblacion 17, P. Burgos St. Batangas City | 4A | Batangas | Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29130 | MACARE-HBDC | gjanlorenzo@gmail.com | 25 |  |  | GXIV-10C | 110024615 | For installation of GeneXpert System. | Installation. | Performed installatiom of GeneXpert System. Performed basic training and troubleshooting. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5995041755425764915/5995041755425764915_signature_28.png | Jan Lorenzo Godoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5995041755425764915/5995041755425764915_signature_31.png
| 2024-08-15 | 2024-07-31 11:06 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29129 | DOH | arleneabad32@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 750 |  | GXIV-4-D-6C | 838288 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993437732601419457/5993437732601419457_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5993437732601419457/5993437732601419457_signature_31.png
| 2024-08-15 | 2024-08-13 10:33 | Naujan Municipal Health Office | Brgy. Santiago, Naujan | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29128 | DOH | baldoza.shane25@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 806 | 2024-08-09 | GXIV | 838226 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module B1 (632561), service required. |  | Xpertcheck kit. | DR#11803 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994963939618913962/5994963939618913962_signature_28.png | Lucena Malicsi | https://www.jotform.com/uploads/macarewecare/222927978372067/5994963939618913962/5994963939618913962_signature_31.png
| 2024-08-15 | 2024-08-15 09:39 | Quezon City General Hospital 2 | Seminary Rd. Project 8, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29127 | MACARE-HBDC | jomarilitonjua@gmail.com | 25 |  |  | GX IV 10C-D | 110023495 | GeneXpert detected but modules does not appear on GX DX Software | Troubleshoot the connection problem Performed preventive maintenance generated report | Checked IP address Ping IP address Correct IP - still not detected Opened GX and tightened all cables Problem solved Checked and cleaned cartridge bays, plunger rod, icore slot and filter | Unit is operational All modules working | Performed on 8/13/24. Jotform error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994958122621082898/5994958122621082898_signature_28.png | Jomari Bryan Litonjua | https://www.jotform.com/uploads/macarewecare/222927978372067/5994958122621082898/5994958122621082898_signature_31.png
| 2024-08-14 | 2024-08-12 15:43 | Davao Reproductive Health and Wellness Center | Jacinto Street, Poblacion District Davao | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29126 | DOH | rhwclab@gmail.com | 27.5 | ASSAY: Xpert HBV VL, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 275 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1236 |  | GX-IV 4D 6C | 838262 | Stuck Cartidge at Module 3 | Removed the stuck cartidge at Module 3, clean module cartidge bay | Perform removal of stuck cartidge | Successfully removed the stuck cartidge |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994454380811752133/5994454380811752133_signature_28.png | Cherry Ann Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5994454380811752133/5994454380811752133_signature_31.png
| 2024-08-14 | 2024-08-14 10:14 | Kidapawan City Health Office | Poblacion Kidapawan | 11 | Cotabato | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29125 | DOH | malufranz5@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2121 |  | GX-IV 4D 6C | 838545 | For Preventive Maintenance and GX Sticker Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Install GX Sticker | Perform Preventive Maintenance and GX Sticker Installation | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994453110811529414/5994453110811529414_signature_28.png | Ma. Luisa Francisco, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5994453110811529414/5994453110811529414_signature_31.png
| 2024-08-14 | 2024-08-14 16:38 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | Cebu City | 7 | Cebu | Blood Bank | NO | REPAIR | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29124 | COMMERCIAL | cduhbloodbank2@gmail.com | 23 |  |  | MCS+ | 14F014 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994343338302100616/5994343338302100616_signature_28.png | Christine Jade T Lagat | https://www.jotform.com/uploads/macarewecare/222927978372067/5994343338302100616/5994343338302100616_signature_31.png
| 2024-08-14 | 2024-08-14 15:58 | Sulu provincial hospital | Sulu provincial hospital | BARMM | Sulu | Laboratory | YES | MAINTENANCE GX Barcode Sticker installation and PM | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia@gmail.com |  | 29123 | DOH | sphcl2016@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1457 | 2024-08-14 | GX 4 | 838838 | Module 1 hardware failed, UPS is defective must replace both defective devices/Parts. | Preventive Maintenence, Troubleshooting for UPS and GX barcode sticker installation | Preventive Maintenence, Troubleshooting for UPS and GX barcode sticker installation | Modules 2,3 and 4 are available while module 1 is currently hardware failed. UpS battered must be replaced. | Modules 2,3 and 4 are available while module 1 is currently hardware failed. UpS battered must be replaced. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994320673529611178/5994320673529611178_signature_28.png | Abdul anam H. Nuh, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5994320673529611178/5994320673529611178_signature_31.png
| 2024-08-14 | 2024-08-14 14:46 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | Tb Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29122 | DOH | sajecopooh20@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1096 |  | 6C GXIV-4D | 838555 | Due for XpertCheck Calibration | #NAME? | #NAME? | Module A1 and A3 passed XPC. Module A2 and A4 failed, for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994292705709431282/5994292705709431282_signature_28.png | Sarah Jean R. Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5994292705709431282/5994292705709431282_signature_31.png
| 2024-08-14 | 2024-08-14 13:37 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29121 | DOH | suzzettekho34@gmail.com | 30 |  |  | GX-4 6 Colors | 838554 | Module B4 found Service Required thru XPC | Reported to Cepheid | GX ID Sticker Installation Preventive Maintenance Error Checking of Module B4 thru XpertCheck Calibration | 3 modules working, 1 module for replacement ( service required ) All GX accessories are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994238760128935512/5994238760128935512_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/5994238760128935512/5994238760128935512_signature_31.png
| 2024-08-14 | 2024-08-14 13:31 | Cebu TB Reference Lab - SRP | SRP cebu city | 7 | Cebu | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29120 | DOH | asmantiquilla.pbsp@gmail.com | 26 |  |  | Gx 4 | 838263 | Due Date for Preventive Maintenance | Preventive Maintenance Clean Fan Filter Plunger rod maintenance | Preventive Maintenance Clean Fan Filter Plunger rod maintenance | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994238110429750455/5994238110429750455_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5994238110429750455/5994238110429750455_signature_31.png
| 2024-08-14 | 2024-08-14 13:27 | Kairos Health Medical Diagnostic | 1123 A. Mabini St, Malate, Manila, 1000 Metro Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29119 | MACARE-HBDC | kairosermita.lab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 52 |  | GX IV 2 10C-D | 110022622 | Unit due for maintenance | Performed preventive maintenance Archived results Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Archived Results | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994227612629121565/5994227612629121565_signature_28.png | Isabelle C. Tesalona | https://www.jotform.com/uploads/macarewecare/222927978372067/5994227612629121565/5994227612629121565_signature_31.png
| 2024-08-14 | 2024-08-13 15:17 | Victor R. Potenciano Medical Center | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE Sticker Placement | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29118 | CONSORTIUM | oro_janet@yahoo.com | 27 |  |  | GXIV | 835600 | For sticker placement | NA | #NAME? | Unit is operational Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994211908544487523/5994211908544487523_signature_28.png | Floriza Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5994211908544487523/5994211908544487523_signature_31.png
| 2024-08-14 | 2024-08-14 11:24 | Manila Doctors Hospital | 667 U.N. Ave., Ermita, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29117 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 19 | ASSAY: Xpert CT/NG, COUNT: 86 ASSAY: Xpert HBV VL, COUNT: 94 ASSAY: Xpert HPV, COUNT: 28 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1297 ASSAY: Xpert HCV VL, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 109 |  | GX IV 10C-D | 110001648 | Unit due for maintenance Module B3 error twice yesterday | Performed preventive maintenance Archived results for technical to check the error Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Checked system log and view results Error 5007 and invalid Gathered data to be checked by technical | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994158558506924578/5994158558506924578_signature_28.png | Ryamin F. Buentiempo | https://www.jotform.com/uploads/macarewecare/222927978372067/5994158558506924578/5994158558506924578_signature_31.png
| 2024-08-14 | 2024-08-14 11:12 | Sulu Sanitarium and General Hospital | Sulu Sanitarium and General Hospital | BARMM | Sulu | Laboratory | YES | MAINTENANCE ["Installation of GX ID sticker and Basic Training for end user","Installation of GX ID sticker and Basic Training for end user"] | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia@gmail.com |  | 29116 | COMMERCIAL | shirleymedtech22@gmail.com | 25 |  | 2024-08-14 | GX 4 | 110015481 | Printer unable to print | Repair printer, Preventive maintenence, Conduct end user training regarding basic troubleshooting and maintenence. Installed GX ID sticker. | Repair printer, Preventive maintenence, Conduct end user training regarding basic troubleshooting and maintenence. Installed GX ID sticker. | All modules are available and printer is working | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994157472426772413/5994157472426772413_signature_28.png | Shirley y.julkanain | https://www.jotform.com/uploads/macarewecare/222927978372067/5994157472426772413/5994157472426772413_signature_31.png
| 2024-08-14 | 2024-08-14 09:40 | Escalante City Health Office | Poblacio | 6 | Negros Occidental | Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29115 | DOH | sheila.negrido@yahoo.com | 23 |  | 2024-08-14 | Gx4 | 838523 | . Install GX ID sticker . Perform Preventive maintenance . Save System log . Save Installation Qualification . Change filter | Preventive maintenance | Preventive maintenance | All module's ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994107485343433103/5994107485343433103_signature_28.png | Sheila Mae Negrido | https://www.jotform.com/uploads/macarewecare/222927978372067/5994107485343433103/5994107485343433103_signature_31.png
| 2024-08-14 | 2024-08-12 14:04 | Bantayan Rural Health Unit | Bantayan Island, Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION Id sticker installation | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 29114 | DOH | mcdelle916bausin@gmail.com | 27 |  |  | Gx 4 | 838509 | Defective module B2 Damaged Fan Filter | ID sticker installation Clean fan filter | ID sticker installation Clean fan filter | Defective module 2 Damage fan filter Need new fan filter |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994105804325265697/5994105804325265697_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/5994105804325265697/5994105804325265697_signature_31.png
| 2024-08-14 | 2024-08-06 09:28 | Overseas Filipino Workers Hospital | San Fernando City, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | tiquifrancisca@gmail.com |  | 29113 | MACARE-HBDC | ofwhlaboratory@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5 |  | GXIV-D 10C | 110018556 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994065729574194562/5994065729574194562_signature_28.png | Francisca Tiqui | https://www.jotform.com/uploads/macarewecare/222927978372067/5994065729574194562/5994065729574194562_signature_31.png
| 2024-08-14 | 2024-08-02 14:31 | Jose B. Lingad Memorial and General Hospital | San Fernando City, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29112 | MACARE-HBDC | jblmgh.is@gmail.com | 28 | ASSAY: Xpert HBV VL, COUNT: 168 ASSAY: Xpert HCV VL, COUNT: 25 ASSAY: Xpert HIV-1 Viral Load, COUNT: 950 |  | GXIV-D 10C | 1100137448 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994065439579873993/5994065439579873993_signature_28.png | Concepcion G. Alfonso | https://www.jotform.com/uploads/macarewecare/222927978372067/5994065439579873993/5994065439579873993_signature_31.png
| 2024-08-14 | 2024-08-02 13:47 | Jose B. Lingad Memorial and General Hospital | San Fernando City, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29111 | MACARE-HBDC | jblmghmpl@gmail.com | 28 | ASSAY: Others, COUNT: 89 ASSAY: Xpert MTB/RIF Ultra, COUNT: 140 ASSAY: Others, COUNT: 10 ASSAY: Xpert SARS-CoV-2, COUNT: 889 |  | GXIV-D 10C | 110012825 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit us operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994064999577829200/5994064999577829200_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/5994064999577829200/5994064999577829200_signature_31.png
| 2024-08-14 | 2024-07-30 11:46 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29110 | MACARE-HBDC | honeyeichel@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 17 ASSAY: Xpert SARS-CoV-2, COUNT: 8106 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Others, COUNT: 940 |  | GXIV-D | 829390 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except module B1, B2 and B3 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994064659574559269/5994064659574559269_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5994064659574559269/5994064659574559269_signature_31.png
| 2024-08-14 | 2024-07-30 11:42 | Bataan General Hospital ad Medical Center | Balanga Bataan | 3 | Bataan | PCR LABORATORY | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 29109 | MACARE-HBDC | honeyeichel@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert SARS-CoV-2, COUNT: 3662 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 ASSAY: Others, COUNT: 1554 |  | GXIV-D 10C | 110017894 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5994064379576751798/5994064379576751798_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5994064379576751798/5994064379576751798_signature_31.png
| 2024-08-13 | 2024-08-09 10:17 | Naguilian District Hospital | Naguilian, La Union | 1 | La Union | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29108 | DOH | marjelynrayray01@gmail.com | 25 |  | 2024-07-19 | GXIV-4-D-6C | 838426 | For Calibration and Preventive Maintenance | #NAME? | Performed Preventive Maintenance and XpertCheck | #NAME? |  | XpertCheck Kits | DR 11731 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993475492608545424/5993475492608545424_signature_28.png | Marjelyn Rayray | https://www.jotform.com/uploads/macarewecare/222927978372067/5993475492608545424/5993475492608545424_signature_31.png
| 2024-08-13 | 2024-08-08 11:43 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | La Union | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29107 | DOH | piaramuj@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 850 |  | GXIV-4-D-6C | 838270 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993474622603292430/5993474622603292430_signature_28.png | PIA JUMAR T. RAMIREZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5993474622603292430/5993474622603292430_signature_31.png
| 2024-08-13 | 2024-08-08 11:42 | Candon City Health Office | Candon, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29106 | DOH | candoncholaboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 628 |  | GXIV-4-D-6C | 838273 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993473602606363836/5993473602606363836_signature_28.png | Michelle Bucalen | https://www.jotform.com/uploads/macarewecare/222927978372067/5993473602606363836/5993473602606363836_signature_31.png
| 2024-08-13 | 2024-08-08 11:39 | Sta. Maria Rural Health Unit | Sta. Maria, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29105 | UNOPS | jomarcaberto24@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 17 |  | GXIV-4-D-10C | 110023500 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993472632609902525/5993472632609902525_signature_28.png | Jomar Caberto | https://www.jotform.com/uploads/macarewecare/222927978372067/5993472632609902525/5993472632609902525_signature_31.png
| 2024-08-13 | 2024-08-08 10:41 | Sinait District Hospital | Sinait, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29104 | DOH | isdhsinaitlaboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 675 | 2024-06-27 | GXIV-4-D-6C | 838277 | For Calibration and Preventive Maintenance | #NAME? | Performed Preventive Maintenance and Calibration | - Unit is Partially Operational - Did not perform XpertCheck on Reported module A4 (210045938) -All remainin modules passed XpertCheck - Unit has been temporarily Relocated due to hospital construction - Current room has no air-conditioning unit - GeneXpert is recommended to be transferred to a room with air-conditioning |  | XpertCheck Kits | DR 11657 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993470932607063508/5993470932607063508_signature_28.png | Claire Marie Baguyo-Diculen | https://www.jotform.com/uploads/macarewecare/222927978372067/5993470932607063508/5993470932607063508_signature_31.png
| 2024-08-13 | 2024-08-08 07:32 | Black Nazarene Hospital | San Nicolas, Ilocos Norte | 1 | Ilocos Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29103 | CONSORTIUM | tbnhi.lab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 30 |  | GXIV-4-D-6C | 835542 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993469842606984984/5993469842606984984_signature_28.png | Carmella Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5993469842606984984/5993469842606984984_signature_31.png
| 2024-08-13 | 2024-08-07 10:51 | Mariano Marcos State University | Batac, Ilocos Norte | 1 | Ilocos Norte | Microbiology Laboratory | NO | ["Attached Service Sticker ","Attached Service Sticker "] | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29102 | MACARE-HBDC | jpkperalta@mmsu.edu.ph | 30 |  |  | GXIV-4-D-10C | 110022474 | To Attach Service Sticker | Attached Service Sticker | Attached Service Sticker | Unit is Operational Unable to perform Preventive Maintenance Lab is under renovation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993467022609952319/5993467022609952319_signature_28.png | James Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/5993467022609952319/5993467022609952319_signature_31.png
| 2024-08-13 | 2024-08-07 08:34 | Laoag City Health Office | Laoag City, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29101 | DOH | denver.mahor@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 722 | 2024-06-27 | GXIV-4-D-6C | 838416 | #NAME? | #NAME? | Performed Calibration and Preventive Maintenance | #NAME? |  | XpertCheck Kits | DR 11659 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993462962604840626/5993462962604840626_signature_28.png | Rodenver D. Mahor | https://www.jotform.com/uploads/macarewecare/222927978372067/5993462962604840626/5993462962604840626_signature_31.png
| 2024-08-13 | 2024-08-02 11:55 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29100 | DOH | allan.vinluan2013@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 988 |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993461372608435644/5993461372608435644_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5993461372608435644/5993461372608435644_signature_31.png
| 2024-08-13 | 2024-08-02 10:10 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29099 | DOH | jeffreyvrumbaoa@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1309 |  | GXIV-4-D-6C | 838473 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993460022609987858/5993460022609987858_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/5993460022609987858/5993460022609987858_signature_31.png
| 2024-08-13 | 2024-08-01 13:37 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29098 | DOH | dianeruiz147@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 426 |  | GXIV-4-D-6C | 838438 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993458582602767026/5993458582602767026_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5993458582602767026/5993458582602767026_signature_31.png
| 2024-08-13 | 2024-08-01 11:35 | Urdaneta Sacred Heart Hospital | Urdaneta, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29097 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 181 |  | GXIV-4-D-6C | 820303 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993456992604947034/5993456992604947034_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5993456992604947034/5993456992604947034_signature_31.png
| 2024-08-13 | 2024-08-01 10:29 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29096 | DOH | serly_reyes@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 728 |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993455322606876044/5993455322606876044_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/5993455322606876044/5993455322606876044_signature_31.png
| 2024-08-13 | 2024-08-01 08:11 | Dagupan City Health Office | Dagupan City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29095 | DOH | thisisformyfamily@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1455 |  | GXIV-4-D-6C | 838286 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993448252607164303/5993448252607164303_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5993448252607164303/5993448252607164303_signature_31.png
| 2024-08-13 | 2024-07-31 11:06 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29094 | DOH | arleneabad32@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 750 |  | GXIV-4-D-6C | 838288 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Partially Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993437732601419457/5993437732601419457_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5993437732601419457/5993437732601419457_signature_31.png
| 2024-08-13 | 2024-07-31 10:23 | San Jacinto Rural Health Unit | San Jacinto, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29093 | DOH | allen.indong@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 270 |  | GXIV-4-D-6C | 838518 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993436692606426803/5993436692606426803_signature_28.png | Alexander Indong | https://www.jotform.com/uploads/macarewecare/222927978372067/5993436692606426803/5993436692606426803_signature_31.png
| 2024-08-13 | 2024-07-31 08:58 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29092 | DOH | rhupozorrubio@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 705 |  | GXIV-4-D-6C | 838467 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993432272604566668/5993432272604566668_signature_28.png | Jinky Grace Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/5993432272604566668/5993432272604566668_signature_31.png
| 2024-08-13 | 2024-08-13 15:18 | Laoang Rural Health Unit | Laoang Northern Samar | 8 | Northern Samar | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29091 | DOH | markiansarmiento727@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1410 | 2024-08-13 | GXIV | 838202 | Due for Calibration | PM , Calibration | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993431967566952030/5993431967566952030_signature_28.png | mark ian sarmiento | https://www.jotform.com/uploads/macarewecare/222927978372067/5993431967566952030/5993431967566952030_signature_31.png
| 2024-08-13 | 2024-07-30 13:29 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 29090 | DOH | gbmmbg@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993429662608831920/5993429662608831920_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5993429662608831920/5993429662608831920_signature_31.png
| 2024-08-13 | 2024-08-13 14:24 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29089 | COMMERCIAL | vsmmcbloodservices@gmail.com | 24 |  |  | MCS+ | 02A090 | Due Preventive Maintenance and check apheresis machine. | Conduct Maintenance in apheresis machine check rubber pumps, check ups, line sensor and bowl optics | Conduct Maintenance in apheresis machine check rubber pumps, check ups, line sensor and bowl optics | Units are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993397934992435129/5993397934992435129_signature_28.png | Charlotte Mae Ygoña, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5993397934992435129/5993397934992435129_signature_31.png
| 2024-08-13 | 2024-08-13 14:22 | Macare In House | 1005 Aurora Grand, Aurora blvd QC | NCR | NCR 2nd - Quezon City | Macare In House | NO | REPAIR | HAEMONETICS | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 29088 | MACARE-HBDC | nfloresca@macare-medicals.com | 26 |  | 2024-08-13 | MCS + | 13B243 | Without CPU board | Replaced and installed CPU board s/n: MM021214873 Performed functionality test | Repair | Operational |  | CPU board s/n: MM021214873 | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993396643614554414/5993396643614554414_signature_28.png | Macare Medicas Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/5993396643614554414/5993396643614554414_signature_31.png
| 2024-08-13 | 2024-08-13 13:36 | San Roque Rural Health Unit | San Roque Northern Samar | 8 | Northern Samar | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 29087 | DOH | lynnjarito@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 752 | 2024-08-13 | GXIV | 838498 | Due for Calibration | PM and Calibration | #NAME? | Unit is operational (Module C3 for service required) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993368884919720030/5993368884919720030_signature_28.png | LYNN N. JARITO | https://www.jotform.com/uploads/macarewecare/222927978372067/5993368884919720030/5993368884919720030_signature_31.png
| 2024-08-13 | 2024-08-13 13:15 | Matanao Rural Health Unit | Poblacion Matanao | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29086 | DOH | gmdomingso13@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 536 |  | GX-IV 4d 6C | 838456 | For Preventive Maintenance and GX Sticker Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install GX Sticker | Perform Preventive Maintenance and GX Sticker Installation | All Modules are operational Damaged Filter, temporarily used gauze since there are no extra filters available | Will report to jotform for request of replacement filters |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993365302115870943/5993365302115870943_signature_28.png | Dominador S. Castro III | https://www.jotform.com/uploads/macarewecare/222927978372067/5993365302115870943/5993365302115870943_signature_31.png
| 2024-08-13 | 2024-08-13 12:02 | Talacogon District Hospital | San Isidro, Talacogon | 13 | Agusan del Sur | RTDL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29085 | DOH | shyratortor@gmail.com | 30 |  |  | GX-4 6 Colors | 838275 | UPS doesn't able to backup | None | GX ID Sticker Installation Preventive Maintenance | All modules are functional Computer and other accessories (except for UPS) are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993314068517879293/5993314068517879293_signature_28.png | Shyra Ericka Calamba | https://www.jotform.com/uploads/macarewecare/222927978372067/5993314068517879293/5993314068517879293_signature_31.png
| 2024-08-13 | 2024-08-13 08:24 | Pinamalayan Municipal Health Office | Pinamalayan, Oriental Mindoro | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 29084 | DOH | mholaboratory.pinamalayan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1138 | 2024-08-09 | GXIV | 838230 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  | Xpertcheck kit | DR#11802 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993192770184813258/5993192770184813258_signature_28.png | Kris Allen A. Mabansag | https://www.jotform.com/uploads/macarewecare/222927978372067/5993192770184813258/5993192770184813258_signature_31.png
| 2024-08-13 | 2024-08-02 11:13 | Pinagbuhatan Health Center | Urbano Velasco Ave., Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29083 | MACARE-HBDC | peachmango318@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 152 |  | GX-IV-D-10C | 110024245 | Module B1 has error 4012 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Checked module B1 Generated reports Disabled module B1 | Unit is operational. Module B1/210127592 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5993184273612058838/5993184273612058838_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/5993184273612058838/5993184273612058838_signature_31.png
| 2024-08-12 | 2024-08-09 10:35 | Isulan Rural Health Unit | Poblacion Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29082 | UNOPS | val.binondo@gmail.com | 24.6 |  |  | GxIV-10C | 110023520 | N/A | Perform Preventive maintenance | Perform and introduce Jotform reporting and troubleshooting GX ID sticker installation | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992552839493358735/5992552839493358735_signature_28.png | Valent Binondo | https://www.jotform.com/uploads/macarewecare/222927978372067/5992552839493358735/5992552839493358735_signature_31.png
| 2024-08-12 | 2024-08-12 14:43 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29081 | DOH | sarah_yangzon@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1009 | 2024-08-12 | GXIV | 838406 | Due for Calibration | PM, Calibration | #NAME? | Unit is operational (Module B4 for service required) |  | Xpertcheck kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992544658206538107/5992544658206538107_signature_28.png | Sarah V Yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5992544658206538107/5992544658206538107_signature_31.png
| 2024-08-12 | 2024-08-12 11:35 | Atimonan Rural Health Unit | Leon Guinto Blvd. , Atimonan | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29080 | DOH | amcl2008@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 795 | 2024-07-30 | GXIV 6C D | 838206 | Unit due for Calibration | Performed preventive maintenance Performed calibration Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck B1, B2 and B4 passed B3 service Required | Unit is operational B3 service Required | Used 5 Xpertcheck cartridges | Xpertcheck kit | Dr11761 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992440553617227262/5992440553617227262_signature_28.png | Ma. Rica Christine Noriesta | https://www.jotform.com/uploads/macarewecare/222927978372067/5992440553617227262/5992440553617227262_signature_31.png
| 2024-08-12 | 2024-08-07 09:47 | St. Elizabeth Hospital Inc. | General Santos City, South Cotabato | 12 | South Cotabato | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29079 | COMMERCIAL | covidsupport@sehi.ph | 21.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 8 |  | GxIV-10C | 110003713 | Module 2,3 and 4 request for replacement due to multiple temperature errors | Perform Genexpert machine Preventive maintenance | Perform Genexpert Cleaning/preventive maintenance GX ID Sticker installation Data collection | 1 operational modules (module 1) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992417269496713838/5992417269496713838_signature_28.png | Sehi | https://www.jotform.com/uploads/macarewecare/222927978372067/5992417269496713838/5992417269496713838_signature_31.png
| 2024-08-12 | 2024-08-08 11:19 | Polomolok Rural Health Unit | Poblacion Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29078 | DOH | jmmirasolrmt@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1776 |  | GxIV-6C | 838265 | Module 1 for Service Required Module 3 multiple temperature errors | Perform Genexpert machine Preventive maintenance Gx ID sticker installation | Perform Genexpert Cleaning/preventive maintenance Perform XPC at module 1 for module troubleshooting Checking of errors Problem/errors reporting Data collection | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992417269495988977/5992417269495988977_signature_28.png | JUNE Michele J. MIRASOL | https://www.jotform.com/uploads/macarewecare/222927978372067/5992417269495988977/5992417269495988977_signature_31.png
| 2024-08-12 | 2024-08-12 10:05 | Macare in House | 1005 Aurora Grand,Aurora blvd QC | NCR | NCR 2nd - Quezon City | Macare in House | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 29077 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | CPU board | C111252810 | CRC bios failed | Replaced and installed battery Performed calibration | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992375843617371927/5992375843617371927_signature_28.png | Macare in House | https://www.jotform.com/uploads/macarewecare/222927978372067/5992375843617371927/5992375843617371927_signature_31.png
| 2024-08-12 | 2024-08-12 10:02 | Macare In house | 1005 Aurora Grand,Aurora blvd QC | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 29076 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | CPU board | C121304219 | CRC bios failed | Replaced and installed battery Performed calibration | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992374133611184168/5992374133611184168_signature_28.png | Macare in House | https://www.jotform.com/uploads/macarewecare/222927978372067/5992374133611184168/5992374133611184168_signature_31.png
| 2024-08-12 | 2024-08-12 09:58 | Macare In House | 1005 Aurora Grand,Aurora blvd QC | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 29075 | MACARE-HBDC | florescanoel77@gmail.com | 28 |  |  | CPU | MM021214873 | CRC bios failed | Replaced and installed CPU Performed calibration | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992372303612703147/5992372303612703147_signature_28.png | Macare In house | https://www.jotform.com/uploads/macarewecare/222927978372067/5992372303612703147/5992372303612703147_signature_31.png
| 2024-08-12 | 2024-08-12 09:51 | Macare in House | 1005 Aurora Grand, Aurora blvd QC | NCR | NCR 2nd - Quezon City | In House | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 29074 | MACARE-HBDC | florescanoel77@gmail.com | 26 |  |  | CPU board | SP071703335G | CRC bios failed | Replace and installed battery Performed calibration | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992369983619547168/5992369983619547168_signature_28.png | Macare In House | https://www.jotform.com/uploads/macarewecare/222927978372067/5992369983619547168/5992369983619547168_signature_31.png
| 2024-08-12 | 2024-08-09 13:40 | Kalinga Health Hub Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29073 | CONSORTIUM | shad.famisaran@accenture.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1205 |  | GXIV | 110008415 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992347423611631643/5992347423611631643_signature_28.png | Shad Famisaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5992347423611631643/5992347423611631643_signature_31.png
| 2024-08-12 | 2024-08-09 13:37 | Kalinga Health Hub Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29072 | CONSORTIUM | shad.famisaran@accenture.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1303 |  | GXIV | 110008420 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992347423618348000/5992347423618348000_signature_28.png | Shad Famisaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5992347423618348000/5992347423618348000_signature_31.png
| 2024-08-12 | 2024-08-09 14:49 | Medilinx (MMC) | Amorsolo st., Makati | NCR | NCR 4th - Makati | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | John.Manzano@sodexo.com |  | 29071 | COMMERCIAL | loiszoemontemayor.medilinx@outlook.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 209 ASSAY: Xpert CT/NG, COUNT: 829 ASSAY: Xpert HPV, COUNT: 780 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3108 ASSAY: Others, COUNT: 22 ASSAY: Xpert HCV VL, COUNT: 85 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1890 |  | GX4 (10C) | 110001647 | Due for maintenance B1 disabled due to 2008 and 5007 errror UPS battery unable to backup current supply. | Checked and test ups Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | UPS need to replace B1 disabled | Genexpert module out of warranty |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992344403614726410/5992344403614726410_signature_28.png | Lois Zoe Montemayor | https://www.jotform.com/uploads/macarewecare/222927978372067/5992344403614726410/5992344403614726410_signature_31.png
| 2024-08-12 | 2024-08-09 10:39 | Tropical Disease Foundation | Amorsolo st., Makati | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | silobo@tdf.org.ph |  | 29070 | CONSORTIUM | krcueno@tdf.org.ph | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1952 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 |  | GX 4D | 812560 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance Attached sticker | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992344123614363310/5992344123614363310_signature_28.png | Matthew Lobo | https://www.jotform.com/uploads/macarewecare/222927978372067/5992344123614363310/5992344123614363310_signature_31.png
| 2024-08-12 | 2024-08-06 12:39 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29069 | COMMERCIAL | gabbyg.carag@gmail.com | 24 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Simulation run Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992330613615753390/5992330613615753390_signature_28.png | Ma Gabriela Grace Frances Carag | https://www.jotform.com/uploads/macarewecare/222927978372067/5992330613615753390/5992330613615753390_signature_31.png
| 2024-08-12 | 2024-08-06 12:05 | Cardinal Santos Medical Center | Wilson St. Greenhills, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 29068 | COMMERCIAL | microbiology.csmc@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 145 ASSAY: Xpert HPV, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1119 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992324273617363709/5992324273617363709_signature_28.png | Ellaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5992324273617363709/5992324273617363709_signature_31.png
| 2024-08-12 | 2024-08-02 16:01 | St. Luke's Medical Center BGC | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29067 | COMMERCIAL | ci.bgc@stlukes.com.ph | 20 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 18 ASSAY: Xpert CT/NG, COUNT: 90 ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert HPV, COUNT: 109 ASSAY: Xpert MTB/RIF Ultra, COUNT: 223 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV-D-10C | 120002383 | Printer wasted pad end of service life | Cleaned printer waste pad | Checked printer Cleaned printer waste pad Performed maintenance: Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. Printer is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992323813617146415/5992323813617146415_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5992323813617146415/5992323813617146415_signature_31.png
| 2024-08-12 | 2024-08-02 15:44 | St. Luke's Medical Center BGC | Rizal Dr. BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 29066 | COMMERCIAL | ci.bgc@stlukes.com.ph | 20 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 8 ASSAY: Xpert HCV VL, COUNT: 1 |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992322663614344475/5992322663614344475_signature_28.png | Alvin Pabalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5992322663614344475/5992322663614344475_signature_31.png
| 2024-08-12 | 2024-08-02 11:13 | Pinagbuhatan Health Center | Urbano Velasco Ave., Pinagbuhatan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29065 | MACARE-HBDC | peachmango318@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 152 |  | GX-IV-D-10C | 110024245 | Module B1 has error 4012 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Checked module B1 Generated reports Disabled module B1 | Unit is operational. Module B1/210127592 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992321933618484070/5992321933618484070_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/5992321933618484070/5992321933618484070_signature_31.png
| 2024-08-12 | 2024-08-09 12:50 | Ospital ng Makati | Sampaguita St., Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | omakproperty@yahoo.com |  | 29064 | MACARE-HBDC | osmakpatho_labs@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV - 4D10C | 110023460 | Unit for Calibration | Performed Maintenance | Performed Maintenance Cleaned Filter Cartridge Bay Plunger Rod Generate Reports | Unit is Operational | Work with Rhea Anne Cañelas, Cienne Valerio |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992320323612900798/5992320323612900798_signature_28.png | Riza Bianca Reovoca | https://www.jotform.com/uploads/macarewecare/222927978372067/5992320323612900798/5992320323612900798_signature_31.png
| 2024-08-12 | 2024-08-08 15:54 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29063 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 269 ASSAY: Others, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 | 2024-07-23 | GX-IV-D-10C | 120000613 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  | XpertCheck | DR#11743 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992320323613601929/5992320323613601929_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/5992320323613601929/5992320323613601929_signature_31.png
| 2024-08-12 | 2024-08-08 15:51 | Marikina City Health Office | Shoe Ave. Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29062 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1008 ASSAY: Others, COUNT: 1 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992320323615511250/5992320323615511250_signature_28.png | Estrelita Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/5992320323615511250/5992320323615511250_signature_31.png
| 2024-08-12 | 2024-08-09 10:55 | Labo District Hospital | Purok 1, Talubatib, Labo | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29061 | DOH | esteves110597@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1384 |  | GX4-6C | 838207 | GX for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Attached QR Sticker on GX Generate reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992046622313820730/5992046622313820730_signature_28.png | Janela Joyce Esteves | https://www.jotform.com/uploads/macarewecare/222927978372067/5992046622313820730/5992046622313820730_signature_31.png
| 2024-08-12 | 2024-08-08 13:50 | Ragay Rural Health Unit | Poblacion Ilaod, Ragay | 5 | Camarines Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29060 | PBSP | phylgomez0115@gmail.com | 40 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 702 |  | GX4-6C | 838560 | Gx for Calibration | XpertCheck Result: B1/860776 - newly replaced module B2/757700 - passed B3/861222 - newly replaced module B4/757448 - passed Cleaned cartridge bays, plunger rod, icore Attached QR sticker Generate report | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992046192313820731/5992046192313820731_signature_28.png | PHYL FRANCES C.GOMEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5992046192313820731/5992046192313820731_signature_31.png
| 2024-08-12 | 2024-08-08 10:36 | San Fernando Rural Health Unit | San Fernando | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29059 | DOH | josephus_garfin@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 837 |  | GX4-6C | 838496 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data Attached QR sticker on Gx | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992045502316453540/5992045502316453540_signature_28.png | Josephus garfin | https://www.jotform.com/uploads/macarewecare/222927978372067/5992045502316453540/5992045502316453540_signature_31.png
| 2024-08-12 | 2024-08-07 16:00 | Calabanga Rural Health Unit | Calabanga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29058 | DOH | chanomode@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 748 |  | GX4-6C | 838398 | GX for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate data Attached QR sticker on Gx | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992044002318245652/5992044002318245652_signature_28.png | Jan Christian Orlin | https://www.jotform.com/uploads/macarewecare/222927978372067/5992044002318245652/5992044002318245652_signature_31.png
| 2024-08-12 | 2024-08-07 12:44 | Tigaon Municipal Health Office | Tigaon | 5 | Camarines Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29057 | DOH | maritesjudar611@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 961 |  | GX4-6C | 838558 | GX for Calibration | XpertCheck Result: B1/862620 - passed B2/653396 - newly replaced module B3/862622 - passed B4/695719 - passed Cleaned cartridge bays, plunger rod, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992042312316341939/5992042312316341939_signature_28.png | Marites Judar | https://www.jotform.com/uploads/macarewecare/222927978372067/5992042312316341939/5992042312316341939_signature_31.png
| 2024-08-12 | 2024-08-06 16:38 | Libon Rural Health Unit | Velasco St., Zone 1, Poblacion, Libon, Albay | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29056 | DOH | zh.surge@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1080 |  | GX4-6C | 838397 | GX for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Data Attached QR sticker on GX | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992041232314560715/5992041232314560715_signature_28.png | Marilyn S. Razal | https://www.jotform.com/uploads/macarewecare/222927978372067/5992041232314560715/5992041232314560715_signature_31.png
| 2024-08-12 | 2024-08-06 14:20 | Iriga City Health Office | Iriga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29055 | UNOPS | mongealyzza@yahoo.com.ph | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 179 |  | GX4-6C | 110023511 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992038362313475809/5992038362313475809_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5992038362313475809/5992038362313475809_signature_31.png
| 2024-08-12 | 2024-08-06 14:05 | Iriga City Health Office | Iriga | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29054 | DOH | mongealyzza@yahoo.com.ph | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1887 |  | GX4-6C | 838391 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Attached QR sticker on GX Generate Data | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992036672319364816/5992036672319364816_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5992036672319364816/5992036672319364816_signature_31.png
| 2024-08-12 | 2024-08-09 13:16 | Vinzons Rural Health Unit | Brgy. Dos, Poblacion Dos, Vinzons | 5 | Camarines Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29053 | DOH | jasminlottidp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 996 |  | GX4-6C | 838576 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Archived GX files Attached QR sticker on Gx Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992035242319715919/5992035242319715919_signature_28.png | Jasmin Lotti A. Dela Peña | https://www.jotform.com/uploads/macarewecare/222927978372067/5992035242319715919/5992035242319715919_signature_31.png
| 2024-08-12 | 2024-08-09 16:58 | Bicol Medical Center | Concepcion, Naga | 5 | Camarines Sur | Laboratory | YES | Sticker attachment | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29052 | MACARE-HBDC | bmc.bacteriology@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 32 ASSAY: Xpert HPV, COUNT: 129 ASSAY: Xpert MTB/RIF Ultra, COUNT: 521 ASSAY: Xpert SARS-CoV-2, COUNT: 737 |  | GX4-6C | 110008458 | QR Sticker attachment on GX | QR Sticker attachment on GX Generate data | QR Sticker attachment on GX | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5992034712315745686/5992034712315745686_signature_28.png | Mary Elaine Tolidanes | https://www.jotform.com/uploads/macarewecare/222927978372067/5992034712315745686/5992034712315745686_signature_31.png
| 2024-08-09 | 2024-08-09 18:04 | Southern Philippines Medical Center | JP Laurel Avenue Bajada, Davao City | 11 | Davao del Sur | ICU | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 29051 | COMMERCIAL | airalynecruz@yahoo.com.ph | 25.6 |  |  | MCS+ | 15HO34G | Error at startup, error message display BIOS CRC Fail | Reset Connections of all the boards Check the voltages | Perform Troubleshooting | Still Error BIOS CRC Failed Unit not operational Request for quotation of RAM Battery |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5990082422429516624/5990082422429516624_signature_28.png | Airalyne Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5990082422429516624/5990082422429516624_signature_31.png
| 2024-08-09 | 2024-08-09 16:18 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 29050 | MACARE-HBDC | laboratory@slec.ph | 22 | ASSAY: Xpert CT/NG, COUNT: 5132 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1497 | 2024-08-07 | Infinity 48s | 809643 | Modules A3 and H4 for replacement. | Replaced defective modules | - Performed preventive maintenance cleaning - Replaced modules A3 (702072 > 754001) and H4 (672198 > 210011662) | Unit is operational | With Harold Toco | 900-0491RS CERTIFIED PRE-OWNED GENEXPERT MODULE (2 MODULES) 300-7388 MODIFIED EJECTOR FOR GENEXPERT INFINITY MODULE (2 PCS) 700-4167 ASSY DOOR GENEXPERT MODULE (2 PCS) | DR# 11791 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5990023512738723903/5990023512738723903_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5990023512738723903/5990023512738723903_signature_31.png
| 2024-08-09 | 2024-08-09 15:03 | Chong Hua Hospital | Cebu City | 7 | Cebu | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 29049 | COMMERCIAL | chhc_bloodbank@chonghua.com.ph | 28 |  |  | MCS + | 02F035 | No Display on the monitor. | Observed machine during the weekend. | Replaced cpu board and Reset . | Restart the unit |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989972670228636730/5989972670228636730_signature_28.png | Anna Blanch | https://www.jotform.com/uploads/macarewecare/222927978372067/5989972670228636730/5989972670228636730_signature_31.png
| 2024-08-09 | 2024-08-09 14:19 | CEBU DOCTORS UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29048 | COMMERCIAL | cduhlabresults2@yahoo.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 77 ASSAY: Xpert CT/NG, COUNT: 205 ASSAY: Xpert HBV VL, COUNT: 181 ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1151 ASSAY: Xpert HCV VL, COUNT: 157 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 |  | GX-IV | 110008524 | FOR SCHEDULED MONTHLY PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989967215255136287/5989967215255136287_signature_28.png | Queenie Valencia | https://www.jotform.com/uploads/macarewecare/222927978372067/5989967215255136287/5989967215255136287_signature_31.png
| 2024-08-09 | 2024-08-09 14:37 | Philippine Red Cross Zamboanga | Philippine Red Cross Zamboanga, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION Installation of barcode sticker and convertion Factors | HAEMONETICS | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 29047 | COMMERCIAL | zamboangacity@redcross.org.ph | 21 |  | 2024-08-09 | MCS + | 01M256 | None | Installation of Barcode Sticker and Conversation Factors. | Installation of Barcode Sticker and Conversation Factors. | Unit is available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989954706857446691/5989954706857446691_signature_28.png | Joan A. Luna | https://www.jotform.com/uploads/macarewecare/222927978372067/5989954706857446691/5989954706857446691_signature_31.png
| 2024-08-09 | 2024-08-09 14:10 | Ciudad Medical Zamboanga | Ciudad Medical Zamboanga, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE Barcode sticker Installation | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 29046 | COMMERCIAL | thairaphoebegrey13@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 348 ASSAY: Xpert HCV VL, COUNT: 29 | 2024-08-09 | GX 4 (6c) | 834128 | Module 2 need to be replaced, Monitor unable to process image properly. | Preventive Maintenence and Barcode sticker installation. Replaced VGA cord of monitor. | Preventive Maintenence and Barcode sticker installation, replaced vga cord. | Module 1 available while module 2 needs to be replaced. Monitor problem solved. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989933506856480118/5989933506856480118_signature_28.png | Shereemer A. Cheong-Ynawat | https://www.jotform.com/uploads/macarewecare/222927978372067/5989933506856480118/5989933506856480118_signature_31.png
| 2024-08-09 | 2024-08-09 12:55 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29045 | DOH | burauenlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 460 | 2024-08-09 | GXIV | 838551 | Due for Calibration | #NAME? | #NAME? | Unit is operational (Module B3 for monitoring - multiple temperature error) |  | Xpertcheck kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989890807519125741/5989890807519125741_signature_28.png | maryann | https://www.jotform.com/uploads/macarewecare/222927978372067/5989890807519125741/5989890807519125741_signature_31.png
| 2024-08-09 | 2024-08-08 16:14 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 29044 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1455 ASSAY: Others, COUNT: 95 ASSAY: Xpert HBV VL, COUNT: 1038 ASSAY: Xpert HPV, COUNT: 125 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2323 ASSAY: Xpert SARS-CoV-2, COUNT: 12615 ASSAY: Xpert HCV VL, COUNT: 136 |  | GX16 D | 110005587 | A1,A2,A4,C2,D2 reported already for replacement | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and attached sticker | 11 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989870817192904564/5989870817192904564_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5989870817192904564/5989870817192904564_signature_31.png
| 2024-08-09 | 2024-08-08 15:37 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 29043 | COMMERCIAL | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Xpert HPV, COUNT: 34 ASSAY: Xpert MTB/RIF Ultra, COUNT: 144 ASSAY: Xpert HCV VL, COUNT: 2 |  | GX16 D | 804476 | Due for maintenance A1,A2,B1,B3,C1,C2,C3,C4 encounter loss comm. B4 temp error 4010,4009,4012 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and attached sticker | A3,A4,B2 operational A1,A2,B1,B3,B4,C1,C2,C3,C4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989870407194626796/5989870407194626796_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5989870407194626796/5989870407194626796_signature_31.png
| 2024-08-09 | 2024-08-08 15:06 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 29042 | MACARE-HBDC | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 43 ASSAY: Others, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 21 ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 46 ASSAY: Xpert HCV VL, COUNT: 5 |  | GX4D | 838524 | Due for maintenance B1,B3,B4 reported already for replacement | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and attached sticker | B1,B3 disabled B2,B4 in use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989869697191645788/5989869697191645788_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5989869697191645788/5989869697191645788_signature_31.png
| 2024-08-09 | 2024-08-08 09:18 | Best Diagnostics Corporation | Masikap St,Diliman QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 29041 | COMMERCIAL | bestdiagnosticacctg@gmail.com | 30 | ASSAY: Xpert CT/NG, COUNT: 119 ASSAY: Xpert HBV VL, COUNT: 32 ASSAY: Xpert MTB/RIF Ultra, COUNT: 14 ASSAY: Xpert HCV VL, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GX4 D | 110008529 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989869157191478623/5989869157191478623_signature_28.png | Darwin Jae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5989869157191478623/5989869157191478623_signature_31.png
| 2024-08-09 | 2024-08-07 15:18 | De los Santos Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@dlsmc.ph |  | 29040 | CONSORTIUM | molecularlab@dlsmc.ph | 23 | ASSAY: Xpert HBV VL, COUNT: 12 ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 313 |  | GX4 (10C) | 110020700 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989868737197846223/5989868737197846223_signature_28.png | Trisha Manzo | https://www.jotform.com/uploads/macarewecare/222927978372067/5989868737197846223/5989868737197846223_signature_31.png
| 2024-08-09 | 2024-08-09 10:11 | Maddela District Hospital | Diduyon, Maddela, Quirino | 2 | Quirino | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29039 | DOH | mdhlaboratory00@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 264 |  | GXIV-6C | 838469 | Module B2 error | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989792886029938433/5989792886029938433_signature_28.png | Jane D. Diaron | https://www.jotform.com/uploads/macarewecare/222927978372067/5989792886029938433/5989792886029938433_signature_31.png
| 2024-08-08 | 2024-08-08 14:40 | MANDAUE CITY HEALTH OFFICE | Mandaue City | 7 | Cebu | Laboratory | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29038 | DOH | mendozarey42@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5468 ASSAY: Xpert HCV VL, COUNT: 39 ASSAY: Xpert HIV-1 Viral Load, COUNT: 233 |  | GX-IV | 838510 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989087448815196433/5989087448815196433_signature_28.png | Rey C. Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5989087448815196433/5989087448815196433_signature_31.png
| 2024-08-08 | 2024-08-08 14:36 | MANDAUE CITY HEALTH OFFICE | Mandaue City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29037 | DOH | mendozarey42@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5468 ASSAY: Xpert HCV VL, COUNT: 39 ASSAY: Xpert HIV-1 Viral Load, COUNT: 233 |  | GX-IV | 838508 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989083995443397194/5989083995443397194_signature_28.png | Rey C. Mendoza,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5989083995443397194/5989083995443397194_signature_31.png
| 2024-08-08 | 2024-08-08 11:13 | Metro Antipolo Hospital and Medical Center | Marcas Highway Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 29036 | COMMERCIAL | mahmci_labcmt@yahoo.com | 22 | ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 223 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989071114916306031/5989071114916306031_signature_28.png | Graziela Marie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/5989071114916306031/5989071114916306031_signature_31.png
| 2024-08-08 | 2024-08-08 11:00 | Buhangin District Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 29035 | DOH | buhanginlaboratory@gmail.com | 26.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 888 |  | GX-IV 4D 6C | 838232 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install GX Sticker for reporting | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5989023232029396035/5989023232029396035_signature_28.png | Catherine B. Deocampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5989023232029396035/5989023232029396035_signature_31.png
| 2024-08-08 | 2024-08-08 11:46 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Serology | NO | MAINTENANCE Pullout | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29034 | MACARE-HBDC | cvmcserology@gmail.com | 26 |  |  | GXIV-6C | 803870 | For Pullout | Pullout Genexpert Pullout Laptop Pullout Barcode Scanner Pullout UPS | Performed Performed Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988990171229477676/5988990171229477676_signature_28.png | Dorothy Pajarillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5988990171229477676/5988990171229477676_signature_31.png
| 2024-08-08 | 2024-08-08 11:02 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29033 | MACARE-HBDC | ruellamorena@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 340 ASSAY: Xpert SARS-CoV-2, COUNT: 2695 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 ASSAY: Others, COUNT: 382 |  | GXIV-6C | 110008457 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988955884815594818/5988955884815594818_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5988955884815594818/5988955884815594818_signature_31.png
| 2024-08-08 | 2024-08-08 10:31 | Batasan Social Hygiene Clinic | 1 IBP Road, Batasan Hills, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29032 | MACARE-HBDC | prestozatricia@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 359 | 2024-07-30 | GXIV 10C D | 110020697 | Unit due for calibration | Performed preventive maintenance Performed calibration Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck All modules passed | Unit is operational All modules working |  | XpertCheck Kit | DR 11764 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988946172717309667/5988946172717309667_signature_28.png | rochelle sumilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5988946172717309667/5988946172717309667_signature_31.png
| 2024-08-08 | 2024-08-08 09:24 | Pandre Burgos Rural Health Unit | Padre Burgos Southern Leyte | 8 | Southern Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29031 | UNOPS | keishaelezze14@gmail.com | 26 |  | 2024-08-08 | GXIV 10C | 110023514 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988896117882750343/5988896117882750343_signature_28.png | KEISHA ELEZZE BATALON | https://www.jotform.com/uploads/macarewecare/222927978372067/5988896117882750343/5988896117882750343_signature_31.png
| 2024-08-07 | 2024-08-06 13:59 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Stem Cell | NO | MAINTENANCE | OTHERS | Noel Floresca |  | florescanoel77@gmail.com |  | Airinspace | 29030 | MACARE-HBDC | lcp.mdctl2024@gmail.com | 27 |  |  | Plasmaire | 2006 -T - 188 | Due for maintenance | Cleaned unit Checked and remove dust of filter Checked program parameters Checked cable connection Performed functionality test | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988250623933954951/5988250623933954951_signature_28.png | Jennifer panuelos | https://www.jotform.com/uploads/macarewecare/222927978372067/5988250623933954951/5988250623933954951_signature_31.png
| 2024-08-07 | 2024-08-07 14:47 | Jones Rural Health Unit | Barangay 1, Jones, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | cheryle_rmt@yahoo.com |  | 29029 | DOH | rhujoneslab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1168 |  | GXIV-6C | 838484 | Module B1 and B4 error | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988231329524053691/5988231329524053691_signature_28.png | Cheryle C. Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5988231329524053691/5988231329524053691_signature_31.png
| 2024-08-07 | 2024-08-07 14:17 | Klinika Eastwood | Klinika Eastwood, City Walk, Bagumbayan, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29028 | MACARE-HBDC | klinikaeastwood@gmail.com | 25 |  | 2024-07-30 | GXIV 10C D | 110020698 | Unit due for calibration | Performed preventive maintenance Performed calibration Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck All modules passed | Unit is operational All modules working |  | Xpertcheck kit | DR11762 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988214747228843776/5988214747228843776_signature_28.png | Mary Claire p baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/5988214747228843776/5988214747228843776_signature_31.png
| 2024-08-07 | 2024-08-07 13:26 | Alicia Rural Health Unit | Magsaysay, Alicia, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29027 | DOH | labalicia.rhu@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 |  | GXIV-6C | 838521 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988179433709855726/5988179433709855726_signature_28.png | Catherine Morante | https://www.jotform.com/uploads/macarewecare/222927978372067/5988179433709855726/5988179433709855726_signature_31.png
| 2024-08-07 | 2024-08-07 12:10 | Ramon Rural Health Unit | Bugallon | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29026 | DOH | rhuramonlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 |  | GXIV-6C | 838480 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988134625792285878/5988134625792285878_signature_28.png | Janette Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5988134625792285878/5988134625792285878_signature_31.png
| 2024-08-07 | 2024-08-07 12:11 | Poro Rural Health Unit | Poro, Cebu | 7 | Cebu | TB Dots | NO | MAINTENANCE GeneXpert ID Sticker Installation | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 29025 | DOH | avilacharmaine6@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 621 |  | 6C GXIV-4D | 838279 | None | None | #NAME? | All modules are functional. | No filter, module B2 for observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988133327482364628/5988133327482364628_signature_28.png | Charmaine Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/5988133327482364628/5988133327482364628_signature_31.png
| 2024-08-07 | 2024-08-07 11:10 | Siargao Island Medical Center | Dapa, Siargao Island | 13 | Surigao del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 29024 | DOH | carlynvirtudazo10@gmail.com | 30 |  |  | GX-4 6 Colors | 838444 | None | None | Preventive Maintenance GX ID Sticker Installation | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988111589028965862/5988111589028965862_signature_28.png | Carlyn Virtudazo | https://www.jotform.com/uploads/macarewecare/222927978372067/5988111589028965862/5988111589028965862_signature_31.png
| 2024-08-07 | 2024-08-07 11:27 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 29023 | DOH | claritomirasol@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 460 | 2024-08-07 | GXIV | 838548 | Due for Calibration | #NAME? | #NAME? | Unit is operational |  | Xpertcheck Kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988107173421317619/5988107173421317619_signature_28.png | MIRASOL P. CLARITO | https://www.jotform.com/uploads/macarewecare/222927978372067/5988107173421317619/5988107173421317619_signature_31.png
| 2024-08-07 | 2024-08-07 10:42 | Pontevedra Rural Health Unit | Pontevedra, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29022 | DOH | pontevedra.laboratory@gmai.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 752 |  | GX-4 | 838491 | None | #NAME? | - Perform Xpertcheck in all 4 modules. - Install Genexpert ID Sticker. - Preventive Maintenance | - 4 modules operational. - All 4 modules Passed Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988094286718497241/5988094286718497241_signature_28.png | Eliseo Pabalinas III | https://www.jotform.com/uploads/macarewecare/222927978372067/5988094286718497241/5988094286718497241_signature_31.png
| 2024-08-07 | 2024-08-07 10:02 | Medi Linx laboratory Inc. / St. Elizabeth Hospital Inc. | General Santos City, South Cotabato | 12 | South Cotabato | Micrology | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29021 | COMMERCIAL | richei.hernandez@medilinx.com.ph | 23.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 298 |  | GxIV-6C | 831313 | Module 1, 3 and 4 requesting for module replacement due to multiple errors | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Cleaning/preventive maintenance Gx ID Sticker installation Data collection | 1 operational module (module 2) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988058206014231950/5988058206014231950_signature_28.png | Riche Grace Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5988058206014231950/5988058206014231950_signature_31.png
| 2024-08-07 | 2024-08-06 15:09 | Valenzuela Social Hygiene Clinic | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29020 | COMMERCIAL | markanthony.manaois@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 181 | 2024-07-23 | GXIV 10C | 110020699 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed calibration |  | XpertCheck Kit | 11737 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988015343617604416/5988015343617604416_signature_28.png | Mark Manaois | https://www.jotform.com/uploads/macarewecare/222927978372067/5988015343617604416/5988015343617604416_signature_31.png
| 2024-08-07 | 2024-08-06 15:09 | Valenzuela Social Hygiene Clinic | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 29019 | COMMERCIAL | markanthony.manaois@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 181 | 2024-07-23 | GXIV 10C | 110020699 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed calibration |  | XpertCheck Kit | 11737 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5988012883616478633/5988012883616478633_signature_28.png | Mark Manaois | https://www.jotform.com/uploads/macarewecare/222927978372067/5988012883616478633/5988012883616478633_signature_31.png
| 2024-08-07 | 2024-08-06 15:30 | Rizal Medical Center | Pasig Blvd. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | princeszvillaran13@gmail.com |  | 29018 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 | ASSAY: Others, COUNT: 3 ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 72 ASSAY: Xpert SARS-CoV-2, COUNT: 1446 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-XVI-D-6C | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules A1, A2 B3, B4, C2, C3 and D2 not operational, for replacement. Modules D1 and D3 not operational, for reporting. | Unit due for XpertCheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5987998303611212621/5987998303611212621_signature_28.png | Jhon Patrick Ariola | https://www.jotform.com/uploads/macarewecare/222927978372067/5987998303611212621/5987998303611212621_signature_31.png
| 2024-08-07 | 2024-08-06 16:20 | Rizal Medical Center | Pasig Blvd. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | princeszvillaran13@gmail.com |  | 29017 | COMMERCIAL | rmcbloodbank@gmail.com | 24 |  |  | MCS+ | 15H045G | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Run simulation Verified calibration parameters Functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5987998303611969627/5987998303611969627_signature_28.png | Mae Sarah C. Alojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5987998303611969627/5987998303611969627_signature_31.png
| 2024-08-06 | 2024-08-06 14:16 | Hi-Precision Diagnostics | NS Amoranto St., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29016 | CONSORTIUM | jose@hi-precision.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 133 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6570 |  | GXIV 6C D | 835601 | Frequent error in Module 4 | Performed Preventive maintenance Checked what type of erros is occurring Archived results to be checked by technical Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Checked system log Cheked details of error Mostly 5007 Archived results from April to August | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5987501125668877256/5987501125668877256_signature_28.png | Joselito Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5987501125668877256/5987501125668877256_signature_31.png
| 2024-08-06 | 2024-08-06 16:01 | TAYASAN RURAL HEALTH UNIT | TAYASAN, NEGROS ORIENTAL | 7 | Negros Oriental | LABORATORY | NO | MAINTENANCE GX ID Sticker Installation | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29015 | UNOPS | porshcabz151@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 47 |  | GX-IV | 110023502 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5987412813595861369/5987412813595861369_signature_28.png | Portia Cabanting | https://www.jotform.com/uploads/macarewecare/222927978372067/5987412813595861369/5987412813595861369_signature_31.png
| 2024-08-06 | 2024-08-06 15:47 | Malapatan Rural Health Unit | Malapatan, Sarangani Province | 12 | Sarangani | Tb-Lab | NO | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29014 | DOH | kamilahnichol@yahoo.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 818 |  | GxIV-6C | 838419 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert Cleaning/preventive maintenance Gx ID sticker installation Checking of errors/problem Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5987403618956645347/5987403618956645347_signature_28.png | KAMILAH NICHOL CHU | https://www.jotform.com/uploads/macarewecare/222927978372067/5987403618956645347/5987403618956645347_signature_31.png
| 2024-08-06 | 2024-08-06 13:17 | Sarangani Health Care Facility | Alabel, Sarangani Province | 12 | Sarangani | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29013 | DOH | ronayadao@yahoo.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 |  | GxIV-6C | 838542 | H/W failed Module 2 multiple temperature errors Module 3 temperature errors | Perform Genexpert machine Preventive maintenance | Perform Genexpert Cleaning/ preventive maintenance Perform XPC at module 3 for module temperature Gx ID sticker installation Problem/errors reporting Data collection | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5987395828955671696/5987395828955671696_signature_28.png | Ronora Yadao | https://www.jotform.com/uploads/macarewecare/222927978372067/5987395828955671696/5987395828955671696_signature_31.png
| 2024-08-06 | 2024-08-06 15:11 | Solano Rural Health Unit | Roxas, Solano, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29012 | DOH | rhusolanolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 682 |  | GXIV-6C | 838201 | CPU Sometimes dont power On | Troubleshoot CPU | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | For Observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5987377064514762368/5987377064514762368_signature_28.png | May Vanessa A. Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5987377064514762368/5987377064514762368_signature_31.png
| 2024-08-06 | 2024-08-06 13:26 | Nueva Vizcaya Provincial Hospital | Almaguer, Bambang, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 29011 | DOH | dumlaojoy4@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1103 |  | GXIV-6C | 838272 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Attached Sticker | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5987318444029417209/5987318444029417209_signature_28.png | Eloisa De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5987318444029417209/5987318444029417209_signature_31.png
| 2024-08-06 | 2024-08-06 10:30 | Project 7 Social Hyguene Clinic | Veterans Village, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 29010 | MACARE-HBDC | michelenaval@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 276 | 2024-07-30 | GX IV 10C-D | 120000577 | Unit due for XpertCheck | Performed preventive maintenance Performed Calibration | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck All modules passed. | Unit is operational All modules working |  | XpertCheck Kit | DR 11763 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5987209179719956677/5987209179719956677_signature_28.png | Michele Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5987209179719956677/5987209179719956677_signature_31.png
| 2024-08-05 | 2024-07-31 10:34 | Lake Sebu Rural Health Unit | Purok Rosas, Lake Sebu, Sou | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29009 | DOH | jennypingkian23@yahoo.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 531 |  | GxIV - 6C | 838569 | N/A | Genexpert machine Preventive maintenance | Performed: Genexpert Cleaning/ preventive maintenance Gx ID sticker installation Checking of errors Data collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986737216861271910/5986737216861271910_signature_28.png | Jenny A. Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/5986737216861271910/5986737216861271910_signature_31.png
| 2024-08-05 | 2024-08-01 11:06 | Tampkan Rural Health Unit | Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 29008 | DOH | jmedtech_lab@yahoo.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 733 |  | GxIV-6C | 838544 | Module 4 for Service Required (H/W failed) Module 2 multiple temperature error | Genexpert Preventive maintenance | Performed: Genexpert Cleaning and Preventive maintenance Checking of errors Errors/problem reporting Gx ID sticker installation Data collection | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986736766868334878/5986736766868334878_signature_28.png | Jovi G. Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/5986736766868334878/5986736766868334878_signature_31.png
| 2024-08-05 | 2024-08-05 10:52 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Oriental | Laboratory | NO | MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 29007 | DOH | cauayanmholab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1113 |  | GX-4 | 838477 | None | #NAME? | #NAME? | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986485334103364719/5986485334103364719_signature_28.png | Doriza Joy Tizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5986485334103364719/5986485334103364719_signature_31.png
| 2024-08-05 | 2024-08-01 10:35 | WVMC rHIVda Confirmatory Laboratory | Iloilo City | 6 | Iloilo | Laboratory | YES | INSTALLATION Machine Transfer | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 29006 | PBSP | blacklion012886@yahoo.com | 24 |  | 2024-08-01 | GX4 | 706910 | GeneXpert Machine Transfer. | Install GeneXpert Machine System. | From molecular laboratory to RHIVda room GeneXpert Machine System Installation. | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986415834329971825/5986415834329971825_signature_28.png | Julie Garanganao | https://www.jotform.com/uploads/macarewecare/222927978372067/5986415834329971825/5986415834329971825_signature_31.png
| 2024-08-05 | 2024-08-05 12:24 | BINDOY DISTRICT HOSPITAL | BINDOY, NEGROS ORIENTAL | 7 | Negros Oriental | LABORATORY | NO | MAINTENANCE GX ID Sticker Installation | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 29005 | DOH | bindoydistricthosp.laboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 702 |  | GX-IV | 838475 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986414323397969484/5986414323397969484_signature_28.png | Lemuel Quirit, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5986414323397969484/5986414323397969484_signature_31.png
| 2024-08-05 | 2024-08-05 10:25 | Pilar Rural Health Unit | Calongay, Pilar, Sorsogon | 5 | Sorsogon | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29004 | DOH | manalangmark88@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 |  | GX4-6C | 838266 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore, fan Provide QR sticker for easy reporting Generate reports | Performed Preventive Maintenance | Unit is Operational | Pm performed last 07/25/24 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986341974715705199/5986341974715705199_signature_28.png | Mark Ramir C. Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5986341974715705199/5986341974715705199_signature_31.png
| 2024-08-05 | 2024-08-02 12:43 | Pili Rural Health Unit | Poblacion, San Isidro, Pili, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29003 | DOH | lfsanao724@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1127 |  | GX4-6C | 838408 | Gx for Calibration | XpertCheck Result: B1/688686 - Passed B2/220005368 - Passed B3/210029760 - Passed B4/662296 - Passed Cleaned cartridge bays, plunger rod, icore, fan filter Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986338824717556065/5986338824717556065_signature_28.png | Liezl F. Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/5986338824717556065/5986338824717556065_signature_31.png
| 2024-08-05 | 2024-08-01 15:06 | San Jose Municipal Health Office | Soledad, San Jose, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29002 | DOH | jhuneric2008@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2245 |  | GX4-6C | 838389 | Gx for Calibration | XpertCheck Result: B1/676102 - newly replaced module B2/607968 - service required B3/739734 - newly replaced module B4/718686 - service required Cleaned cartridge bays, plunger rod, icore Placed QR sticker on Gx for reporting Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational | B2/607968, B4/718686 service required; requesting for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986338534719575578/5986338534719575578_signature_28.png | ROSMAR V FORMALEJO | https://www.jotform.com/uploads/macarewecare/222927978372067/5986338534719575578/5986338534719575578_signature_31.png
| 2024-08-05 | 2024-07-30 14:35 | Virac Municipal Health Office | San Isidro Village, Virac, Catanduanes | 5 | Catanduanes | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 29001 | DOH | limarmariano50@yahoo.com | 28 |  |  | GX4-6C | 838563 | CPU frequent forced restart | Replaced CPU(DELL Optiplex XE2) Cleaned cartridge bays, plunger rod, icore Generate reports | Replaced CPU Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986337954714100279/5986337954714100279_signature_28.png | Ramil Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/5986337954714100279/5986337954714100279_signature_31.png
| 2024-08-05 | 2024-08-05 05:29 | Ramon Magsaysay Rural Health unit | Poblacion Ramon Magsaysay, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE ["Installation of Barcode sticker for reporting ","Installation of Barcode sticker for reporting "] | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 29000 | DOH | ramonmagsaysay@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 532 | 2024-08-01 | GX 4(6C) | 838443 | None | Preventive Maintenence and Installation of barcode sticker for reporting | Preventive Maintenence and Installation of barcode sticker for reporting | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5986173768217066495/5986173768217066495_signature_28.png | Gina Borinaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5986173768217066495/5986173768217066495_signature_31.png
| 2024-08-02 | 2024-08-02 12:25 | Dasmariñas City Health Office II | DBB-E, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28996 | DOH | hazelmilan0777@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2450 | 2024-07-29 | GXIV-D-6C | 838535 | Unit is due for Maintenance. No Backup Power on UPS. | Preventive Maintenance. Troubleshooting. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Intall CPU and Pull-out Backup. Generate Reports. Checked UPS Voltage. Performed Functionality Test. | All Modules are Operational. CPU is Functioning. UPS has Defective Battery. (Low Backup Voltage) | UPS: AT1000190123ST0101 | GX Desktop | DR 11753 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200811441640/5983960200811441640_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200811441640/5983960200811441640_signature_31.png
| 2024-08-02 | 2024-07-31 14:21 | Aaian Hospital and Medical Center | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | agimang@asianhospital.com |  | 28999 | MACARE-HBDC | angelgimang17@gmail.com | 25 |  |  | MCS+ | 01F116 | Unit is due for Maintenance. | Preventive Maintenance. | Performed Simulation Test. Checked Pressure and Optic Values. Checked Pump Rubber. Checked A/D Values. | Values are Good. Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200813069182/5983960200813069182_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200813069182/5983960200813069182_signature_31.png
| 2024-08-02 | 2024-07-31 14:16 | Asian Hospital and Medical Center | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | agimang@asianhospital.com |  | 28997 | MACARE-HBDC | angelgimang17@gmail.com | 25 |  |  | MCS+ | 10B129 | Unit is due for Maintenance. | Preventive Maintenance. | Performed Simulation Test. Checked Pressure and Optic Values. Checked Pump Rubber. Checked A/D Values. | Unit is Operational. Values are Good. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200813393790/5983960200813393790_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200813393790/5983960200813393790_signature_31.png
| 2024-08-02 | 2024-07-30 14:17 | Dr. Jose Fabella Memorial Hospital | Tayuman St, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | trishpamittan@gmail.com |  | 28995 | MACARE-HBDC | fabella.mollabgene@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 118 ASSAY: Xpert MTB/RIF Ultra, COUNT: 31 ASSAY: Xpert SARS-CoV-2, COUNT: 2094 ASSAY: Others, COUNT: 265 |  | GXIV-D-10C | 110015482 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | Disabled Modules: None |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200816505957/5983960200816505957_signature_28.png | Tricia Wella M. Pamittan | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200816505957/5983960200816505957_signature_31.png
| 2024-08-02 | 2024-07-30 14:23 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | trishpamittan@gmail.com |  | 28998 | MACARE-HBDC | fabella.mollabgene@gmail.com | 21 | ASSAY: Others, COUNT: 20 ASSAY: Xpert SARS-CoV-2, COUNT: 502 |  | GXIV-L-6C | 834953 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Generate Reports. | Disabled Modules: A1/ 680659 & A3/ 789441 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200818397670/5983960200818397670_signature_28.png | Tricia Wella M. Pamittan | https://www.jotform.com/uploads/macarewecare/222927978372067/5983960200818397670/5983960200818397670_signature_31.png
| 2024-08-02 | 2024-08-02 13:45 | MLLI-Riverside Medical Center | Bs Aquino drive, Bacolod | 6 | Negros Occidental | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28994 | COMMERCIAL | rjmanito@gmail.com | 24 |  |  | Gx 4 | 110005268 | Module B1 and B4 | Install ID sticker PM Clean | Install ID sticker PM Clean | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983870487921508787/5983870487921508787_signature_28.png | Rumary Joan Manito | https://www.jotform.com/uploads/macarewecare/222927978372067/5983870487921508787/5983870487921508787_signature_31.png
| 2024-08-02 | 2024-08-02 13:19 | Calamba Rural Health Unit | Calamba Misamis Occidental | 10 | Misamis Occidental | Laboratory | YES | INSTALLATION ["Perfomed end user training on GeneXpert operation, cleaning, and basic troubleshooting.","Perfomed end user training on GeneXpert operation, cleaning, and basic troubleshooting."] | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28993 | UNOPS | palubonibn@gmail.com | 27 |  | 2024-06-20 | GX 4(10x) | 110023518 | None | Installation of Genexpert Machine. Perfomed end user training on GeneXpert operation, cleaning, and basic troubleshooting | Installation of Genexpert Machine. Perfomed end user training on GeneXpert operation, cleaning, and basic troubleshooting. | Genexpert Machine Ready to use | None | GX1V-4-D-10C GX-DESKTOP GX-MONITOR GX-BSCAN FC-IT-UPS ACCUIK FC-IT-PRINTER XPRSARS-COV2-10 | 11646 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983867462511135896/5983867462511135896_signature_28.png | Mary-ibn E. Palubon | https://www.jotform.com/uploads/macarewecare/222927978372067/5983867462511135896/5983867462511135896_signature_31.png
| 2024-08-02 | 2024-08-02 12:14 | Panabo City Health Office | Poblacion Panabo | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28992 | DOH | nadineeust67@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2876 |  | GX-IV 6C 4D | 838295 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983840764166818630/5983840764166818630_signature_28.png | Sitti hannah deloso | https://www.jotform.com/uploads/macarewecare/222927978372067/5983840764166818630/5983840764166818630_signature_31.png
| 2024-08-02 | 2024-08-02 11:21 | Nasipit Rural Health Unit | Brgy. Talisay, Nasipit | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28991 | DOH | peterdolotina@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1316 |  | GX-4 6 Colors | 837611 | None | None | Preventive Maintenance and GeneXpert ID Sticker Installation. | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983798689224886049/5983798689224886049_signature_28.png | Peter Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/5983798689224886049/5983798689224886049_signature_31.png
| 2024-08-02 | 2024-08-02 11:16 | The Medical City Clark | 100 Gatwick Gateway, Zone, Clark Freeport, Mabalacat, 2023 Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | tmcclarkbiomed@themedicalcity.com |  | 28990 | COMMERCIAL | tmci_laboratory@themedicalcity.com | 24 | ASSAY: Xpert CT/NG, COUNT: 11 ASSAY: Xpert HBV VL, COUNT: 49 ASSAY: Others, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 179 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GXIV-D | 839235 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit us operational | Machine due for calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983782215416342781/5983782215416342781_signature_28.png | Jessa Rae Remorta | https://www.jotform.com/uploads/macarewecare/222927978372067/5983782215416342781/5983782215416342781_signature_31.png
| 2024-08-02 | 2024-08-02 08:21 | Corazon Locsin Medical Memorial Regional Hospital | Bacolod City | 6 | Negros Occidental | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28989 | COMMERCIAL | alainjuayang@yahoo.com | 28 |  |  | Gx 4 | 110020743 | None | ID Sticker Installation PM Clean fan filter | ID Sticker Installation PM Clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983706671151137380/5983706671151137380_signature_28.png | Alain Juayang | https://www.jotform.com/uploads/macarewecare/222927978372067/5983706671151137380/5983706671151137380_signature_31.png
| 2024-08-01 | 2024-08-01 11:47 | Davao Reproductive Health and Wellness Center | Jacinto Street, Davao | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28988 | MACARE-HBDC | rhwclab@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 254 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1219 ASSAY: Xpert HBV VL, COUNT: 8 |  | GX-IV 6C 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install GX Stickers | Perform Preventive Maintenance | Three Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983266066215851143/5983266066215851143_signature_28.png | Cherry Ann Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5983266066215851143/5983266066215851143_signature_31.png
| 2024-08-01 | 2024-08-01 11:50 | Davao Reproductive Health and Wellness Center | Jacinto Street Davao | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28987 | DOH | rhwclab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 354 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1219 ASSAY: Xpert HBV VL, COUNT: 8 |  | GX-IV 6C 4D | 838262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install GX Sticker | Perform Preventive Maintenance | 2 Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983265606219853930/5983265606219853930_signature_28.png | Cherry Ann Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5983265606219853930/5983265606219853930_signature_31.png
| 2024-08-01 | 2024-07-31 15:07 | The Medical City | Ortigas Avenue, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28986 | COMMERCIAL | moleculardx@themedicalcity.com | 20 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 39 ASSAY: Xpert CT/NG, COUNT: 69 ASSAY: Xpert HBV VL, COUNT: 583 ASSAY: Xpert HPV, COUNT: 460 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5300 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 74 ASSAY: Xpert HIV-1 Viral Load, COUNT: 653 ASSAY: Others, COUNT: 931 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports Archive test | Unit is operational. Module B4/210082262 operational but under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983066103612279601/5983066103612279601_signature_28.png | ashleyespiel@gmail.com | https://www.jotform.com/uploads/macarewecare/222927978372067/5983066103612279601/5983066103612279601_signature_31.png
| 2024-08-01 | 2024-08-01 14:55 | DOST - Philippine Institute of Volcanology and Seismology | C.P. Garcia Avenue, Quezon City | NCR | NCR 2nd - Quezon City | DOST PhilVOLCS | YES | MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com | jeannemyrtia@gmail.com |  | 28985 | COMMERCIAL | valerie.olfindo@phivolcs.dost.gov.ph | 25 |  |  | Stemi 508 | 3942010146 | There is a dark spot in the center of the screen | Cleaned camera adapter | Performed Preventive Maintenance Cleaned Eyepieces Checked moving parts Checked lights Checked camera Performed functionality tests | Unit is operational. Black spot in screen was cleaned and is now clear. Eyepieces have some mold on the outer parts. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983062071122347612/5983062071122347612_signature_28.png | Valerie Shayne Olfindo | https://www.jotform.com/uploads/macarewecare/222927978372067/5983062071122347612/5983062071122347612_signature_31.png
| 2024-08-01 | 2024-08-01 15:13 | Tamparan District Hospital | Tamparan, Lanao del Sur | BARMM | Lanao del Sur | TB Dots | NO | INSTALLATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28984 | DOH | pmamowalas@yahoo.com | 25 |  | 2024-06-20 | 10C GXIV-4D | 110023525 | None | None | #NAME? | Machine is functional. |  |  | DR 11655 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983059171935361913/5983059171935361913_signature_28.png | Pangilamun Mamowalas, Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5983059171935361913/5983059171935361913_signature_31.png
| 2024-08-01 | 2024-08-01 14:11 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | YES | INSTALLATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 28983 | MACARE-HBDC | laboratory@slec.ph | 26 |  |  | GX XVI 10-Color | 110022374 | Unit for installation | Installed unit | Performed unit installation Performed functionality test Generated reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983057999129264818/5983057999129264818_signature_28.png | Khariz Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5983057999129264818/5983057999129264818_signature_31.png
| 2024-08-01 | 2024-08-01 14:29 | Teresita L Jalandoni Provincial Hospital | Silay, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28982 | DOH | mrvillan.pbsp@gmail.com | 29 |  |  | Gx 4 | 838464 | None | ID sticker Installation Pm Installation XDR assay | ID sticker Installation Pm Installation XDR assay | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983033607712080684/5983033607712080684_signature_28.png | Ma. Arna R. Villan | https://www.jotform.com/uploads/macarewecare/222927978372067/5983033607712080684/5983033607712080684_signature_31.png
| 2024-08-01 | 2024-08-01 13:43 | Bacolod CHO | Bacolod City | 6 | Negros Occidental | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28981 | DOH | vetpahilanga@gmail.com | 29 |  |  | Gx 4 | 837612 | Error in module 3 and 4 | Clean Fan Filter PM ID sticker installation | Clean Fan Filter PM ID sticker installation | Defective module 3 and 4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5983029537718878894/5983029537718878894_signature_28.png | Ruby Eyvette Asendente | https://www.jotform.com/uploads/macarewecare/222927978372067/5983029537718878894/5983029537718878894_signature_31.png
| 2024-08-01 | 2024-08-01 13:12 | Baybay City Health Office | Baybay Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28980 | DOH | sajecopooh20@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1007 | 2024-08-01 | GXIV | 838555 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982986244536824594/5982986244536824594_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5982986244536824594/5982986244536824594_signature_31.png
| 2024-08-01 | 2024-08-01 13:03 | Magsaysay Rural Health Unit | Kakilgan St. Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | INSTALLATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28979 | DOH | jkakilgan12@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 389 |  | GX-4 6 colors | 838557 | None | None | GX ID Sticker Installation and Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982981908926666155/5982981908926666155_signature_28.png | Jesusa N. Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/5982981908926666155/5982981908926666155_signature_31.png
| 2024-08-01 | 2024-08-01 12:30 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Mabini, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28978 | UNOPS | bryanmichaelbalunes@gmail.com | 25 |  |  | GXIV-10C | 110023508 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982960770906111358/5982960770906111358_signature_28.png | Bryan Michael Balunes | https://www.jotform.com/uploads/macarewecare/222927978372067/5982960770906111358/5982960770906111358_signature_31.png
| 2024-08-01 | 2024-08-01 11:29 | Talavera General Hospital | Maestrang Kikay, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28977 | MACARE-HBDC | pjgtehlab@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 89 ASSAY: Xpert MTB/RIF Ultra, COUNT: 704 ASSAY: Others, COUNT: 34 ASSAY: Xpert HCV VL, COUNT: 9 ASSAY: Xpert HIV-1 Viral Load, COUNT: 125 |  | GXIV-10C | 110013773 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982925230907035108/5982925230907035108_signature_28.png | John Carlo Cuanan | https://www.jotform.com/uploads/macarewecare/222927978372067/5982925230907035108/5982925230907035108_signature_31.png
| 2024-08-01 | 2024-07-31 15:54 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28974 | MACARE-HBDC | laboratory@loveyourself.com.ph | 27 |  |  | GXIV | 807112 | Unit for maintenance | NA | #NAME? | Unit is operational Module C4 for follow up/reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982894933611661265/5982894933611661265_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/5982894933611661265/5982894933611661265_signature_31.png
| 2024-08-01 | 2024-07-31 16:00 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28976 | MACARE-HBDC | laboratory@loveyourself.com.ph | 27 |  |  | GXIV | 807956 | Unit for maintenance | NA | #NAME? | Unit is operational Module A1 for follow up/reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982894933615413090/5982894933615413090_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/5982894933615413090/5982894933615413090_signature_31.png
| 2024-08-01 | 2024-07-31 15:53 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28975 | PBSP | laboratory@loveyourself.com.ph | 27 |  |  | GXIV | 808649 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982894933615537465/5982894933615537465_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/5982894933615537465/5982894933615537465_signature_31.png
| 2024-08-01 | 2024-07-31 15:47 | To Love Yourself | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28973 | DOH | laboratory@loveyourself.com.ph | 27 |  |  | GXIV | 881895 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982894933618147233/5982894933618147233_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/5982894933618147233/5982894933618147233_signature_31.png
| 2024-08-01 | 2024-07-31 15:45 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 28972 | MACARE-HBDC | nchmbdteam@gmail.com | 26 |  |  | MCS + | 17H632SPG | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982715932147074393/5982715932147074393_signature_28.png | Mary Jane Amores | https://www.jotform.com/uploads/macarewecare/222927978372067/5982715932147074393/5982715932147074393_signature_31.png
| 2024-08-01 | 2024-07-31 12:11 | National Childrens Hospital | E Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 28971 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 43 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2618 ASSAY: Xpert SARS-CoV-2, COUNT: 2368 ASSAY: Xpert SARS-CoV-2, COUNT: 591 |  | GX16 D | 110014171 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982715042145127264/5982715042145127264_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5982715042145127264/5982715042145127264_signature_31.png
| 2024-08-01 | 2024-07-26 15:23 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28970 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 02D166 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982714462142800508/5982714462142800508_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/5982714462142800508/5982714462142800508_signature_31.png
| 2024-08-01 | 2024-07-26 14:36 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank bank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28969 | MACARE-HBDC | lcpbloodbank@gmail.com | 26 |  |  | MCS + | 08H188 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982714092142762553/5982714092142762553_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/5982714092142762553/5982714092142762553_signature_31.png
| 2024-07-31 | 2024-07-31 21:58 | PHO MBL Palawan UNOPS | PGP Compound, South National Hi-way, Brgy. Irawan, Puerto Princesa | 4B | Palawan | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28968 | UNOPS | phomblpalawan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: ASSAY: Xpert SARS-CoV-2, COUNT: | 2024-06-20 | GX IV 10C-D | 110023509 | Unit due for installation | Installed GX Generated report | Unpacked and setup GX and accessories Setup User admin Setup system configuration Installed assays Installed printer driver | unit is operational All modules working |  | GXIV 10C Desktop Package | DR 11636 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982449721017250478/5982449721017250478_signature_28.png | Cyrus Cecilio A. Caabay | https://www.jotform.com/uploads/macarewecare/222927978372067/5982449721017250478/5982449721017250478_signature_31.png
| 2024-07-31 | 2024-07-30 11:17 | Philippine Red Cross - Davao City Chapter | Roxas Boulevard, Davao | 11 | Davao del Sur | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28967 | COMMERCIAL | davaocitybloodservices@gmail.com | 25.8 |  |  | MCS+ | 02J021 | For Preventive Maintenance | Clean Surface Area, Check the parameters of SPM, DPM and Line Sensor Run Cycle Diagnostic Test and Perform Cycle Test | Perform Preventive Maintenance | Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982185814428872790/5982185814428872790_signature_28.png | Joy Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5982185814428872790/5982185814428872790_signature_31.png
| 2024-07-31 | 2024-07-31 13:07 | Mini Forest Health Center ( District C ) | Quezon Boulevard, Davao | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28966 | DOH | miniforestlaboratory@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 423 |  | GX-IV 4D 6C | 838507 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982180434429341644/5982180434429341644_signature_28.png | Joy L. Nio | https://www.jotform.com/uploads/macarewecare/222927978372067/5982180434429341644/5982180434429341644_signature_31.png
| 2024-07-31 | 2024-07-30 14:40 | Rosario Super Health Center | Dr. Sixto Ave., Rosario, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28964 | MACARE-HBDC | eden.cawayan@yahoo.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 282 |  | GX-IV-D-10C | 110024246 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982170224443461518/5982170224443461518_signature_28.png | Eden R. Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5982170224443461518/5982170224443461518_signature_31.png
| 2024-07-31 | 2024-07-30 11:31 | Santolan Super Health Center | Evangelista St., Santolan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28965 | MACARE-HBDC | samandlet17@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 360 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982170234441205193/5982170234441205193_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5982170234441205193/5982170234441205193_signature_31.png
| 2024-07-31 | 2024-07-31 13:35 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 28963 | DOH | lornsabaigar@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 792 | 2024-07-31 | GXIV | 838409 | Due for Calibration | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982168349965266541/5982168349965266541_signature_28.png | lorna c.abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/5982168349965266541/5982168349965266541_signature_31.png
| 2024-07-31 | 2024-07-31 12:16 | Tomas Cabili Health Center | Brgy. Tominobo, Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | ["GeneXpert ID Sticker Installation","GeneXpert ID Sticker Installation"] | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28962 | DOH | mariacielovenezuela@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1196 |  | 6C GXIV-4D | 838294 | Module B1 temperature error. | None | -GeneXpert ID Sticker Installation -Preventive Maintenance -Run XPC in module B1 -Upgrade XpertCheck software version to 1.5b | Module B1 failed, for replacement. All other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982100648993639792/5982100648993639792_signature_28.png | Maria Cielo Venezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5982100648993639792/5982100648993639792_signature_31.png
| 2024-07-31 | 2024-07-31 11:56 | Zamboanga del sur medical center | Zamboanga del sur medical center, Pagadian City | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Installation of barcode sticker for reporting | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 28961 | DOH | limboemillie@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1678 ASSAY: Xpert HIV-1 Viral Load, COUNT: 127 | 2024-07-31 | GX 4 | 838204 | Module 2 and 4 temperature errors | Preventive Maintenence and barcode sticker installation | Preventive Maintenence and barcode sticker installation | Module 1 and 3 are running and available while module 2 and 4 are pending for replacement | Module 2 and 4 are for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982076847314733876/5982076847314733876_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/5982076847314733876/5982076847314733876_signature_31.png
| 2024-07-31 | 2024-07-31 10:18 | Hinaplanon Health Center | Brgy. Hinaplanon, Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | ["GeneXpert ID Sticker Installation","GeneXpert ID Sticker Installation"] | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28960 | DOH | villacortaamy50@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1080 |  | 6C GXIV-4D | 838227 | Module B1 H/W failed. | None | -Preventive Maintenance -GeneXpert ID Sticker Installation -Run XPC in module B1 -Upgrade XpertCheck software to version 1.5b | Module B1 defective, for replacement. All other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5982022823671261956/5982022823671261956_signature_28.png | Amy S. Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/5982022823671261956/5982022823671261956_signature_31.png
| 2024-07-31 | 2024-07-30 10:48 | Norala District Hospital | Poblacion Norala, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28959 | DOH | ndhlab2023@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1078 |  | GxIV-6C | 838543 | N/A | Perform Genexpert machine Preventive maintenance | Genexpert machine Cleaning/preventive maintenance Checking of errors Macare sticker installation Data Collection | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5981989157142407300/5981989157142407300_signature_28.png | Geraldine Magan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5981989157142407300/5981989157142407300_signature_31.png
| 2024-07-31 | 2024-07-31 08:54 | Kauswagan Rural Health Unit | Kauswagan, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | MAINTENANCE GeneXpert ID Sticker Installation | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28958 | DOH | dextertinoy@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 298 |  | 6C GXIV-4D | 838402 | None | None | -Preventive Maintenance -GeneXpert ID Sticker Installation -Upgrade XpertCheck software version to 1.5b | Waiting for module B3 allocation. All other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5981967210037829650/5981967210037829650_signature_28.png | Dexter Tinoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5981967210037829650/5981967210037829650_signature_31.png
| 2024-07-30 | 2024-07-30 15:23 | Puerto Princesa City Health Office - UNOPS | Valencia St., Puerto Princesa, Palawan | 4B | Palawan | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28957 | UNOPS | jocal0324@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: | 2024-06-20 | GXIV 6C D | 110023510 | Unit due for Installation | Installed GX Generated report | Unpacked and setup GX and accessories Setup User admin Setup system configuration Installed assay Installed printer driver | Unit is operational All modules working |  | GXIV Desktop Package | DR 11637 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5981617529734090147/5981617529734090147_signature_28.png | Jocelyn L. Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5981617529734090147/5981617529734090147_signature_31.png
| 2024-07-30 | 2024-07-30 13:41 | St Lukes Medical Center | E Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiologu | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 28956 | COMMERCIAL | megapo@stlukes.com.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 273 ASSAY: Xpert HPV, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3192 |  | GX4 | 839234 | Due for maintenance B3 and B4 disabled failed xpertcheck B1 broken uniframe hinge | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B1 and B2 ready to use B3 and B4 disabled due to failed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5981304601428126583/5981304601428126583_signature_28.png | Cherry Ann M. Pastrana | https://www.jotform.com/uploads/macarewecare/222927978372067/5981304601428126583/5981304601428126583_signature_31.png
| 2024-07-30 | 2024-07-30 14:17 | St Lukes Medical Center | E Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 28955 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 29 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 60 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert HPV, COUNT: 267 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX4 L | 807155 | Due for maintenance C4 disabled due to temp error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | C1,C2,C3 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5981303601427092288/5981303601427092288_signature_28.png | Jeinz Quinnie Alcsla | https://www.jotform.com/uploads/macarewecare/222927978372067/5981303601427092288/5981303601427092288_signature_31.png
| 2024-07-30 | 2024-07-30 14:04 | Mabuhay Rural Health unit | Poblacion mabuhay, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION Basic troubleshooting and Maintenence training for end users | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28954 | UNOPS | moffetgayvirtudazo@gmail.com | 23 |  | 2024-06-20 | GX 4 (10C) | 110023517 | None | Installed Genexpert Machine 10c. Perfomed end user training on GeneXpert operation, cleaning, and basic troubleshooting. ty. | Installed Genexpert Machine 10c. Perfomed end user training on GeneXpert operation, cleaning, and basic troubleshooting. ty. | All modules are available and ready to use | None | GX-D-10C GX-DESKTOP GX-MONITOR GXBCSCAN FC-IT-UPSACCU1K FC-IT-PRINTER XPRSCARS-COV2-10 | 11645 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5981291168007268349/5981291168007268349_signature_28.png | Moffat Gay C. Virtudazo | https://www.jotform.com/uploads/macarewecare/222927978372067/5981291168007268349/5981291168007268349_signature_31.png
| 2024-07-30 | 2024-07-30 13:25 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28953 | DOH | ammegzedd@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1196 | 2024-07-30 | GXIV | 838421 | Unit for Calibration | Preventive Maintenance Calibration | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5981278024419274922/5981278024419274922_signature_28.png | gemma malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5981278024419274922/5981278024419274922_signature_31.png
| 2024-07-30 | 2024-07-23 14:10 | Graman Medical And Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28952 | COMMERCIAL | gramanmolecularlab@gmail.com | 23 |  |  | GXIV 10C | 110008530 | Unit for troubleshooting | Generated reports | - scanned system log for hardware errors (none) - scanned test reports errors (error 5007 found) - generated reports - generated archived tests (2 months worth) - checked cartridge bay, plunger rod, filter | Unit is operational. No hardware problems found. Error/s found is suspected to be sample and/or cartridge related. Found error/s for forwarding to applications specialist for further analysis. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5981218406312342572/5981218406312342572_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5981218406312342572/5981218406312342572_signature_31.png
| 2024-07-30 | 2024-07-30 12:00 | Manila Central University | Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | lab_mcuhospital@yahoo.com |  | 28951 | CONSORTIUM | purchasing@mcufdtmf.org | 27 | ASSAY: Xpert HPV, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 326 |  | GXIV | 110018008 | Unit for maintenance | NA | #NAME? | Unit is operational and in good condition No hardware problems found |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5981218396319845096/5981218396319845096_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5981218396319845096/5981218396319845096_signature_31.png
| 2024-07-29 | 2024-07-29 15:14 | President Ramon Magsaysay Memorial Hospital | Iba, Zambales | 3 | Zambales | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28950 | MACARE-HBDC | prmmhmoleclab@gmail.com | 29 | ASSAY: Xpert SARS-CoV-2, COUNT: 1814 |  | GXIV-D | 838520 | Due for preventive maintenance B3:888979 HW Failed B4:688212 HW Failed Monitor display shaking and freezing | Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational exept modules B23 and B4 Monitor for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5980471629024388997/5980471629024388997_signature_28.png | Angelina Cuaton | https://www.jotform.com/uploads/macarewecare/222927978372067/5980471629024388997/5980471629024388997_signature_31.png
| 2024-07-29 | 2024-07-29 13:39 | Masinloc Rural Health Unit | Masinloc Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28949 | DOH | rodenne@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2351 |  | GXIV-D | 831608 | Stuck cartridge module A2 UPS not Backing up | Repair Preventive maintenance Troubleshooting | Performed stuck cartridge removal Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter UPS troubleshooting Archived test results Performed functionality test Gathered reports | Unit is operational UPS not Backing up | UPS for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5980414229023462057/5980414229023462057_signature_28.png | Rodenn Educala | https://www.jotform.com/uploads/macarewecare/222927978372067/5980414229023462057/5980414229023462057_signature_31.png
| 2024-07-29 | 2024-07-29 12:00 | Southern Philippines Medical Center | JP Laurel Bajada Davao | 11 | Davao del Sur | ICU | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28948 | COMMERCIAL | icu2estuya@gmail.com | 26.5 |  | 2024-07-24 | MCS+ | 15H034G | Centrifuge not working | Replace the defective Centrifuge Assembly with Serial SN: 7562 with Centrifuge Assembly SN: 282 Perform Preventive Maintenance, Calibration and Diagnostic Test Perform Cycle Test | Perform Centrifuge replacement to isolate the problem | Successfully passed the cycle test Unit is Operational | Centrifuge Assembly Service Part Only, Request of Quotation for Centrifuge Assembly | MCS+ Centrifuge Assembly SN: 282 | 11747 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5980370805323101929/5980370805323101929_signature_28.png | Dianne Karla De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5980370805323101929/5980370805323101929_signature_31.png
| 2024-07-29 | 2024-07-19 11:46 | San Isidro Rural Health Unit | San Isidro Northern Samar | 8 | Northern Samar | TB-DOTS | NO | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28947 | UNOPS | crisdiazmedtech@yahoo.com | 24 |  | 2024-07-19 | GXIV - 10C | 110023515 | Unit for Installation | Installation | - Install Genexpert machine SN: 110023515 - CPU SN:CPH1023F47168 - Monitor SN:CN0WMYC5FCC0037KCLJXA05 - UPS SN:AT1K231111ST0005 - Printer SN:CN3B2431VN - Perfomed end user training on GeneXpert operation, cleaning, and basic troubleshooting. | Unit is in good condition |  | Genexpert machine CPU Printer Machine UPS Barcode scanner Monitor | DR 11643 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5971767726328808311/5971767726328808311_signature_28.png | cristina d. diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5971767726328808311/5971767726328808311_signature_31.png
| 2024-07-29 | 2024-07-26 13:56 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28946 | COMMERCIAL | lcpmicrolab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 12 ASSAY: Xpert SARS-CoV-2, COUNT: 3 |  | GX4 L | 804205 | D2,D3 and D4 disabled encounter 5007 error Defective ups Last xpertcheck 6/30/21 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | D1 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5980114976837339603/5980114976837339603_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5980114976837339603/5980114976837339603_signature_31.png
| 2024-07-29 | 2024-07-26 13:52 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28945 | MACARE-HBDC | lcpmicrolab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 408 ASSAY: Xpert SARS-CoV-2, COUNT: 965 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 ASSAY: Xpert HIV-1 Viral Load, COUNT: 113 ASSAY: Others, COUNT: 3 |  | GX4. D | 838837 | Due for maintenance B1 and B4 reported already encounter temp error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B2 and B3 operational B1 and B4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5980114616837024146/5980114616837024146_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5980114616837024146/5980114616837024146_signature_31.png
| 2024-07-29 | 2024-07-26 10:26 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28944 | MACARE-HBDC | avabichuela3216@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 361 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 685 |  | GX16 D | 838655 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5980114106831714938/5980114106831714938_signature_28.png | Arjay V. Abichuela (PHC) | https://www.jotform.com/uploads/macarewecare/222927978372067/5980114106831714938/5980114106831714938_signature_31.png
| 2024-07-29 | 2024-07-26 09:47 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28943 | MACARE-HBDC | eamcmolecularlab@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 18 ASSAY: Xpert HPV, COUNT: 249 ASSAY: Xpert SARS-CoV-2, COUNT: 1281 ASSAY: Xpert SARS-CoV-2, COUNT: 2 |  | GX4(10C) | 110016781 | Due for maintenance | Installed sticker Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5980113656839421485/5980113656839421485_signature_28.png | Aldwin Kristofer M Orpilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5980113656839421485/5980113656839421485_signature_31.png
| 2024-07-29 | 2024-07-26 09:22 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Tbdots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28942 | MACARE-HBDC | noriel_ruiz7@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2433 |  | GX4 (10C) | 110014085 | Due for maintenance | Installed sticker Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5980113346833882419/5980113346833882419_signature_28.png | Noriel S.Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5980113346833882419/5980113346833882419_signature_31.png
| 2024-07-28 | 2024-07-23 12:58 | Central Luzon Doctor's Hospital | Hospital drive Brgy. San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 28941 | COMMERCIAL | pjlsantos@cldh.com.ph | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 ASSAY: Xpert SARS-CoV-2, COUNT: 186 |  | GXIV-D | 839232 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5979838793128584426/5979838793128584426_signature_28.png | Precision Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5979838793128584426/5979838793128584426_signature_31.png
| 2024-07-28 | 2024-07-23 13:01 | Central Luzon Doctor's Hospital | Hospital drive Brgy San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | cldhlaboratory@cldh.com.ph |  | 28940 | COMMERCIAL | pjlsantos@cldh.com.ph | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 ASSAY: Xpert SARS-CoV-2, COUNT: 186 |  | GXII-D | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5979838433123543443/5979838433123543443_signature_28.png | Precision Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5979838433123543443/5979838433123543443_signature_31.png
| 2024-07-28 | 2024-07-28 21:05 | Padada Rural Health Unit | Rizal Street, Poblacion Padada | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28939 | DOH | zhakie.marvz@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 500 |  | GX-IV 4D 6C | 838511 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational | Enduser advised to achieve optimal operating temperature when using the machine |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5979813572211215145/5979813572211215145_signature_28.png | Marvielyn Lusica RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5979813572211215145/5979813572211215145_signature_31.png
| 2024-07-27 | 2024-07-26 11:33 | Isulan Rural Health Unit | Isu | 12 | South Cotabato | Laboratory | NO | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28938 | UNOPS | nancymoscoso18@gmail.com | 25.8 |  |  | Gx4- 10C | 110023520 | N/A | Perform Genexpert machine Installation | Installed: Genexpert machine SN:110023520, with barcode and stand CPU SN: CPH1023F47172 MONITOR SN: CN0WMYC5FCC0037KCL2XA05 UPS : Accupower 1000VA SN:AT1K231111ST0036 PRINTER SN: CN3B2431MH Import Assay: MTB-ULTRA and Hiv Viral load | All modules and accessories are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5978960627148415812/5978960627148415812_signature_28.png | NANCY A. MOSCOSO | https://www.jotform.com/uploads/macarewecare/222927978372067/5978960627148415812/5978960627148415812_signature_31.png
| 2024-07-26 | 2024-07-26 18:27 | NAZARETH HEALTH CENTER | NAZARETH CAGAYAN DE ORO CITY | 10 | Misamis Oriental | LABORATORY | NO | TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28937 | DOH | nanettebtolentino@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 988 |  | GX-IV | 838269 | SYSTEM UNIT DOES NOT FUNCTION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5977993159123182807/5977993159123182807_signature_28.png | Nanette Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/5977993159123182807/5977993159123182807_signature_31.png
| 2024-07-26 | 2024-07-26 14:09 | Calamba Medical Centee | Calamba Laguna | 4A | Laguna | Microbiology | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28936 | COMMERCIAL | maria.resurreccion@medilinx.com.ph | 25 | ASSAY: Xpert HBV VL, COUNT: 55 ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 682 |  | GXII | 812916 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5977836053202621064/5977836053202621064_signature_28.png | Maria Niña Resurreccion | https://www.jotform.com/uploads/macarewecare/222927978372067/5977836053202621064/5977836053202621064_signature_31.png
| 2024-07-26 | 2024-07-26 12:24 | Bagong Pag-asa Health Center | Bagong Pag-asa, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28935 | DOH | louiseanntangolmontenegro@gmail.com | 25 |  | 2024-07-25 | GXIV 6C D | 831597 | Original CPU fixed Back up CPU for pull out, original CPU for installation | Pulled out back up CPU Installed original CPU Generated report | Archived reports Pulled out back up CPU Setup original CPU | Unit is operational All modules working |  | Genexpert Desktop | DR 11751 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5977784904119698471/5977784904119698471_signature_28.png | Louise Ann Montenegro | https://www.jotform.com/uploads/macarewecare/222927978372067/5977784904119698471/5977784904119698471_signature_31.png
| 2024-07-25 | 2024-07-25 18:45 | MAMBAJAO RURAL HEALTH UNIT | MAMBAJAO, CAMUIGIN | 10 | Misamis Oriental | TB DOTS | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28934 | UNOPS | rhumambajaolaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-IV 10C | 110023519 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5977137820216816042/5977137820216816042_signature_28.png | MARJORIE OCLARIT | https://www.jotform.com/uploads/macarewecare/222927978372067/5977137820216816042/5977137820216816042_signature_31.png
| 2024-07-25 | 2024-07-25 14:23 | Veterans Memorial Medical Center | North Ave., Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28933 | MACARE-HBDC | vmmc.moleclab@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 178 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1304 ASSAY: Xpert SARS-CoV-2, COUNT: 110 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GX IV 6C-D | 839150 | GX modules defective To perform additional troubleshooting from cepheid | Performed preventive maintenance Performed additional troubleshooting Generated report | Checked and cleaned cartridge bays, icore slot and filter Ping the IP address Checked voltages, cables gateway board and grounding | Desktop OK GX modules defective | with Maam Issa |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5976981603613322944/5976981603613322944_signature_28.png | Gerard GhanivC. Sidic | https://www.jotform.com/uploads/macarewecare/222927978372067/5976981603613322944/5976981603613322944_signature_31.png
| 2024-07-25 | 2024-07-23 10:18 | Padre Burgos Rural Health Unit | Padre Burgos, Southern Leyte | 8 | Southern Leyte | Tb Dots | NO | INSTALLATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28932 | UNOPS | keishaelezze14@gmail.com | 25 |  |  | 10C GXIV-4D | 110023514 | None | None | #NAME? | Module is installed and functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5976962292167814497/5976962292167814497_signature_28.png | Keisha Elezze C. Batalon | https://www.jotform.com/uploads/macarewecare/222927978372067/5976962292167814497/5976962292167814497_signature_31.png
| 2024-07-25 | 2024-07-25 13:35 | Cong. Lamberto L. Macias Memorial Hospital | Ppoblacion | 7 | Cebu | Laboratory | YES | MAINTENANCE GX ID Sticker Installation | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 28931 | DOH | baldebrinmarquenicole@gmail.com | 23 |  | 2024-07-25 | GX4 | 838482 | .Due for Preventive Maintenance. .GX ID Installation. | Perform Preventive Maintenance | Perform Xpertcheck. Save IQR Save Installation Qualification Save GXC file | .Module B1,B3 fail for XpertCheck .Module B2, B4 passed for XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5976957448416963761/5976957448416963761_signature_28.png | Marque Nicole M. Baldebrin | https://www.jotform.com/uploads/macarewecare/222927978372067/5976957448416963761/5976957448416963761_signature_31.png
| 2024-07-25 | 2024-07-25 09:38 | Socorro Rural Health Center | Brgy. Navarro, Socorro | 13 | Surigao del Norte | RTDL | NO | INSTALLATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28930 | UNOPS | dynamite.tux@gmail.com | 30 |  |  | GX-4 10 Colors | 110023523 | None | None | - Installation of New GeneXpert Machine (Gx-4 10 Colors) including accessories (Desktop, Monitor Display, Barcode scanner with stand, UPS, and printer) - Performed end user training on GeneXpert operation, cleaning and basic troubleshooting. | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5976860093501571219/5976860093501571219_signature_28.png | James P. Anastacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5976860093501571219/5976860093501571219_signature_31.png
| 2024-07-24 | 2024-07-24 14:36 | Claver Rural Health Unit | Brgy. Tayaga, Claver | 13 | Surigao del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28929 | DOH | rhodaruthdalogdog@gmail.com | 29 |  |  | GX-4 6 Colors | 838514 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5976123152427754217/5976123152427754217_signature_28.png | Ruth Rhoda Dalogdog | https://www.jotform.com/uploads/macarewecare/222927978372067/5976123152427754217/5976123152427754217_signature_31.png
| 2024-07-24 | 2024-07-24 14:28 | Surigao City Health Office | City Health Office, City Hall Compound, Brgy. Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | INSTALLATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28928 | UNOPS | jingnep@yahoo.com | 28 |  |  | GX-4 10 Colors | 110023522 | None | None | - Installation of New GeneXpert Machine (GX-4 10 Colors) including GeneXpert Desktop, Monitor Display, Barcode Scanner with Stand, UPS (ACCUPOWER 1000VA) and HP Printer - Performed end user training on GeneXpert operation, cleaning and basic troubleshooting | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5976121342429306100/5976121342429306100_signature_28.png | Jessica I. Nepomucino | https://www.jotform.com/uploads/macarewecare/222927978372067/5976121342429306100/5976121342429306100_signature_31.png
| 2024-07-24 | 2024-07-24 13:19 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28927 | DOH | gprincessrona@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 459 | 2024-07-24 | GXIV | 838551 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5976080020412417269/5976080020412417269_signature_28.png | Princess Geven | https://www.jotform.com/uploads/macarewecare/222927978372067/5976080020412417269/5976080020412417269_signature_31.png
| 2024-07-24 | 2024-07-24 10:09 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28926 | DOH | gmpelaez144@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1342 |  | GxIV-6color | 838528 | Module 1 for Service Required due to multiple temperature error | Genexpert machine Cleaning Perform XPC at module 1 for Troubleshooting Genexpert errors/problem reporting Data collection | Perform Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975975529418871368/5975975529418871368_signature_28.png | GLADIOLA MARIE E. PELAEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5975975529418871368/5975975529418871368_signature_31.png
| 2024-07-24 | 2024-07-23 14:25 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28925 | UNOPS | onehope834@gmail.com | 26 |  |  | GxIV - 10C | 110023524 | N/A | Perform Genexpert machine Installation | Installed: Genexpert machine SN:110023524 CPU SN:CPH1023F46949 Monitor SN:CN0WMYC5FCC00337KCK3XA05 UPS SN:AT1K231111ST0047 PRINTER SN:1001363852 | All Genexpert modules and accessories are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975960579417214081/5975960579417214081_signature_28.png | Emelyn A. Cadiz, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5975960579417214081/5975960579417214081_signature_31.png
| 2024-07-23 | 2024-07-23 17:04 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | CALIBRATION | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28924 | COMMERCIAL | pshclinlab@gmail.com | 25 |  |  | MCS+ | 02D167 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975347950019472668/5975347950019472668_signature_28.png | Roda Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5975347950019472668/5975347950019472668_signature_31.png
| 2024-07-23 | 2024-07-23 17:03 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | CALIBRATION | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28923 | COMMERCIAL | pshclinlab@gmail.com | 25 |  |  | MCS+ | 11H307 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975346680019495663/5975346680019495663_signature_28.png | Roda Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5975346680019495663/5975346680019495663_signature_31.png
| 2024-07-23 | 2024-07-23 13:53 | Makilala Rural Health Unit | P.16 Makilala Poblacion Makilala, Cotabato | 12 | Cotabato | TB-DOTS | NO | INSTALLATION | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28922 | UNOPS | abegailweill17@gmail.com | 25.6 |  | 2024-06-20 | GX-IV 4D 10C | 110023521 | For Installation of GeneXpert 10C Unit | Making sure that the machine is correctly installed within acceptable temperature operating conditions. | Perform Installation: Installed the GeneXpert System Installed the UPS with SN: AT1K231111ST0046 Installed the CPU System Unit with SN:CPH1023F47469 Installed the Monitor With SN: CN0WMYC5FCC0037KCLNXA05 Installed the Barcode Scanner with SN: 22355010554001 Installed the GX-IV 4D 10C with SN:110023521 Installed the Printer with SN: CN3B2431HT Provide Basic Operation Training and Maintenance to Enduser upon turnover | Successfully Installed the GX-IV 4d 10C UNOPS unit |  |  | DR 11651 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975236606127556257/5975236606127556257_signature_28.png | Abegail Publico RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5975236606127556257/5975236606127556257_signature_31.png
| 2024-07-23 | 2024-07-19 15:20 | St. Luke's Medical Center BGC | Rizal Dr., BGC, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 28921 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 15 ASSAY: Xpert CT/NG, COUNT: 56 ASSAY: Xpert HBV VL, COUNT: 7 ASSAY: Xpert HPV, COUNT: 64 ASSAY: Xpert MTB/RIF Ultra, COUNT: 137 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV-D-10C | 120002383 | For network switch disconnection | Removed connection from network switch | Removed connection from network switch Connected to separate CPU Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975230503615215595/5975230503615215595_signature_28.png | Sapla Grace | https://www.jotform.com/uploads/macarewecare/222927978372067/5975230503615215595/5975230503615215595_signature_31.png
| 2024-07-23 | 2024-07-19 15:13 | St. Luke's Medical Center BGC | Rizal Dr. BGC, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 28920 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 |  |  | GX-IV-D-6C | 839231 | Setup | Removed network switch connection | Removed connection from network switch Connected to separate CPU Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A4/210096427 reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975230503617152608/5975230503617152608_signature_28.png | Sapla, Grace | https://www.jotform.com/uploads/macarewecare/222927978372067/5975230503617152608/5975230503617152608_signature_31.png
| 2024-07-23 | 2024-07-23 13:24 | Macare In-house | Aurora Grand, Quezon City | NCR | NCR 2nd - Quezon City | Macare | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28919 | DOH | bagbaguenbrenguiler@gmail.com | 25 |  |  | GXIV-D-6C | 838535 | CPU Not Booting. | Repair | Checked Power Supply. Replaced RAM. Performed Functionality Test. | CPU is Operational. | CPU SN: CPH0520F08551 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975217717119844113/5975217717119844113_signature_28.png | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5975217717119844113/5975217717119844113_signature_31.png
| 2024-07-23 | 2024-07-16 10:31 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta.Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28918 | MACARE-HBDC | covidlabjrrmc@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 336 ASSAY: Xpert SARS-CoV-2, COUNT: 2297 |  | GX16-D-10C | 110013856 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reporta. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214867111911697/5975214867111911697_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214867111911697/5975214867111911697_signature_31.png
| 2024-07-23 | 2024-07-10 10:59 | Tunasan Health Center | Rodriguez St, Tunasan, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28916 | PBSP | gigietan0911@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1973 ASSAY: Xpert HIV-1 Viral Load, COUNT: 257 |  | GXIV-D-6C | 807613 | Module 3 with 5007 Errors. | Preventive Maintenance. XpertCheck. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed XpertCheck on Reported Module. Performed Self-test. Generate Reports. | 3 Modules are Operational. Module 3 Disabled and Failed in XpertCheck. | B3/210100380 Service Required. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797118594797/5975214797118594797_signature_28.png | Geraldine Trazo Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797118594797/5975214797118594797_signature_31.png
| 2024-07-23 | 2024-07-09 11:34 | Taguig City Health Office | Col P. Cruz, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28917 | PBSP | taguigclinicallaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4037 |  | GXIV-D-6C | 814372 | Module 3 with Temperature Error | Module Replacement. | Cleaned Cartridge Bay, Plunger Rod, And Icore Optics. Checked Filter if Clogged. Generate Reports. Performed Self-test. | All Modules are Operational. | B3: 661374->723715 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797119651168/5975214797119651168_signature_28.png | Gina Dumawal | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797119651168/5975214797119651168_signature_31.png
| 2024-07-23 | 2024-07-12 11:20 | De La Sall Health Sciences Institute | Congressional Ave., Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28914 | CONSORTIUM | cevillete@dlsmhsi.edu.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 75 |  | GXIV-D-6C | 812380 | Date and Time Problem in CPU. | Troubleshoot. Preventive Maintenance. | Checked Bios Setup. Replaced CMOS Battery. Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Generate Reports. Performed Self-test. | All Modules are Operational. CPU Problem Resolved. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797111302112/5975214797111302112_signature_28.png | Chlorish Ann Villete | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797111302112/5975214797111302112_signature_31.png
| 2024-07-23 | 2024-07-11 11:25 | Bacoor Rural Health Unit II (Panapaan HC) | Floraville Subd, Panapaan,Bacoor | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28913 | PBSP | tobiasmariacecila@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3066 ASSAY: Xpert HIV-1 Viral Load, COUNT: 144 |  | GXIV-D-6C | 807584 | Unit is due for XpertCheck. | XpertCheck. Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Performed XpertCheck. Generate Reports | 1 Module Disabled. 3 are Operationals. | 1-Service Required, 2- Pass with Error, 3- Pass,4-Newly Replaced |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797112046141/5975214797112046141_signature_28.png | Maria Cecila Tobias | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797112046141/5975214797112046141_signature_31.png
| 2024-07-23 | 2024-07-16 10:25 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta.Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28915 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 730 ASSAY: Xpert SARS-CoV-2, COUNT: 235 |  | GXIV-D-10C | 110014099 | Unit is due for Maintenance | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | 1 Module with Error, 3 Module are Operational. | B1/210069526 for Reporting. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797119599209/5975214797119599209_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5975214797119599209/5975214797119599209_signature_31.png
| 2024-07-22 | 2024-07-22 16:40 | Surigao City Health Office | Brgy. San Juan, Surigao City | 13 | Surigao del Norte | RTDL | NO | INSTALLATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28912 | UNOPS | jingnep@yahoo.com | 23 |  |  | GX-4 10C | 110023522 | Heat Sink not installed in CPU error. | Troubleshooting of cpu error and found heat sink clip not in contact. (Should solderized) | Installation of GeneXpert 10C and its peripherals | CPU prior for replacement or to be solderized |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5974471084276101128/5974471084276101128_signature_28.png | Jessica I. Nepomuceno | https://www.jotform.com/uploads/macarewecare/222927978372067/5974471084276101128/5974471084276101128_signature_31.png
| 2024-07-22 | 2024-07-22 15:29 | Piñan Rural Health Unit | Poblacion South, Piñan | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28911 | UNOPS | ivannagmarchan@gmail.com | 27 |  |  | Gx 4 | 110023516 | None | New Installation Genexpert Basic training end user on operating and maintenace | New Installation Genexpert. Basic training end user on operating and maintenaceNew Installation. | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5974428402315766858/5974428402315766858_signature_28.png | Ivanna G. Marchan | https://www.jotform.com/uploads/macarewecare/222927978372067/5974428402315766858/5974428402315766858_signature_31.png
| 2024-07-22 | 2024-07-22 14:12 | Macare In House | 1005 Aurora Grand Aurora blvd. | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 28910 | COMMERCIAL | florescanoel77@gmail.com | 28 |  |  | Dell Optiplex XE2 | 12151624502 | CPU for repair | Reformat CPU Installed GXDX Test run unit | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5974385403618845324/5974385403618845324_signature_28.png | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5974385403618845324/5974385403618845324_signature_31.png
| 2024-07-22 | 2024-07-16 17:20 | Philippine Red Cross QC | Kalayaan Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28909 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 27 |  |  | MCS + | 02A227 | As per end user they reddish collection in upp. | Checked program parameters Checked and adjust line sensor value. Ajust bowl optics value 25 and 2800 Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5974353683611371771/5974353683611371771_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/5974353683611371771/5974353683611371771_signature_31.png
| 2024-07-19 | 2024-07-19 15:44 | Philippine Heart Center | East Ave., QC | NCR | NCR 2nd - Quezon City | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 28908 | COMMERCIAL | phc_bb@yahoo.com | 28 |  | 2024-07-19 | MCS+ | 11H204 | DPM is out of tolerance. | Performed troubleshooting on DPM. Performed functionality test. | 1. Checked and calibrate DPM parameteer. 2. Run simulation and functionality test. | Unit is operational. | N/A | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5971851408323650224/5971851408323650224_signature_28.png | Lian | https://www.jotform.com/uploads/macarewecare/222927978372067/5971851408323650224/5971851408323650224_signature_31.png
| 2024-07-19 | 2024-07-19 14:33 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28906 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 59 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3056 ASSAY: Xpert HIV-1 Viral Load, COUNT: 472 |  | 6C GXIV-4D | 838285 | Due for monthly PM | #NAME? | -Preventive Maintenance -Upgrade XpertCheck software to 1.5b | Machine is operational. Module B4 still disabled; waiting for allocation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5971811165244012613/5971811165244012613_signature_28.png | Jogeveve Alforque | https://www.jotform.com/uploads/macarewecare/222927978372067/5971811165244012613/5971811165244012613_signature_31.png
| 2024-07-19 | 2024-07-19 14:46 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28907 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2917 ASSAY: Xpert HIV-1 Viral Load, COUNT: 398 |  | 6C GVI-4D | 820206 | Due for monthly PM | #NAME? | -Preventive Maintenance -Upgrade XpertCheck software to 1.5b | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5971811165244467505/5971811165244467505_signature_28.png | Jogeveve Alforque | https://www.jotform.com/uploads/macarewecare/222927978372067/5971811165244467505/5971811165244467505_signature_31.png
| 2024-07-19 | 2024-07-19 14:01 | Candelaria Rural Health Unit | Candelaria, Quezon | 4A | NCR 2nd - Quezon City | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28905 | PBSP | shirleyzrubz@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2259 ASSAY: Xpert HIV-1 Viral Load, COUNT: 107 | 2024-07-09 | GX IV 6C D | 814155 | Back Up laptop for installation while waiting for CPU Unit due for Calibration | Installed backup laptop Performed preventive maintenance Performed Xpertcheck Generated report | Installed back up laptop Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck D1: Passed D2: Failed D3: Passed D4: Passed | Unit is operational Module D2 for reporting | jotform error. Service done on 07/12/24 | Bakcup laptop (Lenovo) Xpertcheck kit | G7090 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5971787153617243218/5971787153617243218_signature_28.png | Shirley Dollente | https://www.jotform.com/uploads/macarewecare/222927978372067/5971787153617243218/5971787153617243218_signature_31.png
| 2024-07-19 | 2024-07-19 13:42 | Macalelon Rural Health Unit | Macalelon, Quezon Province | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28904 | DOH | cheyster7388@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 74 ASSAY: Xpert MTB/RIF Ultra, COUNT: 394 |  | GX IV 6C D | 831636 | Unit due for maintenance | Performed preventive maintenance Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter | Unit is operational Module A4 for reporting | Jotform error. Service done on 7/11/24 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5971779143618060404/5971779143618060404_signature_28.png | Cheysrter O. Agustin | https://www.jotform.com/uploads/macarewecare/222927978372067/5971779143618060404/5971779143618060404_signature_31.png
| 2024-07-19 | 2024-07-19 11:46 | San Isidro Rural Health Unit | San Isidro Northern Samar | 8 | Northern Samar | TB-DOTS | NO | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28903 | UNOPS | crisdiazmedtech@yahoo.com | 24 |  | 2024-07-19 | GXIV - 10C | 110023515 | Unit for Installation | Installation | - Install Genexpert machine SN: 110023515 - CPU SN:CPH1023F47168 - Monitor SN:CN0WMYC5FCC0037KCLJXA05 - UPS SN:AT1K231111ST0005 - Printer SN:CN3B2431VN | Unit is in good condition |  | Genexpert machine CPU Printer Machine UPS Barcode scanner Monitor | DR 11643 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5971767726328808311/5971767726328808311_signature_28.png | cristina d. diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5971767726328808311/5971767726328808311_signature_31.png
| 2024-07-19 | 2024-07-19 11:33 | Cuyegkeng Health Center | Cuyegkeng, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28902 | PBSP | demlhc.pasay@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1371 |  | GX IV 6C D | 805106 | UPS for replacement | Performed preventive maintenance Installed ups Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Checked ups Replaced ups | Unit is operational All modules working | Jotform error. Service made on 07/09/2024 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5971704493613987434/5971704493613987434_signature_28.png | Loraine Cosico | https://www.jotform.com/uploads/macarewecare/222927978372067/5971704493613987434/5971704493613987434_signature_31.png
| 2024-07-18 | 2024-07-18 15:25 | AFPMC AFRIMS | V. Luna, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28901 | MACARE-HBDC | mtvalderama@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 57 ASSAY: Xpert MTB/RIF Ultra, COUNT: 717 ASSAY: Xpert SARS-CoV-2, COUNT: 969 ASSAY: Xpert HCV VL, COUNT: 40 ASSAY: Xpert HIV-1 Viral Load, COUNT: 45 ASSAY: Xpert HIV-1 Viral Load, COUNT: 88 |  | GX IV 6C D | 845462 | Unit due for maintenance | Performed preventive maintenance Performed reliability test B3 for reporting Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test All modules passed reliability test Module B3 system log frequent error 2014 | Unit is operational B3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5970983670578333221/5970983670578333221_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/5970983670578333221/5970983670578333221_signature_31.png
| 2024-07-18 | 2024-07-17 13:41 | Carmen Health Center | Brgy. Carmen, Cagayan de Oro City | 10 | Misamis Oriental | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28900 | DOH | carmenhealthcenterlab@gmail.com | 20 |  |  | GX-4 | 838442 | 1 module found error in temperature (B2) while 1 module for observation. 2 modules are working and functional | None | Preventive Maintenance and Data Gathering | 2 modules for observation 2 modules are working and functional All units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5970963228102420918/5970963228102420918_signature_28.png | Muriel G. Melchor | https://www.jotform.com/uploads/macarewecare/222927978372067/5970963228102420918/5970963228102420918_signature_31.png
| 2024-07-18 | 2024-07-18 14:40 | University of The East Ramon Magsaysay Medical Center | Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28899 | CONSORTIUM | osangcomedador19@gmail.com | 21 |  | 2024-07-16 | GXIV | 847578 | Modules B1, B4 for replacement | Replaced B1, B4 | #NAME? | Unit is operational All modules functio al |  | GeneXpert Module SN: 912409, 912371 | 11724 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5970942554026637008/5970942554026637008_signature_28.png | Osang Comedador | https://www.jotform.com/uploads/macarewecare/222927978372067/5970942554026637008/5970942554026637008_signature_31.png
| 2024-07-18 | 2024-07-18 13:49 | LAPULAPU CITY HEALTH OFFICE | LAPULAPU CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28898 | DOH | tanya_pax@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2184 |  | GX-IV | 837661 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5970912490725064479/5970912490725064479_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5970912490725064479/5970912490725064479_signature_31.png
| 2024-07-18 | 2024-07-18 13:21 | Makati Medical Center | Amorsolo St., Makati City | NCR | NCR 4th - Makati | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomedsupervisor@makatimed.net.ph |  | 28897 | COMMERCIAL | blood.bank@makatimed.net.ph | 28 |  | 2024-07-12 | MCS+ | 97H330 | Due for maintenance and Calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Replaced rollers Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Performed maintenance and Calibration | Operational |  | Rollers | 98419 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5970904771218245951/5970904771218245951_signature_28.png | Makati Medical Center - Blood Bank and Transfusion Services | https://www.jotform.com/uploads/macarewecare/222927978372067/5970904771218245951/5970904771218245951_signature_31.png
| 2024-07-18 | 2024-07-18 12:59 | Tomas Claudio Health Center (District A) | Camus St. Corner E. Quirino Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28896 | DOH | arleen_p_alconera@yahoo.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 580 |  | GX-IV 4D 6C | 838492 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses. Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Only Three Modules are operational, Module 2 disabled due to multiple temperature related errors ; Waiting for parts from Cepheid |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5970885801817716488/5970885801817716488_signature_28.png | Arleen P. Alconera RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5970885801817716488/5970885801817716488_signature_31.png
| 2024-07-18 | 2024-07-18 11:59 | Makati Medical Center | Amorsolo st, Makati City | NCR | NCR 4th - Makati | CRM | NO | MAINTENANCE | OTHERS | Noel Floresca |  | florescanoel77@gmail.com | biomedsupervisor@makatimed.net.ph | Miltenyi | 28895 | COMMERCIAL | AlcorDeeond.Padilla@makatimed.net.ph | 27 |  | 2024-07-12 | Clinimacs | 797 | Due for maintenance | Cleaned unit Checked rollers Checked and test tube sensor Checked and test keypad and lcd display. Checked program parameters Performed functionality test. | Performed maintenance | Operational |  |  | 98420 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5970852111212708210/5970852111212708210_signature_28.png | Alcor Deeond F. Padilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5970852111212708210/5970852111212708210_signature_31.png
| 2024-07-17 | 2024-07-16 15:51 | Pasig Treatment Hub | Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28894 | MACARE-HBDC | pathmedtech@gmail.com | 20 | ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 800 |  | GX-IV-L-6C | 821648 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules A1/733228, A2/734493 and A4/804598 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5970124773612057485/5970124773612057485_signature_28.png | Jhozza Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5970124773612057485/5970124773612057485_signature_31.png
| 2024-07-17 | 2024-07-17 13:22 | Monkayo Rural Health Unit | Poblacion, Monkayo | 11 | Davao de Oro | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28893 | DOH | kendralois730@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1733 |  | GX-IV 6C 4D | 838487 | Lost Communication at Module 1 only Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Reset Gateway Board and GeneXpert Connections Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Troubleshooting | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5970033701069968649/5970033701069968649_signature_28.png | Karen Anuta RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5970033701069968649/5970033701069968649_signature_31.png
| 2024-07-17 | 2024-07-17 10:59 | Talakag Rural Health Unit | Poblacion, Talakag | 10 | Bukidnon | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28892 | DOH | khizajesseltimo@gmail.com | 27 |  |  | GX-4 | 838565 | None | Maintenance | Preventive Maintenance and Data Collection | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5969953170497437899/5969953170497437899_signature_28.png | Khiza Jessel Timo | https://www.jotform.com/uploads/macarewecare/222927978372067/5969953170497437899/5969953170497437899_signature_31.png
| 2024-07-16 | 2024-07-16 14:15 | Nazareth Health Center | Brgy. Nazareth, Cagayan de Oro City | 10 | Misamis Oriental | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28891 | DOH | nanettebtolentino@gmail.com | 28 |  |  | GX-4 | 838269 | B1 found Service Required from XpertCheck Calibration | Maintenance | Preventive Maintenance | 3 modules working, 1 module for replacement (service required) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5969527586226222207/5969527586226222207_signature_28.png | Michelle Therese Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5969527586226222207/5969527586226222207_signature_31.png
| 2024-07-16 | 2024-07-16 18:09 | CANDIJAY COMMUNITY HOSPITAL | CANDIJAY BOHOL | 7 | Bohol | LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28890 | DOH | angelinaredulla@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 760 |  | GX-IV | 838396 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5969340160257101465/5969340160257101465_signature_28.png | ANGELINA O. REDULLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5969340160257101465/5969340160257101465_signature_31.png
| 2024-07-16 | 2024-07-16 10:49 | Metro Antipolo Hospital and Medical Center | Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28889 | COMMERCIAL | mahmci_labcmt@yahoo.com | 20 | ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 210 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5969221611076699381/5969221611076699381_signature_28.png | Graziela Marie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/5969221611076699381/5969221611076699381_signature_31.png
| 2024-07-16 | 2024-07-16 14:48 | Macare In-House | 1005 Aurora Blvd, Project 3, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | gpdmagnaye30@gmail.com |  | 28888 | MACARE-HBDC | macarewecare@gmail.com | 25 |  |  | HP Z240 | 2UA8471VRW | Unit for repair | Reapair CPU | Reaired Unit Performed Functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5969221063613289633/5969221063613289633_signature_28.png | Macare in-House | https://www.jotform.com/uploads/macarewecare/222927978372067/5969221063613289633/5969221063613289633_signature_31.png
| 2024-07-16 | 2024-07-11 13:45 | Tacurong City Health Office | Tacurong City, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28887 | PBSP | cho_clinlab@yahoo.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1610 | 2024-07-11 | Gx-IV (6-color) | 815393 | Module 4 for Service Required | Replace defective module 4 SN:894007 with new module SN: 759943 Genexpert machine Cleaning/preventive maintenance Checking of errors Data collection | Perform Genexpert machine module replacement/repair Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6-color | Cepheid case: 1566169/1566629 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5969106149495254046/5969106149495254046_signature_28.png | CHERLYN A. BOLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5969106149495254046/5969106149495254046_signature_31.png
| 2024-07-16 | 2024-07-16 10:46 | Cagayan de Oro City Health Office | Cagayan de Oro City | 10 | Misamis Oriental | RTDL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28886 | DOH | chocdotbdots@gmail.com | 28 |  |  | GX-4 | 838441 | CPU won't turn on Defective Monitor (Glitching) | Troubleshooting and Preventive Maintenance | Re-seating and Cleaning of the RAM in CPU Changing CMOS Battery Replacing Old Monitor (CNK9430B95) to Tablon Health Center Monitor Unit (CNK9430C26) Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5969083110499377221/5969083110499377221_signature_28.png | Charissa Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5969083110499377221/5969083110499377221_signature_31.png
| 2024-07-16 | 2024-07-10 13:22 | Lingayen Rural Health Unit 2 | Lingayen Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28885 | PBSP | michelledoriarmt@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1069 | 2024-07-03 | GXIV-D | 834629 | B1:784881 B2:785081 For module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational B1: 702701 new module B2: 706465 new module |  | 2 Genexpert module | 49004 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5969004140119487511/5969004140119487511_signature_28.png | Glady DG. Urbano | https://www.jotform.com/uploads/macarewecare/222927978372067/5969004140119487511/5969004140119487511_signature_31.png
| 2024-07-15 | 2024-07-12 13:50 | Hi-precision Diagnostics | N.S. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28883 | CONSORTIUM | pcr@hi-precision.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 127 ASSAY: Xpert HPV, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6523 |  | GX-IV-D-6C | 835601 | Module B4 has error 5007 | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B4 is operational but under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5968429193614743863/5968429193614743863_signature_28.png | Andrei Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/5968429193614743863/5968429193614743863_signature_31.png
| 2024-07-15 | 2024-07-12 13:53 | Hi-Precision Diagnostics | N.S. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28884 | CONSORTIUM | pcr@hi-precision.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 13 ASSAY: Xpert CT/NG, COUNT: 103 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7304 |  | GX-IV-D-6C | 835602 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5968429193618805609/5968429193618805609_signature_28.png | Andrei Mina | https://www.jotform.com/uploads/macarewecare/222927978372067/5968429193618805609/5968429193618805609_signature_31.png
| 2024-07-15 | 2024-07-12 15:31 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28882 | DOH | gbmmbg@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 609 |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5968287747389401742/5968287747389401742_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5968287747389401742/5968287747389401742_signature_31.png
| 2024-07-15 | 2024-06-26 10:31 | Pamplona Uno Health Center | Alabang-Zapote Road, Pamplona Uno, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28881 | PBSP | carolinevecino5821@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1332 |  | GXIV-D-6C | 838541 | Module 1 Service Required in XpertCheck. | Module Replacemet. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Replaced Module 1. Generate Reports. | B4 for Reporting. | B1: 657495 -> 761378 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5968134333614850449/5968134333614850449_signature_28.png | Caroline R. Vecino | https://www.jotform.com/uploads/macarewecare/222927978372067/5968134333614850449/5968134333614850449_signature_31.png
| 2024-07-13 | 2024-07-12 13:30 | Manito Rural Health Unit | Manito, Albay | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28880 | PBSP | rhumanito40@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3027 |  | GX4-6C | 804257 | B3/600725 for module replacement CPU not powering On | Replaced module: B3 - 600725>>210068327 Cleaned cartridge bays, plunger rod, icore Generate reports CPU cleaned and Reset RAM moved RAM to different slot Replaced power cord Checked UPS power output going to CPU plugged CPU directly to AC source checked CPU power button | Replaced B3/600725 module Performed Troubleshooting | Unit is Operational CPU not powering On(needs several attempt of plugged and unplugged from power source) | CPU request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5966797085816978248/5966797085816978248_signature_28.png | Hael Lachet L. Germones | https://www.jotform.com/uploads/macarewecare/222927978372067/5966797085816978248/5966797085816978248_signature_31.png
| 2024-07-13 | 2024-07-11 16:00 | Bicol Medical Center | Concepcion, Naga City | 5 | Camarines Sur | Bacteriology Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28879 | PBSP | nacariorosalie1@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1320 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7478 ASSAY: Xpert SARS-CoV-2, COUNT: 1641 ASSAY: Xpert HIV-1 Viral Load, COUNT: 833 ASSAY: Others, COUNT: 372 |  | GX4-10C | 842225 | A1/885370 for module replacement (service required in Calibration) | Replaced Module: A1 - 885370>>739879 Cleaned cartridge bays, plunger, fan Generate report | Replaced Module A1/885370 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5966789545811170171/5966789545811170171_signature_28.png | Rosalie Nacario, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5966789545811170171/5966789545811170171_signature_31.png
| 2024-07-13 | 2024-07-11 09:33 | Jose Panganiban Rural Health Unit | South Poblacion, Jose Panganiban | 5 | Camarines Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28878 | PBSP | materesitaperes1@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2030 |  | GX4-6C | 808641 | Module B1/667807 - temperature Error, for module replacement | Module replaced: B1/667807>>683577 Cleaned cartridge bays, plunger rod, icore Replaced filter Generate report | Replaced module B1/667807 Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5966787045813821791/5966787045813821791_signature_28.png | Ma. teresita G. Peres | https://www.jotform.com/uploads/macarewecare/222927978372067/5966787045813821791/5966787045813821791_signature_31.png
| 2024-07-13 | 2024-07-09 16:15 | Bicol South Luzon Sub National Reference Laboratory | BRTTH Compound, Legazpi City, Albay | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28877 | PBSP | rbalzaga.pbsp@gmail.com | 23 | ASSAY: Others, COUNT: 723 |  | GX4-6C | 110016954 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5966780545813904706/5966780545813904706_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5966780545813904706/5966780545813904706_signature_31.png
| 2024-07-13 | 2024-07-09 16:09 | Bicol South Luzon Sub National Reference Laboratory | BRTTH Compound, Legazpi City, Albay | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28876 | PBSP | rbalzaga.pbsp@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 347 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2125 ASSAY: Xpert HIV-1 Viral Load, COUNT: 547 |  | GX4-6C | 808636 | B4/644565 for module replacement | Replaced module: B4 - 644565>>684259 Cleaned cartridge bays, plunger rod, icore Generate reports | Replaced B4/644565 module | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5966778915816231352/5966778915816231352_signature_28.png | richard alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5966778915816231352/5966778915816231352_signature_31.png
| 2024-07-12 | 2024-07-12 12:31 | Sta. Rosa City Health Office 1 | 2nd Floor Clinical Laboratory, City Health Office Building B, City Goverment Compound, Rizal Boulevard, Brgy. Market Area, City of Sta. Rosa Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28875 | UNOPS | santarosacho1clinlab@gmail.com | 25 |  |  | GXIV-10C | 110023497 | For installation of GeneXpert System. | Installation. | Peformed installation of GeneXpert System. Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5965737779213351041/5965737779213351041_signature_28.png | ALONA NINA L. ALVARO | https://www.jotform.com/uploads/macarewecare/222927978372067/5965737779213351041/5965737779213351041_signature_31.png
| 2024-07-12 | 2024-07-11 11:31 | San Pedro City Health Office 1 | San Pedro Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28874 | DOH | venerdantos0425@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 62 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2120 | 2024-07-11 | GXIV | 821224 | Unit for PM. For module replacement of module A2. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A2 (604063) in replacement to def. A2 (678381). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5965737779214152447/5965737779214152447_signature_28.png | Vener R. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5965737779214152447/5965737779214152447_signature_31.png
| 2024-07-12 | 2024-07-10 10:07 | Doña Marta HC 10C | Doña Marta, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28873 | PBSP | LBCANICULA@gmail.com | 25 |  |  | GX IV 10C-D | 110020393 | 10C Unit for installation | Insatlled GXIV 10C D Generated report | Prepared the table to be used for installation Setup GX and accessories Setup User administration Installed assays | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5965692862199750061/5965692862199750061_signature_28.png | Lorna Canicula | https://www.jotform.com/uploads/macarewecare/222927978372067/5965692862199750061/5965692862199750061_signature_31.png
| 2024-07-11 | 2024-07-11 13:05 | Bangkal Health Center | Rodriguez Avenue, Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28871 | DOH | mhdbangkalsatlab@gmail.com | 22 |  |  | GX-IV-D-6C | 831875 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. | Filter is worn out. For reporting. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964896503556490823/5964896503556490823_signature_28.png | Danica Adolfo | https://www.jotform.com/uploads/macarewecare/222927978372067/5964896503556490823/5964896503556490823_signature_31.png
| 2024-07-11 | 2024-07-11 13:02 | Bangkal Health Center | Rodriguez Avenue, Makati City | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28872 | PBSP | mhdbangkalsatlab@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4949 | 2024-07-11 | GX-IV-D-6C | 838875 | Module B4 for replacement | Replaced module B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B4 Generated reports | Unit is operational. All modules operational |  | Old: B4/863458 New: B4/793029 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964896503557040055/5964896503557040055_signature_28.png | Danica Adolfo | https://www.jotform.com/uploads/macarewecare/222927978372067/5964896503557040055/5964896503557040055_signature_31.png
| 2024-07-11 | 2024-07-11 11:56 | Buluan District Hospital | Buluan, Maguindanao | BARMM | Maguindanao | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28870 | PBSP | bdhclinicallaboratory@gmail.com | 29.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1652 | 2024-07-11 | Gx-IV (6C) | 815490 | Module 2 for Service Required | Replace defective module 2 SN:910652 with new module SN:711199 Genexpert machine Cleaning/ preventive maintenance Checking of errors Data collection | Perform Genexpert machine module replacement/repair Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert 6-color module | Cepheid case:1562143 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964856390711249482/5964856390711249482_signature_28.png | Hamza Datu C. Acob | https://www.jotform.com/uploads/macarewecare/222927978372067/5964856390711249482/5964856390711249482_signature_31.png
| 2024-07-11 | 2024-07-11 11:31 | Hilongos Rural Health Unit | Hilongos, Leyte | 8 | Leyte | Laboratory | YES | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28869 | PBSP | leovillaflor@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 668 |  | 6C GXIV-4D | 838143 | All modules experiencing Error 2126 | -Reset internal power supply connections -Reset LAN cable connections -Check AC voltage supply of laboratory, 220V -Check GeneXpert and peripheral connections to UPS, stable -Observed machine while specimen running | -Preventive Maintenance -Upgrade XpertCheck Software to version 1.5b | Sample run was completed with no problems occurring despite power fluctuations. All modules are functional. For further observation. | Recommended for AVR usage since laboratory is prone to frequent fluctuations (occurred every 10-20mins) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964848924835605117/5964848924835605117_signature_28.png | Leonides Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5964848924835605117/5964848924835605117_signature_31.png
| 2024-07-11 | 2024-07-10 15:30 | The Medical City | Ortigas Avenue, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28868 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 39 ASSAY: Xpert CT/NG, COUNT: 79 ASSAY: Xpert HBV VL, COUNT: 588 ASSAY: Xpert HPV, COUNT: 444 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5279 ASSAY: Xpert HCV VL, COUNT: 66 ASSAY: Xpert HIV-1 Viral Load, COUNT: 639 ASSAY: Others, COUNT: 794 |  | GX-XVI-D-10C | 110015766 | For maintenance | Performed maintenance. | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292803725543/5964789292803725543_signature_28.png | Kristine Chelle Parilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292803725543/5964789292803725543_signature_31.png
| 2024-07-11 | 2024-07-10 10:46 | Taytay Rural Health Unit 5-B | Ynares Ave, Lupang Arenda, Taytay, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28867 | PBSP | zenresurreccion@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1545 | 2024-07-10 | GX-IV-D-6C | 838136 | Module B3 for replacement | Replaced module B3 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B3 Generated reports | Unit is operational. All modules operational. |  | Old: B3/806713 New:B3/210047087 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292804902321/5964789292804902321_signature_28.png | Zenilynn Bobiles | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292804902321/5964789292804902321_signature_31.png
| 2024-07-11 | 2024-07-10 14:17 | The Medical City | Ortigas Avenue, Pasig City | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28866 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292801476937/5964789292801476937_signature_28.png | Gilian dominique dacanay | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292801476937/5964789292801476937_signature_31.png
| 2024-07-11 | 2024-07-10 13:37 | The Medical City | Ortigas Avenue, Pasig City | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28865 | COMMERCIAL | bloodbank@themedicalcity.com | 20 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292807553579/5964789292807553579_signature_28.png | Tio, Daryl Paolo | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292807553579/5964789292807553579_signature_31.png
| 2024-07-11 | 2024-07-09 10:15 | Old Lower Bicutan Health Center | M.L Quezon St., Bicutan Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28864 | PBSP | ellainemagtoto35@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3089 | 2024-07-09 | GX-IV-D-6C | 838858 | Modules B1, B2 and B3 for replacement | Replaced modules B1, B2 and B3 | Performed maintenance Replaced modules B1, B2 and B3. Generated reports | Unit is operational. All modules operational. |  | Old: B1/863216, B2/863322, B3/863340 New: B1/711219, B2/711524, B3/717660 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292808418262/5964789292808418262_signature_28.png | Ellaine Magtoto | https://www.jotform.com/uploads/macarewecare/222927978372067/5964789292808418262/5964789292808418262_signature_31.png
| 2024-07-11 | 2024-07-11 10:55 | Anahawan District Hospital | Anahawan Southern leyte | 8 | Southern Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28863 | DOH | laboratoryadh@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 471 | 2024-07-11 | GXIV | 838548 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964759559168892461/5964759559168892461_signature_28.png | MARY GRACE SANTIAGO | https://www.jotform.com/uploads/macarewecare/222927978372067/5964759559168892461/5964759559168892461_signature_31.png
| 2024-07-11 | 2024-07-11 09:13 | CEBU DOCTORS' UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28862 | COMMERCIAL | cduhcovidpoct@gmail.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 73 ASSAY: Xpert CT/NG, COUNT: 197 ASSAY: Xpert HBV VL, COUNT: 171 ASSAY: Xpert HPV, COUNT: 19 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1167 ASSAY: Xpert HCV VL, COUNT: 146 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 |  | GX-IV | 110008524 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964698829225332915/5964698829225332915_signature_28.png | Queenie Valencia | https://www.jotform.com/uploads/macarewecare/222927978372067/5964698829225332915/5964698829225332915_signature_31.png
| 2024-07-10 | 2024-07-10 12:50 | Davao City Chest Center | Villa Abrille Street, Davao City | 11 | Davao del Sur | TB-DOTS | YES | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28861 | WHO | dcclabdept@gmail.com | 23.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 141 |  | GX-IV 6C 4D | 804835 | Module B3 and B4 for Repair | Replaced the Defective Module SN:6354866 with new Module SN:807102 Replaced the Defective Module SN: 657440 with new Module SN:862401 Generate IQR, Assay Stat and System Log | Perform Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5964009127179168771/5964009127179168771_signature_28.png | Crizz valerey bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5964009127179168771/5964009127179168771_signature_31.png
| 2024-07-10 | 2024-07-10 11:52 | Mangatarem District Hospital | Casilagan, Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | INSTALLATION Relocation | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28860 | PBSP | sweetsellveras1@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1711 | 2024-07-10 | GXIV-4-D-6C | 807598 | For Installation of GeneXpert System | #NAME? | Performed Installation of GeneXpert System | Unit is Fully Operational | GeneXpert System originally from Western Pangasinan District Hospital | GeneXpert System | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963935304682887691/5963935304682887691_signature_28.png | Sweet Sell C. Veras | https://www.jotform.com/uploads/macarewecare/222927978372067/5963935304682887691/5963935304682887691_signature_31.png
| 2024-07-10 | 2024-07-10 09:33 | Dr. Amado Diaz Provincial Foundation Hospital | Midsayap, North Cotabato | 12 | Cotabato | Laboratory | NO | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28859 | MACARE-HBDC | charlenemae_amatac@yahoo.com | 25.8 |  |  | Gx-IV (10C) | 110024249 | N/A | Installed: Genexpert machine SN: 110024249 CPU: CPH0224F49301 Monitor: Dell SN:CN0WMYC5FCC00398G2YXA05 | Perform Genexpert machine Installation | All modules operational and accessories |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963860197316187486/5963860197316187486_signature_28.png | Charlene Mae A. Amatac | https://www.jotform.com/uploads/macarewecare/222927978372067/5963860197316187486/5963860197316187486_signature_31.png
| 2024-07-10 | 2024-07-10 08:53 | Tarlac City Health Center 2 | Brgy. Matatalaib, Tarlac City | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28858 | DOH | clarizannebeltran1996@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 518 | 2024-07-04 | GXIV-D | 831917 | UPS for replacement | Repair Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed UPS replacement Performed functionality test Gathered reports | Unit is operational |  |  | 49051 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963823065324931445/5963823065324931445_signature_28.png | Clariz Anne M. Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5963823065324931445/5963823065324931445_signature_31.png
| 2024-07-09 | 2024-07-09 14:59 | Zamboanga City Health Office | Pettit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28857 | PBSP | floresteresitaj25@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 243 | 2024-07-09 | GX 4 | 818902 | Module 1 and 3 for replacement | Replaced module 714798 and 714711 with Module 732575 and 737061 respectively. Conducted Preventive Maintenence. | Module Replacement and Preventive Maintenence | All module are available. UPS for replacement and monitor | UPS and monitor are for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963248790003359820/5963248790003359820_signature_28.png | Teresita Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5963248790003359820/5963248790003359820_signature_31.png
| 2024-07-09 | 2024-07-09 14:59 | General Tinio Rural Health Unit | Poblacion, General Tinio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28856 | DOH | jadelyn_pajarillaga@yahoo.com | 24 |  |  | GXIV-6C | 831695 | CPU not booting | Troubleshoot CPU | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963180056015213298/5963180056015213298_signature_28.png | Jadelyn B. Pajarillaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5963180056015213298/5963180056015213298_signature_31.png
| 2024-07-09 | 2024-07-09 13:32 | San Miguel District Hospital | San Miguel, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28855 | PBSP | smdh_laboratory@yahoo.com | 27 |  | 2024-07-09 | GXIV | 807604 | Module A4 for replacement | Replaced A4 | #NAME? | Unit is operational |  | GeneXpert Module SN: 681162 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963153173115259457/5963153173115259457_signature_28.png | Florencia Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5963153173115259457/5963153173115259457_signature_31.png
| 2024-07-09 | 2024-07-09 12:18 | Belmonte Health Center | Sobriedad St. Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28854 | PBSP | maricar.lira@gmail.com | 25 |  | 2024-07-03 | GXIV | 808619 | Unit for PM. For module replacement of module B2. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B2 (784427) in replacement to def. B2 (210059617). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963121606252285156/5963121606252285156_signature_28.png | Maricar Lira | https://www.jotform.com/uploads/macarewecare/222927978372067/5963121606252285156/5963121606252285156_signature_31.png
| 2024-07-09 | 2024-07-09 12:20 | Belmonte Health Center | Sobriedad St. Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28853 | PBSP | maricar.lira@gmail.com | 25 |  | 2024-07-03 | GXIV | 814376 | Unit for PM. For module replacement of module A1. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A1 (759993) in replacement to def. A1 (707949). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963121606253401093/5963121606253401093_signature_28.png | Maricar Lira | https://www.jotform.com/uploads/macarewecare/222927978372067/5963121606253401093/5963121606253401093_signature_31.png
| 2024-07-09 | 2024-07-04 10:26 | Candon City Health Office | Mc Arthur Highway, Candon | 1 | Ilocos Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28852 | DOH | candoncholaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 599 | 2024-06-27 | GXIV | 838273 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck kit. | DR#11658 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963121606257903004/5963121606257903004_signature_28.png | Marites B. Cortes | https://www.jotform.com/uploads/macarewecare/222927978372067/5963121606257903004/5963121606257903004_signature_31.png
| 2024-07-09 | 2024-07-09 13:19 | Gapan Rural Health Unit II | San Nicolas, Gapan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | emmaliwaggutierrez@gmail.com |  | 28851 | PBSP | krystellebautista90@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 801 |  | GXIV-6C | 814182 | Module 1 Error UPS not working | Performed Module Replacement B1(690683 - 714204) Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963118416013081356/5963118416013081356_signature_28.png | Krystelle Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5963118416013081356/5963118416013081356_signature_31.png
| 2024-07-09 | 2024-07-09 12:54 | Ormoc City Health Office | Ormoc City, Leyte | 8 | Leyte | Laboratory | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28850 | PBSP | theresesaladaga@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2850 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 |  | 6C GXIV-4D | 804253 | None | #NAME? | -Preventive Maintenance -Install delivered fan filter -Upgrade XpertCheck Software to version 1.5b | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963115815759927209/5963115815759927209_signature_28.png | Ma. Therese A. Saladaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5963115815759927209/5963115815759927209_signature_31.png
| 2024-07-09 | 2024-07-05 12:31 | West Visayas State University Medical Center | Iloilo | 6 | Iloilo | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28849 | PBSP | tvanrlujan@gmail.com | 29 |  |  | Gx 4 | 804699 | Defective Module b4 210068572 | New module install B4 710004 Clean Fan Filter Gather data IQR | New module install B4 710004 Clean Fan Filter Gather data IQR | Sytems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963115134225854046/5963115134225854046_signature_28.png | Van Ryan R. Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5963115134225854046/5963115134225854046_signature_31.png
| 2024-07-09 | 2024-07-09 11:49 | Manila Doctors Hospital | United Nations Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28848 | PBSP | lab.bacteriology@maniladoctors.com.ph | 19 | ASSAY: Xpert CT/NG, COUNT: 96 ASSAY: Xpert HBV VL, COUNT: 89 ASSAY: Xpert HPV, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1276 ASSAY: Xpert HCV VL, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 | 2021-06-23 | GX IV 10C-D | 110001648 | Unit due for maintenance | Performed preventive maintenance Performed Xpertcheck Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck All modules passed | Unit is operational All modules working |  | XpertCheck Kit | lot 1001120421 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963075030413319105/5963075030413319105_signature_28.png | Ryamin F. Buentiempo | https://www.jotform.com/uploads/macarewecare/222927978372067/5963075030413319105/5963075030413319105_signature_31.png
| 2024-07-09 | 2024-07-09 10:32 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB-Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28847 | DOH | sajecopooh20@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1272 |  | 6C GXIV-4D | 838555 | None | None | -Preventive Maintenance -Upgrade XpertCheck Software to version 1.5b | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963019777591845578/5963019777591845578_signature_28.png | Sarah Jean R. Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5963019777591845578/5963019777591845578_signature_31.png
| 2024-07-09 | 2024-07-09 10:18 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | TROUBLESHOOTING | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28846 | COMMERCIAL | pshclinlab@gmail.com | 20 |  |  | MCS+ | 11H307 | CUFF RELATED ISSUE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5963009274222181393/5963009274222181393_signature_28.png | Roda Fatima Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5963009274222181393/5963009274222181393_signature_31.png
| 2024-07-09 | 2024-07-09 09:58 | Tuburan Rural Health Unit | Poblacion | 7 | Cebu | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 28845 | PBSP | marielgracia20@gmail.com | 23 |  | 2024-07-01 | Hx4 | 815485 | Module B2 disable due to error reported error Code 2120. | Preventive Maintenance | . Perform Preventive Maintenance . Cleaning all modules and cables inside the macine. . Cleaning filter. . Plunger rod maintenance | All module's functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962999324229968686/5962999324229968686_signature_28.png | Mareil Gracia | https://www.jotform.com/uploads/macarewecare/222927978372067/5962999324229968686/5962999324229968686_signature_31.png
| 2024-07-09 | 2024-06-19 13:00 | Bicol Region General Hospital and Geraitric Medical Center | Cabusao, Camarines Sur | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28844 | PBSP | jlvdvera_rmt@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 35 |  | GX4-10C | 110019624 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962656536491037685/5962656536491037685_signature_28.png | Jessie Lowell de vera | https://www.jotform.com/uploads/macarewecare/222927978372067/5962656536491037685/5962656536491037685_signature_31.png
| 2024-07-08 | 2024-07-05 14:50 | Binangonan Rural Health Unit | Osmena St., Binangonan, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28843 | PBSP | dhayanara.zapanta@gmail.com | 24 |  |  | GX-IV-D-6C | 807605 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. | Unit from Binangonan RHU I |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962360643613072123/5962360643613072123_signature_28.png | Dhayanara Zapanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5962360643613072123/5962360643613072123_signature_31.png
| 2024-07-08 | 2024-07-05 14:44 | Binangonan Rural Health Unit III | Osmena St., Binangonan, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28842 | PBSP | dhayanara.zapanta@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 535 | 2024-07-05 | GX-IV-D-6C | 846045 | Module B4 for replacement | Replaced module B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B4 Generated reports | Unit is operational. All modules operational. |  | Old: B4/210100341 New: B4/666188 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962360643618004830/5962360643618004830_signature_28.png | Dhayanara Zapanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5962360643618004830/5962360643618004830_signature_31.png
| 2024-07-08 | 2024-07-05 14:48 | Binangonan Rural Health Unit | Osmena St., Binangonan, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28841 | PBSP | dhayanara.zapanta@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2274 |  | GX-IV-D-6C | 838199 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. | Unit from Binangonan RHU I |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962360643617194325/5962360643617194325_signature_28.png | Dhayanara Zapanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5962360643617194325/5962360643617194325_signature_31.png
| 2024-07-08 | 2024-07-04 13:20 | Sudipen Rural Health Unit | Poblacion, Sudipen, La Union | 1 | La Union | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28840 | DOH | piaramuj@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 | 2024-06-27 | GXIV-D-6C | 838270 | Unit is due for XpertCheck. | Calibration. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed XpertCheck. Performed Self-test. Generate Reports. | Module 2 with Temperature Error but Unavailable in XpertCheck. Module 1 Passed. Module 3 Neely Replaced Module. Module 4 Service Required. |  | XperrCheck Kit | DR 11657 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962337133617084917/5962337133617084917_signature_28.png | Pia Jumar T. Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/5962337133617084917/5962337133617084917_signature_31.png
| 2024-07-08 | 2024-07-05 08:51 | Vitali Lying Center | Vitali, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION MAINTENANCE Replacement of System Unit | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28839 | PBSP | vitalihealthdistrict@gmail.com | 28 |  | 2024-07-05 | GX 4 | 838140 | System unit unable to turn on. | Replaced System unit and conducted Preventive maintenance | Replaced System unit and conducted Preventive maintenance | All modules are available and system unit is working well. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962308770009251271/5962308770009251271_signature_28.png | Carmelo E. Umandal | https://www.jotform.com/uploads/macarewecare/222927978372067/5962308770009251271/5962308770009251271_signature_31.png
| 2024-07-08 | 2024-07-03 15:38 | President Ramon Magsaysay Memorial Hospital | Iba Zambales | 3 | Zambales | PCR LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28838 | PBSP | prmmhmoleclab@gmail.com | 21 | ASSAY: Xpert SARS-CoV-2, COUNT: |  | GXIV-D | 834622 | B4:784824 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational B4:781528new module UPS for replacement (defective) | Defective UPS |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962126773614991395/5962126773614991395_signature_28.png | Bianca Yzabel Yap | https://www.jotform.com/uploads/macarewecare/222927978372067/5962126773614991395/5962126773614991395_signature_31.png
| 2024-07-08 | 2024-07-05 14:59 | Esperanza Rural Health Unit | Poblacion-Iligan Rd., Poblacion, Esperanza | 5 | Masbate | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28837 | DOH | floramarmoncada@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1392 |  | GX4-6C | 838268 | GX4-6C for Calibration | XpertCheck Result: B1/210068697 - service required B2/879219 - newly replaced modulr B3/210068696 - service required B4/668421 - passed Cleaned cartridge bays, plunger rod, icore, replaced fan filter Generate report | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962099956496614691/5962099956496614691_signature_28.png | Floramar M. Antipolo | https://www.jotform.com/uploads/macarewecare/222927978372067/5962099956496614691/5962099956496614691_signature_31.png
| 2024-07-08 | 2024-07-04 16:58 | Aroroy Rural Health Unit | Aroroy, Masbate | 5 | Masbate | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28836 | PBSP | morenorhenaluz@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 369 |  | GX4-10C | 110019625 | Frequent 5007 error | Performed sample test run(aborted) Collected data(gxx file) Cleaned cartridge bays, plunger rod, icore Generate report | Performed Troubleshooting Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962099616496191117/5962099616496191117_signature_28.png | Rhena M. Etcobanez | https://www.jotform.com/uploads/macarewecare/222927978372067/5962099616496191117/5962099616496191117_signature_31.png
| 2024-07-08 | 2024-07-04 15:22 | Milagros Rural Health Unit | Milagros, Masbate | 5 | Masbate | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28835 | PBSP | gracemarymesa1029@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1455 |  | GX4-6C | 838244 | Monitor grid lines appearing on screen | Firmly connect GVA cable to monitor and CPU Collected data(video) of grid lines appearing on screen Cleaned cartridge bays, plunger rod, icore Generate report | Performed Troubleshooting Performed Preventive Maintenance | Unit is Operational Grid lines on monitor(still functional) | Monitor request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962097056498334019/5962097056498334019_signature_28.png | Mary Grace C.Mesa | https://www.jotform.com/uploads/macarewecare/222927978372067/5962097056498334019/5962097056498334019_signature_31.png
| 2024-07-08 | 2024-07-06 12:08 | Philippine Red Cross (Laguna chapter) | P.Guevarra st.,Sta Cruz Laguna | 4A | Laguna | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28834 | MACARE-HBDC | lagunablood@redcross.org.ph | 27 |  |  | MCS | 02J046 | Pinch Valve error Loose valve tubing | Re tighten tubings Checked leakages Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge,weigher Performed functionality test Performed simulation running | Perfomed troubleshooting | Operational | Work with Monsour Doctor |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5962045590319380634/5962045590319380634_signature_28.png | John Paul Carmona | https://www.jotform.com/uploads/macarewecare/222927978372067/5962045590319380634/5962045590319380634_signature_31.png
| 2024-07-07 | 2024-07-05 13:30 | Sta. Lucia Health Center | Sta. Lucia, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28833 | DOH | josiedingle2004@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1968 |  | GX IV 6C D | 831901 | Keyboard keys were hard to press and sometimes it gets stucked when pressed | Replaced keyboard Performed preventive maintenance | Installed new keyboard Checked and cleaned cartridge bays, plunger rod, icore slot and filter Generated report | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5961666771016688194/5961666771016688194_signature_28.png | Josephine Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5961666771016688194/5961666771016688194_signature_31.png
| 2024-07-07 | 2024-07-05 09:40 | NGC Super Health Center | Pilot Drive, Mangahan, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28832 | DOH | tmcolisao@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 666 |  | GX IV 6C-D | 831586 | Unit due for maintenance | Performed Preventive maintenance Performed reliability test Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot, and filter Performed reliability test Module A1 failed reliability test. Error 2035 Informed the end user to monitor module A1 | unit is operational all modules working Module A1 for monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5961661571012757917/5961661571012757917_signature_28.png | Mila N Baldovizo | https://www.jotform.com/uploads/macarewecare/222927978372067/5961661571012757917/5961661571012757917_signature_31.png
| 2024-07-07 | 2024-07-05 15:24 | Urdaneta City Health Office | Alexander Street, Urdaneta City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28831 | PBSP | Urdcho1lab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1411 | 2024-07-05 | GXIV-4-D-6C | 805109 | For Installation/Relocation | #NAME? | Performed Installation/Relocation | Unit is Fully Operational | From Urdaneta District Hospital | GeneXpert System | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5961361960625638951/5961361960625638951_signature_28.png | REVILYN C. GERONA | https://www.jotform.com/uploads/macarewecare/222927978372067/5961361960625638951/5961361960625638951_signature_31.png
| 2024-07-07 | 2024-07-05 10:18 | UMINGAN COMMUNITY HOSPITAL | Umingan Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28830 | DOH | jeffreyvrumbaoa@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1275 |  | GXIV-4-D-6C | 838473 | Due for Preventive Maintenance | #NAME? | Performed Preventive Maintenance Performed functionality test | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5961360830624795235/5961360830624795235_signature_28.png | Jeffrey V. Rumbaoa | https://www.jotform.com/uploads/macarewecare/222927978372067/5961360830624795235/5961360830624795235_signature_31.png
| 2024-07-05 | 2024-07-05 15:52 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28829 | MACARE-HBDC | clstan1998@gmail.com | 28 |  |  | MCS + | 11H204 | DPM out of tolerance | Checked unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Adjust DPM and SPM value Performed functionality test Performed simulation running | Performed trubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959882372395282754/5959882372395282754_signature_28.png | Cayle Lester S. Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5959882372395282754/5959882372395282754_signature_31.png
| 2024-07-05 | 2024-07-05 11:41 | St lukes Medical Center Global City | Global City | NCR | NCR 4th - Taguig | Laboratory | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28828 | COMMERCIAL | ci.bgc@stlukes.com.ph | 25 |  |  | GX 4 D(10) | 120002383 | For installation | Installed unit | Performed installation Import ADF files | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959881022395558505/5959881022395558505_signature_28.png | Sapla, Ma Grace Angelica | https://www.jotform.com/uploads/macarewecare/222927978372067/5959881022395558505/5959881022395558505_signature_31.png
| 2024-07-05 | 2024-07-04 14:37 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28827 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1425 ASSAY: Others, COUNT: 89 ASSAY: Xpert HBV VL, COUNT: 1032 ASSAY: Xpert HPV, COUNT: 101 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2331 ASSAY: Xpert SARS-CoV-2, COUNT: 13588 ASSAY: Xpert HCV VL, COUNT: 134 |  | GX16 D | 110005587 | A1 reported already C2 reported already D2,A2,A4 reported already confirmed for replacement | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | 5 modules not in use A1.A2.A4.C2.D2 11 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959838442399911724/5959838442399911724_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5959838442399911724/5959838442399911724_signature_31.png
| 2024-07-05 | 2024-07-04 14:04 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28826 | COMMERCIAL | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Others, COUNT: 6 ASSAY: Xpert HPV, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 147 ASSAY: Xpert HCV VL, COUNT: 7 |  | GX16 D | 804476 | Due for maintenance, A1 A2,B1,B3,C1,C2,C3,C4 disabled due of continous error loss communication error B4 temp error 4009,4012,4017,4010 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | A1 A2,B1,B3,C1,C2,C3,C4 disabled due of continous error loss communication error B4 temp error 4009,4012,4017,4010 A3,A4,B2 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959838042393724739/5959838042393724739_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5959838042393724739/5959838042393724739_signature_31.png
| 2024-07-05 | 2024-07-04 13:35 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28825 | MACARE-HBDC | molecularlabnkti@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 47 ASSAY: Others, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 18 ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 44 |  | GX4 D | 838524 | Due for maintenance B1,B3,B4 reported already confirmed for replacement | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B1 and B3 disabled B2 and B4 in use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959837574315785851/5959837574315785851_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5959837574315785851/5959837574315785851_signature_31.png
| 2024-07-05 | 2024-07-05 15:52 | Kasibu Rural Health Unit | LGU Compound, Kasibu, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28824 | PBSP | mhokasibulaboratory@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 306 |  | GXII6C | 834802 | Xpertcheck error 3005 | Reinstalling xpertcheck | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959759841614060427/5959759841614060427_signature_28.png | Halal Dumangeng | https://www.jotform.com/uploads/macarewecare/222927978372067/5959759841614060427/5959759841614060427_signature_31.png
| 2024-07-05 | 2024-07-05 09:56 | Ospital ng Lungsod ng San Jose Del Monte | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28823 | PBSP | arvalerio21@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1785 |  | GXIV | 807624 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Module B3 for follow up report |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959756730816609919/5959756730816609919_signature_28.png | Arnold Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5959756730816609919/5959756730816609919_signature_31.png
| 2024-07-05 | 2024-07-05 11:40 | Meycauayan City Health Unit VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28822 | DOH | ronamarieiris_25@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1696 ASSAY: Xpert HIV-1 Viral Load, COUNT: 118 | 2024-07-05 | GXIV | 831904 | Module A3 for replacement | Replaced A3 | #NAME? | Unit is operational |  | GeneXpert Module SN: 604539 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959756290813149066/5959756290813149066_signature_28.png | Rona Marie Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5959756290813149066/5959756290813149066_signature_31.png
| 2024-07-05 | 2024-06-21 11:21 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28821 | PBSP | bmcgenexpertlab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 1859 |  | GXIV | 834623 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959755930816772442/5959755930816772442_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5959755930816772442/5959755930816772442_signature_31.png
| 2024-07-05 | 2024-07-03 16:19 | Detoxicare Molecular Diagnostics Laboratory | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | detoxicare.clinical@gmail.com |  | 28820 | CONSORTIUM | detoxicare.clinical@gmail.com | 29 |  |  | GXIV 10C | 110018312 | Module B1 for troubleshooting | NA | #NAME? | Unit is operational Module B1 for reporting Average room temperature is 29C (30C threshold) Room average temperature is recommended below 30C to prevent temperature related errors |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959754980813336525/5959754980813336525_signature_28.png | Cyrill Tataro | https://www.jotform.com/uploads/macarewecare/222927978372067/5959754980813336525/5959754980813336525_signature_31.png
| 2024-07-05 | 2024-07-05 11:40 | Meycauayan City Health Unit VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28818 | DOH | ronamarieiris_25@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1696 ASSAY: Xpert HIV-1 Viral Load, COUNT: 118 | 2024-07-05 | GXIV | 831904 | Module A3 for replacement | Replaced A3 | #NAME? | Unit is operational |  | GeneXpert Module SN: 604539 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959754710813760794/5959754710813760794_signature_28.png | Rona Marie Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5959754710813760794/5959754710813760794_signature_31.png
| 2024-07-05 | 2024-07-05 09:56 | Ospital ng Lungsod ng San Jose Del Monte | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28819 | PBSP | arvalerio21@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1785 |  | GXIV | 807624 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Module B3 for follow up report |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959754710817963532/5959754710817963532_signature_28.png | Arnold Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5959754710817963532/5959754710817963532_signature_31.png
| 2024-07-05 | 2024-07-05 12:11 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | ON-SITE ASSISTANCE Relocation of GeneXpert ,Relocation of GeneXpert | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | loreyleinela@gmail.com |  | 28817 | DOH | dy_rosemarie71@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1902 |  | GX-IV 6C 4D | 838401 | For Relocation of GeneXpert Machine | Relocate the Machine back to the Newly Renovated Laboratory | Perform Relocation of The GeneXpert Machine | Module B1 and B2 : Disabled waiting for parts from Cepheid Only Module B3 and B4 are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959699227919322468/5959699227919322468_signature_28.png | Rosemarie Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/5959699227919322468/5959699227919322468_signature_31.png
| 2024-07-05 | 2024-07-05 13:55 | Felipe Abrigo Memorial Hospital | Guiuan Eastern Samar | 8 | Eastern Samar | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28816 | PBSP | famh_guiuan@yahoo.com | 28 |  | 2024-07-05 | GXIV | 815271 | Due for Calibration | n/a | #NAME? | Unit is operational |  | xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959686420815004389/5959686420815004389_signature_28.png | rogelio bagania | https://www.jotform.com/uploads/macarewecare/222927978372067/5959686420815004389/5959686420815004389_signature_31.png
| 2024-07-05 | 2024-07-05 13:36 | BAIS DISTRICT HOSPITAL | BAIS NEGROS ORIENTAL | 7 | Negros Oriental | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28815 | PBSP | laboratorybdh@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 577 |  | GX-IV | 838857 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959673908453634485/5959673908453634485_signature_28.png | EVAN JUN C. PACURIB | https://www.jotform.com/uploads/macarewecare/222927978372067/5959673908453634485/5959673908453634485_signature_31.png
| 2024-07-05 | 2024-07-05 13:08 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28814 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1946 |  | 6C GXIV-4D | 838561 | Due for PM | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959660093071764885/5959660093071764885_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5959660093071764885/5959660093071764885_signature_31.png
| 2024-07-05 | 2024-07-05 11:53 | La Castellana Rural Health Unit | La Castellana, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 28813 | DOH | milouepascual@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 791 |  | GX-4 | 838483 | Due for Xpertcheck/Calibration. | #NAME? | - Perform Xpertcheck in all 4 modules. - Install Dell 21" LCD monitor. | #NAME? | Diaabled module B3 due to runtime errors. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959659729053798748/5959659729053798748_signature_28.png | Mira Luna E. Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/5959659729053798748/5959659729053798748_signature_31.png
| 2024-07-05 | 2024-07-05 11:37 | Cardinal Santos Medical Center | Wilson St., San Juan | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 28812 | COMMERCIAL | robileanchin@gmail.com | 24 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959605146511506518/5959605146511506518_signature_28.png | Robi Lean Ching | https://www.jotform.com/uploads/macarewecare/222927978372067/5959605146511506518/5959605146511506518_signature_31.png
| 2024-07-05 | 2024-07-05 11:06 | Eastern Pangasinan District Hospital | Tayug Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28811 | PBSP | epdhlaboraory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 613 | 2024-07-03 | GXIV-D | 815394 | B4:696024 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is fully operational |  | 1 genexpert module | 49022 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959593191714183734/5959593191714183734_signature_28.png | Fergyl Escaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5959593191714183734/5959593191714183734_signature_31.png
| 2024-07-05 | 2024-07-05 11:06 | Cardinal Santos Medical Center | Wilson St., San Juan City | NCR | NCR 2nd - San Juan | Microbiology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 28810 | COMMERCIAL | egatchalian@csmc.ph | 24 | ASSAY: Xpert HBV VL, COUNT: 141 ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1095 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports Generated gxx file | Unit is operational. All modules operational. | Module B3 has error 1002 but under observation. Module B4 has error 2014 also under observation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959589656514633901/5959589656514633901_signature_28.png | Ellaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5959589656514633901/5959589656514633901_signature_31.png
| 2024-07-05 | 2024-07-04 15:42 | Marikina City Health Office | Shoe Ave., Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28809 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1676 ASSAY: Xpert HIV-1 Viral Load, COUNT: 330 | 2024-07-04 | GX-IV-D-6C | 805211 | Module A1 for replacement | Replaced module A1 | Replaced module A1 Functionality test Generated reports | Unit is operational. All modules operational. |  | Old: A1/650238 New: A1/692508 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959550680082517139/5959550680082517139_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5959550680082517139/5959550680082517139_signature_31.png
| 2024-07-05 | 2024-07-03 12:06 | Southern Philippines Medical Center | JP Laurel Avenue Bajada Davao City | 11 | Davao del Sur | RHIVDA, POPUP LABORATORY | NO | INSTALLATION | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28808 | UNOPS | spmcrhivda@gmail.com | 24.6 |  | 2024-06-20 | GX-IV 10C 4D | 110023506 | For Installation of GeneXpert | Clean Surface Area Check the line voltages Install the UPS with SN: AT1K231111ST0024 Install the GeneXpert CPU System Unit with SN: CPH1023F46916 Install the GeneXpert 10 color Machine with SN:110023506 Install the assays to be performed Setup Enduser Account | Perform Installation | Unit Successfully Installed Machine is operational |  |  | DR 11649 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959511256321056075/5959511256321056075_signature_28.png | Flordeliza B Go RMT MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/5959511256321056075/5959511256321056075_signature_31.png
| 2024-07-05 | 2024-07-05 08:20 | Bacolod City Health Office | Bacolod City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 28807 | PBSP | jacabiao.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 7040 |  | GX-4 | 804838 | Defective module 720028/C1 | Cleaned optics, module bay areas, pkunger rod, filter. | Replaced module C1 with new module SN 860168. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959482758653991801/5959482758653991801_signature_28.png | Jenzel Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5959482758653991801/5959482758653991801_signature_31.png
| 2024-07-05 | 2024-07-05 08:10 | Potrero Health Center | Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28806 | PBSP | ailyn.santos19@yahoo.com | 25 |  | 2024-07-05 | GXIV | 804715 | Modules B2, B3, B4 for replacement | Replaced modules B2, B3, B4 | #NAME? | Unit is operational |  | GeneXpert Module SN: 210006959, 749017, 790993 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959481085611282837/5959481085611282837_signature_28.png | Ailyn Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5959481085611282837/5959481085611282837_signature_31.png
| 2024-07-05 | 2024-07-05 06:53 | MEDI LINX LABORATORY INC. | St. Elizabeth Hospital Inc. General Santos City, South Cotabato | 12 | South Cotabato | Micrology | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28805 | COMMERCIAL | sofia.maclang@medilinx.com.ph | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 268 |  | Gx-IV (6-color) | 831313 | Module 1 and 3 multiple temperature errors Module 4 Ultrasonic Problem | Genexpert machine Cleaning/preventive maintenance Checking of errors Genexpert Problem/errors reporting | Perform Genexpert machine Preventive maintenance | 1 operational modules (module 2) Module 1, 3, and 4 requesting for module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5959439867147414680/5959439867147414680_signature_28.png | Sofia Maclang, Unit Head - Main Test ll | https://www.jotform.com/uploads/macarewecare/222927978372067/5959439867147414680/5959439867147414680_signature_31.png
| 2024-07-04 | 2024-07-04 18:32 | AYUNGON RURAL HEALTH UNIT | AYUNGON NEGROS ORIENTAL | 7 | Negros Oriental | LABORATORY | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28804 | PBSP | kesheiadgantalao@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-IV | 805275 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958985177814170537/5958985177814170537_signature_28.png | Kesheia Gantalao, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5958985177814170537/5958985177814170537_signature_31.png
| 2024-07-04 | 2024-07-04 15:03 | Tagum City Health Office | Mabini Street, Tagum City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28803 | DOH | tagumcholab.mabini@gmail.com | 26.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1828 |  | GX-IV 6C 4D | 838553 | Modules B1, B2 and B4 encounter multiple temperature related errors | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform Xpertcheck Calibration on Modules 1, 2, and 4 | Perform Preventive Maintenance and Functionality Test of the Modules | Modules 1 and 4 For Service Required Module 2 Passed Module 2 for under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958908468348422198/5958908468348422198_signature_28.png | Cristina Tiara A. Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5958908468348422198/5958908468348422198_signature_31.png
| 2024-07-04 | 2024-07-04 16:14 | Baggao Rural Health Unit | San Jose, Baggao, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28802 | DOH | baggaorhulaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 484 |  | GXIV-6C | 838531 | Module 3 Error, CPU sometimes not powering on, Due for Calibration | Performed Calibration (B3/862506 - Service Required) (B1 , B4 - PASSED) Performed Troubleshooting on CPU (Defective Power Supply) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958906767111509241/5958906767111509241_signature_28.png | Jan Anthony Eslabra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958906767111509241/5958906767111509241_signature_31.png
| 2024-07-04 | 2024-07-04 14:42 | Sen. Gerardo M. Roxas Memorial District Hospital | Dao, Capiz | 6 | Capiz | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28801 | PBSP | bermendirowena@gmail.com | 28 |  |  | Gx 4 | 815280 | Defective Module B2 700369 | New Module B2 installed 609798 Clean Fan and Fan Filter Gather data IQR Xpertcheck module B3 | New Module B2 installed 609798 Clean Fan and Fan Filter Gather data IQR | Systems Operational Xpertcheck B3 service require |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958904465391376421/5958904465391376421_signature_28.png | Rowena L. Bermendi | https://www.jotform.com/uploads/macarewecare/222927978372067/5958904465391376421/5958904465391376421_signature_31.png
| 2024-07-04 | 2024-07-04 11:39 | Marikina City Health Office | Shoe Ave., Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28800 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 402 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110020361 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958853688571908150/5958853688571908150_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958853688571908150/5958853688571908150_signature_31.png
| 2024-07-04 | 2024-07-04 14:46 | Amang Rodriguez Memorial Medical Center | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28799 | PBSP | pathologymole.armmc@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 2548 ASSAY: Others, COUNT: 13 | 2024-07-04 | GX-XVI-D-6C | 844113 | Module A3 for replacement | Replaced module A3 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module A3 Generated reports | Unit is operational. All modules operational. |  | Old: A3/891761 New: A3/865095 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958853528573624742/5958853528573624742_signature_28.png | Racquel Verzosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5958853528573624742/5958853528573624742_signature_31.png
| 2024-07-04 | 2024-07-04 14:01 | Amang Rodriguez Memorial Medical Center | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28798 | PBSP | maricel_c27@yahoo.cim | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4487 |  | GX-IV-D-6C | 807982 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958853388579080115/5958853388579080115_signature_28.png | Maricel S Calderon | https://www.jotform.com/uploads/macarewecare/222927978372067/5958853388579080115/5958853388579080115_signature_31.png
| 2024-07-04 | 2024-07-04 14:46 | Amang Rodriguez Memorial Medical Center | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28797 | PBSP | pathologymole.armmc@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 2548 ASSAY: Others, COUNT: 13 | 2024-07-04 | GX-XVI-D-6C | 844113 | Module A3 for replacement | Replaced module A3 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module A3 Generated reports | Unit is operational. All modules operational. |  | Old: A3/891761 New: A3/865095 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852988573280018/5958852988573280018_signature_28.png | Racquel Verzosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852988573280018/5958852988573280018_signature_31.png
| 2024-07-04 | 2024-07-04 11:41 | Marikina City Health Office | Shoe Ave., Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28796 | MACARE-HBDC | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 224 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852918573138581/5958852918573138581_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852918573138581/5958852918573138581_signature_31.png
| 2024-07-04 | 2024-07-04 11:35 | Marikina City Health Office | 805211 | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28795 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1675 ASSAY: Xpert HIV-1 Viral Load, COUNT: 330 |  | GX-IV-D-6C | 805211 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852918578588426/5958852918578588426_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852918578588426/5958852918578588426_signature_31.png
| 2024-07-04 | 2024-07-04 14:01 | Amang Rodriguez Memorial Medical Center | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28792 | PBSP | maricel_c27@yahoo.cim | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4487 |  | GX-IV-D-6C | 807982 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852908575480499/5958852908575480499_signature_28.png | Maricel S Calderon | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852908575480499/5958852908575480499_signature_31.png
| 2024-07-04 | 2024-07-04 11:39 | Marikina City Health Office | Shoe Ave., Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28793 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 402 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110020361 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852908577801544/5958852908577801544_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852908577801544/5958852908577801544_signature_31.png
| 2024-07-04 | 2024-07-04 11:38 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28794 | DOH | amymitra22@gmail.com | 24 |  |  | GX-IV-D-6C | 831588 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852918571272209/5958852918571272209_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852918571272209/5958852918571272209_signature_31.png
| 2024-07-04 | 2024-07-04 14:12 | Philippine Orthopedic Center | Banawe, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28791 | PBSP | cmblanceta@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1884 | 2024-07-03 | GX IV 6C-D | 803088 | Module B1 for replacement | Performed Preventive maintenance Replaced module B1 Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced module B1 210092883 -> 807091 | Unit is operational All modules working |  | Gx module 6C | DR 49013 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852275947185592/5958852275947185592_signature_28.png | Claire Marie B. Lanceta | https://www.jotform.com/uploads/macarewecare/222927978372067/5958852275947185592/5958852275947185592_signature_31.png
| 2024-07-04 | 2024-07-04 11:39 | Marikina City Health Office | Shoe Ave., Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28790 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 402 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110020361 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847887382644449/5958847887382644449_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847887382644449/5958847887382644449_signature_31.png
| 2024-07-04 | 2024-07-04 11:41 | Marikina City Health Office | Shoe Ave., Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28788 | MACARE-HBDC | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 224 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847877387708225/5958847877387708225_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847877387708225/5958847877387708225_signature_31.png
| 2024-07-04 | 2024-07-04 14:01 | Amang Rodriguez Memorial Medical Center | Sumulong Highway, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28789 | PBSP | maricel_c27@yahoo.cim | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4487 |  | GX-IV-D-6C | 807982 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847887388123965/5958847887388123965_signature_28.png | Maricel S Calderon | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847887388123965/5958847887388123965_signature_31.png
| 2024-07-04 | 2024-07-04 11:38 | Marikina City Health Office | Shoe Ave, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28787 | DOH | amymitra22@gmail.com | 24 |  |  | GX-IV-D-6C | 831588 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847867386460517/5958847867386460517_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847867386460517/5958847867386460517_signature_31.png
| 2024-07-04 | 2024-07-04 11:35 | Marikina City Health Office | 805211 | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28786 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1675 ASSAY: Xpert HIV-1 Viral Load, COUNT: 330 |  | GX-IV-D-6C | 805211 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847857387261948/5958847857387261948_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847857387261948/5958847857387261948_signature_31.png
| 2024-07-04 | 2024-07-03 15:42 | Teresa Rural Health Unit | Teresa, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28785 | DOH | teresa_rhu@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 960 |  | GX-IV-D-6C | 820710 | GeneXpert unable to power on. | Checked UPS | Checked connections Checked AVR and UPS output Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational but without UPS. UPS for replacement. | UPS has no output voltage. GeneXpert System currently connected to the mains. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847787386805430/5958847787386805430_signature_28.png | Rosbenilda Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847787386805430/5958847787386805430_signature_31.png
| 2024-07-04 | 2024-07-04 11:30 | Marikina City Health Office | Shoe Ave., Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28784 | MACARE-HBDC | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 896 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847777388936513/5958847777388936513_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847777388936513/5958847777388936513_signature_31.png
| 2024-07-04 | 2024-07-02 11:02 | Ospital ng Makati | Sampaguita St. Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 28783 | MACARE-HBDC | osmakpatho_labs@yahoo.com.ph | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847767389867764/5958847767389867764_signature_28.png | Cristina Lee | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847767389867764/5958847767389867764_signature_31.png
| 2024-07-04 | 2024-07-02 15:58 | St. Luke's Medical Center - BGC | Rizal Dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 28782 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 |  |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module C4/210096427 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847747388642343/5958847747388642343_signature_28.png | Ma. Mikhaela Rapadas | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847747388642343/5958847747388642343_signature_31.png
| 2024-07-04 | 2024-07-02 15:53 | St. Luke's Medical Center - BGC | Rizal dr., BGC, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 28781 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 192 ASSAY: Xpert CT/NG, COUNT: 1030 ASSAY: Xpert HBV VL, COUNT: 89 ASSAY: Xpert HPV, COUNT: 1079 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2745 ASSAY: Xpert HCV VL, COUNT: 59 ASSAY: Xpert HIV-1 Viral Load, COUNT: 70 |  | GX-IV-L-6C | 808411 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B3/702387 not operational, under observation. Module B4/68862 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847737387110585/5958847737387110585_signature_28.png | Ma. Mikhaela Rapadas | https://www.jotform.com/uploads/macarewecare/222927978372067/5958847737387110585/5958847737387110585_signature_31.png
| 2024-07-04 | 2024-07-04 14:04 | TAYASAN RURAL HEALTH UNIT | POB. TAYASAN NEGROS ORIENTAL | 7 | Negros Oriental | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28780 | UNOPS | porshcabz151@yahoo.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-IV 10C | 110023502 | NEW INSTALLATION | COLLECT IQR | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958827103326081086/5958827103326081086_signature_28.png | Portia Cabanting | https://www.jotform.com/uploads/macarewecare/222927978372067/5958827103326081086/5958827103326081086_signature_31.png
| 2024-07-04 | 2024-07-04 12:42 | Pamplona Rural Health Unit | Centro, Pamplona, Cagayan | 2 | Cagayan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28779 | UNOPS | roxasmatthew659@yahoo.com.ph | 28 |  | 2024-06-20 | GXIV-10C | 110023503 | For Installation | Installed Genexpert Installed PC Installed Printer Installed Drivers Installed UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Package | 11630 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958776373819464255/5958776373819464255_signature_28.png | Wicki Zena Gerona | https://www.jotform.com/uploads/macarewecare/222927978372067/5958776373819464255/5958776373819464255_signature_31.png
| 2024-07-04 | 2024-07-04 12:22 | Pamplona Rural Health Unit | Centro, Pamplona, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28778 | PBSP | wickizena@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 557 |  | GXIV-6C | 814177 | Module B1 Valve not Aligned | Manually aligned Valve Checked Wirings (broken) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module B1 for Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958764865716756121/5958764865716756121_signature_28.png | Wicki Zena Gerona | https://www.jotform.com/uploads/macarewecare/222927978372067/5958764865716756121/5958764865716756121_signature_31.png
| 2024-07-04 | 2024-07-04 11:25 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Blood Bank | NO | INSTALLATION ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28777 | COMMERCIAL | bloodbank@drmc.doh.gov.ph | 26.5 |  |  | MCS+ | 21M389SPG | Newly Installed Haemonetics MCS+ | Install Diagnostic Card and perform Diagnostic Tests Checked the Cuff pressure Checked the Centrifuge speed at 3000rpm to 7000 rpm Checked the Valves if working properly Checked the Donor Lights Perform the Cycle Test to ensured the functionality of the Machine | Perform Diagnostics and Cycle Test | Successfully finished the Cycle Test Machine is operational and functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958746823412057452/5958746823412057452_signature_28.png | Christine C. Porquiado | https://www.jotform.com/uploads/macarewecare/222927978372067/5958746823412057452/5958746823412057452_signature_31.png
| 2024-07-04 | 2024-07-04 11:28 | Masbate City Health Office | City Hall Compound, Brgy. Kalipay, Masbate City | 5 | Masbate | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28776 | PBSP | desrmt2017@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1329 |  | GX4-10C | 808633 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod,icore, fan filter Generate reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958734504228779445/5958734504228779445_signature_28.png | Desiree Mae D. Dela Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5958734504228779445/5958734504228779445_signature_31.png
| 2024-07-04 | 2024-07-04 11:24 | Masbate City Health Office | City Hall Compound, Brgy. Kalipay, Masbate City | 5 | Masbate | Laboratory | NO | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28775 | UNOPS | desrmt2017@gmail.com | 29 |  |  | GX4-10C | 110023512 | GX4-10C for Installation | Installed GX4-10C(110023512), CPU(CPH1023F47089), Monitor, UPS(AT1000), Printer Generate reports | Installed Gx4-10C set | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958733594227694208/5958733594227694208_signature_28.png | Desiree Mae D.Dela Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5958733594227694208/5958733594227694208_signature_31.png
| 2024-07-04 | 2024-07-04 10:31 | Poro Rural Health Unit | Poro, Camotes, Cebu | 7 | Cebu | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28774 | DOH | avilacharmaine6@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 625 |  | 6C GXIV-4D | 838279 | Due for PM | #NAME? | -Preventive Maintenance -Archive Tests -Upgrade XpertCheck 1.5a to 1.5b | All modules are functional. | No filter (to be delivered) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958702796818354813/5958702796818354813_signature_28.png | Charmaine Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/5958702796818354813/5958702796818354813_signature_31.png
| 2024-07-04 | 2024-07-04 10:28 | Jose Dalman Rural Health Unit | Jose Dalman Rural Health Unit | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE Replaced System Unit and repaired printer | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28773 | PBSP | crizzycristina@yahoo.com | 23 |  | 2024-07-03 | GX 2 (6c) | 834827 | System unit unable to turn on. | Replaced System uni, repaired printer and conducted Preventive Maintenence | Replaced System uni, repaired printer and conducted Preventive Maintenence | All modules are available. Note that uninterruptible Power supply (UPS) is not functioning properly and must be replaced emedietly. | Note that uninterruptible Power supply (UPS) is not functioning properly and must be replaced emedietly. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958702587767595949/5958702587767595949_signature_28.png | MA CRISTINA P MARIÑO | https://www.jotform.com/uploads/macarewecare/222927978372067/5958702587767595949/5958702587767595949_signature_31.png
| 2024-07-04 | 2024-07-04 09:49 | Philippine Children's Medical Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | gehbedayos@gmail.com |  | 28772 | PBSP | kate_july27@yahoo.com | 28 |  | 2024-06-25 | GX4 D | 807672 | For module replacement B1 685665 - 732262 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed module replacement | Unit operational | For pull out defective module 685665 | Genexpert 6 color module | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958679686916986269/5958679686916986269_signature_28.png | Katherine Tabaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5958679686916986269/5958679686916986269_signature_31.png
| 2024-07-04 | 2024-07-04 09:37 | Masambong Health Center | Capoaa, Brgy. Masambong, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28771 | PBSP | michellebonrostro@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1985 | 2024-06-25 | GX IV 6C-D | 814160 | Module B1 for replacement | Performed preventive maintenance Replaced module B1 Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced module B1 690588 -> 869848 | Unit is operational All modules working |  | Gx Module 6C | DR 48924 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958667123241655302/5958667123241655302_signature_28.png | Michelle Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/5958667123241655302/5958667123241655302_signature_31.png
| 2024-07-04 | 2024-07-03 13:06 | Infanta Rural Health Unit | Poblacion, Infanta Pangasinan | 1 | Pangasinan | Laboratory | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28770 | UNOPS | charissealmandres8@gmail.com | 30 |  | 2024-06-20 | GXIV-4-D-10C | 110023501 | For Installation of GeneXpert System | #NAME? | Performed Installation of GeneXpert System | Unit is Fully Operational |  | GeneXpert System | DR 11629 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958665360626636569/5958665360626636569_signature_28.png | Charisse A. Alferez | https://www.jotform.com/uploads/macarewecare/222927978372067/5958665360626636569/5958665360626636569_signature_31.png
| 2024-07-04 | 2024-07-03 08:53 | Dasol Community Hospital | Dasol, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28769 | DOH | jeremypulido0521@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 660 | 2024-04-30 | GXIV-4-D-6C | 838287 | For Calibration and Preventive Maintenance | #NAME? | Performed Calibration and Preventive Maintenance | #NAME? | Work with Engineer Venzon | XpertCheck Kits | DR 11448 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958661160623171123/5958661160623171123_signature_28.png | Jeremy Pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/5958661160623171123/5958661160623171123_signature_31.png
| 2024-07-04 | 2024-07-04 08:32 | Catmon Health Center | Catmon, Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28768 | DOH | mariathes0218@gmail.com | 25 |  |  | GXIV | 831903 | CPU for troubleshooting | Installed components under DOH onto one system | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958648562415513921/5958648562415513921_signature_28.png | Clara Gloria | https://www.jotform.com/uploads/macarewecare/222927978372067/5958648562415513921/5958648562415513921_signature_31.png
| 2024-07-04 | 2024-07-04 08:45 | Sta. Maria Rural Health Unit | Poblacion West, Sta. Maria | 1 | Lanao del Sur | Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28767 | UNOPS | jomarcaberto24@gmail.com | 25 |  |  | GXIV-10C | 110023500 | For installation of GeneXpert System. | Installation. | Performed installation of GeneXpert System. Verified outlet voltage. Performed GebeXpert system maintenance training. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958633454225161947/5958633454225161947_signature_28.png | Jomar Caberto | https://www.jotform.com/uploads/macarewecare/222927978372067/5958633454225161947/5958633454225161947_signature_31.png
| 2024-07-04 | 2024-07-03 11:40 | Philippine Red Cross (Laguna Chapter) | P.Guevarra,Sta Cruz Laguna | 4A | Laguna | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28766 | MACARE-HBDC | lagunablood@redcross.org.ph | 26 |  | 2024-07-02 | MCS + | 02J046 | Self test failed yellow valve | Replaced and instaaled valve manifold Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Performed repair | Operational | Work with Monsour Doctor | Valve manifold | 11697 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958611926914181952/5958611926914181952_signature_28.png | Phil. Red Cross Laguna Chapter | https://www.jotform.com/uploads/macarewecare/222927978372067/5958611926914181952/5958611926914181952_signature_31.png
| 2024-07-03 | 2024-07-03 14:00 | Mlang Rural Health Unit | Poblacion, Mlang, North Cotabato | 12 | Cotabato | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28765 | PBSP | kathlynlacabe7@gmail.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1164 | 2024-07-03 | Gx-IV (6-color) | 815434 | Module 2 for Service Required | Replace defective module 2 SN:696344 with new module SN:628510 Genexpert machine Cleaning/preventive maintenance Checking of errors Data collection | Perform Genexpert machine module replacement/repair Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6-color | Cepheid case: 1553561 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958084567708551262/5958084567708551262_signature_28.png | Kathlyn mae y lacabe | https://www.jotform.com/uploads/macarewecare/222927978372067/5958084567708551262/5958084567708551262_signature_31.png
| 2024-07-03 | 2024-07-03 16:52 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28764 | MACARE-HBDC | lisandrogacias@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 2695 ASSAY: Xpert MTB/RIF Ultra, COUNT: 340 ASSAY: Xpert SARS-CoV-2, COUNT: 218 |  | GXIV-6C | 110008457 | Due for Maintenance | Performed Maintenance (Cleaned Icore) (Cleaned Cartridge Bay) (Cleaned Filter) (Cleaned Plunger Rod) Performed Functionality Test | Gathered Reports | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958063390801304310/5958063390801304310_signature_28.png | Lisandro Gacias | https://www.jotform.com/uploads/macarewecare/222927978372067/5958063390801304310/5958063390801304310_signature_31.png
| 2024-07-03 | 2024-07-03 16:30 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Serology | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28763 | MACARE-HBDC | cvmcserology@gmail.com | 25 | ASSAY: Others, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 144 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 30 ASSAY: Xpert HIV-1 Viral Load, COUNT: 447 |  | GXIV-6C | 803870 | Due for Maintenance | Performed Maintenance (Cleaned Filter) (Cleaned Icore) (Cleaned Cartridge Bay) (Cleaned Plunger Rod) | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958049985712134542/5958049985712134542_signature_28.png | Dorothy Pajarillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5958049985712134542/5958049985712134542_signature_31.png
| 2024-07-03 | 2024-06-28 10:54 | General Trias City Health Office | Hospital Drive, General Trias, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | chogentri.lab@gmail.com |  | 28761 | DOH | mackchico00@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3594 |  | GXIV-D-6C | 821340 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Generate Reports. Performed Self-test. | 3 Modules are Operational. 1 Disabled for Reporting. | A3/ 601653 for reporting. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958037336402062951/5958037336402062951_signature_28.png | Mack Alfred L. Chico | https://www.jotform.com/uploads/macarewecare/222927978372067/5958037336402062951/5958037336402062951_signature_31.png
| 2024-07-03 | 2024-07-02 11:26 | San Isidro Health Center | Dr. A. Santos Ave, San Isidro, Parañaque | NCR | NCR 4th - Parañaque | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28762 | PBSP | herminadoa@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2031 ASSAY: Xpert HIV-1 Viral Load, COUNT: 76 |  | GXIV-D-6C | 838192 | Module 3 Service Required in XpertCheck. System with UPS Problem. | Repair. UPS Troubleshooting. | Cleaned and Replaced Module 3. Performed Self-test. Generate Reports. Checked UPS Output Voltage and Backup Function. | 2 Modules are Operational. 2 for Retry in XpertCheck. UPS Backup Voltage is Low. | Recommendation for UPS Replacement. B3: 210091236 -> 749978 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958037336408566191/5958037336408566191_signature_28.png | Amelia G. Herminado | https://www.jotform.com/uploads/macarewecare/222927978372067/5958037336408566191/5958037336408566191_signature_31.png
| 2024-07-03 | 2024-06-28 11:01 | General Trias City Health Office | Hospital Drive, General Trias, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | chogentri.lab@gmail.com |  | 28759 | PBSP | mackchico00@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3594 |  | GXIV-D-6C | 804691 | Unit is due for Maintenance. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Generate Reports. Performed Self-test. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958037336401911082/5958037336401911082_signature_28.png | Mack Alfred L. Chico | https://www.jotform.com/uploads/macarewecare/222927978372067/5958037336401911082/5958037336401911082_signature_31.png
| 2024-07-03 | 2024-07-03 12:44 | Abra Provincial Hospital | Abra-Ilocos Road, Calaba, Bangued, Abra | CAR | Abra | Molecular | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28760 | UNOPS | aphpathologydepartment@gmail.com | 25 |  | 2024-06-20 | GXIV-D-10C | 110023499 | Unit is due for Installation. | Installation. | Checked Outlet Voltage. Installed GeneXpert System. Performed Self-test. Generate Reports. | GeneXpert System Installed and Functional. |  | GeneXpert System | DR 11627 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958037336408006826/5958037336408006826_signature_28.png | Riverlie A Valera | https://www.jotform.com/uploads/macarewecare/222927978372067/5958037336408006826/5958037336408006826_signature_31.png
| 2024-07-03 | 2024-07-03 16:01 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28758 | DOH | iegparayno@gmail.com | 25 |  |  | MCS+ | 15HO336 | Due for Maintenance | Performed Maintenance Adjust Line Sensor | Check line sensor Check rollers Check values Performed Test Run | Unit is Operational | Due for Calibration, Rollers need Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958033805712161430/5958033805712161430_signature_28.png | Reggie Parayno | https://www.jotform.com/uploads/macarewecare/222927978372067/5958033805712161430/5958033805712161430_signature_31.png
| 2024-07-03 | 2024-07-03 12:14 | Binalbagan Municipal Health Office | Binalbagan, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 28757 | DOH | ntpbinalbagan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 957 |  | GX-4 R2 | 838280 | Due doe Xpertcheck on July 31, 2024. | #NAME? | #NAME? | All modules passed Xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958006273413518206/5958006273413518206_signature_28.png | Xerxanne A. Galilea | https://www.jotform.com/uploads/macarewecare/222927978372067/5958006273413518206/5958006273413518206_signature_31.png
| 2024-07-03 | 2024-07-03 13:14 | Quirino Memorial Medical Center | JP Rizal, Project 4, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | giandesilvamagnaye@gmail.com |  | 28756 | PBSP | carlopalomera58@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 21 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3764 | 2024-06-25 | GX IV 6C-D | 815484 | Module B3 for replacement | Performed preventive maintenance Replaced module B3 Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced module B3 754303 -> 878880 | unit is operational all modules working |  | GX module 6C | DR 48925 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5958002273616279093/5958002273616279093_signature_28.png | Ma. Regina Catalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5958002273616279093/5958002273616279093_signature_31.png
| 2024-07-03 | 2024-07-03 14:44 | Davao Regional Medical Center Blood Bank | Apokon, Tagum City | 11 | Davao del Norte | Blood Bank | NO | INSTALLATION | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28755 | COMMERCIAL | bloodbank@drmc.doh.gov.ph | 26.9 |  | 2024-06-14 | MCS+ | 21M389SPG | For Installation of Haemonetics MCS+ | Set up Haemonetics Machine System Assemble the Trolley for Haemonetics Check the Line Voltages Setup the AVR with SN:2379PSX809 and UPS with SN: AT1K240323ST0046 Assemble the Haemonetics Machine with SN:21M389SPG Setup Parameters for UPP Configuration | Perform Installation | Unit Successfully Installed Unit is Operational |  |  | DR 11607 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957995031416421225/5957995031416421225_signature_28.png | Christine C. Porquiado | https://www.jotform.com/uploads/macarewecare/222927978372067/5957995031416421225/5957995031416421225_signature_31.png
| 2024-07-03 | 2024-07-01 11:43 | Bagumbayan Rural Health Unit | Poblacion Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28753 | DOH | marizalenoid@yahoo.com | 27.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 901 ASSAY: Xpert MTB/RIF Ultra, COUNT: 905 |  | Gx-IV (6-color) | 838261 | Module 2,3 and 4 stuck cartridge | Repair module stuck cartridge Genexpert Cleaning/preventive maintenance Checking of errors Perform actual run at module 2,3 and 4 to test module functionality Data collection | Perform Genexpert machine module repair Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957955822672217398/5957955822672217398_signature_28.png | Ma. Riza R. Dionela | https://www.jotform.com/uploads/macarewecare/222927978372067/5957955822672217398/5957955822672217398_signature_31.png
| 2024-07-03 | 2024-07-02 16:08 | Dr.Amado Diaz Provincial Foundation Hospital | Midsayap, North Cotabato | 12 | Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28754 | PBSP | charlenemaeamatac526@gmail.com | 27.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1038 ASSAY: Xpert HIV-1 Viral Load, COUNT: 108 | 2024-07-03 | Gx-IV | 807623 | Module 2 for Service Required | Replace defective module 1 SN:729721 with new module SN:210036596 Genexpert Cleaning/preventive maintenance Checking of errors Data collections | Perform Genexpert machine module replacement/repair Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6-color | Cepheid case: 1550281 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957955822677406131/5957955822677406131_signature_28.png | CHARLENE MAE A. AMATAC | https://www.jotform.com/uploads/macarewecare/222927978372067/5957955822677406131/5957955822677406131_signature_31.png
| 2024-07-03 | 2024-07-03 13:44 | Cabagan Rural Health Unit | Centro, Cabagan, Isabela | 2 | Isabela | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28752 | PBSP | rhuCLcabagan2022@gmail.com | 29 |  |  | GXIV-6C | 805104 | CPU not booting | Replaced Cmos Battery | Installed Genexpert Installed UPS Installed Printer Installed PC Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | From Southern Isabela Medical Center |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957950403819072304/5957950403819072304_signature_28.png | Millet Joy Jampas | https://www.jotform.com/uploads/macarewecare/222927978372067/5957950403819072304/5957950403819072304_signature_31.png
| 2024-07-03 | 2024-07-03 12:24 | San Francisco Rural Health Unit | San Francisco, Camotes, Cebu | 7 | Cebu | TB Dots | YES | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28751 | PBSP | ejoey22@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 634 |  | 6C GXIV-2D | 835058 | #NAME? | #NAME? | #NAME? | Module B1 H/W fail; failed XPC. Module B2 is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957923814022340284/5957923814022340284_signature_28.png | Josiecora Albarracin | https://www.jotform.com/uploads/macarewecare/222927978372067/5957923814022340284/5957923814022340284_signature_31.png
| 2024-07-03 | 2024-07-03 10:44 | Caloocan City Health Office | Sangandaan, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28750 | PBSP | cruzernitamay@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2363 ASSAY: Xpert HIV-1 Viral Load, COUNT: 141 ASSAY: Xpert HIV-1 Viral Load, COUNT: 355 |  | GX IV 6C D | 814154 | Packed 6C machine requested to be used while waiting for relacation | Unpacked 6C machine Merged 2 machine in one cpu Generated report | Unpacked 6C machine CPU waiting for replacement Merged to 805271 through network switch for the meantime Checked and cleaned cartridge bays, plunger rod, icore slot, and filter | Unit is operational B1 waiting for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957845128347902612/5957845128347902612_signature_28.png | ernita may b. cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5957845128347902612/5957845128347902612_signature_31.png
| 2024-07-03 | 2024-07-03 10:28 | Caloocan City Health Office | Sangandaan, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28749 | PBSP | cruzernitamay@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2363 ASSAY: Xpert HIV-1 Viral Load, COUNT: 141 ASSAY: Xpert HIV-1 Viral Load, COUNT: 355 | 2024-06-25 | GXIV 6C D | 805271 | Module A2 and A3 for replacement | Performed preventive maintenance Replaced modules Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced modules A2: 757103 -> 889857 A3: 620656 -> 210000207 | unit is operational All modules working |  | GX Modules 6C | DR 48932 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957838198344317050/5957838198344317050_signature_28.png | ernita may b. cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5957838198344317050/5957838198344317050_signature_31.png
| 2024-07-03 | 2024-07-02 08:17 | E. Rodriguez Health Center | Pajo St, Project 2, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28748 | DOH | beccaramilo07@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 58 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2189 |  | GX IV 6C | 831848 | Unit due for maintenance Module A4 frequent error | Performed Preventive maintenance Performed reliability test Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Checked system log Temp error on module A4 A4 for reporting | Unit is operational 3 modules working 1 module for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957747992325922554/5957747992325922554_signature_28.png | Rebecca Joy Ramilo | https://www.jotform.com/uploads/macarewecare/222927978372067/5957747992325922554/5957747992325922554_signature_31.png
| 2024-07-02 | 2024-07-02 17:32 | Southern Philippines Medical Center | JP. Laurel Avenue, Bajada Davao City | 11 | Davao del Sur | RHIVDA | NO | INSTALLATION | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28747 | UNOPS | spmcrhivda@gmail.com | 24.7 |  | 2024-06-20 | GX-IV 10C | 110023505 | For Installation | Installed the GeneXpert Machine Serial Number :110023505 With CPU System Unit SN: CPH1123F47508 Install Assay for HIV VL, MTB RIF ULTRA Enroll Enduser Account Generate IQR | Perform Installation | Successfully Installed the GeneXpert Machine System |  |  | DR 11648 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957225146029456192/5957225146029456192_signature_28.png | Flordeliza B. Go RMT MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/5957225146029456192/5957225146029456192_signature_31.png
| 2024-07-02 | 2024-07-02 14:45 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28746 | UNOPS | chd.laureanjocabase@gmail.com | 28 |  |  | Gx4 10c | 110023513 | None | New Genexpert Machine Installations Gather data IQR Basic Operation training of End-User | New Genexpert Machine Installations Gather data IQR Basic Operation training of End-User | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957157569916936744/5957157569916936744_signature_28.png | Laurean Jo F. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5957157569916936744/5957157569916936744_signature_31.png
| 2024-07-02 | 2024-07-02 15:23 | Cauayan City Health Office | Cabaruan, Cauayan City, Isabela | 2 | Isabela | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28745 | UNOPS | cho1lab14@gmail.com | 25 |  | 2024-06-20 | GXIV-10C | 110023504 | N/a | N/a | Installed Genexpert Installed PC Installed Printer Installed UPS Installed Monitor Installed Drivers Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Genexpert Package | 11631 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957145973818580359/5957145973818580359_signature_28.png | Arlyn Daggao | https://www.jotform.com/uploads/macarewecare/222927978372067/5957145973818580359/5957145973818580359_signature_31.png
| 2024-07-02 | 2024-07-02 14:26 | Iriga City Health Office | Sta. Cruz Sur, Iriga City, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28744 | DOH | mongealyzza@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1799 |  | GX4-6C | 838391 | GX DX version 6.4 | Install and update GX DC Cleaned cartridge bays, plunger rod, icore Generate report | Update GX DX Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957125212323645173/5957125212323645173_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5957125212323645173/5957125212323645173_signature_31.png
| 2024-07-02 | 2024-07-02 14:20 | Iriga City Health Office | Sta. Cruz Sur, Iriga City, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28743 | UNOPS | mongealyzza@yahoo.com.ph | 25 |  |  | GX4-10C | 110023511 | GX4-10C for Installation | Installed GeneXpert machine(110023511), CPU(CPH1023F46954), monitor, UPS, Printer Generate report | Installed GX4-10C set | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957124332326597689/5957124332326597689_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5957124332326597689/5957124332326597689_signature_31.png
| 2024-07-02 | 2024-07-02 14:26 | Jaro Rural Health Unit | Jaro Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28742 | PBSP | jocelynpagpaguitan@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 626 | 2024-07-02 | GXIV - 2D | 834817 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957112532124505007/5957112532124505007_signature_28.png | Jocelyn C. Pagpaguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5957112532124505007/5957112532124505007_signature_31.png
| 2024-07-02 | 2024-07-02 13:45 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | REPAIR | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28741 | COMMERCIAL | cduhlabresults2@yahoo.com | 22 |  | 2024-07-01 | MCS+ | 14F014 | #NAME? | #NAME? | #NAME? | #NAME? |  | DISPLAY CABLE | 97529 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957087009916909125/5957087009916909125_signature_28.png | CHRISTINE JADE LAGAT | https://www.jotform.com/uploads/macarewecare/222927978372067/5957087009916909125/5957087009916909125_signature_31.png
| 2024-07-02 | 2024-07-02 12:36 | Angadanan Rural Health Unit | Centro 3, Angadanan, Isabela | 2 | Isabela | Laboratory | NO | INSTALLATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | rhuangadananlaboratory@gmail.com |  | 28740 | DOH | rhuangadananlaboratory@gmail.com | 27 |  |  | GXIV-6C | 807458 | Cmos Battery | Replaced CMOS Battery | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957051081819633545/5957051081819633545_signature_28.png | Julie Ann Dalupang | https://www.jotform.com/uploads/macarewecare/222927978372067/5957051081819633545/5957051081819633545_signature_31.png
| 2024-07-02 | 2024-07-02 12:32 | Tondo Medical Center | H. Lopez Blvd., Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28739 | PBSP | microscan@tmc.doh.gov.ph | 25 |  |  | GXIV | 844121 | Unit for Preventive Maintenance. Shuffled module assignments. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed re-assignments of modules. Generated reports. | Unit is operational. For reporting of modules B2 and C3 due to temperature errors. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957050962992825313/5957050962992825313_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5957050962992825313/5957050962992825313_signature_31.png
| 2024-07-02 | 2024-07-02 11:55 | Lianga Rural Health Unit | Poblacion, Lianga | 13 | Surigao del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28738 | PBSP | gyvxs@yahoo.com | 25 |  |  | GX-2 | 835174 | Module B2 found Service Required from XPC | Module Replacement | Replacing Module B2 with SN: 788906 to SN: 691828 | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957020147872758938/5957020147872758938_signature_28.png | Gyver I. Sistona | https://www.jotform.com/uploads/macarewecare/222927978372067/5957020147872758938/5957020147872758938_signature_31.png
| 2024-07-02 | 2024-05-14 13:49 | Bulakan Rural Health Unit | Bulakan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28737 | DOH | l_gemino@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1194 ASSAY: Xpert HIV-1 Viral Load, COUNT: 97 |  | GXIV | 831641 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Module A1 under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5957009420816811738/5957009420816811738_signature_28.png | Lilibeth Gemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5957009420816811738/5957009420816811738_signature_31.png
| 2024-07-01 | 2024-07-01 13:48 | Victorias City Health Office | Victorias City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 28736 | PBSP | dharwinliquez23@gmail.com | 25 |  |  | GX-4 | 838141 | Defective nodules 691294/B1 and 692687/B2 | Preventive maintenance - cleaned optics, module bays. plunger rods, filter. | Replaced modules B1 and B2 with new modules SN 692459 and 696137 respectively. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5956228635767836816/5956228635767836816_signature_28.png | Dharwin Liquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5956228635767836816/5956228635767836816_signature_31.png
| 2024-07-01 | 2024-07-01 12:47 | Talavera General Hospital | Maestrang Kikay, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28735 | MACARE-HBDC | pjgtehlab@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 81 ASSAY: Xpert MTB/RIF Ultra, COUNT: 720 ASSAY: Xpert CT/NG, COUNT: 29 ASSAY: Xpert HCV VL, COUNT: 11 ASSAY: Xpert HIV-1 Viral Load, COUNT: 138 | 2024-07-27 | GXIV-10C | 110013773 | Module B1 and B2 error 1002 | Performed Calibration (Module B2 - Service Required) (Module B1, B3, B4 - Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | 11691 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5956214147759276943/5956214147759276943_signature_28.png | Sharey R. Capati | https://www.jotform.com/uploads/macarewecare/222927978372067/5956214147759276943/5956214147759276943_signature_31.png
| 2024-07-01 | 2024-07-01 12:11 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28734 | DOH | marteetingzon20@gmail.com | 26 |  | 2024-07-01 | GXIV 6C | 838421 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5956198933312549683/5956198933312549683_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5956198933312549683/5956198933312549683_signature_31.png
| 2024-07-01 | 2024-06-28 14:56 | Cainta City Health Office | Karangalan Dr., Cainta, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28733 | PBSP | parconriza521@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 959 ASSAY: Xpert HIV-1 Viral Load, COUNT: 117 | 2024-06-28 | GX-IV-D-6C | 838182 | Module B4 for replacement | Replaced module B4 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B4 Generated reports | Unit is operational. All modules operational. |  | Old: B4/807137 New: B4/897559 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5956063433614598845/5956063433614598845_signature_28.png | Riza Parcon | https://www.jotform.com/uploads/macarewecare/222927978372067/5956063433614598845/5956063433614598845_signature_31.png
| 2024-07-01 | 2024-06-28 11:11 | San Pedro Rural Health Unit 2 | Brgy. Langgam, San Pedro Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28731 | DOH | lenimahalin@gmail.com | 25 |  | 2024-06-28 | GXIV | 821189 | Unit for preventive maintenance. For module replacement of module A1. For troubleshooting of CPU. | Preventive maintenance, module replacement and troubleshooting. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Performed troubleshooting of CPU. Generated reports. | Unit is operational. CPU (HP Z240 SN: 2UA8471W4T) not operational and for replacement due to continous booting. |  | GX module A1 (210018940) in replacement to defective A1 (784169). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5956062783615532777/5956062783615532777_signature_28.png | Celenita Mahalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5956062783615532777/5956062783615532777_signature_31.png
| 2024-07-01 | 2024-06-27 10:29 | Philippine General Hospital | Taft Ave., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28730 | PBSP | pghmolecularlab@gmail.com | 25 |  |  | GXIV | 834624 | Unit for preventive maintenance. For troubleshooting of monitor. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed troubleshooting of monitor Generated reports. | Unit is operational. Monitor HP ProDisplay 19A (CNK94302DX) has a malfunctioning controls. Monitor for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5956062783618371885/5956062783618371885_signature_28.png | Kleante bautro | https://www.jotform.com/uploads/macarewecare/222927978372067/5956062783618371885/5956062783618371885_signature_31.png
| 2024-07-01 | 2024-06-27 10:31 | Philippine General Hospital | Taft Ave., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28732 | COMMERCIAL | pghmolecularlab@gmail.com | 25 |  |  | GX16 | 844975 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5956062793612864527/5956062793612864527_signature_28.png | Kleante bautro | https://www.jotform.com/uploads/macarewecare/222927978372067/5956062793612864527/5956062793612864527_signature_31.png
| 2024-07-01 | 2024-06-28 15:39 | Puerto Princesa City Health Office | Puerto Princesa Palawan | 4B | Palawan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28729 | PBSP | jocal0324@gmail.com | 29 |  |  | GXIV 6C L | 804272 | Troubleshoot module B1 for stuck cartridge Unit due for calibration | Performed preventive maintenance Module B1 troubleshooting Tried to perform Xpertcheck | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Tried removing stuck cartridge but nothing happens Cartridge still stuck on B1 Gathered report for B1 to be reported since it's been on H/W failed since May Tried Xpertcheck but failed to do so because of error 2031 Can't perform Xpertcheck for reporting | unit is operational 3 modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5956053393618860072/5956053393618860072_signature_28.png | Jocelyn L. Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5956053393618860072/5956053393618860072_signature_31.png
| 2024-07-01 | 2024-06-28 10:04 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28728 | PBSP | crmctblab@gmail.com | 21.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 28 ASSAY: Xpert HBV VL, COUNT: 7 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 | 2024-06-27 | Gx-IV (6-color) | 807463 | Module 3 for Service Required | Replace defective module 3 SN:210100874 with new module SN:639321 Genexpert machine Cleaning/preventive maintenance Checking of errors Data collection | Perform Genexpert machine module repair Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6-color | Cepheid case:1554610/1554694 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5955727357142583234/5955727357142583234_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5955727357142583234/5955727357142583234_signature_31.png
| 2024-06-30 | 2024-06-30 18:47 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28727 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 11J219 | Error 413 Valve Problem "purple valve test failure" Display Monitor input key Draw,Prime, and Return, etc malfunction. | Replaced temporarily the defective pneumatic assembly with service pneumatic assy and advance cpu board while waiting for the purchase order of the parts to avoid interruptions of their apheresis procedures. | Replaced temporarily the defective pneumatic assembly with service pneumatic assy and advance cpu board while waiting for the purchase order of the parts to avoid interruptions of their apheresis procedures. | Unit is operational with service parts of pneumatic assembly and advance cpu board. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5955539041223852841/5955539041223852841_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/5955539041223852841/5955539041223852841_signature_31.png
| 2024-06-29 | 2024-06-28 12:25 | Philippine Red Cross Laguna Chapter | P Guevarra St., Sta Cruz, Laguna | 4A | Laguna | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28726 | MACARE-HBDC | lagunablood@gmail.org.ph | 26 |  |  | MCS + | 02J046 | Error display yellow valve selt test failed | Checked and reset CPU board,Motor board,Safety board, Checked and bypass manually yellow pinch valve Interchange white and yellow valve position but error still occur yellow valve test failed error. Checked valve manifold Checked and test compressor ok Checked connectors Performed system reset multiple times but error still occur | Performed troubleshooting | Unit unoperational Suspect defective parts valve manifold | For additional troubleshooting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5954212510932256011/5954212510932256011_signature_28.png | Redcross Laguna Chapter | https://www.jotform.com/uploads/macarewecare/222927978372067/5954212510932256011/5954212510932256011_signature_31.png
| 2024-06-29 | 2024-06-29 05:51 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Bloodbank | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28725 | DOH | chaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1972 |  | GX4 | 838283 | UPS noisy relay during charging Due for Monthly Preventive Maintenance Clean up probe, Icore, cartridge docking bay, and filter | Clean up probe, Icore, cartridge docking bay, and filter Generate GXX files, IQR, Sys Log, and Assay Statistics | Due for Monthly Preventive Maintenance Clean up probe, Icore, cartridge docking bay, and filter Generate GXX files, IQR, Sys Log, and Assay Statistics | UPS is still working but noisy relay. All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5954210583493731742/5954210583493731742_signature_28.png | Lucille Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5954210583493731742/5954210583493731742_signature_31.png
| 2024-06-28 | 2024-06-28 12:57 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | loreyleinela@gmail.com |  | 28724 | DOH | dy_rosemarie71@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1928 |  | GX-IV 4D | 838401 | Due to for Preventive Maintenance Module Error at B2 and B4 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Install XpertCheck version 1.5B Perform Xpertcheck Calibration for Module 2 and Module 4 | Perform Preventive Maintenance and XpertCheck Calibration for Error Checking | Module B1:Disable - Waiting for replacement parts Module B2 Retry and For Service Required Module B3 Passed the Calibration | Module B4 For under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953699234627958851/5953699234627958851_signature_28.png | Rosemarie Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/5953699234627958851/5953699234627958851_signature_31.png
| 2024-06-28 | 2024-06-28 11:45 | Notre Dame De Chartres Hospital | Baguio City | CAR | Benguet | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28723 | COMMERCIAL | ndchpathology@gmail.com | 25 |  |  | MCS + | 10M103 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational | Work with Engr. Harold Toco |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953659198711004720/5953659198711004720_signature_28.png | Kenneth Michael Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/5953659198711004720/5953659198711004720_signature_31.png
| 2024-06-28 | 2024-06-28 09:26 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | RTDL | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28722 | PBSP | benguetphortdl@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2602 |  | GXIV-4-D-6C | 807621 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953658408713052442/5953658408713052442_signature_28.png | Frannie Calixto | https://www.jotform.com/uploads/macarewecare/222927978372067/5953658408713052442/5953658408713052442_signature_31.png
| 2024-06-28 | 2024-06-28 09:11 | Benguet Provincial Health Office | Km. 5 Benguet General Hospital Compound, La Trinidad, Benguet | CAR | Benguet | TB DOTS | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28721 | UNOPS | benguetphortdl@gmail.com | 29 |  | 2024-06-20 | GXIV-4-D-10C | 110023498 | For Installation of GeneXpert System | #NAME? | Performed Installation of GeneXpert System | Unit is Fully Operational |  | GeneXpert System | DR 11626 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953657508719638569/5953657508719638569_signature_28.png | Frannie G. Calixto | https://www.jotform.com/uploads/macarewecare/222927978372067/5953657508719638569/5953657508719638569_signature_31.png
| 2024-06-28 | 2024-06-27 10:32 | Baguio Health Services Office | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28720 | DOH | sherryannemarrero@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2265 |  | GXIV-4-D-6C | 838281 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953656268714528812/5953656268714528812_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5953656268714528812/5953656268714528812_signature_31.png
| 2024-06-28 | 2024-06-27 09:48 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28719 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 20 | ASSAY: Xpert HBV VL, COUNT: 132 ASSAY: Xpert SARS-CoV-2, COUNT: 207 ASSAY: Xpert HCV VL, COUNT: 25 ASSAY: Xpert HIV-1 Viral Load, COUNT: 23 |  | GXIV-4-D-10C | 110016112 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953655228719557192/5953655228719557192_signature_28.png | Rhea Annetonetye Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5953655228719557192/5953655228719557192_signature_31.png
| 2024-06-28 | 2024-06-27 09:24 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | HACT | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com | hact@bghmc.doh.gov.ph |  | 28718 | MACARE-HBDC | balbahazel@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 470 |  | GXIV-4-D-6C | 821632 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is partially operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953654428719916248/5953654428719916248_signature_28.png | Hazel Gwyneth T. Balba | https://www.jotform.com/uploads/macarewecare/222927978372067/5953654428719916248/5953654428719916248_signature_31.png
| 2024-06-28 | 2024-06-27 09:07 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28717 | PBSP | dmgonzales.pbsp@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 247 ASSAY: Xpert MTB/RIF Ultra, COUNT: 944 |  | GXIV-4-D-6C | 805212 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953653628713224845/5953653628713224845_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5953653628713224845/5953653628713224845_signature_31.png
| 2024-06-28 | 2024-06-27 09:01 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28716 | PBSP | dmgonzales.pbsp@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 65 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1169 ASSAY: Others, COUNT: 37 |  | GXIV-4-D-10C | 120000171 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953652728719973395/5953652728719973395_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5953652728719973395/5953652728719973395_signature_31.png
| 2024-06-28 | 2024-06-27 08:10 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28715 | PBSP | dmgonzales.pbsp@gmail.com | 23 |  | 2024-06-27 | GXIV-4-D-6C | 834685 | For Calibration and Preventive Maintenance | #NAME? | Performed Calibration and Preventive Maintenance | #NAME? |  | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953651248716515974/5953651248716515974_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5953651248716515974/5953651248716515974_signature_31.png
| 2024-06-28 | 2024-06-28 13:56 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28714 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 |  |  | 6C GXIV | 820206 | #NAME? | #NAME? | #NAME? | All modules are functional. | Machine not used for 2 months for lack of cartridge. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953640836225130470/5953640836225130470_signature_28.png | Christine Ivy Paula Agtuca | https://www.jotform.com/uploads/macarewecare/222927978372067/5953640836225130470/5953640836225130470_signature_31.png
| 2024-06-28 | 2024-06-28 14:00 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28713 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 |  |  | 6C GXIV | 838285 | Due for monthly PM | #NAME? | #NAME? | Module B4 disabled for replacement. All other modules are functional. | Machine not used for 2 months due to lack of cartridge. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953640466229718841/5953640466229718841_signature_28.png | Christine Ivy Paula Agtuca | https://www.jotform.com/uploads/macarewecare/222927978372067/5953640466229718841/5953640466229718841_signature_31.png
| 2024-06-28 | 2024-06-26 13:17 | Detoxicare Molecular Diagnostics Laboratory Inc. | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28711 | CONSORTIUM | mikaela.chester@detoxicare.com.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 234 |  | GXIV | 110018312 | Unit for maintenance | NA | #NAME? | Unit is operational Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953639149101922812/5953639149101922812_signature_28.png | Maria Mikaela Chester | https://www.jotform.com/uploads/macarewecare/222927978372067/5953639149101922812/5953639149101922812_signature_31.png
| 2024-06-28 | 2024-06-28 13:25 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28712 | UNOPS | bmclab_micro@yahoo.com | 35 |  |  | GXIV 10C | 110023507 | Unit for installation | NA | #NAME? | Unit is operational Unit is not recommended for use until air conditioning unit is repaired/installed (Room temperature exceeds optimal ambient temperature of below 30C) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953639149102127875/5953639149102127875_signature_28.png | Mary Joy Marquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5953639149102127875/5953639149102127875_signature_31.png
| 2024-06-28 | 2024-06-25 12:31 | Guiguinto Rural Health Unit | Guiguinto, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28710 | PBSP | joie_calica@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 460 ASSAY: Xpert HIV-1 Viral Load, COUNT: 179 | 2024-06-25 | GXIV | 814165 | Unit for maintenance Module B3(696755) for replacement | Replaced module B3 | #NAME? | Unit os operational |  | GeneXpert Module SN: 716345 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953639149102427219/5953639149102427219_signature_28.png | Ligaya Calica | https://www.jotform.com/uploads/macarewecare/222927978372067/5953639149102427219/5953639149102427219_signature_31.png
| 2024-06-28 | 2024-06-27 16:49 | Southern Palawan Provincial Hospital | Brookes Point, Palawan | 4B | Palawan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28709 | PBSP | spphlaboratory2023@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1855 | 2024-04-29 | GXIv 6C L | 804271 | Modules B3 and B4 for replacement | Performed preventive maintenance Replaced modules b3 and b4 Troubleshoot error on module b4 Problem still persists Module B4 for reporting | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced module B3 and B4 B3: 761307 -> 711217 B4: 707921 -> 702990 Module B4 error 4017 upon installation Perform slef test; problem persists Troubleshoot module by switching to other module slot Error still persists Module B4 for replacement | Unit is operational 3 modules working B4 for reporting |  | gX modules 6C | PL15740 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953597017512401922/5953597017512401922_signature_28.png | Mariah Paula Bolen | https://www.jotform.com/uploads/macarewecare/222927978372067/5953597017512401922/5953597017512401922_signature_31.png
| 2024-06-28 | 2024-06-27 10:10 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28708 | PBSP | lcpmicrolab@gmail.com | 40 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6120 ASSAY: Xpert SARS-CoV-2, COUNT: 1029 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Others, COUNT: 2 |  | GX 16 L | 801080 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953592762558144759/5953592762558144759_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5953592762558144759/5953592762558144759_signature_31.png
| 2024-06-28 | 2024-06-27 09:53 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28707 | COMMERCIAL | lcpmicrolab@gmail.com | 29 |  |  | GX4 L | 804205 | D2,D3,D4 disabled error 5007 Last xpertcheck 6.30.21 Def.UPS | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | D2,D3,D4 disabled D1 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953591832558872737/5953591832558872737_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5953591832558872737/5953591832558872737_signature_31.png
| 2024-06-28 | 2024-06-27 09:49 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28706 | MACARE-HBDC | lcpmicrolab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 423 ASSAY: Xpert SARS-CoV-2, COUNT: 1162 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 113 ASSAY: Others, COUNT: 3 |  | GX4 D | 838837 | Due for maintenance B1 and B4 reported already encounter temp error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B1 and B4 disabled B2 and B3 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953591162553603433/5953591162553603433_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5953591162553603433/5953591162553603433_signature_31.png
| 2024-06-28 | 2024-06-27 08:53 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Tb dots (koica) | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28705 | PBSP | nilabis.pbsp@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 356 |  | GX4 L | 707450 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953590622555467266/5953590622555467266_signature_28.png | Norman Charles Labis | https://www.jotform.com/uploads/macarewecare/222927978372067/5953590622555467266/5953590622555467266_signature_31.png
| 2024-06-28 | 2024-06-27 08:46 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Tb dots (koica) | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28704 | PBSP | nilabis.pbsp@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 741 |  | GX4 L | 704584 | Due for maintenance, defective lan cable | Performed troubleshooting Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Troubleshooting and maintenance | Operational | Replaced temporary lan cable provided of end user |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953590022559435933/5953590022559435933_signature_28.png | Norman Charles Labis | https://www.jotform.com/uploads/macarewecare/222927978372067/5953590022559435933/5953590022559435933_signature_31.png
| 2024-06-28 | 2024-06-27 12:44 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 28703 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 59 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert HPV, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 | 2024-06-25 | GX4 L | 807155 | For module replacement C1 654545 - 210101947 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | 3 module operational |  | Genexpert 6 color modules | 11681 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953589662552313661/5953589662552313661_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/5953589662552313661/5953589662552313661_signature_31.png
| 2024-06-28 | 2024-06-28 12:36 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28702 | DOH | lornsabaigat@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 832 | 2024-06-28 | GXIV 6C | 838409 | n/a | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953587425717908866/5953587425717908866_signature_28.png | Loran C. Abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/5953587425717908866/5953587425717908866_signature_31.png
| 2024-06-28 | 2024-06-28 12:19 | Antipolo City Health Office | M. Santos St. Antipolo City | 4A | Rizal | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28701 | DOH | acho.laboratory@gmail.com | 23 |  |  | GX-IV-D-10C | 110023496 | For installation | Installed the GeneXpert System | Installed the GeneXpert System -GeneXpert unit -Workstation -UPS -Printer -Barcode scanner with stand Functionality test Conduct end-user training | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953582049812555358/5953582049812555358_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5953582049812555358/5953582049812555358_signature_31.png
| 2024-06-28 | 2024-06-27 10:16 | Rizal Medical Center | Pasig Blvd., Pasig City | NCR | NCR 2nd - Pasig | Molecular Laboratory | NO | REPAIR | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28700 | PBSP | molecularlab@rmc.doh.gov.ph | 22 | ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4300 ASSAY: Xpert SARS-CoV-2, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 86 | 2024-06-27 | GX-IV-D-6C | 814374 | Module B3 for replacement | Replaced module B3 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B3 Checked module B4 Generated reports | Unit is operational. Modules B1, B2 and B3 operational. Module B4/645415 not operational, for reporting | Module B4 has torn ultrasonic horn cable. | Old: B3/210068672 New: B3/705609 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953575409811990068/5953575409811990068_signature_28.png | Jhon Patrick Ariola | https://www.jotform.com/uploads/macarewecare/222927978372067/5953575409811990068/5953575409811990068_signature_31.png
| 2024-06-28 | 2024-06-27 10:22 | Rizal Medical Center | Pasig Blvd., Pasig City | NCR | NCR 2nd - Pasig | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28699 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 | ASSAY: Others, COUNT: 2 ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 11 ASSAY: Xpert SARS-CoV-2, COUNT: 1443 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-XVI-D-6C | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules A1, A2, B3, B4, C2, C3, D2 and D4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953575409817407973/5953575409817407973_signature_28.png | Jhon Patrick Ariola | https://www.jotform.com/uploads/macarewecare/222927978372067/5953575409817407973/5953575409817407973_signature_31.png
| 2024-06-28 | 2024-06-28 11:51 | NORTHERN MINDANAO MEDICAL CENTER | CAGAYAN DE ORO, MISAMIS ORIENTAL | 10 | Misamis Oriental | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28698 | PBSP | nmmc.gxmtbrif@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 346 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7683 ASSAY: Xpert SARS-CoV-2, COUNT: 7682 ASSAY: Xpert HIV-1 Viral Load, COUNT: 628 |  | GX-XVI | 844118 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE REPLACEMENT: MODULE A2 DEF SN:688804 / NEW SN: 710021 MODULE B3 DEF SN:683687 / NEW SN: 735112 MODULE B4 DEF SN:210102261 / NEW SN: 793590 MODULE C1 DEF SN:766096 / NEW SN:806405 MODULE C2 DEF SN:693907 / NEW SN:874960 MODULE C3 DEF SN:210063174 / NEW SN:881066 MODULE C4 DEF SN:906389 / NEW SN:896365 MODULE D1 DEF SN:691194 / NEW SN:903893 MODULE D2 DEF SN:677101 / NEW SN:210036382 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953566463537770450/5953566463537770450_signature_28.png | Rione Frances Lomongo | https://www.jotform.com/uploads/macarewecare/222927978372067/5953566463537770450/5953566463537770450_signature_31.png
| 2024-06-28 | 2024-06-24 11:40 | Toledo City Health Office | Toledo City | 7 | Cebu | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 28697 | PBSP | loriemaediapen@yahoo.com | 23 |  | 2024-06-24 | Gx4 | 808640 | Module B3 Hardware fail. | Replace module defective. | Replace Module B3 with new module serial number 629652. | Module B1, 2,3 ok to use module B4 need to replace due to error's. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953540362167508843/5953540362167508843_signature_28.png | Lorie Mae C. Diapen | https://www.jotform.com/uploads/macarewecare/222927978372067/5953540362167508843/5953540362167508843_signature_31.png
| 2024-06-27 | 2024-06-27 21:42 | Calbayog City Health Office | Calbayog Samar | 8 | Samar | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28696 | PBSP | aaron_noel122770sd@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1976 | 2024-06-27 | GXIV 6C | 804265 | Module B2/632085 defective | #NAME? | - Module replacement B2/632085 into 619845 - Preventive maintenance (Cleaned plunger rod, Cartridge bay, iCore and Fan Filter) - Performed functionality test - Generate IQR and System log | Unit is operational |  | 1 Genexpert Module 6C | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953051363024055173/5953051363024055173_signature_28.png | Ma. Leonora S. Dejapin | https://www.jotform.com/uploads/macarewecare/222927978372067/5953051363024055173/5953051363024055173_signature_31.png
| 2024-06-27 | 2024-06-27 21:28 | Marawi City Health Office | Marawi City | BARMM | Lanao del Sur | LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28695 | PBSP | faridahmamao@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 507 |  | GX-IV | 838293 | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953044643414498739/5953044643414498739_signature_28.png | Faridah M. Busara, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5953044643414498739/5953044643414498739_signature_31.png
| 2024-06-27 | 2024-06-27 21:20 | Marawi City Health Office | Marawi City | BARMM | Lanao del Sur | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28694 | DOH | faridahmamao@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1249 |  | GX-IV | 838260 | -FOR PREVENTIVE MAINTENANCE -MODULE B4 ERROR 1002 | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953040723419862714/5953040723419862714_signature_28.png | Faridah M. Busara, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5953040723419862714/5953040723419862714_signature_31.png
| 2024-06-27 | 2024-06-27 21:16 | Amai Pakpak Medical Center RTDL | Marawi City | BARMM | Lanao del Sur | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28693 | WHO | rbarco2007@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2036 ASSAY: Xpert HIV-1 Viral Load, COUNT: 32 |  | GX-IV | 847303 | #NAME? | #NAME? | -PERFORM CLEANING OF GX WORKSTATION AND PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B1 DEF SN:911082 REPLACEMENT WITH NEW SN:210059851 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953036279177075555/5953036279177075555_signature_28.png | Raisah P. Barco | https://www.jotform.com/uploads/macarewecare/222927978372067/5953036279177075555/5953036279177075555_signature_31.png
| 2024-06-27 | 2024-06-27 21:10 | Amai Pakpak Medical Center RTDL | Marawi City | BARMM | Lanao del Sur | LABORATORY | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28692 | PBSP | rbarco2007@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 810 |  | GX-IV | 807457 | #NAME? | #NAME? | -PERFORM CLEANING OF GENEXPERT WORKSTATION -PERFORM GENEXPERT PREVENTIVE MAINTENANCE PROCEDURE -PERFORM GATEWAY BOARD REPLACEMENT AND SERIAL NUMBER CONFIG -PERFORM MODULE C1 DEF SN: 655108 REPLACEMENT WITH NEW SN:634832 -INSTALL NEW PRINTER DRIVER | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5953033689178592740/5953033689178592740_signature_28.png | Raisah P. Barco | https://www.jotform.com/uploads/macarewecare/222927978372067/5953033689178592740/5953033689178592740_signature_31.png
| 2024-06-27 | 2024-06-27 12:55 | Taal Rural Health Unit | Tierra Alta, Taal, Batangas | 4A | Batangas | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28691 | PBSP | hanetmanny@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1918 |  | GXIV-D-6C | 805259 | Module 4 Failed in XpertCheck. | Repair. | Replaced Module 4. Performed Self-test. Generate Reports. | New Module with 4004 Error (H/W Failed). 3 Modules are Operational. | B4: 761622 -> 673300 for reporting. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952942116911193514/5952942116911193514_signature_28.png | Juanita B Balbacal | https://www.jotform.com/uploads/macarewecare/222927978372067/5952942116911193514/5952942116911193514_signature_31.png
| 2024-06-27 | 2024-06-27 11:06 | San Luis Rural Health Unit | Poblacion, San Luis, Batangas | 4A | Batangas | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28690 | PBSP | ilao.rochell@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 60 |  | GXIV-D-6C | 838131 | Module 1 Failed in XpertCheck. | Repair | Performed Self-test. Replaced Module 1. Generate Reports. | 3 Modules are Operational, 1 Disabled Module. | B1: 806301 -> 679767 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952941706915228015/5952941706915228015_signature_28.png | Rochell C. Ilao | https://www.jotform.com/uploads/macarewecare/222927978372067/5952941706915228015/5952941706915228015_signature_31.png
| 2024-06-27 | 2024-06-26 15:51 | Ospital ng Muntinlupa | Civic Drive, FCC, Alabang, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | osmunmicro@yahoo.com |  | 28689 | DOH | osmunclinlab@gmail.com | 26 |  | 2024-06-20 | GXIV-D-10C | 110023494 | Unit is due for Installation. | Installation. | Checked Outlet Voltage. Installed GeneXpert System. Performed Self-test. Generate Reports. | GX System is Functional. |  | GeneXpert System | DR 11624 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952941236913340109/5952941236913340109_signature_28.png | Claire P. Bersamina | https://www.jotform.com/uploads/macarewecare/222927978372067/5952941236913340109/5952941236913340109_signature_31.png
| 2024-06-27 | 2024-06-27 14:37 | Balayan Rural Health Unit | Brgy Gumamela, Balayan, Batangas | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28688 | PBSP | veronica.dalida@yahoo.com | 25 |  |  | GXIV-D-6C | 838451 | CPU with Hard Drive Problem. Unit is due for XpertCheck. | XpertCheck. CPU Replacement. | Replaced CPU. Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed XpertCheck. Generate Reports. | 2 Modules Newly Replaced, 1 Service Required, 1 Passed | B4: 861386 - Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952911391837982132/5952911391837982132_signature_28.png | Veronica M. Dalida | https://www.jotform.com/uploads/macarewecare/222927978372067/5952911391837982132/5952911391837982132_signature_31.png
| 2024-06-27 | 2024-06-27 15:01 | Cotabato Sanitarium and General Hospital | Sultan Kudarat, Maguindanao | BARMM | Maguindanao | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28687 | PBSP | normiaando.tahir@gmail.com | 21.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1257 | 2024-06-27 | Gx-IV (6C) | 815277 | Module 1 for Service Required | Replace defective module 1 SN:801637 with new module SN:210074471 Genexpert Cleaning/Preventive maintenance Checking of errors Data collection | Perform Genexpert machine module repair Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6-color | Cepheid case:1542566/1542591 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952858357276873969/5952858357276873969_signature_28.png | Normia A. Tahir | https://www.jotform.com/uploads/macarewecare/222927978372067/5952858357276873969/5952858357276873969_signature_31.png
| 2024-06-27 | 2024-06-27 15:53 | Dr. Paulini J. Garcia Memorial Research and Medical Center | Mabini, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28686 | UNOPS | bryanmichaelbalunes@gmail.com | 25 |  | 2024-06-20 | GXIV-10C | 110023508 | N/A | N/a | Installed UPS Installed Genexpert Installed Computer Installed Printer Installed Monitor | Unit is Operational |  |  | 11633 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952845022251809002/5952845022251809002_signature_28.png | Bryan Michael Balunes | https://www.jotform.com/uploads/macarewecare/222927978372067/5952845022251809002/5952845022251809002_signature_31.png
| 2024-06-27 | 2024-06-27 14:30 | HIMAMAYLAN CITY HEALTH OFFICE | Himamaylan City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 28685 | PBSP | fayemariejuguan@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1100 |  | GX-4 | 838138 | Defective Module 673308/B2 and 806690/B3. | Preventive Maintenance - optics cleaning, filter cleabing, module bay cleabing, plunger rod maintenance. | Replaced module B2 and B3 with new modules SN 661690 and 664077 respectively. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952794028572181959/5952794028572181959_signature_28.png | Faye Marie Juguan | https://www.jotform.com/uploads/macarewecare/222927978372067/5952794028572181959/5952794028572181959_signature_31.png
| 2024-06-27 | 2024-06-27 13:55 | Talaingod Rural Health Unit | JBL Brgy. Sto Niño, Talaingod | 11 | Davao del Norte | TB DOTS | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28684 | PBSP | denniseolayan@gmail.com | 24.5 |  |  | GX-IV 4D | 838859 | System Unit CPU not working For Preventive Maintenance and XpertCheck Calibration | Replaced the Defective CPU System unit SN:CPH0420F07731 with new CPU System Unit SN:CPH1023F46029 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR Install XPERTCHECK 1.5B and run XpertCheck on four modules | Perform CPU System Unit replacement Perform Preventive Maintenance and XpertCheck Calibration | Module B1 and B2:Passed Module B3: retry and Passed Module B4: Retry and Skipped ;no available xpc cartidge Module 4 for under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952772214624131385/5952772214624131385_signature_28.png | Christine Dennise Olayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5952772214624131385/5952772214624131385_signature_31.png
| 2024-06-27 | 2024-06-27 13:45 | Legazpi City Health Office | Rizal st., Legazpi City | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28683 | WHO | janzchel_neneh10@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1245 |  | GX4-6C | 846354 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate report | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952766288425087323/5952766288425087323_signature_28.png | Janzhel marie chan | https://www.jotform.com/uploads/macarewecare/222927978372067/5952766288425087323/5952766288425087323_signature_31.png
| 2024-06-27 | 2024-06-21 15:31 | Baras Medicare & Community Hospital | Bagong Sirang, Baras | 5 | Catanduanes | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28682 | PBSP | malustesorerormt@gmail.com | 21 |  |  | GX4-10C | 110020684 | GX modules not detected | Checked UTP cable connections checked/adjust network settings Checked power cable Opened Gx and reset cable connections Cleaned cartridge bays, plunger rod Generate report | Performed Troubleshooting | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952762368424391383/5952762368424391383_signature_28.png | Malu S.Tesorero,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5952762368424391383/5952762368424391383_signature_31.png
| 2024-06-27 | 2024-06-19 15:03 | Libmanan Rural Health Unit ll | San Isidro, Libmanan | 5 | Camarines Sur | Laboratory | NO | REPAIR | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28681 | PBSP | amyquiapo1972@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1261 |  | GX4-6C | 805260 | Monitor for replacement | Set up and replaced with new Monitor (CN-0WMYC5-FCC00-36T-FK4X REV A05) Generate report | Replaced Monitor | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952761878426489968/5952761878426489968_signature_28.png | Amy Quiapo | https://www.jotform.com/uploads/macarewecare/222927978372067/5952761878426489968/5952761878426489968_signature_31.png
| 2024-06-27 | 2024-06-27 10:27 | Don Emilio Del Valle Memorial Hospital | Ubay, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28680 | PBSP | Laboratory@devmh.doh.gov.ph | 24 |  |  | Gx 4 | 807603 | Defective Module A2 734887 | New module A2 647493 Clean Fan Filter Clean Fan Gather data IQR | New module A2 647493 Clean Fan Filter Clean Fan Gather data IQR | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952653096216497393/5952653096216497393_signature_28.png | Roger Lagura RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5952653096216497393/5952653096216497393_signature_31.png
| 2024-06-27 | 2024-06-27 10:27 | Don Emilio Del Valle Memorial Hospital | Ubay, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28679 | PBSP | Laboratory@devmh.doh.gov.ph | 24 |  |  | Gx 4 | 807603 | Defective Module A2 734887 | New module A2 647493 Clean Fan Filter Clean Fan Gather data IQR | New module A2 647493 Clean Fan Filter Clean Fan Gather data IQR | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952651576213158619/5952651576213158619_signature_28.png | Roger Lagura RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5952651576213158619/5952651576213158619_signature_31.png
| 2024-06-27 | 2024-06-27 10:27 | Don Emilio Del Valle Memorial Hospital | Ubay, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28678 | PBSP | Laboratory@devmh.doh.gov.ph | 24 |  |  | Gx 4 | 807603 | Defective Module A2 734887 | New module A2 647493 Clean Fan Filter Clean Fan Gather data IQR | New module A2 647493 Clean Fan Filter Clean Fan Gather data IQR | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952651576215102477/5952651576215102477_signature_28.png | Roger Lagura RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5952651576215102477/5952651576215102477_signature_31.png
| 2024-06-27 | 2024-06-27 10:27 | Don Emilio Del Valle Memorial Hospital | Ubay, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28677 | PBSP | Laboratory@devmh.doh.gov.ph | 24 |  |  | Gx 4 | 807603 | Defective Module A2 734887 | New module A2 647493 Clean Fan Filter Clean Fan Gather data IQR | New module A2 647493 Clean Fan Filter Clean Fan Gather data IQR | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952649416216728384/5952649416216728384_signature_28.png | Roger Lagura RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5952649416216728384/5952649416216728384_signature_31.png
| 2024-06-27 | 2024-06-27 09:44 | Philippine General Hospital - UPM | Taft Ave. Manila | NCR | NCR 1st - Manila | Microscopy | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28676 | PBSP | agllacuna@up.edu.ph | 25 |  |  | GXIV | 838145 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952626153846686854/5952626153846686854_signature_28.png | Allyana G. Llacuna | https://www.jotform.com/uploads/macarewecare/222927978372067/5952626153846686854/5952626153846686854_signature_31.png
| 2024-06-27 | 2024-06-27 09:49 | Philippine General Hospital - UPM | Taft Ave., Manila | NCR | NCR 1st - Manila | Microscopy | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28675 | PBSP | agllacuna@up.edu.ph | 25 |  |  | GXIV | 802511 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952625603841529917/5952625603841529917_signature_28.png | Allyana G. Llacuna | https://www.jotform.com/uploads/macarewecare/222927978372067/5952625603841529917/5952625603841529917_signature_31.png
| 2024-06-27 | 2024-06-27 09:37 | Don Carlos Rural Health Unit | Poblacion, Don Carlos | 10 | Bukidnon | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28674 | PBSP | doncarlosrhu@gmail.com | 30 |  |  | GX-4 6 Colors | 807625 | Old Monitor not working. | None | Replacing LG Monitor with SN: 504NDQA1N308 with New Dell Monitor E222HS with SN: CN-0WMYC5-FCC00-36T-FK8X-A05 | All modules and units are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952619513192503581/5952619513192503581_signature_28.png | Gilbert E. Loro | https://www.jotform.com/uploads/macarewecare/222927978372067/5952619513192503581/5952619513192503581_signature_31.png
| 2024-06-27 | 2024-06-26 14:15 | Mauban Molecular Laboratory | Signal St. Mauban, Quezon | 4A | Quezon | Laboratory | NO | REPAIR | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28673 | MACARE-HBDC | rhumauban.laboratory@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 227 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1260 |  | GX-IV-D-10C | 110015487 | Stuck cartridge on module B2 | Removed stuck cartridge | Manual removal of stuck cartridge Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952607081035021370/5952607081035021370_signature_28.png | HAZEL JOY G. PARDINWA | https://www.jotform.com/uploads/macarewecare/222927978372067/5952607081035021370/5952607081035021370_signature_31.png
| 2024-06-27 | 2024-06-26 15:10 | Mauban Rural Health Unit | Mauban, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28672 | PBSP | rhumauban.laboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 232 |  | GX-IV-D-6C | 814375 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B4/620410 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952607031031039011/5952607031031039011_signature_28.png | HAZEL JOY G. PARDINES | https://www.jotform.com/uploads/macarewecare/222927978372067/5952607031031039011/5952607031031039011_signature_31.png
| 2024-06-27 | 2024-06-25 12:29 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28671 | MACARE-HBDC | peachmango318@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX-IV-D-10C | 110024245 | For Relocation | Installed the GeneXpert System | Installed the GeneXpert System: -GeneXpert unit -Workstation -Barcode scanner with stand -UPS -Printer Refresher training for basic operation and troubleshooting | Unit is operational. All modules operational. | Transferred from Santolan SHC. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952607031031443225/5952607031031443225_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/5952607031031443225/5952607031031443225_signature_31.png
| 2024-06-27 | 2024-06-25 12:29 | Pinagbuhatan TB Microscopy Laboratory | Urbano Velasco Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28670 | MACARE-HBDC | peachmango318@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX-IV-D-10C | 110024245 | For Relocation | Installed the GeneXpert System | Installed the GeneXpert System: -GeneXpert unit -Workstation -Barcode scanner with stand -UPS -Printer Refresher training for basic operation and troubleshooting | Unit is operational. All modules operational. | Transferred from Santolan SHC. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952606781034409774/5952606781034409774_signature_28.png | Arlen Delgado | https://www.jotform.com/uploads/macarewecare/222927978372067/5952606781034409774/5952606781034409774_signature_31.png
| 2024-06-27 | 2024-06-26 14:15 | Mauban Molecular Laboratory | Signal St. Mauban, Quezon | 4A | Quezon | Laboratory | NO | REPAIR | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28669 | MACARE-HBDC | rhumauban.laboratory@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 227 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1260 |  | GX-IV-D-10C | 110015487 | Stuck cartridge on module B2 | Removed stuck cartridge | Manual removal of stuck cartridge Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952606781037191076/5952606781037191076_signature_28.png | HAZEL JOY G. PARDINWA | https://www.jotform.com/uploads/macarewecare/222927978372067/5952606781037191076/5952606781037191076_signature_31.png
| 2024-06-27 | 2024-06-26 15:10 | Mauban Rural Health Unit | Mauban, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28668 | PBSP | rhumauban.laboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 232 |  | GX-IV-D-6C | 814375 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B4/620410 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952606761033770328/5952606761033770328_signature_28.png | HAZEL JOY G. PARDINES | https://www.jotform.com/uploads/macarewecare/222927978372067/5952606761033770328/5952606761033770328_signature_31.png
| 2024-06-27 | 2024-06-26 14:53 | San Mateo Super Health Center | Guitnang bayan San Mateo | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28667 | PBSP | wilhelminanartatez@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 641 |  | GX4 D | 838161 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952593600802068434/5952593600802068434_signature_28.png | Wilhelmina Nartatez | https://www.jotform.com/uploads/macarewecare/222927978372067/5952593600802068434/5952593600802068434_signature_31.png
| 2024-06-27 | 2024-06-26 14:34 | San Mateo Super Health Center | Guitnang bayan,San Mateo | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28666 | PBSP | wilhelminanartatez@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1710 | 2024-06-21 | GX4 D | 808616 | For module replacement A2 621048 - 698618 Due for maintenance | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | Genexpert 6 color module | 48859 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952593080807659368/5952593080807659368_signature_28.png | Wilhelmina Nartatez | https://www.jotform.com/uploads/macarewecare/222927978372067/5952593080807659368/5952593080807659368_signature_31.png
| 2024-06-27 | 2024-06-27 07:54 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | YES | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28665 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 11J219 | Error 413 Valve Problem "purple valve test failure" Display Monitor input key Draw,Prime, and Return, etc malfunction. | Replaced temporarily the pneumatic assembly while waiting for the purchase order of the parts to avoid interruptions of their apheresis procedures. - troubleshoot cable, display monitor | Replaced temporarily the pneumatic assembly while waiting for the purchase order of the parts to avoid interruptions of their apheresis procedures. | Pneumatics assembly already replaced. Only the display monitor input keyboard problem. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952556394221510971/5952556394221510971_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/5952556394221510971/5952556394221510971_signature_31.png
| 2024-06-27 | 2024-06-26 10:56 | Calumpang Rural Health Unit | Brgy. Calumpang General Santos City | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28664 | DOH | jan19.cam@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1469 |  | Gx-IV (6C) | 838504 | N/A | Genexpert Cleaning/Preventive maintenance Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952506719492626052/5952506719492626052_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5952506719492626052/5952506719492626052_signature_31.png
| 2024-06-27 | 2024-06-26 10:56 | Calumpang Rural Health Unit | Brgy. Calumpang General Santos City | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28663 | DOH | jan19.cam@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1469 |  | Gx-IV (6C) | 838504 | N/A | Genexpert Cleaning/Preventive maintenance Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952505927146629142/5952505927146629142_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5952505927146629142/5952505927146629142_signature_31.png
| 2024-06-26 | 2024-06-26 20:52 | Eastern Visayas Medical Center | Brgy.93 Bagacay, Tacloban City | 8 | Leyte | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28662 | PBSP | merl_agner62@yahoo.com.sg | 25 |  | 2024-06-25 | GXIV-6C | 808631 | Module B3/685252 defective | #NAME? | - Module replacement B3/685252 into 661605 - Performed preventive maintenance (Cleaned plunger rod, cartridge bay, iCore and Fan Filter) - Functionality test - Generate IQR and System log | Unit is operational |  | 1 Genexpert Module 6C | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952158343023468361/5952158343023468361_signature_28.png | Ma. Merlina Vistal | https://www.jotform.com/uploads/macarewecare/222927978372067/5952158343023468361/5952158343023468361_signature_31.png
| 2024-06-26 | 2024-06-26 17:55 | Iligan Society of Internists | Iligan City | 10 | Lanao del Norte | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28661 | PBSP | dayolaermajessica@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3227 ASSAY: Xpert HIV-1 Viral Load, COUNT: 124 |  | GX-IV | 707842 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952050774421374090/5952050774421374090_signature_28.png | Erma Jessica Dayola | https://www.jotform.com/uploads/macarewecare/222927978372067/5952050774421374090/5952050774421374090_signature_31.png
| 2024-06-26 | 2024-06-26 17:47 | Iligan Society of Internists | Iligan City | 10 | Lanao del Norte | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28660 | PBSP | dayolaermajessica@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3227 ASSAY: Xpert HIV-1 Viral Load, COUNT: 124 |  | GX-IV | 805090 | #NAME? | #NAME? | #NAME? | EQUIPMENT OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5952048304421391688/5952048304421391688_signature_28.png | Erma Jessica Dayola | https://www.jotform.com/uploads/macarewecare/222927978372067/5952048304421391688/5952048304421391688_signature_31.png
| 2024-06-26 | 2024-02-16 12:56 | Bicol Region General Hospital and Geriatric Medical Center | San Pedro, Cabusoa | 5 | Camarines Norte | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28659 | PBSP | jlvdvera_rmt@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1851 |  | GX4-6C | 807601 | For Preventive Maintenance UPS shuts down when having power outage | Cleaned Cartridge bays, plunger rod, icore Checked dependability of UPS when unplugged to AC power source(UPS was charged for an hour) Checked power connections going to UPS tested power output of UPS Generate reports | Performed Preventive Maintenance Performed Troubleshooting | Unit is operational UPS defective backup battery | UPS request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838802519417118770/5838802519417118770_signature_28.png | JESSIE LOwell de vera | https://www.jotform.com/uploads/macarewecare/222927978372067/5838802519417118770/5838802519417118770_signature_31.png
| 2024-06-26 | 2024-06-26 14:57 | Malay Primary Care Facility | Caticlan, Malay, Aklan | 6 | Aklan | Tb-Dots | NO | TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28658 | DOH | mhomalayclinicallaboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1016 |  | 6C GXIV | 838457 | Module B2 frequent H/W error | None | #NAME? | All modules are functional. Module B2 passed XPC. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951951250496560938/5951951250496560938_signature_28.png | May Soledad | https://www.jotform.com/uploads/macarewecare/222927978372067/5951951250496560938/5951951250496560938_signature_31.png
| 2024-06-26 | 2024-06-26 14:52 | Mexico Rural Health Unit IV | Pandacaqui, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28657 | DOH | annbathan79@gmail.com | 25 |  |  | GXIV-6C | 831902 | UPS not working CPU not powering on | Troubleshoot UPS Check Fuse Trouble shoot CPU Check Bios Battery | Performed Maintenance Gathered Reports Performed Functionality Test | Unit not Functional | UPS and Computer for Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951944212118039596/5951944212118039596_signature_28.png | Mary Ann D. Carreon | https://www.jotform.com/uploads/macarewecare/222927978372067/5951944212118039596/5951944212118039596_signature_31.png
| 2024-06-26 | 2024-06-26 13:42 | Arayat Rural Health Unit | Arayat, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28656 | PBSP | baylonrochelle06@yahoo.com | 22 |  |  | GXIV-6C | 812790 | All Modules error 2126 | Reset Gatewayboard | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | All Modules for Observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951899492118075821/5951899492118075821_signature_28.png | Rochelle Baylon | https://www.jotform.com/uploads/macarewecare/222927978372067/5951899492118075821/5951899492118075821_signature_31.png
| 2024-06-26 | 2024-06-26 09:57 | Navotas Medical and Wellness Center | M. Naval St. Navotas City | NCR | NCR 3rd - Navotas | Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28655 | PBSP | alejandrinoknowell@gmail.com | 25 |  |  | GXIV-10C | 110020388 | Unit for installation. | Installation. | Performed installation of GeneXpert system. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951881036608399068/5951881036608399068_signature_28.png | Noel Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5951881036608399068/5951881036608399068_signature_31.png
| 2024-06-26 | 2024-06-26 09:22 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | Tb-Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28654 | DOH | sanjoseantiquegxsite@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 733 ASSAY: Xpert HIV-1 Viral Load, COUNT: 85 |  | 6C GXIV | 838412 | Due for PM | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951752021507154131/5951752021507154131_signature_28.png | Oscar Delano Salazar, IV | https://www.jotform.com/uploads/macarewecare/222927978372067/5951752021507154131/5951752021507154131_signature_31.png
| 2024-06-26 | 2024-06-26 09:30 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | Tb-Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28653 | PBSP | sanjoseantiquegxsite@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 733 ASSAY: Xpert HIV-1 Viral Load, COUNT: 85 | 2024-06-07 | 10C GXIV | 120000591 | Module A1 defective | -Replace defective module A1/220013550 with module SN: 860815 | #NAME? | All modules are functional. |  | 900-0491RS | 48687 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951751901503205873/5951751901503205873_signature_28.png | Oscar Delano Salazar, IV | https://www.jotform.com/uploads/macarewecare/222927978372067/5951751901503205873/5951751901503205873_signature_31.png
| 2024-06-26 | 2024-06-25 07:55 | Tondo Health Center | Juan Luna St. Gagalangin Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28652 | DOH | malorinalontoc@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1713 |  | GXIV | 831620 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951749412819195155/5951749412819195155_signature_28.png | Lontoc, ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5951749412819195155/5951749412819195155_signature_31.png
| 2024-06-26 | 2024-06-25 08:14 | Juan Posadas Health Center | Rodriguez St., Balut Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28651 | PBSP | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 485 |  | GXIV | 805060 | Unit for preventive maintenance. For checking of barcode scanner. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed troubleshooting of barcode scanner. Generated reports. | Unit is operational. Barcode scanner sometimes unable to detect and connection loss reoccurs after several minutes. Barcode scanner for reporting. | Barcode Scanner (Symbol SN: 1324600508421) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951749312817034854/5951749312817034854_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/5951749312817034854/5951749312817034854_signature_31.png
| 2024-06-26 | 2024-06-20 11:36 | San Pablo Colleges Medical Center | San Pablo Laguna | 4A | Laguna | Molecular laboratory | NO | Relocation | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28650 | CONSORTIUM | spcmc.molecularlab@gmail.com | 25 |  |  | GXIV-10C | 110024552 | Unit for relocation. | Relocation and installation. | Performed relocation of GeneXpert to another room. Performed installation. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951748652814706765/5951748652814706765_signature_28.png | Mark Lenard Moreno | https://www.jotform.com/uploads/macarewecare/222927978372067/5951748652814706765/5951748652814706765_signature_31.png
| 2024-06-26 | 2024-06-20 10:07 | Alaminos Rural Health Unit | Alaminos Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28649 | PBSP | irahrobles@gmail.com | 25 |  | 2024-06-20 | GXIV | 838447 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951748342814022427/5951748342814022427_signature_28.png | Iralyn M. Robles | https://www.jotform.com/uploads/macarewecare/222927978372067/5951748342814022427/5951748342814022427_signature_31.png
| 2024-06-26 | 2024-06-25 07:55 | Tondo Health Center | Juan Luna St. Gagalangin Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28648 | DOH | malorinalontoc@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1713 |  | GXIV | 831620 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951745632814338103/5951745632814338103_signature_28.png | Lontoc, ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5951745632814338103/5951745632814338103_signature_31.png
| 2024-06-26 | 2024-06-20 11:36 | San Pablo Colleges Medical Center | San Pablo Laguna | 4A | Laguna | Molecular laboratory | NO | Relocation | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28647 | CONSORTIUM | spcmc.molecularlab@gmail.com | 25 |  |  | GXIV-10C | 110024552 | Unit for relocation. | Relocation and installation. | Performed relocation of GeneXpert to another room. Performed installation. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951745632812196879/5951745632812196879_signature_28.png | Mark Lenard Moreno | https://www.jotform.com/uploads/macarewecare/222927978372067/5951745632812196879/5951745632812196879_signature_31.png
| 2024-06-26 | 2024-06-25 08:14 | Juan Posadas Health Center | Rodriguez St., Balut Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28646 | PBSP | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 485 |  | GXIV | 805060 | Unit for preventive maintenance. For checking of barcode scanner. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed troubleshooting of barcode scanner. Generated reports. | Unit is operational. Barcode scanner sometimes unable to detect and connection loss reoccurs after several minutes. Barcode scanner for reporting. | Barcode Scanner (Symbol SN: 1324600508421) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951745632814568783/5951745632814568783_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/5951745632814568783/5951745632814568783_signature_31.png
| 2024-06-26 | 2024-06-25 14:37 | Quezon City Jail | Payatas QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28645 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1417 |  | GX4 D | 803089 | For maintenance and for replacement of UPS | Replaced and installed UPS 1kva Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | UPS installation Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951714093613814975/5951714093613814975_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5951714093613814975/5951714093613814975_signature_31.png
| 2024-06-25 | 2024-06-25 14:03 | Dr. Cornelio T. Martinez SR. Memorial Hospital | Kiamba, Sarangani | 12 | Sarangani | Laboratory | YES | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | dcmmh_lab@yahoo.com |  | 28644 | PBSP | martineznoemi11281965@gmail.com | 29.1 |  |  | Gx-IV (10C) | 120000498 | Defective UPS upon receiving/delivered on site | Installed: Genexpert machine SN:120000498 Cpu SN: *CPH0623F42165* Monitor SN: V5FTG679 Printer SN: CNB1RBK7RQ UPS (defective) Import assay MTB-ULTRA and MTB-XDR | Perform Genexpert machine Installation | 4 operational modules all accessories functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951242184214678106/5951242184214678106_signature_28.png | Noemi S. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5951242184214678106/5951242184214678106_signature_31.png
| 2024-06-25 | 2024-06-25 15:47 | Cabanatuan City Health Office | Kapt Pepe, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28643 | DOH | cholab2021@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2369 |  | GXIV-6C | 831592 | Module A3 Hardware Failure | Gathered Reports for Reporting | Performed Maintenance Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951115415619762878/5951115415619762878_signature_28.png | Maricris So | https://www.jotform.com/uploads/macarewecare/222927978372067/5951115415619762878/5951115415619762878_signature_31.png
| 2024-06-25 | 2024-06-25 12:58 | Quezon City General Hospital | Seminary Rd, Project 8, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28642 | UNOPS | jomarilitonjua@gmail.com | 27 |  | 2024-06-20 | GX IV 10C D | 110023495 | Unit due for installation | Installed GX and accessories Gave training to end user Gave exam | setup GX and all its accessories Setup user administration Trained end user anout troubleshooting and maintenance | Unit is operational All modules working |  | GX machine | DR 11625 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951070377618706919/5951070377618706919_signature_28.png | Jomari Bryan Litonjua | https://www.jotform.com/uploads/macarewecare/222927978372067/5951070377618706919/5951070377618706919_signature_31.png
| 2024-06-25 | 2024-06-25 13:56 | Quezon City General Hospital | Seminary Rd. Project 8, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28641 | MACARE-HBDC | treatmenthub.qcgh@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 63 ASSAY: Xpert CT/NG, COUNT: 1 | 2024-06-13 | GX Module 10C D | 110012766 | Unit due for calibration | Performed Preventive maintenance Performed Xpertcheck | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck | unit is operational All modules working |  | DR 11601 | DR 11601 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951046359213590614/5951046359213590614_signature_28.png | Elena G. Ventura | https://www.jotform.com/uploads/macarewecare/222927978372067/5951046359213590614/5951046359213590614_signature_31.png
| 2024-06-25 | 2024-06-25 13:38 | Davao Regional Medical Center - RHIVDA Section | Apokon, Tagum City | 11 | Davao del Norte | RHIVDA | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28640 | SHELL FOUNDATION | nimezsheilamarie@gmail.com | 23.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 59 ASSAY: Xpert SARS-CoV-2, COUNT: 29 ASSAY: Xpert HIV-1 Viral Load, COUNT: 797 |  | GX-IV 4D 10Colors | 120000127 | Module B2 for Service Required Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Replaced the Defective Module SN: 220011054 with new Module SN: 210069469 | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5951037724617523426/5951037724617523426_signature_28.png | Sheila Marie Nimez RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5951037724617523426/5951037724617523426_signature_31.png
| 2024-06-25 | 2024-06-25 11:53 | Gilbert Teodoro Memorial Hospital | Camiling Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28639 | PBSP | almabellapdancel3@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 770 |  | GXIV-D | 812838 | C2:605738 for module replacement | Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit us operational C2:702757 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950972555125859925/5950972555125859925_signature_28.png | Alma Bella P. Dancel | https://www.jotform.com/uploads/macarewecare/222927978372067/5950972555125859925/5950972555125859925_signature_31.png
| 2024-06-25 | 2024-06-25 11:09 | Estevez Memorial Hospital, Inc | Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City, 4500 Albay | 5 | Albay | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28638 | COMMERCIAL | reylovendino83@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 62 ASSAY: Xpert MTB/RIF Ultra, COUNT: 53 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 7 |  | GX4-10C | 110017942 | A2/220009328 temperature Error 2008 error for A1, A4 GX set shuts down after power outage | Checked system log: A2/220009328 error/code - 1001, 2014 Checked/tested the dependability of UPS when unplugged from AC (UPS charged for 3hrs) Checked power output/connections going to UPS Cleaned cartridge bays, plunger rod, Icore Generate reports | Performed Troubleshooting Performed Preventive Maintenance | Unit is Operational UPS lasts for a few seconds after disconnected from AC power source | A2/220009328 sufficient error collected for replacement, UPS request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950952495117263013/5950952495117263013_signature_28.png | Lovendino, Rey | https://www.jotform.com/uploads/macarewecare/222927978372067/5950952495117263013/5950952495117263013_signature_31.png
| 2024-06-25 | 2024-06-20 14:09 | Philippine Red Cross - Sta. Cruz | Pedro Guevara Ave, Sta. Cruz, Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28637 | COMMERCIAL | xernan.dawinan@redcross.org.ph | 24 |  |  | MCS+ | 02J046 | For relocation | Installed Haemonetics. | Installed Haemonetics system. Performed maintenance Simulation run Verified calibration parameters Functionality test | Unit is operational. | Unit from PRC Calamba |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950911965013351515/5950911965013351515_signature_28.png | Xernan Dawinan | https://www.jotform.com/uploads/macarewecare/222927978372067/5950911965013351515/5950911965013351515_signature_31.png
| 2024-06-25 | 2024-06-25 08:23 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28636 | DOH | riazarizzamae@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 896 |  | GXIV-6C | 831637 | Due for Maintenance Filter worn out | Replaced Filter | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950845875618858514/5950845875618858514_signature_28.png | Jazer M. Palasigue | https://www.jotform.com/uploads/macarewecare/222927978372067/5950845875618858514/5950845875618858514_signature_31.png
| 2024-06-25 | 2024-06-25 08:23 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28635 | DOH | riazarizzamae@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 896 |  | GXIV-6C | 831637 | Due for Maintenance Filter worn out | Replaced Filter | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950845715619650199/5950845715619650199_signature_28.png | Jazer M. Palasigue | https://www.jotform.com/uploads/macarewecare/222927978372067/5950845715619650199/5950845715619650199_signature_31.png
| 2024-06-24 | 2024-06-24 21:06 | AYUNGON RURAL HEALTH UNIT | AYUNGON, NEGROS ORIENTAL | 7 | Negros Oriental | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28634 | PBSP | kesheiadgantalao@gmail.com | 28 |  |  | GX-IV | 805275 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950440951514828140/5950440951514828140_signature_28.png | Kesheia Gantalao, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5950440951514828140/5950440951514828140_signature_31.png
| 2024-06-24 | 2024-06-24 14:58 | Southern Philippines Medical Center | JP Laurel Avenue, Bajada Davao City | 11 | Davao del Sur | TB-HIV | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28633 | PBSP | spmcrhivda@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2761 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1805 |  | GX-IV 4D | 808298 | Module B3 encounter multiple temperature related errors For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Replaced the Defective Module SN: 644429 with new Module SN:604900 | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950217558326180742/5950217558326180742_signature_28.png | Flordeliza Go | https://www.jotform.com/uploads/macarewecare/222927978372067/5950217558326180742/5950217558326180742_signature_31.png
| 2024-06-24 | 2024-06-24 11:40 | Toledo City Health Office | Toledo City | 7 | Cebu | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 28632 | PBSP | loriemaediapen@yahoo.com | 23 |  | 2024-06-24 | Gx4 | 808640 | Module B3 Hardware fail. | Replace module defective. | Replace Module B3 with new module serial number 629652. | Module B1, 2,3 ok to use module B4 need to replace due to error's. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950099708327647704/5950099708327647704_signature_28.png | Lorie Mae C. Diapen | https://www.jotform.com/uploads/macarewecare/222927978372067/5950099708327647704/5950099708327647704_signature_31.png
| 2024-06-24 | 2024-06-22 09:29 | Calamba Medical Center | CMC Ave., Calamba, Laguna | 4A | Laguna | Laboratory | NO | Relocation | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28630 | COMMERCIAL | cmc.rtpcr@gmail.com | 20 | ASSAY: Xpert HBV VL, COUNT: 50 ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 665 |  | GXII-D-6C | 812916 | GeneXpert System for Relocation | Relocation. | Performed Self-test. Checked Cables. Checked Outlet Voltage at Micro. Generate Reports. Relocate GX System. | GX System Relocated fro Molecular to Micro. All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950065523615944144/5950065523615944144_signature_28.png | Andrea Gail B. Orgesvik | https://www.jotform.com/uploads/macarewecare/222927978372067/5950065523615944144/5950065523615944144_signature_31.png
| 2024-06-24 | 2024-06-21 12:28 | Alfonso Rural Health Unit | Burgos St., Alfonso, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28631 | DOH | abrigojan90@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 873 |  | GXIV-D-6C | 821279 | Unit is due for Maintenance | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Module are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950065523617773856/5950065523617773856_signature_28.png | Jan Karmel Abrigo | https://www.jotform.com/uploads/macarewecare/222927978372067/5950065523617773856/5950065523617773856_signature_31.png
| 2024-06-24 | 2024-06-20 14:48 | SACCL | Quiricada St., Sta.Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28629 | MACARE-HBDC | estrellesb@yahoo.com | 28 | ASSAY: Xpert HCV VL, COUNT: 64 |  | GXIV-D-6C | 839147 | Unit is due for Maintenance | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950065523618058886/5950065523618058886_signature_28.png | Blessie Estrelles | https://www.jotform.com/uploads/macarewecare/222927978372067/5950065523618058886/5950065523618058886_signature_31.png
| 2024-06-24 | 2024-06-14 12:10 | De La Salle Health Sciences Institute | Congressional Road, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angiepabruada@gmail.com |  | 28628 | MACARE-HBDC | acpabruada@dlsmhsi.edu.ph | 28 |  |  | GXIV-L-10C | 110023457 | Unit is due for Installation. | Installation. | Checked Output Voltage. Install Gx System. Generate Reports. | System is Functional. No UPS and Printer. | Recommendation to have UPS (minimum of 1kVA). |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950065523618948320/5950065523618948320_signature_28.png | Angelita C. Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5950065523618948320/5950065523618948320_signature_31.png
| 2024-06-24 | 2024-06-24 09:35 | Barili Rural Health Unit | Poblacion, Barili Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28627 | PBSP | barilimainrhu@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 842 |  | 6C GXIV | 834693 | Module B3 SN # 786755failed module during calibration | Module Replacement of B3 SN # 786755failed module during calibration Pulled out defective module, clean Icore, clean plunger and cartridge bay | Module Replacement of B3 SN # 786755failed module during calibration Pulled out defective module, clean Icore, clean plunger and cartridge bay Generate IQR, System, Gxx files, Assay Statistics | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5950025751114852788/5950025751114852788_signature_28.png | Ma. Grace Cui, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5950025751114852788/5950025751114852788_signature_31.png
| 2024-06-24 | 2024-06-21 14:33 | RAKKK PROPHET MEDICAL CENTER 2 | Gumaca, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28626 | PBSP | rpmcimoleclab20@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 166 |  | GX IV 6C-D | 808268 | Unit due for maintenance | Perfromed Preventive maintenance | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Generated report | unit is operational all modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5949992593611711823/5949992593611711823_signature_28.png | Katrina Rose F. Aureada | https://www.jotform.com/uploads/macarewecare/222927978372067/5949992593611711823/5949992593611711823_signature_31.png
| 2024-06-24 | 2024-06-21 11:25 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Tbdots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 28625 | PBSP | roseannparnes@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1938 ASSAY: Xpert SARS-CoV-2, COUNT: 871 |  | GX4 D | 803087 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5949963453613964498/5949963453613964498_signature_28.png | Rose Annn Parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5949963453613964498/5949963453613964498_signature_31.png
| 2024-06-24 | 2024-06-21 11:18 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 28624 | MACARE-HBDC | roseannparnes@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1938 ASSAY: Xpert SARS-CoV-2, COUNT: 871 |  | GX4 D (10C) | 110014085 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5949963063619231822/5949963063619231822_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/5949963063619231822/5949963063619231822_signature_31.png
| 2024-06-24 | 2024-06-21 10:06 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 28623 | MACARE-HBDC | eamcmolecularlab@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HPV, COUNT: 244 ASSAY: Xpert SARS-CoV-2, COUNT: 1243 |  | GX4 D (10C) | 110016781 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5949962833619598327/5949962833619598327_signature_28.png | East Avenue Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5949962833619598327/5949962833619598327_signature_31.png
| 2024-06-24 | 2024-06-21 09:17 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 28622 | MACARE-HBDC | bloodbankeamc@gmail.com | 28 |  |  | MCS + | 15H044G | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5949962373612675821/5949962373612675821_signature_28.png | Jocelyn Rote | https://www.jotform.com/uploads/macarewecare/222927978372067/5949962373612675821/5949962373612675821_signature_31.png
| 2024-06-24 | 2024-06-20 15:04 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 28621 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2752 ASSAY: Xpert SARS-CoV-2, COUNT: 3050 ASSAY: Others, COUNT: 340 |  | Gx 16D (10C) | 110014171 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5949962113618822253/5949962113618822253_signature_28.png | Jessica Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5949962113618822253/5949962113618822253_signature_31.png
| 2024-06-24 | 2024-06-20 14:54 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 28620 | PBSP | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 18 ASSAY: Xpert SARS-CoV-2, COUNT: 24 |  | GX4 D | 803109 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5949961793614982063/5949961793614982063_signature_28.png | Jessica Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5949961793614982063/5949961793614982063_signature_31.png
| 2024-06-23 | 2024-06-13 14:47 | Tarlac City Health Center 3 | Brgy. Sa Miguel, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28619 | DOH | daganlea572@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 556 |  | GXIV-D | 831697 | A1:633848 error 1002 A3:757252 and A4:608004 for additional Cepheid troubleshooting | Preventive maintenance Cepheid troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed Cepheid troubleshooting Performed functionality test Gathered reports | Unit is operational except A1 for reporting A3 and A4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5949556990112301019/5949556990112301019_signature_28.png | Lea Dagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5949556990112301019/5949556990112301019_signature_31.png
| 2024-06-23 | 2024-06-13 13:23 | Ramos Rural Health Unit | Ramos Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28618 | DOH | jadeeh8143@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 346 |  | GXIV-D | 831721 | Due for preventive maintenance A2:757555 error 4012 | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational A2:757555 for reporting error 4012 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5949556590116452026/5949556590116452026_signature_28.png | Emily Jade Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/5949556590116452026/5949556590116452026_signature_31.png
| 2024-06-22 | 2024-06-21 11:34 | Bolinao Rural Health Unit | Germinal, Bolinao, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28617 | DOH | ruralhealthunitlaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 573 | 2024-04-30 | GXIV-4-D-6C | 838278 | #NAME? | #NAME? | #NAME? | #NAME? |  | XpertCheck Kits | DR 11449 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5948768489029579346/5948768489029579346_signature_28.png | Alma Sotelo Sicat | https://www.jotform.com/uploads/macarewecare/222927978372067/5948768489029579346/5948768489029579346_signature_31.png
| 2024-06-22 | 2024-06-20 10:44 | Urdaneta Sacred Heart Hospital | Urdaneta, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28616 | COMMERCIAL | annabellelcatbagan@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 |  | GXII-2D-6C | 820303 | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5948767529023124273/5948767529023124273_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5948767529023124273/5948767529023124273_signature_31.png
| 2024-06-22 | 2024-06-20 09:21 | San Manuel Rural Health Unit | San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28615 | DOH | serly_reyes@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 842 | 2024-05-02 | GXIV-4-D-6C | 838516 | For Calibration and Preventive Maintenance | #NAME? | Performed Calibration and Preventive Maintenance | - Unit is Partially Operational - Module B4 (862532) Failed XpertCheck - Module B1, B2 and B3 Passed XpertCheck |  | XpertCheck Kits | DR 11454 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5948766619022262172/5948766619022262172_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/5948766619022262172/5948766619022262172_signature_31.png
| 2024-06-22 | 2024-06-22 15:41 | Kasibu Rural Health Unit | Municipal Compound, Kasibu, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28614 | PBSP | apollojoselkimayong@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 310 |  | GXII-6C | 834802 | Error 2126 | Put Module B1, B2 to Module B3, B4 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5948515404327271094/5948515404327271094_signature_28.png | Apollo Josel D. Kimayong | https://www.jotform.com/uploads/macarewecare/222927978372067/5948515404327271094/5948515404327271094_signature_31.png
| 2024-06-22 | 2024-06-21 14:30 | RAKKK PROPHET MEDICAL CENTER 3 | Gumaca, Quezon | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28613 | PBSP | rpmcimoleclab20@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 69 ASSAY: Xpert SARS-CoV-2, COUNT: 7 | 2024-04-29 | GXIV 6C D | 834625 | Unit due for calibration | Performed preventive maintenance Performed Xpertcheck Generated report | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck B1: Passed B2: Passed B3: Sevice Required B4: Passed | Unit is operational B3 for reporting |  | Xpertcheck kit | DR 47730 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5948426773748105924/5948426773748105924_signature_28.png | Katrina Rose F. Aureada | https://www.jotform.com/uploads/macarewecare/222927978372067/5948426773748105924/5948426773748105924_signature_31.png
| 2024-06-21 | 2024-06-19 17:03 | Cebu TB Reference Lab | South Road Property-Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28612 | PBSP | asmantiquilla.pbsp@gmail.com | 22 |  |  | Gx4 | 120000180 | Module Replacement B2 220011618 & B4 220011638 | New module Installed B2 875317 & B4 694507 Clean Fan Filter Gather data IQR | New module Installed B2 875317 & B4 694507 Clean Fan Filter Gather data IQR | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5947679899118625978/5947679899118625978_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5947679899118625978/5947679899118625978_signature_31.png
| 2024-06-21 | 2024-06-21 14:36 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28611 | PBSP | laureanjo@yahoo.com | 24 |  |  | Gx4 | 804840 | Genexpert Machine cannot be turned on | Install temporary spare power supply from cebu office Replace new fuse Gather IQR Request new AVR for another genexpert machine | Install temporary spare power supply from cebu office Replace new fuse Gather IQR | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5947679039115474330/5947679039115474330_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5947679039115474330/5947679039115474330_signature_31.png
| 2024-06-21 | 2024-06-21 16:09 | Lasam District Hospital | IBJ, Lasam, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | verbellmanuela36@gmail.com |  | 28610 | DOH | lasamdh.laboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 474 | 2024-06-21 | GXIV-6c | 838530 | Module B2 Hardware Failure Due for Calibration | Performed Calibration (B1,B3,B4 - Passed) (B2 - Service Required) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5947669168897267016/5947669168897267016_signature_28.png | Verbell M.Manuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5947669168897267016/5947669168897267016_signature_31.png
| 2024-06-21 | 2024-06-21 15:36 | Baras Medicare & Community Hospital | Bagong Sirang, Baras | 5 | Catanduanes | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28609 | PBSP | malustesorerormt@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 721 |  | GX4-6C | 835215 | B1/788023 - high temperature, burned surface at entrance of icore inside cartridge bay | Cleaned cartridge bays, plunger rod Generate reports | Performed Preventive Maintenance | Unit is Operational | B1/788023 - sufficient error collected to request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5947649786314171623/5947649786314171623_signature_28.png | Malu S. Tesorero,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5947649786314171623/5947649786314171623_signature_31.png
| 2024-06-21 | 2024-06-21 14:57 | Davao Chest Center | Villa Abrille Street, Poblacion District | 11 | Davao del Sur | TB-DOTS | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28608 | WHO | dcclabdept@gmail.com | 22.9 |  |  | GX-IV 4D | 804835 | Defective CPU System unit For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Replaced the Defective System Unit CPU SN: 2UA3520VTL with new System Unit CPU SN:CPH1223F48485 Run XpertCheck Calibration on four Modules | Perform System Unit replacement Perform Preventive Maintenance and XpertCheck Calibration | New System Unit CPU installed Module B1 and B2 passed the Calibration Module B3 and B4 skipped and retest required - no available XPC Cartidge left | Module B3 and B4 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5947631255419932108/5947631255419932108_signature_28.png | Maria Lisa Paras RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5947631255419932108/5947631255419932108_signature_31.png
| 2024-06-21 | 2024-06-21 13:56 | RAKKK Prophet Medical Center (1) | Brgy. Rosario, Maharlika Hiway, Gumaca, Quezon | 4A | Quezon | Molecular Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28607 | PBSP | rpmcimoleclab20@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 465 ASSAY: Xpert SARS-CoV-2, COUNT: 53 | 2024-04-29 | GXIV 6C D | 838871 | Monitor for replacement Unit due for Calibration | Replaced monitor Performed Xpertcheck | Replaced monitor HP -> Dell Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck B1: Passed B2: Service Required B3: Passed B4: Passed Generated report | Unit is operational Module B2 Service Required |  | Dell monitor Xpertcheck Kit | DR 48688, DR 47730 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5947605978129905959/5947605978129905959_signature_28.png | Katrina Rose F. Aureada | https://www.jotform.com/uploads/macarewecare/222927978372067/5947605978129905959/5947605978129905959_signature_31.png
| 2024-06-21 | 2024-06-21 13:17 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28606 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 660 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 17 |  | GXIV | 110008530 | Unit for maintenance | NA | #NAME? | Unit is operational No issues found |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5947569480715337579/5947569480715337579_signature_28.png | Maria Elizabeth De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5947569480715337579/5947569480715337579_signature_31.png
| 2024-06-21 | 2024-06-21 12:13 | Santa Ana Rural Health Unit | Centro, Santa Ana, Cagayan | 2 | Cagayan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28605 | PBSP | shena_22780@yahoo.com | 29 |  |  | GXII-6C | 834839 | N/a | N/a | Installed Genexpert Installed PC Installed Printer Installed Ups Performed Maintenance Gathered Reports | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5947529382615006101/5947529382615006101_signature_28.png | Shena Marie Semana | https://www.jotform.com/uploads/macarewecare/222927978372067/5947529382615006101/5947529382615006101_signature_31.png
| 2024-06-21 | 2024-06-21 10:11 | Alfonso Ponce Enrile Memorial District Hospital | Gonzaga, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28604 | DOH | apemdh_laboratory@yahoo.com | 2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1532 | 2024-04-23 | GXIV-6C | 838474 | Barcode Scanner not working Due for Calibration | Replaced Barcode Scanner Performed Calibration (2 Modules Passed) B3 - Service Required B4(Newly Replaced) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Barcode Scanner | 11434 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5947462602616577735/5947462602616577735_signature_28.png | Regina Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/5947462602616577735/5947462602616577735_signature_31.png
| 2024-06-20 | 2024-06-20 16:16 | Cagayan Valley Regional Blood Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | baclig.katherine@yahoo.com |  | 28603 | DOH | nvbspchdii2022@gmail.com | 28 |  |  | MCS+ | 19G146SPG | Due for Calibration | Performed Calibration Adjust the Line Sensor Adjust the Weigher Adjust the DPM Adjust the SPM Adjust Bowl Optics | Performed Maintenance Check the Line Sensor Check the Weigher Check the DPM Check the SPM Check the Bowl Optics Performed Functionality Test(2x) | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946814863027732469/5946814863027732469_signature_28.png | Katherine I. Baclig | https://www.jotform.com/uploads/macarewecare/222927978372067/5946814863027732469/5946814863027732469_signature_31.png
| 2024-06-20 | 2024-06-20 15:54 | St Lukes Extension Clinic | J. Bocobo St, Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 28602 | MACARE-HBDC | laboratory@slec.ph | 22 | ASSAY: Xpert CT/NG, COUNT: 4884 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1713 | 2024-06-14 | GX Infinity 48s | 809643 | Kiosk PC flagged by customer's IT system for outdated OS | Restored database from old Kiosk PC Performed functionality test Generated reports | Replaced Kiosk PC | Unit is operational |  | GeneXpert Desktop | DR# 11610 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946799060617981462/5946799060617981462_signature_28.png | MA. SARAH R. SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5946799060617981462/5946799060617981462_signature_31.png
| 2024-06-20 | 2024-06-20 15:23 | Marulas 3S Health Center | Marulas, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28601 | DOH | melissamartin0828@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1555 |  | GXIV | 831696 | Monitor for troubleshooting | Replaced monitor by a back up unit temporarily | - checked monitor's functionality (defective controls) - generated reports thru a back up monitor - performed self test | Unit is operational Monitor for replacement / repair SN: CNK83224PP |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946777737687744704/5946777737687744704_signature_28.png | Melissa Martin Galit | https://www.jotform.com/uploads/macarewecare/222927978372067/5946777737687744704/5946777737687744704_signature_31.png
| 2024-06-20 | 2024-06-20 15:23 | Bayawan CHO | Bayawan City | 7 | Negros Oriental | TB DOTS | YES | INSTALLATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com | 06emerson23@gmail.com |  | 28600 | PBSP | lobella_belliet@yahoo.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2070 |  | GX4 | 808282 | N/A | Replaced module B1 with new module SN 220001406 | Module Replacement for B1 w/ SN 607064 | All modules are now functional and in good condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946777514819776170/5946777514819776170_signature_28.png | Lobella Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5946777514819776170/5946777514819776170_signature_31.png
| 2024-06-20 | 2024-06-20 13:57 | Catmon Health Center | Catmon, Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28599 | PBSP | mariathes0218@gmail.com | 26 |  |  | GXIV-10C | 110020392 | Unit for installation | NA | #NAME? | Unit is operational and ready to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946744913629589930/5946744913629589930_signature_28.png | Clara Gloria | https://www.jotform.com/uploads/macarewecare/222927978372067/5946744913629589930/5946744913629589930_signature_31.png
| 2024-06-20 | 2024-06-20 13:24 | San Francisco Rural Health Unit | Brgy. Poblacion, San Francisco, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28598 | DOH | micoveLARDE13@GMAIL.COM | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 376 | 2024-04-25 | GX IV 6C - D | 831698 | Module A2 for replacement | Performed preventive maintenance Replaced module A2 | Checked and cleaned cartridge bays, plunger rod, icore slot, and filter Replaced module A2 757206 -> 210042259 Generated report | Unit is operational All modules working |  | GX module 6C | DR 47631 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946732371424657966/5946732371424657966_signature_28.png | Michael Francis Velarde | https://www.jotform.com/uploads/macarewecare/222927978372067/5946732371424657966/5946732371424657966_signature_31.png
| 2024-06-20 | 2024-06-20 13:42 | Salcedo Rural Health Unit | Salcedo Eastern Samar | 8 | Eastern Samar | TB-DOTS | YES | INSTALLATION REPAIR | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 28597 | PBSP | angelinegarcia_23@yahoo.com | 28 |  | 2024-06-20 | GXIV - 2D | 835061 | - Module B1/787858 and B2/787899 defective - Unit for installation | #NAME? | -Install new module in B1/760731 and B2/783221 - Install other devices CPU - SN/CPH1119F04303 Monitor - SN/CNK92102RD Printer - SN/VNC7703319 Keyboard/Mouse and barcode scanner - No UPS included | Unit is operational |  | Genexpert machine CPU Monitor Printer | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946718558713458337/5946718558713458337_signature_28.png | Angeline G. Bandoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5946718558713458337/5946718558713458337_signature_31.png
| 2024-06-20 | 2024-06-19 15:17 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28595 | DOH | rsjalova1974@yahoo.com | 25 |  |  | GXIV | 831956 | Unit for maintenance and troubleshooting | #NAME? | #NAME? | Unit is operational C1: passed C2: passed C3: insufficient cartridge C4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946711234526320609/5946711234526320609_signature_28.png | Rosevi Jalova | https://www.jotform.com/uploads/macarewecare/222927978372067/5946711234526320609/5946711234526320609_signature_31.png
| 2024-06-20 | 2024-06-19 15:20 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28596 | PBSP | rsjalova1974@yahoo.com | 25 |  |  | GXIV | 805062 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946711234528955740/5946711234528955740_signature_28.png | Rosevi Jalova | https://www.jotform.com/uploads/macarewecare/222927978372067/5946711234528955740/5946711234528955740_signature_31.png
| 2024-06-20 | 2024-06-20 12:57 | Solana Municipal Health Office | Solana, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28594 | PBSP | daisyjane2094@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 188 |  | GXII-6C | 835090 | Module B2 error Module B1 Hardware Failure | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Unit is Operational |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946689970322157831/5946689970322157831_signature_28.png | Daisy Jane P. Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5946689970322157831/5946689970322157831_signature_31.png
| 2024-06-20 | 2024-06-20 11:27 | Piat Municipal Health Office | Poblacion 2, Piat, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | sentrongsigla1999@yahoo.com.ph |  | 28593 | DOH | marieatalipsalas@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 87 | 2024-03-19 | GXIV-6C | 838532 | Due for Calibration | Performed Calibration (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | 11314 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946645474216389034/5946645474216389034_signature_28.png | Marie Atalip Salas | https://www.jotform.com/uploads/macarewecare/222927978372067/5946645474216389034/5946645474216389034_signature_31.png
| 2024-06-20 | 2024-06-20 10:50 | Davao Occidental General Hospital | Brgy. Lacaron, Malita | 11 | Davao Occidental | TB-DOTS | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28592 | PBSP | carelpapelera@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 86 |  | GX-IV 4D 10C | 110020262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install HIV VL Assay Install XpertCheck 1.5b | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946623819644860979/5946623819644860979_signature_28.png | Nazel John F. Velasco, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5946623819644860979/5946623819644860979_signature_31.png
| 2024-06-20 | 2024-06-20 10:07 | Alaminos Rural Health Unit | Alaminos Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28591 | PBSP | irahrobles@gmail.com | 25 |  | 2024-06-20 | GXIV | 838447 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946604656346289408/5946604656346289408_signature_28.png | Iralyn M. Robles | https://www.jotform.com/uploads/macarewecare/222927978372067/5946604656346289408/5946604656346289408_signature_31.png
| 2024-06-20 | 2024-06-20 10:22 | Tuao Rural Health Unit | Centro 2, Tuao, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28590 | DOH | annrollaine@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 229 | 2024-03-19 | GXIV-6C | 838465 | Due for Calibration | Performed Calibration (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | 11313 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946601580118896464/5946601580118896464_signature_28.png | Rollaine Ann Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5946601580118896464/5946601580118896464_signature_31.png
| 2024-06-20 | 2024-06-20 08:18 | Malita Rural Health Unit | Poblacion, Malita | 11 | Davao Occidental | TB-DOTS | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | mariastellarmt@gmail.com |  | 28589 | PBSP | rhumalitalaboratory@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1843 |  | GX-IV 4D | 805258 | Module 3 encounter multiple temperature related errors For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log Rearrange the GeneXpert position for suitable cooling position | Perform Preventive Maintenance and Module Troubleshooting | Module 3 reported to Cepheid for further Assessment Filter was damaged due to aging; will request new filters Only 3 Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946559509642413911/5946559509642413911_signature_28.png | Maria Stella E. Solitana | https://www.jotform.com/uploads/macarewecare/222927978372067/5946559509642413911/5946559509642413911_signature_31.png
| 2024-06-20 | 2024-06-19 14:31 | St lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 28588 | COMMERCIAL | megapo@stlukes.com.ph | 24 | ASSAY: Xpert CT/NG, COUNT: 286 ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 57 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3217 |  | GX4 D | 839234 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B1,B2 operational B3,B4 disabled due to failed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946423161199800489/5946423161199800489_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5946423161199800489/5946423161199800489_signature_31.png
| 2024-06-20 | 2024-06-19 13:05 | St lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 28587 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 59 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert HPV, COUNT: 217 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 | 2024-06-19 | GX4 L | 807155 | Due for xpertcheck, C1 encounter temp error 2014 | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | C3 and C4 retry xpertcheck lack of cartridge. C1 failed xpertcheck C2 passed |  | Xpertcheck-CE-5 | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946422531194045271/5946422531194045271_signature_28.png | Bonnie Lana, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5946422531194045271/5946422531194045271_signature_31.png
| 2024-06-20 | 2024-06-19 10:59 | Quezon Medical Center | Provincial Health Office, Brgy XI, Lucena City | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28586 | PBSP | qmcrtdl2015@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 153 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1981 ASSAY: Xpert HIV-1 Viral Load, COUNT: 56 | 2024-04-29 | GXIV 6C D | 807462 | Unit due for calibration | Unpacked Machine and accessories Performed preventive maintenance Performed Xpertcheck | Removed machine and accessories from box Setup machine and accessories Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck on D3 and D4 D1: New D2: New D3: service Required D4: Service Required | Unit is operational D3: Service Required D4: Service Required |  | Xpertcheck kit | DR 47729 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946251065317030748/5946251065317030748_signature_28.png | Jo-Ann P. Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5946251065317030748/5946251065317030748_signature_31.png
| 2024-06-19 | 2024-06-19 15:11 | Lower Bicutan Health Center | M.L Quezon St., Bicutan, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28585 | PBSP | ellainemagtoto35@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3172 |  | GX-IV-D-6C | 838858 | Modules B1, B2 and B3 has errors 1001/1002/4009. | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules B1/863216, B2/863322, B3/863340 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5946071018589695330/5946071018589695330_signature_28.png | Ellaine Magtoto | https://www.jotform.com/uploads/macarewecare/222927978372067/5946071018589695330/5946071018589695330_signature_31.png
| 2024-06-19 | 2024-06-19 15:47 | Mallig Super Health Center | Mallig, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28584 | PBSP | shcmalliglab1@gmial.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 933 |  | GXIV-6C | 834686 | Module B3 Error | Performed Module Replacement B3(210092911 - 692259) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945928283412620707/5945928283412620707_signature_28.png | Hanna Jane C. Respicio | https://www.jotform.com/uploads/macarewecare/222927978372067/5945928283412620707/5945928283412620707_signature_31.png
| 2024-06-19 | 2024-06-19 15:44 | LAMBUNAO RURAL HEALTH UNIT | LAMBUNAO, ILOILO | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28583 | PBSP | gabatanelson@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2447 |  | GX-IV | 815279 | #NAME? | #NAME? | -PERFORM CLEANING OF GENEXPERT WORKSTATION -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B1 REPLACEMENT: DEF SN:793566 WITH NEW SN:210060891 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945924799743130482/5945924799743130482_signature_28.png | Nelson Gabata, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5945924799743130482/5945924799743130482_signature_31.png
| 2024-06-19 | 2024-06-19 13:44 | Tumauini District Hospital | Lingaling, Tumauini, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28582 | MACARE-HBDC | kayecatembung2911@gmail.com | 28 |  | 2024-06-13 | GXIV-10C | 110013709 | Due for Calibration | Performed Calibration (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  | 11602 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945863213418131941/5945863213418131941_signature_28.png | Mylene Claravall | https://www.jotform.com/uploads/macarewecare/222927978372067/5945863213418131941/5945863213418131941_signature_31.png
| 2024-06-19 | 2024-06-19 13:31 | Cebu TB reference lab | SRP Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28581 | PBSP | asmantiquilla.pbsp@gmail.com | 20 |  |  | Gx 4 | 808620 | Defective Module D1, D3 & D4 Defective monitor | New Installation Module D1,D3 & D4 New monitor Installed Clean Fan Filter | New Installation Module D1,D3 & D4 New monitor Installed Clean Fan Filter | Systems operational UPS defective |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945848279718259005/5945848279718259005_signature_28.png | Allan Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5945848279718259005/5945848279718259005_signature_31.png
| 2024-06-19 | 2024-06-19 13:02 | Delfin Albano Memorial Hospital | Ragan Sur, Delfin Albano, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28580 | DOH | maryrosebacanirmt@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 946 |  | GXIV-6C | 838208 | Due for calibration | Performed Xpertcheck 2 Modules Passed(2 modules newly replaced) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945829040252987304/5945829040252987304_signature_28.png | Mary Rose Bacani | https://www.jotform.com/uploads/macarewecare/222927978372067/5945829040252987304/5945829040252987304_signature_31.png
| 2024-06-19 | 2024-06-19 12:51 | Abuyog District Hospital | Abuyog Leyte | 8 | Leyte | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28579 | PBSP | erwindelacruz089@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 992 | 2024-06-19 | GXIV | 815274 | n/a | #NAME? | #NAME? | Unit is operational | All module passed for calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945821083215479320/5945821083215479320_signature_28.png | Erwin B. De La Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5945821083215479320/5945821083215479320_signature_31.png
| 2024-06-19 | 2024-06-19 09:40 | Cainta City Health Office | Karangalan Dr. Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28578 | PBSP | ballesterrichard18@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 949 ASSAY: Xpert HIV-1 Viral Load, COUNT: 117 |  | GX-IV-D-6C | 838182 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B4/807137 waiting for replacement module. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945809956795985756/5945809956795985756_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5945809956795985756/5945809956795985756_signature_31.png
| 2024-06-19 | 2024-06-19 09:35 | Cainta City Health Office | Karangalan Dr., Cainta, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28577 | DOH | ballesterrichard18@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 763 ASSAY: Xpert HIV-1 Viral Load, COUNT: 74 | 2024-06-19 | GX-IV-D-6C | 831623 | Module A1 for replacement | Replaced module A1 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module A1 Generated reports | Unit is operational. All modules operational. |  | Old: A1/802425 New: A1/693718 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945809876794215134/5945809876794215134_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5945809876794215134/5945809876794215134_signature_31.png
| 2024-06-19 | 2024-06-19 09:38 | Cainta City Health Office | Karangalan Dr. Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28576 | PBSP | ballesterrichard18@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1169 |  | GX-IV-D-6C | 804687 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945809776793360002/5945809776793360002_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5945809776793360002/5945809776793360002_signature_31.png
| 2024-06-19 | 2024-06-19 11:27 | The Medical City | Ortigas Ave., Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | Relocation | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28575 | MACARE-HBDC | moleculardx@themedicalcity.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 802630 | Unit for relocation | NA | #NAME? | Unit is operational Temp. Errors found on module 4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945774393615632807/5945774393615632807_signature_28.png | Fritzie Mae T. Perante | https://www.jotform.com/uploads/macarewecare/222927978372067/5945774393615632807/5945774393615632807_signature_31.png
| 2024-06-19 | 2024-06-19 11:16 | Sto Tomas Rural Health Unit | Bolinao, Santo Tomas, Isabela | 2 | Isabela | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28574 | MACARE-HBDC | gloangelier@gmail.com | 28 |  |  | GXIV-10c | 110024248 | N/a | N/a | Installed Genexpert Installed PC Installed Drivers Installed UPS Installed Printer Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945770938414922541/5945770938414922541_signature_28.png | Glo Angelie Bagain | https://www.jotform.com/uploads/macarewecare/222927978372067/5945770938414922541/5945770938414922541_signature_31.png
| 2024-06-19 | 2024-06-19 09:12 | Mary Johnston Hospital | J. Nolasco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | mchjyclm@gmail.com |  | 28573 | CONSORTIUM | mjhlaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2140 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 |  | GXIV | 804702 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945690401421940769/5945690401421940769_signature_28.png | Micah Joy C. Talabong | https://www.jotform.com/uploads/macarewecare/222927978372067/5945690401421940769/5945690401421940769_signature_31.png
| 2024-06-18 | 2024-06-18 17:36 | Badian Rural Health Unit | Poblacion, Badian Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28572 | PBSP | avenzr@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1191 |  | GX4 | 807588 | B1 SN# 739696 failed during calibration. Need additional dust filter. | Pulled out B1 SN# 739696 and replace new module SN # 694164 . Generate IQR, GXX data, System log, Assay Statistics | Pulled out B1 SN# 739696 and replace new module SN # 694164 . Generate IQR, GXX data, System log, Assay Statistics | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945128941117006526/5945128941117006526_signature_28.png | Avenizer Absin, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5945128941117006526/5945128941117006526_signature_31.png
| 2024-06-18 | 2024-06-18 15:05 | Calabanga Rural Health Unit | Calabanga, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28571 | DOH | martyarc11@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 807 |  | GX4-6C | 838398 | GX for Calibration | XpertCheck Result: B1/807565 - passed B2/737294 - newly replaced module B3/731543 - newly replaced module B4/807261 - newly replaced module Cleaned cartridge bays, plunger rod, icore Generate report | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5945052313715158387/5945052313715158387_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/5945052313715158387/5945052313715158387_signature_31.png
| 2024-06-18 | 2024-06-18 12:37 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28570 | DOH | burauenlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 452 | 2024-06-18 | GXIV | 838551 | n/a | #NAME? | #NAME? | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944950029311367774/5944950029311367774_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5944950029311367774/5944950029311367774_signature_31.png
| 2024-06-18 | 2024-06-18 11:39 | Bicol Medical Center | Concepcion, Naga City | 5 | Camarines Sur | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28569 | MACARE-HBDC | mongeshiela@gmail.com | 25 |  |  | MCS+ | 15H027G | MCS+ request for Preventive Maintenance | Cleaned/disinfect Centrifuge, rotor pumps, valves, surface of MCS+ Checked power connections Checked parameters | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944948197612431364/5944948197612431364_signature_28.png | Shiela P. Monge R.N. | https://www.jotform.com/uploads/macarewecare/222927978372067/5944948197612431364/5944948197612431364_signature_31.png
| 2024-06-18 | 2024-06-18 11:33 | Bicol Medical Center | Concepcion, Naga City | 5 | Camarines Sur | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28568 | MACARE-HBDC | mongeshiela@gmail.com | 25 |  |  | MCS+ | 15H031G | MCS+ request for Preventive Maintenance | Cleaned/disinfect Centrifuge, rotor pumps, valves, surface of MCS+ Checked power connections Checked parameters | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944947457612798115/5944947457612798115_signature_28.png | Shiela P. Monge R.N. | https://www.jotform.com/uploads/macarewecare/222927978372067/5944947457612798115/5944947457612798115_signature_31.png
| 2024-06-18 | 2024-06-18 11:49 | Ilagan City Health Office | Ilagan City, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28567 | PBSP | cho1ilagan.laboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2519 |  | GXIV-6C | 808635 | Module B4 valve not aligned Module B2 error Module B3 Temperature error | Performed Module Replacement B2(678888 - 675119) Performed Troubleshooting | Peformed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module B4 for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944924293324333527/5944924293324333527_signature_28.png | April Jenneth Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/5944924293324333527/5944924293324333527_signature_31.png
| 2024-06-18 | 2024-06-18 10:41 | Saragani Rural Health Unit | Poblacion Mabila, Saragani | 11 | Davao Occidental | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28566 | DOH | lovitosquennie26@gmail.com | 30.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 343 |  | GX-IV 4d | 837614 | Stuck Cartidge at Module B3 Due for Preventive Maintenance | Remove Stuck Cartidge Reset network connections Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log Archived Data test results | Perform Preventive Maintenance and Module Troubleshooting | All Modules are operational | High temperature due to no ACU, reported to higherups for assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944880048613272051/5944880048613272051_signature_28.png | QUENNIE C. LOVITOS | https://www.jotform.com/uploads/macarewecare/222927978372067/5944880048613272051/5944880048613272051_signature_31.png
| 2024-06-18 | 2024-06-13 15:11 | Dr. Jose Rizal Memorial Hospital | Dapitan, Lawaan | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28565 | PBSP | lovinoantonieta@gmail.comc | 23 |  |  | Gx 4 | 804698 | Defective module A3 640864 | New module installed A3 210086464 Clean Fan Gather data pre IQR and Post IQR | New module installed A3 210086464 Clean Fan Gather data pre IQR and Post IQR | Sytems operational |  |  |  | Raymond Laurito | https://www.jotform.com/uploads/macarewecare/222927978372067/5944831291112413298/5944831291112413298_signature_28.png | Antonieta B. Lovino | https://www.jotform.com/uploads/macarewecare/222927978372067/5944831291112413298/5944831291112413298_signature_31.png
| 2024-06-18 | 2024-06-18 08:52 | ALIMODIAN RURAL HEALTH UNIT | PLAZA LIBERTAD STREET, ALIMODIAN, ILOILO | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28564 | PBSP | rosemarie.rmt2@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 939 |  | GX-II | 834841 | FOR MODULE B1 AND MODULE B2 REPLACEMENT | #NAME? | -PERFORM CLEANING OF GX WORKSTATION AND PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE REPLACEMENT: B1 DEF SN:787348, NEW SN:631282 B2 DEF SN:787350, NEW SN:675964 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944815394579196667/5944815394579196667_signature_28.png | Rosemarie Ani | https://www.jotform.com/uploads/macarewecare/222927978372067/5944815394579196667/5944815394579196667_signature_31.png
| 2024-06-18 | 2024-06-06 12:30 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28563 | MACARE-HBDC | ljf7403@yahoo.com | 25 |  | 2024-06-06 | MCS+ | 13C270 | Unit not powering on | Replaced fuse | #NAME? | Unit is operational |  | 2.5A, 250V Fuse (2 pcs) | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944808933612695527/5944808933612695527_signature_28.png | Leonila flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5944808933612695527/5944808933612695527_signature_31.png
| 2024-06-18 | 2024-06-06 13:05 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28562 | CONSORTIUM | guelaquino11@gmail.com | 23 | ASSAY: Xpert C. difficile BT, COUNT: 32 ASSAY: Xpert CT/NG, COUNT: 79 ASSAY: Xpert MTB/RIF Ultra, COUNT: 644 ASSAY: Others, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 28 |  | GXIV | 847578 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Modules B1, B4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944789993611795692/5944789993611795692_signature_28.png | Miguel Joseph C. Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/5944789993611795692/5944789993611795692_signature_31.png
| 2024-06-18 | 2024-06-11 13:21 | Malolos Rural Health Unit VI | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28561 | PBSP | sandiego.redarizza@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1969 |  | GXIV | 838862 | Module B2 for troubleshooting Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Module B2 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944788623611272972/5944788623611272972_signature_28.png | Reda San Diego | https://www.jotform.com/uploads/macarewecare/222927978372067/5944788623611272972/5944788623611272972_signature_31.png
| 2024-06-17 | 2024-06-17 14:07 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28560 | DOH | marteetingzon20@gmail.com | 23 |  | 2024-06-11 | GXIV | 838421 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5944138913583364989/5944138913583364989_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5944138913583364989/5944138913583364989_signature_31.png
| 2024-06-16 | 2024-06-16 19:34 | Chong Hua Hospital- Fuente | Capitol Site, Fuente, Cebu City | 7 | Cebu | Bloodbank | YES | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28559 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 11J219 | Error 413 Valve Problem "purple valve test failure" | Defective Pneumatic Assembly for replacement (out of warranty) Need a quotation for Pneumatic assembly | Trouble problem check valve, check pneumatic assembly | MCS+ is standby waiting for replacement. Pneumatic assembly problem. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5943472713501643526/5943472713501643526_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/5943472713501643526/5943472713501643526_signature_31.png
| 2024-06-15 | 2024-06-15 23:42 | Dr. Jorge P. Royeca Hospital | General Santos City, South Cotabato | 12 | South Cotabato | Tb-Lab | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28558 | PBSP | elsigh76@gmail.com | 30.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 107 |  | Gx-IV (10C) | 110020386 | N/A | Genexpert machine Cleaning/Preventive maintenance Checking of errors Data collection for new Installed Genexpert machine | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5942756659495175496/5942756659495175496_signature_28.png | Elsie R. Dacua | https://www.jotform.com/uploads/macarewecare/222927978372067/5942756659495175496/5942756659495175496_signature_31.png
| 2024-06-15 | 2024-06-11 13:52 | Dr. Jorge P. Royeca Hospital | General Santos City, South Cotabato | 12 | South Cotabato | Tb-Lab | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28557 | PBSP | elsigh76@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1102 |  | Gx-IV (6C) | 838130 | N/A | Genexpert machine Cleaning/Preventive maintenance Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5942753079496157897/5942753079496157897_signature_28.png | Elsie R. Dacua | https://www.jotform.com/uploads/macarewecare/222927978372067/5942753079496157897/5942753079496157897_signature_31.png
| 2024-06-15 | 2024-06-14 10:26 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | avabichuela3216@gmail.com |  | 28556 | PBSP | molecularlab@phc.gov.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1273 ASSAY: Others, COUNT: 17 ASSAY: Xpert SARS-CoV-2, COUNT: 2671 |  | GX4 D | 834443 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5942100963415005416/5942100963415005416_signature_28.png | Arjay V. Abichuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5942100963415005416/5942100963415005416_signature_31.png
| 2024-06-15 | 2024-06-14 10:07 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | avabichuela3216@gmail.com |  | 28555 | MACARE-HBDC | molecularlab@phc.gov.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 283 ASSAY: Xpert SARS-CoV-2, COUNT: 829 |  | GX16 D | 838655 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5942100643411703999/5942100643411703999_signature_28.png | Arjay V. Abichuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5942100643411703999/5942100643411703999_signature_31.png
| 2024-06-15 | 2024-06-14 09:05 | Best Diagnostic Corp | Masikap st.,Diliman QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28554 | COMMERCIAL | bestdx@yahoo.com | 30 | ASSAY: Xpert CT/NG, COUNT: 148 ASSAY: Xpert HBV VL, COUNT: 45 ASSAY: Xpert MTB/RIF Ultra, COUNT: 13 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert CT/NG, COUNT: 148 ASSAY: Xpert HBV VL, COUNT: 45 ASSAY: Xpert MTB/RIF Ultra, COUNT: 13 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX4 D | 110008529 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5942100253413999487/5942100253413999487_signature_28.png | Best Diagnostic Corp. | https://www.jotform.com/uploads/macarewecare/222927978372067/5942100253413999487/5942100253413999487_signature_31.png
| 2024-06-15 | 2024-06-13 14:45 | Biopath QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Lab | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28553 | COMMERCIAL | biopathqc@biopath.ph | 29 | ASSAY: Xpert CT/NG, COUNT: 34 ASSAY: Xpert HPV, COUNT: 5 ASSAY: Xpert SARS-CoV-2, COUNT: 30 |  | GX16 | 839306 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | 14 modules operational A3 disabled H/W failed C2 failed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5942099843419150414/5942099843419150414_signature_28.png | Biopath Clinical Diagnostics, Inc-E. Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5942099843419150414/5942099843419150414_signature_31.png
| 2024-06-15 | 2024-06-13 11:43 | Philippine Red Cross tower | Edsa,Mandaluyong | NCR | NCR 2nd - Mandaluyong | Bloodbank | NO | REPAIR CALIBRATION | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28552 | MACARE-HBDC | bloodcollection.tower@gmail.com | 26 |  |  | MCS + | 08B049 | CRC bios failed/ rollers slight worn out and yellowish color | Replaced CPU board Performed calibration Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Repair and calibration | Rollers need to replace Unit ready to use | Pull out defective CPU board |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5942099513411367235/5942099513411367235_signature_28.png | Ardel brecinio | https://www.jotform.com/uploads/macarewecare/222927978372067/5942099513411367235/5942099513411367235_signature_31.png
| 2024-06-14 | 2024-06-14 15:51 | Klinika - Eastwood | L3, Citywalk 1, Bagumbayan, Q.C. | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Harold Toco |  | htoco@macare-medicals.com | clairebaptista32786@gmail.com |  | 28551 | MACARE-HBDC | klinikaeastwoodlaboratory@gmail.com | 28 |  | 2024-06-14 | GXIV 10C | 110020698 | For installation. | Unit installation done. Setup printer. Check unit funstionality. | Performed unit installation. | Unit is operational. | Recommended for wall partition between the zinc and the unit. Needs small table for the printer. Also needs heavy duty extention cord. | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941614196229087695/5941614196229087695_signature_28.png | Mary Claire P Baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/5941614196229087695/5941614196229087695_signature_31.png
| 2024-06-14 | 2024-06-14 15:29 | Josefina Belmonte Duran Albay Provincial Hospital | Tuburan, Ligao City | 5 | Albay | Laboratory | NO | REPAIR | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28550 | PBSP | slrempillo@gmail.com | 26 | ASSAY: , COUNT: 804 |  | GX4-6C | 814371 | Monitor for replacement | Replaced Monitor Cleaned cartridge bays, plunger rod Generate reports | Replaced monitor Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941598908103229447/5941598908103229447_signature_28.png | Lirio llorin | https://www.jotform.com/uploads/macarewecare/222927978372067/5941598908103229447/5941598908103229447_signature_31.png
| 2024-06-14 | 2024-06-14 11:07 | Pilar Rural Health Unit | Calongay, Pilar, Sorsogon | 5 | Sorsogon | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28549 | DOH | manalangmark88@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 723 |  | GX4-6C | 838266 | Printer not printing UPS no power | Printer - Reinstalled Printer driver, checked ink cartridge, performed print test UPS - Checked power cord. Charged for 1hr. Tested the UPS status when unplugged. Checked output voltage. | Performed troubleshooting for UPS & Printer | Unit is Operational Printer is operational UPS - drained/defective battery | UPS - battery is defective, request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941593328106346676/5941593328106346676_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5941593328106346676/5941593328106346676_signature_31.png
| 2024-06-14 | 2024-06-14 14:56 | Navotas Medical and Wellness Center | M. Naval St. Navotas City | NCR | NCR 3rd - Navotas | Laboratort | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | alejandrinoknowell@gmail.com |  | 28548 | DOH | navotascitylab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2017 | 2024-06-14 | GXIV | 831626 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module B4 (210102146) unable to Xpertcheck because it is new and installed last 02/22/2024. Modules B1, B2 and B3 passed Xpertcheck. |  | Xpertcheck kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941587869614920015/5941587869614920015_signature_28.png | Jade Renette Cesista | https://www.jotform.com/uploads/macarewecare/222927978372067/5941587869614920015/5941587869614920015_signature_31.png
| 2024-06-14 | 2024-06-14 14:32 | Soccsksargen General Hospital | Surallah, South Cotabato | 12 | South Cotabato | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28547 | PBSP | johnellamarisbautista@gmail.com | 27.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1228 | 2024-06-14 | Gx-IV (6C) | 815432 | N/A | Perform Genexpert annual Xpertcheck/Calibration Genexpert Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Xpertcheck | 4 operational modules | All modules passed the Xpertcheck | Xpertcheck CE-5 cartridge | 1001075092 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941575202927679549/5941575202927679549_signature_28.png | Johnella Maris Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5941575202927679549/5941575202927679549_signature_31.png
| 2024-06-14 | 2024-06-14 14:44 | Soccsksargen General Hospital | Surallah, South Cotabato | 12 | South Cotabato | Tb-Lab | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28546 | PBSP | johnellamarisbautista@gmail.com | 26.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 940 ASSAY: Others, COUNT: 37 |  | Gx-IV (10C) | 120000172 | N/A | Genexpert machine Cleaning/Preventive maintenance Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941574252924437218/5941574252924437218_signature_28.png | Johnella Maris R. Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5941574252924437218/5941574252924437218_signature_31.png
| 2024-06-14 | 2024-06-13 12:02 | Tacurong City Health Office | Tacurong City, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28545 | PBSP | mjgppalabrica@gmail.com | 28.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1597 | 2024-06-13 | Gx-IV | 815393 | Module 4 for Service Required | Perform genexpert annul Xpertcheck/Calibration Genexpert Cleaning Checking of errors Problem/errors reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert Xpertcheck | 3 operational modules | Module 1,2, and 3 passed the Xpertcheck | Xpertcheck CE-5 cartridge | Lot#: 1001252536 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941560052924105377/5941560052924105377_signature_28.png | Mary June Grace Villacruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5941560052924105377/5941560052924105377_signature_31.png
| 2024-06-14 | 2024-06-14 14:05 | Arayat Rural Health Unit I | Arayat, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | kamanlapaz@pbsp.org.ph |  | 28544 | PBSP | baylonrochelle06@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1397 |  | GXIV-6C | 812790 | UPS not powering on Due for Calibration | Performed Calibration (All Modules Passed) Performed UPS Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941546151316778596/5941546151316778596_signature_28.png | Rochelle Baylon | https://www.jotform.com/uploads/macarewecare/222927978372067/5941546151316778596/5941546151316778596_signature_31.png
| 2024-06-14 | 2024-06-14 12:52 | Basilan General Hospital | Basilan, Isabela City | BARMM | Basilan | Laboratory | YES | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 28543 | PBSP | enellaventura@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 496 | 2024-06-14 | GX 4 | 707515 | Module 4 for replacement | Replaced Module 662321 with Module 877080 and Preventive Maintenence | Module Replacement and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941537194915420982/5941537194915420982_signature_28.png | Allene A. Ventura | https://www.jotform.com/uploads/macarewecare/222927978372067/5941537194915420982/5941537194915420982_signature_31.png
| 2024-06-14 | 2024-06-11 11:02 | Asian Hospital And Medical Center | Civic Dr., Filinvest, Alabang, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | rpescador@asianhospital.com |  | 28542 | COMMERCIAL | agalang@asisnhospital.com | 22 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Others, COUNT: 69 ASSAY: Xpert HBV VL, COUNT: 21 ASSAY: Xpert HPV, COUNT: 144 ASSAY: Xpert MTB/RIF Ultra, COUNT: 917 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 | 2024-06-10 | GXIV-D-10C | 110014996 | Unit is due for XpertCheck | XpertCheck. Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, Icore Optics and Filter. Installed XpertCheck Software. Performed XpertCheck. Performed Self-test. Generate Reports. | All Modules Passed. |  | XpertCheck Kit. | 11593 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941485886411493397/5941485886411493397_signature_28.png | Alyssa R. Galang | https://www.jotform.com/uploads/macarewecare/222927978372067/5941485886411493397/5941485886411493397_signature_31.png
| 2024-06-14 | 2024-06-11 13:56 | University of Perpetual Helps Dalta Medical Center | Alabang-Zapote Road, Las Piñas City | NCR | NCR 4th - Las Piñas | Blood Bank | NO | CALIBRATION | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | uphdmc_lab@yahoo.com |  | 28541 | COMMERCIAL | neressleanoratienza@yahoo.com.ph | 24 |  |  | MCS+ | 11J220 | Unit is due for Calibration. | Calibration | Performed Simulation Runs. Checked A/D Values. Checked and Modify Calibration Values. Checked Pump Rollers. Cleaned Surface Unit. | Unit is Operational and Calibration Values Passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941485886414670400/5941485886414670400_signature_28.png | Neress Atienza | https://www.jotform.com/uploads/macarewecare/222927978372067/5941485886414670400/5941485886414670400_signature_31.png
| 2024-06-14 | 2024-06-06 12:30 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28540 | MACARE-HBDC | ljf7403@yahoo.com | 25 |  | 2024-06-06 | MCS+ | 13C270 | Unit not powering on | Replaced fuse | #NAME? | Unit is operational |  | 2.5A, 250V Fuse (2 pcs) | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941480824682247543/5941480824682247543_signature_28.png | Leonila flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5941480824682247543/5941480824682247543_signature_31.png
| 2024-06-14 | 2024-06-05 14:27 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28539 | MACARE-HBDC | sah.genexpertlaboratory@gmail.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1118 ASSAY: Xpert SARS-CoV-2, COUNT: 1884 |  | GXIV | 847728 | Unit for maintenance | NA | #NAME? | Unit is operational Room temp exceeds max threshold (30C) Unit is not recommended for use until room temp. is ideally under 30C to prevent module damage |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941480524686265666/5941480524686265666_signature_28.png | Ren Joshua Tumbado | https://www.jotform.com/uploads/macarewecare/222927978372067/5941480524686265666/5941480524686265666_signature_31.png
| 2024-06-14 | 2024-06-05 14:36 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28538 | MACARE-HBDC | sahlabmedicine@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 189 |  | GXIV | 847728 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941480044687184561/5941480044687184561_signature_28.png | Bon Angelo Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5941480044687184561/5941480044687184561_signature_31.png
| 2024-06-14 | 2024-06-06 12:30 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28537 | MACARE-HBDC | ljf7403@yahoo.com | 25 |  | 2024-06-06 | MCS+ | 13C270 | Unit not powering on | Replaced fuse | #NAME? | Unit is operational |  | 2.5A, 250V Fuse (2 pcs) | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941479194684940371/5941479194684940371_signature_28.png | Leonila flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5941479194684940371/5941479194684940371_signature_31.png
| 2024-06-14 | 2024-06-06 13:05 | University of The East Ramon Magsaysay Medical Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28536 | CONSORTIUM | guelaquino11@gmail.com | 23 | ASSAY: Xpert C. difficile BT, COUNT: 32 ASSAY: Xpert CT/NG, COUNT: 79 ASSAY: Xpert MTB/RIF Ultra, COUNT: 644 ASSAY: Others, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 28 |  | GXIV | 847578 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Modules B1, B4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941478814688183657/5941478814688183657_signature_28.png | Miguel Joseph C. Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/5941478814688183657/5941478814688183657_signature_31.png
| 2024-06-14 | 2024-06-11 13:21 | Malolos Rural Health Unit VI | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28535 | PBSP | sandiego.redarizza@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1969 |  | GXIV | 838862 | Module B2 for troubleshooting Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Module B2 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941478564681217932/5941478564681217932_signature_28.png | Reda San Diego | https://www.jotform.com/uploads/macarewecare/222927978372067/5941478564681217932/5941478564681217932_signature_31.png
| 2024-06-14 | 2024-03-20 12:06 | Our Lady of Lourdes Hospital | Sta. Mesa, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | jcmasaganda@ollh.ph |  | 28534 | CONSORTIUM | ljmalora@ollh.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 111 ASSAY: Xpert MTB/RIF Ultra, COUNT: 387 | 2024-03-07 | GXIV | 840974 | Unit for calibration and maintenance | Performed XpertCheck | #NAME? | Unit is operational All modules passed calibration |  | XpertCheck Kit Lot: 1000410214 | 11272 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941475424684712974/5941475424684712974_signature_28.png | Loureine Joy Obrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5941475424684712974/5941475424684712974_signature_31.png
| 2024-06-14 | 2024-06-14 11:56 | Our Lady of Lourdes Hospital | Sta. Mesa, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | jcmasaganda@ollh.ph |  | 28533 | CONSORTIUM | leslie.cabanilla@medilinx.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 444 |  | GXIV | 840974 | Unit for maintenance | NA | #NAME? | Unit is operational Unit is in good working condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941475144681869640/5941475144681869640_signature_28.png | Leslie Jhon D. Cabanilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5941475144681869640/5941475144681869640_signature_31.png
| 2024-06-14 | 2024-06-14 09:50 | Coron Rural Health Unit | Brgy. 2, Coron, Palawan | 4B | Palawan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28532 | PBSP | lycking@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 594 | 2024-04-30 | GX IV 6C-D | 807612 | Unit due for calibration | Performed Preventive maintenance Performed Xpertcheck | checked and cleaned cartridge bays, plunger rod, icore slot, and filter Filter is worn out Performed Xpertcheck B1, B2, B4 - Passed B3 - Service required Generated report | Unit is operational B1: Passed B2: Passed B3: Service required B4: Passed Filter for reporting |  | XpertCheck kit | DR 48191 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941408287896847276/5941408287896847276_signature_28.png | Lyca Torrible-Abaroa | https://www.jotform.com/uploads/macarewecare/222927978372067/5941408287896847276/5941408287896847276_signature_31.png
| 2024-06-14 | 2024-06-13 10:50 | Culion Sanitarium General Hospital | Robert St., Brgy. Tiza, Culion Palawan | 4B | Palawan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28531 | PBSP | evecastillocalinawan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 392 ASSAY: Xpert SARS-CoV-2, COUNT: 1509 | 2023-10-24 | GX II 6C-D | 835101 | Unit due for XpertCheck | Setup GX and accessories Performed Preventive maintenance Performed Xpertcheck | Unit packed by cepheid Unpacked gx and accessories to perform XpertCheck Checked and cleaned cartridge bays, plunger, rod, icore slot, and filter Performed Xpertcheck Module A1, Insufficient cartridge Module A2, failed Generated report | Unit is operational A1: insufficient cartridge for XpertCheck A2: Failed XpertCheck |  | XpertCheck Kit | Lot 1000725070 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941380477894537378/5941380477894537378_signature_28.png | Rose Eve C. Calinawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5941380477894537378/5941380477894537378_signature_31.png
| 2024-06-14 | 2024-06-14 08:45 | Malay Primary Care Facility | Sitio Tabon, Caticlan, Malay, Aklan | 6 | Aklan | Clinical Laboratory | NO | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28530 | DOH | mhomalayclinicallaboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 983 |  | GXIV-4D | 838457 | #NAME? | #NAME? | -Preventive Maintenance -Reinstallation -Archive tests -Upgrade XpertCheck software to version 1.5b | All modules are functional. Module B2 has history of temperature error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5941360550121374011/5941360550121374011_signature_28.png | May Soledad | https://www.jotform.com/uploads/macarewecare/222927978372067/5941360550121374011/5941360550121374011_signature_31.png
| 2024-06-13 | 2024-06-13 12:34 | Occidental Mindoro Provincial Hospital | Mamburao, Occidental Mindoro | 4B | Occidental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28528 | MACARE-HBDC | larryamita74@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 782 ASSAY: Xpert SARS-CoV-2, COUNT: 394 ASSAY: Xpert HIV-1 Viral Load, COUNT: 61 | 2024-06-13 | GX-XVI-D-6C | 845005 | Module B2 for replacement | Replaced module B2 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B2 Generated reports | Unit is operational. Modules A1/903695, A3/655470, B1/902969, B3/902978, B4/902977, C2/903601, C4/902997, D1/902984, D3/903602, D4/902994 not operational. For reporting. |  | Old: B2/902976 New: B2/890344 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940745758841377553/5940745758841377553_signature_28.png | Larry Amita | https://www.jotform.com/uploads/macarewecare/222927978372067/5940745758841377553/5940745758841377553_signature_31.png
| 2024-06-13 | 2024-06-13 12:41 | Occidental Mindoro Provincial Hospital | Mamburao, Occidental Mindoro | 4B | Occidental Mindoro | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28529 | PBSP | larryamita74@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 165 ASSAY: Xpert HIV-1 Viral Load, COUNT: 7 |  | GX-IV-D-10C | 110019620 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940745758842646095/5940745758842646095_signature_28.png | Larry Amita | https://www.jotform.com/uploads/macarewecare/222927978372067/5940745758842646095/5940745758842646095_signature_31.png
| 2024-06-13 | 2024-06-13 14:33 | Roxas City Health Office | Roxas City, Capiz | 6 | Capiz | TB Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28527 | PBSP | ruby_camacho@ymail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1864 ASSAY: Xpert HIV-1 Viral Load, COUNT: 138 | 2024-06-13 | GXIV-4D | 805333 | Module B3 defective | Replace defective module B3/210091312 with module SN: 210079550 | #NAME? | All modules are functional. |  | 900-0491RS GeneXpert Module 6-Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940700051439364792/5940700051439364792_signature_28.png | Ma. Victoria So | https://www.jotform.com/uploads/macarewecare/222927978372067/5940700051439364792/5940700051439364792_signature_31.png
| 2024-06-13 | 2024-06-04 13:29 | Obando Rural Health Unit | Obando, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28526 | PBSP | romenvill@gmail.com | 25 |  |  | GXIV | 805252 | Unit for maintenance Unit for check up | Cleaned filter | #NAME? | Unit is operational Cepheid Engineer performed XpertCheck on May Modules 2 and 3 failed Modules 4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940682901309725928/5940682901309725928_signature_28.png | Rosalina Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5940682901309725928/5940682901309725928_signature_31.png
| 2024-06-13 | 2024-06-13 13:53 | Meycauayan City Health Unit VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28524 | DOH | ronamarieiris_25@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1674 ASSAY: Xpert HIV-1 Viral Load, COUNT: 138 | 2024-06-13 | GXIV | 831904 | Module A2(758217) for replacement UPS for troubleshooting | Replaced A2 by 688602 | #NAME? | Unit is operational UPS is recommended for replacement / repair SN: AT1000181128ST0080 |  | GeneXpert Module SN: 688602 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940682271301299106/5940682271301299106_signature_28.png | Rona Marie Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5940682271301299106/5940682271301299106_signature_31.png
| 2024-06-13 | 2024-06-13 11:45 | Bagumbong Health Center | Bagumbong, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28525 | DOH | jenilyn_duenas@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2799 |  | GXIV | 831907 | UPS for troubleshooting | NA | #NAME? | Unit is operational UPS is recommended for replacement / repair SN: AT1000180910ST0399 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940682271303807376/5940682271303807376_signature_28.png | Jenilyn Duenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5940682271303807376/5940682271303807376_signature_31.png
| 2024-06-13 | 2024-06-13 12:48 | Davao del Norte Provincial Hospital Samal | Brgy. Peñaplata, Island Garden City of Samal | 11 | Davao del Norte | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28523 | PBSP | allensoriano92@yahoo.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX-IV 4D | 815383 | Module B2 for service required | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Replaces the Defective Module SN:608788 with new Module SN:220000589 | Perform Module Replacement and Preventive Maintenance | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940680010962073302/5940680010962073302_signature_28.png | Allen Glenn B. Soriano RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5940680010962073302/5940680010962073302_signature_31.png
| 2024-06-13 | 2024-06-13 12:40 | Manito Rural Health Unit | Purok 2, It-Ba, Manito, Albay | 5 | Albay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28522 | PBSP | famelafatima2022@gmail.com | 32 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3268 |  | GX4-6C | 804257 | Gx for Calibration | XpertCheck Result: B1/884721 - passed B2/779135 - newly replaced module B3/600735 - service required B4/631003 - passed Cleaned cartridge bays, plunger rod, icore Generate report | Performed XpertCheck Performed Preventive Maintenance | Unitvis operational | B3/600735 - service required/for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940648387719574203/5940648387719574203_signature_28.png | Famela Fatima A. Magdaong | https://www.jotform.com/uploads/macarewecare/222927978372067/5940648387719574203/5940648387719574203_signature_31.png
| 2024-06-13 | 2024-06-13 12:30 | Laguna Chest Center | J. De Leon St., Sta Cruz Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28521 | PBSP | haizelmaganokumalao@rocketmail.com | 25 |  |  | GXIV | 807568 | Unit for installation. | Installation and training. | Performed installation. Performed basic training to enduser. Generated report. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940631069325890592/5940631069325890592_signature_28.png | Haizel M. Kumalao | https://www.jotform.com/uploads/macarewecare/222927978372067/5940631069325890592/5940631069325890592_signature_31.png
| 2024-06-13 | 2024-06-11 12:04 | Santolan Super Health Center | Evangelista Ave, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | For packing | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28518 | MACARE-HBDC | samandlet17@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 116 |  | GX-IV-D-10C | 110024245 | GeneXpert System for packing | Packed the GeneXpert System | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Packed the GeneXpert System: -GeneXpert unit -Workstation with keyboard, mouse and monitor -barcode scanner with stand -printer -UPS Generated reports | Unit is operational and ready for pickup. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940624286588288938/5940624286588288938_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5940624286588288938/5940624286588288938_signature_31.png
| 2024-06-13 | 2024-06-11 12:19 | Santolan Super Health Center | Evangelista Ave, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28519 | PBSP | samandlet17@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3028 |  | GX-IV-D-6C | 838898 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit os operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940624296586174533/5940624296586174533_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5940624296586174533/5940624296586174533_signature_31.png
| 2024-06-13 | 2024-06-11 12:08 | Santolan Super Health Center | Evangelista Ave, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28520 | MACARE-HBDC | samandlet17@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 100 |  | GX-IV-D-10C | 110024247 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940624296589023428/5940624296589023428_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5940624296589023428/5940624296589023428_signature_31.png
| 2024-06-13 | 2024-06-13 10:39 | Misamis Occidental Provincial Hospital | Poblacio | 10 | Misamis Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 28517 | PBSP | mophlaboratory2017@yahoo.com | 24 |  | 2024-06-13 | Gx 6c | 807618 | Module A4 hardware fail with serial number 210059819. | Replace module | .Replace module A4 with new serial number 210007029. . Perform preventive maintenance | All modules are functional ok to use. Ups not ok to use not charging. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940560031819772754/5940560031819772754_signature_28.png | Mary Auxilium D. Omndam | https://www.jotform.com/uploads/macarewecare/222927978372067/5940560031819772754/5940560031819772754_signature_31.png
| 2024-06-13 | 2024-06-13 10:06 | Culion Sanitarium General Hospital 10C | Robert St., Brgy. Tiza, Culion, Palawan | 4B | Palawan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28516 | PBSP | evecastillocalinawan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 535 ASSAY: Xpert SARS-CoV-2, COUNT: 1509 |  | GX IV 10C - D | 110013925 | UPS for troubleshooting | Troubleshoot UPS Performed Preventive Maintenance | Checked UPS UPS does not turn on No output voltage Temporarily uses the UPS of packed GX machine UPS for replacement Checked and cleaned cartridge bays, plunger rod, icore slot and filter | Unit is operational UPS needs replacement ASAP since power fluctuates most of the time |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940556963114773569/5940556963114773569_signature_28.png | Rose Eve C. Calinawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5940556963114773569/5940556963114773569_signature_31.png
| 2024-06-13 | 2024-06-13 09:58 | Culion Sanitarium General Hospital | Robert St., Brgy. Tiza, Culion Palawan | 4B | Palawan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28515 | MACARE-HBDC | evecastillocalinawan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 669 ASSAY: Xpert SARS-CoV-2, COUNT: 3425 | 2024-05-01 | GX IV 6C-D | 847725 | Modules B2 and B3 for replacement Unit due for XpertCheck | Performed Module replacement Performed Xpertcheck Performed PM | Replaced Modules B2: 210068598 -> 621192 B3: 619092 -> 633356 Performed Xpertcheck on modules B1 and B4 B1 passed; B4 failed Checked and cleaned cartridge bays, plunger rod, icore slot, and filter Generated report | Unit is operational 3 modules working |  | GX modules 6C XpertCheck Kit | PL 15724 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940555113112086191/5940555113112086191_signature_28.png | Rose Eve C. Calinawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5940555113112086191/5940555113112086191_signature_31.png
| 2024-06-13 | 2024-06-11 14:07 | MMC (Medilinx) | Amorsolo st., Makati City | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28514 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 195 ASSAY: Xpert CT/NG, COUNT: 834 ASSAY: Xpert HPV, COUNT: 808 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3039 ASSAY: Others, COUNT: 18 ASSAY: Xpert HCV VL, COUNT: 85 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1853 |  | GX 4 (10C) | 110001647 | As per end user they error 5007 in mtb ultra assay | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Troubleshooting | B1 and B3 for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940545650999527259/5940545650999527259_signature_28.png | Medi Linx Laboratory Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/5940545650999527259/5940545650999527259_signature_31.png
| 2024-06-13 | 2024-06-13 09:23 | Oroquieta City Health Office | Poblacion | 10 | Misamis Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 28513 | PBSP | oroquietacho.rtdl@gmail.com | 24 |  | 2024-06-13 | Gx4 6c | 834438 | Module B2 Xpertcheck fail. | .Replace module .Preventive Maintenance | .Replace module B2 with new serial number 210019009. .Perform preventive maintenance | Ok to use module B2,B3 under observation module B4 and B1 due to errors. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5940514421817270205/5940514421817270205_signature_28.png | Keneth Jane Lamoren | https://www.jotform.com/uploads/macarewecare/222927978372067/5940514421817270205/5940514421817270205_signature_31.png
| 2024-06-12 | 2024-06-12 16:33 | Chong Hua Hospital- Fuente | Capitol Site, Fuente, Cebu City | 7 | Cebu | Bloodbank | NO | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28512 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | MCS+ | 02F035 | Centrifuge cover lock problem during procedure of the platelet collection. | Aligned the Centrifuge cover lock switch aligned the switch and cover of the centrifuge. | Performed diagnostic of the Centrifuge cover lock switch aligned the switch and cover of the centrifuge. | MCS+ unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5939909041112046030/5939909041112046030_signature_28.png | Chong Hua Bloodbank | https://www.jotform.com/uploads/macarewecare/222927978372067/5939909041112046030/5939909041112046030_signature_31.png
| 2024-06-11 | 2024-06-11 15:41 | Lamitan city health office | Lamitan city | BARMM | Basilan | Laboratory | NO | INSTALLATION | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28511 | PBSP | lcholaboratory@gmail.com | 25 |  | 2024-06-11 | Gx 4 6c | 807460 | System unit for replacement. | Replaced System unit and Preventive Maintenence | Replaced System unit and Preventive Maintenence | All modules and system unit are in good condition | All modules and system unit are in good condition | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5939019908313045784/5939019908313045784_signature_28.png | Criselita A. Sojor | https://www.jotform.com/uploads/macarewecare/222927978372067/5939019908313045784/5939019908313045784_signature_31.png
| 2024-06-11 | 2024-06-11 14:47 | Lamitan city health office | Lamitan city | BARMM | Basilan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28510 | PBSP | lcholaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 | 2024-06-11 | GX 4 10c | 110020679 | Modules unable to read | Reset network and Preventive Maintenence | Reset network and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5939019908319081917/5939019908319081917_signature_28.png | Criselita A. Sojor | https://www.jotform.com/uploads/macarewecare/222927978372067/5939019908319081917/5939019908319081917_signature_31.png
| 2024-06-11 | 2024-06-11 15:27 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | rHivda | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28509 | MACARE-HBDC | jblmgh.is@gmail.com | 28 | ASSAY: Xpert HBV VL, COUNT: 126 ASSAY: Xpert HCV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 760 |  | GXIV-10c | 110013748 | Due for Calibration | Performed Calibration (ALL MODULES PASSED) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5939003882329031262/5939003882329031262_signature_28.png | Lerdie Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5939003882329031262/5939003882329031262_signature_31.png
| 2024-06-11 | 2024-06-06 16:32 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28508 | MACARE-HBDC | jblmgh.hpv@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 89 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 899 |  | GXIV-10c | 110012825 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938988722322884203/5938988722322884203_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/5938988722322884203/5938988722322884203_signature_31.png
| 2024-06-11 | 2024-05-29 14:18 | Region II Trauma and Medical Center | Bayombong, Nueva Vizcaya | 2 | Nueva Vizcaya | Bacteriology | NO | MAINTENANCE Disassemble | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28507 | PBSP | cecillepabance@gmail.com | 27 |  |  | GXIV-6c | 807458 | For Pullout | Disassemble Genexpert Packed Genexpert | Performed Maintenance Gathered Reports Performed Funtionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938988602324136249/5938988602324136249_signature_28.png | CECILLE A. HERMOSO | https://www.jotform.com/uploads/macarewecare/222927978372067/5938988602324136249/5938988602324136249_signature_31.png
| 2024-06-11 | 2024-06-11 14:58 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28506 | MACARE-HBDC | ejustincarlos@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 89 ASSAY: Xpert MTB/RIF Ultra, COUNT: 17 ASSAY: Others, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 899 |  | GXIV-10c | 110012825 | Due for Calibration | Performed Calibration (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938986692327833824/5938986692327833824_signature_28.png | Earvin Justin K. Carlos | https://www.jotform.com/uploads/macarewecare/222927978372067/5938986692327833824/5938986692327833824_signature_31.png
| 2024-06-11 | 2024-06-11 14:40 | CHONG HUA HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28505 | COMMERCIAL | chh_microbiology@chonghua.com.ph | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 26 ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 120 ASSAY: Xpert HBV VL, COUNT: 309 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1835 ASSAY: Xpert HCV VL, COUNT: 124 ASSAY: Xpert HIV-1 Viral Load, COUNT: 380 ASSAY: Others, COUNT: XDR - 1 |  | GX-IV | 110020777 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938976666115224050/5938976666115224050_signature_28.png | Matini E. Canedo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5938976666115224050/5938976666115224050_signature_31.png
| 2024-06-11 | 2024-06-11 14:21 | Cebu Doctors' University Hospital | CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28504 | COMMERCIAL | cduhcovidpoct@gmail.com | 18 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 69 ASSAY: Xpert CT/NG, COUNT: 180 ASSAY: Xpert HBV VL, COUNT: 165 ASSAY: Xpert HPV, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1122 ASSAY: Xpert HCV VL, COUNT: 139 ASSAY: Xpert HIV-1 Viral Load, COUNT: 15 |  | GX-IV | 110008524 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938971991119129109/5938971991119129109_signature_28.png | Alyssa C. Sayson, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5938971991119129109/5938971991119129109_signature_31.png
| 2024-06-11 | 2024-06-11 14:09 | Las Nieves Rural Health Unit | Brgy. Poblacion, Las Nieves | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28503 | WHO | Daisermt02@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 247 |  | GX-4 6 Colors | 847338 | Stuck Cartridge at Module B1 | None | Removing Stuck Cartridge Preventive Maintenance | All modules are functional All units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938957197414250669/5938957197414250669_signature_28.png | Daisy Siega | https://www.jotform.com/uploads/macarewecare/222927978372067/5938957197414250669/5938957197414250669_signature_31.png
| 2024-06-11 | 2024-06-11 13:39 | Sta. Lucia Health Center | Sta. Lucia Ave., Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28502 | DOH | josiedingle2004@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX IV 6C-D | 831901 | UPS for troubleshooting | Checked UPS Performed preventive maintenance | Troubleshoot UPS Checked Input Voltage from AVR 220V Checked Output Voltage 0V UPS does not work UPS needs replacement Temporarily plugged GX, CPU and monitor directly to AVR Checked and cleaned cartridge bays, plunger rod, icore slot and filter Some keys on keyboard are hard to press Generated report | Unit is operational All modules working UPS for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938945547919900441/5938945547919900441_signature_28.png | Josephine Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5938945547919900441/5938945547919900441_signature_31.png
| 2024-06-11 | 2024-06-11 13:28 | Kiblawan Rural Health Unit | Magsaysay Street, Poblacion, Kiblawan | 11 | Davao del Sur | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28501 | PBSP | chachindeasis0220@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 385 |  | GX-IV 4D | 835219 | For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log Install Xpc Version 1.5b Run XpertCheck on 2 Module | Perform Preventive Maintenance and XpertCheck Calibration | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938934297718191320/5938934297718191320_signature_28.png | Charice Dara P. De Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5938934297718191320/5938934297718191320_signature_31.png
| 2024-06-11 | 2024-06-11 13:28 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | rHivda | NO | MAINTENANCE Disassemble | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28500 | MACARE-HBDC | jblmgh.is@gmail.com | 28 |  |  | GXIV-6c | 812571 | Due for Pull out | Disassembled Genexpert Packed Genexpert | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938932582323215668/5938932582323215668_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5938932582323215668/5938932582323215668_signature_31.png
| 2024-06-11 | 2024-06-11 13:04 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28499 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2226 |  | GXIV-4D | 838561 | Due for monthly PM | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938920869129309804/5938920869129309804_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5938920869129309804/5938920869129309804_signature_31.png
| 2024-06-11 | 2024-06-11 12:16 | Bulan Rural Health Unit | Brgy Aquino, C. Gotladera st., Bulan, Sorsogon | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28498 | PBSP | bing.gonzales33@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2952 |  | GX4-6C | 807983 | B4/690369 for replacement | Replaced module: B4 - 690369>>793624 Cleaned cartridge bays, plunger rod, icore Generate reports | Replaced module B4 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938891499301801267/5938891499301801267_signature_28.png | Claribel G. Nacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5938891499301801267/5938891499301801267_signature_31.png
| 2024-06-11 | 2024-06-06 15:18 | Sta. Cruz Municipal Health Office | Sta. Cruz, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28497 | PBSP | stacruzrhudots@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 903 | 2024-06-06 | GXIV-4-D-6C | 808622 | #NAME? | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 660878 >>>632355 - Generated Reports | #NAME? | #NAME? |  | GeneXpert Module 632355 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938889839779190998/5938889839779190998_signature_28.png | Donna Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5938889839779190998/5938889839779190998_signature_31.png
| 2024-06-11 | 2024-06-05 11:51 | Agno Rural Health Unit | Agno, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28496 | PBSP | renahumildearenas@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 324 | 2024-06-05 | GXIV-2D-6C | 834837 | #NAME? | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 210057605 >>> 770928 - Performed XpertCheck - Generated Reports | #NAME? | - Unit is Operational - Did not perform XpertCheck on Newly Replaced Module B2 (770928) - Module B1 passed XpertCheck |  | GeneXpert Modules XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938888419713061400/5938888419713061400_signature_28.png | Maria Cris Rena Arenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5938888419713061400/5938888419713061400_signature_31.png
| 2024-06-11 | 2024-06-04 14:02 | Pangasinan Provincial Hospital | Bulingit, San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28495 | PBSP | judzfm@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 159 ASSAY: Others, COUNT: 2 |  | GXIV-4-D-10C | 120000484 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938887349715485291/5938887349715485291_signature_28.png | Judy Ann Mabunga | https://www.jotform.com/uploads/macarewecare/222927978372067/5938887349715485291/5938887349715485291_signature_31.png
| 2024-06-11 | 2024-06-11 11:13 | Lipa City Health Office | CM Recto Ave., Lipa Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28494 | PBSP | mylaolayao@yahoo.com | 25 |  | 2024-06-11 | GXIV | 805329 | Unit for PM. For module replacement of modules B1, B2 and B3. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B1 (804678), B2 (805004) and B3 (862375) in replacement to defective B1 (649040) B2 (649071) and B3 (210041477). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938863348321801394/5938863348321801394_signature_28.png | Myla Olayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5938863348321801394/5938863348321801394_signature_31.png
| 2024-06-11 | 2024-06-11 11:18 | Lipa City Health Office | CM Recto Ave., Lipa Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28493 | PBSP | mylaolayao@yahoo.com | 25 |  | 2024-06-11 | GXIV | 808642 | Unit for PM. For module replacement of module C2. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module C2 (800647) in replacement to defective C2 (600607). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938863358326071201/5938863358326071201_signature_28.png | Myla Olayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5938863358326071201/5938863358326071201_signature_31.png
| 2024-06-11 | 2024-06-11 10:09 | Commonwealth Health Center 10C | Katuparan St., Brgy. Commonwealth, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28492 | DOH | commonwealthclinical@gmail.com | 25 |  | 2024-05-28 | GX IV 10C-D | 110020279 | 10C unit for installation | Installed GX and accessories | Setup GX and accessories Setup system Configuration Setup User admin Installed ADF Changed Time and Date Generated report | Unut is operational all modules working |  | GX IV 10C HP Printer Intex UPS | 48598 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938823195269619470/5938823195269619470_signature_28.png | Danica Laisa Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5938823195269619470/5938823195269619470_signature_31.png
| 2024-06-10 | 2024-06-07 13:01 | Lucban Rural Health Unit | Lucban, Quezon | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28491 | PBSP | cudaninanniemarie@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 377 |  | GX-IV-D-6C | 838147 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | All modules operational. Module B2/806828 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938087733613577780/5938087733613577780_signature_28.png | Annie Marie Cudanin | https://www.jotform.com/uploads/macarewecare/222927978372067/5938087733613577780/5938087733613577780_signature_31.png
| 2024-06-10 | 2024-06-10 13:49 | Talisay City City Health Unit | Cansojong, Talisay City Cebu | 7 | Cebu | Tb laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28490 | PBSP | marydelgella@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1840 |  | GX4 | 815384 | Due for annual calibration Due for Preventive Maintenance | Performed machine maintenance clean ICore, plunger, cartridge bay, filter Conduct calibration in genexpert machine | Performed machine maintenance clean ICore, plunger, cartridge bay, filter Conduct calibration in genexpert machine | All module passed and genexpert machine is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938080181116313756/5938080181116313756_signature_28.png | Marydel Gella, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5938080181116313756/5938080181116313756_signature_31.png
| 2024-06-10 | 2024-06-07 11:56 | Buluan District Hospital | Buluan, Maguindanao | BARMM | Maguindanao | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28489 | PBSP | bdhclinicallaboratory@gmail.com | 28.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1627 |  | Gx-IV | 815490 | Module 2 and 4 for Service Required | Genexpert machine Cleaning/Preventive maintenance Genexpert machine annual Xpertcheck/Calibration Checking of errors Problem/errors reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules | Module 1 and 3 passed the Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938025907312993947/5938025907312993947_signature_28.png | BDH LABORATORY | https://www.jotform.com/uploads/macarewecare/222927978372067/5938025907312993947/5938025907312993947_signature_31.png
| 2024-06-10 | 2024-06-10 12:08 | Giporlos Rural Health Unit | Giporlos Eastern Samar | 8 | Eastern Samar | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28488 | PBSP | ningrediang@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 909 | 2024-06-10 | GXIV -2D | 835060 | Module B1/787877 defective | #NAME? | - Module replacement B1/787877 into B1/713211 - Performed preventive maintenance - Generate reports | Unit is operational |  | 1 Genexpert module 6c | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5938019725894895619/5938019725894895619_signature_28.png | Luningning R. Abuda | https://www.jotform.com/uploads/macarewecare/222927978372067/5938019725894895619/5938019725894895619_signature_31.png
| 2024-06-10 | 2024-06-04 13:55 | Pangasinan Provincial Hospital | Bulingit, San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28487 | PBSP | judzfm@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2493 | 2024-06-04 | GXIV-4-D-6C | 808283 | #NAME? | - Assembled GeneXpert System - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 210067929 with New Module 688593 - Performed XpertCheck - Generated Reports - Disassembled and Repackaged GeneXpert System | #NAME? | - Unit is Operational - Did not run XpertCheck on newly replaced Module A4 688593 - Modules A1, A2 and A3 passed XpertCheck |  | - GeneXpert Module 688593 - XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5937979759712958578/5937979759712958578_signature_28.png | Judy Mabunga | https://www.jotform.com/uploads/macarewecare/222927978372067/5937979759712958578/5937979759712958578_signature_31.png
| 2024-06-10 | 2024-06-04 08:46 | San Jacinto Rural Health Unit | San Jacinto, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28486 | DOH | allen.indong@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 298 | 2024-05-02 | GXIV-4-D-6C | 838518 | #NAME? | #NAME? | #NAME? | #NAME? |  | XpertCheck Kits | DR 11453 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5937978379718733725/5937978379718733725_signature_28.png | Alexander Indong | https://www.jotform.com/uploads/macarewecare/222927978372067/5937978379718733725/5937978379718733725_signature_31.png
| 2024-06-10 | 2024-06-10 08:43 | Danao City Health Office | Danao City, Cebu | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28485 | PBSP | chockenriquez@gmail.com | 29 |  |  | Gx IV | 808286 | Due date for Site Visit/PM | Gathering data IQR, System log & assay stats Clean Fan Filter | Gathering data IQR, System log & assay stats Clean Fan Filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5937901981341876662/5937901981341876662_signature_28.png | Nicanor Enriquez Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5937901981341876662/5937901981341876662_signature_31.png
| 2024-06-08 | 2024-06-07 13:54 | Payatas B Health Center | Bulacan st., Payatas QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28484 | DOH | racquelmrocero@gmail.com | 37 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2374 |  | GX4 | 831899 | Due for Maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5936060590929178442/5936060590929178442_signature_28.png | Racquel M.Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5936060590929178442/5936060590929178442_signature_31.png
| 2024-06-08 | 2024-06-07 10:20 | Philippine Children's Medical Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | mjcordova02@yahoo.com |  | 28483 | PBSP | kate_july27@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2668 | 2024-06-07 | GX 4 D | 807672 | Due for maintenance and xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance and xpert check | B2 newly replaced B1 failed xpert check B3,B4 passed xpertcheck |  | Xpertcheck kits | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5936060200929729897/5936060200929729897_signature_28.png | Katherine Tabaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5936060200929729897/5936060200929729897_signature_31.png
| 2024-06-07 | 2024-06-07 14:58 | Capitol Medical Center | Scout Magbanua, Quezon Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28482 | CONSORTIUM | princeosenarmt@gmail.com | 24 |  |  | GXIV-D-10C | 110024553 | Unit is due for Installation. | Installation. | Checked Output Voltage of Outlet and UPS. Install GeneXpert System. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935530648219417437/5935530648219417437_signature_28.png | Prince Henry Osena | https://www.jotform.com/uploads/macarewecare/222927978372067/5935530648219417437/5935530648219417437_signature_31.png
| 2024-06-07 | 2024-06-07 14:50 | San Francisco Rural Health Unit | Brgy. 4, San Francisco, Agusan del Sur | 13 | Agusan del Sur | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28481 | PBSP | periabrasangelica@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 986 |  | GX-2 6 Colors | 835078 | Module B2 found error in temperature | Module Replacement Calibration Preventive Maintenance | Replacing Module B2 with SN: 696050 with SN: 784340 Preventive Maintenance XPC Calibration | All modules passed in XPC All modules are working and functional All units are working and functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935529648115137126/5935529648115137126_signature_28.png | Ma. Angelica Ga | https://www.jotform.com/uploads/macarewecare/222927978372067/5935529648115137126/5935529648115137126_signature_31.png
| 2024-06-07 | 2024-06-07 14:24 | Arsenia J. Maximo Health Center | Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28480 | PBSP | sheryllopez310@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 97 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1963 ASSAY: Xpert HIV-1 Viral Load, COUNT: 123 |  | GX IV 6C-D | 814162 | Unit due for maintenance | Performed Preventive maintenance Performed reliability test | Checked and cleaned cartridge bays, plunger rod, icore slot, and filter Performed reliability test All modules passed Module B1 for monitoring, frequent error 5007 Archived results Generated report | Unit is operational Module B1 for cheling by technical dept. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935523393618076814/5935523393618076814_signature_28.png | Sheryl Lopez' | https://www.jotform.com/uploads/macarewecare/222927978372067/5935523393618076814/5935523393618076814_signature_31.png
| 2024-06-07 | 2024-06-06 13:04 | Calaca Rural Health Unit | Dela Salud St., Calaca, Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28479 | DOH | Calacacityhealthlaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1993 |  | GXIV-D-6C | 821300 | Moduled 3 and 4 with Temperature Error. | Module Replacement. Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Replaced Modules 3 and 4. Performed Self-test. Generate Reports. | All Modules are Operational. | A3: 732771 -> 897123 , A4:731307->909058 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519778213454505/5935519778213454505_signature_28.png | Baldwin Catibog | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519778213454505/5935519778213454505_signature_31.png
| 2024-06-07 | 2024-05-29 12:37 | Moonwalk Health Center | St. Francis St., San Agustin Village, Moonwalk, Parañaque | NCR | NCR 4th - Parañaque | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28478 | PBSP | almaaurea@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1557 |  | GXIV-D-6C | 805113 | Unit is due for XpertCheck. CPU has Date and Time Problem. | XpertCheck. Troubleshooting. Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod, Icore Optics. Checked Filter if Clogged. Performed Selt-test. Performed XpertCheck. Replaced Bios Battery. Generate Reports. | All Modules Passed. CPU Problem Resolved. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519528214329765/5935519528214329765_signature_28.png | Alma Aurea Limson | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519528214329765/5935519528214329765_signature_31.png
| 2024-06-07 | 2024-06-04 09:44 | Caridad Health Center | Dela Cruz St, Caridad, Cavite City, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28477 | DOH | seyerlanie85@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1371 |  | GXIV-D-6C | 821227 | Module 4 with Temperature Error. | Module Replacement. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Replaced Module 4. Performed Self-test. Generate Reports. | All Modules are Operational. | A4: 730140 -> 884033 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519528217012559/5935519528217012559_signature_28.png | Melanie Reyed | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519528217012559/5935519528217012559_signature_31.png
| 2024-06-07 | 2024-06-05 15:09 | Silang Rural Health Unit | J. Rizal St., Silang, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28476 | DOH | silangrhu@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1869 |  | GXIV-D-6C | 820708 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | 2 Modules were Disabled. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519528212632724/5935519528212632724_signature_28.png | Benessie Madlangsakay | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519528212632724/5935519528212632724_signature_31.png
| 2024-06-07 | 2024-05-29 09:29 | San Isidro Health Center | Dr. A. Santos Ave., San Isidro, Parañaque | NCR | NCR 4th - Parañaque | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28475 | PBSP | herminadoa@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2063 ASSAY: Xpert HIV-1 Viral Load, COUNT: 84 |  | GXIV-D-6C | 838192 | Unit is due for XpertCheck | XpertCheck. Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked if Filter is Clogged. Performed Self-test. Performed XpertCheck. Generate Reports. | 2 Retry, 1 Service Required and 1 Passed. | Retry: B1/ 784193, B2/ 788537. Service Required: B3/ 210091236 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519528217414063/5935519528217414063_signature_28.png | Amelia G. Herminado | https://www.jotform.com/uploads/macarewecare/222927978372067/5935519528217414063/5935519528217414063_signature_31.png
| 2024-06-07 | 2024-06-07 14:13 | Silay City Health Office | Silay City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 28474 | PBSP | gutierrezdaviddior@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3599 |  | GX4 6 Color | 805093 | Modules A2, A3 and A4 failed in XPC. While module A1 is not due for calibration. | Save files(GXC file, IQR, System Log and Assay Stats) for reporting. | Run Xpert check calibration for 3 modules( A2, A3 and A4) | Module A1 is in good condition while module A2, A3 & A4 need service. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935504487218914144/5935504487218914144_signature_28.png | David Dior V. Gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5935504487218914144/5935504487218914144_signature_31.png
| 2024-06-07 | 2024-06-07 13:32 | SILIMAN UNIVERSITY MEDICAL CENTER | DUMAGUETE CITY | 7 | Negros Oriental | MOLECULAR LABORATORY | YES | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com | badayos@gmail.com |  | 28473 | COMMERCIAL | clab@sumcfi.org | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert C. difficile BT, COUNT: 17 ASSAY: Xpert CT/NG, COUNT: 20 ASSAY: Xpert HBV VL, COUNT: 137 ASSAY: Xpert HPV, COUNT: 92 ASSAY: Xpert MTB/RIF Ultra, COUNT: 571 ASSAY: Xpert SARS-CoV-2, COUNT: 125 ASSAY: Xpert HCV VL, COUNT: 19 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 ASSAY: Others, COUNT: RSV 10 | 2024-05-30 | GX IV | 821471 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE REPLACEMENT: MODULE A1 DEF SN:792139 , NEW SN:220016262 MODULE A4 DEF SN:210070900 , NEW SN:220022646 | EQUIPMENT IS OPERATIONAL |  | GENEXPERT MODULE | SI 98123 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935499677094053896/5935499677094053896_signature_28.png | George Guevarra | https://www.jotform.com/uploads/macarewecare/222927978372067/5935499677094053896/5935499677094053896_signature_31.png
| 2024-06-07 | 2024-06-07 13:29 | Daram Rural Health Unit | Daram Rural Health Unit | 8 | Samar | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28472 | PBSP | nenitaamistoso@gmail.com | 26 |  | 2024-06-07 | GXIV 6C | 838868 | -Module B2/863402 defective | #NAME? | - Performed preventive maintenance - Module replacement B2/863402 into 883824 - Generate reports | Unit is operational |  | 1 Genexpert module 6c | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935484280027788679/5935484280027788679_signature_28.png | Nenita Amistoso | https://www.jotform.com/uploads/macarewecare/222927978372067/5935484280027788679/5935484280027788679_signature_31.png
| 2024-06-07 | 2024-06-07 13:01 | Davao Reproductive Health and Wellness Center | Jacinto Street, Poblacion District | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28471 | MACARE-HBDC | rhwclab@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 197 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1172 ASSAY: Xpert HBV VL, COUNT: 9 |  | GX-IV 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Only Three Modules are operational Module 3 disabled; reported to Cepheid for further assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935463476956578333/5935463476956578333_signature_28.png | Cherry Ann Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5935463476956578333/5935463476956578333_signature_31.png
| 2024-06-07 | 2024-06-07 13:04 | Davao Reproductive Health and Wellness Center | Jacinto Street, Poblacion District | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28470 | DOH | rhwclab@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 197 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1172 ASSAY: Xpert HBV VL, COUNT: 9 |  | GX-IV 4D | 838262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Only Two Modules are operational Modules C2 and C4 disabled, reported to Cepheid for further assessments |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935463036954222767/5935463036954222767_signature_28.png | Cherry Ann Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5935463036954222767/5935463036954222767_signature_31.png
| 2024-06-07 | 2024-06-07 12:09 | Central Luzon Doctor's Hospital | Hospital drive Brgy. Sa Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | gmgalamay@cldh.com.ph |  | 28469 | COMMERCIAL | cldhlaboratory@cldh.com.ph | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 366 |  | GXIV-D | 830841 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935435904748136630/5935435904748136630_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/5935435904748136630/5935435904748136630_signature_31.png
| 2024-06-07 | 2024-06-07 12:05 | Central Luzon Doctor's Hospital | Hospital drive Brgy. San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | gmgalamay@cldh.com.ph |  | 28468 | COMMERCIAL | cldhlaboratory@cldh.com.ph | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 ASSAY: Xpert SARS-CoV-2, COUNT: 366 |  | GXIV-D | 839232 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational excet B4:738203 error(2014) Temperature error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935434764743523448/5935434764743523448_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/5935434764743523448/5935434764743523448_signature_31.png
| 2024-06-07 | 2024-06-05 11:23 | Taytay Rural Health Unit - 5B | Ynares Ave Taytay, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28466 | PBSP | zenresurreccion@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1477 |  | GX-IV-D-6C | 838136 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Module B1 newly replaced and operational. Modules B2, and B4 passed and operational. Module B3/806713 service required and not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935433031871782393/5935433031871782393_signature_28.png | Zenilynn Bobiles | https://www.jotform.com/uploads/macarewecare/222927978372067/5935433031871782393/5935433031871782393_signature_31.png
| 2024-06-07 | 2024-06-05 13:45 | Pasig City General Hospital | Eusebio St. Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28465 | PBSP | pcghlaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2993 |  | GX-IV-D-6C | 805010 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935433031873733634/5935433031873733634_signature_28.png | Norma Angelique Valencia,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5935433031873733634/5935433031873733634_signature_31.png
| 2024-06-07 | 2024-06-06 12:13 | General Nakar Rural Health Unit | General Nakar, Quezon | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28467 | PBSP | buendichol@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 449 |  | GX-IV-D-6C | 838180 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | All modules passed. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935433031879604424/5935433031879604424_signature_28.png | Lydelyn B.Costumbrado | https://www.jotform.com/uploads/macarewecare/222927978372067/5935433031879604424/5935433031879604424_signature_31.png
| 2024-06-07 | 2024-06-07 12:14 | Ipil Rural Health Unit | Ipil, Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28464 | PBSP | rheasolana@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1526 ASSAY: Xpert HIV-1 Viral Load, COUNT: 60 | 2024-06-07 | GX 4 (6c) | 808615 | None | Preventive Maintenence and Site visit | Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935431624759529808/5935431624759529808_signature_28.png | Rheamay E. Solana | https://www.jotform.com/uploads/macarewecare/222927978372067/5935431624759529808/5935431624759529808_signature_31.png
| 2024-06-07 | 2024-06-07 11:51 | Binan City Health Office 1 | Binan Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28463 | PBSP | evalexaa29@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2426 | 2024-06-07 | GXIV | 812780 | Unit for PM. For module replacement of module B2. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B2 (732611) in replacement to defective B2 (799008). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935418707808310804/5935418707808310804_signature_28.png | Alexandra Mae Evangelista | https://www.jotform.com/uploads/macarewecare/222927978372067/5935418707808310804/5935418707808310804_signature_31.png
| 2024-06-07 | 2024-06-07 11:31 | Tarlac City Health Center X | Brgy. Tibag, Talac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28462 | DOH | jayteecornelius@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 869 |  | GXIV-D | 831587 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit us operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935405650223455475/5935405650223455475_signature_28.png | Jaytee Cornelius Bachiller | https://www.jotform.com/uploads/macarewecare/222927978372067/5935405650223455475/5935405650223455475_signature_31.png
| 2024-06-07 | 2024-06-07 11:29 | Dr. George T. Hofer Medical Center | Purok Santan, Barangay tenan, ipil Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | CALIBRATION MAINTENANCE Packed Genexpert Machine | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 28461 | PBSP | exelrose.remotigue@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 880 | 2024-06-06 | GX 2 (6c) | 834831 | None | Preventive Maintenence, Xpertcheck Calibration and Packed Genexpert Machine with other accessories | Preventive Maintenence, Xpertcheck Calibration and Packed Genexpert Machine with other accessories | Genexpert Machine and other accessories are packed. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935405248102418764/5935405248102418764_signature_28.png | Exel Rose Cainglet | https://www.jotform.com/uploads/macarewecare/222927978372067/5935405248102418764/5935405248102418764_signature_31.png
| 2024-06-07 | 2024-06-07 11:28 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Mabini, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28460 | PBSP | bryanmichaelbalunes@gmail.com | 21 | ASSAY: Others, COUNT: 11 |  | GXIV-10c | 110020311 | UPS not working | Troubleshoot UPS Replaced Power Cable Replace Fuse | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935404362523959951/5935404362523959951_signature_28.png | Bryan Michael Balunes | https://www.jotform.com/uploads/macarewecare/222927978372067/5935404362523959951/5935404362523959951_signature_31.png
| 2024-06-07 | 2024-06-07 11:06 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Mabini, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Bacteriology | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28459 | PBSP | bryanmichaelbalunes@gmail.com | 29 |  |  | GXXVI-6c | 844111 | UPS cannot backup Due for Calibration | Troubleshoot UPS Performed Calibration (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935392972522517407/5935392972522517407_signature_28.png | Bryan Michael Balunes | https://www.jotform.com/uploads/macarewecare/222927978372067/5935392972522517407/5935392972522517407_signature_31.png
| 2024-06-07 | 2024-06-07 10:29 | Corazon Locsin Montelibano M.R.H | Lacson St., Bacolod City | 6 | Negros Occidental | Molecular Laboratory | NO | Transfer of Location | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 28458 | MACARE-HBDC | alainjuayang@yahoo.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1006 |  | GX4 10 Color | 110020743 | NONE | Pulled out the machines and its accessories together with their two BioMed engineers and transfer to their Molecular Lab. | Transfer of their newly procurred GeneXpert IV 10 C machine from Immuno/Sero Dept. to Molecular Laboratory. | GeneXpert Machine and its accessories are in good working conditions. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935376706019198424/5935376706019198424_signature_28.png | Alain Juayang | https://www.jotform.com/uploads/macarewecare/222927978372067/5935376706019198424/5935376706019198424_signature_31.png
| 2024-06-07 | 2024-06-07 10:20 | Tarlac City Health Center 2 | Brgy. Matatalaib Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28457 | DOH | clarizannebeltran1996@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 470 |  | GXIV-D | 831917 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935374325613251825/5935374325613251825_signature_28.png | Clariz Anne Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5935374325613251825/5935374325613251825_signature_31.png
| 2024-06-07 | 2024-06-06 11:06 | Sta Ignacia Rural Health Unit | Sta. Ignacia, Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28456 | DOH | jdi0973@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 504 |  | GXIV-D | 831909 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935359895614465214/5935359895614465214_signature_28.png | Jess D. Ibarra | https://www.jotform.com/uploads/macarewecare/222927978372067/5935359895614465214/5935359895614465214_signature_31.png
| 2024-06-07 | 2024-06-06 10:16 | Gerona Rural Health Unit | Brgy. Pinasling, Gerona, Talac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28455 | DOH | huwakie15@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 |  | GXIV-D | 831693 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit us operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935359585615035897/5935359585615035897_signature_28.png | Artemio Nool III | https://www.jotform.com/uploads/macarewecare/222927978372067/5935359585615035897/5935359585615035897_signature_31.png
| 2024-06-07 | 2024-06-06 09:07 | Victoria Rural Health Unit | Brgy. Bulo Victoria, Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28454 | DOH | diazrmt@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 756 |  | GXIV-D | 831669 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935359325615815455/5935359325615815455_signature_28.png | Alma Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5935359325615815455/5935359325615815455_signature_31.png
| 2024-06-07 | 2024-06-07 09:42 | Wenceslao Dela Paz Health Center | Culiat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28453 | DOH | ameliafclemente2001@yahoo.com | 29 | ASSAY: Xpert HPV, COUNT: 211 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2973 |  | GX IV 6C-D | 831876 | Unit due for maintenance | Performed Preventive maintenance Performed reliability Test Checked UPS | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test All modules passed UPS does not back up Generated report | Unit is operational All modules working UPS does not seotch to battery mode Needs UPS replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5935351653611530997/5935351653611530997_signature_28.png | Felicidad R. De Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5935351653611530997/5935351653611530997_signature_31.png
| 2024-06-06 | 2024-06-05 11:56 | St. Elizabeth Hospital Inc. | General Santos | 12 | South Cotabato | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28452 | COMMERCIAL | covidsupport@sehi.ph | 20.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 53 ASSAY: Xpert SARS-CoV-2, COUNT: 8 |  | Gx-IV (10C) | 110003713 | Module 2, 3, and 4 requested for replacement due to multiple temperature error | Genexpert machine Cleaning/Preventive maintenance Checking of errors/problem Problem/errors reporting Data collection | Perform Genexpert machine Preventive maintenance | 1 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934956314311505211/5934956314311505211_signature_28.png | Norjanah Antap | https://www.jotform.com/uploads/macarewecare/222927978372067/5934956314311505211/5934956314311505211_signature_31.png
| 2024-06-06 | 2024-06-05 11:35 | St.Elizabeth Hospital Inc | General Santos City | 12 | South Cotabato | Micrology | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28451 | COMMERCIAL | therese.cariaga@medilinx.com.ph | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 268 |  | Gx-IV | 831313 | Module 1 and 3 multiple temperature errors Module 4 ultrasonic problem | Genexpert machine Cleaning/ Preventive maintenance Checking of errors Problem/errors reporting Data collection | Perform Genexpert machine Preventive maintenance | 1 operational module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934955674314927596/5934955674314927596_signature_28.png | Therese Christia Diana Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5934955674314927596/5934955674314927596_signature_31.png
| 2024-06-06 | 2024-06-06 10:53 | Lake Sebu Rural Health Unit | Poruk Rosa's, Brgy. Poblacion, Lake Sebu | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28450 | DOH | jennypingkian23@yahoo.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 631 |  | Gx-IV | 838569 | Module 2 stuck cartridge | Repair module 2 stuck cartridge Genexpert machine Cleaning Checking of errors Problem/errors reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine module repair | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934954544316570967/5934954544316570967_signature_28.png | Jenny A. Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/5934954544316570967/5934954544316570967_signature_31.png
| 2024-06-06 | 2024-06-06 16:32 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28449 | MACARE-HBDC | jblmgh.hpv@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 89 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 899 |  | GXIV-10c | 110012825 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934737507727675338/5934737507727675338_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/5934737507727675338/5934737507727675338_signature_31.png
| 2024-06-06 | 2024-06-06 16:58 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28448 | PBSP | jblmgh.hpv@gmail.com | 25 |  |  | GXIV-6c | 804763 | Due for Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Combined with another GXIV-6c | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934736705413656594/5934736705413656594_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/5934736705413656594/5934736705413656594_signature_31.png
| 2024-06-06 | 2024-06-06 16:53 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28447 | PBSP | jblmgh.hpv@gmail.com | 25 |  |  | GXIV-6c | 834445 | Always Lost Communication | Reset Gatewayboard | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | For observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934734785413159064/5934734785413159064_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/5934734785413159064/5934734785413159064_signature_31.png
| 2024-06-06 | 2024-06-06 16:48 | Dr. George T. Hofer Medical Center | Purok santa, barangay tenant, ipil Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | INSTALLATION Site assessment | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28446 | PBSP | exelrose.remotigue@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-06-06 | GX 4 (10c) | 110020394 | None | Genexpert Machine 10c Installation and Site assessment | Genexpert Machine 10c Installation and Site assessment | All modules are ready to use. Uninterrupted power supply is broken and will be replaced | Uninterrupted power supply is broken and will be replaced | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934732508103349216/5934732508103349216_signature_28.png | Exel Rose Cainglet | https://www.jotform.com/uploads/macarewecare/222927978372067/5934732508103349216/5934732508103349216_signature_31.png
| 2024-06-06 | 2024-06-06 16:28 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28445 | PBSP | jblmgh.hpv@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Others, COUNT: 1096 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 |  | GXIV-10c | 110016957 | Module 3 error | Performed Module Replacement B3(210084051 - 898071) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934731305411902505/5934731305411902505_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/5934731305411902505/5934731305411902505_signature_31.png
| 2024-06-06 | 2024-06-06 14:49 | San Pablo Colleges Medical Cente | Maharlika Highway, San Pablo | 4A | Laguna | Molecular laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28444 | CONSORTIUM | templo.richard@yahoo.com | 25 |  |  | GXIV-10C | 110024552 | Unit for installation. | Installation. | Performed installation of GeneXpert. Performed maintenance training with Sir Noel Vicente. Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934695188424362714/5934695188424362714_signature_28.png | Richard Templo | https://www.jotform.com/uploads/macarewecare/222927978372067/5934695188424362714/5934695188424362714_signature_31.png
| 2024-06-06 | 2024-06-06 14:35 | Dr. Elvira Lagrosa Health Center (Cuyegkeng) | R. Layug, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28443 | PBSP | demlhc.pasay@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1416 |  | GX IV 6C-D | 805106 | Unit due for Preventive Maintenance | Performed Preventive maintenance Performed reliability test | Checked and cleaned cartridge bays, plunger rod, icore slot, and filter Performed Reliability test All modules passed Generated report | Unit is operational All modules working UPS not switching to battery, needs replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934690183614105851/5934690183614105851_signature_28.png | Georgette Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5934690183614105851/5934690183614105851_signature_31.png
| 2024-06-06 | 2024-06-06 14:33 | Baao Rural Health Unit | Del Rosario | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28442 | PBSP | demibreboneria@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 567 |  | GX4-6C | 808639 | GX4 for Calibration | XpertCheck Result: A1/667540 - passed A2/210005380 - passed A3/210008802 - passed A4/210009609 - passed Cleaned cartridge bays, plunger rod, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934680716413955576/5934680716413955576_signature_28.png | Demi ann breboneria | https://www.jotform.com/uploads/macarewecare/222927978372067/5934680716413955576/5934680716413955576_signature_31.png
| 2024-06-06 | 2024-06-06 14:07 | Philippine Red Cross (PRC Tower) | Edsa,Mandaluyong City | NCR | NCR 2nd - Mandaluyong | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28441 | MACARE-HBDC | marnelli.manuel@gmail.com | 27 |  |  | MCS + | 08B049 | Encounter CRC bios failed | Checked unit, Checked and measured battery voltage output of CPU board 0 volts reading. Performed CPU board reset. | Performed troubleshooting | Unit unoperational For replacement of CPU board |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934642871313578149/5934642871313578149_signature_28.png | Marnelli Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/5934642871313578149/5934642871313578149_signature_31.png
| 2024-06-06 | 2024-06-06 13:17 | Hagonoy Rural Health Unit | Poblacion, Hagonoy | 11 | Davao del Sur | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28440 | WHO | rhuhagonoy@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 466 |  | GX-IV 4D | 847256 | For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install the new version of XPC Run XpertCheck on 4 Modules | Perform Preventive Maintenance Perform Xpertcheck Calibration | All Modules Passed the Calibration Machine is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934607725422968152/5934607725422968152_signature_28.png | ARNEL CRESAR G. NEROSA | https://www.jotform.com/uploads/macarewecare/222927978372067/5934607725422968152/5934607725422968152_signature_31.png
| 2024-06-06 | 2024-06-06 11:55 | Samar Provincial Hospital | Catbalogan Samar | 8 | Samar | TB-DOTS | YES | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28439 | PBSP | janicedytan@gmail.com | 26 |  | 2024-06-06 | GXIV 10C | 120000737 | n/a | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934560341223318720/5934560341223318720_signature_28.png | Janice Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5934560341223318720/5934560341223318720_signature_31.png
| 2024-06-06 | 2024-06-06 09:35 | Klinika Bernardo (QC Health Department) | Kalayaan Ave., Quezon City | NCR | NCR 2nd - Quezon City | laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28438 | PBSP | klinikabernardolaboratory@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 14 ASSAY: Xpert HBV VL, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 394 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1337 |  | GX IV 6C-D | 808618 | GX reuested for transfer to another room | Transferred GX Performed Reliability Test Performed Preventive Maintenance | Setup GX to new location Check and cleaned cartridge bays, plunger rod, icore slot, and filter Performed Reliability test Archived past results Generated report | Unit is operational all modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934477762045366017/5934477762045366017_signature_28.png | msileen reyna | https://www.jotform.com/uploads/macarewecare/222927978372067/5934477762045366017/5934477762045366017_signature_31.png
| 2024-06-06 | 2024-06-06 08:50 | Carrascal Rural Health Unit | Poblacion, Carrascal | 13 | Surigao del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28437 | PBSP | rhucarrascal2016@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1185 |  | GX-2 6 Colors | 835098 | B2 found temperature error | Module Replacement Preventive Maintenance | Replacing Module B2 with SN: 787993 to SN: 760645 Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934445680916021665/5934445680916021665_signature_28.png | Roela O. Garay | https://www.jotform.com/uploads/macarewecare/222927978372067/5934445680916021665/5934445680916021665_signature_31.png
| 2024-06-06 | 2024-06-06 08:35 | Leyte Provincial Hospital | Pawing palo leyte | 8 | Leyte | Laboratory | YES | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28436 | PBSP | lphlabdeptcovid19@gmail.com | 27 |  | 2024-06-03 | GXIV 10C | 110018056 | n/a | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5934435664753079291/5934435664753079291_signature_28.png | Kareen Tampil | https://www.jotform.com/uploads/macarewecare/222927978372067/5934435664753079291/5934435664753079291_signature_31.png
| 2024-06-05 | 2024-06-05 14:56 | Quezon City Jail | Payatas QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28435 | PBSP | carleecruz0@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1417 |  | GX IV D | 803089 | UPS unable to power on due to burn out power supply of ups | Checked voltage output of main outlet 230v, Checked and test power cord continuity. Checked power supply | Troubleshooting of ups | UPS defective Genexpert and CPU operational | For replacement of UPS |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933899775897364063/5933899775897364063_signature_28.png | Carlee Mae L Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5933899775897364063/5933899775897364063_signature_31.png
| 2024-06-05 | 2024-06-05 15:58 | Vinzons RHU | Vinzons | 5 | Camarines Norte | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28434 | DOH | bardonamabel@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 971 |  | GX4-6C | 838576 | 2037, 2005 Error | Cleaned icore Check collected data test result Generate report | Performed Troubleshooting | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933840099726527689/5933840099726527689_signature_28.png | Amabel A. Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/5933840099726527689/5933840099726527689_signature_31.png
| 2024-06-05 | 2024-06-05 15:35 | Philippine Orthopedic Center | Banawe, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28433 | PBSP | cmblanceta@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1925 | 2024-04-30 | GX IV 6C-D | 803088 | Unit due for maintenance | Performed Preventive maintenance Performed Xpertcheck | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck B1: Failed B2: Passed B3: Passed B4: Passed Generated report | unit is operational B1 for reporting |  | XpertCheck | DR 48180 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933834646274433348/5933834646274433348_signature_28.png | Claire Marie B. Lanceta | https://www.jotform.com/uploads/macarewecare/222927978372067/5933834646274433348/5933834646274433348_signature_31.png
| 2024-06-05 | 2024-06-04 13:43 | Doña Marta Health Center | Doña Marta, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | CALIBRATION | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28432 | DOH | donamartaclinicallaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1431 | 2024-05-24 | GX IV 6C-D | 831700 | Unit due for calibration | Performed Preventive maintenance Performed Xpertcheck | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck All modules passed Generated report | unit is operational All modules working |  | Xpertcheck kit | PL 015867 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933821136279604489/5933821136279604489_signature_28.png | Tessa Rose Tabugo | https://www.jotform.com/uploads/macarewecare/222927978372067/5933821136279604489/5933821136279604489_signature_31.png
| 2024-06-05 | 2024-06-05 14:35 | Madrid District Hospital | Madrid, Surigao del Sur | 13 | Surigao del Sur | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28431 | DOH | madrid_dh@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1038 |  | GX-4 6 colors | 838439 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933795232322945383/5933795232322945383_signature_28.png | Rochelle J. Pueblas | https://www.jotform.com/uploads/macarewecare/222927978372067/5933795232322945383/5933795232322945383_signature_31.png
| 2024-06-05 | 2024-06-05 14:15 | Camarines Norte Provincial Hospital | Brgy. 5. Bagasbas. Daet | 5 | Camarines Norte | Clinical Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28430 | PBSP | cnphlaboratory@gmail.com | 28 |  |  | GX4-6C | 807600 | GX4-6C 807600 for Installation GX for Calibration | Installed GX4-10C, CPU , Monitor, UPS, Printer Checked/Adjust GX information details XpertCheck Result: B1/623932 - passed B2/680631 - passed B3/662501 - passed B4/210101552 - passed Cleaned cartridge bays, plunger rod, icore Generate reports | Installed Gx set Performed XpertCheck Performed Preventive Maintenance | Unit is operational | Gx4-6C(807600) relocation from Basud MHO to CNPH |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933787635815533746/5933787635815533746_signature_28.png | Ernesto A. Gapoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5933787635815533746/5933787635815533746_signature_31.png
| 2024-06-05 | 2024-06-05 13:03 | Bato Rural Health Unit | Bato, Leyte | 8 | Leyte | Laboratory | NO | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28429 | PBSP | batodiagnosticcenterlaboratory@gmail.com | 15 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 22 |  | GXIV-4D | 815267 | None | None | #NAME? | All modules are functional. | Installed by enduser, checked by Cepheid. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933740726648197709/5933740726648197709_signature_28.png | Grace C. Quijano | https://www.jotform.com/uploads/macarewecare/222927978372067/5933740726648197709/5933740726648197709_signature_31.png
| 2024-06-05 | 2024-06-05 12:18 | Paquibato District Health Center | Malabog, Paquibato Davao City | 11 | Davao del Sur | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28428 | PBSP | paquibatohclab@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 435 |  | GX-IV 4D | 838538 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Install the new version of XPC Run XPC on 4 Modules | Perform Preventive Maintenance and XpertCheck Calibration | Module B1, Retry and For Service Required Module B2:Passed Module B3, B4: Retry and Skipped due to out of spare XPC Cartidge Modules B2, B3 and B4 are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933709570319492540/5933709570319492540_signature_28.png | John Mark Mataganas | https://www.jotform.com/uploads/macarewecare/222927978372067/5933709570319492540/5933709570319492540_signature_31.png
| 2024-06-05 | 2024-06-05 11:25 | Ospital ning Capas | Capas, Talac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28427 | DOH | onclaboratory@gmail.com | 34 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1778 |  | GXIV-D | 831591 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933683732609902251/5933683732609902251_signature_28.png | Afrilyn Roque | https://www.jotform.com/uploads/macarewecare/222927978372067/5933683732609902251/5933683732609902251_signature_31.png
| 2024-06-05 | 2024-06-05 11:11 | Ciudad Medical Zamboanga | Mayor, Vitaliano Agan Ave, Zamboanga, 7000 Zamboanga del Sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | thairaphoebegrey13@gmail.com |  | 28426 | COMMERCIAL | ciudadlabmed@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 62 ASSAY: Xpert MTB/RIF Ultra, COUNT: 328 ASSAY: Xpert HCV VL, COUNT: 30 | 2024-05-23 | GX 4 (6c) | 834128 | Module 1 and Module 2 hardware failed | Replaced module 781101 with Module 210118430 | Preventive Maintenence and Module Replacement | Module 1 is available and ready to use while module 2 must be replaced. | None | 900-0508 | 98076 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933672368128025200/5933672368128025200_signature_28.png | Shereemer Cheing-Ynawat | https://www.jotform.com/uploads/macarewecare/222927978372067/5933672368128025200/5933672368128025200_signature_31.png
| 2024-06-05 | 2024-06-05 10:57 | St. Luke's Medical Center | Rizal Drive cor 32nd St., 5th Ave.Taguig | NCR | NCR 4th - Taguig | Blood Bank | NO | INSTALLATION | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 28425 | COMMERCIAL | jkcidiesca@stlukes.com.ph | 28 |  | 2024-05-23 | MCS+ | 09C125 | None for unit installation. | Performed unit installation..Performed functionality test. | Performed unit installation. | Unit is operational and ready to use. | N/A | None | DR num 11546 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933657863852569266/5933657863852569266_signature_28.png | Kenneth Idiesca | https://www.jotform.com/uploads/macarewecare/222927978372067/5933657863852569266/5933657863852569266_signature_31.png
| 2024-06-05 | 2024-06-05 10:26 | Tarlac City Health Center 3 | Brgy. Sn Miguel, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28424 | DOH | daganlea572@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 558 |  | GXIV-D | 831697 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933638762609287368/5933638762609287368_signature_28.png | Audrey Capiendo | https://www.jotform.com/uploads/macarewecare/222927978372067/5933638762609287368/5933638762609287368_signature_31.png
| 2024-06-05 | 2024-05-28 11:34 | Eastern Pangasinan District Hospital | Tayug, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28423 | PBSP | epdhlaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 454 | 2024-05-28 | GXIV-4-D-6C | 815394 | For Calibration and Preventive Maintenance | #NAME? | Performed Calibration and Preventive Maintenance | #NAME? |  | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5931163443453789340/5931163443453789340_signature_28.png | Fergyl Escaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5931163443453789340/5931163443453789340_signature_31.png
| 2024-06-05 | 2024-05-28 15:55 | Mapanas Rural Health Unit | Mapanas Northern Samar | 8 | Northern Samar | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28422 | PBSP | redendelmonte5@gmail.com | 29 |  | 2024-05-28 | GXIV - 2D | 834822 | Module B2/787222 defective | #NAME? | - Replaced module B2/787222 into 640942 - Performed preventive maintenance (Cleaned plunger rod, cartridge bay,l and Filter) - Generate reports | Unit is operational |  | GX 6c SN/640942 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933403461026136760/5933403461026136760_signature_28.png | Reden B. Delmonte | https://www.jotform.com/uploads/macarewecare/222927978372067/5933403461026136760/5933403461026136760_signature_31.png
| 2024-06-04 | 2024-06-04 11:48 | Sultan Kudarat Provincial Hospital | Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28421 | PBSP | hazyljoy_05@yahoo.com | 28.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3113 | 2024-06-04 | Gx-IV | 807405 | Module 2 multiple error: 4004 and error: 2006 | Genexpert machine Cleaning Perform annual Xpertcheck/Calibration Checking of errors Problem/errors reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules | All modules passed the Xpertcheck/Calibration | Xpertcheck CE-5 | Lot #: 1001075092 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933227467318561958/5933227467318561958_signature_28.png | Hazyl Joy Quilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5933227467318561958/5933227467318561958_signature_31.png
| 2024-06-04 | 2024-06-04 12:03 | Sultan Kudarat Provincial Hospital | Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | YES | MAINTENANCE Import additional Assay | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28420 | PBSP | hazyljoy_05@yahoo.com | 28.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 150 |  | Gx-IV | 110013958 | Module 1 for module placement UPS defective upon receiving | Genexpert machine preventive maintenance Import HIV viral load Assay Problem/errors Reporting Data collection | Perform Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933225494312727181/5933225494312727181_signature_28.png | Hazyl Joy Quilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5933225494312727181/5933225494312727181_signature_31.png
| 2024-06-04 | 2024-06-04 12:37 | San Isidro Rural Health Unit | San Isidro, Leyte | 8 | Leyte | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28419 | PBSP | kchinky26@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 404 |  | GXII | 834816 | #NAME? | -Perform preventive maintenance; plunger rod, optic slot, fan filter, cartridge bay -Archive test results -Refresh enduser on maintenance procedures | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933184379115887425/5933184379115887425_signature_28.png | Ana Cariza V. Singson | https://www.jotform.com/uploads/macarewecare/222927978372067/5933184379115887425/5933184379115887425_signature_31.png
| 2024-06-04 | 2024-06-04 17:56 | Bicol Medical Center | Concepcion Pequeña, Naga City | 5 | Camarines Sur | Bacteriology Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28418 | PBSP | bmc.bacteriology@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1464 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7398 ASSAY: Xpert SARS-CoV-2, COUNT: 1934 ASSAY: Xpert HIV-1 Viral Load, COUNT: 865 ASSAY: Others, COUNT: 182 |  | GX16-6C | 842225 | GX16 for Calibration | XpertCheck Result: A1/885370 - service required A2/702996 - new module A3/807255 - passed A4/889656 - passed B1/210068158 - passed B2/669505 - new module B3/210013422 - new module B4/645181 - passed C1/693486 - passed C2/789699 - new module C3/652013 - passed C4/790018 - new module D1/909656 - new module D2/210092932 - passed D3/210006124 - new module D4/210092857 - passed | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933059227021734506/5933059227021734506_signature_28.png | Nathalie Grace Reamico | https://www.jotform.com/uploads/macarewecare/222927978372067/5933059227021734506/5933059227021734506_signature_31.png
| 2024-06-04 | 2024-06-04 14:25 | Libmanan RHU ll | San Isidro, Camarines Sur | 5 | Albay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28417 | PBSP | amyquiapo1972@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1269 |  | GX4-6C | 805260 | GX4 for Calibration | XpertCheck Result: B1/744281 - passed B2/916366 - new module B3/693766 - passed B4/736878 - passed Cleaned cartridge bays, plunger rod, icore | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933041617022124096/5933041617022124096_signature_28.png | Amy Quiapo | https://www.jotform.com/uploads/macarewecare/222927978372067/5933041617022124096/5933041617022124096_signature_31.png
| 2024-06-04 | 2024-06-04 17:20 | Gamay Rural Health Unit | Gamay Northern Samar | 8 | Northern Samar | Laboratory | NO | INSTALLATION CALIBRATION | CEPHEID | John Paul Labutap |  | zexuzahmpaul22@gmail.com |  |  | 28416 | PBSP | teberioluz4@gmail.com | 28 |  | 2024-06-04 | GXIV | 807566 | unit for installation and Calibration | #NAME? | Install Genexpert machine - Genexpert SN/807566 - CPU SN/2UA3481GQJ - printer SN/505NTRL2F694 - UPS model RPT-1025AP | Unit is operational - Module B2 service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5933026803649179209/5933026803649179209_signature_28.png | Luz J. Teberio | https://www.jotform.com/uploads/macarewecare/222927978372067/5933026803649179209/5933026803649179209_signature_31.png
| 2024-06-04 | 2024-06-04 15:02 | St Lukes Medical Center Global City | 32nd st.,cor 5tn ave,Rizal drive,Bonifacio Global City Taguig | NCR | NCR 4th - Taguig | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 28415 | COMMERCIAL | florescanoel77@gmail.com | 28 |  |  | MCS + | 09C125 | For maintenance and calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Performed maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932939227026084207/5932939227026084207_signature_28.png | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5932939227026084207/5932939227026084207_signature_31.png
| 2024-06-04 | 2024-05-30 12:30 | Molino III City Health Unit | Astro Drive, Sansan, Molino III, Bacoor, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28414 | DOH | joanrnd@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1959 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 |  | GXIV-D-6C | 821221 | Module 3 with Temperature Error. | Module Replacement. Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked if Filter is Clogged. Replaced Module 3. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932890358318313414/5932890358318313414_signature_28.png | Jonalyn Aranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5932890358318313414/5932890358318313414_signature_31.png
| 2024-06-04 | 2024-06-03 14:26 | Quirino Mermorial Medical Center | JP Rizal, Project 4, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28413 | PBSP | carlopalomera58@gmail.com | 29 | ASSAY: Xpert CT/NG, COUNT: 21 ASSAY: Others, COUNT: 10 ASSAY: Xpert HIV-1 Viral Load, COUNT: 48 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3710 ASSAY: Xpert HPV, COUNT: 13 |  | GX IV 6C-D | 815484 | Site visit for Data to Care | Performed PM Introduced Dr. Xavier to End User | Data To care conducted interview to end user | unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932889785219334397/5932889785219334397_signature_28.png | Carlo Palomera | https://www.jotform.com/uploads/macarewecare/222927978372067/5932889785219334397/5932889785219334397_signature_31.png
| 2024-06-04 | 2024-06-04 12:51 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28412 | DOH | buhanginlaboratory@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 939 |  | GX-IV 4D | 838232 | Stuck Cartidge at Module B1 | Remove the stuck cartidge at Module B1 | Perform Module Troubleshooting | Successfully removed the Stuck Cartidge All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932879332026631474/5932879332026631474_signature_28.png | Catherine B. Deocampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5932879332026631474/5932879332026631474_signature_31.png
| 2024-06-04 | 2024-06-04 13:06 | Lianga Rural Health Unit | Purok 2, Brgy. Poblacion, Lianga | 13 | Surigao del Sur | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28411 | PBSP | gyvxs@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 780 |  | GX-2 6 colors | 835174 | Module B2 found Service Required thru XPC Calibration | Service Required Module was reported to Cepheid | Replacing Module B1 with SN: 210068498 with SN: 861624 XpertCheck Calibration of module B2 Preventive Maintenance | All modules are functional All units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932878160919094125/5932878160919094125_signature_28.png | Gyver I. Sistona | https://www.jotform.com/uploads/macarewecare/222927978372067/5932878160919094125/5932878160919094125_signature_31.png
| 2024-06-04 | 2024-06-04 10:42 | Toril Health Center | Toril, Davao City | 11 | Davao del Sur | TB-DOTS | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28410 | PBSP | leabuelis@yahoo.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1734 |  | GX-IV 4d | 838865 | Stuck cartidge on Module B2 | Remove the stuck cartidge on Module B2 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log | Perform Module Troubleshooting and Preventive Maintenance | Successfully removed the cartidge All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932878142021387374/5932878142021387374_signature_28.png | Lea Buelis RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5932878142021387374/5932878142021387374_signature_31.png
| 2024-06-04 | 2024-06-04 11:36 | Taguig-Pateros District Hospitla | East Service Road, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28409 | DOH | kristinejane.kjd25@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3667 ASSAY: Xpert HIV-1 Viral Load, COUNT: 454 | 2024-06-04 | GX-IV-D-6C | 831638 | Module A2 for replacement. | Replaced module A2 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module A2 Generated reports | Unit is operational. All modules operational. |  | Old: A2/670055 New: A2/681512 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932831804615823660/5932831804615823660_signature_28.png | Kristine Jane Base | https://www.jotform.com/uploads/macarewecare/222927978372067/5932831804615823660/5932831804615823660_signature_31.png
| 2024-06-04 | 2024-06-04 11:25 | Guimaras Provincial Health Office | Jordan Guimaras | 6 | Guimaras | TB DOTS | YES | INSTALLATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 28408 | PBSP | mylene210@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX 4 10 Colors | 120000501 | N/A | Save back up files and packed the old GeneXpert IV machine together with the accessories(UPS, Computer and Barcode Scanner) for pull-out. Installed also the MTB RIF Ultra and/MTB XDR Assays. | Installed their new GeneXpert IV 10 color machine for MTB XDR under PBSP. Installed also their new printer and other accessories( UPS, Computer and Barcode Scanner). | GeneXpert IV 10 color machine is in good condition together with the UPS, Computer, Printer and Barcode Scanner. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932814339815242637/5932814339815242637_signature_28.png | Mylene G. Publico | https://www.jotform.com/uploads/macarewecare/222927978372067/5932814339815242637/5932814339815242637_signature_31.png
| 2024-06-04 | 2024-06-04 10:05 | Gat Andres Bonifacio Memorial Medical Center | Delpan St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28407 | PBSP | laboratory.gabmmc@gmail.com | 25 |  |  | GXIV | 805103 | Unit for PM. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932773499221898097/5932773499221898097_signature_28.png | Shana Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/5932773499221898097/5932773499221898097_signature_31.png
| 2024-06-04 | 2024-05-29 13:36 | Tropical Disease Foundation | Amorsolo st., Makati City | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28406 | PBSP | aageronimo@tdf.org.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1805 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 |  | GX4 | 838134 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932698567024666910/5932698567024666910_signature_28.png | Anthony Geronimo | https://www.jotform.com/uploads/macarewecare/222927978372067/5932698567024666910/5932698567024666910_signature_31.png
| 2024-06-04 | 2024-05-29 13:31 | Tropical Disease Foundation | Amorsolo st.,Makati City | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28405 | CONSORTIUM | aageronimo@tdf.org.ph | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1801 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2024-05-27 | GX4 D | 812560 | Due for xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed xpertcheck and maintenance | C3 newly replaced C1,C2,C4 passed xpertcheck |  | Xpertcheck | 11564 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932698287024726457/5932698287024726457_signature_28.png | Anthony Geronimo | https://www.jotform.com/uploads/macarewecare/222927978372067/5932698287024726457/5932698287024726457_signature_31.png
| 2024-06-04 | 2024-05-29 09:30 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28404 | PBSP | lcpmicrolab@gmail.com | 38 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6028 ASSAY: Xpert SARS-CoV-2, COUNT: 1397 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 ASSAY: Others, COUNT: 2 |  | GX16 L | 801080 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932697917024385705/5932697917024385705_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5932697917024385705/5932697917024385705_signature_31.png
| 2024-06-04 | 2024-05-29 08:58 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28403 | COMMERCIAL | lcpmicrolab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 84 ASSAY: Xpert SARS-CoV-2, COUNT: 9 |  | GX4 L | 804205 | D2,D3,D4 disabled error 5007 last Xpertcheck 6.30.21 Defective ups | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | D1 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932697677022216175/5932697677022216175_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5932697677022216175/5932697677022216175_signature_31.png
| 2024-06-04 | 2024-05-29 08:33 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28402 | MACARE-HBDC | lcpmicrolab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 455 ASSAY: Xpert SARS-CoV-2, COUNT: 1323 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 ASSAY: Xpert HIV-1 Viral Load, COUNT: 106 ASSAY: Others, COUNT: 2 |  | GX4D | 838837 | Due for maintenance,B4 module reported already | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | 3 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932697417026404476/5932697417026404476_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5932697417026404476/5932697417026404476_signature_31.png
| 2024-06-03 | 2024-06-03 18:33 | Dalaguete City Health Office | Poblacion, Dalaguete Cebu | 7 | Cebu | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28401 | PBSP | ajreef5517@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1669 |  | GX4 | 815464 | Due for calibration and preventive maintenance of genexpert machine | Conduct preventive maintenance of genexpert machine Clean Icore, plunger, valve, cartridge bay and filter Generate IQR, System log, Assay Statistics Report and Gxx files Need additional filter for genexpert machine | Conduct preventive maintenance of genexpert machine Clean Icore, plunger, valve, cartridge bay and filter Generate IQR, System log, Assay Statistics Report and Gxx files | Module B1,B2 and B3 passed during calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932203541115222645/5932203541115222645_signature_28.png | Allyza Rivera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5932203541115222645/5932203541115222645_signature_31.png
| 2024-06-03 | 2024-06-03 18:28 | Kapatagan Rural Health Unit | Kapatagan , Lanao del Norte | 10 | Cebu | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28400 | DOH | jhayrajaculbe@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 703 |  | GX4 | 838224 | Module 4 Serial Number 807289 Hardware failure. | Calibrate module 4 Clean up ICore, plunger, valve, cartridge bay , and filter. Generate IQR, System log, GXX files, and Assay statistics | Calibrate module 4 Clean up ICore, plunger, valve, cartridge bay , and filter. Generate IQR, System log, GXX files, and Assay statistics | Module 4 failed during calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932199901113287940/5932199901113287940_signature_28.png | Jhayra Jaculbe, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5932199901113287940/5932199901113287940_signature_31.png
| 2024-06-03 | 2024-06-03 18:21 | MAYOR HILARION A RAMIRO SR MEDICAL CENTER | Maningcol, Ozamiz-Oroquieta National Highway 7200 Ozamis City Misamis Occidental · | 7 | Cebu | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28399 | PBSP | tyrosine.c@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1770 |  | Gx4 | 804765 | Calibrate module 1 Serial Number 660897 | Calibrate module 1 Serial Number 660897 Need a separate room for genexpert machine alone to maintain room temperature | Calibrate module 1 Serial Number 660897 Clean ICore, plunger, and cartridge bay, and filter | Module 1 passed during calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932196641116177952/5932196641116177952_signature_28.png | Don Felipe Jr. Saren, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5932196641116177952/5932196641116177952_signature_31.png
| 2024-06-03 | 2024-06-03 18:07 | Tangub City Health Office | Poblacion, Tangub , Misamis Occidental | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28398 | PBSP | chotangubcity@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 937 |  | GX4 | 815268 | Defective B3 # 653356 and B4 # 695611 | Replaced B3 with new module Serial # 756860 and B4 Serial # 718097 Generate IQR, System Log, Gxx files archive, and Assay Statistics Calibrate genexpert machine after module replacement | Replaced B3 with new module Serial # 756860 and B4 Serial # 718097 Generate IQR, System Log, Gxx files archive, and Assay Statistics Calibrate genexpert machine after module replacement | All modules are operational During Calibration Module B1 and B2 failed during calibration and need for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932192281118905825/5932192281118905825_signature_28.png | Vivian Malig-on | https://www.jotform.com/uploads/macarewecare/222927978372067/5932192281118905825/5932192281118905825_signature_31.png
| 2024-06-03 | 2024-06-03 14:48 | Balamban Rural Health Unit | Poblacion | 7 | Cebu | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 28397 | DOH | comclimaco.mildred@yahoo.com | 23 |  | 2024-06-03 | Gx4 | 83 | Due for XpertCheck. Ups not charging | Perform XpertCheck. | Perform XpertCheck for the GeneXpert Machine. | . Module B1, B2 passed for XpertCheck. . Module B3 Sr.# 807124, B4 Sr.# 806389 failed for XpertCheck. | Ups not charging not ok to use for GeneXpert Machine. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932069334267977711/5932069334267977711_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5932069334267977711/5932069334267977711_signature_31.png
| 2024-06-03 | 2024-05-29 10:54 | Victor R. Potenciano Medical Center | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28396 | CONSORTIUM | oro_janet@yahoo.com | 28 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 839 |  | GXIV | 835600 | Unit for maintenance | NA | #NAME? | Unit is operational and is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932057233618593664/5932057233618593664_signature_28.png | Jasmin Lardizabal | https://www.jotform.com/uploads/macarewecare/222927978372067/5932057233618593664/5932057233618593664_signature_31.png
| 2024-06-03 | 2024-05-28 14:20 | San Miguel District Hospital | San Miguel, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28395 | PBSP | smdh_laboratory@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1373 |  | GXIV | 807604 | Unit for calibration | NA | #NAME? | Unit is operational A1: passed A2: passed A3: passed A4: service required A4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932054063613731482/5932054063613731482_signature_28.png | Florencia Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5932054063613731482/5932054063613731482_signature_31.png
| 2024-06-03 | 2024-05-30 12:17 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28394 | PBSP | bmclab_micro@yahoo.com | 35 | ASSAY: Xpert HBV VL, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2361 ASSAY: Xpert HCV VL, COUNT: 26 ASSAY: Xpert HIV-1 Viral Load, COUNT: 568 | 2024-05-30 | GXIV | 805112 | Module B4 (210068588) for replacement | Replaced B4 | - replaced module B4 by 784650 - checked and cleaned cartridge bay, plunger rod, filter - generated reports - performed self test | Unit is operational Room temperature exceeds 35C Unit is not recommended for use until proper air conditioning system is installed/repaired to prevent damage on GX |  | GeneXpert Module SN: 784650 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932053743611436921/5932053743611436921_signature_28.png | Alyssa Mae Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/5932053743611436921/5932053743611436921_signature_31.png
| 2024-06-03 | 2024-05-29 14:03 | Kalinga Health Hub Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28392 | CONSORTIUM | pearl.balderama@innovationsch.org | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1288 |  | GXIV | 110008420 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932053743617328684/5932053743617328684_signature_28.png | Shad Famisaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5932053743617328684/5932053743617328684_signature_31.png
| 2024-06-03 | 2024-05-29 14:08 | Kalinga Health Hub Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28393 | CONSORTIUM | pearl.balderama@innovationsch.org | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1120 |  | GXIV | 110008415 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932053743619494614/5932053743619494614_signature_28.png | Shad Famisaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5932053743619494614/5932053743619494614_signature_31.png
| 2024-06-03 | 2024-06-03 13:53 | Davao del Norte Provincial Hospital-Kapalong Zone | Poblacion, Kapalong | 11 | Davao del Norte | TB-DOTS | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28391 | PBSP | lot2salcedo@yahoo.com | 22.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4464 |  | GX-IV 4D | 815270 | Module C3 for Replacement Due for Preventive Maintenance Due for XpertCheck Calibration | Replaced the Defective Module SN: 210062831 with new Module SN: 688263 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck on Module C1, C2, C4 | Perform Module Replacement Perform Preventive Maintenance Perform Xpertcheck Calibration | All Modules Passed the Calibration Module C3 Skipped, new Module Machine is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5932050172781678560/5932050172781678560_signature_28.png | Leona S. Olila RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5932050172781678560/5932050172781678560_signature_31.png
| 2024-06-03 | 2024-06-03 12:37 | ILOG RURAL HEALTH UNIT AND BIRTHING FACILITY | Dancalan, Ilog, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 28390 | PBSP | villavicencioemyrose@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1781 ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 |  | GX-4 | 807455 | UPS Defective - No backup power during simulated power interruption. - UPS model: 5E-1100I-USBC - UPS SN: PA64J48DE9 (GeneXpert, CPU and Monitor temporarily directly connected to wall outlet) | #NAME? | - Install 6c unit - End user training on operation, maintenance, troubleshooting. | - All 4 modules are operational - Defective UPS - No AVR |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5931990026787365904/5931990026787365904_signature_28.png | Emy Rose V. Garanganao | https://www.jotform.com/uploads/macarewecare/222927978372067/5931990026787365904/5931990026787365904_signature_31.png
| 2024-06-03 | 2024-05-30 11:28 | Kalayaan Rural Health Unit | Kalayaan, Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28388 | DOH | rhukalayaanlab@gmaol.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 247 |  | GX-IV-D-6C | 821292 | For Bios battery replacement. | Replaced CPU bios battery | Replaced CPU bios battery Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5931841193618978286/5931841193618978286_signature_28.png | Leslee Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5931841193618978286/5931841193618978286_signature_31.png
| 2024-06-03 | 2024-05-30 14:16 | Baras Rural Health Unit | J.P Rizal St. Baras, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28389 | PBSP | mesu.baras@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 650 |  | GX-IV-D-6C | 815389 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. All modules operational. Modules B1, B2 and B3 passed. Module B4 newly replaced. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5931841203612270921/5931841203612270921_signature_28.png | MICHELLE MENDOZA-AVELA | https://www.jotform.com/uploads/macarewecare/222927978372067/5931841203612270921/5931841203612270921_signature_31.png
| 2024-06-02 | 2024-05-30 11:11 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28387 | DOH | arleneabad32@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 735 |  | GXIV-4-D-6C | 838288 | For Preventive Maintenance | Performed Preventive Maintenance | For Preventive Maintenance | - Unit is partially operational - Module B1 (808218) For Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5931165453455841316/5931165453455841316_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5931165453455841316/5931165453455841316_signature_31.png
| 2024-06-02 | 2024-05-29 09:06 | Lingayen Rural Health Unit | Domalandan Ceter, Lingayen, Pang | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28386 | PBSP | michelledoriarmt@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1191 | 2024-05-29 | GXIV-4-D-6C | 834629 | For Calibration and Preventive Maintenance | #NAME? | Performed Calibration and Preventive Maintenance | - Unit is Partially Operational - Module B1 (784881) and B2 (785081) failed XpertCheck; For Reporting - Module B3 and B4 passed XpertCheck |  | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5931164573456209200/5931164573456209200_signature_28.png | Glady DG. Urbano | https://www.jotform.com/uploads/macarewecare/222927978372067/5931164573456209200/5931164573456209200_signature_31.png
| 2024-06-02 | 2024-05-28 11:34 | Eastern Pangasinan District Hospital | Tayug, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28385 | PBSP | epdhlaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 454 | 2024-05-28 | GXIV-4-D-6C | 815384 | For Calibration and Preventive Maintenance | #NAME? | Performed Calibration and Preventive Maintenance | #NAME? |  | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5931163443453789340/5931163443453789340_signature_28.png | Fergyl Escaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5931163443453789340/5931163443453789340_signature_31.png
| 2024-06-02 | 2024-05-28 08:33 | Bayambang District Hospital | Bical Norte, Bayambang, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28384 | PBSP | bevzneri@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1468 |  | GXIV-4-D-6C | 812830 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5931161453451888891/5931161453451888891_signature_28.png | Beverly Neri | https://www.jotform.com/uploads/macarewecare/222927978372067/5931161453451888891/5931161453451888891_signature_31.png
| 2024-06-02 | 2024-05-28 15:55 | Mapanas Rural Health Unit | Mapanas Northern Samar | 8 | Northern Samar | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28383 | PBSP | redendelmonte5@gmail.com | 29 |  | 2024-05-28 | GXIV - 2D | 834822 | Module B2/787222 defective | #NAME? | - Replaced module B2/787222 into 640942 - Performed preventive maintenance (Cleaned plunger rod, cartridge bay,l and Filter) - Generate reports | Unit is operational |  | GX 6c SN/640942 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5931018182518446190/5931018182518446190_signature_28.png | Reden B. Delmonte | https://www.jotform.com/uploads/macarewecare/222927978372067/5931018182518446190/5931018182518446190_signature_31.png
| 2024-06-01 | 2024-06-01 17:03 | Davao Oriental Provincial Medical Center | National Highway, Mati City | 11 | Davao Oriental | Microbiology | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28382 | PBSP | ofeliamendiola3@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1482 |  | GX-IV 10C | 110020275 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5930422803719582687/5930422803719582687_signature_28.png | Ofelia Mendiola, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5930422803719582687/5930422803719582687_signature_31.png
| 2024-05-31 | 2024-05-31 13:47 | Tampakan Rural Health Unit | Poblacion, Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28381 | DOH | lotisglegada@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 845 |  | Gx-IV | 838544 | Module 3 waiting for replacement module | Genexpert machine Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5929503424314057369/5929503424314057369_signature_28.png | Lotis G. Legada | https://www.jotform.com/uploads/macarewecare/222927978372067/5929503424314057369/5929503424314057369_signature_31.png
| 2024-05-31 | 2024-05-31 12:40 | Victorias CHO | Victorias city, negros occidental | 6 | Negros Occidental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymin899@gmail.com |  |  | 28380 | PBSP | vincent_llamoleonor@yahoo.com | 29 |  |  | Gx 4 | 838141 | Due for xpertcheck | Xpertcheck B1,B2 Service require B3,B4 passed Perform xpertcheck Clean fan filter | Xpertcheck B1,B2 Service require B3,B4 passed | Need for module replacement B1 and B2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5929408260515736125/5929408260515736125_signature_28.png | Vincent Leonor | https://www.jotform.com/uploads/macarewecare/222927978372067/5929408260515736125/5929408260515736125_signature_31.png
| 2024-05-31 | 2024-05-31 11:57 | MATI CITY HEALTH OFFICE | Mati Davao Oriental | 11 | Davao Oriental | LABORATORY | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28379 | PBSP | jaslimjay@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1020 |  | GX-II | 835221 | FOR MODULE B1 REPLACEMENT | #NAME? | -PERFORM CLEANING OF GENEXPERT EQUIPMENT AND WORKSTATION -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE B1 REPLACEMENT: DEF SN: 210068172 WITH NEW SN: 683668 -PERFORM MODULE B2 CALIBRATION, PASSED | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5929377462966231198/5929377462966231198_signature_28.png | JASPER JAY LIM | https://www.jotform.com/uploads/macarewecare/222927978372067/5929377462966231198/5929377462966231198_signature_31.png
| 2024-05-31 | 2024-05-31 11:32 | Surigao del Norte Provincial Hospital | National Highway, Sta. Cruz, Bad-as, Placer | 13 | Surigao del Norte | RTDL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28378 | PBSP | biyoyojayl@gmail.com | 29 |  |  | GX-4 6 colors | 815435 | B2 found temperature error | Reported to Cepheid | Preventive Maintenance Troubleshooting Data Collection | 3 modules working, 1 modules for replacment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5929357939428067113/5929357939428067113_signature_28.png | Jay Biyoyo | https://www.jotform.com/uploads/macarewecare/222927978372067/5929357939428067113/5929357939428067113_signature_31.png
| 2024-05-31 | 2024-05-31 11:10 | Orani Rural Health Unit | Orani Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28377 | DOH | oranirhu@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV-D | 831668 | A3:757003 for module replacement | Module replacement Preventive maintenance | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational A3:635093 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5929353921011895733/5929353921011895733_signature_28.png | Belinda Guinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5929353921011895733/5929353921011895733_signature_31.png
| 2024-05-31 | 2024-05-31 08:57 | San Remigio Rural Health Unit | Poblacion | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 28376 | WHO | rodacanetecamay@gmail.com | 23 |  | 2024-05-31 | Gx4 | 847333 | Preventive maintenance. UPS Defective date Septem 26,2023 | Perform preventive maintenance. | . Clean plunger rod. . Clean filter . Save System Log file . Save Installation Qualification Report | . All modules are functional ok to use . GeneXpert Machine not connected in UPS directly to 220V outlet. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5929278419614027459/5929278419614027459_signature_28.png | Roda Camay | https://www.jotform.com/uploads/macarewecare/222927978372067/5929278419614027459/5929278419614027459_signature_31.png
| 2024-05-30 | 2024-05-30 18:58 | Tondo Medical Center | Honorio Lopez Blvd, Tondo, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28375 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 393 ASSAY: Xpert SARS-CoV-2, COUNT: 47 |  | GX XVI 6C-D | 844121 | CPU not turning on | Troubleshoot CPU Troubleshoot GXIV Performed PM | Checked CPU CPU beeps Tightened RAM Tightened SSD Problem persists Swapped GX XVI and GX IV's position temporarily 1 GX IV with power but not detected by software Tightened gateway board and power supply Gx IV detected, problem solved Performed PM Generated report Generated report | GX units working but 1 CPU for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928764101712007559/5928764101712007559_signature_28.png | Analyn N.Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5928764101712007559/5928764101712007559_signature_31.png
| 2024-05-30 | 2024-05-30 15:46 | M. Icasiano Health Center | Plaza Hugo St., Sta. Ana, Manila | NCR | NCR 1st - Manila | laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28374 | DOH | fabroslorrainelucila@gmail.com | 25 |  |  | GX IV 6C-D | 831684 | Unit due for maintenance | Performed Preventive maintenance Performed reliability test | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Reliability test All modules passed Generated report | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928663388125204200/5928663388125204200_signature_28.png | Lorraine Lucila L.Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5928663388125204200/5928663388125204200_signature_31.png
| 2024-05-30 | 2024-05-30 16:02 | Southern Philippines Medical Center | JP Laurel, Davao City | 11 | Davao del Sur | ICU2 | NO | TROUBLESHOOTING | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28373 | COMMERCIAL | icu2estuya@gmail.com | 25.8 |  |  | MCS+ | 15H034G | Centrifuge related problem | Reset connections Flush the centrifuge to check if there is blockage | Perform Machine Troubleshooting Perform board isolation and power supply isolation | For Centrifuge problem isolation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928662078416356857/5928662078416356857_signature_28.png | Dianne Karla de Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5928662078416356857/5928662078416356857_signature_31.png
| 2024-05-30 | 2024-05-30 15:07 | Pedro Gil Health Center | Plaza Hugo St., Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28372 | PBSP | fabroslorrainelucila@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 846 | 2024-05-24 | GX IV 6C-D | 812829 | Modules B2, B3 and B4 for Replacement Unit due for calibration | Perfoemed PM Replaced Modules B2: 210100154 -> 210039590 B3: 885647 -> 210062452 B4: 705885 -> 210062519 Performed Xpertcheck on B1 | Checked and cleaned cartridge bays, plunger rod, icore slot, and filter Replaced modules B2, B3 and B4 Performed Xpertcheck on B1 Generated report | Unit is operational All modules working Filter for replacement |  | 3 GX Modules 6C XpertCheck Kit | PL 0158861 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928643778128048715/5928643778128048715_signature_28.png | Lorraine Lucila L.Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5928643778128048715/5928643778128048715_signature_31.png
| 2024-05-30 | 2024-05-30 13:47 | Albuera Rural Health Unit | Albuera, Leyte | 8 | Leyte | TB-Dots | YES | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28371 | PBSP | mariaana.barte@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 349 |  | GXII | 834818 | Due for XpertCheck | #NAME? | #NAME? | All modules passed calibration and are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928615851132549600/5928615851132549600_signature_28.png | Maria Ana S. Barte | https://www.jotform.com/uploads/macarewecare/222927978372067/5928615851132549600/5928615851132549600_signature_31.png
| 2024-05-30 | 2024-05-30 14:44 | Dinagat District Hospital | Brgy. Escolta, Dinagat, Dinagat Islands | 13 | Dinagat Islands | RTDL | NO | INSTALLATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28370 | PBSP | idhona@yahoo.com | 25 |  |  | GX-4 | 110013918 | None | None | Installation of New GeneXpert 4 Modules 10 Colors and peripherals. Dissambling Old GeneXpert 4 6 Colors and Repacked | All modules are functional. CPU and monitor are working Printer is working Barcode Scanner is working UPS not working. (Using old UPS) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928611566698265786/5928611566698265786_signature_28.png | Idhona N. Tillo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5928611566698265786/5928611566698265786_signature_31.png
| 2024-05-30 | 2024-05-30 14:23 | Silay City Health Office | Silay City, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28369 | PBSP | gutierrezdaviddior@gmail.com | 29 |  |  | Gx 4 | 805093 | Defective module A1 606944 | New module installed Sn 210006956 New installation location Silay City Health Office Gathering data IQR, System log Clean Fan Filter and Fan | New module installed Sn 210006956 New installation location Silay City Health Office Gathering data IQR, System log Clean Fan Filter and Fan | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928597325834393933/5928597325834393933_signature_28.png | David dior gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5928597325834393933/5928597325834393933_signature_31.png
| 2024-05-30 | 2024-05-30 12:51 | Caloocan City Health Office | A. Mabini St., Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28368 | PBSP | caloocancentrallaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2251 ASSAY: Xpert HIV-1 Viral Load, COUNT: 141 ASSAY: Xpert HIV-1 Viral Load, COUNT: 335 | 2024-05-24 | GX IV 6C-D | 805271, 814154 | Unite due for calibration | Performed PM Performed Xpertcheck | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck Modules A2, A3 and B1 Service required B2 lost comm, lack of cartridge Generated report | Unit is operational A1: Passed A2: Service required A3: Service required A4: New Module B1: Service required B2: Lost comm needs XpertCheck kit B3: Passed B4: Passed |  | XpertCheck kit | PL 015858 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928543869325025117/5928543869325025117_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5928543869325025117/5928543869325025117_signature_31.png
| 2024-05-30 | 2024-05-30 12:13 | Caloocan City Health Office 10C | A. Mabini St., Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28367 | PBSP | caloocancentrallaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-05-24 | GX IV 10C-D | 120000489 | 10C unit due for Installation 6C unit for packing | Setup and installed 10C machine Packed 6C machine | Unpacked 10C machine Setup GX and accessories Changed time Setup user admin Setup system config Packed 6C unit Generated report | Unit is operational All modules working |  | GX 10C Intex UPS HP Printer | DR 48548 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928525169327907524/5928525169327907524_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5928525169327907524/5928525169327907524_signature_31.png
| 2024-05-30 | 2024-05-30 10:03 | Corazon Locsin Montelibano M.R.H | Lacson St., Bacolod City | 6 | Negros Occidental | TB DOTS | YES | TROUBLESHOOTING | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 28366 | DOH | gallegojennybelle@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1149 ASSAY: Xpert HIV-1 Viral Load, COUNT: 325 |  | GX4 | 821219 | Their concerns is not with the GeneXpert machine but with their UPS(Uninterrupted Power Supply) battery cannot sustain anymore supply of power during brown out or fluctuation of electricity in the area. | Conduct PM and save files. | Conduct troubleshooting and suggested to request for a higher capacity of UPS and AVR to 2KVA or higher. | GeneXpert Machine modules are in good conditions. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928451435614423076/5928451435614423076_signature_28.png | Jennybell Gallego | https://www.jotform.com/uploads/macarewecare/222927978372067/5928451435614423076/5928451435614423076_signature_31.png
| 2024-05-30 | 2024-05-30 10:10 | SACCL | Quiricada St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | ciudadanosophiapatrice@gmail.com |  | 28365 | MACARE-HBDC | robethsonfp@gmail.com | 28 |  |  | GXIV-L-6C | 821257 | Modules have 5007 Errors. | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Performed Reliability Test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928444353848178303/5928444353848178303_signature_28.png | Robethson Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5928444353848178303/5928444353848178303_signature_31.png
| 2024-05-29 | 2024-05-27 13:34 | Liloan RHU | Poblacion Liloan, Cebu | 7 | Cebu | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28364 | PBSP | vickcabahugmt@gmail.com | 35 |  |  | Gx4 | 835189 | Damaged Fan Filter Room area cold temp is not enough for genexpert machine | Clean fan filter Plunger rod maintenance Gathering data of genexpert machine | Clean fan filter Plunger rod maintenance Gathering data of genexpert machine | Systems operational/rooam area cold temperature is not enough |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5928051140815575587/5928051140815575587_signature_28.png | Vick P. Cabahug | https://www.jotform.com/uploads/macarewecare/222927978372067/5928051140815575587/5928051140815575587_signature_31.png
| 2024-05-29 | 2024-05-29 16:34 | Mabalacat Rural Health Unit | Old Municipal Hall Mac Arthur Highway, Poblacion Mabalacat City Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28363 | PBSP | perezriyadh@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1138 ASSAY: Xpert HIV-1 Viral Load, COUNT: 49 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GXII-D | 835104 | B2:715593 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational B2:635095 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927811803694878316/5927811803694878316_signature_28.png | Riyadh Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5927811803694878316/5927811803694878316_signature_31.png
| 2024-05-29 | 2024-05-29 16:05 | Dr. Rafael S. Tumbokon Memorial Hospital | Kalebo Aklan | 10 | Aklan | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 28362 | PBSP | delfinteodia@yahoo.com | 23 |  | 2024-05-29 | GX4 10C | 110020390 | New Installation for 10c GeneXpert Machine | Install GeneXpert System. | . Install GeneXpert Machine. . Install CPU . Install LCD Monitor . Install HP printer | All are functional ok to use. | New UPS not charging not ok to use. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927801052969061099/5927801052969061099_signature_28.png | Teodia P. Delfin | https://www.jotform.com/uploads/macarewecare/222927978372067/5927801052969061099/5927801052969061099_signature_31.png
| 2024-05-29 | 2024-05-29 16:05 | Caraga TB Reference Laboratory | Rizal St. Brgy. Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28361 | PBSP | abpabellan.pbsp@gmail.com | 27 |  |  | GX-4 10 Colors | 110016959 | None | None | Preventive Maintenance. XpertCheck Calibration | All modules and units and functional All modules passed in XPC Calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927793351215595839/5927793351215595839_signature_28.png | Andrew B. Pabellan | https://www.jotform.com/uploads/macarewecare/222927978372067/5927793351215595839/5927793351215595839_signature_31.png
| 2024-05-29 | 2024-05-29 14:32 | Pagsanjan Rural Health Unit | Pagsanjan, Laguna | 4A | Laguna | Laboratory | NO | For packing | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28360 | PBSP | dotspagsanjan@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 545 ASSAY: Xpert HIV-1 Viral Load, COUNT: 174 |  | GX-IV-D-6C | 807568 | GeneXpert system for packing | Packed the GeneXpert System | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports Packed the GeneXpert system: -GeneXpert unit -Workstation -Printer -Barcode scanner with stand -UPS | Unit is operational. UPS operational but has no backup power, recommend to replace. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927775913487989562/5927775913487989562_signature_28.png | Lei Arnie Abrigo | https://www.jotform.com/uploads/macarewecare/222927978372067/5927775913487989562/5927775913487989562_signature_31.png
| 2024-05-29 | 2024-05-29 14:37 | Pagsanjan Rural Health Unit | Pagsanjan, Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28359 | PBSP | iamleiarnie@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 229 ASSAY: Xpert HIV-1 Viral Load, COUNT: 49 |  | GX-IV-D-10C | 110013962 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927775903485371358/5927775903485371358_signature_28.png | Lei Arnie Abrigo | https://www.jotform.com/uploads/macarewecare/222927978372067/5927775903485371358/5927775903485371358_signature_31.png
| 2024-05-29 | 2024-05-29 14:18 | Region II Trauma and Medical Center | Bayombong, Nueva Vizcaya | 2 | Nueva Vizcaya | Bacteriology | NO | MAINTENANCE Disassemble | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28358 | PBSP | cecillepabance@gmail.com | 27 |  |  | GXIV-6c | 807458 | For Pullout | Disassemble Genexpert Packed Genexpert | Performed Maintenance Gathered Reports Performed Funtionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927759400526549771/5927759400526549771_signature_28.png | CECILLE A. HERMOSO | https://www.jotform.com/uploads/macarewecare/222927978372067/5927759400526549771/5927759400526549771_signature_31.png
| 2024-05-29 | 2024-05-29 14:50 | Region II Trauma and Medical Center | Bayombong, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28357 | PBSP | cecillepabance@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 880 ASSAY: Xpert SARS-CoV-2, COUNT: 76 |  | GXIV-10c | 120000537 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927748840525610883/5927748840525610883_signature_28.png | CECILLE A. HERMOSO | https://www.jotform.com/uploads/macarewecare/222927978372067/5927748840525610883/5927748840525610883_signature_31.png
| 2024-05-29 | 2024-05-29 14:37 | Region II Trauma and Medical Center | Bayombong, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28356 | PBSP | cecillepabance@gmail.com | 27 |  |  | GXIV-6c | 834621 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927741080521409278/5927741080521409278_signature_28.png | CECILLE A. HERMOSO | https://www.jotform.com/uploads/macarewecare/222927978372067/5927741080521409278/5927741080521409278_signature_31.png
| 2024-05-29 | 2024-05-29 11:49 | Sta. Cruz Rural Health Unit | Pio Valenzuela St. Sta. Cruz, Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28355 | DOH | jezzaleal@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1462 | 2024-05-29 | GX-IV-D-6C | 820229 | Module D2 for replacement | Replaced module D2 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module D2 Generated reports | Unit is operational. All modules operational. |  | Old: D2/722095 New: D2/883154 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927724869326118407/5927724869326118407_signature_28.png | Jezza N. Van patten | https://www.jotform.com/uploads/macarewecare/222927978372067/5927724869326118407/5927724869326118407_signature_31.png
| 2024-05-29 | 2024-05-29 12:28 | A.H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28354 | PBSP | dinajaen1990@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4140 |  | GX XVI | 839292 | Unit for calibration | NA | #NAME? | Unit is operational A1: passed A2: passed A3: passed A4: service required B1: passed B2: passed B3: passed B4: passed C1: passed C2: passed C3: service required C4: passed D1: passed D2: service required D3: passed D4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927682955815106113/5927682955815106113_signature_28.png | Dina Jaen | https://www.jotform.com/uploads/macarewecare/222927978372067/5927682955815106113/5927682955815106113_signature_31.png
| 2024-05-29 | 2024-05-29 12:32 | A.H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28353 | PBSP | dinajaen1990@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1109 |  | GXIV | 805067 | Unit for calibration | NA | #NAME? | Unit is operational A1: service required A2: passed A3: passed A4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927682955815961355/5927682955815961355_signature_28.png | Dina Jaen | https://www.jotform.com/uploads/macarewecare/222927978372067/5927682955815961355/5927682955815961355_signature_31.png
| 2024-05-29 | 2024-05-29 12:11 | Rosario Rural Health Unit | Rosario Batangas | 4A | Batangas | Laboratort | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28352 | DOH | liberty_yutuc@yahoo.com | 25 |  | 2024-05-29 | GXIV | 821421 | Unit for PM. For module replacement of module A3. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A3 (797805) in replacement to defective A3 (783674). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927667036517431234/5927667036517431234_signature_28.png | Liberty Ilagan Yutuc | https://www.jotform.com/uploads/macarewecare/222927978372067/5927667036517431234/5927667036517431234_signature_31.png
| 2024-05-29 | 2024-05-29 12:11 | Rosario Rural Health Unit | Rosario Batangas | 4A | Batangas | Laboratort | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28351 | DOH | liberty_yutuc@yahoo.com | 25 |  | 2024-05-29 | GXIV | 821421 | Unit for PM. For module replacement of module A3. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A3 (797805) in replacement to defective A3 (783674). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927667036516912737/5927667036516912737_signature_28.png | Liberty Ilagan Yutuc | https://www.jotform.com/uploads/macarewecare/222927978372067/5927667036516912737/5927667036516912737_signature_31.png
| 2024-05-29 | 2024-05-29 12:08 | Gov. Benjamin T. Romualdez General Hospital and Schistosomiasis Center | Palo, Leyte | 8 | Leyte | TB-Dots | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28350 | PBSP | mbesplago.pbsp@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2595 | 2024-04-26 | GXIV R1 | 707521 | #NAME? | -Replaced defective module B4/615648 with module SN: 210035570 -Run XpertCheck on all three modules | #NAME? | All modules are functional. Modules passed XpertCheck. |  | 900-0491RS GeneXpert 6-Color Module | 47670 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927655173485513452/5927655173485513452_signature_28.png | Maria Chrissa Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5927655173485513452/5927655173485513452_signature_31.png
| 2024-05-29 | 2024-05-29 11:42 | Quirino Mermorial Medical Center | JP Rizal, Project 4, Quezon City | NCR | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28349 | PBSP | carlopalomera58@gmail.com | 29 | ASSAY: Xpert CT/NG, COUNT: 22 ASSAY: Others, COUNT: 10 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3705 ASSAY: Xpert HIV-1 Viral Load, COUNT: 49 |  | GX IV 6C-D | 815484 | Unit due for maintenance | Performed Preventive maintenance Performed Reliability test | Checked and cleaned cartridge bays, plunger rod, icore slot, and filter Performed reliability test module B3 many errors, for reporting Disabled module B3 Generated report | Unit is operational Module B3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927643312124050316/5927643312124050316_signature_28.png | carlo palomera | https://www.jotform.com/uploads/macarewecare/222927978372067/5927643312124050316/5927643312124050316_signature_31.png
| 2024-05-29 | 2024-05-29 11:43 | Oras District Hospital | Oras Eastern Samar | 8 | Eastern Samar | TB-DOTS | YES | Disassemble | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28348 | PBSP | espesopriscilla@gmail.com | 28 |  | 2024-05-29 | GXIV - 2D | 835081 | Unit for pullout | #NAME? | - Disassembled unit Genexpert SN/835081 CPU SN/CPH119F04303 LCD Monitor SN/CNK92102RD Printer SN/VNC7703319 Barcode scanner Keyboard/Mouse | Unit for module replacement B1 and B2 | Unit is ready to pullout |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927638371025083653/5927638371025083653_signature_28.png | Priscilla Espeso | https://www.jotform.com/uploads/macarewecare/222927978372067/5927638371025083653/5927638371025083653_signature_31.png
| 2024-05-29 | 2024-05-29 11:34 | Oras District Hospital | Oras Eastern Samar | 8 | Eastern Samar | TB-DOTS | YES | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28347 | PBSP | espesopriscilla@gmail.com | 27 |  | 2024-05-29 | GXIV | 804256 | unit for installation | Installation | -Install Genexpert unit SN/804256 - CPU SN/2UA3450YRR - Monitor SN/308NDPHJL568 -UPS SN/4018631904 - Printer SN/VN6N47660 - Generate reports | Unit is operational |  | Genexpert machine Monitor Printer | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927632421022263381/5927632421022263381_signature_28.png | Priscilla Espeso | https://www.jotform.com/uploads/macarewecare/222927978372067/5927632421022263381/5927632421022263381_signature_31.png
| 2024-05-29 | 2024-05-29 11:30 | Minalin Rural Health Unit | Minalin Pamanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28346 | PBSP | bernadetteann0318@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 466 |  | GXIV-D | 834688 | Due for preventive maintenance UPS long beeing sound problem | Preventive maintenance UPS check-u and troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed UPS troubleshooting Checked for faulty fuse Performed functionality test Gathered reports | Unit is operational UPS is operational ( for observation) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927630668882218545/5927630668882218545_signature_28.png | Rizzie Niturada | https://www.jotform.com/uploads/macarewecare/222927978372067/5927630668882218545/5927630668882218545_signature_31.png
| 2024-05-29 | 2024-05-28 11:09 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28345 | DOH | wcpatron11x4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1815 | 2024-05-28 | GXIV | 831534 | Unit for preventive maintenance. For module replacement of module.A2. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A2 (616338) in replacement to defective module A2 (711157). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927616916511784464/5927616916511784464_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5927616916511784464/5927616916511784464_signature_31.png
| 2024-05-29 | 2024-05-29 10:46 | Batangas Medical Center | Kumintang Ibaba, Batangas | 4A | Batangas | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28344 | PBSP | batmc.tbculturelab@gmail.com | 25 |  | 2024-05-29 | GXIV | 814153 | Unit for PM. For installation of new CPU. | PM and installation of CPU. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed installation of CPU. Generated reports. | Unit is operational. |  | Dell CPU (CPH1023F46024) in replacement to defective Lenovo CPU (CPH0622F38432). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927616916514321001/5927616916514321001_signature_28.png | Marienan E. Marasigan | https://www.jotform.com/uploads/macarewecare/222927978372067/5927616916514321001/5927616916514321001_signature_31.png
| 2024-05-29 | 2024-05-29 10:37 | Batangas Medical Center | Kumintang Ibaba, Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28343 | PBSP | batmc.tbculturelab@gmail.com | 25 |  | 2024-05-29 | GX16 | 844120 | Unit for PM. For module replacement of module A1 and C2. For checking of monitor. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. For reporting of defective monitor HP P19A (CNK94309WP). |  | GX module A1 (668724) and C2 (685644) in replacement to defective A1 (644387) and C2 (892210). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927616916517381718/5927616916517381718_signature_28.png | Marienan E. Marasigan | https://www.jotform.com/uploads/macarewecare/222927978372067/5927616916517381718/5927616916517381718_signature_31.png
| 2024-05-29 | 2024-05-29 10:57 | Southern Isabela Medical Center | Zamora, Santiago City, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE Disassemble | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28342 | PBSP | simcmicrobiology@yahoo.com | 28 |  |  | GXIV-6c | 805104 | For Disassemble | Disassembled Genexpert System Packed Genexpert System | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Unit is ready for pickup |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927610351145237180/5927610351145237180_signature_28.png | Jannell T.Valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/5927610351145237180/5927610351145237180_signature_31.png
| 2024-05-29 | 2024-05-29 10:54 | Southern Isabela Medical Center | Zamora, Santiago City, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28341 | PBSP | simcmicrobiology@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 342 |  | GXIV-10c | 120000738 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927610211141951677/5927610211141951677_signature_28.png | Jannell T.Valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/5927610211141951677/5927610211141951677_signature_31.png
| 2024-05-29 | 2024-05-29 10:13 | Cotabato Regional and Medical Center | Sinsuat Ave, Rosary Heights Cotabato City | BARMM | Maguindanao | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28340 | PBSP | crmctblab@gmail.com | 25.8 |  |  | GX-IV 4D | 807607 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Perform Xpertcheck Calibration on Modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules are operational All Modules Passed the Calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927582445748256335/5927582445748256335_signature_28.png | Velchie Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5927582445748256335/5927582445748256335_signature_31.png
| 2024-05-29 | 2024-05-28 18:12 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | crmctblab@gmail.com |  | 28339 | PBSP | tindocvelche@gmail.com | 24.6 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 ASSAY: Xpert HPV, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 126 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 | 2024-05-28 | Gx-XVI | 839307 | Module A1 and A2 for Service Required | Replace defective module 2 SN:870832 with new module SN: 909867 Replace defective module 2 SN:870834 with new module SN: 904969 Genexpert machine Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine module replacement/repair | 16 operational modules |  | Genexpert machine 6-color SN: 909867 and SN: 904969 | Cepheid case: 1498439 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927580967917007889/5927580967917007889_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5927580967917007889/5927580967917007889_signature_31.png
| 2024-05-29 | 2024-05-28 17:00 | Cotabato Regional and Medical Center | Sinsuat Ave, Rosary Heights Cotabato City | BARMM | Maguindanao | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28338 | PBSP | crmctblab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 ASSAY: Xpert HPV, COUNT: 1 |  | GX-IV 4D | 807463 | Due for Preventive Maintenance and XpertCheck Calibration Error on Module D1 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck on Modules Reset network connections | Perform Preventive Maintenance and XpertCheck Calibration Perform Module Troubleshooting | Module D3 For Service Required Modules D1, D2, D4 passed the Calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927576855746072930/5927576855746072930_signature_28.png | Velchie Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5927576855746072930/5927576855746072930_signature_31.png
| 2024-05-29 | 2024-05-28 16:57 | Cotabato Regional and Medical Center | Sinsuat Ave, Rosary Heights Cotabato City | BARMM | Maguindanao | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28337 | PBSP | crmctblab@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 9 ASSAY: Others, COUNT: 230 |  | GX-IV 10 4D | 110016958 | For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform Xpertcheck Calibration | Run XpertCheck and Perform Preventive Maintenance | All Modules passed the Calibration All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927576045745779981/5927576045745779981_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5927576045745779981/5927576045745779981_signature_31.png
| 2024-05-28 | 2024-05-24 16:23 | Quezon Medical Center | Provincial Health Office, Barangay XI, Lucena City | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28336 | PBSP | qmcrtdl2015@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 747 ASSAY: Xpert HPV, COUNT: 89 | 2024-04-29 | GXIV 10C - D | 110019627 | Unit due for calibration | Performed Preventive maintenance Performed Xpertcheck | Checked and cleaned cartridge bays, plunger rod icore slot and filter Performed Xpertcheck on modules B1, B2 and B3. B4 newly replaced All modules passed Generated report | Unit is operational all modules working |  | XpertCheck kit | DR 47700 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927131193683894566/5927131193683894566_signature_28.png | JoAnn P. Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5927131193683894566/5927131193683894566_signature_31.png
| 2024-05-28 | 2024-05-24 16:20 | Quezon Medical Center 10C | Provincial Health Office, Barangay XI, Lucena City | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28335 | PBSP | qmcrtdl2015@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 849 |  | 804692 | 804692 | Module A1 stuck cartridge | Removed stuck cartridge Checked and cleaned cartridge bays, plunger rod, icore slot and filter Installed UPS Generated report | Removed stuck cartridge Performed Preventive maintenance Installed UPS | unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927127623686326428/5927127623686326428_signature_28.png | JoAnn P. Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5927127623686326428/5927127623686326428_signature_31.png
| 2024-05-28 | 2024-05-28 17:42 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28334 | PBSP | crmcmolecularlab@gmail.com | 23.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 12986 ASSAY: Others, COUNT: 356 | 2024-05-28 | Gx-IV | 834441 | Module 2 for Service Required | Replace defective module 2 SN:210100872 with new module SN: 802585 Genexpert machine Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine module replacement/repair | 4 operational modules |  | Genexpert module 6-color SN 210100872 | Cepheid case : 1532113 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927001365218730308/5927001365218730308_signature_28.png | Darianne Mae Tubilan | https://www.jotform.com/uploads/macarewecare/222927978372067/5927001365218730308/5927001365218730308_signature_31.png
| 2024-05-28 | 2024-05-28 17:51 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28333 | PBSP | crmcmolecularlab@gmail.com | 23.1 | ASSAY: Xpert SARS-CoV-2, COUNT: 12986 | 2024-05-28 | Gx-IV | 834442 | Module 2 H/W failed | Replace defective module 2 SN:867625 with new module SN: 210052543 Genexpert machine Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine module replacement/repair | 4 operational modules |  | Genexpert module 6-color SN:210052543 | Cepheid case: 1499268 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5927000245214259930/5927000245214259930_signature_28.png | Darianne Mae Tubilan | https://www.jotform.com/uploads/macarewecare/222927978372067/5927000245214259930/5927000245214259930_signature_31.png
| 2024-05-28 | 2024-05-28 17:55 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28332 | DOH | crmcmolecularlab@gmail.com | 22.8 | ASSAY: Xpert SARS-CoV-2, COUNT: 12986 |  | Gx-IV | 837662 | Module 2 and 4 multiple errors 1002 | Genexpert machine Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926999585211117502/5926999585211117502_signature_28.png | Darianne Mae Tubilan | https://www.jotform.com/uploads/macarewecare/222927978372067/5926999585211117502/5926999585211117502_signature_31.png
| 2024-05-28 | 2024-05-28 17:55 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28331 | DOH | crmcmolecularlab@gmail.com | 22.8 | ASSAY: Xpert SARS-CoV-2, COUNT: 12986 |  | Gx-IV | 837662 | Module 2 and 4 multiple errors 1002 | Genexpert machine Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926999475215089660/5926999475215089660_signature_28.png | Darianne Mae Tubilan | https://www.jotform.com/uploads/macarewecare/222927978372067/5926999475215089660/5926999475215089660_signature_31.png
| 2024-05-28 | 2024-05-28 15:45 | Philippine Red Cross QC chapter | Kalayaan ave QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28330 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 26 |  | 2024-05-28 | MCS + | 02A227 | CRC bios failed | Delivered and replaced CPU board Checked Calibration parameters Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running Pull out defective CPU board | Repair | Operational |  | CPU board | 11566 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926961726347139069/5926961726347139069_signature_28.png | Philippine Red Cross Qiezon City Chapter | https://www.jotform.com/uploads/macarewecare/222927978372067/5926961726347139069/5926961726347139069_signature_31.png
| 2024-05-28 | 2024-05-28 10:19 | Quezon City Jail | Payatas QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28329 | PBSP | carleecruz0@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1400 |  | GX4D | 803089 | GXDX unable to boot sql error display | Reinstall gxdx 4.8 version Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Troubleshooting and maintenance | Unit operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926961396342477649/5926961396342477649_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5926961396342477649/5926961396342477649_signature_31.png
| 2024-05-28 | 2024-05-27 15:52 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28328 | MACARE-HBDC | bloodbankphc@gmail.com | 28 |  |  | MCS + | 20168 | As per end iser they encounter dpm out of tolerance | Checked unit Cheked and test DPM Chekecd programs parameters Adjust DPM value Performed functionality test Test run unit | Troubleshooting | Operational After performed simulation test no error encountered |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926960946345464093/5926960946345464093_signature_28.png | Jeric Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5926960946345464093/5926960946345464093_signature_31.png
| 2024-05-28 | 2024-05-27 15:20 | Philippine Heart Center | East Aveneue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28327 | MACARE-HBDC | bloodbankphc@gmail.com | 26 |  |  | MCS+ | 11H204 | As per end user they encounter DPM is out of tolerance | Checked unit Cheked and test DPM Chekecd programs parameters Adjust DPM value Performed functionality test Test run unit | Troubleshooting | Operational After performing simulation run no error encountered |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926960306348961210/5926960306348961210_signature_28.png | Jeric Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5926960306348961210/5926960306348961210_signature_31.png
| 2024-05-28 | 2024-05-28 16:18 | Samboan Rural Health Unit | Poblacion, Samboan Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28326 | DOH | sheneannesenagan@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 452 |  | GX4 | 838513 | Due for Preventive Maintenance | Conduct preventive maintenance; Clean up Icore, Plunger, cartridge bay , and filter Generate IQR, GXX files, System log, Assay Statistics Check up UPS and AVR | Conduct preventive maintenance; Clean up Icore, Plunger, cartridge bay , and filter Generate IQR, GXX files, System log, Assay Statistics Check up UPS and AVR | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926937574345880761/5926937574345880761_signature_28.png | Shene Anne A. Senagan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5926937574345880761/5926937574345880761_signature_31.png
| 2024-05-28 | 2024-05-28 15:35 | Santiago City Health Office | San Andres, Santiago City, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28325 | DOH | cholabstgo@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2378 |  | GXIV-6c | 831718 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926914126993950103/5926914126993950103_signature_28.png | Marichelle Zapanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5926914126993950103/5926914126993950103_signature_31.png
| 2024-05-28 | 2024-05-28 14:17 | The Medical City | Ortigas Avenue, Pasig | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28324 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 21 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926867782512470461/5926867782512470461_signature_28.png | analyn Ponferrada | https://www.jotform.com/uploads/macarewecare/222927978372067/5926867782512470461/5926867782512470461_signature_31.png
| 2024-05-28 | 2024-05-28 13:35 | The Medical City | Ortigas Avenue, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28323 | COMMERCIAL | pedelacruz@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Simulation run Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926843002511216759/5926843002511216759_signature_28.png | Perlita De La Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5926843002511216759/5926843002511216759_signature_31.png
| 2024-05-28 | 2024-05-28 13:28 | Quirino Province Medical Center | Mangandingay, Cabarroguis, Quirino | 2 | Quirino | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Packing | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28322 | PBSP | ylbumidang0817@gmail.com | 28 |  |  | GXIV-6c | 814166 | Cpu not powering on | Troubleshoot wirings Troubleshoot bios battery | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is not Operational | CPU for Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926839656518038007/5926839656518038007_signature_28.png | Yvone Bumidang | https://www.jotform.com/uploads/macarewecare/222927978372067/5926839656518038007/5926839656518038007_signature_31.png
| 2024-05-28 | 2024-05-28 13:33 | Quirino Province Medical Center | Mangandingay, Cabarroguis, Quirino | 2 | Quirino | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28321 | PBSP | ylbumidang0817@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 170 |  | GXIV-10c | 120000609 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926839531814427912/5926839531814427912_signature_28.png | Yvone Bumidang | https://www.jotform.com/uploads/macarewecare/222927978372067/5926839531814427912/5926839531814427912_signature_31.png
| 2024-05-28 | 2024-05-28 12:35 | Iba Rural Health Unit | Iba Zambales | 3 | Zambales | Laboratory | NO | Machine Disassembly and Packing | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28320 | PBSP | garadoshanna@gmail.com | 34 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2142 ASSAY: Xpert HIV-1 Viral Load, COUNT: 83 |  | GXIV-D | 805070 | Unit for disassembly and packing | Preventive maintenance Disassembly and packing | Performed preventive maintenance Performed functionality test Gathered reports Performed machine Disassembly and Packing | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926827273785219418/5926827273785219418_signature_28.png | Hanna C. Garados | https://www.jotform.com/uploads/macarewecare/222927978372067/5926827273785219418/5926827273785219418_signature_31.png
| 2024-05-28 | 2024-05-28 12:15 | The Medical City | Ortigas Ave, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28319 | COMMERCIAL | moleculardx@themedicalcity.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 40 ASSAY: Xpert CT/NG, COUNT: 81 ASSAY: Xpert HBV VL, COUNT: 586 ASSAY: Xpert HPV, COUNT: 442 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5157 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 71 ASSAY: Xpert HIV-1 Viral Load, COUNT: 689 ASSAY: Others, COUNT: 711 |  | GX-XVI-D-10c | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926816492514807669/5926816492514807669_signature_28.png | Fritzie Mae T. Perante | https://www.jotform.com/uploads/macarewecare/222927978372067/5926816492514807669/5926816492514807669_signature_31.png
| 2024-05-28 | 2024-05-28 12:31 | The Medical City | Ortigas Ave, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28318 | COMMERCIAL | moleculardx@themedicalcity.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Others, COUNT: 11 ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7 ASSAY: Others, COUNT: 10 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GX-IV-L-6C | 802630 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B4/210071314 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926815792514096591/5926815792514096591_signature_28.png | Fritzie Mae T. Perante | https://www.jotform.com/uploads/macarewecare/222927978372067/5926815792514096591/5926815792514096591_signature_31.png
| 2024-05-28 | 2024-05-28 12:09 | Baliwag Rural Health Unit | Baliwag, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28317 | PBSP | baliwagrhulab.003@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV | 838891 | Unit for calibration | NA | #NAME? | Unit is operational B1: passed B2: passed B3: passed B4: passed Filter is worn out Filter for request |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926810300424264621/5926810300424264621_signature_28.png | Erwin Valenzuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5926810300424264621/5926810300424264621_signature_31.png
| 2024-05-28 | 2024-05-28 12:30 | Calaanan Health Center | Nadurata St. Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28316 | DOH | rosaleslorima0917@gmail.com | 25 |  | 2024-05-28 | GXIV | 831892 | Unit for PM. For module replacement of module A4. | PM and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A4 (793401) in replacement to defective A4 (758230). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926799956512541768/5926799956512541768_signature_28.png | Lorina Rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/5926799956512541768/5926799956512541768_signature_31.png
| 2024-05-28 | 2024-05-28 11:27 | Iba Super Health Center | Brgy. Bangatalinga Iba, Zambales | 3 | Zambales | Laboratory | NO | INSTALLATION | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28315 | PBSP | garadoshanna@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 |  | GXIV-D 10C | 110020681 | Installation of 10C machine | Installation and setup | Performed installation and setup of Genexpert Machine Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926762706997564379/5926762706997564379_signature_28.png | Hanna C. Gaados | https://www.jotform.com/uploads/macarewecare/222927978372067/5926762706997564379/5926762706997564379_signature_31.png
| 2024-05-28 | 2024-05-28 10:30 | National Tuberculosis Reference Laboratory | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | roeusvincent@gmail.com |  | 28314 | PBSP | RoutineNTRL@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2590 |  | GXIV-D-6C | 807606 | Modules 2 and 4 for Cepheid Additional Troubleshooting. | Troubleshooting. Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked if Filter is Clogged. Get Gateway Board Information. Run Reliability Test on Modules 2 and 4. Performed Self-test. Generate Reports. | 1 Module is Operational. Modules 1 and 2 with Temperature Related Errors for Reporting. Module 4 with Ultrasonic Horn Problem for Reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926753228642824977/5926753228642824977_signature_28.png | Vaeda Paula Ortiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5926753228642824977/5926753228642824977_signature_31.png
| 2024-05-28 | 2024-05-28 10:54 | National Tuberculosis Reference Laboratory | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | roeusvincent@gmail.com |  | 28313 | PBSP | RoutineNTRL@gmail.com | 32 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 55 ASSAY: Others, COUNT: 2094 |  | GXIV-D-10C | 848244 | Unit has 5006 and 5007 Errors. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. | Room Temperature is Not Ideal for GX Operation. Recommendation to Relocate on Colder Room. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926753228644291061/5926753228644291061_signature_28.png | vaeda paula d. ortiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5926753228644291061/5926753228644291061_signature_31.png
| 2024-05-28 | 2024-05-28 10:38 | National Tuberculosis Reference Laboratory | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | roeusvincent@gmail.com |  | 28312 | PBSP | RoutineNTRL@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2590 |  | GXIV-D-6C | 707512 | Modules with Temperature Errors. | Preventive Maintenance. Troubleshooting. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Fan. Performed Self-test. Generate Reports. | 3 Modules are Operational. 1 Module with Temperature Error for Reporting. Unit has Defective Fan. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926753228646271520/5926753228646271520_signature_28.png | vaeda paula d. ortiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5926753228646271520/5926753228646271520_signature_31.png
| 2024-05-28 | 2024-05-28 10:30 | NORTHERN MINDANAO MEDICAL CENTER | CAGAYAN DE ORO, MISAMIS ORIENTAL | 10 | Misamis Oriental | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28311 | PBSP | nmmc.gxmtbrif@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 364 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7770 ASSAY: Xpert SARS-CoV-2, COUNT: 8030 ASSAY: Xpert HIV-1 Viral Load, COUNT: 647 |  | GX-XVI | 844118 | #NAME? | #NAME? | -PERFORM CLEANING OF GENEXPERT EQUIPMENT AND WORKSTATION -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE REPLACEMENT: MODULE A4 DEF SN: 683869 WITH NEW SN : 717338 MODULE D4 DEF SN: 656996 WITH NEW SN : 758394 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926730486226538685/5926730486226538685_signature_28.png | Rione Frances Lomongo | https://www.jotform.com/uploads/macarewecare/222927978372067/5926730486226538685/5926730486226538685_signature_31.png
| 2024-05-28 | 2024-05-28 09:44 | Metro Antipolo Hospital and Medical Center | Marikina - Infanta Highway Antipolo, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28310 | COMMERCIAL | mahmci_laboratory@yahoo.com | 200 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 204 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926709810512384178/5926709810512384178_signature_28.png | Sarah Almeria | https://www.jotform.com/uploads/macarewecare/222927978372067/5926709810512384178/5926709810512384178_signature_31.png
| 2024-05-28 | 2024-05-28 09:05 | Toledo City Health Office - PPMD | Poblacion Toledo City | 7 | Cebu | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 28309 | PBSP | tchcl6038@gmail.com | 23 |  | 2024-05-28 | GX4 | 808640 | Module B3 with serial number 682924 hardware fail. | Perform preventive maintenance for the GeneXpert System. | .Save Sytem Log file .Save Installation Qualification .Save GXC files. Perform preventive maintenance. | Module B1, 2 and 4 ok to use. Module B3 not ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926682650017383037/5926682650017383037_signature_28.png | Mary june elemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5926682650017383037/5926682650017383037_signature_31.png
| 2024-05-27 | 2024-05-23 13:21 | Bagumbayan Rural Health Unit | Poblacion, Bagumbayan,Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28308 | DOH | marizalenoid@yahoo.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 914 |  | Gx-IV | 838261 | CPU hard start up problem | Replace CPU power supply CPU problem troubleshooting | CPU troubleshooting and Repair | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926323735021785794/5926323735021785794_signature_28.png | MA. RIZA R. DIONELA | https://www.jotform.com/uploads/macarewecare/222927978372067/5926323735021785794/5926323735021785794_signature_31.png
| 2024-05-27 | 2024-05-27 16:30 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Tb laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28307 | COMMERCIAL | vsmmcbloodservices@gmail.com | 24 |  |  | MCS+ | 02A090 | Due for Calibration on MCS+ | Conduct Calibration for MCS+ Preventive Maintenance in Transducers,Valve, Power Supply, Rubber pump | Conduct Calibration for MCS+ Preventive Maintenance in Transducers,Valve, Power Supply, Rubber pump | MCS+ is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926080361112805372/5926080361112805372_signature_28.png | Charlotte Mae Ygoña, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5926080361112805372/5926080361112805372_signature_31.png
| 2024-05-27 | 2024-05-27 16:19 | Dr. JOSE RIZAL MEMORIAL HOSPITAL | Dr. Jose Rizal Memorial Hospital, Dapitan City | 9 | Zamboanga del Norte | Laboratory | NO | MAINTENANCE Packed Genexpert Machine | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28306 | PBSP | lovinoantonieta@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1488 | 2024-05-27 | Gx 4 (6c) | 804698 | Module 3 Temperature error | Preventive Maintenence and Packed Genexpert Machine and other peripherals. | Preventive Maintenence and Packed Genexpert Machine and other peripherals. | Module 1,2 and 4 are all available while module 3 pending for replacement | No printer packed due to printer is already broken and unable to work properly. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926074918378428119/5926074918378428119_signature_28.png | Antonieta B. Lovino | https://www.jotform.com/uploads/macarewecare/222927978372067/5926074918378428119/5926074918378428119_signature_31.png
| 2024-05-27 | 2024-05-27 16:03 | Dr. jose rizal Memorial Hospital | Dr. Jose Rizal Memorial Hospital, Dapitan City | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION Site assessment | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28305 | PBSP | lovinoantonieta@gmail.com | 23 |  | 2024-05-27 | GX 4 10c | 110020389 | None | Installation of Genexpert Machine GX4, 10c modules. | Installation of Genexpert Machine GX4, 10c modules. | All modules are available. Brand new Uninterruptible power supply unable to turn on. | Uninterruptible power supply (UPS), model IT-M2000A is Broken and unable to turn on. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926070053027645424/5926070053027645424_signature_28.png | Antonieta B. Lovino | https://www.jotform.com/uploads/macarewecare/222927978372067/5926070053027645424/5926070053027645424_signature_31.png
| 2024-05-27 | 2024-05-27 14:40 | Panopdopan District Hospital | Lamut, Ifugao | CAR | Ifugao | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28304 | PBSP | ilamawan.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 937 |  | GXIV-6c | 807617 | Module 4 error | Performed Module Replacement A4(672908 - 685260) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926020123551023926/5926020123551023926_signature_28.png | Iolanthe Dawn L. Amawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5926020123551023926/5926020123551023926_signature_31.png
| 2024-05-27 | 2024-05-27 14:15 | San Jorge Rural Health Unit | San Jorge Samar | 8 | Samar | Laby | YES | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28303 | PBSP | johnfrancisortiz23@gmail.com | 35 |  | 2024-05-27 | GXIV | 808291 | n/a | Installation | Installed - Genexpert machine SN 808291 - CPU SN/CPH0520F08484 - LCD monitor SN/V5TL380 - UPS model IT-M200A - Printer SN/VNF8R45696 Generate reports | Unit is operational | room temperature above 30°c due to air-conditioner not working properly | - GX SN 808291 - CPU SN/CPH0520F08484 - LCD monitor SN/V5TL380 - UPS model IT-M200A - Printer SN/VNF8R45696 - Mouse/Keyboard - Barcode scanner | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5926009519531099565/5926009519531099565_signature_28.png | John Francis C. Ortiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5926009519531099565/5926009519531099565_signature_31.png
| 2024-05-27 | 2024-05-27 13:28 | Kananga Rural Health Unit | Kananga, Leyte | 8 | Leyte | TB-Dots | YES | INSTALLATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28302 | PBSP | retulla_den@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 722 | 2024-05-27 | GXII | 835085 | #NAME? | #NAME? | #NAME? | All modules are functional. |  | Dell Flat Panel Monitor SN: CN0WMYC5FCC0034RCACXA05 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925984615266963199/5925984615266963199_signature_28.png | Dennis Retulla | https://www.jotform.com/uploads/macarewecare/222927978372067/5925984615266963199/5925984615266963199_signature_31.png
| 2024-05-27 | 2024-05-24 09:40 | Ilocos Sur District Hospital- Narvacan | Paratong, Zone 4, Narvacan | 1 | Ilocos Sur | TB DOTS | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28301 | PBSP | lfm11141977@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1141 | 2024-05-24 | GXIV-4-D-6C | 807578 | For Installation and Calibration | #NAME? | Installed GeneXpert Performed Calibration | - Unit is Operational - Unable to perform XpertCheck on Newly Installed Module D4 705557 - Performed XpertCheck on Remaining Modules - Modules D1, D2 and D3 passed XpertCheck | GeneXpert from Gabriela Silang Hospital | #NAME? | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925982383455255702/5925982383455255702_signature_28.png | Lorely Miranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5925982383455255702/5925982383455255702_signature_31.png
| 2024-05-27 | 2024-05-23 08:01 | Dingras District Hospital | Dingras, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28300 | PBSP | kenacudapas1022@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 756 | 2024-05-23 | GXIV-4-D-6C | 807597 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 662121 with New Module 642641 - Generated Reports | Performed Module Replacement | #NAME? |  | GeneXpert Module 642641 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925981623455838416/5925981623455838416_signature_28.png | KENADINE CUDAPAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5925981623455838416/5925981623455838416_signature_31.png
| 2024-05-27 | 2024-05-21 12:23 | Ilocos Training and Regional Medical Center | Mc. Arthur Highway, San Fernando City, La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28299 | PBSP | itrmc.mbl@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 133 ASSAY: Xpert HPV, COUNT: 230 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3254 ASSAY: Xpert SARS-CoV-2, COUNT: 798 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GXIV-4-D-6C | 707524 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925980563454292948/5925980563454292948_signature_28.png | Russell Alili | https://www.jotform.com/uploads/macarewecare/222927978372067/5925980563454292948/5925980563454292948_signature_31.png
| 2024-05-27 | 2024-05-21 11:48 | Ilocos Training and Regional Medical Center | Mc. Arthur Highway, San Fernando City, La Union | 1 | La Union | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28298 | PBSP | itrmc.mbl@gmail.com | 27 |  | 2024-05-21 | GXIV-4-D-10C | 848249 | #NAME? | #NAME? | #NAME? | - Unit is operational - Did not perform XpertCheck on Newly Replaced Module B2 210056220 - Remaining 3 Modules Passed XpertCheck |  | CPU CPH0923F44834 XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925979673455579636/5925979673455579636_signature_28.png | Russell Alili | https://www.jotform.com/uploads/macarewecare/222927978372067/5925979673455579636/5925979673455579636_signature_31.png
| 2024-05-27 | 2024-05-27 12:43 | Santa Cruz Rural Health Unit | Poblacion Sta. Cruz, Davao del Sur | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28297 | DOH | chendalagan72@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 513 |  | GX-IV 4D | 838505 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational | Module 1 encounter temperature related error, for observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925948411515594989/5925948411515594989_signature_28.png | Dorothy A. Dalagan RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5925948411515594989/5925948411515594989_signature_31.png
| 2024-05-27 | 2024-05-24 09:07 | Imus City Health Office II | Bluebell, Bayan Luma, Imus | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28294 | DOH | amielynnmangulabnan1217@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1280 ASSAY: Xpert HIV-1 Viral Load, COUNT: 44 |  | GXIV-D-6C | 821284 | Unit is due for Maintenance | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925882803616778452/5925882803616778452_signature_28.png | Amielynn Mangulabnan | https://www.jotform.com/uploads/macarewecare/222927978372067/5925882803616778452/5925882803616778452_signature_31.png
| 2024-05-27 | 2024-05-23 13:12 | Taal Rural Health Unit | Tierra Alta, Taal, Batangas | 4A | Batangas | laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28295 | PBSP | hanetmanny@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1882 |  | GXIV-D-6C | 805259 | Time and Date CPU Problem. Unit is due for XpertCheck. | Troubleshooting. XpertCheck. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter. Performed Self-test. Performed XpertCheck. Replaced Bios Battery of CPU. Generate Reports. | With Wornout Filter. 3 Modules Passed and 1 Service Required. Date and Time Problem Resolved. | XC: A4/ 761622 - Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925882803617044575/5925882803617044575_signature_28.png | Juanita B Balbacal | https://www.jotform.com/uploads/macarewecare/222927978372067/5925882803617044575/5925882803617044575_signature_31.png
| 2024-05-27 | 2024-05-23 10:52 | San Luis Rural Health Unit | Poblacion, San Luis, Batangas | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28296 | PBSP | rhoda_villanueva05@yahoo.com | 25 |  |  | GXIV-D-6C | 838131 | Defective CPU. Unit is due for XpertCheck. | CPU Replacement. XpertCheck. | Backup Tests. Install CPU Replacement. Pullout Backup Laptop. Cleaned Cartridge Bay, Plunger Rod and Icor Optics. Checked Filter if Clogged. Performed XpertCheck. Generate Reports. | 3 Modules Passed and 1 Service Required in XpertCheck. | XC: B1/ 806301 - Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925882813616352336/5925882813616352336_signature_28.png | Rhoda Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5925882813616352336/5925882813616352336_signature_31.png
| 2024-05-27 | 2024-05-24 16:31 | Bicol South Luzon Sub National Reference Laboratory | BRTTH Compound, Legazpi City | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28293 | PBSP | rbalzaga.pbsp@gmail.com | 20 |  |  | GX4-6C | 808636 | B4/644565 error/code 2126 Continuous error msg in msg log | Reset cable connections inside Gx Generate report | Performed Troubleshooting | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925851490015411675/5925851490015411675_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5925851490015411675/5925851490015411675_signature_31.png
| 2024-05-27 | 2024-05-15 13:01 | Kamuning Health Center | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28292 | DOH | vicvic.chua25@yahoo.com | 27 | ASSAY: Xpert HPV, COUNT: 220 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1496 |  | GXIV | 831612 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925793263619763833/5925793263619763833_signature_28.png | Mary Rose Perejon | https://www.jotform.com/uploads/macarewecare/222927978372067/5925793263619763833/5925793263619763833_signature_31.png
| 2024-05-27 | 2024-05-03 14:07 | Plaridel Rural Health Unit | Plaridel, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28291 | DOH | Yellei0908@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-05-03 | GXIV | 831711 | Module A2 for replacement | Replaced module A2 by 684364 | #NAME? | Unit is operational |  | GemeXpert Module SN: 684364 | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925792983612285457/5925792983612285457_signature_28.png | Ahlelei Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5925792983612285457/5925792983612285457_signature_31.png
| 2024-05-27 | 2024-05-07 11:40 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28290 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1288 |  | GXIV | 831691 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925792663618859134/5925792663618859134_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5925792663618859134/5925792663618859134_signature_31.png
| 2024-05-27 | 2024-05-14 13:49 | Bulakan Rural Health Unit | Bulakan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28289 | DOH | l_gemino@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1194 ASSAY: Xpert HIV-1 Viral Load, COUNT: 97 |  | GXIV | 831641 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Module A1 under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925792403613819064/5925792403613819064_signature_28.png | Lilibeth Gemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5925792403613819064/5925792403613819064_signature_31.png
| 2024-05-27 | 2024-05-15 13:01 | Kamuning Health Center | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28288 | DOH | vicvic.chua25@yahoo.com | 27 | ASSAY: Xpert HPV, COUNT: 220 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1496 |  | GXIV | 831612 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925791993619747416/5925791993619747416_signature_28.png | Mary Rose Perejon | https://www.jotform.com/uploads/macarewecare/222927978372067/5925791993619747416/5925791993619747416_signature_31.png
| 2024-05-27 | 2024-05-16 13:12 | San Rafael Rural Health Unit | San Rafael, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28287 | DOH | tatsdayrit25@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 894 |  | GXIV | 831692 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925791163613647337/5925791163613647337_signature_28.png | Joana Manelle Dayrit | https://www.jotform.com/uploads/macarewecare/222927978372067/5925791163613647337/5925791163613647337_signature_31.png
| 2024-05-27 | 2024-05-17 12:32 | Meycauayan City Health Center VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28286 | DOH | ronamarieiris_25@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1716 ASSAY: Xpert HIV-1 Viral Load, COUNT: 138 |  | GXIV | 831904 | Temperature error on A3 | NA | #NAME? | Unit is operational A3 is under observation A3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925790943611926982/5925790943611926982_signature_28.png | Rona Marie Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5925790943611926982/5925790943611926982_signature_31.png
| 2024-05-27 | 2024-05-14 15:38 | Marilao Rural Health Unit | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28285 | DOH | jmjenilo27@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1546 ASSAY: Xpert HIV-1 Viral Load, COUNT: 110 |  | GXIV | 831897 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5925790103618526446/5925790103618526446_signature_28.png | Christopher Dilag | https://www.jotform.com/uploads/macarewecare/222927978372067/5925790103618526446/5925790103618526446_signature_31.png
| 2024-05-24 | 2024-05-24 22:20 | Aloguinsan Rural Health Unit | Aloguinsan, Poblacion Malolos | 7 | Cebu | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28284 | PBSP | limjennifer82@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 857 |  | gx2 | 834812 | Due for yearly calibration. | Conduct Preventive Maintenance on genexpert machine Conduct calibration on four modules Generate IQR, System, Gxx files, Assay Statistics | Conduct Preventive Maintenance on genexpert machine Conduct calibration on four modules Generate IQR, System, Gxx files, Assay Statistics | All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923698991118272112/5923698991118272112_signature_28.png | Jennifer L. Sayson, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5923698991118272112/5923698991118272112_signature_31.png
| 2024-05-24 | 2024-05-24 22:09 | Barili Rural Health Unit | Poblacion, Barili Cebu | 7 | Cebu | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28283 | PBSP | barilimainrhu@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 911 |  | Gx4 | 834693 | Due for Annual Calibrations | Conduct Preventive Maintenance on genexpert machine Conduct calibration on four modules Generate IQR, System, Gxx files, Assay Statistics Need additional filter | Conduct Preventive Maintenance on genexpert machine Conduct calibration on four modules Generate IQR, System, Gxx files, Assay Statistics | All modules passed except module 3 Serial Number 786755 failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923692581114344222/5923692581114344222_signature_28.png | Ma. Grace Cui, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5923692581114344222/5923692581114344222_signature_31.png
| 2024-05-24 | 2024-05-24 14:16 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28282 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15H030G | Due for maintenance and calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Performed maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923597501295530094/5923597501295530094_signature_28.png | Ivan | https://www.jotform.com/uploads/macarewecare/222927978372067/5923597501295530094/5923597501295530094_signature_31.png
| 2024-05-24 | 2024-05-24 15:25 | Philippine Blood Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28281 | MACARE-HBDC | strobecatcher@gmail.com | 26 |  |  | MCS+ | 15H028G | Due for maintenance and calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Performed maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923597001293284047/5923597001293284047_signature_28.png | Ivan | https://www.jotform.com/uploads/macarewecare/222927978372067/5923597001293284047/5923597001293284047_signature_31.png
| 2024-05-24 | 2024-05-24 15:08 | Binangonan Rural Health Unit III | Pritil St. Binangonan, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28280 | PBSP | dhayanara.zapanta@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 575 |  | GX-IV-D-6C | 846045 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | All modules operational. Module B4 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923548196182120032/5923548196182120032_signature_28.png | Jenalyn Antiporda | https://www.jotform.com/uploads/macarewecare/222927978372067/5923548196182120032/5923548196182120032_signature_31.png
| 2024-05-24 | 2024-05-24 16:49 | Agusan del Norte Provincial Health Office | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | INSTALLATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28279 | PBSP | ronniebalansag74@gmail.com | 27 |  |  | GX-4 | 807461 | None | None | Installation of GeneXpert 6-colors and peripherals. | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923501249224015567/5923501249224015567_signature_28.png | Ronnie M. Balansag | https://www.jotform.com/uploads/macarewecare/222927978372067/5923501249224015567/5923501249224015567_signature_31.png
| 2024-05-24 | 2024-05-24 15:59 | Margosatubig Regional Hospital | Margosatubig Regional Hospital, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | INSTALLATION | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28278 | PBSP | cherylpuludo@gmail.com | 27 |  | 2024-05-24 | GX 4 / 10c | 110019600 | None | Installation of Genexpert Machine GX 4 (10c). | Installation of Genexpert Machine GX 4 (10c). | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923471853561017102/5923471853561017102_signature_28.png | Cheryl Pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/5923471853561017102/5923471853561017102_signature_31.png
| 2024-05-24 | 2024-05-24 09:44 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28277 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  |  | MCS + | 11H204 | Due for maintenance and calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Adjust bowl optics value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Performed calibration and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923449689428200624/5923449689428200624_signature_28.png | Philippine Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5923449689428200624/5923449689428200624_signature_31.png
| 2024-05-24 | 2024-05-24 10:21 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28276 | MACARE-HBDC | phc_bb@yahoo.com | 26 |  |  | MCS + | 20168 | Due for maintenance and calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Performed maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923449299423667967/5923449299423667967_signature_28.png | Philippine Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5923449299423667967/5923449299423667967_signature_31.png
| 2024-05-24 | 2024-05-24 13:56 | Binangonan Rural Health Unit | Manila East Rd, Binangonan, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28275 | PBSP | devera.dexter@gmail.com | 24 |  |  | GX-IV-D-6C | 807605 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | All modules passed. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447282574825527/5923447282574825527_signature_28.png | Dexter P. De Vera | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447282574825527/5923447282574825527_signature_31.png
| 2024-05-24 | 2024-05-24 13:52 | Binangonan Rural Health Unit | Manila East Rd, Binangonan, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28274 | PBSP | devera.dexter@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2250 |  | GX-IV-D-6C | 838199 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | All modules operational. Modules B1, B3 and B4 passed. Module B2 newly replaced. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447282575025548/5923447282575025548_signature_28.png | Dexter P. De Vera | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447282575025548/5923447282575025548_signature_31.png
| 2024-05-24 | 2024-05-24 09:41 | Marikina City Health Office | Shoe Ave, Marikina | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28273 | PBSP | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1697 ASSAY: Xpert HIV-1 Viral Load, COUNT: 345 |  | GX-IV-D-6C | 805211 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Checked module A1 Generated reports | Unit is operational. Module A1/650238 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447162575721895/5923447162575721895_signature_28.png | Lou Tomas | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447162575721895/5923447162575721895_signature_31.png
| 2024-05-24 | 2024-05-24 09:37 | Marikina City Health Office | Shoe Ave., Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28271 | MACARE-HBDC | estyaragon2019@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447152571580885/5923447152571580885_signature_28.png | Lou Tomas | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447152571580885/5923447152571580885_signature_31.png
| 2024-05-24 | 2024-05-24 09:50 | Marikina City Health Office | Shoe Ave, Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28272 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 294 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 |  | GX-IV-D-10C | 110020361 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447152573529761/5923447152573529761_signature_28.png | Lou Tomas | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447152573529761/5923447152573529761_signature_31.png
| 2024-05-24 | 2024-05-24 09:48 | Marikina City Health Office | Shoe Ave, Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28269 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 779 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447152575489442/5923447152575489442_signature_28.png | Lou Tomas | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447152575489442/5923447152575489442_signature_31.png
| 2024-05-24 | 2024-05-24 09:44 | Marikina City Health Office | Shoe Ave, Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28270 | DOH | estyaragon2019@gmail.com | 24 |  |  | GX-IV-D-6C | 831588 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447152578706179/5923447152578706179_signature_28.png | Lou Tomas | https://www.jotform.com/uploads/macarewecare/222927978372067/5923447152578706179/5923447152578706179_signature_31.png
| 2024-05-24 | 2024-05-24 14:08 | Bagong Silang Ph. 1 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28268 | PBSP | alcaydeirene07@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2780 |  | GXIV | 838856 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: passed B2: passed B3: newly replaced B4: newly replaced |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923432596725501906/5923432596725501906_signature_28.png | Irene Alcayde | https://www.jotform.com/uploads/macarewecare/222927978372067/5923432596725501906/5923432596725501906_signature_31.png
| 2024-05-24 | 2024-05-23 12:49 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28266 | DOH | laboratory@loveyourself.com.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 427 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1097 | 2024-05-23 | GXIV | 831895 | Module A2 for replacement | Replaced module A2 by 756790 | #NAME? | Unit is operational |  | GeneXpert Module SN: 756790 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923432596728602686/5923432596728602686_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/5923432596728602686/5923432596728602686_signature_31.png
| 2024-05-24 | 2024-05-23 15:08 | Maysan 3S Health Center | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28267 | PBSP | neldimatera@gmail.com | 25 |  |  | GXIV | 815381 | GeneXpert Workstation not booting up | Replaced BIOS battery Reset internal connections | #NAME? | Unit is not operational GX desktop not booting up Desktop is recommended for repair/replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923432596729084357/5923432596729084357_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/5923432596729084357/5923432596729084357_signature_31.png
| 2024-05-24 | 2024-05-24 14:21 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28265 | DOH | bmccrcl2022@gmail.com | 27 |  |  | GX-4 6 colors | 838403 | None | None | Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923420024508931803/5923420024508931803_signature_28.png | Kristine Pelletero | https://www.jotform.com/uploads/macarewecare/222927978372067/5923420024508931803/5923420024508931803_signature_31.png
| 2024-05-24 | 2024-05-24 14:15 | Butuan Medical Center | KM 5, Baan, Butuan City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28264 | PBSP | bmccrcl2022@gmail.com | 28 |  |  | GX-4 6 colors | 805099 | Module B2 and B4 found error in temperature | Reported to Cepheid | Troubleshooting of Module B2 using XPC Data Collection Preventive Maintenance | 2 modules working and 2 modules for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923419404504551525/5923419404504551525_signature_28.png | Kristine Pelletero | https://www.jotform.com/uploads/macarewecare/222927978372067/5923419404504551525/5923419404504551525_signature_31.png
| 2024-05-24 | 2024-05-23 15:24 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | YES | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28263 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 53 ASSAY: Others, COUNT: 5 ASSAY: Xpert HBV VL, COUNT: 27 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 41 ASSAY: Xpert HCV VL, COUNT: 9 |  | GX4 D | 838524 | Due for maintenance B1,B3,B4 reported already confirmed for replacement | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | B2 operational B1,B3,B4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923404789427377456/5923404789427377456_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5923404789427377456/5923404789427377456_signature_31.png
| 2024-05-24 | 2024-05-23 14:54 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28262 | COMMERCIAL | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Others, COUNT: 12 ASSAY: Xpert HPV, COUNT: 42 ASSAY: Xpert MTB/RIF Ultra, COUNT: 178 ASSAY: Xpert HCV VL, COUNT: 12 |  | GX16 D | 804476 | Due for maintenance, A1 A2,B1,B3,C1,C2,C3,C4 disaabled due of continous error loss communication error B4 temp error 4009,4012,4017,4010 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | 9 modules disable due to error 3 module operational A3,A4,B2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923404379421891922/5923404379421891922_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5923404379421891922/5923404379421891922_signature_31.png
| 2024-05-24 | 2024-05-23 13:54 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28261 | MACARE-HBDC | molecularlabnkti@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1375 ASSAY: Others, COUNT: 75 ASSAY: Xpert HBV VL, COUNT: 1004 ASSAY: Xpert HPV, COUNT: 90 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2298 ASSAY: Xpert SARS-CoV-2, COUNT: 15034 ASSAY: Xpert HCV VL, COUNT: 126 |  | GX16 D | 110005587 | A1 encounter temp error 1002/4017 C2 encounter temp error 4010 D2,A2,A4 reported already confirmed for replacement For maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | 11 module operational 5 modules disable due to temp error | A1 and C2 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923403889422392352/5923403889422392352_signature_28.png | Ma. Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5923403889422392352/5923403889422392352_signature_31.png
| 2024-05-24 | 2024-05-24 09:51 | Binan City Health Office 2 | Binan Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28260 | DOH | eloicyparra@gmail.com | 25 |  | 2024-05-24 | GXIV | 821278 | Unit for preventive maintenance. For module replacement of module | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A1 (627235) in replacement to defective A1 (783692). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923381155125614353/5923381155125614353_signature_28.png | Elizabeth de Pedro | https://www.jotform.com/uploads/macarewecare/222927978372067/5923381155125614353/5923381155125614353_signature_31.png
| 2024-05-24 | 2024-05-24 09:49 | San Pedro RHU 2 | Brgy. Langgam, San Pedro Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28259 | DOH | miraquely@gmail.com | 25 |  | 2024-05-24 | GXIV | 821189 | Unit for preventive maintenance. For module replacement of module A4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A4 (619830) in replacement to defectivr A4 (883941). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923378445129595467/5923378445129595467_signature_28.png | MIRAQUEL GUIAO YUZON | https://www.jotform.com/uploads/macarewecare/222927978372067/5923378445129595467/5923378445129595467_signature_31.png
| 2024-05-24 | 2024-05-24 13:19 | Dingalan Community Hospital | Caragsacan, Dingalan, Aurora | 3 | Aurora | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28258 | DOH | pajarilloerwin@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 613 |  | GXIV-6c | 831621 | Due for Quarterly Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923374485717517806/5923374485717517806_signature_28.png | Erwin Pajarillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5923374485717517806/5923374485717517806_signature_31.png
| 2024-05-24 | 2024-05-24 12:14 | Margosatubig Regional Hospital | Margosatubig Regional Hospital, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 28257 | PBSP | cherylpuludo@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2069 ASSAY: Xpert HIV-1 Viral Load, COUNT: 77 | 2024-05-24 | GX 4 | 805121 | Module 3 and 4 H/W failed | Replaced module 614393 and 627052 with Module 769480 and 606863 respectively. Preventive Maintenance. | Replaced module 614393 and 627052 with Module 769480 and 606863 respectively. Preventive Maintenance. | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923365323565777430/5923365323565777430_signature_28.png | Cheryl Pulido | https://www.jotform.com/uploads/macarewecare/222927978372067/5923365323565777430/5923365323565777430_signature_31.png
| 2024-05-24 | 2024-05-24 12:02 | Northern Samar Provincial Hospital | Catarman Northern Samar | 8 | Northern Samar | Laboratory | YES | MAINTENANCE Pullout | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28256 | PBSP | dinahada70@gmail.com | 27 |  | 2024-05-24 | GXIV 6C | 807566 | n/a | #NAME? | - Performed preventive maintenance - Disassembled unit and packed Genexpert SN/807566 CPU SN/2UA381GQJ UPS SN/40022022002 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923362288854493655/5923362288854493655_signature_28.png | Dinah F. Ada,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5923362288854493655/5923362288854493655_signature_31.png
| 2024-05-24 | 2024-05-24 11:14 | Southern Philippines Medical Center | JP Laurel, Davao City | 11 | Davao del Sur | TB-HIV | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28255 | PBSP | spmcrhivda@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2683 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1728 |  | GX-IV 4D | 808298 | Module B1 reported for multiple temperature related error Module B3, encounter temperature related errors For Preventive Maintenance | Replaced the Defective Module SN:210068532 with new Module SN: 725035 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck on Module B3 | Perform Module Replacement Perform Preventive Maintenance Perform Xpertcheck for error verification on Module B3 | Module B3 for Service Required, temporarily Disabled Only Three Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923357232659036608/5923357232659036608_signature_28.png | Heart Hellenic S. Descallar | https://www.jotform.com/uploads/macarewecare/222927978372067/5923357232659036608/5923357232659036608_signature_31.png
| 2024-05-24 | 2024-05-24 12:46 | Northern Samar Provincial Hospital | Catarman Northern Samar | 8 | Northern Samar | Laboratory | YES | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28254 | PBSP | dinahada70@gmail.com | 25 |  | 2024-05-24 | GXIV 10C | 12000480 | n/a | Installation | - install new Genexpert 10c SN: 120000480 - CPU SN/CPH0523F42151 - LCD Monitor SN/V5FTG709 - Printer SN/CNB1RBK7P2 - UPS SN/SKPH2KV23910070 | Unit is Operational |  | - Gexpert SN: 120000480 - CPU SN/CPH0523F42151 - LCD Monitor SN/V5FTG709 - Printer SN/CNB1RBK7P2 - UPS SN/SKPH2KV23910070 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923356688858474487/5923356688858474487_signature_28.png | Dinah F. Ada | https://www.jotform.com/uploads/macarewecare/222927978372067/5923356688858474487/5923356688858474487_signature_31.png
| 2024-05-24 | 2024-05-23 10:53 | Dumalinao Rural Health Unit | Purok talisay barangay pagasa dumalinao ZDS | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 28253 | WHO | ryannjamesbalasabas@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 511 | 2024-05-23 | GX 4 | 847250 | Module 1 hardware failed | Replaced module 909829 with Module 701981 | Preventive Maintenence and Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923328643568009371/5923328643568009371_signature_28.png | RYANN JAMES BALASABAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5923328643568009371/5923328643568009371_signature_31.png
| 2024-05-24 | 2024-05-24 11:45 | Gabaldon Rural Health Unit | South Poblacion, Gabaldon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28252 | PBSP | chinablue_12@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 274 |  | GXII-6C | 834804 | Module 1 and 2 error | Performed Module Replacement B1(787168 - 676771) B2(787169 - 677244) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923319892515642568/5923319892515642568_signature_28.png | Helen Grace Parungao | https://www.jotform.com/uploads/macarewecare/222927978372067/5923319892515642568/5923319892515642568_signature_31.png
| 2024-05-24 | 2024-05-24 11:01 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28251 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 76 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3359 ASSAY: Xpert HIV-1 Viral Load, COUNT: 472 |  | GXIV-4D | 838285 | Due for PM | #NAME? | #NAME? | Still waiting for module A4 replacement. Other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923298478178927842/5923298478178927842_signature_28.png | Sarah Nicole Aurea Merino | https://www.jotform.com/uploads/macarewecare/222927978372067/5923298478178927842/5923298478178927842_signature_31.png
| 2024-05-24 | 2024-05-24 11:09 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28250 | DOH | vsmmdeptpatho.microbiology@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 41 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3273 ASSAY: Xpert HIV-1 Viral Load, COUNT: 429 |  | GXIV-4D | 820206 | Due for PM | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923298478176098560/5923298478176098560_signature_28.png | Sarah Nicole Aurea Merino | https://www.jotform.com/uploads/macarewecare/222927978372067/5923298478176098560/5923298478176098560_signature_31.png
| 2024-05-24 | 2024-05-24 09:46 | Calamba City Health Office | Baretto St. Calamba Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28249 | DOH | calambatbdots@gmail.com | 25 |  |  | GXIV | 838225 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923254831953021256/5923254831953021256_signature_28.png | Lily P.Banalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5923254831953021256/5923254831953021256_signature_31.png
| 2024-05-24 | 2024-05-24 09:42 | Calamba City Health Office | Baretto St. Calamba Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28248 | PBSP | calambatndots@gmail.com | 25 |  |  | GXIV | 807570 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923254311953857082/5923254311953857082_signature_28.png | Lily P.Banalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5923254311953857082/5923254311953857082_signature_31.png
| 2024-05-24 | 2024-05-24 09:59 | President Ramon Magsaysay Memorial Hospital | Iba Zambales | 3 | Zambales | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28247 | MACARE-HBDC | prmmhmoleclab@gmail.com | 29 | ASSAY: Xpert SARS-CoV-2, COUNT: 2108 |  | GXIV-D | 838520 | Due for maintenance B4: H/W Failed | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational except B4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5923253708316642115/5923253708316642115_signature_28.png | Bianca Yzabel Yap | https://www.jotform.com/uploads/macarewecare/222927978372067/5923253708316642115/5923253708316642115_signature_31.png
| 2024-05-23 | 2024-05-23 21:02 | Minglanilla Rural Health Unit | Pider 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28246 | DOH | brendagonzagaparedes@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1350 |  | GX4 | 838512 | Reset module 4 due to self-test error 4009. | Reset module 4 due to self-test error 4009. Generate IQR, System log, Assay Statistics | Reset module 4 due to self-test error 4009. Generate IQR, System log, Assay Statistics | All module are available for test. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922788121112404141/5922788121112404141_signature_28.png | Brenda Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/5922788121112404141/5922788121112404141_signature_31.png
| 2024-05-23 | 2024-05-23 15:48 | Masinloc Rural Health Unit | Masinloc Zambales | 3 | Zambales | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28245 | DOH | rodenne@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1208 |  | GXIV-D | 831608 | A3: 636534 error 2014 Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational A3: 6363534 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922786048128506531/5922786048128506531_signature_28.png | Rodenn Eduvala | https://www.jotform.com/uploads/macarewecare/222927978372067/5922786048128506531/5922786048128506531_signature_31.png
| 2024-05-23 | 2024-05-22 13:57 | Balanga Rural Health Unit | Samaguita St. Brgy. San Jose Balanga City, Bataan | 3 | Bataan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28244 | PBSP | cho.melaniealonzo@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 732 | 2024-04-29 | GXII-D | 835102 | CPU start up error Date and Time not syncing Due for calibration | Repair( CMOS replacement) Xpertcheck | Performed CMOS battery replacement Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed calibration Performed functionality test Gathered reports | Unit is operational 2 modules passed calibration |  | 2 genexpert kit | 47695 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922785628127824941/5922785628127824941_signature_28.png | Melanie Alonzo | https://www.jotform.com/uploads/macarewecare/222927978372067/5922785628127824941/5922785628127824941_signature_31.png
| 2024-05-23 | 2024-05-23 15:08 | Manito Rural Health Unit | Purok 2, It-Ba, Manito, Albay | 5 | Albay | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | rhumanito40@gmail.com |  | 28243 | PBSP | haelgermones18@gmail.com | 30 |  |  | GX4-6C | 804257 | GX4-6C for Installation | Installed/set-up GX4-6C(804257), Monitor, CPU, UPS, Printer, Barcode Scanner Checked/adjust user information Generate report | Installed Gx4-6C | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922651428916081601/5922651428916081601_signature_28.png | Hael Lachet L. Germones | https://www.jotform.com/uploads/macarewecare/222927978372067/5922651428916081601/5922651428916081601_signature_31.png
| 2024-05-23 | 2024-05-23 10:31 | Tondo Health Center | N. Zamora St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28242 | DOH | malorinalontoc@gmail.com | 25 |  |  | GXIV | 831620 | For troubleshooting of UPS. | Troubleshooting. | Performed troubleshooting of UPS. Generated reports. | Unit is operational. UPS Accu power (AT1000180910ST0206) not backing up, for replacement of batttery or UPS. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922616749517187187/5922616749517187187_signature_28.png | Lorina Lontoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5922616749517187187/5922616749517187187_signature_31.png
| 2024-05-23 | 2024-05-22 16:06 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | princeszvillaran13@gmail.com |  | 28241 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 24 | ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 917 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-XVI-D-6C | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned: - cartridge bay -plunger rod -icore -filter Checked module C3 Generated reports | Unit is operational. Modules A1/860941, A2/860940, B3/702534, B4/860923, C2/668748, D2/620837, D4/661946 and C3/662448, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922596903462069306/5922596903462069306_signature_28.png | Justine Lorenz Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5922596903462069306/5922596903462069306_signature_31.png
| 2024-05-23 | 2024-05-23 14:53 | Bais District Hospital | Bry. Poblacion | 7 | Negros Oriental | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 28240 | PBSP | laboratorybdh@gmail.com | 23 |  |  | Gx4 | 838857 | Module B3 with serial number 878813 hardware fail. | Replace module B3. | Replace module B3 with new serial number 706241 | All modules ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922567134103154798/5922567134103154798_signature_28.png | EVAN JUN C. PACURIB | https://www.jotform.com/uploads/macarewecare/222927978372067/5922567134103154798/5922567134103154798_signature_31.png
| 2024-05-23 | 2024-05-23 14:31 | Calbayog City Health Office | Calbayog Samar | 8 | Samar | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28239 | PBSP | aaron_noel122770sd@yahoo.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2005 | 2024-05-23 | GXIV | 804265 | Due for Calibration | #NAME? | #NAME? | Unit is Operational - Module B2 for service required - UPS not working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922564884816496982/5922564884816496982_signature_28.png | Ma. Leonora S. Dejapin | https://www.jotform.com/uploads/macarewecare/222927978372067/5922564884816496982/5922564884816496982_signature_31.png
| 2024-05-23 | 2024-05-23 14:04 | Masambong Health Center | 8 Malac St., Masambong, Quezon City, Metro Manila | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28238 | PBSP | ichellebonrostro@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2032 | 2024-04-30 | GX IV 6C-D | 814160 | Unit due for XpertCheck | Performed Preventive Maintenance Performed Xpertcheck | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpertcheck on modules B1, B2 and B3 B4 newly replaced Generated report | Unit is operational B1: Failed B2: Passed B3: Passed B4: Newly replaced |  | XpertCheck Kit | DR 48175 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922555221323128290/5922555221323128290_signature_28.png | Michelle Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/5922555221323128290/5922555221323128290_signature_31.png
| 2024-05-23 | 2024-05-23 12:19 | Western Pangasinan District Hospital | Alaminos Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE 6C Machine Disassembly | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28237 | PBSP | wpdhlab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1912 |  | GXIV-D | 87598 | Due for Calibration Machine Disassembly | Preventive maintenance Xpertcheck Disassembly and packing | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed xpertcheck Performed functionality test Gathered reports Performed machine Disassembly and Packing | Unit is operational Module A2 and A3 unsupported for xpertcheck Module A1 abd A4 passed calibration | Machine is packed and ready for pickup |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922510270817986183/5922510270817986183_signature_28.png | Czarina Catabay | https://www.jotform.com/uploads/macarewecare/222927978372067/5922510270817986183/5922510270817986183_signature_31.png
| 2024-05-23 | 2024-05-22 10:53 | Kabacan Rural Health Unit | Poblacion, Kabacan North Cotabato | 12 | Cotabato | Tb-Lab x | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28236 | PBSP | yayo_kmsi@yahoo.com | 28.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 742 | 2024-05-22 | Gx-IV (2modules) | 835182 | Module 2 for Service Required | Replace defective module 2 SN:788903 with new module SN: 779279 Genexpert machine Cleaning Checking of errors Data collection | Perform Genexpert machine module replacement/repair Perform Genexpert machine Preventive maintenance | 2 operational modules |  | Genexpert module 6-color SN:779279 | SN:779279 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922508613476769074/5922508613476769074_signature_28.png | Joan A Sangines | https://www.jotform.com/uploads/macarewecare/222927978372067/5922508613476769074/5922508613476769074_signature_31.png
| 2024-05-23 | 2024-05-23 12:35 | Valentina G. Plaza Memorial Hospital | San Teodoro, Bunawan | 13 | Agusan del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28235 | PBSP | rolandoivanlagapa@gmail.com | 23.35 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4743 |  | GX-4 6 colors | 815436 | A3 found temperature errors Ventillation not recommended | Module Replacement | Replacing Module A3 with SN: 870956 to SN: 210045941 Preventive Maintenance | All modules and units are functional | Ventillation of the site not recommended. Air-condtioning unit is required. System IQR temperature must range (18 to 25 degrees celsius). |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922502227715537894/5922502227715537894_signature_28.png | Rolando P. Lagapa | https://www.jotform.com/uploads/macarewecare/222927978372067/5922502227715537894/5922502227715537894_signature_31.png
| 2024-05-23 | 2024-05-23 12:39 | San Isidro Rural Health Unit | Batobato, San Isidro Davao Oriental | 11 | Davao Oriental | TB-DOTS | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28234 | PBSP | kwoc1510@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 220 |  | GX-IV 2D | 835222 | Module B1 for service required Due for Preventive Maintenance | Replaced the Defective Module SN:788659 with new Module SN: 713886 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck on Module B2 for error verification | Perform Module Repair Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922496935212658020/5922496935212658020_signature_28.png | Kjell Wenn Cabusao | https://www.jotform.com/uploads/macarewecare/222927978372067/5922496935212658020/5922496935212658020_signature_31.png
| 2024-05-23 | 2024-05-23 11:51 | Tinambacan Main Health Center | Tinambacan Calbayog City | 8 | Samar | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28233 | PBSP | porlaressteph@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 429 | 2024-05-23 | GXIV | 834824 | -Module B1/787208 defective | #NAME? | - Module replacement B1/787208 into B1/716897 - Performed preventive maintenance - Cleaned plunger rod, cartridge bay and filter - Generate reports | Unit is operational |  | 1 Genexpert 6 color module SN: 716897 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922468453849339916/5922468453849339916_signature_28.png | Stephanie Jis Rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/5922468453849339916/5922468453849339916_signature_31.png
| 2024-05-23 | 2024-05-23 11:53 | Guihulngan City Health Office | Poblacion Guihulngan | 7 | Negros Oriental | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 28232 | PBSP | kian.macalua06@gmail.com | 23 |  | 2024-05-23 | Gx4 | 815463 | Module B3 with serial number 732888 Hardware fail. | Replace new module | Replace module B3 with new serial number 624674 | Module B2, 3,4 ok to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922461009838787879/5922461009838787879_signature_28.png | Hyacinth S. Pedillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5922461009838787879/5922461009838787879_signature_31.png
| 2024-05-23 | 2024-05-23 11:31 | Bayawan City Health Office | Bayawan City | 7 | Negros Oriental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 28231 | PBSP | lobella_belliet@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2245 |  | GX4 | 808282 | Module B1 with SN 607064 failed in XPC. | Conduct PM and save files(IQR, System Log, Assay Stats and GXC File) | Run Xpert check calibration for 4 modules. Module B1 need service while module B2, B3 & B4 passed in XPC. | Module B1 need service while modules B2, B3 and B4 passed in XPC. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922452037162027580/5922452037162027580_signature_28.png | Lobella T. Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5922452037162027580/5922452037162027580_signature_31.png
| 2024-05-23 | 2024-05-23 10:53 | Dumalinao Rural Health Unit | Purok talisay barangay pagasa dumalinao ZDS | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 28230 | WHO | ryannjamesbalasabas@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 511 | 2024-05-23 | GX 4 | 847250 | Module 1 hardware failed | Replaced module 909829 with Module 701981 | Preventive Maintenence and Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922435768667592585/5922435768667592585_signature_28.png | RYANN JAMES BALASABAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5922435768667592585/5922435768667592585_signature_31.png
| 2024-05-23 | 2024-05-23 10:53 | Dumalinao Rural Health Unit | Purok talisay barangay pagasa dumalinao ZDS | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 28229 | WHO | ryannjamesbalasabas@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 511 | 2024-05-23 | GX 4 | 847250 | Module 1 hardware failed | Replaced module 909829 with Module 701981 | Preventive Maintenence and Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922435308666422293/5922435308666422293_signature_28.png | RYANN JAMES BALASABAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5922435308666422293/5922435308666422293_signature_31.png
| 2024-05-23 | 2024-05-23 10:55 | Bogo City Health Office | Bogo City, Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28228 | PBSP | bogocholab@gmail.com | 26 |  |  | Gx 4 | 120000494 | Defective UPS To be reported by the End User | New Genexpert Machine Installation Install MTB ultra assay Gathering data, IQR | New Genexpert Machine Installation | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922424892118641994/5922424892118641994_signature_28.png | Reggie V. Dahili | https://www.jotform.com/uploads/macarewecare/222927978372067/5922424892118641994/5922424892118641994_signature_31.png
| 2024-05-23 | 2024-05-23 10:03 | Bagong Pag-Asa Health Center | 28 Rd 9, Bago Bantay, Quezon City, Metro Manila | NCR | NCR 2nd - Quezon City | Labo | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28227 | DOH | louiseanntangolmontenegro@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1397 |  | GX IV 6C-D | 831597 | Printer not working Unit due maintenance | Troubleshoot printer Performed PM Performed reliability test | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Checked Printer Installed driver Replaced Ink cartridges Performed reliability test All modules passed Generated report | Unit is operational. All modules and printer working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922399101323350642/5922399101323350642_signature_28.png | LOUISE ANN T. MONTENEGRO | https://www.jotform.com/uploads/macarewecare/222927978372067/5922399101323350642/5922399101323350642_signature_31.png
| 2024-05-23 | 2024-05-23 09:58 | Parola Health Center | Parola, Tondo Manila | NCR | NCR 1st - Manila | Laboatory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28226 | DOH | olartecharlie06@gmail.com | 25 |  | 2024-05-23 | GXIV | 831871 | Unit for preventive maintenance. For module replacement of module B3. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B3 (612920) in replacement to defective B3 (661116). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5922389346428259250/5922389346428259250_signature_28.png | Charlie olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5922389346428259250/5922389346428259250_signature_31.png
| 2024-05-22 | 2024-05-22 21:28 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb laboratory | YES | TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28225 | PBSP | chaska.lp14@gmail.com | 35 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX4 | 110020263 | UPS defective during installation | Connect to the existing UPS and power up the unit | Connect to the existing UPS and power up the unit | Need new UPS for the new genexpert machine |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921939229419078170/5921939229419078170_signature_28.png | Lucille Potoy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5921939229419078170/5921939229419078170_signature_31.png
| 2024-05-22 | 2024-05-22 21:00 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28224 | DOH | chaska.lp14@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2156 |  | GX4 | 838283 | Due for Monthly Preventive Maintenance Clean up probe, Icore, cartridge docking bay, and filter | Due for Monthly Preventive Maintenance Clean up probe, Icore, cartridge docking bay, and filter | Due for Monthly Preventive Maintenance Clean up probe, Icore, cartridge docking bay, and filter Generate GXX files, IQR, Sys Log, and Assay Statistics | All units of module are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921932549417743994/5921932549417743994_signature_28.png | Lucille Potoy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5921932549417743994/5921932549417743994_signature_31.png
| 2024-05-22 | 2024-05-22 20:45 | Badian Rural Health Unit | Poblacion, Badian Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28223 | PBSP | avenzr@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 1219 |  | 807588 | 210093378 | Module 4 SN # 210093378 Hardware failure | Replaced Module 4 SN # 210093378 Generate Installation Qualification Report, System Log, Gxx files, and Assay Statistics Calibrate Genexpert machine A1 - Failed, A2/A3 - Passed, - Area of the laboratory required additional airconditioning unit to maintain desired room temperature. | Replaced Module 4 SN # 210093378 Generate Installation Qualification Report, System Log, Gxx files, and Assay Statistics - Calibrate Genexpert machine A1 - Failed, A2/A3 - Passed, | All Units are operational. A1 - Failed during calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921915949412734000/5921915949412734000_signature_28.png | Avenizer Absin, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5921915949412734000/5921915949412734000_signature_31.png
| 2024-05-22 | 2024-05-22 15:07 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28222 | MACARE-HBDC | roseannparnes@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1573 | 2024-05-21 | GX4 D(10C) | 110014085 | For module replacement b4 684160 - 619325 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | Unit operational |  | 900 -0650rs | 11541 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921831501291274878/5921831501291274878_signature_28.png | Rose Ann Parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5921831501291274878/5921831501291274878_signature_31.png
| 2024-05-22 | 2024-05-22 16:38 | Dinalupihan Rural Health Unit | Dinalupihan Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28221 | PBSP | dianalandicho87@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2244 ASSAY: Xpert HIV-1 Viral Load, COUNT: 49 ASSAY: Xpert HIV-1 Viral Load, COUNT: 41 | 2024-04-29 | GXIV-D | 807561 | Due for Calibration UPS not Backing Up | Xpertcheck Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed xpertcheck Performed functionality test Gathered reports | Unit is operational UPS for replacement All modules passed calibration |  | 4 xpertcheck kits | 47697 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921774208691423210/5921774208691423210_signature_28.png | Diana Landicho | https://www.jotform.com/uploads/macarewecare/222927978372067/5921774208691423210/5921774208691423210_signature_31.png
| 2024-05-22 | 2024-05-22 15:31 | Marihatag District Hospital | Pagbahanan II, Poblacion, Marihatag | 13 | Surigao del Sur | Laboratory | NO | INSTALLATION | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28220 | PBSP | marihatagdh@gmail.com | 30 |  |  | GX-4 | 804255 | None | None | Installation of GeneXpert Machine from Adela Serra Ty Memorial Medical Center | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921748384657540695/5921748384657540695_signature_28.png | Precious Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/5921748384657540695/5921748384657540695_signature_31.png
| 2024-05-22 | 2024-05-22 15:46 | Doña Marta Health Center | 39 Don C Revilla St., Pasay, Metro Manila | NCR | NCR 4th - Pasay | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28219 | DOH | LBCANICULA@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1389 |  | GX IV 6C-D | 831700 | Unit due for maintenance | Performed Preventive Maintenance Performed reliability test | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Filter for reporting Performed reliability test All modules passed reliability test Generated report | Unit is operational All modules working Filter for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921745847498217260/5921745847498217260_signature_28.png | Lorna Canicula | https://www.jotform.com/uploads/macarewecare/222927978372067/5921745847498217260/5921745847498217260_signature_31.png
| 2024-05-22 | 2024-05-22 15:49 | Governor Celestino Gallares Memorial Hospital | TAGBILARAN, BOHOL | 7 | Bohol | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28218 | DOH | cyrilxoxocagoco@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2388 ASSAY: Xpert HIV-1 Viral Load, COUNT: 207 |  | GX-IV | 820228 | #NAME? | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | -PERFORM CLEANING OF GX WORKSTATION AND EQUIPMENT -PERFORM MODULE B1 DEF SN 600529 REPLACEMENT WITH NEW SN 616526 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921742553417003416/5921742553417003416_signature_28.png | Cyril Cagoco, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5921742553417003416/5921742553417003416_signature_31.png
| 2024-05-22 | 2024-05-22 14:52 | Amang Rodriguez Memorial Medical Center | Sumulong Highway, Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28217 | PBSP | maricel_c27@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4565 |  | GX-IV-D-6C | 807982 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Run XpertCheck Generated reports | Unit is operational. Modules B1 and B4 passed. Modules B2 and B3 not calibrated, out of cartridge due to retry. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921733560539607408/5921733560539607408_signature_28.png | Maricel Calderon | https://www.jotform.com/uploads/macarewecare/222927978372067/5921733560539607408/5921733560539607408_signature_31.png
| 2024-05-22 | 2024-05-22 15:40 | Amang Rodriguez Memorial Medical Center | Sumulong Highway, Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28216 | PBSP | pathologymole.armmc@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 2802 |  | GX-XVI-D-6C | 844113 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A3/891761 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921733560539857632/5921733560539857632_signature_28.png | Racquel Verzosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5921733560539857632/5921733560539857632_signature_31.png
| 2024-05-22 | 2024-05-22 15:13 | Matnog Medicare Hospital | Matnog, Sorsogon | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28215 | PBSP | mtmh@sorsogon.gov.ph | 26 |  |  | GX4-6C 2placer | 835200 | CPU for replacement | CPU Replacement: CPH0923F44832 Cleeaned Cartridge bays, plunger rod, icore Generate reports | Replaced CPU Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921718774115066998/5921718774115066998_signature_28.png | Mernadith Garcera | https://www.jotform.com/uploads/macarewecare/222927978372067/5921718774115066998/5921718774115066998_signature_31.png
| 2024-05-22 | 2024-05-22 14:46 | Nabunturan Rural Health Unit | Poblacion, Nabunturan Davao de Oro | 11 | Davao de Oro | TB-DOTS | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28214 | WHO | mbacerraiii@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1582 |  | GX-IV 4D | 847253 | Error on Module B3 UPS Not working | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Replaced the Defective Module SN:684339 with new Module SN:708660 Check the line voltage of the PSU and reset it connection | Perform Module Replacement Troubleshooting the UPS Perform Preventive Maintenance | All Modules are operational UPS Still not having an output voltage | Error Documented and sent to authorities for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921712705029958282/5921712705029958282_signature_28.png | Manuel Bacerra III | https://www.jotform.com/uploads/macarewecare/222927978372067/5921712705029958282/5921712705029958282_signature_31.png
| 2024-05-22 | 2024-05-22 14:50 | Almanza Uno Health Center | Alabang-Zapote Road, Almanza Uno, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28213 | PBSP | milcerpl@gmail.com | 24 |  |  | GXIV-D-10C | 120000486 | Unit is due for Installation. | Installation. | Checked Outlet Voltage. Checked Stability of the Table. Installed GeneXpert System. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921704290615918098/5921704290615918098_signature_28.png | Milcer P. Lladoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5921704290615918098/5921704290615918098_signature_31.png
| 2024-05-22 | 2024-05-22 10:41 | San Dionisio Health Center | Pelaez St., San Dionisio, Parańaque | NCR | NCR 4th - Parañaque | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28212 | DOH | marubi1019@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2017 |  | GXIV-D-6C | 831898 | Modules 2 and 4 with Temperature Error. | Module Replacement. | Replaced Modules 2 and 4. Performed Self-test. Cleaned Back Fan of Gx Unit. Generate Reports. | All Modules are Operational. | No UPS. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921692850614716115/5921692850614716115_signature_28.png | Fatima Joyce Tresvalles | https://www.jotform.com/uploads/macarewecare/222927978372067/5921692850614716115/5921692850614716115_signature_31.png
| 2024-05-22 | 2024-05-22 13:24 | Manila Central University | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | lab_mcuhospital@yahoo.com |  | 28211 | CONSORTIUM | purchasing@mcufdtmf.org | 22 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 309 |  | GXIV | 110018008 | Unit for maintenance | NA | #NAME? | Unit is operational No issues found |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921658864228129587/5921658864228129587_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5921658864228129587/5921658864228129587_signature_31.png
| 2024-05-22 | 2024-05-22 12:27 | Bataan Genaral Hospital and Medical Center | Balanga Bataan | 3 | Bataan | PCR Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 28210 | PBSP | mark10venzon@gmail.com | 26 | ASSAY: Xpert HBV VL, COUNT: 17 ASSAY: Xpert SARS-CoV-2, COUNT: 10282 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Others, COUNT: 179 |  | GXIV-D | 834433 | C1:686924 C4: 664590 For module replacement | Module replacement Preventive maintenance | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational C1: 914919(new) C4: 784107(new) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921645507969742420/5921645507969742420_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5921645507969742420/5921645507969742420_signature_31.png
| 2024-05-22 | 2024-05-22 11:55 | Sumilang Super Health Center | Dr. Garcia St. Sumilang, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28209 | PBSP | maricor.obnamia05@gmail.com | 20 |  |  | GX-IV-D-10C | 120000495 | For installation | Installed the GeneXpert System | Installed the GeneXpert System: -GeneXpert unit -Workstation -Barcode Scanner -Printer -UPS Functionality test Generated reports | Unit is operational. UPS not operational, for replacement. | UPS SN: SKPHU2K230910081 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921604635032632286/5921604635032632286_signature_28.png | Maria Corazon Obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/5921604635032632286/5921604635032632286_signature_31.png
| 2024-05-22 | 2024-05-22 11:36 | Potrero Health Center | Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28206 | PBSP | ailyn.santos19@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2295 ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 |  | GXIV | 804715 | #NAME? | NA | #NAME? | Unit is operational B2, B4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921594366124799652/5921594366124799652_signature_28.png | Ailyn Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5921594366124799652/5921594366124799652_signature_31.png
| 2024-05-22 | 2024-05-21 11:20 | San Jose City Health Center VI | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28208 | PBSP | theadelrosario1@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2551 ASSAY: Xpert HIV-1 Viral Load, COUNT: 129 |  | GXIV | 807620 | CPU for repair Unit for calibration Unit for maintenance | #NAME? | #NAME? | Unit is operational C1: retry C2: passed C3: passed C4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921594376124720836/5921594376124720836_signature_28.png | Thea Del Rosario | https://www.jotform.com/uploads/macarewecare/222927978372067/5921594376124720836/5921594376124720836_signature_31.png
| 2024-05-22 | 2024-05-21 11:26 | San Jose Del Monte City Health Center VI | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28207 | PBSP | theadelrosario1@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2655 ASSAY: Xpert HIV-1 Viral Load, COUNT: 154 |  | GXIV | 814161 | Unit for calibration Unit for maintenance | Cleaned filter | #NAME? | Unit is operational B1: passed B2: passed B3: passed B4: retry |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921594376128300709/5921594376128300709_signature_28.png | Thea Del Rosario | https://www.jotform.com/uploads/macarewecare/222927978372067/5921594376128300709/5921594376128300709_signature_31.png
| 2024-05-22 | 2024-05-22 11:34 | Belmonte Health Center | Sobriedad St. Sampaloc Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28204 | PBSP | maricar.lira@gmail.com | 25 |  | 2024-05-22 | GXIV | 808619 | Unit for preventive maintenance. For installation of new monitor. | Preventive maintenance and installation of monitor. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed installation of monitor. Generated reports. | Unit is operational. |  | Dell monitor (CN-0WMYC5-FCC00-34R-ALFX) in replacement to defective HP monitor (CNK629119F). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921590280966229695/5921590280966229695_signature_28.png | Maricar Lira | https://www.jotform.com/uploads/macarewecare/222927978372067/5921590280966229695/5921590280966229695_signature_31.png
| 2024-05-22 | 2024-05-22 11:42 | Belmonte Health Center | Sobriedad St. Sampaloc Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28205 | PBSP | maricar.lira@gmail.com | 25 |  |  | GXIV | 814376 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921590280969628710/5921590280969628710_signature_28.png | Maricar Lira | https://www.jotform.com/uploads/macarewecare/222927978372067/5921590280969628710/5921590280969628710_signature_31.png
| 2024-05-22 | 2024-05-22 10:45 | Cebu TB reference Lab 838882 | SRP cebu city | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28203 | PBSP | asmantiquilla.pbsp@gmail.com | 28 |  |  | Gx4 | 838882 | Due date for Xpertcheck calibration | Xpertcheck calibration for module B1,B2 &B4 Gather xpertcheck data | Xpertcheck calibration for module B1,B2 &B4 Gather xpertcheck data | All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921569524315680515/5921569524315680515_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5921569524315680515/5921569524315680515_signature_31.png
| 2024-05-22 | 2024-05-22 10:52 | Cebu TB Reference Lab 808620 | SRP - Cebu City | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28202 | PBSP | asmantiquilla.pbsp@gmail.com | 28 |  |  | Gx4 | 808620 | Due Date for Xpertcheck Calibration for Module B1,B2,B3 and B4 | Xpertcheck Calibration for Module B1,B2,B3 and B4 Gathering data of Xpertcheck calibration | Xpertcheck Calibration for Module B1,B2,B3 and B4 Gathering data of Xpertcheck calibration | Module B1, B3 and B4 Service require |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921562004316379502/5921562004316379502_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5921562004316379502/5921562004316379502_signature_31.png
| 2024-05-22 | 2024-05-22 09:34 | Tondo Medical Center | H. Lopez Blvd. Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28201 | PBSP | microscan@tmc.doh.gov.ph | 25 |  |  | GX16 | 844121 | Unit for preventive maintenance. | Preventive Maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921557110963836855/5921557110963836855_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5921557110963836855/5921557110963836855_signature_31.png
| 2024-05-22 | 2024-05-22 09:36 | Tondo Medical Center | H. Lopez Blvd. Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28199 | PBSP | microscan@tmc.doh.gov.ph | 25 |  |  | GXIV | 834430 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921557110965448189/5921557110965448189_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5921557110965448189/5921557110965448189_signature_31.png
| 2024-05-22 | 2024-05-22 09:16 | Tondo Medical Center | H. Lopez Blvd. Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28200 | DOH | microscan@tmc.doh.gov.ph | 25 |  | 2024-05-01 | GXIV | 831872 | Unit for Preventive maintenance. For module replacement of module A4. | Preventive maintenanxe and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A4 (686765) in replacement to defective A4 (688318). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921557110965855500/5921557110965855500_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5921557110965855500/5921557110965855500_signature_31.png
| 2024-05-22 | 2024-05-22 08:35 | Tuburan Rural Health Unit | Pobalacion Tuburan Cebu | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 28198 | PBSP | tuburanlaboratoryrhu1@gmail.com | 23 |  | 2024-05-22 | Gx4 | 815485 | Module B2 with serial number 717505 no calibration data. Last calibration 05/09/2024 Module B1, 3,4 passed the calibration. | Save all files. | .Perform preventive maintenance for the GeneXpert machine. .Save Installation Qualification .Save GXC file .Save System Log | Module B1, B3., B4 ok to use Module B2 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921491363559718883/5921491363559718883_signature_28.png | MARIEL R. GRACIA | https://www.jotform.com/uploads/macarewecare/222927978372067/5921491363559718883/5921491363559718883_signature_31.png
| 2024-05-22 | 2024-05-21 14:01 | De los Santos Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28197 | CONSORTIUM | molecularlab@dlsmc.ph | 30 | ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 227 |  | GX4D (10C) | 110020700 | Transfer from microbiology to molecular lab | Transfer unit Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Transfer and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921460063617197971/5921460063617197971_signature_28.png | Mary Grace Cabales | https://www.jotform.com/uploads/macarewecare/222927978372067/5921460063617197971/5921460063617197971_signature_31.png
| 2024-05-22 | 2024-05-21 16:10 | Dr. Amado Diaz Provincial Foundation Hospital | Midsayap, North Cotabato | 12 | Cotabato | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28196 | PBSP | dadpfh@yahoo.com | 34.1 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 118 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1030 | 2024-05-21 | Gx-IV | 807623 | Module 1 for Service Required Genexpert machine temperature higher or not within the acceptable range of 15°C to 30°C | Replace defective module B2 SN: 673498 with new module SN: 716967 Perform Genexpert Xpertcheck/Calibration Genexpert Cleaning Checking of errors Data collection | Genexpert machine Preventive maintenance Genexpert module replacement/repair Genexpert machine Xpertcheck/Calibration | 4 operational modules |  | Genexpert module 6-color SN:716967 | Module SN:716967 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921422267289145475/5921422267289145475_signature_28.png | Ermariz P. Jaducana | https://www.jotform.com/uploads/macarewecare/222927978372067/5921422267289145475/5921422267289145475_signature_31.png
| 2024-05-21 | 2024-05-16 08:58 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28195 | DOH | rhupozorrubio@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 663 |  | GXIV-4-D-6C | 838467 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | - Unit is partially operational - Module B1 (861502) has already been reported |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921024852011581043/5921024852011581043_signature_28.png | Jinky Grace Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/5921024852011581043/5921024852011581043_signature_31.png
| 2024-05-21 | 2024-05-15 11:11 | Rosales Rural Health Unit | Pine St., Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28194 | DOH | allan.vinluan2013@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1003 |  | GXIV-4-D-6C | 838472 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921023552011345132/5921023552011345132_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5921023552011345132/5921023552011345132_signature_31.png
| 2024-05-21 | 2024-05-15 10:07 | Umingan Community Hospital | Burgos St. Poblacion West, Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28193 | DOH | jeffreyvrumbaoa@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1194 |  | GXIV-4-D-6C | 838473 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921020502016514128/5921020502016514128_signature_28.png | Jeffrey Rumbaoa | https://www.jotform.com/uploads/macarewecare/222927978372067/5921020502016514128/5921020502016514128_signature_31.png
| 2024-05-21 | 2024-05-14 10:33 | San Manuel Rural Health Unit | Ambassador Ramos Guiset Sur, San Manuel, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28192 | DOH | serly_reyes@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 844 |  | GXIV-4-D-6C | 838516 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5921019462016272717/5921019462016272717_signature_28.png | Serly Galamgam | https://www.jotform.com/uploads/macarewecare/222927978372067/5921019462016272717/5921019462016272717_signature_31.png
| 2024-05-21 | 2024-05-21 15:04 | Bayugan City Health Office | Brgy. Poblacion, Bayugan | 13 | Agusan del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28191 | PBSP | lourdescayao@gmail.com | 21.26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3191 |  | GX-4 6 colors | 815391 | B2 an B4 reported temperature error. | Module Replacements | Replacing B2 with SN: 901636 to SN: 641028 Replacing B4 with SN: 612924 to SN: 773241 Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920902832124137916/5920902832124137916_signature_28.png | Lourdes Diola | https://www.jotform.com/uploads/macarewecare/222927978372067/5920902832124137916/5920902832124137916_signature_31.png
| 2024-05-21 | 2024-05-21 16:19 | St. Luke's Medical Center - Global City | Rizal Dr. BGC, Taguig | NCR | NCR 4th - Taguig | Laboratory | YES | REPAIR | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 28190 | COMMERCIAL | vrrdy@stluke.com.ph | 24 |  | 2024-05-16 | GX-IV-D-6C | 839231 | Module C1 for replacement | Replaced module C1 | Performed maintenance Replaced module C1 Functionality test Generated reports | Unit is operational. All modules operational. |  | Old: C1/866691 New: C1/210124694 | SI# 098031 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920892105294229139/5920892105294229139_signature_28.png | Vincent-Raquiza Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/5920892105294229139/5920892105294229139_signature_31.png
| 2024-05-21 | 2024-05-21 15:48 | St. Luke's Medical Center - Global City | Rizal Dr. BGC, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 28189 | COMMERCIAL | vrrdy@stluke.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 179 ASSAY: Xpert CT/NG, COUNT: 1055 ASSAY: Xpert HBV VL, COUNT: 77 ASSAY: Xpert HPV, COUNT: 1019 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2629 ASSAY: Xpert HCV VL, COUNT: 44 ASSAY: Xpert HIV-1 Viral Load, COUNT: 50 |  | GX-IV-L-6C | 808411 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module B4/688662 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920892105293895634/5920892105293895634_signature_28.png | Vincent Raquiza Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/5920892105293895634/5920892105293895634_signature_31.png
| 2024-05-21 | 2024-05-21 15:57 | Philippine Red Cross QC chapter | Kalayaan st., QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28188 | COMMERCIAL | prcqc.bloodservices@gmail.com | 26 |  |  | MCS + | 02A227 | CRC bios failed | Performed CPU board reset. Enter bypass mode to access calibration parameters but unable to change the value. encounter auto restart and crc bios failed appear on display. Restart multiple time but error still occur | Troubleshooting | Unit unoperational For replacement of cpu board |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920878948014502084/5920878948014502084_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/5920878948014502084/5920878948014502084_signature_31.png
| 2024-05-21 | 2024-05-21 14:50 | Lanuza Health Center | Isagani St. Sta Cruz | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28187 | PBSP | kikokendz_060830@yahoo.com | 25 |  | 2024-05-21 | GXIV | 805111 | Unit for preventive maintenance. Unit for module replacement. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B2 (692028), B3 (694168) and B4 (714257). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920840110841673783/5920840110841673783_signature_28.png | Eva jone | https://www.jotform.com/uploads/macarewecare/222927978372067/5920840110841673783/5920840110841673783_signature_31.png
| 2024-05-21 | 2024-05-21 14:46 | Dimasalang Health Center | Isagani St. Sta Cruz | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28186 | PBSP | kikokendz_060830@yahoo.com | 25 |  | 2024-05-21 | GXIV | 707072 | Unit for preventive maintenance. For module replacement of modules B1, B2, B3 and B4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX modules B1 (687294), B2 (688240, B3 (688926) and B4 (689004). | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920839400849637642/5920839400849637642_signature_28.png | Maria joy cleofas m. Abenales | https://www.jotform.com/uploads/macarewecare/222927978372067/5920839400849637642/5920839400849637642_signature_31.png
| 2024-05-21 | 2024-05-21 14:42 | Mexico RHU I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28185 | DOH | jasperdizon20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 534 |  | GXIV-6C | 831877 | Due for Quarterly Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920834627618683299/5920834627618683299_signature_28.png | Jasper Dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5920834627618683299/5920834627618683299_signature_31.png
| 2024-05-21 | 2024-05-21 14:35 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28184 | DOH | burauenlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 457 | 2024-05-21 | GXIV | 838551 | Fan Filter | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920829744964099569/5920829744964099569_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5920829744964099569/5920829744964099569_signature_31.png
| 2024-05-21 | 2024-05-21 14:06 | Josefina Belmonte Duran Albay Provincial Hospital | Tuburan, Ligao City, Albay | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | salvadorlyndonrempillo@gmail.com |  | 28183 | PBSP | salvadorlyndonrempillo@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 689 |  | GX4-6C | 814371 | B4/210092937 for module replacement | Replaced Module: B4 - 210092937>>711296 Cleaned cartridge bays, Plunger rod, Icore Generate reports | Replaced Module 210092937 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920825484023850137/5920825484023850137_signature_28.png | Lirio Llorin | https://www.jotform.com/uploads/macarewecare/222927978372067/5920825484023850137/5920825484023850137_signature_31.png
| 2024-05-21 | 2024-05-21 14:23 | Philippine Red Cross Zamboanga | Pettit baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 28182 | COMMERCIAL | nimfalinao2006@yahoo.com | 0 |  | 2024-05-21 | MCS + | 01M256 | Haemonetics machine unable to turn on | Replaced Power cord of the machine | Troubleshoot and replaced power cord of the machine. | Available and running | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920821283315977617/5920821283315977617_signature_28.png | Nimfa A. Linao | https://www.jotform.com/uploads/macarewecare/222927978372067/5920821283315977617/5920821283315977617_signature_31.png
| 2024-05-21 | 2024-05-21 14:23 | Sison Rural Health Unit | Sison Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28181 | PBSP | sisonrhu@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 391 | 2024-04-26 | GXIV-D | 834684 | B3:784437 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational B3:885220 new module |  | 1 genexpert module | 47661 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920820377632063298/5920820377632063298_signature_28.png | Jeannette O. Alcaide | https://www.jotform.com/uploads/macarewecare/222927978372067/5920820377632063298/5920820377632063298_signature_31.png
| 2024-05-21 | 2024-05-21 12:30 | Cardinal Santos Medical Center | Wilson St. San Juan | NCR | NCR 2nd - San Juan | Laboratory | YES | REPAIR | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 28180 | COMMERCIAL | microbiology.csmc@gmail.com | 24 | ASSAY: Xpert HBV VL, COUNT: 132 ASSAY: Xpert HPV, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1095 | 2024-05-16 | GX-IV-D-10C | 110003714 | Module B1 and B4 for replacement | Replaced modules B1 and B4 | Performed maintenance Replaced modules B1 and B4 Functionality test Generated reports | Unit is operational. All modules operational. |  | Old: B1/210022152, B4/210048235 New: B1/220023301, B4/210113679 | SI# 098028, 098027 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920778060943848324/5920778060943848324_signature_28.png | Ellaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5920778060943848324/5920778060943848324_signature_31.png
| 2024-05-21 | 2024-05-21 11:55 | Cadiz City Health Office | Cabahug Street, Cadiz City | 6 | Negros Occidental | TB DOTS | YES | INSTALLATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 28179 | PBSP | zhevebacomo@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1817 |  | GX4 | 807614 | Defective module for D2 with SN 808098. | Conduct PM and save files (IQR, System Log and assay statistics). | Replaced the defective module for D2 with new module SN 784326. | All modules are now functional and in good running condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920749668812767317/5920749668812767317_signature_28.png | Zherlyn Eve Bacomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5920749668812767317/5920749668812767317_signature_31.png
| 2024-05-21 | 2024-05-21 12:01 | Don Juan C. Cojuangco Memorial Health Center | Cuyapo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28178 | DOH | jolenerodico09171998@gmsil.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV-D | 831672 | Due for maintenance UPS no power/not charging | Preventive maintenance UPS Troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational Ups for replacement | UPS SN:AT1000181128ST0186 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920745556402853370/5920745556402853370_signature_28.png | John Esteban | https://www.jotform.com/uploads/macarewecare/222927978372067/5920745556402853370/5920745556402853370_signature_31.png
| 2024-05-21 | 2024-05-21 12:07 | Camalig Rural Health Unit | Brgy. 2, Camalig, Albay | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28177 | PBSP | leilanigrageda@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1032 |  | GX4-6C 2placet | 835187 | Date and time reset when booting PC | Replaced CMOS Battery of PC for date and time Cleaned cartridge bays, plunger rod, icore Generate report | Performed Troubleshooting Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920742484021748150/5920742484021748150_signature_28.png | Leilani M. Grageda | https://www.jotform.com/uploads/macarewecare/222927978372067/5920742484021748150/5920742484021748150_signature_31.png
| 2024-05-21 | 2024-05-21 11:28 | Chinese General Hospital | Blumentritt, Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28176 | CONSORTIUM | ktringler@gmail.com | 25 | ASSAY: Xpert C. difficile BT, COUNT: 19 ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Xpert HBV VL, COUNT: 69 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1625 ASSAY: Xpert SARS-CoV-2, COUNT: 378 ASSAY: Xpert HCV VL, COUNT: 29 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GXIV | 835604 | Unit for relocation from bacteriology to molecular laboratory. Unit for preventive maintenance. | Relocation, installation and preventive maintenance. | Performed relocation and installation of Genexpert. Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920716740319552442/5920716740319552442_signature_28.png | KAREN RINGLER | https://www.jotform.com/uploads/macarewecare/222927978372067/5920716740319552442/5920716740319552442_signature_31.png
| 2024-05-21 | 2024-05-21 10:59 | Davao TB Reference Laboratory | DOH Compound, JP Laurel Davao City | 11 | Davao del Sur | TB-DOTS | YES | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28175 | WHO | dtrlmark.pbsp@gmail.com | 29.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1870 ASSAY: Others, COUNT: 386 |  | GX-IV 4D 10C | 848248 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log | Perform Preventive Maintenance | All Modules are operational | Other Assay : XDR, High Room Temp, end user advise to regulate room temperature optimal for GeneXpert Usages |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920701406381725668/5920701406381725668_signature_28.png | Joseph Mark B. Anque | https://www.jotform.com/uploads/macarewecare/222927978372067/5920701406381725668/5920701406381725668_signature_31.png
| 2024-05-21 | 2024-05-21 10:44 | Davao TB Reference Laboratory | DOH Compound, JP Laurel Davao City | 11 | Davao del Sur | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28174 | PBSP | dtrlmark.pbsp@gmail.com | 29.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1077 |  | GX-IV 4L R1 | 707513 | Modules B2, B3 and B4 disabled by end-user due to Multiple Temperature Error | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Replace the defective Module SN:661183 with new Module SN:711446 Replaced the Defective Module SN: 610307 with new Module SN: 719867 Generate IQR, Assay Stat and System Log Archived Data test results | Perform Module Replacement, Troubleshooting and Preventive Maintenance | Module B2 and B3 successfully replaced Module Module B4 disabled and error datas will be sent to Cepheid for further Assessment | High Room Temperature, End-user is advised to regulate the optimal room temperature |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920695276384057663/5920695276384057663_signature_28.png | Joseph Mark B. Anque | https://www.jotform.com/uploads/macarewecare/222927978372067/5920695276384057663/5920695276384057663_signature_31.png
| 2024-05-21 | 2024-05-21 10:27 | Eversly Childs Sanitarium | Jagobiao, mandaue City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28173 | PBSP | alimurungjane@gmail.com | 28 |  |  | Gx 4 | 805270 | Defective module A1 644504 | New module Installed A1 678236 Clean Fan Filter Gather data IQR, system log | New module Installed A1 678236 Clean Fan Filter Gather data IQR, system log | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920685162125941113/5920685162125941113_signature_28.png | Juanita v. Alimurung | https://www.jotform.com/uploads/macarewecare/222927978372067/5920685162125941113/5920685162125941113_signature_31.png
| 2024-05-20 | 2024-05-20 16:10 | Veterans Memorial Medical Center | North Avenue, Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Mole | NO | TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28172 | MACARE-HBDC | vmmc.moleclab@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 182 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1576 ASSAY: Xpert SARS-CoV-2, COUNT: 110 ASSAY: Xpert HCV VL, COUNT: 17 |  | GX IV 6C-D | 839150 | Modules not detected in GeneXpert DX Software since lightning yesterday | Troubleshoot GeneXpert Checked UPS Checked GX Fuse Checked gateway board Checked modules Generated report | Outlet power too low 195V but UPS regulates the power to 225V UPS is working GX fuse intact Tried switching gateway board still problem persists Tried using the modules for office Module from office detected. Modules in GX still not detected Tried switching positions Still module from office is detected Gateway still works Lightning surge hits the network cables and flow to the machine Hence, all modules damaged | UPS Ok Desktop OK Gateway OK Power Supply OK 4 GX Modules defective | Needs proper grounding of data cables and network cables and avoid putting cables on tree branches, as they are more likely to be hit by lightning. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920033581371443106/5920033581371443106_signature_28.png | Gerard GhanivC. Sidic | https://www.jotform.com/uploads/macarewecare/222927978372067/5920033581371443106/5920033581371443106_signature_31.png
| 2024-05-20 | 2024-05-20 14:39 | Makati Medical Center (Medilinx) | Amorsolo st., Makati City | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28171 | COMMERCIAL | reinaluz.bactol@medilinx.com.ph | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 187 ASSAY: Xpert CT/NG, COUNT: 800 ASSAY: Xpert HPV, COUNT: 787 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3042 ASSAY: Others, COUNT: 17 ASSAY: Xpert HCV VL, COUNT: 81 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1912 |  | GX4 D(10C) | 110001647 | As per end user they encounter a 5007 error in B1 210076418 and B3 210061783 in mtb rif ultra | Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Backup and purge nxx files. Test run unit | Troubleshooting | B1 and B3 for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920022601614766938/5920022601614766938_signature_28.png | Maria Reinaluz Bactol | https://www.jotform.com/uploads/macarewecare/222927978372067/5920022601614766938/5920022601614766938_signature_31.png
| 2024-05-20 | 2024-05-20 16:09 | Estevez Memorial Hospital, Inc. | 4PQR+67P, Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City, 4500 Albay | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28170 | COMMERCIAL | melosept22@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 24 |  | Gx4-10C | 110017942 | Gx for Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5920021798916162268/5920021798916162268_signature_28.png | Estevez memorial hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/5920021798916162268/5920021798916162268_signature_31.png
| 2024-05-20 | 2024-05-15 15:02 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 28169 | MACARE-HBDC | mbalquen@gmail.com | 28 |  |  | MCS+ | 17H632SPG | As per end user they encounter error message 108 network configuration | Checked unit Checked parameters for network configuration. Test run unit | Troubleshooting | Endorse operator to press the YES once the error message 108 appear.for further troubleshooting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919927021611933813/5919927021611933813_signature_28.png | Mary Jane Amores | https://www.jotform.com/uploads/macarewecare/222927978372067/5919927021611933813/5919927021611933813_signature_31.png
| 2024-05-20 | 2024-05-15 14:09 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 28168 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2919 ASSAY: Xpert SARS-CoV-2, COUNT: 3688 ASSAY: Others, COUNT: 163 |  | GX16 D(10C) | 110014171 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919926771612717357/5919926771612717357_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5919926771612717357/5919926771612717357_signature_31.png
| 2024-05-20 | 2024-05-15 09:36 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28167 | MACARE-HBDC | roseannparnes@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1472 |  | GX4 D (10C) | 110014085 | Due for maintenance B4 module reported already | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | 3 modules operational 1 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919926281612852489/5919926281612852489_signature_28.png | Rose Ann Parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5919926281612852489/5919926281612852489_signature_31.png
| 2024-05-20 | 2024-05-15 08:50 | East Avenue Medical Center | East Avenue,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 28166 | MACARE-HBDC | eamcmolecular@gmail.com | 22 | ASSAY: Xpert CT/NG, COUNT: 31 ASSAY: Xpert HPV, COUNT: 244 ASSAY: Xpert SARS-CoV-2, COUNT: 1220 |  | GX4 D(10C) | 110016781 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919925961619764396/5919925961619764396_signature_28.png | East Avenue Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5919925961619764396/5919925961619764396_signature_31.png
| 2024-05-20 | 2024-05-17 12:39 | Talisay Rural Health Unit | Poblacion 3, Talisay Batangas | 4A | Batangas | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28165 | DOH | rodriguezadett@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 972 |  | GXIV | 820201 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919914563611212414/5919914563611212414_signature_28.png | Adettlou A.Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5919914563611212414/5919914563611212414_signature_31.png
| 2024-05-20 | 2024-04-19 13:44 | Hagonoy Rural Health Unit | Hagonoy, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28164 | PBSP | lanielyncastro@yahoo.com | 27 |  |  | GXIV | 814158 | CPU for troubleshooting | Replaced BIOS battery | #NAME? | Unit is not operational Unit for follow up reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919902123615903146/5919902123615903146_signature_28.png | Lanielyn Castro Abadam | https://www.jotform.com/uploads/macarewecare/222927978372067/5919902123615903146/5919902123615903146_signature_31.png
| 2024-05-20 | 2024-05-03 14:07 | Plaridel Rural Health Unit | Plaridel, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28163 | DOH | Yellei0908@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-05-03 | GXIV | 831711 | Module A2 for replacement | Replaced module A2 by 684364 | #NAME? | Unit is operational |  | GemeXpert Module SN: 684364 | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919862223616239993/5919862223616239993_signature_28.png | Ahlelei Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5919862223616239993/5919862223616239993_signature_31.png
| 2024-05-20 | 2024-05-07 11:40 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28162 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1288 |  | GXIV | 831691 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919861933616861251/5919861933616861251_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5919861933616861251/5919861933616861251_signature_31.png
| 2024-05-20 | 2024-05-14 13:49 | Bulakan Rural Health Unit | Bulakan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28161 | DOH | l_gemino@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1194 ASSAY: Xpert HIV-1 Viral Load, COUNT: 97 |  | GXIV | 831641 | Unit for maintenance | Cleaned filter | #NAME? | Unit is operational Module A1 under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919861633614304098/5919861633614304098_signature_28.png | Lilibeth Gemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5919861633614304098/5919861633614304098_signature_31.png
| 2024-05-20 | 2024-05-14 15:38 | Marilao Rural Health Unit | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28160 | DOH | jmjenilo27@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1546 ASSAY: Xpert HIV-1 Viral Load, COUNT: 110 |  | GXIV | 831897 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919859583616190226/5919859583616190226_signature_28.png | Christopher Dilag | https://www.jotform.com/uploads/macarewecare/222927978372067/5919859583616190226/5919859583616190226_signature_31.png
| 2024-05-20 | 2024-05-15 13:01 | Kamuning Health Center | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28159 | DOH | vicvic.chua25@yahoo.com | 27 | ASSAY: Xpert HPV, COUNT: 220 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1496 |  | GXIV | 831612 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919854993617271572/5919854993617271572_signature_28.png | Mary Rose Perejon | https://www.jotform.com/uploads/macarewecare/222927978372067/5919854993617271572/5919854993617271572_signature_31.png
| 2024-05-20 | 2024-05-16 13:12 | San Rafael Rural Health Unit | San Rafael, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28158 | DOH | tatsdayrit25@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 894 |  | GXIV | 831692 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919854673614545070/5919854673614545070_signature_28.png | Joana Manelle Dayrit | https://www.jotform.com/uploads/macarewecare/222927978372067/5919854673614545070/5919854673614545070_signature_31.png
| 2024-05-20 | 2024-05-14 09:04 | Urdaneta Sacred Heart Hospital | Mc Arthur Highway, San Vincente, Urdaneta City | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28157 | COMMERCIAL | annabellelcatbagan@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 190 |  | GXII-2D | 820303 | For Calibration and Preventive Maintenance | #NAME? | Performed Calibration and Preventive Maintenance | - Unit is partially operational - Module A1(720229) passed Retry of XpertCheck - Module A2(720228) failed last XpertCheck; For Reporting | Replacing of Module A2(720228) is Recommended |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919838912018433669/5919838912018433669_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5919838912018433669/5919838912018433669_signature_31.png
| 2024-05-20 | 2024-05-17 12:32 | Meycauayan City Health Center VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28156 | DOH | ronamarieiris_25@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1716 ASSAY: Xpert HIV-1 Viral Load, COUNT: 138 |  | GXIV | 831904 | Temperature error on A3 | NA | #NAME? | Unit is operational A3 is under observation A3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919834793618670569/5919834793618670569_signature_28.png | Rona Marie Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5919834793618670569/5919834793618670569_signature_31.png
| 2024-05-20 | 2024-05-17 12:02 | Dasmariñas City Health Office 2 | Brgy Sta.Cruz, DBB-E, Dasmariñas City, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28155 | DOH | hazelmilan0777@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1971 | 2024-05-02 | GXIV-D-6C | 838535 | Unit is due for XpertCheck | XpertCheck. Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked if Filter is Clogged and Replaced it. Performed XpertCheck. Performed Self-test. Generate Reports. | All Modules are Operational. | 2 Modules Passed, 2 Newly Replaced. | XpertCheck Kit | DR 11455 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919782553617932076/5919782553617932076_signature_28.png | Hazel M. Tabucol | https://www.jotform.com/uploads/macarewecare/222927978372067/5919782553617932076/5919782553617932076_signature_31.png
| 2024-05-20 | 2024-05-17 14:31 | Angono Rural Health Unit | Angono, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28154 | DOH | siennablanco@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1164 |  | GX-IV-D-6C | 819864 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919744043612339229/5919744043612339229_signature_28.png | Sienna Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5919744043612339229/5919744043612339229_signature_31.png
| 2024-05-20 | 2024-05-17 10:48 | Cainta City Health Office | Karangalan Dr., Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28153 | DOH | ballesterrichard18@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 ASSAY: Xpert HIV-1 Viral Load, COUNT: 87 |  | GX-IV-D-6C | 831623 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Module A1/802425 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5919743873614046440/5919743873614046440_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5919743873614046440/5919743873614046440_signature_31.png
| 2024-05-17 | 2024-05-17 15:47 | Cordlife Medical Philippines | UP-Ayala Technohub, Commonwealth Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CYTIVA | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 28152 | CONSORTIUM | patricia.perfecto@cordlife.com | 25 |  |  | Smartmax | 5131.3196 | Unit unable to boot | Reset BIOS settings on CPU board directly Performed functionality test Checked calibration parameters Tested temperature settings Tested air bags Tested pumps | Reset BIOS settings | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5917418395636097923/5917418395636097923_signature_28.png | Patricia Perfecto | https://www.jotform.com/uploads/macarewecare/222927978372067/5917418395636097923/5917418395636097923_signature_31.png
| 2024-05-17 | 2024-05-17 13:39 | Sariaya Rural Health Unit | Brgy. Poblacion 4, Sariaya, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28151 | DOH | smcl_2012@yahoo.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2035 | 2024-04-25 | GX IV 6C-D | 831890 | Module A4 for replacement | Checked and cleaned cartridge bays, plungger rod icore slot and filter Replaced module A4 Performed reliability test Generated report | Replaced Module A4 210091100 -> 611836 Performed Preventive maintenance | Unit is operational All modules working |  | GX module 6C | DR 0047630 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5917345583313955195/5917345583313955195_signature_28.png | rhea parreño | https://www.jotform.com/uploads/macarewecare/222927978372067/5917345583313955195/5917345583313955195_signature_31.png
| 2024-05-17 | 2024-05-17 12:41 | Southern Philippines Medical Center | JP Laurel, Davao City | 11 | Davao del Sur | ICU2 | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28150 | COMMERCIAL | icu2estuya@gmail.com | 24.6 |  |  | MCS+ | 15H034G | Machine suddenly forcing shutdown | Reset all board connection Check voltages Replaces the PSU with a spare PSU to isolate the issue Reset Connections | Perform troubleshooting | Machine is still forcing shutdown Machine is not operational | Waiting for spareparts from Manila to Isolate the Issue |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5917308766248593087/5917308766248593087_signature_28.png | Noel N. Ligan | https://www.jotform.com/uploads/macarewecare/222927978372067/5917308766248593087/5917308766248593087_signature_31.png
| 2024-05-17 | 2024-05-17 11:02 | Cuenca Rural Health Unit | Cuenca Municipal Compound, Cuenca Batangas | 4A | Batangas | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28149 | DOH | ligayamijares50@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1176 |  | GXIV | 821297 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5917248671501290579/5917248671501290579_signature_28.png | Ligaya R. Mijares | https://www.jotform.com/uploads/macarewecare/222927978372067/5917248671501290579/5917248671501290579_signature_31.png
| 2024-05-17 | 2024-05-15 09:16 | Kawit Rural Health Unit | Brgy Combalay, Binakayan, Kawit , Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28147 | DOH | eleanor.perona@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 877 |  | GXIV-D-6C | 820705 | GX System has No Power | Troubleshoot | Checked Outlet Voltage. Checked UPS Output Voltage. Cleaned Plunger Rod, Cartridge Bay and Icore Optics. Checked and Cleaned Filter. Performed Self-test. Generate Reports. | UPS is defective and needs to be replaced. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5917065237612022753/5917065237612022753_signature_28.png | Eleanor C.Perona | https://www.jotform.com/uploads/macarewecare/222927978372067/5917065237612022753/5917065237612022753_signature_31.png
| 2024-05-17 | 2024-05-15 11:29 | Rosario Rural Health Unit | Scout Torillo, Rosario, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28148 | DOH | arlenepenaarmijo@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1082 | 2024-05-02 | GXIV-D-6C | 838537 | Unit is due for XpertCheck. | XpertCheck | Cleaned Cartridge Bay, Plunger Rod,and Icore Optics. Checked Filter if Clogged. Performed XpertCheck. Performed Self-test. Generate Reports. | 1 Module Reported, 2 Passed and 1 Service Required in XpertCheck. | B3/732613 Failed | XpertCheck Kit | DR 11456 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5917065247611736318/5917065247611736318_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/5917065247611736318/5917065247611736318_signature_31.png
| 2024-05-16 | 2024-05-16 14:51 | Sta. Lucia Health Center | Sta. Lucia, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28146 | DOH | josiedingle2004@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2055 | 2024-05-13 | GX IV 6C-D | 831901 | Unit due fo calibration | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed xpertcheck on modules A1, A2 and A3 All 3 modules passed Module A4 new Archived results Generated report | Performed preventive maintenance Performed Xpertcheck | Unit is operational All modules working |  | XpertCheck Kit | Lot 06103 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916525735617151811/5916525735617151811_signature_28.png | Josephine L. Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5916525735617151811/5916525735617151811_signature_31.png
| 2024-05-16 | 2024-05-16 13:17 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28145 | DOH | buhanginlaboratory@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 946 |  | GX-IV 4D | 838232 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are working, machine is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916466579717437085/5916466579717437085_signature_28.png | Catherine B. Deocampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5916466579717437085/5916466579717437085_signature_31.png
| 2024-05-16 | 2024-05-16 11:45 | Nasipit Rural Health Unit | District 1 Tiga, Brgy. Talisay, Nasipit | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28144 | DOH | peterdolotina@gmail.com | 28.275 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1462 |  | GX-4 6 colors | 837611 | Filter not reusable | Filter reported to Jotform | Preventive Maintenance | All modules are functional All units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916428497957947967/5916428497957947967_signature_28.png | Peter T. Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/5916428497957947967/5916428497957947967_signature_31.png
| 2024-05-16 | 2024-05-16 11:49 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 28143 | DOH | lornsabaigat@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 834 | 2024-05-16 | GXIV | 838409 | n/a | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916407825311016043/5916407825311016043_signature_28.png | LORNA C.ABAIGAR | https://www.jotform.com/uploads/macarewecare/222927978372067/5916407825311016043/5916407825311016043_signature_31.png
| 2024-05-16 | 2024-05-16 11:09 | Chinese General Hospital | Blumentritt St. Manila | NCR | NCR 1st - Manila | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28142 | COMMERCIAL | cghmcbloodbank2016@gmail.com | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916394576426633863/5916394576426633863_signature_28.png | Joan Escolano | https://www.jotform.com/uploads/macarewecare/222927978372067/5916394576426633863/5916394576426633863_signature_31.png
| 2024-05-16 | 2024-05-16 11:08 | Santolan Super Health Center | Evangelista Ave., Santolan, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | Replacement | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28141 | MACARE-HBDC | samandlet17@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 9 | 2024-05-15 | GX-IV-D-10C | 110024245 | UPS for replacement | Replaced defective UPS | Installed new UPS Connected the GeneXpert System to the UPS Functionality test Checked voltage Generated reports | Unit is operational. |  | Old UPS SN: AT1K230830ST0024 New UPS SN: AT1K230830ST0021 | DR# 11498, under DR# 11478 and SI# 097968 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916389422511861105/5916389422511861105_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5916389422511861105/5916389422511861105_signature_31.png
| 2024-05-16 | 2024-05-16 09:54 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28140 | DOH | riazarizzamae@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 903 |  | GXIV-6C | 831637 | Module 1 error | Performed Module Replacement A1(756662 - 643503) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916343980414086553/5916343980414086553_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5916343980414086553/5916343980414086553_signature_31.png
| 2024-05-16 | 2024-05-16 09:54 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28139 | DOH | riazarizzamae@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 903 |  | GXIV-6C | 831637 | Module 1 error | Performed Module Replacement A1(756662 - 643503) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916343840418146410/5916343840418146410_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5916343840418146410/5916343840418146410_signature_31.png
| 2024-05-16 | 2024-05-16 09:38 | FEU - Dr. Nicanor Reyes Medical Foundation | Regalado Avenue, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28138 | CONSORTIUM | kmtejoso@feu-nrmf.edu.ph | 26 | ASSAY: Xpert HPV, COUNT: 62 ASSAY: Xpert MTB/RIF Ultra, COUNT: 422 |  | GX IV 6C-D | 840975 | Unit due for maintenance | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test All modules passed Generated report | Performed Preventive Maintenance Performed reliability test | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916341745619770094/5916341745619770094_signature_28.png | Eddyline Pecaoco | https://www.jotform.com/uploads/macarewecare/222927978372067/5916341745619770094/5916341745619770094_signature_31.png
| 2024-05-16 | 2024-05-16 08:33 | Bongabon District Hospital | Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28137 | DOH | tsantuajr0128@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 527 |  | GXIV-6C | 831562 | Due for Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5916290865316780214/5916290865316780214_signature_28.png | Joyce Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5916290865316780214/5916290865316780214_signature_31.png
| 2024-05-15 | 2024-05-15 12:40 | Bagumbayan Rural Health Unit | Poblacion Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28136 | DOH | larrozarey27@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 894 |  | Gx-IV | 838261 | CPU problem CPU hard start up problem | Troubleshoot CPU problem Genexpert machine Cleaning Checking of errors Problem reporting Data Collection | Perform Genexpert machine Preventive maintenance Perform CPU troubleshooting | 4 operational modules | Requesting for CPU power supply replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5915830155024310466/5915830155024310466_signature_28.png | Rey D Larroza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5915830155024310466/5915830155024310466_signature_31.png
| 2024-05-15 | 2024-05-15 12:30 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28135 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 03J030 | Unit for preventive maintenance. Unit for calibration. | Preventive maintenance and calibration. | Performed preventive maintenance. Checked and cleaned rotor pumps and the unit. Performed calibration. Verfied and adjusted calibration parameters. Performed simulation running. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5915605402561564328/5915605402561564328_signature_28.png | Lallaine Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5915605402561564328/5915605402561564328_signature_31.png
| 2024-05-15 | 2024-05-15 12:26 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28134 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. Unit for calibration. | Preventive maintenance and calibration. | Performed preventive maintenance. Checked and cleaned rotor pumps and the unit. Performed calibration. Verfied and adjusted calibration parameters. Performed simulation running. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5915605402564545863/5915605402564545863_signature_28.png | Lalaine ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5915605402564545863/5915605402564545863_signature_31.png
| 2024-05-15 | 2024-05-15 13:04 | Manila Med - Medical Center Manila | 850 United Nations Ave., Paco, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28133 | CONSORTIUM | microbiologylab.manilamed@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1097 |  | GX IV 6C-D | 835540 | Unit due for maintenance | Checked and cleaned cartridge bays, plunger rod, icoreslot and filter Performed reliability test All modules passed Generated report | Performed Preventive Maintenance Performed reliability test | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5915597115611003198/5915597115611003198_signature_28.png | FRANCES LEVINA VELARDE | https://www.jotform.com/uploads/macarewecare/222927978372067/5915597115611003198/5915597115611003198_signature_31.png
| 2024-05-15 | 2024-05-15 13:10 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28132 | DOH | marteetingzon20@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1164 | 2024-05-15 | GXIV | 838421 | Fan Filter | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5915594883811314570/5915594883811314570_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5915594883811314570/5915594883811314570_signature_31.png
| 2024-05-15 | 2024-05-15 11:11 | Kalinga Health Hub - Marikina | J.P. Rizal St. Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28131 | CONSORTIUM | lorenzo.cruz@innovationsch.org | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1118 |  | GX-IV-D-6C | 110008418 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5915526098484366313/5915526098484366313_signature_28.png | Lorenzo Bryll acruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5915526098484366313/5915526098484366313_signature_31.png
| 2024-05-15 | 2024-05-15 10:27 | Manila Doctors Hospital | United Nations Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28130 | COMMERCIAL | lab.bacteriology@maniladoctors.com.ph | 20 | ASSAY: Xpert CT/NG, COUNT: 106 ASSAY: Xpert HBV VL, COUNT: 92 ASSAY: Xpert HPV, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1243 ASSAY: Xpert HCV VL, COUNT: 10 ASSAY: Xpert HIV-1 Viral Load, COUNT: 124 |  | GX IV 10C-D | 110001648 | Unit due for maintenance | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test. All modules passed Generated report | Performed Preventive Maintenance Performed Relaiability Test | Unit is operational all modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5915502057408299036/5915502057408299036_signature_28.png | Apple Mae Cabildo | https://www.jotform.com/uploads/macarewecare/222927978372067/5915502057408299036/5915502057408299036_signature_31.png
| 2024-05-15 | 2024-05-14 10:04 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28129 | DOH | laureanjo@yahoo.com | 28 |  |  | Gx IV | 838527 | Monthly Preventive maintenance | Clean Fan Filter Clean Module Bay Plunger Maintenance Gather data IQR, System log, assay stats | Clean Fan Filter Clean Module Bay Plunger Maintenance Gather data IQR, System log, assay stats | Systems operational Module D4 for MR |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5915472491113622621/5915472491113622621_signature_28.png | Laurean jo Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5915472491113622621/5915472491113622621_signature_31.png
| 2024-05-14 | 2024-05-14 11:04 | Norala District Hospital | Poblacion Norala, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28128 | DOH | rezielrds0709@gmail.com | 23.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1091 |  | Gx-IV | 838543 | N/A | Genexpert machine Cleaning Checking of errors Data collection: IQR, S.LOG, ASSAY stat. and Archive test | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5915043765027897857/5915043765027897857_signature_28.png | Reziel D. Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/5915043765027897857/5915043765027897857_signature_31.png
| 2024-05-14 | 2024-05-14 14:38 | Cabanatuan City Health Office | Kapt Pepe, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28127 | DOH | cholab2021@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2369 |  | Genexpert IV | 831592 | Due for Calibration | Performed Calibration (3 modules Passed, A1 module newly MR) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5914781521504480112/5914781521504480112_signature_28.png | Maricris So | https://www.jotform.com/uploads/macarewecare/222927978372067/5914781521504480112/5914781521504480112_signature_31.png
| 2024-05-14 | 2024-05-14 14:24 | LAPU LAPU CITY HEALTH OFFICE | LAPULAPU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28126 | DOH | tanya_pax@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2161 |  | GX-IV | 837661 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5914772434433172688/5914772434433172688_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5914772434433172688/5914772434433172688_signature_31.png
| 2024-05-14 | 2024-05-14 13:34 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Lanoratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28125 | DOH | wcpatron11x4@gmail.com | 25 |  |  | GXIV | 831534 | Unit for preventive maintenance. UPS for checking. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed troubleshooting of UPS. Generated reports. | Unit is operational. Accupower UPS (AT1000180910ST0538) not backing up and is for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5914745749494457963/5914745749494457963_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5914745749494457963/5914745749494457963_signature_31.png
| 2024-05-14 | 2024-05-14 13:20 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28124 | DOH | ccmclaboratory01@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2147 |  | GXIV-4D | 838561 | Due for PM. | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5914735521583301569/5914735521583301569_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5914735521583301569/5914735521583301569_signature_31.png
| 2024-05-14 | 2024-05-14 11:58 | Monkayo Rural Health Unit | Poblacion Monkayo | 11 | Davao de Oro | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28123 | DOH | kendralois730@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1642 |  | GX-IV 4D | 838487 | Due for Preventive Maintenance Monitor Display Problem | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Reset Cables of the monitor replug and plug. Retry to adjust monitor settings. Still display issue problem | Perform Preventive Maintenance Monitor Troubleshooting | All Modules are working, Machine is Operational Monitor LCD Display problem | Monitor Display Error documented and sent to Jotform for further assessment. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5914688713319660710/5914688713319660710_signature_28.png | Karen S. Anuta | https://www.jotform.com/uploads/macarewecare/222927978372067/5914688713319660710/5914688713319660710_signature_31.png
| 2024-05-14 | 2024-05-14 10:57 | Talacogon District Hospital | Purok Bayabas, San Isidro, Talacogon | 13 | Agusan del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28122 | DOH | tdh.labdept@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1302 |  | GX-4 | 838275 | No problems reported | Activities done | Preventive Maintenance | All modules are functional All units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5914654519113637871/5914654519113637871_signature_28.png | Stella Rica Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5914654519113637871/5914654519113637871_signature_31.png
| 2024-05-14 | 2024-05-14 11:02 | Calaanan Health Center | Nadurata St. Caloocan | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28121 | DOH | rosaleslorina0917@gmail.com | 25 |  |  | GXIV | 831892 | Unit for preventive maintenance. UPS not functioning. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed trouble shooting of UPS. Generated reports. | Unit is operational. Accupower UPS (AT1K230830ST0014) is not operational and for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5914653388621315840/5914653388621315840_signature_28.png | Lorina Rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/5914653388621315840/5914653388621315840_signature_31.png
| 2024-05-13 | 2024-05-09 15:55 | Iriga City Health Office | Sta. Cruz Sur, Iriga City | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28120 | DOH | monge@yahoo.com.ph | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1620 |  | GX4-6C | 838391 | Gx for Preventive Maintenance Data collection for B2 & B4 | Cleaned cartridge bays, plunger rod, icore Generate reports Collected B2&B4 gxx file Charged UPS & tested for backup power dependability. Checked power connections. Checked power output. | Performed Preventive Maintenance Collected gxx for b2& b4 | Unit is operational | UPS no backup power: cant store power when plugged off from AC. request for replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913913407424893108/5913913407424893108_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5913913407424893108/5913913407424893108_signature_31.png
| 2024-05-13 | 2024-05-10 13:52 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28119 | DOH | dianeruiz147@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 382 |  | GXIV-4-D-6C | 838438 | #NAME? | #NAME? | #NAME? | - Unit is partially operational - Module B3 (806441) has already been reported. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913870392184720628/5913870392184720628_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5913870392184720628/5913870392184720628_signature_31.png
| 2024-05-13 | 2024-05-10 11:36 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28118 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 187 |  | GXII-2D | 820303 | #NAME? | #NAME? | #NAME? | - Unit is partially operational - Module A1 (720229) For Retry of XpertCheck next site visit - Module A2(720228) Failed XpertCheck. For Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913868802184682076/5913868802184682076_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5913868802184682076/5913868802184682076_signature_31.png
| 2024-05-13 | 2024-05-08 11:27 | Naguilian District Hospital | Naguilian, La Union | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28117 | DOH | marjelynrayray01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 952 |  | GXIV-4-D-6C | 838426 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Fully Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913855462181644573/5913855462181644573_signature_28.png | Marjelyn B. Rayray | https://www.jotform.com/uploads/macarewecare/222927978372067/5913855462181644573/5913855462181644573_signature_31.png
| 2024-05-13 | 2024-05-07 09:51 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28116 | DOH | gbmmbg@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 636 |  | GXIV-4-D-6C | 838437 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913854522183585001/5913854522183585001_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5913854522183585001/5913854522183585001_signature_31.png
| 2024-05-13 | 2024-05-07 08:11 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28115 | DOH | thisisformyfamily@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1559 |  | GXIV-4-D-6C | 838286 | For Preventive Maintenance | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913852582189641419/5913852582189641419_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5913852582189641419/5913852582189641419_signature_31.png
| 2024-05-13 | 2024-05-13 12:08 | Southern Philippines Medical Center | JP Laurel, Davao City | 11 | Davao del Sur | ICU 2 | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28114 | COMMERCIAL | icu2estuya@gmail.com | 23.7 |  |  | MCS+ | 15H034G | Machine not functional | Reset the connection of the PCB Boards Test the voltage of the power supply | Troubleshoot the MCS+ Machine | Still the Machine is not operational Waiting for PSU from Headoffice for problem Isolation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913839680121145534/5913839680121145534_signature_28.png | Dianne Karla de Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5913839680121145534/5913839680121145534_signature_31.png
| 2024-05-13 | 2024-05-07 11:40 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28113 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1288 |  | GXIV | 831691 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913796183618463746/5913796183618463746_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5913796183618463746/5913796183618463746_signature_31.png
| 2024-05-13 | 2024-05-03 14:07 | Plaridel Rural Health Unit | Plaridel, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28112 | DOH | Yellei0908@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2024-05-03 | GXIV | 831711 | Module A2 for replacement | Replaced module A2 by 684364 | #NAME? | Unit is operational |  | GemeXpert Module SN: 684364 | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913795863615331180/5913795863615331180_signature_28.png | Ahlelei Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5913795863615331180/5913795863615331180_signature_31.png
| 2024-05-13 | 2024-05-07 11:40 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28111 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1288 |  | GXIV | 831691 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913795433615989307/5913795433615989307_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5913795433615989307/5913795433615989307_signature_31.png
| 2024-05-13 | 2024-05-09 13:37 | Bustos Rural Health Unit | Bustos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28110 | DOH | ronald_galvez2002@yahoo.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 744 |  | GXIV | 831874 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913795213614448527/5913795213614448527_signature_28.png | Ronald Galvez | https://www.jotform.com/uploads/macarewecare/222927978372067/5913795213614448527/5913795213614448527_signature_31.png
| 2024-05-13 | 2024-05-09 13:37 | Bustos Rural Health Unit | Bustos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28109 | DOH | ronald_galvez2002@yahoo.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 744 |  | GXIV | 831874 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913794823611339119/5913794823611339119_signature_28.png | Ronald Galvez | https://www.jotform.com/uploads/macarewecare/222927978372067/5913794823611339119/5913794823611339119_signature_31.png
| 2024-05-13 | 2024-05-09 15:32 | National Children's Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 28108 | MACARE-HBDC | arellanojohneric09@gmail.com | 26 |  |  | MCS + | 17H632SPG | As per end user they encounter system problem error and blurr image in display | Checked unit Checked program parameters Checked and reset cable Performed simulation running | Troubleshooting | During simulation running no error encounter unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913697163612426094/5913697163612426094_signature_28.png | JOHN ERIC ARELLANO | https://www.jotform.com/uploads/macarewecare/222927978372067/5913697163612426094/5913697163612426094_signature_31.png
| 2024-05-12 | 2024-05-08 10:00 | General Santos City Health Office | General Santos City, South Cotabato | 12 | South Cotabato | Social hygiene Clinic | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28107 | SHELL FOUNDATION | dontingzon@gmail.com | 26.3 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 235 |  | Gx-IV 10-C | 120000147 | N/A | Genexpert machine Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5913378355915574691/5913378355915574691_signature_28.png | Jonnel Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5913378355915574691/5913378355915574691_signature_31.png
| 2024-05-10 | 2024-05-10 17:26 | Southern Philippines Medical Center | JP. Laurel Davao City | 11 | Davao del Sur | ICU2 | NO | TROUBLESHOOTING | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28106 | COMMERCIAL | icu2estuya@gmail.com | 22.6 |  |  | MCS+ | 15H034G | Machine Suddenly off during procedure | Check Centrifuge, check voltages across the machine and the Centrifuge Resetting Board connection | Diagnose each part | Unit not operational Waiting for parts from Manila to Isolate the diagnosis | Machine is unavailable to use. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911437295868018437/5911437295868018437_signature_28.png | Dianne karla | https://www.jotform.com/uploads/macarewecare/222927978372067/5911437295868018437/5911437295868018437_signature_31.png
| 2024-05-10 | 2024-05-10 16:59 | Santolan Super Health Center | Evangelista Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28105 | MACARE-HBDC | samandlet17@gmail.com | 23 |  |  | GX-IV-D-10C | 110024245 | For installation | Installed the GeneXpert System | Installed the GeneXpert System: -GeneXpert unit -Workstation -UPS -Barcode Scanner -Printer Functionality test Generated IQR | Unit is operational. UPS not operational, for replacement. | UPS SN: AT1K230830ST0024 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911410076248060991/5911410076248060991_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5911410076248060991/5911410076248060991_signature_31.png
| 2024-05-10 | 2024-05-10 13:58 | Santolan Super Health Center | Evangelista Ave. | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28104 | MACARE-HBDC | samandlet17@gmail.com | 23 |  |  | GX-IV-D-10C | 110024247 | For installation | Installed the GeneXpert System | Installed the GeneXpert System: -GeneXpert unit -Workstation -Barcode Scanner -Printer -UPS Functionality test Generated IQR | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911407706249166031/5911407706249166031_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5911407706249166031/5911407706249166031_signature_31.png
| 2024-05-10 | 2024-05-10 15:39 | Veterans Memorial Medical Center | North Ave., Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | biomedunit.vmmc@gmail.com |  | 28103 | CONSORTIUM | vmmc.moleclab@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 178 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1575 ASSAY: Xpert SARS-CoV-2, COUNT: 110 ASSAY: Xpert HCV VL, COUNT: 17 |  | GX IV 6C-D | 839150 | GX unit due for Preventive Maintenance | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test All modules passed Generated report | Performed Preventive Maintenance Performed reliability test | Unit is operational All modules working | Unit for maintenance every 1st Friday of the month |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911365365683726156/5911365365683726156_signature_28.png | Sheila Bermido | https://www.jotform.com/uploads/macarewecare/222927978372067/5911365365683726156/5911365365683726156_signature_31.png
| 2024-05-10 | 2024-05-10 15:41 | Cordlife Medical Philippines | UP-Ayala Technohub, Commonwealth Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CYTIVA | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 28102 | COMMERCIAL | patrick.baco@cordlife.com | 25 |  |  | SmartMax | 5131.3196 | Unit unable to boot, suspected drained CPU BIOS battery | Replaced CPU BIOS battery | Replaced CPU BIOS battery | Unit is not operational, further troubleshooting required with additional tools. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911364568028387816/5911364568028387816_signature_28.png | Patrick Baco | https://www.jotform.com/uploads/macarewecare/222927978372067/5911364568028387816/5911364568028387816_signature_31.png
| 2024-05-10 | 2024-05-10 15:32 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28101 | COMMERCIAL | cduhlabresults2@yahoo.com | 25 |  |  | MCS+ | 14F014 | Preventive Maintenance pump roller, ups, check valve and compressor sounds | Preventive Maintenance pump roller, ups, check valve and compressor sounds | Preventive Maintenance pump roller, ups, check valve and compressor sounds | Machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911357661119392469/5911357661119392469_signature_28.png | Ericka Jane Corro, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5911357661119392469/5911357661119392469_signature_31.png
| 2024-05-10 | 2024-05-10 13:25 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Road, Pamplona, Las Piñas City | NCR | NCR 4th - Las Piñas | Blood Bank | YES | REPAIR | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | noel.feranil67@gmail.com |  | 28100 | MACARE-HBDC | neressleanoratienza@yahoo.com.ph | 25 |  | 2024-05-08 | MCS+ | 11J220 | Unit has Blood Pump Problem. | Repair/Parts Replacement. | Removed and Replaced Defective Blood Pump Motor Assy. Checked Diagnostic and A/D Values. Performed Simulation Test. | Unit is Operational. | Motor Assy: 230705 | Motor Assembly | SI 097969 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911355550149154292/5911355550149154292_signature_28.png | Neress Atienza | https://www.jotform.com/uploads/macarewecare/222927978372067/5911355550149154292/5911355550149154292_signature_31.png
| 2024-05-10 | 2024-05-10 15:19 | Vinzons Rural Health Unit | FO, Balce St., Brgy. Dos, Poblacion, Vinzons | 5 | Camarines Norte | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com1 |  |  | 28099 | DOH | jasminlottidp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 908 |  | GX4-6C | 838576 | Frequent Date and Time reset | Replaced CPU Date and Time battery Generate report | Performed Troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911350480012018777/5911350480012018777_signature_28.png | Jasmin Lotti A. Dela Peña | https://www.jotform.com/uploads/macarewecare/222927978372067/5911350480012018777/5911350480012018777_signature_31.png
| 2024-05-10 | 2024-05-10 15:07 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank Service Unit | NO | MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28098 | COMMERCIAL | vsmmcbloodservices@gmail.com | 25 |  |  | MCS+ | 02A090 | Preventive Maintenance pump roller, ups, check valve and compressor sounds | Preventive Maintenance pump roller, ups, check valve and compressor sounds | Preventive Maintenance pump roller, ups, check valve and compressor sounds | Machine is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911342451116080507/5911342451116080507_signature_28.png | Charlotte Mae Ygoña, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5911342451116080507/5911342451116080507_signature_31.png
| 2024-05-10 | 2024-05-10 14:47 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28097 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 23 |  |  | MCS+ | 02F035 | Preventive Maintenance pump roller, ups, check valve and compressor sounds | Preventive Maintenance pump roller, ups, check valve and compressor sounds | Preventive Maintenance pump roller, ups, check valve and compressor sounds | Machine operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911330011111570510/5911330011111570510_signature_28.png | Dave Gabini | https://www.jotform.com/uploads/macarewecare/222927978372067/5911330011111570510/5911330011111570510_signature_31.png
| 2024-05-10 | 2024-05-10 14:33 | Chong Hua Hospital- Fuente | FUNTE OSMENA, CEBU | 7 | Cebu | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 28096 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 23 |  |  | MCS | 11J219 | Check pump roller, check valve, check ups | Check pump roller, check valve, check ups | Check pump roller, check valve, check ups | Machine operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911327771115599081/5911327771115599081_signature_28.png | Dave Gabini | https://www.jotform.com/uploads/macarewecare/222927978372067/5911327771115599081/5911327771115599081_signature_31.png
| 2024-05-10 | 2024-05-10 13:35 | The Medical City South Luzon | Sta. Rosa Laguna | 4A | Laguna | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28095 | COMMERCIAL | tmcsl-molecular@themedicalcity.com | 25 |  | 2024-05-10 | GXIV | 839233 | Unit for prevenntive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Modules B3 (866650) and B4 (866802), Service Required. Modules B1 (866519) and B2 (866649), Passed. |  | Xpertcheck kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911304951618352253/5911304951618352253_signature_28.png | Roxanne D. Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5911304951618352253/5911304951618352253_signature_31.png
| 2024-05-10 | 2024-05-10 13:15 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY, CEBU | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28094 | DOH | mendozarey42@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5929 ASSAY: Xpert HCV VL, COUNT: 25 ASSAY: Xpert HIV-1 Viral Load, COUNT: 241 |  | GX-IV | 838510 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911275391117960215/5911275391117960215_signature_28.png | Rey C. Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5911275391117960215/5911275391117960215_signature_31.png
| 2024-05-10 | 2024-05-10 13:09 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY, CEBU | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28093 | DOH | mendozarey42@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5929 ASSAY: Xpert HCV VL, COUNT: 25 ASSAY: Xpert HIV-1 Viral Load, COUNT: 241 |  | GX-IV | 838508 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911273281118904871/5911273281118904871_signature_28.png | Rey C. Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5911273281118904871/5911273281118904871_signature_31.png
| 2024-05-10 | 2024-05-10 12:37 | Makati Medical Center | Amorsolo st., Makati City | NCR | NCR 4th - Makati | CRM | YES | MAINTENANCE | OTHERS | Noel Floresca |  | florescanoel77@gmail.com | troxoctubre@gmail.com | Miltenyi | 28092 | COMMERCIAL | AlcorDeeond.Padilla@makatimed.net.ph | 28 |  | 2024-05-09 | Gentle MACS Octo Dissociator | 1261 | Due for Maintenance | Cleaned unit Performed functionality test Checked and test connector Checked and test display panel Checked program parameters Test run unit ok. | Performed maintenance | Operational |  |  | 2987 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911261250217457806/5911261250217457806_signature_28.png | Alcor Deeond F. Padilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5911261250217457806/5911261250217457806_signature_31.png
| 2024-05-10 | 2024-05-10 12:37 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28091 | COMMERCIAL | cmc.rtpcr@gmail.com | 25 |  |  | GXII | 812916 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911251931611590212/5911251931611590212_signature_28.png | Andrea Gail B. Orgesvik | https://www.jotform.com/uploads/macarewecare/222927978372067/5911251931611590212/5911251931611590212_signature_31.png
| 2024-05-10 | 2024-05-07 12:31 | Virac Municipal Health Office | Virac, Catanduanes | 5 | Catanduanes | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28090 | DOH | limarmariano50@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1822 |  | GX4-6C | 838563 | Gx for Calibration CPU restarting while running Gx test Monitor grid/blue lines covering half the screen | XpertCheck Result: B1/730980 - new module B2/210030593 - new module B3/736617 - service required B4/785028 - new module Checked AVR power output Checked ups power output Checked/Tighten power cable connection Checked/cleaned CPU RAM Collect data Cleaned cartridge bays, plunger rod, icore Generate report | Performed XpertCheck Performed Preventive Maintenance Performed Troubleshooting | Unit is operational Monitor defective/not operational CPU problem still persists | CPU Problem still persists, request for replacement, Monitor request for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911228181115726701/5911228181115726701_signature_28.png | Ramil A. Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/5911228181115726701/5911228181115726701_signature_31.png
| 2024-05-10 | 2024-05-10 11:18 | Bicol Medical Center | Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28089 | MACARE-HBDC | josefreyes0715@gmail.com | 22 |  |  | MCS+ | 11J221 | Retrieve defective LCD ribbon cable Performed Preventive Maintenance | Retrieve defective lcd ribbon cable Cleaned/disinfect mcs+ surface, centrifuge bowl, rotor pumps, valves | Retrieve Defective component for replacement | Unit is not operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911227251114944050/5911227251114944050_signature_28.png | Josef D. Reyes, RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5911227251114944050/5911227251114944050_signature_31.png
| 2024-05-10 | 2024-05-10 11:23 | Bicol Medical Center | Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28088 | MACARE-HBDC | josefreyes0715@gmail.com | 22 |  |  | MCS+ | 15H031G | MCS+ for Preventive Maintenance | Cleaned/disinfect mcs+ surface, centrifuge bowl, valves, rotor pumps Checked parameters | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911226631114387280/5911226631114387280_signature_28.png | Josef D. Reyes, RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5911226631114387280/5911226631114387280_signature_31.png
| 2024-05-10 | 2024-05-10 11:26 | Bicol Medical Center | Naga | 5 | Camarines Sur | Blood Bank | YES | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28087 | MACARE-HBDC | josefreyes0715@gmail.com | 22 |  |  | MCS+ | 15H027G | MCS+ for Preventive Maintenance | Cleaned/disinfect mcs+ surface, centrifuge bowl, rotor pumps, valves Checked parameters | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911225871116481231/5911225871116481231_signature_28.png | Josef D. Reyes, RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5911225871116481231/5911225871116481231_signature_31.png
| 2024-05-10 | 2024-05-10 11:03 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Womens Health Center | NO | MAINTENANCE TROUBLESHOOTING | MEDGYN | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomedtech@cmc.ph |  | 28086 | COMMERCIAL | sweetlove@yahoo.com | 25 |  |  | AL-106 | 118AAGN0006 | Unit for preventive maintenance. For troubleshooting of recording console. | Preventive maintenance and troubleshooting. | Performed preventive maintenance. Cleaned unit and lens. Performed functionality testing. Performed troubleshooting on recording console. Performed actual recording. Generated reports. | Unit is operational. Recording console's AV input is loosed and only working on a certain fixed position. Endorsed to biomed the issue. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5911203861618674918/5911203861618674918_signature_28.png | Dulc amor tmodong | https://www.jotform.com/uploads/macarewecare/222927978372067/5911203861618674918/5911203861618674918_signature_31.png
| 2024-05-09 | 2024-05-09 15:40 | SACCL | Quiricada St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | Relocation | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28085 | MACARE-HBDC | sacclkiteval@gmail.com | 28 |  | 2024-05-08 | GXIV-D-6C | 839147 | Unit has Lost Communication Errors. | Relocation | Checked Output Voltage. Add Voltage Regulator. Relocate GeneXpert System. Generate Reports. | All Modules are Operational. |  | Automatic Voltage Regulator | DR 11482 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910514117217278812/5910514117217278812_signature_28.png | Melanie Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5910514117217278812/5910514117217278812_signature_31.png
| 2024-05-09 | 2024-05-09 15:26 | Rosario Super Health Center | Dr. Sixto Avenue | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28084 | MACARE-HBDC | eden.cawayan@yahoo.com | 25 |  |  | GX-IV-D-10C | 110024246 | For installation | Installed the GeneXpert System | Installed the GeneXpert System: -GeneXpert unit -Workstation -UPS -Printer Functionality test Generated IQR | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910494273691052950/5910494273691052950_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5910494273691052950/5910494273691052950_signature_31.png
| 2024-05-09 | 2024-05-09 12:07 | Philippine General Hospital | Taft Ave. Ermita Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28083 | MACARE-HBDC | pghmolecularlab@gmail.com | 25 |  |  | GX16 | 844975 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910478631614285144/5910478631614285144_signature_28.png | Allan canseco | https://www.jotform.com/uploads/macarewecare/222927978372067/5910478631614285144/5910478631614285144_signature_31.png
| 2024-05-09 | 2024-05-09 14:34 | Biopath QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28082 | COMMERCIAL | biopathqc@biopath.ph | 29 | ASSAY: Xpert CT/NG, COUNT: 50 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 43 |  | GX16 D | 839306 | Due for maintenance,C2 disable failed xpertcheck,A3 HW failed | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | 14 module operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910466023211614661/5910466023211614661_signature_28.png | Biopath Clinical Diagnostics, Inc-E. Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5910466023211614661/5910466023211614661_signature_31.png
| 2024-05-09 | 2024-05-09 14:42 | CEBU DOCTORS' UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | MOLECULAR LAB | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 28081 | COMMERCIAL | cduhcovidpoct@gmail.com | 18 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 66 ASSAY: Xpert CT/NG, COUNT: 167 ASSAY: Xpert HBV VL, COUNT: 154 ASSAY: Xpert HPV, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1110 ASSAY: Xpert HCV VL, COUNT: 131 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 |  | GX-IV | 110008524 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910465631112692116/5910465631112692116_signature_28.png | Queenie Valencia | https://www.jotform.com/uploads/macarewecare/222927978372067/5910465631112692116/5910465631112692116_signature_31.png
| 2024-05-09 | 2024-05-09 13:19 | Ospital ning Capas | Capas Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28080 | DOH | onclaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1676 |  | GXIV-D | 831591 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910416227594963430/5910416227594963430_signature_28.png | Afrilyn Roque | https://www.jotform.com/uploads/macarewecare/222927978372067/5910416227594963430/5910416227594963430_signature_31.png
| 2024-05-09 | 2024-05-09 12:04 | Pulilan Rural Health Unit | Pulilan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28079 | DOH | celsptcruz19@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1097 |  | GXIV | 831915 | Unit for maintenance | Temp. errors found on module A3 Advised end user to use external fan for air circulation | #NAME? | Unit is operational A3 under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910370871426205069/5910370871426205069_signature_28.png | Cecille Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5910370871426205069/5910370871426205069_signature_31.png
| 2024-05-09 | 2024-05-09 11:39 | UP-ICHHD | Pefro Gil St, Ermita Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28078 | MACARE-HBDC | dcsantos9@up.edu.ph | 28 |  |  | GXIV | 808015 | Unit for preventive maintenance. For troubleshooting of GX as per instruction of cepheid. | Preventive maintenance and troubleshooting. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910357901618824064/5910357901618824064_signature_28.png | Dana Kamylle Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5910357901618824064/5910357901618824064_signature_31.png
| 2024-05-09 | 2024-05-09 11:48 | Pulilan Rural Health Unit | Pulilan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28077 | DOH | celsptcruz19@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1097 |  | GXIV | 831915 | Unit for maintenance | Found temperature errors on module A3 | #NAME? | Unit is operational Module A3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910364899954085923/5910364899954085923_signature_28.png | Cecille Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5910364899954085923/5910364899954085923_signature_31.png
| 2024-05-09 | 2024-05-09 11:13 | Tarlac City Health Center 3 | Brgy. San Miguel Tarlac City | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28076 | DOH | daganlea572@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 551 |  | GXIV-D | 831697 | Due for preventive and calibration | Preventive maintenance Calibration | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test result Performed xpertcheck Performed functionality test Gathered reports | Unit is operational All modules passed calibration UPS: AT1000181128ST0132 for replacement | Used 4 xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910351977593568759/5910351977593568759_signature_28.png | Lea A. Dagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5910351977593568759/5910351977593568759_signature_31.png
| 2024-05-09 | 2024-05-09 11:20 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28075 | DOH | sarah_yangzon@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1073 | 2024-05-09 | GXIV | 838406 | n/a | PM | - Preventive maintenance (Cleaned plunger rod, iCore, Cartridge Bay nad Fan filter) - | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910345329123562558/5910345329123562558_signature_28.png | Sarah Yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5910345329123562558/5910345329123562558_signature_31.png
| 2024-05-09 | 2024-05-09 11:06 | Claver Rural Health Unit | Brgy. Poblacion, Claver | 13 | Surigao del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 28074 | DOH | rhodaruthdalogdog@yahoo.com.ph | 25 |  |  | GX-4 | 838514 | Filter not reusable. | Filter to be reported to Cepheid | Preventive Maintenance | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910337926309010390/5910337926309010390_signature_28.png | Joyce E. Andoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5910337926309010390/5910337926309010390_signature_31.png
| 2024-05-09 | 2024-05-09 10:02 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR CALIBRATION | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28073 | COMMERCIAL | cchdserology@gmail.com | 22 |  |  | Gx 4 | 829419 | Defective module 3 and 4 | Clean fan filter New module 3 installed sn 861590 Gathering data of genexpert machine IQR, SYSTEM LOG | Clean fan filter New module 3 installed sn 861590 Gathering data of genexpert machine IQR, SYSTEM LOG | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5910314111119884372/5910314111119884372_signature_28.png | Owen joshua L. briones | https://www.jotform.com/uploads/macarewecare/222927978372067/5910314111119884372/5910314111119884372_signature_31.png
| 2024-05-08 | 2024-05-08 11:11 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28072 | MACARE-HBDC | kaye11ren08@gmail.com | 25 |  | 2024-05-03 | GXIV-10C | 110001650 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module A1 (210077560) retry then service required and A3 (210010363) is service required. Module A4 (210048255) passed. Result of A2 (210077583) is Retry, but there's no spare Xpertcheck to continue. |  | Xpertcheck kit c/o pbsp. | PL-015868 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909678849819540774/5909678849819540774_signature_28.png | Carina Moncada | https://www.jotform.com/uploads/macarewecare/222927978372067/5909678849819540774/5909678849819540774_signature_31.png
| 2024-05-08 | 2024-05-08 13:44 | Hermosa Rural Health Unit | Hermosa Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28071 | DOH | kryztnangela@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 543 ASSAY: Xpert HIV-1 Viral Load, COUNT: 36 |  | GXIV-D | 831644 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909577097592808413/5909577097592808413_signature_28.png | Mary Ann Patambang | https://www.jotform.com/uploads/macarewecare/222927978372067/5909577097592808413/5909577097592808413_signature_31.png
| 2024-05-08 | 2024-05-08 09:44 | San Vicente Municipal Health Office | San Vicente Municipal Puerto Princesa North Road, Poblacion, San Vicente, 5309 Palawan | 4B | Palawan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28070 | MACARE-HBDC | santosdollethorecel69@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX IV 10C-L | 120002441 | GX Unit for installation | Unbox GX and all accessories Setup GX Setup user admin Setup System configuration Installed printer driver Performed PM Generated report | Setup and Installed GX | Uni is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909562200612410532/5909562200612410532_signature_28.png | Dollethorecel Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5909562200612410532/5909562200612410532_signature_31.png
| 2024-05-08 | 2024-05-08 12:24 | Calubian Rural Health Unit | Calubian, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28069 | DOH | mjcabardo06@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 928 |  | GXIV-4D | 838399 | Due for PM | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909517094085816386/5909517094085816386_signature_28.png | Mighty Joy V. Cabardo | https://www.jotform.com/uploads/macarewecare/222927978372067/5909517094085816386/5909517094085816386_signature_31.png
| 2024-05-08 | 2024-05-08 11:57 | Porac Rural Health Unit | Porac Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28068 | DOH | rhulabporac@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 845 |  | GXIV-D | 831911 | Due for maintenance UPS for troubleshooting | Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed UPS Troubleshooting Performed functionality test Gathered reports | Unit is operational UPS:AT1000181128ST0159 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909504977599076344/5909504977599076344_signature_28.png | Marife Lansangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5909504977599076344/5909504977599076344_signature_31.png
| 2024-05-08 | 2024-05-08 10:39 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28067 | DOH | abigailgsula@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1310 |  | GXIV-D | 831912 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909454667598802285/5909454667598802285_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5909454667598802285/5909454667598802285_signature_31.png
| 2024-05-08 | 2024-05-08 10:19 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28066 | DOH | laguitannataniel@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1137 |  | GXIV | 831845 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909449089954367055/5909449089954367055_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5909449089954367055/5909449089954367055_signature_31.png
| 2024-05-08 | 2024-05-08 10:12 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28065 | DOH | laguitannataniel@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1137 |  | GXIV | 831677 | Unit for maintenance Unit's desktop is not booting properly | End user switched GeneXpert to other GeneXpert desktop unit (Defective desktop is already boxed upon visit) | #NAME? | Unit is operational Desktop for reporting (SN: 2UA8471RT1) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909449089952093536/5909449089952093536_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5909449089952093536/5909449089952093536_signature_31.png
| 2024-05-08 | 2024-05-07 10:50 | Ospital ng Makati | Sampaguita St. | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakproperty@yahoo.com |  | 28064 | MACARE-HBDC | osmakpatho_labs@yahoo.com.ph | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 |  | GX-IV-D-10C | 110023460 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909442398131868393/5909442398131868393_signature_28.png | Lily Rose Barras | https://www.jotform.com/uploads/macarewecare/222927978372067/5909442398131868393/5909442398131868393_signature_31.png
| 2024-05-08 | 2024-05-08 10:06 | Tondo Medical Center | H. Lopez Blvd. Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28063 | DOH | microscan@tmc.doh.gov.ph | 25 |  |  | GXIV | 831872 | Unit for relocation to molecular lab. | Installation. | Performed installation. Rearranged cables. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909436771614005506/5909436771614005506_signature_28.png | Analyn N.Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5909436771614005506/5909436771614005506_signature_31.png
| 2024-05-08 | 2024-05-07 14:00 | Malungon Rural Health Unit | Poblacion Malungon, Sarangani | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 28062 | DOH | ulbenarioe@yahoo.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 875 |  | Gx-IV | 838546 | Request Filter | Genexpert machine Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules | Requesting filter |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909427130688991511/5909427130688991511_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5909427130688991511/5909427130688991511_signature_31.png
| 2024-05-08 | 2024-05-08 08:46 | Macabebe Rural Health Unit | Macabebe Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28061 | DOH | salvamari13@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1091 |  | GXIV-D | 831893 | Due for preventive maintenance UPS not backing up for more than 5 1 minute | Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed UPS troubleshooting Performed functionality test Gathered reports | Unit is operational US:AT1000181128ST0187 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909387227596904885/5909387227596904885_signature_28.png | Salva Mari Sese | https://www.jotform.com/uploads/macarewecare/222927978372067/5909387227596904885/5909387227596904885_signature_31.png
| 2024-05-08 | 2024-05-08 07:29 | Apalit Rural Health Unit | Apalit Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28060 | DOH | duenasblesilda@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 961 ASSAY: Xpert HIV-1 Viral Load, COUNT: 44 |  | GXIV-D | 831900 | Due for maintenance UPS troubleshooting A4:674002 troubleshooting | Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results UPS troubleshooting Performed functionality test Gathered reports | Unit is operational A4:674002 for reporting (temperature error) UPS:AT1000181128ST0220 For replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5909349997595944781/5909349997595944781_signature_28.png | Blesilda Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5909349997595944781/5909349997595944781_signature_31.png
| 2024-05-07 | 2024-05-07 14:35 | Mini Forest Health Center (District C) | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 28059 | DOH | miniforestlaboratory@gmail.com | 25.8 |  |  | GX-IV 4D | 838507 | Due for Preventive Maintenance Relocation of GeneXpert System | Setup and reinstall to new and optimal position of the GeneXpert System Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Transfer the GeneXpert System to newly renovated laboratory Perform Preventive Maintenance | All Modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908735641114346387/5908735641114346387_signature_28.png | Joy L. Nio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5908735641114346387/5908735641114346387_signature_31.png
| 2024-05-07 | 2024-05-07 14:31 | SACCL | Quiricada St., Sta.Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28058 | MACARE-HBDC | blessieestrelles@gmail.com | 28 |  |  | GXIV-D-6C | 839147 | Unit is due for Installation. | Installation | Installed GX System. Checked Output Voltages. Performed Self-test. Generate Reports. | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908731025316746957/5908731025316746957_signature_28.png | Blessie Estrelles | https://www.jotform.com/uploads/macarewecare/222927978372067/5908731025316746957/5908731025316746957_signature_31.png
| 2024-05-07 | 2024-05-07 13:49 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28057 | DOH | malitchelllanos@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1325 |  | GXIV-4D | 838555 | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908705129928893235/5908705129928893235_signature_28.png | Malitchel B. Llanos | https://www.jotform.com/uploads/macarewecare/222927978372067/5908705129928893235/5908705129928893235_signature_31.png
| 2024-05-07 | 2024-05-07 13:25 | Bureau of Quarantine | Intramuros Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28056 | MACARE-HBDC | dklo.lab@quarantine.doh.gov.ph | 25 |  |  | GX16 | 844418 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908693953311024442/5908693953311024442_signature_28.png | Dulce Karmela Lo | https://www.jotform.com/uploads/macarewecare/222927978372067/5908693953311024442/5908693953311024442_signature_31.png
| 2024-05-07 | 2024-05-07 12:22 | Central Luzon Doctor's Hospital | Hospital dr. Brgy. San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28055 | COMMERCIAL | cldhlaboratory@cldh.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 ASSAY: Xpert SARS-CoV-2, COUNT: 563 |  | GXII-D | 830841 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908676163111498609/5908676163111498609_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/5908676163111498609/5908676163111498609_signature_31.png
| 2024-05-07 | 2024-05-07 12:20 | Central Luzon Doctor's Hospital | Hospital dr. Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28054 | COMMERCIAL | cldhlaboratory@cldh.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 ASSAY: Xpert SARS-CoV-2, COUNT: 563 |  | GXIV-D | 839232 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908674243111372227/5908674243111372227_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/5908674243111372227/5908674243111372227_signature_31.png
| 2024-05-07 | 2024-05-07 11:36 | Puerto Princesa City Health Office | DOTS Laboratory City Health Office, Valencia St., Puerto Princesa City | 4B | Palawan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28053 | PBSP | jocal0324@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1331 | 2024-05-06 | GX IV 6C-L | 804272 | Laptop due for replacement | Performed PM Replaced Laptop Archived results Setup new laptop Generated report | Replaced Laptop CNU330BSRQ -> CPH1023F46638 | unit is operational 3modules working 1 module for reporting |  | GX Laptop | DR 47194 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908627753573434624/5908627753573434624_signature_28.png | Jocelyn Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5908627753573434624/5908627753573434624_signature_31.png
| 2024-05-07 | 2024-05-07 11:11 | Tarlac City Health Center X | Brgy. Tibag Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28052 | DOH | rodelie220@gmail.com | 30.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 889 |  | GXIV-D | 831587 | Due for maintenance UPS Troubleshooting | Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed UPS Troubleshooting Performed functionality test Gathered reports | Unit is operational UPS for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908610853119619951/5908610853119619951_signature_28.png | Rodelie C. Verayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5908610853119619951/5908610853119619951_signature_31.png
| 2024-05-07 | 2024-05-07 11:10 | Navotas Medical and Wellness Center | M. Naval St. Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28051 | DOH | navotascitylab@gmail.com | 25 |  |  | GXIV | 831626 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908608444613804139/5908608444613804139_signature_28.png | Jade Renette Cesista | https://www.jotform.com/uploads/macarewecare/222927978372067/5908608444613804139/5908608444613804139_signature_31.png
| 2024-05-07 | 2024-05-07 10:45 | Puerto Princesa City Health Office XVI | DOTS Laboratory City Health Office, Valencia St., Puerto Princesa City | 4B | Palawan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28050 | MACARE-HBDC | jocal0324@gmail.com | 34 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX XVI 10C-L | 110021518 | GX XVI unit for replacement | Performed PM Setup User Administration Setup System Configuration Generated Report | Installed GX XVI to Puerto Princesa CHO | Unit is installed All 16 modules working | Room temp too high reaching 36°C. Ideal temp each module 22°C. AC needs replacement, bigger size at least 2HP to reach the ideal temperature of the machine |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908601063573499944/5908601063573499944_signature_28.png | Jocelyn Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5908601063573499944/5908601063573499944_signature_31.png
| 2024-05-07 | 2024-05-07 09:46 | Sta. Ignacia Rural Health Unit | Sta. Ignacia Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28049 | DOH | jdi0973@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 502 |  | GXIV-D | 831909 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908573063111033972/5908573063111033972_signature_28.png | Jess D. Ibarra Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5908573063111033972/5908573063111033972_signature_31.png
| 2024-05-07 | 2024-05-07 09:06 | Gerona Rural Health Unit | Brgy. Pinasling Gerona Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28048 | DOH | huwakie15@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 810 |  | GXIV-D | 831693 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908533618699933363/5908533618699933363_signature_28.png | Arremio Nool III | https://www.jotform.com/uploads/macarewecare/222927978372067/5908533618699933363/5908533618699933363_signature_31.png
| 2024-05-07 | 2024-05-07 08:16 | Tarlac City Health Center 2 | Brgy. Matatalaib Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28047 | DOH | clarizannebeltran1996@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 418 |  | GXIV-D | 831917 | Due for preventive maintenance UPS not backing up | Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Performed UPS Troubleshooting Gathered reports | Unit is operational UPS SN: AT1000181128ST0144 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5908504998309659831/5908504998309659831_signature_28.png | Clariz Anne Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5908504998309659831/5908504998309659831_signature_31.png
| 2024-05-06 | 2024-05-06 14:07 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 28046 | DOH | ammegzedd@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1161 | 2024-05-06 | GXIV | 838421 | n/a | #NAME? | #NAME? | Unit is operational - Module B2 disabled(defective) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5907867873618971750/5907867873618971750_signature_28.png | gemma malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5907867873618971750/5907867873618971750_signature_31.png
| 2024-05-06 | 2024-05-06 12:58 | Nagpayong Super Health Center | Nagpayong | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28045 | DOH | mnoreend@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1773 | 2024-04-30 | GX-IV-D-6C | 831905 | Module B1 for replacement | Replaced module B1 | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced module B1 Generated reports | All modules operational. |  | Old: B1/758209 New: B1/210071881 | G 6994 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5907843069126653562/5907843069126653562_signature_28.png | Ma.Noreen De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5907843069126653562/5907843069126653562_signature_31.png
| 2024-05-06 | 2024-04-30 14:17 | Batis Health Center | F. Manalo St | NCR | NCR 2nd - San Juan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28044 | DOH | almacenaio16@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1886 ASSAY: Xpert HIV-1 Viral Load, COUNT: 92 | 2024-04-30 | GX-IV-D-6C | 831598 | Modules A1 and A4 for replacement. | Replaced modules A1 and A4. | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Replaced modules A1 and A4 Generated reports. | Unit is operational. All modules operational. |  | Old: A1/210091071, A4/879974 New: A1/652487, A4/665175 | G 6991 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5907662523613089760/5907662523613089760_signature_28.png | Rosario Almacen | https://www.jotform.com/uploads/macarewecare/222927978372067/5907662523613089760/5907662523613089760_signature_31.png
| 2024-05-06 | 2024-04-30 11:06 | The Medical City | Ortigas Ave. | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28043 | COMMERCIAL | moleculardx@themedicalcity.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 41 ASSAY: Xpert CT/NG, COUNT: 82 ASSAY: Xpert HBV VL, COUNT: 562 ASSAY: Xpert HPV, COUNT: 426 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4975 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 70 ASSAY: Xpert HIV-1 Viral Load, COUNT: 675 ASSAY: Xpert SARS-CoV-2, COUNT: 656 | 2024-04-25 | GX-XVI-D-10C | 110015767 | Module C2 for replacement. | Replaced module C2 | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Replaced module C2 Generated report | Unit is operational. All modules operational. |  | Old: C2/210081841 New: C2/210094758 | DR# 11441 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5907660803617720776/5907660803617720776_signature_28.png | Kristine Parilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5907660803617720776/5907660803617720776_signature_31.png
| 2024-05-06 | 2024-04-30 11:31 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 28042 | COMMERCIAL | moleculardx@themedicalcity.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 6 ASSAY: Others, COUNT: 11 ASSAY: Xpert CT/NG, COUNT: 12 ASSAY: Xpert HBV VL, COUNT: 108 ASSAY: Xpert HPV, COUNT: 68 ASSAY: Xpert MTB/RIF Ultra, COUNT: 570 ASSAY: Others, COUNT: 28 ASSAY: Xpert HCV VL, COUNT: 14 ASSAY: Xpert HIV-1 Viral Load, COUNT: 88 |  | GX-IV-L-6C | 802630 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated report | Unit is operational. Module B4/210071314 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5907658853619584339/5907658853619584339_signature_28.png | Kristine Parilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5907658853619584339/5907658853619584339_signature_31.png
| 2024-05-05 | 2024-05-02 13:47 | Alabang Health Center (Bayanan Health Center) | Baywalk, Bayanan, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28041 | DOH | cathycarasig@yahoo.com | 27 | ASSAY: Xpert HPV, COUNT: 112 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1913 | 2024-04-29 | GXIV-D-6C | 831913 | Module 2 with Temperature Error. | Module Replacement | Replaced Module 2. Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational. | A2:692649->693531 , Machine Temporarily Trasferred to Bayanan Health Center | GX Module 6-C | G 6988 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5907130098318910915/5907130098318910915_signature_28.png | Catherine carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/5907130098318910915/5907130098318910915_signature_31.png
| 2024-05-05 | 2024-05-03 09:34 | Calaca Rural Health Unit | Dela Salud St., Calaca, Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28039 | DOH | loidamarasigan123@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1991 | 2024-04-29 | GXIV-D-6C | 821300 | Module 2 with Temperature Error. | Module Replacement. | Replaced Module 2. Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Replaced Filter. Performed Self-test. Generate Reports. | 2 Modules are Available. 2 are Disabled(with Temp Errors). | A2: 730076-> 680107 | GeneXpert Module 6-C | G 6989 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5907130098311348001/5907130098311348001_signature_28.png | Loida c. Marasigan | https://www.jotform.com/uploads/macarewecare/222927978372067/5907130098311348001/5907130098311348001_signature_31.png
| 2024-05-05 | 2024-05-02 07:59 | Molino 3 City Health Unit | Molino Road, Molino, Bacoor, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28038 | DOH | joanrnd@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1937 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 |  | GXIV-D-6C | 821221 | Unit is due for Maintenance | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | 3 Modules are Available, 1 Disabled. | B3/784355 with Temperature Error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5907130098313451205/5907130098313451205_signature_28.png | Jonalyn Aranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5907130098313451205/5907130098313451205_signature_31.png
| 2024-05-05 | 2024-05-02 11:05 | Almanza Uno Health Center | Alabang-Zapote Road, Almanza Uno, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 28040 | DOH | milcerpl@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1435 ASSAY: Xpert HIV-1 Viral Load, COUNT: 280 | 2024-04-29 | GXIV-D-6C | 831918 | Modules 2, 3 and 4 with Temperature Error. | Module Replacement. | Replaced Modules 2,3,4. Cleaned Cartridge Bay and Plunger Rod. Replaced Filter. Performed Self-test. Generate Reports. | All Modules are Operational. | B2:635070->701189, B3:210052509->702149, B4:758518->697937 | GX Module 6-C | G 6987 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5907130098315408076/5907130098315408076_signature_28.png | Milcer P. Lladoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5907130098315408076/5907130098315408076_signature_31.png
| 2024-05-03 | 2024-05-03 15:13 | Casimiro A. Ynares Sr.Memorial Hospital | Becco rd ,burgos Rodriguez Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28037 | DOH | caysmhlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 641 | 2024-04-30 | GX4 D | 831642 | For Module replacement A1 756792 - 747003 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | 900-0491rs genexpert module | G6997 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905429407919448598/5905429407919448598_signature_28.png | John Patrick Cabreros | https://www.jotform.com/uploads/macarewecare/222927978372067/5905429407919448598/5905429407919448598_signature_31.png
| 2024-05-03 | 2024-05-03 16:21 | Cordlife Medical Philippines | UP-Ayala Technohub, Commonwealth Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CYTIVA | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 28036 | COMMERCIAL | patricia.perfecto@cordlife.com | 25 |  |  | Smartmax | 5131.3196 | Unit not booting up | Performed functionality test Checked internal voltages | Performed troubleshooting | Unit is not operational, CPU BIOS battery is drained from normal use and requires replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905340756119956710/5905340756119956710_signature_28.png | Patricia Hannah Perfecto | https://www.jotform.com/uploads/macarewecare/222927978372067/5905340756119956710/5905340756119956710_signature_31.png
| 2024-05-03 | 2024-05-03 16:12 | San Jose Municipal Health Office | New Municipal Hall, San Jose Batangas | 4A | Batangas | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28035 | DOH | lab.sanjoserhu@gmail.com | 25 |  |  | GXIV | 821420 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905335634652088252/5905335634652088252_signature_28.png | Dolly Jeane Nitro | https://www.jotform.com/uploads/macarewecare/222927978372067/5905335634652088252/5905335634652088252_signature_31.png
| 2024-05-03 | 2024-05-03 15:48 | Krus na Ligas Health Center | 89 LT. J. Francisco St. Krus na Ligas | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Rhea Anne Cañelas |  | rheaannecanelas65@gmail.com |  |  | 28034 | PBSP | maryroseperejon@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1412 | 2024-04-29 | GXIV-4D-6C | 831719 | Modules A1, A2, and A4 | Replaced Modules A1, A2 and A4 | Performed preventive maintenance Performed functionality test Cleaned ICORE, plunger rod, cartridge bay and filter. | Unit is operational with 4 modules working. | Work with Carlos Origenes/ Gateway board for reporting due to corrosion on one component. | Module SN: 730714,675115,654343 | G 6981 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905323458829893437/5905323458829893437_signature_28.png | Mary Rose D. Perejon | https://www.jotform.com/uploads/macarewecare/222927978372067/5905323458829893437/5905323458829893437_signature_31.png
| 2024-05-03 | 2024-05-03 14:26 | San Luis Rural Health Unit | San Luis, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28033 | DOH | mohan_santos@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 352 |  | Genexpert IV | 831714 | Due for Preventive Maintenance | Performed Maintenance | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905272939031155683/5905272939031155683_signature_28.png | Mohan S. Pangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5905272939031155683/5905272939031155683_signature_31.png
| 2024-05-03 | 2024-05-03 13:50 | Bangkal Health Center | Bangkal, Makati City | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28032 | DOH | mhdbangkalsatlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4306 | 2024-04-30 | GX IV 6C-D | 831875 | Module A3 and A4 for replacement | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced modules A3 and A4 Generated report | Replaced modules A3: 739620 -> 612647 A4: 616042 -> 210012721 | unit is operational 7 out of 8 modules working | Module B4 for reporting | Gx modules 6C | G 6992 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905253287627876060/5905253287627876060_signature_28.png | blessed joy serrano | https://www.jotform.com/uploads/macarewecare/222927978372067/5905253287627876060/5905253287627876060_signature_31.png
| 2024-05-03 | 2024-05-03 10:37 | Rosario Reyes Health Center | 627 San Andres St., Malate, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28031 | DOH | moninabatil16@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1319 | 2024-04-30 | GX IV 6C-D | 831843 | Module A3 for replacement | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced module A3 Generated report | Replaced module A3 694064 -> 210034101 | Unit is operational All modules working | GX due for XpertCheck. End user asking when | Gx Module 6C | G 6993 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905245617623606953/5905245617623606953_signature_28.png | Monina Batil | https://www.jotform.com/uploads/macarewecare/222927978372067/5905245617623606953/5905245617623606953_signature_31.png
| 2024-05-03 | 2024-05-03 13:24 | Candaba Rural Health Unit | Poblacion, Candaba, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28030 | DOH | rochelannlontoc@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 563 |  | GXIV-6C | 831873 | Module A3 error 2035 | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905238317525236252/5905238317525236252_signature_28.png | Dr jun tecson | https://www.jotform.com/uploads/macarewecare/222927978372067/5905238317525236252/5905238317525236252_signature_31.png
| 2024-05-03 | 2024-05-03 13:29 | Orani Rural Health Unit | Brgy. Mulawin Orani Bataan | 3 | Bataan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28029 | DOH | oranirhu@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 554 |  | GXIV-D | 831668 | A2:710695 for module replacement Due For preventive maintenance | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational A2:210052194 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905235735217206292/5905235735217206292_signature_28.png | Gabriel Mathew Magpantay | https://www.jotform.com/uploads/macarewecare/222927978372067/5905235735217206292/5905235735217206292_signature_31.png
| 2024-05-03 | 2024-05-03 13:03 | Philippine Red Cross - Cagayan de Oro Chapter | Capitol Compound, Cagayan de Oro City | 10 | Misamis Oriental | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28028 | COMMERCIAL | mrbc2@redcross.org.ph | 25.6 |  |  | MCS+ | 02H080 | Due for Preventive Maintenance and Calibration | Clean Surface Area Check the SPM, DPM, Weigher, Optic Bowl and Line Sensor Parameters if it is in the acceptable range Check Rollers Run Cycle test for functionality | Perform Preventive Maintenance and Calibration | Machine is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905222244749311037/5905222244749311037_signature_28.png | Phoenix James Dahili | https://www.jotform.com/uploads/macarewecare/222927978372067/5905222244749311037/5905222244749311037_signature_31.png
| 2024-05-03 | 2024-05-03 12:43 | Nagcarlan Rural Health Unit | Nagcarlan Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 28027 | DOH | zekcoronado@gmail.com | 25 |  | 2024-04-29 | GXIV | 821163 | Unit for Preventive maintenance. For module replacement of modules B1, B2 and B3. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B1 (636101), B2 (639428) and B3 (639633) in replacement to defective B1 (892641), B2 (627110) and B3 (862106). | G6979 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905217727719462650/5905217727719462650_signature_28.png | Jojo B. Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5905217727719462650/5905217727719462650_signature_31.png
| 2024-05-03 | 2024-05-03 12:39 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28026 | DOH | luzvimindanavidad2@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1095 |  | GXIV-4D | 838297 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905211525864982251/5905211525864982251_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5905211525864982251/5905211525864982251_signature_31.png
| 2024-05-03 | 2024-05-03 12:07 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 28025 | COMMERCIAL | marcievguevarra@stlukes.com.ph | 26 | ASSAY: Xpert CT/NG, COUNT: 290 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 109 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3228 ASSAY: Xpert HCV VL, COUNT: 5 |  | GX4 D | 839234 | Due for maintenance, last xpertcheck march 2023 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | B1 ,B2 operational B3 and B4 disabled failed xpert check last march 2023 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905191452021394027/5905191452021394027_signature_28.png | Marcie V. Guevarra | https://www.jotform.com/uploads/macarewecare/222927978372067/5905191452021394027/5905191452021394027_signature_31.png
| 2024-05-03 | 2024-05-03 10:43 | Pilar Rural Health Unit | Pilar Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28024 | DOH | aimna_24@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 626 | 2024-04-25 | GXIV-D | 831613 | A2:754415 for module replacement Due for preventive maintenance | Preventive maintenance Module replacement | Performed surface cleaning Performed module replacement Performed cleaning of plunger rod, cartridge bay, icore and filter Performed functionality test Gathered reports | Unit is operational A2:677004 (new module) |  | 1 genexpert module | 47629 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905161065218483463/5905161065218483463_signature_28.png | Marciana I. Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5905161065218483463/5905161065218483463_signature_31.png
| 2024-05-03 | 2024-05-03 11:16 | St.Lukes Medical Center | E. Rodriguez ,QC | NCR | NCR 2nd - Quezon City | Cellular immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 28023 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 30 | ASSAY: Xpert HBV VL, COUNT: 59 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert HPV, COUNT: 151 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX4 D | 807155 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905158932025272406/5905158932025272406_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/5905158932025272406/5905158932025272406_signature_31.png
| 2024-05-03 | 2024-05-02 15:55 | Makati Medical Center (Medilinx) | Amorsolo st,Makati | NCR | NCR 4th - Makati | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | John.Manzano@sodexo.com |  | 28022 | COMMERCIAL | reinaluz.bactol@medilinx.com.ph | 30 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 186 ASSAY: Xpert CT/NG, COUNT: 773 ASSAY: Xpert HPV, COUNT: 796 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2955 ASSAY: Others, COUNT: 15 ASSAY: Xpert HCV VL, COUNT: 78 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1993 | 2024-04-25 | GX4 (10C) | 110001647 | Due for maintenance and form module replacement B4 210066943 - 210094674 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance and module replacement | Operational | Room ambient temp too high | Genexpert 10C module | 11443 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5905065383619355072/5905065383619355072_signature_28.png | Maria Reinaluz Bactol | https://www.jotform.com/uploads/macarewecare/222927978372067/5905065383619355072/5905065383619355072_signature_31.png
| 2024-05-02 | 2024-04-30 09:17 | Mangatarem Rural Health Unit | Gen Luna, Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28021 | DOH | slasheiji1988new@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 661 | 2024-04-04 | GXIV-4-D-6C | 838284 | For Calibration | #NAME? | Performed Calibration | #NAME? |  | XpertCheck Kits | DR 11359 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904546402655334646/5904546402655334646_signature_28.png | Ken Lee Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/5904546402655334646/5904546402655334646_signature_31.png
| 2024-05-02 | 2024-05-02 13:59 | Ramos Rural Health Unit | Ramos Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28020 | DOH | jadeeh8143@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 341 |  | GXIV-D | 831721 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cartridge bay cleaning, icore, filter and plunger rod. Archived test results Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904497165409553658/5904497165409553658_signature_28.png | Emily Jade Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/5904497165409553658/5904497165409553658_signature_31.png
| 2024-05-02 | 2024-05-02 16:47 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 28019 | MACARE-HBDC | laboratory@slec.ph | 29 |  |  | GX4 10C | 110016786 | Unit for XpertCheck | Performed maintenance cleaning Performed XpertCheck procedure Generated Reports | Performed XpertCheck | Unit is operational, all modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904495681817811436/5904495681817811436_signature_28.png | Khariz Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5904495681817811436/5904495681817811436_signature_31.png
| 2024-05-02 | 2024-05-02 15:33 | National Children"s Hospital | E. Rodriquez Ave., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | TROUBLESHOOTING Unit check up | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com | biomed@nch.doh.gov.ph |  | 28018 | COMMERCIAL | arellanojohneric09@gmail.com | 28 |  | 2024-05-02 | MCS+ | 17H632SPG | Latham bowl breakage on their last procedure. | Monitor functionality and simulation run. Functionlity test = completed with no error. | Performed on site troubleshooting. Run funcionality test and simulation run with layham bowl installed in the centrifuge. | Unit is operational. | N/A | None | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904463790525230041/5904463790525230041_signature_28.png | John Eric Arellano | https://www.jotform.com/uploads/macarewecare/222927978372067/5904463790525230041/5904463790525230041_signature_31.png
| 2024-05-02 | 2024-04-30 13:09 | Catmon Health Center | Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28015 | DOH | mariathes0218@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3282 | 2024-04-30 | GXIV | 831903 | Module C3 for replacement | #NAME? | #NAME? | Unit is operational |  | GeneXpert Module SN: 661625 Fan Filter | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904403847707110705/5904403847707110705_signature_28.png | Clara Gloria | https://www.jotform.com/uploads/macarewecare/222927978372067/5904403847707110705/5904403847707110705_signature_31.png
| 2024-05-02 | 2024-04-30 10:52 | Meycauayan City Health Center VI | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28017 | DOH | ronamarieiris_25@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1702 ASSAY: Xpert HIV-1 Viral Load, COUNT: 157 | 2024-04-30 | GXIV | 831904 | Module A4 for replacement | Replaced module A4 (909140) by 616331 | #NAME? | Unit is operational |  | GeneXpert Module SN: 616331 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904403857704188061/5904403857704188061_signature_28.png | Rona Marie Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5904403857704188061/5904403857704188061_signature_31.png
| 2024-05-02 | 2024-05-02 14:20 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 28016 | DOH | laboratory@loveyourself.com.ph | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 449 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1088 | 2024-05-02 | GXIV | 831895 | Modules A1, A3 for replacement | - replaced A1, A3 by 758228, 758231 respectively - cleaned clogged fan filter due to nearby construction | #NAME? | Unit is operational |  | GeneXpert Module SN: 758228, 758231 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904403857709971806/5904403857709971806_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5904403857709971806/5904403857709971806_signature_31.png
| 2024-05-02 | 2024-05-02 14:14 | Taguig Pateros District Hospital | East Service Road, Western Bicutan, Taguig City | NCR | NCR 4th - Taguig | laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28014 | DOH | kristinejane.kjd25@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3541 ASSAY: Xpert HIV-1 Viral Load, COUNT: 567 | 2024-04-30 | GX IV 6C-D | 831638 | Modules A1 and A4 for replacement | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced A1 and A4 Generated report | Replaced nodules A1: 710693 -> 691267 A4: 754638 -> 693385 | Unit is operational 3 modules working A2 for reporting |  | GX modules 6C | G 6995 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904403412327477373/5904403412327477373_signature_28.png | Kristine Jane D. Base | https://www.jotform.com/uploads/macarewecare/222927978372067/5904403412327477373/5904403412327477373_signature_31.png
| 2024-05-02 | 2024-05-02 13:44 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | Check up machine | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 28013 | DOH | laureanjo@yahoo.com | 28 |  |  | Gx 4 | 838527 | Detects High Temp in IQR | Clean fan filter Clean module bay Module D4 for Module Replacement | Clean fan filter Clean module bay | Systems operational Detect high temperature in IQR | Laboratory's aircon is incapable of maintaining the right temperature for genexpert machines |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904388430612819276/5904388430612819276_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5904388430612819276/5904388430612819276_signature_31.png
| 2024-05-02 | 2024-05-02 09:29 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28012 | DOH | arleneabad32@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 733 | 2024-04-04 | GXIV-4-D-6C | 838288 | For Calibration | #NAME? | Performed Calibration | - Unit is fully operational - Did not perform XpertCheck on Newly Replaced Module B2(739420) - Module B4 (698768) failed XpertCheck, For Reporting - Module B1 (808218) and B3(635767) Passed XpertCheck |  | XpertCheck Kits | DR 11360 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904384272657879435/5904384272657879435_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5904384272657879435/5904384272657879435_signature_31.png
| 2024-05-02 | 2024-04-30 09:17 | Mangatarem Rural Health Unit | Gen Luna, Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 28011 | DOH | slasheiji1988new@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 661 | 2024-04-04 | GXIV-4-D-6C | 838284 | For Calibration | #NAME? | Performed Calibration | - Unit is fully operational - |  | XpertCheck Kits | DR 11359 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904382402653028083/5904382402653028083_signature_28.png | Ken Lee Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/5904382402653028083/5904382402653028083_signature_31.png
| 2024-05-02 | 2024-04-30 14:07 | National Kidney and Transplant Institute | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 28010 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1381 ASSAY: Others, COUNT: 72 ASSAY: Xpert HBV VL, COUNT: 1011 ASSAY: Xpert HPV, COUNT: 71 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2319 ASSAY: Xpert SARS-CoV-2, COUNT: 15566 ASSAY: Xpert HCV VL, COUNT: 115 |  | GX16 | 110005587 | For additional troubleshooting A1 module s/n896977. | Performed cleaning of filter. Restart system. Performed self test multiple times | Troubleshooting | A1,A2,A4,D2 reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904363573324744183/5904363573324744183_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5904363573324744183/5904363573324744183_signature_31.png
| 2024-05-02 | 2024-05-02 11:52 | Makati Medical Center | Amorsolo st., Makati | NCR | NCR 4th - Makati | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | troxoctubre@gmail.com |  | 28009 | COMMERCIAL | blood.bank@makatimed.net.ph | 27 |  | 2024-04-11 | MCS + | 97H330 | Malfunction of blue pinch valve | Delivered and installed of pinch valve Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running. | Repair and replacement of pinch valve | Operational |  | Pinch valve assy | 97793/11377 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904346283325707076/5904346283325707076_signature_28.png | Life An Arit | https://www.jotform.com/uploads/macarewecare/222927978372067/5904346283325707076/5904346283325707076_signature_31.png
| 2024-05-02 | 2024-05-02 11:43 | Victoria Rural Health Unit | Brgy. Bulo, Victoria Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 28008 | DOH | diazrmt@yahoo.com | 36 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 758 | 2024-04-25 | GXIV-D | 831669 | A3:699002 for module replacement Due for maintenance UPS troubleshooting | Preventive maintenance Module replacement Troubleshooting | Performed surface cleaning Performed module replacement Performed UPS troubleshooting Performed functionality test Gathered reports | Unit is operational New module SN:210028945 Defective UPS | UPS:AT1000181128ST063 | 1 genexpert module: 210028945 | 47632 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904322039011364193/5904322039011364193_signature_28.png | Alma D. Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5904322039011364193/5904322039011364193_signature_31.png
| 2024-05-02 | 2024-05-02 10:15 | Marikina City Health Office | Marikina Health City Center, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28007 | DOH | mcho.laboratory102018@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 1634 ASSAY: Xpert HIV-1 Viral Load, COUNT: 352 | 2024-04-30 | GX IV 6C-D | 831588 | Module B2 for replacement | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced module B2 Generated report | Replqced module B2 630566 -> 688096 | Unit is operational All modules working |  | Gx module 6C | G 6996 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5904259842321271226/5904259842321271226_signature_28.png | RONICA ANA TADTAD | https://www.jotform.com/uploads/macarewecare/222927978372067/5904259842321271226/5904259842321271226_signature_31.png
| 2024-04-30 | 2024-04-30 15:47 | Philippine Red Cross Davao Chapter | Roxas St. Davao City | 11 | Davao del Sur | Blood Centre | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 28006 | COMMERCIAL | davaocitybloodservices@gmail.com | 24.8 |  |  | MCS+ | 02J021 | Due for Preventive Maintenance and Calibration | Clean Surface Area Check DPM, SPM, and Line Sensor Values Run Cycle Test for Functionality | Perform Preventive Maintenance and Calibration | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902732284003502700/5902732284003502700_signature_28.png | Mars Panes | https://www.jotform.com/uploads/macarewecare/222927978372067/5902732284003502700/5902732284003502700_signature_31.png
| 2024-04-30 | 2024-04-30 09:17 | Atang dela Rama Health Center | Pampanga St. Gagalangin Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdarpya.macare@gmail.com |  |  | 28005 | DOH | monskie_22@yahoo.com | 25 |  |  | GXIV | 831644 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902729859813278834/5902729859813278834_signature_28.png | Monaliza Balcos | https://www.jotform.com/uploads/macarewecare/222927978372067/5902729859813278834/5902729859813278834_signature_31.png
| 2024-04-30 | 2024-04-30 15:39 | Commonwealth Health Center | Katuparan St., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 28004 | DOH | commonwealthclinical@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 114 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2849 | 2024-04-30 | GX IV 6C-D | 831615 | Module A2 for replacement | Checcked and cleaned cartridge bays, icore slots, plunger rod and filter Replaced module A2 Generated report | Replaced Module A2 643802 -> 687061 | Unit is operational All modules working |  | GCIV 6C -D | G 6990 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902724584691408517/5902724584691408517_signature_28.png | Danica Laisa Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5902724584691408517/5902724584691408517_signature_31.png
| 2024-04-30 | 2024-04-30 15:01 | Philippine Red Cross - Ormoc Chapter | Ormoc City | 8 | Leyte | Laboratory | NO | MAINTENANCE | HAEMONETICS | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 28003 | COMMERCIAL | ormocbloodservice@redcross.org.ph | 28 |  |  | MCS+ | 15H043 | None | None | #NAME? | Unit is functional. Calibration parameters are within range. Rubber rollers are still intact and functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902706116865527149/5902706116865527149_signature_28.png | Roselyn Pura | https://www.jotform.com/uploads/macarewecare/222927978372067/5902706116865527149/5902706116865527149_signature_31.png
| 2024-04-30 | 2024-04-30 15:08 | Estevez Memorial Hospital, Inc | Don Juan S. Estevez Street, Guevarra Subdivision, Legazpi City, Albay | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 28002 | COMMERCIAL | reylovendino83@gmail.com | 25 |  |  | GX4-10C | 110017942 | CPU CPH0522F36751 for replacement | Replaced CPU: CPH0522F36751>>CPH1023F45919 Cleaned cartridge bays, plunger rod, icore Generate reports | Replaced Defective CPU Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902704861109731723/5902704861109731723_signature_28.png | Rey Lovendino | https://www.jotform.com/uploads/macarewecare/222927978372067/5902704861109731723/5902704861109731723_signature_31.png
| 2024-04-30 | 2024-04-30 14:32 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 28001 | MACARE-HBDC | jblmgh.id@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 89 ASSAY: Xpert HCV VL, COUNT: 8 ASSAY: Xpert HIV-1 Viral Load, COUNT: 588 |  | GXIV-10c | 110013748 | Due for Preventive Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902684013227768693/5902684013227768693_signature_28.png | Criselle medina | https://www.jotform.com/uploads/macarewecare/222927978372067/5902684013227768693/5902684013227768693_signature_31.png
| 2024-04-30 | 2024-04-30 14:15 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 28000 | DOH | burauenlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 471 | 2024-04-30 | GXIV | 838551 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902672577373864825/5902672577373864825_signature_28.png | mary ann c. sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5902672577373864825/5902672577373864825_signature_31.png
| 2024-04-30 | 2024-04-30 10:56 | UST Hospital | Dapitan St., Manila | NCR | NCR 1st - Manila | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27998 | MACARE-HBDC | usth_bloodbank@yahoo.com | 24 |  |  | MCS+ | 02F069 | Unit is due for Maintenance | Preventive Maintenance | Performed Simulation. Checked A/D Values and Parameters. Cleaned Surface Area. Checked Valves and Rollers. Cleaned Sensors. | All Values Passed. Unit is in Good Condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902672508418863780/5902672508418863780_signature_28.png | Patricia Candelaria | https://www.jotform.com/uploads/macarewecare/222927978372067/5902672508418863780/5902672508418863780_signature_31.png
| 2024-04-30 | 2024-04-30 14:10 | Carmona Rural Health Unit | Loyola St, Carmona, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27999 | DOH | irishzenreyes2197@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1507 |  | GXIV-D-6C | 821216 | Unit is due for Maintenance. Module 1 with 5007 Error. | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Available. | A1/ 782522 for Observation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902672518415741411/5902672518415741411_signature_28.png | Irishzen reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5902672518415741411/5902672518415741411_signature_31.png
| 2024-04-30 | 2024-04-30 13:26 | OFW Hospital | City of San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27997 | MACARE-HBDC | tiquifrancisca58@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 17 |  | GXIV-10c | 110018556 | Due for Preventive Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902644433222052787/5902644433222052787_signature_28.png | Francisca Tiqui | https://www.jotform.com/uploads/macarewecare/222927978372067/5902644433222052787/5902644433222052787_signature_31.png
| 2024-04-30 | 2024-04-30 12:57 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | REPAIR MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27996 | COMMERCIAL | pshclinlab@gmail.com | 23 |  |  | MCS+ | 02D167 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902624588453248009/5902624588453248009_signature_28.png | Roda Fatima Flores,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5902624588453248009/5902624588453248009_signature_31.png
| 2024-04-30 | 2024-04-30 12:41 | Southern Philippines Medical Center | JP. Laurel, Davao City | 11 | Davao del Sur | ICU 2 | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | hkdrdsales@gmail.com |  | 27995 | COMMERCIAL | icu2estuya@gmail.com | 25.6 |  |  | MCS+ | 15H034G | Due for Preventive Maintenance and Calibration | Clean Surface Area Check the values for SPM, DPM and Line Sensor if is it on range | Perform Preventive Maintenance and Calibration | Machine is working and operational | Maybe there is a debris underneath the centrifuge. Documented the errors further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902619632014666390/5902619632014666390_signature_28.png | Fernando p. Hinong II | https://www.jotform.com/uploads/macarewecare/222927978372067/5902619632014666390/5902619632014666390_signature_31.png
| 2024-04-30 | 2024-04-30 11:37 | Cebu TB Reference Lab | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27994 | DOH | asmantiquilla.pbsp@gmail.com | 28 |  |  | Gx 4 | 838263 | Due date for Site Visit | Site Visit Clean Fan Filter Clean module bay | Site Visit Clean Fan Filter Clean module bay | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902577427141144664/5902577427141144664_signature_28.png | Allan S. Mantiquila | https://www.jotform.com/uploads/macarewecare/222927978372067/5902577427141144664/5902577427141144664_signature_31.png
| 2024-04-30 | 2024-04-30 11:03 | Molave Rural Health unit | Molave, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Data Gathering and site inspection | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27993 | DOH | aldrinacain@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 765 | 2024-04-30 | GX 4 6colors | 838233 | None | Preventive Maintenence, data gathering and site inspection | Preventive Maintenence, data gathering and site inspection | All modules are all available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902567187374029775/5902567187374029775_signature_28.png | Aldrin llanes Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/5902567187374029775/5902567187374029775_signature_31.png
| 2024-04-30 | 2024-04-30 11:07 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27992 | MACARE-HBDC | jblmgh.hpv@gmail.com | 28 | ASSAY: Xpert HBV VL, COUNT: 105 ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert SARS-CoV-2, COUNT: 487 ASSAY: Xpert HCV VL, COUNT: 24 ASSAY: Xpert HIV-1 Viral Load, COUNT: 573 |  | GXIV | 812571 | Due for Preventive Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902560832011020246/5902560832011020246_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/5902560832011020246/5902560832011020246_signature_31.png
| 2024-04-30 | 2024-04-30 11:03 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27991 | MACARE-HBDC | jblmgh.hpv@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 89 ASSAY: Xpert SARS-CoV-2, COUNT: 899 |  | GXIV-10C | 110012825 | Due for Preventive Maintenance | Performed Maintenance | Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902560472028754031/5902560472028754031_signature_28.png | Aurita Cudia | https://www.jotform.com/uploads/macarewecare/222927978372067/5902560472028754031/5902560472028754031_signature_31.png
| 2024-04-30 | 2024-04-26 08:20 | Tondo Medical Center | H. Lopez Blvd. Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27990 | PBSP | microscan@tmc.doh.gov.ph | 25 |  |  | GX16 | 844121 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. For reporting of module B2 (210063240) and C4 (618126) due to temperature error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902490286766228460/5902490286766228460_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5902490286766228460/5902490286766228460_signature_31.png
| 2024-04-30 | 2024-04-30 09:00 | Tondo Health Center | Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27989 | DOH | ma.lorinamanera@gmail.com | 25 |  | 2024-04-29 | GXIV | 831620 | Unit for preventive maintenance. For module replacement of module A4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A4 (659578) in replacement to defective A4 (644057). | G6980 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902489576761113625/5902489576761113625_signature_28.png | Ma. Lorina M. Lontoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5902489576761113625/5902489576761113625_signature_31.png
| 2024-04-30 | 2024-04-26 08:18 | Tondo Medical Center | H. Lopez Blvd. Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27986 | PBSP | microscan@tmc.doh.gov.ph | 25 |  |  | GXIV | 834430 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902489566763297176/5902489566763297176_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5902489566763297176/5902489566763297176_signature_31.png
| 2024-04-30 | 2024-04-26 08:20 | Tondo Medical Center | H. Lopez Blvd. Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27987 | PBSP | microscan@tmc.doh.gov.ph | 25 |  |  | GX16 | 844121 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. For reporting of module B2 (210063240) and C4 (618126) due to temperature error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902489566763670801/5902489566763670801_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5902489566763670801/5902489566763670801_signature_31.png
| 2024-04-30 | 2024-04-26 08:16 | Tondo Medical Center | H. Lopez Blvd. Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27988 | DOH | microscan@tmc.doh.gov.ph | 25 |  |  | GXIV | 831872 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. For reporting of module A4 (688318) due to temperature error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5902489566764603831/5902489566764603831_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5902489566764603831/5902489566764603831_signature_31.png
| 2024-04-29 | 2024-04-16 11:00 | Morong Rural Health Unit | Morong | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27985 | DOH | morongrhu@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 548 |  | GX-IV-D-6C | 821213 | For maintenance | Performed maintenence | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901826543611089786/5901826543611089786_signature_28.png | Necitas v. Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5901826543611089786/5901826543611089786_signature_31.png
| 2024-04-29 | 2024-04-18 14:49 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27984 | DOH | gemmponce11@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 89 |  | GX-IV-D-6C | 821292 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Checked voltage Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901826123614045707/5901826123614045707_signature_28.png | Gemma Ponce | https://www.jotform.com/uploads/macarewecare/222927978372067/5901826123614045707/5901826123614045707_signature_31.png
| 2024-04-29 | 2024-04-19 11:45 | Pasig Treatment Hub | Dr. Garcia St. | NCR | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27983 | MACARE-HBDC | pathmedtech@gmail.com | 19 | ASSAY: Xpert CT/NG, COUNT: 18 ASSAY: Xpert HIV-1 Viral Load, COUNT: 752 |  | GX-IV-L-6C | 821648 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Modules A1/733228, A2/734493 and A4/681602 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901825213612440751/5901825213612440751_signature_28.png | Jhozza Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5901825213612440751/5901825213612440751_signature_31.png
| 2024-04-29 | 2024-04-19 13:52 | Nagpayong Super Health Center | Nagpayong | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27982 | DOH | mnoreend@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1772 |  | GX-IV-D-6C | 831905 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Module B1/758209 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901824773616315323/5901824773616315323_signature_28.png | Ma.Noreen DeLeon | https://www.jotform.com/uploads/macarewecare/222927978372067/5901824773616315323/5901824773616315323_signature_31.png
| 2024-04-29 | 2024-04-16 13:32 | Teresa Rural Health Unit | Teresa | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27981 | DOH | roseannemaetamayo@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 954 |  | GX-IV-D-6C | 820710 | For maintenance | Performed maintenence | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901821863619607240/5901821863619607240_signature_28.png | Roseanne Mae Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5901821863619607240/5901821863619607240_signature_31.png
| 2024-04-29 | 2024-04-18 13:28 | Sta. Cruz Rural Health Unit | Sta. Cruz | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27980 | DOH | jezzaleal@gmail.com | 18 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1469 |  | GX-IV-D-6C | 820229 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Module D2/722095 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901820783613661419/5901820783613661419_signature_28.png | Jezza Nequinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5901820783613661419/5901820783613661419_signature_31.png
| 2024-04-29 | 2024-04-16 13:32 | Teresa Rural Health Unit | Teresa | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27979 | DOH | roseannemaetamayo@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 954 |  | GX-IV-D-6C | 820710 | For maintenance | Performed maintenence | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901819783617378338/5901819783617378338_signature_28.png | Roseanne Mae Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5901819783617378338/5901819783617378338_signature_31.png
| 2024-04-29 | 2024-04-18 13:28 | Sta. Cruz Rural Health Unit | Sta. Cruz | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27978 | DOH | jezzaleal@gmail.com | 18 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1469 |  | GX-IV-D-6C | 820229 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Module D2/722095 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901817903618675341/5901817903618675341_signature_28.png | Jezza Nequinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5901817903618675341/5901817903618675341_signature_31.png
| 2024-04-29 | 2024-04-23 10:20 | Detoxicare Molecular Diagnostics Laboratory Inc. | Shaw, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27977 | CONSORTIUM | mikaela.chester@detoxicare.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV 10C | 110018512 | Unit for installation Unit from Detoxicare Boni | NA | #NAME? | Unit is operational | with Engr. Doctor |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901811503612206308/5901811503612206308_signature_28.png | Maria Mikaela Chester | https://www.jotform.com/uploads/macarewecare/222927978372067/5901811503612206308/5901811503612206308_signature_31.png
| 2024-04-29 | 2024-04-26 12:37 | Marulas 3S Health Station | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27974 | DOH | melissamartin0828@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1478 |  | GXIV | 831696 | Unit for maintenance | NA | #NAME? | Unit is operational Computer monitor for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901810873615392042/5901810873615392042_signature_28.png | Melissa Martin Galit | https://www.jotform.com/uploads/macarewecare/222927978372067/5901810873615392042/5901810873615392042_signature_31.png
| 2024-04-29 | 2024-04-25 12:15 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27975 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 98C002 | Unit for calibration and maintenance | NA | - checked rotor roller pumps - checked calibration parameter - cuff: 72 (offset) , 1596 (full scale) - weigher: 135 (offset), 1588 (full scale) - spm: 2027 (offset), 3232 (full scale) - dpm: 2004 (offset), 3200 (full scale) - performed self test | Unit is operational Calibration certificate to follow |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901810873616185180/5901810873616185180_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/5901810873616185180/5901810873616185180_signature_31.png
| 2024-04-29 | 2024-04-26 14:12 | Punturin 3S Health Station | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27976 | DOH | lukegonzaga09@gmail.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1426 |  | GXIV | 831914 | Unit for maintenance | NA | - checked and cleaned cartridge bay, plunger rod, - generated reports - performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901810873617861511/5901810873617861511_signature_28.png | Judelyn Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5901810873617861511/5901810873617861511_signature_31.png
| 2024-04-29 | 2024-04-25 11:50 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27973 | MACARE-HBDC | marnelli.manuel@gmail.com | 25 |  |  | MCS+ | 08M461 | Unit for calibration and maintenance | NA | - checked rotor roller pumps (good) - checked calibration parameters (within acceptable range) - cuff: 32 (offset), 2124 (full scale) - weigher: 100 (offset), 1598 (full scale) - dpm: 2040 (offset), 3265 (full scale) - spm: 2039 (offset), 3260 (full scale) - performed self test | Unit is operational Calibration certificate to follow |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901810873618410779/5901810873618410779_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/5901810873618410779/5901810873618410779_signature_31.png
| 2024-04-29 | 2024-04-29 12:26 | Baybay City Health Office | Baybay Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27972 | DOH | sajecopooh20@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1120 | 2024-04-29 | GXIV | 838555 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901741868227409415/5901741868227409415_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5901741868227409415/5901741868227409415_signature_31.png
| 2024-04-29 | 2024-04-25 14:15 | Jala-Jala Rural Health Unit | Jala-Jala | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | kreszapaula64@yahoo.com |  | 27971 | DOH | jjrhu@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 401 | 2024-04-25 | GX-IV-D-6C | 820184 | For maintenance For UPS replacement | Performed maintenance Replaced UPS | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Replaced UPS Generated reports | Unit is operational. |  | Old UPS SN: AT1000180910ST0582 New UPS SN: 40353382308 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901712593613094109/5901712593613094109_signature_28.png | paula de guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5901712593613094109/5901712593613094109_signature_31.png
| 2024-04-29 | 2024-04-26 11:05 | San Nicolas Health Center | Asuncion St, San Nicolas, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27969 | DOH | guadex7@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1539 |  | GXIV-D-6C | 831618 | Unit is due for Maintenance | Preventive Maintenance. | Cleaned Cartridge Bay, Plunger Rod and Icore. Checked Filter if Clogged. Performed Self-test. Generate Reports | All Modules are Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901683053612953685/5901683053612953685_signature_28.png | Guadalupe Cube | https://www.jotform.com/uploads/macarewecare/222927978372067/5901683053612953685/5901683053612953685_signature_31.png
| 2024-04-29 | 2024-04-25 14:04 | Bagong Barangay Health Center | West Zamora, Pandacan, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27968 | DOH | thea.peneyra@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 613 |  | GXIV-D-6C | 831694 | Stuck Cartridge on Module 1 Unit is due for Maintenance | Troubleshooting Preventive Maintenance | Removed Stuck Cartridge. Cleaned Cartridge Bay, Plunger Rod and Icore. Checked Filter if Clogged. Performed Self-test. Generate Reports. | All Modules are Operational | A1/210013103 for Observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901683053615501674/5901683053615501674_signature_28.png | Maria Anthea Peneyra | https://www.jotform.com/uploads/macarewecare/222927978372067/5901683053615501674/5901683053615501674_signature_31.png
| 2024-04-29 | 2024-04-26 12:33 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | materials.fabella@gmail.com |  | 27967 | MACARE-HBDC | fabella.mollabgene@gmail.com | 22 | ASSAY: Xpert SARS-CoV-2, COUNT: 1944 ASSAY: Others, COUNT: 8 |  | GXIV-L-6C | 834953 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Checked and Cleaned Filter. Performed Self-test. Generate Reports. | 2 Modules were Disabled. | A2/ 789440 for Observation(with 5007 error) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901683053619086940/5901683053619086940_signature_28.png | Rallion gio paguia | https://www.jotform.com/uploads/macarewecare/222927978372067/5901683053619086940/5901683053619086940_signature_31.png
| 2024-04-29 | 2024-04-26 12:25 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | materials.fabella@gmail.com |  | 27970 | MACARE-HBDC | fabella.mollabgene@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 102 ASSAY: Xpert MTB/RIF Ultra, COUNT: 16 ASSAY: Xpert SARS-CoV-2, COUNT: 3942 ASSAY: Others, COUNT: 21 |  | GXIV-D-10C | 110015482 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Optics. Cleaned Filter. Performed Self-test. Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901683063611762006/5901683063611762006_signature_28.png | Rallion gio paguia | https://www.jotform.com/uploads/macarewecare/222927978372067/5901683063611762006/5901683063611762006_signature_31.png
| 2024-04-29 | 2024-04-29 10:07 | Pioduran Rural Health Unit | Caratagan, Pioduran, Albay | 5 | Albay | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27966 | DOH | mitzijc@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 728 |  | GX4-6C | 838264 | Gx for Calibration | XpertCheck Result: B1/807820 - passed B2/807832 - passed B3/807835 - passed B4/807770 - passed | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901659188421419258/5901659188421419258_signature_28.png | Mitzi Cedeño | https://www.jotform.com/uploads/macarewecare/222927978372067/5901659188421419258/5901659188421419258_signature_31.png
| 2024-04-29 | 2024-04-29 09:55 | Zamboanga del sur medical center | Zamboanga del sur medical center | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | limboemillie@gmail.com |  | 27965 | DOH | niagnozia54@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1833 ASSAY: Xpert HIV-1 Viral Load, COUNT: 159 | 2024-04-29 | GX 6 colors | 838204 | Module 2 and Module 4 pending for replacement | Preventive Maintenence, Site Inspection and data gathering. | Preventive Maintenence, Site Inspection and data gathering. | Module 2 and 4 for replacement, module 1 and 3 are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901656837016223967/5901656837016223967_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/5901656837016223967/5901656837016223967_signature_31.png
| 2024-04-29 | 2024-04-26 15:23 | Labo District Hospital | Talobatib, Labo, Camarines Norte | 5 | Camarines Norte | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27964 | DOH | esteves110597@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1401 |  | GX4-6C | 838207 | Gx for Calibration | XpertCheck Result: B1/807308 - passed B2/878474 - new module B3/719806 - new module B4/807323 - passed Cleaned cartridge bays, plunger rod, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901655268429334454/5901655268429334454_signature_28.png | JANELA JOYCE ESTEVES | https://www.jotform.com/uploads/macarewecare/222927978372067/5901655268429334454/5901655268429334454_signature_31.png
| 2024-04-29 | 2024-04-26 13:33 | Vinzons Rural Health Unit | Brgy. Dos, Vinzons | 5 | Camarines Norte | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27963 | DOH | bardonamabel@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 881 |  | GX4-6C | 838576 | Gx for Calibration | XpertCheck Result: B1/793494 - passed B2/862906 - passed B3/210091325 - new module B4/675496 - passed | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5901654858421984863/5901654858421984863_signature_28.png | Amabel A.Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/5901654858421984863/5901654858421984863_signature_31.png
| 2024-04-28 | 2024-04-23 10:55 | Baguio Health Services Office | T. Alonzo, Baguio City | CAR | Benguet | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27962 | DOH | sherryannemarrero@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1878 | 2024-04-04 | GXIV-4-D-6C | 838281 | Due for Calibration and Preventive Maintenance | #NAME? | Performed Preventive Maintenance | - Unit is fully operational - Unable XpertCheck on newly installed modules B1 (678653) and B2 ( 651808) - Modules B3 (716721) and B4(210005712) passed XpertCheck |  | XpertCheck Kits | DR 11363 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5900942562652743739/5900942562652743739_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5900942562652743739/5900942562652743739_signature_31.png
| 2024-04-26 | 2024-04-26 14:01 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | avabichuela3216@gmail.com |  | 27961 | MACARE-HBDC | molecularlab@phc.gov.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 327 ASSAY: Xpert SARS-CoV-2, COUNT: 1127 |  | GX16D | 838655 | For data gathering and PM | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Performed pms | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899499397918116280/5899499397918116280_signature_28.png | PHILIPPINE HEART CENTER | https://www.jotform.com/uploads/macarewecare/222927978372067/5899499397918116280/5899499397918116280_signature_31.png
| 2024-04-26 | 2024-04-26 12:12 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Tbdots | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27960 | MACARE-HBDC | nelsonmiranda791@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1240 |  | GX4 (10c) | 110014085 | B4 module error continous booting, | Fixed twist cable Reset home to home position After performed troubleshooting error still occur | Troubleshooting | B4 unoperational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899498897912843507/5899498897912843507_signature_28.png | Nelson p. Miranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5899498897912843507/5899498897912843507_signature_31.png
| 2024-04-26 | 2024-04-26 09:11 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27959 | COMMERCIAL | lcpmicrolab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 112 ASSAY: Xpert SARS-CoV-2, COUNT: 13 |  | GX4 L | 804205 | Last xpertcheck 6.30.2021 D2,D3,D4 disabled due to 5007 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance and troubleshooting | D1 operational,D2,D3,D4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899498347911519391/5899498347911519391_signature_28.png | Laurence laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5899498347911519391/5899498347911519391_signature_31.png
| 2024-04-26 | 2024-04-26 09:02 | Lung Center of the Philippines | Quezon Avenue, QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27958 | MACARE-HBDC | lcpmicrolab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 456 ASSAY: Xpert SARS-CoV-2, COUNT: 1385 ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 ASSAY: Xpert HIV-1 Viral Load, COUNT: 96 |  | GX4 D | 838837 | B4 reported already,B2 and B3 module encounter invalid error in HIV,B1 disabled due to intermittent temp error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance and troubleshooting | B1 and B4 disabled, B2 and B3 for observation Last xpertcheck 7.25.23 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899497677913636609/5899497677913636609_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5899497677913636609/5899497677913636609_signature_31.png
| 2024-04-26 | 2024-04-26 17:59 | Governor Celestino Gallares Memorial Hospital | TAGBILARAN, BOHOL | 7 | Bohol | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27957 | DOH | cyrilxoxocaguco@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2206 ASSAY: Xpert HIV-1 Viral Load, COUNT: 197 |  | GX-IV | 820228 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | MODULE B1 FOR REPORTING |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899352095328608073/5899352095328608073_signature_28.png | Cyril Caguco | https://www.jotform.com/uploads/macarewecare/222927978372067/5899352095328608073/5899352095328608073_signature_31.png
| 2024-04-26 | 2024-04-26 15:25 | Tomas Claudio Health Center | Camus St. Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27956 | DOH | arleen_p_alconera@yahoo.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 644 |  | GX-IV 4D | 838492 | Multiple temperature related error to Module B2 | Run functionality test on Module B2 | Perform Module Troubleshooting | Module B2, frequent error on temperature related error Error Sent to Cepheid for further Assessment | Module B2 temporarily disable |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899260632011380892/5899260632011380892_signature_28.png | Arleen P. Alconera | https://www.jotform.com/uploads/macarewecare/222927978372067/5899260632011380892/5899260632011380892_signature_31.png
| 2024-04-26 | 2024-04-26 13:41 | Davao Reproductive Health and Wellness Center | Jacinto St. Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27955 | MACARE-HBDC | rhwclab@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 144 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1044 ASSAY: Xpert HBV VL, COUNT: 9 |  | GX-IV 4D | 839236 | Due for Preventive Maintenance Error at Module B4 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Troubleshooting | Module 4 for under observation since the error encounter is Cartridge Related Error 2037 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899229802012568251/5899229802012568251_signature_28.png | Ellaine Malano, RMT. | https://www.jotform.com/uploads/macarewecare/222927978372067/5899229802012568251/5899229802012568251_signature_31.png
| 2024-04-26 | 2024-04-26 13:49 | Davao Reproductive Health and Wellness Center | Jacinto St. Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27954 | DOH | rhwclab@gmail.com | 26.5 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1006 ASSAY: Xpert MTB/RIF Ultra, COUNT: 144 ASSAY: Xpert HBV VL, COUNT: 9 |  | GX-IV 4d | 838262 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Module B2 and B4 disabled due to multiple errors | Module B2 and B4 waiting for replacement parts |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899228802019239767/5899228802019239767_signature_28.png | Elaine Malano RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5899228802019239767/5899228802019239767_signature_31.png
| 2024-04-26 | 2024-04-26 14:07 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27952 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 74 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3135 ASSAY: Xpert HIV-1 Viral Load, COUNT: 472 |  | GXIV-4D | 838285 | None | None | #NAME? | Waiting for module A4 allocation. Other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899215469063946396/5899215469063946396_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5899215469063946396/5899215469063946396_signature_31.png
| 2024-04-26 | 2024-04-26 14:03 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27953 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 101 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3986 ASSAY: Xpert HIV-1 Viral Load, COUNT: 429 |  | GXIV-4D | 820206 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899215469068543843/5899215469068543843_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5899215469068543843/5899215469068543843_signature_31.png
| 2024-04-26 | 2024-04-26 13:30 | Orani Rural Health Unit | Orani Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27951 | DOH | oranirhu@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 566 |  | GXIV-D | 831688 | UPS not backing up Due for preventive maintenance | Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of cartridge bay, icore, filter and plunger rod UPS troubleshooting Gathered reports Performed functionality test | Unit is operational excet for UPS UPS:AT1000181128ST0196 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899204104722121565/5899204104722121565_signature_28.png | Belinda C. Guinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5899204104722121565/5899204104722121565_signature_31.png
| 2024-04-26 | 2024-04-26 12:59 | West Dadiangas Rural Health Unit | General Santos City, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27950 | DOH | celinejean16@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1313 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 |  | Gx-IV | 838570 | Module 1 H/W failed due to temperature error | Perform Xpertcheck at module 1 for troubleshooting Genexpert Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert module troubleshooting | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899195393824614351/5899195393824614351_signature_28.png | Celine Jean Villanueva, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5899195393824614351/5899195393824614351_signature_31.png
| 2024-04-26 | 2024-04-26 11:26 | Abucay Rural Health Unit | Abucay Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27949 | DOH | parreraledelina22176@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 542 |  | GXIV-D | 831878 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cartridge bay, icor,filter and plunger rod Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899117032319770531/5899117032319770531_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5899117032319770531/5899117032319770531_signature_31.png
| 2024-04-26 | 2024-04-26 09:02 | San Fernando Rural Health Unit | San Fernando, Camarines Sur | 5 | Albay | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27948 | DOH | josephus_garfin@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 |  | GX4-6C | 838496 | Gx for Calibration | XpertCheck Result: B1/722557 - new module B2/723641 - new module B3/725516 - new module B4/725516 - passed Cleaned cartridge bays, plunger rod, icore Generate report | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5899038257512175511/5899038257512175511_signature_28.png | Josephus garfin | https://www.jotform.com/uploads/macarewecare/222927978372067/5899038257512175511/5899038257512175511_signature_31.png
| 2024-04-25 | 2024-04-25 15:41 | Santa Rosa Rural Health Unit | Santa Rosa, Nueva Ecija | 3 | Nueva Ecija | Lahoratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27947 | DOH | laicasalvador.08@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 347 |  | Genexpert GXIV | 831599 | Monitor on and off | Troubleshoot the monitor cables | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898404695515593531/5898404695515593531_signature_28.png | Salvador, Maria Isabel M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5898404695515593531/5898404695515593531_signature_31.png
| 2024-04-25 | 2024-04-25 15:01 | Catigbian District Hospital | CATIGBIAN BOHOL | 7 | Bohol | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27946 | DOH | rheranes@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 661 |  | GX-IV | 838494 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898379751057629163/5898379751057629163_signature_28.png | Rheazel Ranes | https://www.jotform.com/uploads/macarewecare/222927978372067/5898379751057629163/5898379751057629163_signature_31.png
| 2024-04-25 | 2024-04-25 14:54 | CONG. SIMEON G. TORIBIO MEMORIAL HOSPITAL | CARMEN BOHOL | 7 | Bohol | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27945 | DOH | Himaya.vale@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1068 |  | GX-IV | 838394 | #NAME? | #NAME? | -PERFORM BARCODE SCANNER REPLACEMENT: DEF SN : 20112010552690 NEW SN : 1409700507304 -PERFORM CLEANING AND PREVENTIVE MAINTENANCE PROCEDURE -REPLACE FAN FILTER | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898376471054103717/5898376471054103717_signature_28.png | HIMAYA VALE,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5898376471054103717/5898376471054103717_signature_31.png
| 2024-04-25 | 2024-04-25 14:46 | TEODORO B. GALAGAR DISTRICT HOSPITAL | JAGNA BOHOL | 7 | Bohol | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27944 | DOH | applecpalma09@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 862 |  | GX-IV | 838292 | -FOR PREVENTIVE MAINTENANCE -MODULE B2 H/W FAIL ERROR 4004 | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | MODULE B2 FOR REPORTING |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898372181059479066/5898372181059479066_signature_28.png | Therese Arjelin Palma | https://www.jotform.com/uploads/macarewecare/222927978372067/5898372181059479066/5898372181059479066_signature_31.png
| 2024-04-25 | 2024-04-25 13:48 | Cabanatuan City Health Office | Kapt. Pepe, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27943 | DOH | cholab2021@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2360 |  | Genexpert GXIV | 831592 | Due for Calibration | Archive Results | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898339720716006203/5898339720716006203_signature_28.png | Maricris So | https://www.jotform.com/uploads/macarewecare/222927978372067/5898339720716006203/5898339720716006203_signature_31.png
| 2024-04-25 | 2024-04-25 13:08 | Magsaysay Rural Health Unit | Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27942 | DOH | jkakilgan12@gmail.com | 30 |  |  | GX-4 | 838557 | None | None | Preventive Maintenance Data Collection | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898322134614101248/5898322134614101248_signature_28.png | Jesusa N. Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/5898322134614101248/5898322134614101248_signature_31.png
| 2024-04-25 | 2024-04-25 11:25 | Pililla Rural Health Unit | Pililla | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27941 | DOH | tbdotspililla@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 903 |  | GX-IV-D-6C | 821293 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898316562282668839/5898316562282668839_signature_28.png | Pililla RTDL | https://www.jotform.com/uploads/macarewecare/222927978372067/5898316562282668839/5898316562282668839_signature_31.png
| 2024-04-25 | 2024-04-25 11:55 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 27940 | MACARE-HBDC | agimang@asianhospital.com | 22 |  |  | MCS+ | 10B129 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Unit's Surface Area. Checked A/D and Diagnostic Values. Cleaned Optics and Sensors. Cleaned Pump Rollers. Checked Centrifuge Function. Performed Simulation Run. | Unit is Operational. All Values are good. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898281118754089288/5898281118754089288_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5898281118754089288/5898281118754089288_signature_31.png
| 2024-04-25 | 2024-04-25 11:49 | Asian Hospital and Medical Center | Civic Dr, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angelgimang17@gmail.com |  | 27939 | MACARE-HBDC | agimang@asianhospital.com | 22 |  |  | MCS+ | 01F116 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Unit's Surface Area. Checked A/D Values. Cleaned Optics and Sensors. Cleaned Pump Rollers. Checked Centrifuge Function. Performed Simulation Run. | Unit is Operational. All Values are good. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898281098759882232/5898281098759882232_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5898281098759882232/5898281098759882232_signature_31.png
| 2024-04-25 | 2024-04-25 10:23 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27938 | MACARE-HBDC | molecularpathology@asianhospital.com | 23 | ASSAY: Others, COUNT: 62 ASSAY: Xpert HPV, COUNT: 124 ASSAY: Xpert MTB/RIF Ultra, COUNT: 765 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GXIV-D-10C | 110014996 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Checked Filter if Clogged Performed Self-test Generate Reports | All Modules are Operational | B2/210076074 for Observation with Temp Error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898281088754579065/5898281088754579065_signature_28.png | Trish Faye Pastor | https://www.jotform.com/uploads/macarewecare/222927978372067/5898281088754579065/5898281088754579065_signature_31.png
| 2024-04-25 | 2024-04-23 11:39 | Caridad Health Center | Dela Cruz St, Cavite City, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27937 | DOH | seyerlanie85@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1167 |  | GXIV-D-6C | 821227 | GX System No Power | Troubleshooting | Checked Output Voltage of UPS Checked Backup Function of UPS Cleaned Cartridge Bay and Plunger Rod Replaced Clogged Filter Performed Selft-test Generate Reports | 3 Modules are Operational, 1 Disabled UPS is Defective | A4/730140 Disabled. UPS SN:AT1000180910ST0370 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898280998753502313/5898280998753502313_signature_28.png | Melanie Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5898280998753502313/5898280998753502313_signature_31.png
| 2024-04-25 | 2024-04-24 10:01 | General Trias Social Hygiene Clinic | Hospital Rd, General Trias, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27936 | MACARE-HBDC | chogentri.shc@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 500 |  | GXIV-D-10C | 110016785 | Unit is due for XpertCheck | XpertCheck Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Icore Checked Filter if Clogged Performed Self-test Performed XpertCheck Generate Reports | All Modules are Operational | All Modules Passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898280938755293501/5898280938755293501_signature_28.png | Alberto Lagunilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5898280938755293501/5898280938755293501_signature_31.png
| 2024-04-25 | 2024-04-23 09:34 | De La Salle Health Sciences Institute | Congressional Rd, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27935 | CONSORTIUM | acpabruada@dlsmhsi.edu.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2638 | 2024-04-15 | GX16-D-6C | 842166 | Module B4 Failed in XpertCheck | Module Replacement | Replaced Module 4. Performed Self-test. Checked if Filter is Clogged. Generate Reports. | All Modules are Operational | B4: 880257-> 676646 | Genexpert Module | 11399 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898280888756770399/5898280888756770399_signature_28.png | Angelita C. Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5898280888756770399/5898280888756770399_signature_31.png
| 2024-04-25 | 2024-04-24 11:15 | Rosario Rural Health Unit | Scout Torillo St, Poblacion, Rosario, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27934 | DOH | arlenepenaarmijo@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1044 |  | GXIV-D-6C | 838537 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay and Plunger Rod Replaced Clogged Filter Performed Self-test Generate Reports | 3 Modules are Operational, 1 Defective Monitor with Screen Problem | B2 with Lost Comm Problem |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898280878751154263/5898280878751154263_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/5898280878751154263/5898280878751154263_signature_31.png
| 2024-04-25 | 2024-04-25 10:45 | San Pedro City Health Office 1 | San Pedro Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27933 | DOH | venersantos0425@gmail.com | 25 |  |  | GXIV | 821224 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898231960507026789/5898231960507026789_signature_28.png | Solita L.Trinidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5898231960507026789/5898231960507026789_signature_31.png
| 2024-04-25 | 2024-04-24 15:10 | Batis Health Center | F. Manalo St. | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27932 | DOH | almacenaio16@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1891 ASSAY: Xpert HIV-1 Viral Load, COUNT: 98 |  | GX-IV-D-6C | 831598 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports | Unit is operational. Modules A1/210091071 and A4/879974 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898230490337224065/5898230490337224065_signature_28.png | Rosario Almacen | https://www.jotform.com/uploads/macarewecare/222927978372067/5898230490337224065/5898230490337224065_signature_31.png
| 2024-04-25 | 2024-04-25 09:23 | San Pedro Rural Health Unit 2 | Brgy. Langgam, San Pedro Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27931 | DOH | leniemahalin@gmail.com | 27 |  |  | GXIV | 821189 | Unit for preventive maitenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5898179940509767183/5898179940509767183_signature_28.png | Lenita Mahalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5898179940509767183/5898179940509767183_signature_31.png
| 2024-04-24 | 2024-04-24 17:17 | CANDIJAY COMMUNITY HOSPITAL | CANDIJAY BOHOL | 7 | Bohol | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27930 | DOH | angelinaredulla@yahoo.com.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 953 |  | GX-IV | 838396 | #NAME? | #NAME? | -PERFORM CLEANING OF GENEXPERT WORKSTATION -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE B2 SN 860993 CALIBRATION, SERVICE REQUIRED | EQUIPMENT IS OPERATIONAL | MODULE B2 FOR SERVICE |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897596814336027437/5897596814336027437_signature_28.png | ANGELINA O. REDULLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5897596814336027437/5897596814336027437_signature_31.png
| 2024-04-24 | 2024-04-24 17:13 | GARCIA MEMORIAL PROVINCIAL HOSPITAL | TALIBON, BOHOL | 7 | Bohol | TB DOTS | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27929 | DOH | ybanezrosalina19@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1486 |  | GX-IV | 838203 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897594234338930022/5897594234338930022_signature_28.png | ROSALINA A. YBANEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5897594234338930022/5897594234338930022_signature_31.png
| 2024-04-24 | 2024-04-24 14:06 | Cardinal Santos Medical Center | Wilson St. | NCR | NCR 2nd - San Juan | Microbiology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27928 | COMMERCIAL | microbiology.csmc@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 126 ASSAY: Xpert HPV, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1084 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned: -cartridge bay -plunger rod -icore -filter Generated reports Generated gxx file | Unit is operational. Modules B1/210022152 and B4/210048235 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897487717461879637/5897487717461879637_signature_28.png | Ellaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5897487717461879637/5897487717461879637_signature_31.png
| 2024-04-24 | 2024-04-24 11:53 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27927 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 23 | ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 917 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-XVI-D-6C | 838627 | For maintenance | Performed maintenance. | Performed maintenance Cleaned -cartridge bay -plunger rod -icore -filter Generated reports Generated gxx file Performed additional troubleshooting for module C3. | Unit is operational. Modules A1/860941, A2/860940, B3/702534, B4/860923, C2/668748, D2/620837 and D4/661946 not operational. Reported already. Module C3/662448 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897487717464053212/5897487717464053212_signature_28.png | Justine Lorenz Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5897487717464053212/5897487717464053212_signature_31.png
| 2024-04-24 | 2024-04-23 14:29 | Tropical Disease Foundation | Amorsolo st.,Makati City | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27926 | COMMERCIAL | aageronimo@tdf.org.ph | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 2238 | 2024-04-18 | GX4 D R1 | 705175 | Due for Maintenance and Xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance and Xpertcheck | A1 686966 module reported already confirmed defective for quotation replacement. A2,A3,A4 passed in xpertcheck. |  | Xpertcheck kits | 97851 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897482203618000512/5897482203618000512_signature_28.png | Anthony Geronimo | https://www.jotform.com/uploads/macarewecare/222927978372067/5897482203618000512/5897482203618000512_signature_31.png
| 2024-04-24 | 2024-04-24 11:07 | ICHHD | Pedro Gil St. Ermita Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27925 | MACARE-HBDC | dcsantos9@up.edu.ph | 25 |  |  | GXIV | 808015 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897472967919175024/5897472967919175024_signature_28.png | Dana Kamylle Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5897472967919175024/5897472967919175024_signature_31.png
| 2024-04-24 | 2024-04-24 12:08 | Valenzuela Social Hygiene Clinic | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27924 | COMMERCIAL | valenzuelacity.shc@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 134 |  | GXIV 10C | 110020699 | Unit for maintenance | NA | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897417821284501353/5897417821284501353_signature_28.png | Marlette R. Delino | https://www.jotform.com/uploads/macarewecare/222927978372067/5897417821284501353/5897417821284501353_signature_31.png
| 2024-04-24 | 2024-04-24 10:47 | Pilar Rural Health Unit | Calongay, Pilar, Sorsogon | 5 | Sorsogon | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27923 | DOH | manalangmark88@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 |  | GX4-6C | 838266 | GX for Calibration | XpertCheck Result: B1/807683 - passed B2/650253 - new module B3/880428 - new module B4/615874 - passed Cleaned cartridge bays, plunger rod, icore Generate report | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897382126618783737/5897382126618783737_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5897382126618783737/5897382126618783737_signature_31.png
| 2024-04-24 | 2024-04-24 11:04 | UP-NIH | NIH Central Lab, Ermita Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27922 | MACARE-HBDC | vvluzentales@up.edu.ph | 25 |  | 2024-04-22 | GX16-10C | 110014186 | Unit for Preventive Maintenance. For module replacement of module B4. | Preventive maintenance and module replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B4 (887676) in replacement to defectivr B4 (210070221). | DR#11431 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897376813822704664/5897376813822704664_signature_28.png | Vivienne V Luzentales | https://www.jotform.com/uploads/macarewecare/222927978372067/5897376813822704664/5897376813822704664_signature_31.png
| 2024-04-24 | 2024-04-19 09:37 | Victor R. Potenciano Medical Center | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27921 | CONSORTIUM | oro_janet@yahoo.com | 25 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 833 | 2024-04-16 | GXIV | 835600 | Modules B1, B4 for replacement | Replaceed modules B1, B4 | #NAME? | Unit is operational |  | GeneXpert Module SN: 714862, 757409 | 11414 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897351137222068919/5897351137222068919_signature_28.png | Janet Oro | https://www.jotform.com/uploads/macarewecare/222927978372067/5897351137222068919/5897351137222068919_signature_31.png
| 2024-04-24 | 2024-04-24 10:08 | Bagumbong Health Center | Bagumbong, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27920 | DOH | jenilyn_duenas@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2833 |  | GXIV | 831907 | Unit for maintenance Intermittent errors on modules UPS for check up | Performed maintenance | #NAME? | Unit is operational UPS for reporting (SN: AT1000180910ST0399) UPS is recommended for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5897350027225925609/5897350027225925609_signature_28.png | Jenilyn Duenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5897350027225925609/5897350027225925609_signature_31.png
| 2024-04-23 | 2024-04-23 15:22 | Cotabato Regional and Medical Center - Molecular Laboratory | Sinsuat Avenue, Cotabato City | BARMM | Maguindanao | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27919 | DOH | crmcmolecularlab@gmail.com | 23.6 | ASSAY: Xpert SARS-CoV-2, COUNT: 13394 |  | GX-IV 4D | 837662 | Error at Module 4 Install new assay for 4Plex | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Import Assay for 4plex | Perform Preventive Maintenance and Module Troubleshooting Install new assay | Module 2 and Module 4 report for module replacement Successfully installed the new assay |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5896694050313232751/5896694050313232751_signature_28.png | Chareasse love yu utto | https://www.jotform.com/uploads/macarewecare/222927978372067/5896694050313232751/5896694050313232751_signature_31.png
| 2024-04-23 | 2024-04-23 13:38 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27918 | DOH | wcpatron11x4@gmail.com | 25 |  |  | GXIV | 831534 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5896606368316090757/5896606368316090757_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5896606368316090757/5896606368316090757_signature_31.png
| 2024-04-23 | 2024-04-23 13:42 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27917 | PBSP | wcpatron11x4@gmail.com | 25 |  |  | GXIV | 812561 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5896605848314205355/5896605848314205355_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5896605848314205355/5896605848314205355_signature_31.png
| 2024-04-23 | 2024-04-23 11:05 | Cotabato Regional and Medical Center - Blood Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Blood Center | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Jesson Tupas |  | jessontupas.macaremed@gmail.com | hkdrdsales@gmail.com |  | 27916 | COMMERCIAL | crmcbloodcenter2016@gmail.com | 24.7 |  |  | MCS+ Heamonetics | 15H042G | Line Sensor not within acceptable range; Acceptable Range: 2866-3276, Actual Value; 2598-2600 | Perform Calibration Perform cycle test using kits Cleaning Parts Replacement (changing rollers) Problem reporting | Perform MCS+ Heamonetics Calibration | Line Sensor for monitoring due to line sensor value out of range Line Sensor recommendation for replacement | Line Sensor recommendation for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5896559088211669688/5896559088211669688_signature_28.png | Rhea L. Macion | https://www.jotform.com/uploads/macarewecare/222927978372067/5896559088211669688/5896559088211669688_signature_31.png
| 2024-04-23 | 2024-04-23 09:33 | Bataan Genaral Hospital and Medical Center | Manahan St. Tenejero Balanga City, Bataan | 3 | Bataan | PCR | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27915 | MACARE-HBDC | honeyeichel@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 29 ASSAY: Xpert SARS-CoV-2, COUNT: 3436 ASSAY: Xpert HCV VL, COUNT: 8 ASSAY: Xpert HIV-1 Viral Load, COUNT: 59 | 2024-02-14 | GXIV-D 10C | 110017894 | Due for Calibration | Xpertcheck Assay installation Preventive maintenance | Performed Suraface Cleaning Performed cleaning of catridge bay, icore, filter and plunger rod Installation of xpertcheck application Installation of assay(RSV) Archived test results Performed xpertcheck Performed functionality test Gathered reports | Unit is operational, all modules passed calibration B1: 210010244 unsuported for xpertcheck(newly replaced module) |  | 3 xpertcheck kits | 11221 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5896553582871505139/5896553582871505139_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5896553582871505139/5896553582871505139_signature_31.png
| 2024-04-23 | 2024-04-23 11:24 | Minglanilla Rural Health Unit | Pide 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27914 | DOH | brendagonzagaparedes@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1606 |  | GX4 | 838512 | Hardware failure in Module 4 SN# 86234 due to 1Error 4009 and Error 4012 | Adjust room temperature and direct the aircon to the genexpert machine. Generate data IQR, Sys log, Assay Statistics and Archive data files. Observed Module 4 for a week. Cut-off mode of the Air-con too long to recover its desired set-temperature. | Adjust room temperature and direct the aircon to the genexpert machine. Generate data IQR, Sys log, Assay Statistics and Archive data files. Observed Module 4 for a week. | Module 4 for observation and expert check calibration after 1 week. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5896524681113039583/5896524681113039583_signature_28.png | Brenda Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/5896524681113039583/5896524681113039583_signature_31.png
| 2024-04-23 | 2024-04-23 09:18 | Detoxicare Molecular Diagnostics Laboratory Inc. | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27913 | CONSORTIUM | detoxicare.clinical@gmail.com | 25 |  |  | GXIV 10C | 110018312 | Unit for transfer | NA | #NAME? | Unit is operational Unit for transfer | with Engr. Doctor |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5896479735326753988/5896479735326753988_signature_28.png | Regene Mendezabal | https://www.jotform.com/uploads/macarewecare/222927978372067/5896479735326753988/5896479735326753988_signature_31.png
| 2024-04-23 | 2024-04-23 10:13 | Parola Health Center | Parola St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27912 | DOH | olartecharlie@gmail.com | 25 |  |  | GXIV | 831871 | Unit for preventive maintenance. | Preventive maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5896478862166142244/5896478862166142244_signature_28.png | Charlie Olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5896478862166142244/5896478862166142244_signature_31.png
| 2024-04-22 | 2024-04-11 14:34 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 27911 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2999 ASSAY: Xpert SARS-CoV-2, COUNT: 4284 ASSAY: Others, COUNT: 10 | 2024-03-25 | GX16 | 110014171 | Due for xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | C3,B3,D1 newly replaced module |  | Xpertcheck kits | 11331 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5889707453614855825/5889707453614855825_signature_28.png | Jessica Faye Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5889707453614855825/5889707453614855825_signature_31.png
| 2024-04-22 | 2024-04-19 14:28 | Libon Rural Health Unit | Libon, Albay | 5 | Albay | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27910 | DOH | zh.surge@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1209 |  | GX4-6C | 838397 | GX for Calibration | XpertCheck Result: B1/210039783 - Passed B2/860937 - service required B3/860996 - Passed B4/210026807 - Passed Cleaned cartridge bays, plunger rod, icore Generate Report | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational | B2/860937 - service required/for module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5893197613117105049/5893197613117105049_signature_28.png | Cedric V. Ravago | https://www.jotform.com/uploads/macarewecare/222927978372067/5893197613117105049/5893197613117105049_signature_31.png
| 2024-04-22 | 2024-04-19 11:13 | East Avenue Medical Center | East Avenue,QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27909 | MACARE-HBDC | roseannparnes@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1161 | 2024-04-16 | GX4 D (10C) | 110014085 | For module replacement B1 210069919-887287,B3 210069924-718708,B4 210077551-684160 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and Maintenance | Operational |  | Genexpert 10c module | 11411 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895583533613558611/5895583533613558611_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/5895583533613558611/5895583533613558611_signature_31.png
| 2024-04-22 | 2024-04-19 10:20 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 27908 | MACARE-HBDC | eamcmolecularlab@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 31 ASSAY: Xpert HPV, COUNT: 244 ASSAY: Xpert SARS-CoV-2, COUNT: 1209 | 2024-04-16 | GX4- D (10C) | 110016781 | For module replacement B4 210084572 - 210069864 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | Genexpert 10c module | 11410 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895582693614179873/5895582693614179873_signature_28.png | East Avenue Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5895582693614179873/5895582693614179873_signature_31.png
| 2024-04-22 | 2024-04-19 09:09 | National Kidney and Transplant Institute | East Avenue,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27907 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 |  |  | GX16 | 110005587 | A1 encounter error 1002,A2 error 4009 ,4010,4017, A4 error 4017, 4009 4012,D2 error 4009,4010,4017 | Gathered data | Troubleshooting and gathered data | A1,A2,A4,D2 disabled, 12 module operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895581913619681011/5895581913619681011_signature_28.png | Danielle Cortes | https://www.jotform.com/uploads/macarewecare/222927978372067/5895581913619681011/5895581913619681011_signature_31.png
| 2024-04-22 | 2024-03-21 14:25 | University of The East - Ramon Magsaysay Memorial Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27906 | MACARE-HBDC | ljf7403@yahoo.com | 25 |  |  | MCS+ | 13C270 | Unit for calibration and maintenance | Performed calibration | #NAME? | Unit is operational All parameters are within range |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895568613613940141/5895568613613940141_signature_28.png | Leonila Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5895568613613940141/5895568613613940141_signature_31.png
| 2024-04-22 | 2024-03-20 15:18 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27905 | MACARE-HBDC | sah.genexpertlaboratory@gmail.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1165 ASSAY: Xpert SARS-CoV-2, COUNT: 2167 |  | GXIV | 847728 | Unit for troubleshooting | Cleaned filter | #NAME? | - Unit is operational - Room temperature is recommended to be below 30 degrees Celsius (Current ave. temp. is 33 deg.) to avoid temperature errors on modules - Installation of additional air flow machine is recommended (i.e. electric fan pointed towards the back of GeneXpert) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895568193619389698/5895568193619389698_signature_28.png | Yollymel Mary Grejarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5895568193619389698/5895568193619389698_signature_31.png
| 2024-04-22 | 2024-03-21 13:36 | University of The East - Ramon Magsaysay Memorial Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27904 | CONSORTIUM | lpcomendador@uerm.edu.ph | 25 | ASSAY: Xpert C. difficile BT, COUNT: 39 ASSAY: Xpert CT/NG, COUNT: 81 ASSAY: Xpert MTB/RIF Ultra, COUNT: 740 ASSAY: Others, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 25 | 2024-03-07 | GXIV | 847578 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed calibration |  | XpertCheck Kit Lot: 1000410214 | 11273 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895567533614373988/5895567533614373988_signature_28.png | Osang Comendador | https://www.jotform.com/uploads/macarewecare/222927978372067/5895567533614373988/5895567533614373988_signature_31.png
| 2024-04-22 | 2024-04-17 15:51 | Lung Center of the Philippines | Quezon Ave,QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27903 | MACARE-HBDC | lcpbloodbank@gmail.com | 28 |  |  | MCS + | 08H188 | Due for maintenance and calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895567253613161732/5895567253613161732_signature_28.png | Lung Center Of the Philippine Blood Bank Section | https://www.jotform.com/uploads/macarewecare/222927978372067/5895567253613161732/5895567253613161732_signature_31.png
| 2024-04-22 | 2024-03-21 14:25 | University of The East - Ramon Magsaysay Memorial Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27902 | MACARE-HBDC | ljf7403@yahoo.com | 25 |  |  | MCS+ | 13C270 | Unit for calibration and maintenance | Performed calibration | #NAME? | Unit is operational All parameters are within range |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895567033619321215/5895567033619321215_signature_28.png | Leonila Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5895567033619321215/5895567033619321215_signature_31.png
| 2024-04-22 | 2024-04-17 14:49 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27901 | MACARE-HBDC | lcpbloodbank@gmail.com | 28 |  |  | MCS + | 02D166 | Due for Maintenance and Calibration Defective LED bulb white pinch valve | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Maintenance and Calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895566703617832101/5895566703617832101_signature_28.png | Lung Center of the Philippines Blood Bank Section | https://www.jotform.com/uploads/macarewecare/222927978372067/5895566703617832101/5895566703617832101_signature_31.png
| 2024-04-22 | 2024-04-19 09:37 | Victor R. Potenciano Medical Center | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27900 | CONSORTIUM | oro_janet@yahoo.com | 25 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 833 | 2024-04-16 | GXIV | 835600 | Modules B1, B4 for replacement | Replaceed modules B1, B4 | #NAME? | Unit is operational |  | GeneXpert Module SN: 714862, 757409 | 11414 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895566613611080806/5895566613611080806_signature_28.png | Janet Oro | https://www.jotform.com/uploads/macarewecare/222927978372067/5895566613611080806/5895566613611080806_signature_31.png
| 2024-04-22 | 2024-04-19 13:44 | Hagonoy Rural Health Unit | Hagonoy, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27899 | PBSP | lanielyncastro@yahoo.com | 27 |  |  | GXIV | 814158 | CPU for troubleshooting | Replaced BIOS battery | #NAME? | Unit is not operational Unit for follow up reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895566273617015774/5895566273617015774_signature_28.png | Lanielyn Castro Abadam | https://www.jotform.com/uploads/macarewecare/222927978372067/5895566273617015774/5895566273617015774_signature_31.png
| 2024-04-22 | 2024-04-19 13:44 | Hagonoy Rural Health Unit | Hagonoy, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27898 | PBSP | lanielyncastro@yahoo.com | 27 |  |  | GXIV | 814158 | CPU for troubleshooting | Replaced BIOS battery | #NAME? | Unit is not operational Unit for follow up reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895565653614741111/5895565653614741111_signature_28.png | Lanielyn Castro Abadam | https://www.jotform.com/uploads/macarewecare/222927978372067/5895565653614741111/5895565653614741111_signature_31.png
| 2024-04-22 | 2024-04-18 13:28 | Sta. Cruz Rural Health Unit | Sta. Cruz | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27897 | DOH | jezzaleal@gmail.com | 18 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1469 |  | GX-IV-D-6C | 820229 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Module D2/722095 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895560923614031077/5895560923614031077_signature_28.png | Jezza Nequinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5895560923614031077/5895560923614031077_signature_31.png
| 2024-04-22 | 2024-04-19 13:52 | Nagpayong Super Health Center | Nagpayong | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27896 | DOH | mnoreend@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1772 |  | GX-IV-D-6C | 831905 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Module B1/758209 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895560813616634127/5895560813616634127_signature_28.png | Ma.Noreen DeLeon | https://www.jotform.com/uploads/macarewecare/222927978372067/5895560813616634127/5895560813616634127_signature_31.png
| 2024-04-22 | 2024-04-19 11:45 | Pasig Treatment Hub | Dr. Garcia St. | NCR | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27895 | MACARE-HBDC | pathmedtech@gmail.com | 19 | ASSAY: Xpert CT/NG, COUNT: 18 ASSAY: Xpert HIV-1 Viral Load, COUNT: 752 |  | GX-IV-L-6C | 821648 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Modules A1/733228, A2/734493 and A4/681602 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895560813619550360/5895560813619550360_signature_28.png | Jhozza Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5895560813619550360/5895560813619550360_signature_31.png
| 2024-04-22 | 2024-04-18 14:49 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27894 | DOH | gemmponce11@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 89 |  | GX-IV-D-6C | 821292 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Checked voltage Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5895560813611044007/5895560813611044007_signature_28.png | Gemma Ponce | https://www.jotform.com/uploads/macarewecare/222927978372067/5895560813611044007/5895560813611044007_signature_31.png
| 2024-04-19 | 2024-04-19 14:28 | Libon Rural Health Unit | Libon, Albay | 5 | Albay | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27893 | DOH | zh.surge@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1209 |  | GX4-6C | 838397 | GX for Calibration | XpertCheck Result: B1/210039783 - B2/860937 - service required B3/860996 - B4/210026807 - Cleaned cartridge bays, plunger rod, icore Generate Report | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational | B2/860937 - service required/for module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5893197613117105049/5893197613117105049_signature_28.png | Cedric V. Ravago | https://www.jotform.com/uploads/macarewecare/222927978372067/5893197613117105049/5893197613117105049_signature_31.png
| 2024-04-19 | 2024-04-19 12:26 | Sta Rosa Ciry Health Office 2 | Kabesang Moldes St. Sta Rosa | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27892 | DOH | annie.soliman@yahoo.com | 25 |  |  | GXIV | 821156 | Unit for Preventive maintenance. Broken door on module A4. | Preventive maintenance and repair. | Performed preventive maintenance. Performed repair on module door of A4. Performed reliability test. Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5893108192621154270/5893108192621154270_signature_28.png | Santa Rosa CHOII | https://www.jotform.com/uploads/macarewecare/222927978372067/5893108192621154270/5893108192621154270_signature_31.png
| 2024-04-19 | 2024-04-19 11:43 | Meycauayan City Health Center VI | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27891 | DOH | rocelann.caranto14@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1706 |  | GXIV | 831904 | Gathering additional data for reporting | NA | - Generated new reports - Generated .gxx file (Jan. to Apr, 2024) | Unit is operational Additional data for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5893076131225322784/5893076131225322784_signature_28.png | Rocel Ann Caranto | https://www.jotform.com/uploads/macarewecare/222927978372067/5893076131225322784/5893076131225322784_signature_31.png
| 2024-04-19 | 2024-04-19 11:10 | Maddela District Hospital | Diduyon, Maddela, Quirino | 2 | Quirino | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27890 | DOH | mdhlaboratory00@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 324 | 2024-04-19 | Genexpert GXIV | 838469 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration (B1=Service Required) (B2=Passed) (B3=Passed) (B4=Passed) Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5893060053313592154/5893060053313592154_signature_28.png | Alice A. Maranion | https://www.jotform.com/uploads/macarewecare/222927978372067/5893060053313592154/5893060053313592154_signature_31.png
| 2024-04-18 | 2024-04-18 15:58 | MAYOR HILARION A RAMIRO SR MEDICAL CENTER | Maningcol, Ozamiz-Oroquieta National Highway 7200 Ozamis City Misamis Occidental · | 10 | Misamis Oriental | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27889 | DOH | mharsmicrobio@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1795 |  | GX4 | 834437 | Pulled out cartridges in Module 1 SN 784173 and Replaced the Module after cartridge pulled-out. | Pulled out cartridges in Module 1 SN 784173; Module replacement in Module 1 with new module SN # 699652 | Pulled out stocked cartridges in Module 1 SN 784173 and Replaced the Module after cartridge pulled-out. | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892372101113814506/5892372101113814506_signature_28.png | Gretchen Lanas, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5892372101113814506/5892372101113814506_signature_31.png
| 2024-04-18 | 2024-04-18 14:19 | Datu Odin Sinsuat District Hospital | Datu Odin Sinsuat, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27888 | DOH | rowenasalazar1234@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 691 |  | Gx-IV | 838205 | Module 2 multiple temperature error | Perform Xpertcheck at module 2 for troubleshooting Genexpert machine Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance Perform troubleshooting at module 2 | 3 operational modules | Module 2 for service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892331344329551487/5892331344329551487_signature_28.png | Rowena B. Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5892331344329551487/5892331344329551487_signature_31.png
| 2024-04-18 | 2024-04-16 11:40 | Columbio Rural Health Unit | Poblacion Columbio, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27887 | DOH | rochelle.toledo0119@yahoo.com | 28.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 551 |  | Gx-IV | 838556 | N/A | Genexpert Cleaning Checking of errors Data Collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892303374328278181/5892303374328278181_signature_28.png | Rochelle S. Toledo,Rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5892303374328278181/5892303374328278181_signature_31.png
| 2024-04-18 | 2024-04-18 10:00 | Philippine Red Cross Olongapo | Magsaysay St. Olongapo City Zambales | 3 | Zambales | Blood bank | NO | INSTALLATION MAINTENANCE | HAEMONETICS | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27886 | MACARE-HBDC | jona.liza@redcross.org.ph | 25 |  |  | MCS+ | 03AO43 | Setup and installation of Haemonetics | Installation | Performed installation and setup of machine Performed configuration Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892291786115584859/5892291786115584859_signature_28.png | Jona Mae Liza | https://www.jotform.com/uploads/macarewecare/222927978372067/5892291786115584859/5892291786115584859_signature_31.png
| 2024-04-18 | 2024-03-22 14:55 | Detoxicare Molecular Diagnostics Inc. | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27885 | CONSORTIUM | imeehabaluyas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 131 ASSAY: Xpert MTB/RIF Ultra, COUNT: 131 | 2024-03-07 | GXIV 10C | 110018312 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed |  | XpertCheck Kit | 11269 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892287559329925661/5892287559329925661_signature_28.png | Regene Mendezabal | https://www.jotform.com/uploads/macarewecare/222927978372067/5892287559329925661/5892287559329925661_signature_31.png
| 2024-04-18 | 2024-04-18 13:42 | Ramon Rural Health Unit | Laboratory | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27884 | DOH | rhuramonlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 | 2024-03-19 | Genexpert GXIV | 838480 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration (All Modules Passed) Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | 11311 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892287152617095855/5892287152617095855_signature_28.png | Janette T. Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5892287152617095855/5892287152617095855_signature_31.png
| 2024-04-18 | 2024-04-18 13:23 | Manila Central University | Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | lab_mcuhospital@yahoo.com |  | 27883 | CONSORTIUM | purchasing@mcufdtmf.org | 24 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 295 | 2024-04-16 | GXIV 10C | 110018008 | Module B2 for replacement | Replaced module B2 | #NAME? | Unit is operational All 4 modules are available |  | GeneXpert Module SN: 630195 | 11412 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892287059321406030/5892287059321406030_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5892287059321406030/5892287059321406030_signature_31.png
| 2024-04-18 | 2024-04-18 13:16 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27882 | DOH | rhucalbiga2018@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 818 | 2024-04-18 | GXIV | 83409 | Module B4/861165 (H/W failed) | #NAME? | #NAME? | Unit is operational - Module B4 for service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892269046322347168/5892269046322347168_signature_28.png | vanessa mabag | https://www.jotform.com/uploads/macarewecare/222927978372067/5892269046322347168/5892269046322347168_signature_31.png
| 2024-04-18 | 2024-04-18 12:58 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27881 | DOH | buhanginlaboratory@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 921 |  | GX-IV 4D | 838232 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perfrom Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892262393716530131/5892262393716530131_signature_28.png | Catherine Deocampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5892262393716530131/5892262393716530131_signature_31.png
| 2024-04-18 | 2024-04-18 12:17 | Alicia Rural Health Unit | Magsaysay, Alicia, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27880 | DOH | labalicia.rhu@gmail.com | 25 |  | 2024-04-18 | Genexpert GXIV | 838521 | Module 4 error | Performed Calibration (B1, B2 B4 = Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module B3 reported to Cepheid | Xpertcheck kit | 11312 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892235096319722358/5892235096319722358_signature_28.png | Catherine A. Morante | https://www.jotform.com/uploads/macarewecare/222927978372067/5892235096319722358/5892235096319722358_signature_31.png
| 2024-04-18 | 2024-04-18 11:41 | Bagong Barrio Health Center | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27879 | DOH | cruzernitamay@gmail.com | 27 |  |  | GXIV | 831906 | CPU not booting up | Reset RAM | #NAME? | Unit is not operational CPU is for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892211479011031159/5892211479011031159_signature_28.png | Ernita May Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5892211479011031159/5892211479011031159_signature_31.png
| 2024-04-18 | 2024-04-18 11:35 | Chinese General Hospital | Blumentritt St. Sta Cruz Manila | NCR | NCR 1st - Manila | Bacteriology | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 27878 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 |  |  | GXIV-10C | 120002026 | For installation of GeneXpert system. | Installation. | Performed installation of GeneXpert system. Loaded assays. Generated report. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892211220223501694/5892211220223501694_signature_28.png | Maria Odessa R. Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/5892211220223501694/5892211220223501694_signature_31.png
| 2024-04-18 | 2024-04-18 10:50 | Jones Rural Health Unit | Brgy 1, Jones, Isabela | 2 | Isabela | ene | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27877 | DOH | Cheryle_rmt@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1531 | 2024-04-18 | Genexpert GXIV | 838484 | Module 1 and 4 Hardware Failure | Performed Calibration (All Modules Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 1 and 4 for observation | Xpertcheck kit | 11309 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892185817514746341/5892185817514746341_signature_28.png | Cheryle Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5892185817514746341/5892185817514746341_signature_31.png
| 2024-04-18 | 2024-04-18 09:53 | Chinese General Hospital | Blumentritt St. Sta Cruz Manila | NCR | NCR 1st - Manila | Bacteriology | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 27876 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 |  | 2024-04-15 | GXIV | 835604 | Unit for Preventive Maintenance. For module replacement of module B3. | Preventive Maintenance and Module Replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. Module B2 for reporting due to temperature error. |  | GX module B3 (711381) in replacement to defective B3 (685502). | DR#11397 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892148410227578911/5892148410227578911_signature_28.png | Maria Odessa Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/5892148410227578911/5892148410227578911_signature_31.png
| 2024-04-18 | 2024-04-18 08:44 | Hi Precision Diagnostic Center Inc. | N.S. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27875 | CONSORTIUM | jose@hi-precision.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 30 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6771 | 2024-04-15 | GXIV | 835602 | Unit for Preventive Maintenance. For module replacement of modules A3 and A4. | Preventive Maintenance and Module Replacement. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX modules A3 (693796) and A4 (619553) in replacement to defective A3 (904172) and A4 (210028982). | DR#11400 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5892108274619726311/5892108274619726311_signature_28.png | Joselito T. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5892108274619726311/5892108274619726311_signature_31.png
| 2024-04-17 | 2024-03-20 09:59 | Philippine Red Cross | Port Area, Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27874 | COMMERCIAL | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 03J030 | Unit for Preventive Maintenance | Preventive Maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891531083614542804/5891531083614542804_signature_28.png | Lalaine Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5891531083614542804/5891531083614542804_signature_31.png
| 2024-04-17 | 2024-03-20 10:02 | Philippine Red Cross | Port Area, Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27873 | COMMERCIAL | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for Preventive Maintenance. | Preventive Maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891530963613774555/5891530963613774555_signature_28.png | Lalaine ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5891530963613774555/5891530963613774555_signature_31.png
| 2024-04-17 | 2024-04-16 11:00 | Morong Rural Health Unit | Morong | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27872 | DOH | morongrhu@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 548 |  | GX-IV-D-6C | 821213 | For maintenance | Performed maintenence | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891510753614386134/5891510753614386134_signature_28.png | Necitas v. Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5891510753614386134/5891510753614386134_signature_31.png
| 2024-04-17 | 2024-04-17 16:08 | Medical Center Manila | 850 United Nations Avenue, Brgy. 674 Zone 73 Dist. V Paco Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Rhea Anne Cañelas |  | rheaannecanelas65@gmail.com |  |  | 27871 | CONSORTIUM | microbiologylab.manilamed@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1066 | 2024-04-15 | GXIV-4D-6C | 835540 | Module A1 for Module Replacement | Performed Preventive Maintenance Replaced Module A1 Generated Report | Replaced Module A1 789353->893812 | Unit is operational. All modules are working. | Work with Gian Paolo Magnaye | GeneXpert Module 6C | DR 11396 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891509510989687762/5891509510989687762_signature_28.png | Keziah Mae Barit | https://www.jotform.com/uploads/macarewecare/222927978372067/5891509510989687762/5891509510989687762_signature_31.png
| 2024-04-17 | 2024-04-17 15:11 | Tomas Claudio Health Center | Camus St. Corner E. Quirino Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27870 | DOH | arleen_p_alconera@yahoo.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 613 |  | GX-IV 4D | 838492 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891482653711804979/5891482653711804979_signature_28.png | PATRICIA KAYE C. ESTRADA | https://www.jotform.com/uploads/macarewecare/222927978372067/5891482653711804979/5891482653711804979_signature_31.png
| 2024-04-17 | 2024-04-17 15:14 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Animal Development Laboratory | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27869 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 30 |  |  | Stemi DV4 | 2004000887 | For check-up | None | Checked functionality Checked lenses | Unit is operational but mold contamination has affected the clarity very much | Due to mold contamination on the lenses, unit can be cleaned and requested for preventive maintenance but it is recommended for replacement of new unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891474733619438780/5891474733619438780_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891474733619438780/5891474733619438780_signature_31.png
| 2024-04-17 | 2024-04-17 15:10 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Fish Biology Laboratory | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27868 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | Stemi DV4 | 3919017542 | For check-up | None | Checked functionality Checked lenses | Unit is operational but lenses are contaminated with mold | Due to mold contamination on the lenses, unit can be cleaned and requested for preventive maintenance but it is recommended for replacement of new unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891472653617459541/5891472653617459541_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891472653617459541/5891472653617459541_signature_31.png
| 2024-04-17 | 2024-04-17 15:07 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 260 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27867 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 29 |  |  | KF2 ICS | 466208 | For check-up | None | Checked functionality Checked lenses | Unit is operational | Recommended for preventive maintenance. 40x objective for cleaning |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891469813611783365/5891469813611783365_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891469813611783365/5891469813611783365_signature_31.png
| 2024-04-17 | 2024-04-17 15:05 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 260 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27866 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 29 |  |  | KF2 ICS | 466143 | For check-up | None | Checked functionality Checked lenses | Unit is operational | Recommended for preventive maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891468613617243056/5891468613617243056_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891468613617243056/5891468613617243056_signature_31.png
| 2024-04-17 | 2024-04-17 15:03 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 260 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27865 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 29 |  |  | KF2 ICS | 466145 | For check-up | None | Checked functionality Checked lenses | Unit is operational | Recommended for preventive maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891467143611684510/5891467143611684510_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891467143611684510/5891467143611684510_signature_31.png
| 2024-04-17 | 2024-04-17 15:00 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27864 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | Primo Star | 3116008048 | For check-up | None | Checked functionality Checked lenses | Unit is operational Coarse adjust knob has resistance when turning Fine adjust knob looses focus | Recommended for preventive maintenance and repair of knobs |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891465823616342712/5891465823616342712_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891465823616342712/5891465823616342712_signature_31.png
| 2024-04-17 | 2024-04-17 14:57 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27863 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | Primo Star | 3116007573 | For check-up | None | Checked functionality Checked lenses | Unit is operational Coarse adjust knob has resistance when turning Fine adjust knob looses focus Specimen clip left and right adjustment knob is loose 40x objective for cleaning | Recommended for preventive maintenance and repair of knobs |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891464213611225079/5891464213611225079_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891464213611225079/5891464213611225079_signature_31.png
| 2024-04-17 | 2024-04-17 14:53 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27862 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | Primo Star | 3116007590 | For check-up | None | Checked functionality Checked lenses | Unit is operational Coarse adjust knob has resistance when turning Fine adjust knob looses focus | Recommended for preventive maintenance and repair of knobs |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891462393615159543/5891462393615159543_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891462393615159543/5891462393615159543_signature_31.png
| 2024-04-17 | 2024-04-17 10:03 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27861 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 ICS | 466209 | For check-up | None | Checked functionality Checked lenses | Unit does not power on No knob for light source adjustment | This product is no longer supported so repairing the unit is not possible. Recommended for replacement of new unit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891459853613172384/5891459853613172384_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891459853613172384/5891459853613172384_signature_31.png
| 2024-04-17 | 2024-04-17 14:29 | Solano Rural Health Unit | Roxas, Solano, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27860 | DOH | rhusolanolab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 560 | 2024-03-19 | Genexpert IV | 838201 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration (All Modules Passed) Performed Functionality Test | Unit is Operational |  | Xpertcheck kit | 11310 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891450823525550079/5891450823525550079_signature_28.png | Roderick Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5891450823525550079/5891450823525550079_signature_31.png
| 2024-04-17 | 2024-04-17 13:44 | Kauswagan Rural Health Unit -Tb Dots | Kauswagan Lanao del Norte | 10 | Lanao del Norte | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27859 | DOH | dextertinoy@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 326 |  | GX4 | 838402 | Module 3 Hardware Failure during processing error 2014 module 3 | Conduct Calibration in Module 3 Performed preventive maintenance on genexpert machine | Conduct Calibration in Module 3 Performed preventive maintenance on genexpert machine | Module 3 failed during calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891421546958111544/5891421546958111544_signature_28.png | Dexter Tinoy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5891421546958111544/5891421546958111544_signature_31.png
| 2024-04-17 | 2024-04-17 13:12 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27858 | DOH | suzzettekho34@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX-4 | 838554 | Module B4 for Observation (Existing Error 1002) Temperature Error | None | Preventive Maintenance | 3 modules working, 1 module disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891406639952801096/5891406639952801096_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/5891406639952801096/5891406639952801096_signature_31.png
| 2024-04-17 | 2024-04-17 13:15 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27857 | DOH | marteetingzon20@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1220 | 2024-04-17 | GXIV | 838421 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891404299228821907/5891404299228821907_signature_28.png | Marter Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5891404299228821907/5891404299228821907_signature_31.png
| 2024-04-17 | 2024-04-17 12:48 | Nueva Vizcaya Provincial Hospital | Almaguer, Bambang, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27856 | DOH | dumlaojoy4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1316 | 2024-04-17 | Genexpert IV | 838272 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration (B3 = Passed) (B4 = Passed) | Unit is Operational | 2 modules passed | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891391282122911245/5891391282122911245_signature_28.png | Eloisa De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5891391282122911245/5891391282122911245_signature_31.png