| 2024-04-17 | 2024-04-17 12:09 | West Visayas State University Medical Center | Jaro, Iloilo City | 6 | Iloilo | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27855 | MACARE-HBDC | mc-molecular@wvsu.edu.ph | 4 | ASSAY: Xpert HPV, COUNT: 7 |  | GX4 | 110021716 | None | Conduct Preventive Maintenance and save files for IQR, System Log, Assays Stats and Archive Test for 1 patient. | Transfer of their newly acquired GeneXpert Machine 10 color to Molecular Laboratory and guide the new user for their username and password. | All modules are in good condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891365793745980885/5891365793745980885_signature_28.png | Karla Danielle Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/5891365793745980885/5891365793745980885_signature_31.png
| 2024-04-17 | 2024-04-17 10:01 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27854 | MACARE-HBDC | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 ICS | 466149 | For check-up | None | Checked functionality Checked lenses | Unit is operational Coarse and fine adjust knobs have resistance when turning | This product is no longer supported so repairing the knobs is not possible. Recommended for replacement of new unit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891285873611606175/5891285873611606175_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891285873611606175/5891285873611606175_signature_31.png
| 2024-04-17 | 2024-04-17 09:52 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27853 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 KTR | 454545 | For check-up | None | Checked functionality Checked lenses | Unit is operational Coarse and fine adjust knobs have resistance when turning | This product is no longer supported so repairing the knobs is not possible. Recommended for replacement of new unit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891280603616334438/5891280603616334438_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891280603616334438/5891280603616334438_signature_31.png
| 2024-04-17 | 2024-04-17 09:48 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27852 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 ICS | 466154 | For check-up | None | Checked functionality Checked lenses | Unit is operational One eyepiece has a crack | Recommended for preventive maintenance. Eyepiece may not be available for this model |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891278393614584208/5891278393614584208_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891278393614584208/5891278393614584208_signature_31.png
| 2024-04-17 | 2024-04-17 09:45 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27851 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 ICS | 466150 | For check-up | None | Checked functionality Checked lenses | Unit is operational No specimen clip on the stage Light source adjust knob is loose | Recommended for preventive maintenance. Loose knob may not be repaired since it is integrated into the power supply of the microscope which is sealed. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891276973616176337/5891276973616176337_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891276973616176337/5891276973616176337_signature_31.png
| 2024-04-17 | 2024-04-17 09:42 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27850 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 ICS | 466146 | For check-up | None | Checked functionality Checked lenses | Unit is operational Light source adjustment knob is loose | Recommended for preventive maintenance. Loose knob may not be repaired since it is integrated into the power supply of the microscope which is sealed. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891275413618773258/5891275413618773258_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891275413618773258/5891275413618773258_signature_31.png
| 2024-04-17 | 2024-04-17 09:33 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27849 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 ICS | 466142 | For check-up | None | Checked functionality Checked lenses | Unit is functional Lenses are contaminated with mold 5x Objective needs cleaning | Due to mold contamination on the lenses, unit can be cleaned and requested for preventive maintenance but it is recommended for replacement of new unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891273293618323387/5891273293618323387_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891273293618323387/5891273293618323387_signature_31.png
| 2024-04-17 | 2024-04-17 09:25 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27848 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 ICS | 466144 | For check-up | None | Checked functionality Checked lenses | Unit does not power on Lenses are contaminated with mold Coarse and fine adjust knobs have resistance when turning | This product is no longer supported so repairing the knobs is not possible. Recommended for replacement of new unit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891267463612941156/5891267463612941156_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891267463612941156/5891267463612941156_signature_31.png
| 2024-04-17 | 2024-04-17 09:16 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27847 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 ICS | 466206 | For check-up | None | Checked functionality Checked lenses | Working | Recommended for preventive maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891263333613000653/5891263333613000653_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891263333613000653/5891263333613000653_signature_31.png
| 2024-04-17 | 2024-04-17 08:54 | University of the Philippines Diliman - Institute of Biology | UP Institute of Biology Building, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27846 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF2 ICS | 466207 | For check-up | None | Checked functionality Checked lenses | Unit does not power on Lenses are contaminated with mold | Due to mold contamination on the lenses, unit can be cleaned and requested for preventive maintenance but it is recommended for replacement of new unit. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5891251373611321322/5891251373611321322_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5891251373611321322/5891251373611321322_signature_31.png
| 2024-04-16 | 2024-04-16 11:53 | San Dionisio Health Center | A. Bonifacio St., Parañaque | NCR | NCR 4th - Parañaque | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27844 | DOH | editadplaton@yahoo.com | 27 | ASSAY: Xpert HPV, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1944 |  | GXIV-D | 831898 | System has No Power | Troubleshooting | Checked UPS Voltage Output and Fuse Cleaned Cartridge Bay and Plunger Rod Checked Filter if Clogged Performed Self-test Generate Reports | UPS is Defective. Not Charging and No Backup Function | Modules 2 and 4 for Reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5890834568319372938/5890834568319372938_signature_28.png | Edita D.Platojn | https://www.jotform.com/uploads/macarewecare/222927978372067/5890834568319372938/5890834568319372938_signature_31.png
| 2024-04-16 | 2024-04-12 09:00 | De La Salle Health Sciences Institute | Congressional Rd., Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27845 | CONSORTIUM | mgrioflorido@my.dlshsi.edu.ph | 20 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 68 | 2024-04-08 | GX16-8-D-10C | 110012951 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod and Icore Optics Checked Filters if Clogged Performed Self-test Performed XpertCheck Generate Reports | All Modules are Operational | All Modules Passed | XpertCheck Kits | DR 11262 & 11375 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5890834578311834916/5890834578311834916_signature_28.png | Ma. Cristina Rioflorido | https://www.jotform.com/uploads/macarewecare/222927978372067/5890834578311834916/5890834578311834916_signature_31.png
| 2024-04-16 | 2024-04-16 13:06 | Bernadette A. Bautista Medical Clinic | Holy Spirit, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27843 | CONSORTIUM | richeclevia@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 31 | 2024-04-15 | GX IV 6C-D | 835603 | Module b3 for replacement | Performed PM Replaced Module b3 Genereated report | Replaced Module B3 792309 -> 680909 | Unit is operational All modules working | work with Maam Rhea | GX module 6C | DR 11395 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5890551181016149058/5890551181016149058_signature_28.png | richard eclevia | https://www.jotform.com/uploads/macarewecare/222927978372067/5890551181016149058/5890551181016149058_signature_31.png
| 2024-04-16 | 2024-04-16 12:57 | Panabo City Health Office | Panabo Poblacion | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | nadineeust67@gmail.com |  | 27842 | DOH | chopanabolaboratory@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2424 |  | GX-IV | 838295 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5890550554525233082/5890550554525233082_signature_28.png | Louella Jane G. Lopez, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5890550554525233082/5890550554525233082_signature_31.png
| 2024-04-16 | 2024-04-04 10:14 | Cebu City Health Department | Cebu City | 7 | Cebu | Laboratory | YES | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 27841 | COMMERCIAL | iamanne1111@gmail.com | 23 | ASSAY: Xpert HCV VL, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 638 | 2024-04-04 | GX4 | 829419 | Module A1, A2 Error Code 2120,2097,2024 | Perform Maintenance. | Perform Preventive Maintenance for all modules. | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5890434805247552882/5890434805247552882_signature_28.png | Eleanor Chavez | https://www.jotform.com/uploads/macarewecare/222927978372067/5890434805247552882/5890434805247552882_signature_31.png
| 2024-04-15 | 2024-04-11 14:34 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 27840 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2999 ASSAY: Xpert SARS-CoV-2, COUNT: 4284 ASSAY: Others, COUNT: 10 | 2024-03-25 | GX16 | 11014171 | Due for xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | C3,B3,D1 newly replaced module |  | Xpertcheck kits | 11331 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5889707453614855825/5889707453614855825_signature_28.png | Jessica Faye Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5889707453614855825/5889707453614855825_signature_31.png
| 2024-04-15 | 2024-04-05 14:55 | Philippine Red Cross QC | Malakas st , QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27839 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 28 |  | 2024-04-04 | MCS + | 02A227 | Encounter self test error non valital during self test. For CPU board replacement | Cleaned unit Installed and replaced CPU board Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Repair amd calibration | Operational | Pull out defective CPU board | CPU Board | 11364 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5889706983618341483/5889706983618341483_signature_28.png | Lambert John Cordero | https://www.jotform.com/uploads/macarewecare/222927978372067/5889706983618341483/5889706983618341483_signature_31.png
| 2024-04-15 | 2024-04-15 11:44 | Tarlac City Health Center 2 | Matatalaib Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27838 | DOH | clarizannebeltran1996@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 383 |  | GXIV-D | 831917 | Due for maintenance | Preventive maintenance | Performed preventive maintenance Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5889692822877625649/5889692822877625649_signature_28.png | Clariz Anne M. Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5889692822877625649/5889692822877625649_signature_31.png
| 2024-04-15 | 2024-04-15 11:41 | Silliman University Medical Center | Dumaguete City, Philippines | 7 | Negros Oriental | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@protonmail.com |  |  | 27837 | COMMERCIAL | sumcmoleclab@gmail.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 6 ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert C. difficile BT, COUNT: 16 ASSAY: Xpert CT/NG, COUNT: 17 ASSAY: Xpert HBV VL, COUNT: 130 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert HPV, COUNT: 66 ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 ASSAY: Xpert SARS-CoV-2, COUNT: 134 ASSAY: Xpert HCV VL, COUNT: 23 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 10 |  | GX-4 R2 | 821471 | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5889628845966720314/5889628845966720314_signature_28.png | SUMCFI Molecular Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/5889628845966720314/5889628845966720314_signature_31.png
| 2024-04-12 | 2024-04-12 11:30 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27836 | DOH | allan.vinluan2013@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1047 | 2024-04-04 | GXIV-4-D-6C | 838472 | For Calibration and Preventive Maintenance | #NAME? | Performed Calibration and Preventive Maintenance | Unit is operational | Did not run XpertCheck on newly replaced module B2. Module B1 failed XpertCheck, for reporting. Modules B3 and B4 passed XpertCheck | XpertCheck Kits | DR 11362 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5887167547396115297/5887167547396115297_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5887167547396115297/5887167547396115297_signature_31.png
| 2024-04-12 | 2024-04-12 10:17 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27835 | DOH | jeffreyvrumbaoa@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1120 |  | GXIV-4-D-6C | 838473 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5887166287392552756/5887166287392552756_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/5887166287392552756/5887166287392552756_signature_31.png
| 2024-04-12 | 2024-04-12 14:48 | International Organization for Migration | Trafalgar Plaza | NCR | NCR 4th - Makati | Laboratory | YES | CALIBRATION MAINTENANCE | CARL ZEISS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | jsabalboro@iom.int |  | 27834 | IOM | rplan@iom.int | 22 |  |  | Primostar iLED | 3134032159 | For calibration | Performed maintenance/calibration | Performed maintenance/calibration Cleaned microscope body Cleaned objectives Cleaned eyepiece Cleaned microscope stage Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5887143013937286294/5887143013937286294_signature_28.png | Rosamer Anne S Plan | https://www.jotform.com/uploads/macarewecare/222927978372067/5887143013937286294/5887143013937286294_signature_31.png
| 2024-04-12 | 2024-04-12 14:53 | International Organization for Migration | Trafalgar Plaza | NCR | NCR 4th - Makati | Laboratory | YES | CALIBRATION MAINTENANCE | CARL ZEISS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | jsabalboro@iom.int |  | 27833 | IOM | rplan@iom.int | 28 |  |  | Primostar iLED | 3134004145 | For maintenance/calibration | Performed maintenance/calibration | Performed maintenance/calibration Cleaned microscope body Cleaned objectives Cleaned eyepiece Cleaned microscope stage Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5887142383933465989/5887142383933465989_signature_28.png | Rosamet Anne S. Plan | https://www.jotform.com/uploads/macarewecare/222927978372067/5887142383933465989/5887142383933465989_signature_31.png
| 2024-04-12 | 2024-04-12 12:26 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27832 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2039 |  | GXIV-4D | 838561 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5887053845857553957/5887053845857553957_signature_28.png | Hannah Jane D. Resuento | https://www.jotform.com/uploads/macarewecare/222927978372067/5887053845857553957/5887053845857553957_signature_31.png
| 2024-04-12 | 2024-04-12 10:27 | Calaanan Health Center | N. Nadurata St. Caloocan | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27831 | DOH | rosaleslorina0917@gmail.com | 25 |  |  | GXIV | 831892 | Unit for Preventive Maintenance. | Preventive Maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5886987683441827476/5886987683441827476_signature_28.png | Lorina Rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/5886987683441827476/5886987683441827476_signature_31.png
| 2024-04-12 | 2024-04-12 10:14 | St. Luke's Medical Center - Global City | Rizal Drive. | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 27830 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 174 ASSAY: Xpert CT/NG, COUNT: 1098 ASSAY: Xpert HBV VL, COUNT: 72 ASSAY: Xpert HPV, COUNT: 979 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2534 ASSAY: Xpert HCV VL, COUNT: 32 ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 | 2024-04-08 | GX-IV-D-L | 808411 | Module B1 for replacement | Replaced module B1 | Performed maintenance Replaced module B1 Generated reports | Unit is operational. Module B4/688662 not operational, reported already. |  | Old: B1/666350 New: B1/652045 | DR# 11373 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5886983689414259941/5886983689414259941_signature_28.png | Vincent Raquiza Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/5886983689414259941/5886983689414259941_signature_31.png
| 2024-04-11 | 2024-04-11 21:52 | Cebu Velez General Hospital | F. Ramos St. Ranudo St. Cebu City | 7 | Cebu | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | microbiovelez@gmail.com |  | 27829 | WHO | johmicroslee@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 56 |  | GX4 | 110017599 | Due for Calibration and preventive maintenance. | Performed maintenance and calibration. | Performed maintenance and calibration | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5886530891119844546/5886530891119844546_signature_28.png | John Lee, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5886530891119844546/5886530891119844546_signature_31.png
| 2024-04-11 | 2024-04-11 13:48 | The Medical City | Ortigas Avenue | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27828 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 20 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Run simulation Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5886247240992089596/5886247240992089596_signature_28.png | Gracia Angara | https://www.jotform.com/uploads/macarewecare/222927978372067/5886247240992089596/5886247240992089596_signature_31.png
| 2024-04-11 | 2024-04-11 11:53 | The Medical City | Ortigas Avenue | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27827 | COMMERCIAL | pedelacruz@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5886179360992099625/5886179360992099625_signature_28.png | Perlita De La Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5886179360992099625/5886179360992099625_signature_31.png
| 2024-04-11 | 2024-04-11 11:07 | Philippine Red Cross QC | Kalayaan qve QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | INSTALLATION ON-SITE ASSISTANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27826 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 28 |  |  | MCS + | 02A227 | For transfer new location @ kalayaan ave qc | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Transfer and re installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5886144495112765874/5886144495112765874_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/5886144495112765874/5886144495112765874_signature_31.png
| 2024-04-09 | 2024-04-04 10:01 | Polomolok Rural Health Unit | Poblacion, Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27825 | DOH | jmmirasolrmt@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1850 |  | Gx-IV | 838265 | N/A | Genexpert Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5884262394313677532/5884262394313677532_signature_28.png | June Michele Mirasol | https://www.jotform.com/uploads/macarewecare/222927978372067/5884262394313677532/5884262394313677532_signature_31.png
| 2024-04-08 | 2024-04-08 11:29 | Poro Rural Health Unit | Poro, Camotes Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27824 | DOH | avilacharmaine6@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 584 |  | GX4 | 838279 | Stucked module | Pulled out module and bring office for repair and replacement of module door in Module 1 | Pulled out module and bring office for repair and replacement of module door in Module 1 | Module 1 need to be repair and pulled stucked cartridge |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883562641111224438/5883562641111224438_signature_28.png | Charmaigne Avila, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5883562641111224438/5883562641111224438_signature_31.png
| 2024-04-08 | 2024-04-02 10:58 | Asian Hospital and Medical Center | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | agimang@asianhospital.com |  | 27823 | MACARE-HBDC | agimang@asianhospital.com | 24 |  |  | MS+ | 10B129 | Unit has No Power | Troubleshooting | Checked Voltage Outlet Checked Power Cord Replaced Defective Fuse Performed Functionality Test | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883552133615859259/5883552133615859259_signature_28.png | Angel Danica Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5883552133615859259/5883552133615859259_signature_31.png
| 2024-04-08 | 2024-04-08 09:41 | Philippine Red Cross QC chapter | Malakas st., QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27822 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 27 |  | 2024-04-04 | MCS | 02A227 | Encounter self test error non valital during self test. Defective CPU board | Cleaned unit Replaced and installed CPU board Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Repair and calibration | Operational | Pull out defective CPU board | CPU Board | 11364 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883499063619659607/5883499063619659607_signature_28.png | Lambert John Cordero | https://www.jotform.com/uploads/macarewecare/222927978372067/5883499063619659607/5883499063619659607_signature_31.png
| 2024-04-08 | 2024-04-08 09:32 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 27821 | MACARE-HBDC | arellanojohneric09@gmail.com | 28 |  | 2024-04-04 | MCS + | 17H632SPG | CRC Bios failed defective CPU board | Cleaned unit Replaced and installed CPU board Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Repair and calibration | Operational | Pull out defective parts | CPU Board | 11365 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883493963618833323/5883493963618833323_signature_28.png | John Eric Arellano | https://www.jotform.com/uploads/macarewecare/222927978372067/5883493963618833323/5883493963618833323_signature_31.png
| 2024-04-08 | 2024-04-04 14:13 | Marikina City Health Office | Shoe Avenue. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27820 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 136 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883408716715307595/5883408716715307595_signature_28.png | Francesca Herrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5883408716715307595/5883408716715307595_signature_31.png
| 2024-04-08 | 2024-04-05 10:13 | St. Luke's Medical Center- Global City | Rizal Dr. | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 27819 | COMMERCIAL | ci.bgc@stlukes.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 173 ASSAY: Xpert CT/NG, COUNT: 1109 ASSAY: Xpert HPV, COUNT: 966 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2511 ASSAY: Xpert HCV VL, COUNT: 35 ASSAY: Xpert HIV-1 Viral Load, COUNT: 32 |  | GX-IV-L-6C | 808411 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Modules B1 and B4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883407656716118800/5883407656716118800_signature_28.png | NORSHAINA UNGAD | https://www.jotform.com/uploads/macarewecare/222927978372067/5883407656716118800/5883407656716118800_signature_31.png
| 2024-04-08 | 2024-04-05 13:52 | Cardinal Santos Medical Center | Wilson St. | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27818 | COMMERCIAL | eileenjoie.eltanal@yahoo.com | 22 | ASSAY: Xpert HBV VL, COUNT: 120 ASSAY: Xpert HPV, COUNT: 25 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1068 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Modules B1 and B4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883404043478622672/5883404043478622672_signature_28.png | Eileen Eltanal | https://www.jotform.com/uploads/macarewecare/222927978372067/5883404043478622672/5883404043478622672_signature_31.png
| 2024-04-08 | 2024-04-05 10:20 | St. Luke's Medical Center- Global City | Rizal Dr. | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 27817 | COMMERCIAL | ci.bgc@stlukes.com.ph | 25 |  |  | GX-IV-D-6C | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Module C1 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883404003473919970/5883404003473919970_signature_28.png | NORSHAINA UNGAD | https://www.jotform.com/uploads/macarewecare/222927978372067/5883404003473919970/5883404003473919970_signature_31.png
| 2024-04-08 | 2024-04-04 13:51 | Marikina City Health Office | Shoe Avenue | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27816 | MACARE-HBDC | estyaragon2019@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 653 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953471059736/5883402953471059736_signature_28.png | Francesca Herrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953471059736/5883402953471059736_signature_31.png
| 2024-04-08 | 2024-04-04 11:15 | Metro Antipolo Hospital and Medical Center | Marocs Highway, Antipolo | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27815 | COMMERCIAL | mahmci_labcmt@yahoo.com | 23 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 204 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GX-IV-D-6C | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953471312614/5883402953471312614_signature_28.png | Graziela Marie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953471312614/5883402953471312614_signature_31.png
| 2024-04-08 | 2024-04-05 10:13 | St. Luke's Medical Center- Global City | Rizal Dr. | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 27813 | COMMERCIAL | ci.bgc@stlukes.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 173 ASSAY: Xpert CT/NG, COUNT: 1109 ASSAY: Xpert HPV, COUNT: 966 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2511 ASSAY: Xpert HCV VL, COUNT: 35 ASSAY: Xpert HIV-1 Viral Load, COUNT: 32 |  | GX-IV-L-6C | 808411 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Modules B1 and B4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953474615896/5883402953474615896_signature_28.png | NORSHAINA UNGAD | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953474615896/5883402953474615896_signature_31.png
| 2024-04-08 | 2024-04-04 13:48 | Marikina City Health Office | Shoe Ave. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27814 | DOH | estyaragon2019@gmail.co | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1526 ASSAY: Xpert HIV-1 Viral Load, COUNT: 308 |  | GX-IV-D-6C | 831588 | For maintenance | Performed maintenence | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Module B2/630566 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953475464507/5883402953475464507_signature_28.png | Francesca Herrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953475464507/5883402953475464507_signature_31.png
| 2024-04-08 | 2024-04-04 14:13 | Marikina City Health Office | Shoe Avenue. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27812 | MACARE-HBDC | estyaragon2019@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 136 |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953476635538/5883402953476635538_signature_28.png | Francesca Herrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5883402953476635538/5883402953476635538_signature_31.png
| 2024-04-05 | 2024-04-05 15:20 | Poro Rural Health Unit | Poro, Camotes Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27811 | DOH | avilacharmaine6@gmail.com | 26 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 587 |  | GX4 | 838279 | Stucked up module in Module 1 SN # 807772 Conduct calibration in genexpert machine | Pulled out stucked up cartridge in Module 1. Conduct calibration in genexpert machine | Pulled out stucked up cartridge in Module 1. Conduct calibration in genexpert machine | All units are operational . Module 3 are failed during calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5881112651874599637/5881112651874599637_signature_28.png | Charmaigne Avila, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5881112651874599637/5881112651874599637_signature_31.png
| 2024-04-05 | 2024-04-05 13:24 | CEBU CITY HEALTH OFFICE | CEBU CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27810 | COMMERCIAL | jonnifercodoyceniza@gmail.com | 20 | ASSAY: Xpert HCV VL, COUNT: 73 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1839 |  | GX-IV | 829419 | #NAME? | #NAME? | #NAME? | EQUIPMENT OPERATIONAL, MODULE A3 AND A4 FOR SERVICE |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5881041111112428068/5881041111112428068_signature_28.png | Jonnifer Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5881041111112428068/5881041111112428068_signature_31.png
| 2024-04-05 | 2024-04-05 12:50 | Mindanao central sanitarium | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 27809 | DOH | ybioamie@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2581 | 2024-04-05 | GX 4 | 821214 | None | All modules undergone Xpertcheck calibration | Preventive Maintenence and Xpertcheck calibration | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5881039898221813245/5881039898221813245_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5881039898221813245/5881039898221813245_signature_31.png
| 2024-04-05 | 2024-04-05 09:24 | Samal Super Health Center | Tabing Ilog Samal Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27808 | DOH | vickytholits2002@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 529 |  | GXIV-D | 831729 | Machine relocation | Preventive maintenance Relocation | Performed preventive maintenance Perforned machine disassembly Relocate machine to new facility (Samal RHU to Samal Super Health Center) Installation and setup of machine Gahered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880898431521281524/5880898431521281524_signature_28.png | Maria Victoria Banzuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5880898431521281524/5880898431521281524_signature_31.png
| 2024-04-05 | 2024-04-05 08:45 | Bantayan Rural Health Unit | Bantayan, Cebu | 7 | Cebu | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27807 | DOH | mcdelle916bausin@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1281 |  | GXIV-4D | 838509 | #NAME? | None | #NAME? | #NAME? | Used four XPC cartridges. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880882070443631315/5880882070443631315_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/5880882070443631315/5880882070443631315_signature_31.png
| 2024-04-04 | 2024-04-04 15:54 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27806 | COMMERCIAL | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert CT/NG, COUNT: 6 ASSAY: Others, COUNT: 14 ASSAY: Xpert HPV, COUNT: 42 ASSAY: Xpert MTB/RIF Ultra, COUNT: 189 ASSAY: Xpert HCV VL, COUNT: 13 |  | GX16-D (12) | 804476 | Due for maintenance 9 module disabled 3 module in operational | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | 3 modules operational 9 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880413520319601982/5880413520319601982_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5880413520319601982/5880413520319601982_signature_31.png
| 2024-04-04 | 2024-04-04 15:31 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molcular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27805 | MACARE-HBDC | molecularlabnkti@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 50 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Others, COUNT: 10 ASSAY: Xpert HBV VL, COUNT: 31 ASSAY: Xpert HPV, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 43 ASSAY: Xpert HCV VL, COUNT: 13 |  | GX4 | 838524 | Due for maintenance B1,B3,B4 reported already for replacement | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | B2 operational B1 B3 disable B4 not in use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880412680318617759/5880412680318617759_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5880412680318617759/5880412680318617759_signature_31.png
| 2024-04-04 | 2024-04-04 14:37 | National Kidney and Transplant Institute | East Avenue,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27804 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1337 ASSAY: Others, COUNT: 62 ASSAY: Xpert HBV VL, COUNT: 969 ASSAY: Xpert HPV, COUNT: 63 ASSAY: Xpert MTB/RIF, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2363 ASSAY: Xpert SARS-CoV-2, COUNT: 15391 ASSAY: Xpert HCV VL, COUNT: 110 |  | GX16 | 110005587 | Due for maintenance, A2,A4,C4 reported for replacement,D2 encounter temp error 4009,4010,4017, | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | 12 module operational 4 module disable due to temp error | D2 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880412000319738341/5880412000319738341_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5880412000319738341/5880412000319738341_signature_31.png
| 2024-04-04 | 2024-04-04 07:51 | Casimiro Ynares Memorial Hospital | Baeco st., burgos Rodriguez Rizal | 4A | Rizal | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27803 | DOH | caysmhlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 629 |  | GX4 | 831642 | Encounter continous temp error 1001 in module A1 Stuck cartridge A1 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Fixed stuck cartridge Test run unit | Troubleshooting | A1 reported already A2,A3,A4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880411350315851445/5880411350315851445_signature_28.png | Aileen Docto | https://www.jotform.com/uploads/macarewecare/222927978372067/5880411350315851445/5880411350315851445_signature_31.png
| 2024-04-04 | 2024-04-03 15:29 | Commonwealth Health Center | Katuparan st, Commonwealth QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27802 | DOH | commonwealthclinical@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 114 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2901 |  | GX4 | 831615 | As per end user they encounter 4004 valve motion error in module A2 A2 module twist cable | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Fixed twist cable Test run unit | Troubleshooting | After troubleshooting is performed A2 error 4004 still occur A1,A3,A4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880410940319359782/5880410940319359782_signature_28.png | Danica Laisa T. Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5880410940319359782/5880410940319359782_signature_31.png
| 2024-04-04 | 2024-04-04 14:21 | UP-NIH | Pedro Gil St. Ermita Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27801 | COMMERCIAL | nihcentrallab@gmail.com | 25 |  |  | GX16-10C | 110014186 | Unit for PlPreventive maintenance Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module B4 (210070221), service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880225257426595067/5880225257426595067_signature_28.png | Mary Grace Hernaez | https://www.jotform.com/uploads/macarewecare/222927978372067/5880225257426595067/5880225257426595067_signature_31.png
| 2024-04-04 | 2024-04-04 12:57 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27800 | COMMERCIAL | gramanmolecularlab@gmail.com | 22 | ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 592 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 15 | 2024-04-03 | GXIV | 110008530 | Module B4 (210036037) for replacement | Replaced module B4 | - Replaced module B4 by SN: 210089719 - Checked and cleaned cartridge bay, plunger rod, filter - Generated reports - Performed module self test | Unit is operational All 4 modules are in good condition |  | GeneXpert Module SN: 210036037 | 11357 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880192371022268790/5880192371022268790_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5880192371022268790/5880192371022268790_signature_31.png
| 2024-04-04 | 2024-04-04 12:15 | Allied Care Expert Medical Center | Tagbilaran, Bohol | 7 | Bohol | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27799 | CONSORTIUM | laboratory@acemcbohol.com | 27 |  |  | Gx4 | 840969 | Due date for xpertcheck calibration | Performing xpertcheck calibration Gathering genxpert data Clean fan filter | Performing xpertcheck calibration Gathering genxpert data Clean Fan filter | Systems operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880155201424829522/5880155201424829522_signature_28.png | Dianne Dawn Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5880155201424829522/5880155201424829522_signature_31.png
| 2024-04-04 | 2024-04-04 11:17 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | msff-tondo-screeningsup@paris.msf.org |  | 27798 | COMMERCIAL | msff-tondo-labtech@paris.msf.org | 25 |  | 2024-04-03 | GX16 | 110016343 | Unit for PM. For module replacement of modules A4, B2, B4, C2 and C3. | PM and module replacement | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX modules A4 (898763), B2 (210067405), B4 (210013968), C2 (210032686) and C3 (914567) in replacement to defective A4 (210076794), B2 (210083060), B4 (210083059), C2 (210081539) and C4 (210070572). | DR#11356 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880108227427376641/5880108227427376641_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/5880108227427376641/5880108227427376641_signature_31.png
| 2024-04-04 | 2024-04-04 10:43 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27797 | DOH | claritomirasol@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 | 2024-04-04 | GXIV | 838548 | n/a | #NAME? | - Performed preventive maintenance (Cleaned Cartridge bay, i-Core, plunger rod and fan filter) - Performed functionality test - Generate reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880085123226595084/5880085123226595084_signature_28.png | Mirasol Clarito | https://www.jotform.com/uploads/macarewecare/222927978372067/5880085123226595084/5880085123226595084_signature_31.png
| 2024-04-04 | 2024-04-03 13:46 | Sarangani Health Care Facility | Alabel Sarangani | 12 | Sarangani | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27796 | DOH | ronayadao@yahoo.com | 24.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 689 |  | Gx-IV | 838542 | Module 3 temperature error | Perform Xpertcheck at module 3 for troubleshooting Genexpert cleaning Reporting of errors/problem Data collection | Perform Genexpert machine Preventive maintenance Perform module Troubleshooting | 3 operational modules | Module 3 passed Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5880052085917450212/5880052085917450212_signature_28.png | Ronora Yadao | https://www.jotform.com/uploads/macarewecare/222927978372067/5880052085917450212/5880052085917450212_signature_31.png
| 2024-04-03 | 2024-03-21 13:19 | Sariaya Rural Health Unit | Sariaya | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27795 | DOH | marielinnetabril@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2098 |  | GX-IV-D-6C | 831890 | Stuck cartridge on module A1 | Removed stuck cartridge. | Manual removal of stuck cartridge. Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports. | Unit is operational. Module A4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5879189053618659145/5879189053618659145_signature_28.png | Marie Linnet Abril | https://www.jotform.com/uploads/macarewecare/222927978372067/5879189053618659145/5879189053618659145_signature_31.png
| 2024-04-03 | 2024-03-22 13:22 | Camp Capinpin Army Station Hospital | Camp Capinpin | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27794 | DOH | cathialpha@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 139 |  | GX-IV-D-6C | 831896 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5879188303616873493/5879188303616873493_signature_28.png | Anne Catherine I Alpha | https://www.jotform.com/uploads/macarewecare/222927978372067/5879188303616873493/5879188303616873493_signature_31.png
| 2024-04-03 | 2024-03-27 11:09 | Cardinal Santos Medical Center | Wilson St. Greenhills | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27793 | COMMERCIAL | colomakath14@gmail.com | 20 |  |  | MCS+ | 08M461 | Red valve not correctly loaded. SPM error. | Checked valves and SPM for leak. | Simulation Run Performed maintenance Checked valve Checked rotor pumps Checked SPM Verified calibration parameters Functionality test | Unit is operational. No leak found. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5879187943613309649/5879187943613309649_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/5879187943613309649/5879187943613309649_signature_31.png
| 2024-04-03 | 2024-04-02 11:32 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27792 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 24 | ASSAY: Xpert HPV, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 352 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GX-XVI-D-6C | 838627 | Modules A3, A4, B1, B2, C3, C4 and D3 for checking. | Performed maintenance. | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | Unit is operational. Modules A1, A2, B3, B4, C2, D2 and D4 not operational. Reported already. Modules A3, A4, B1, B2, C3, C4 and D3 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5879139173618038557/5879139173618038557_signature_28.png | Justine lorenz Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5879139173618038557/5879139173618038557_signature_31.png
| 2024-04-03 | 2024-04-02 13:08 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27791 | COMMERCIAL | rmcbloodbank@gmail.com | 24 |  |  | MCS+ | 15H045G | For maintenance | Calibrated the unit | Performed maintenance Run simulation Checked rotor pumps Verified calibration parameters Calibrated the unit | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5879139173611019103/5879139173611019103_signature_28.png | Ferdinand B. Barroga | https://www.jotform.com/uploads/macarewecare/222927978372067/5879139173611019103/5879139173611019103_signature_31.png
| 2024-04-02 | 2024-04-02 14:28 | Bicol Medical Center | Naga | 5 | Camarines Sur | Blood Bank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27790 | MACARE-HBDC | jericomd729@gmail.com | 25 |  |  | MCS+ | 15H031G | Kit breakage inside Centrifuge | Removed chipped particles of kit inside centrifuge Checked parameters Checked power cord Checked fuse | Performed Troubleshooting | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5878488284916011359/5878488284916011359_signature_28.png | Jerico R. Cabaña | https://www.jotform.com/uploads/macarewecare/222927978372067/5878488284916011359/5878488284916011359_signature_31.png
| 2024-04-02 | 2024-04-02 14:20 | General Tinio Rural Health Unit | Poblacion, General Tinio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27789 | DOH | jadelyn_pajarillaga@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 863 | 2024-04-02 | Genexpert IV | 831695 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5878482649218392682/5878482649218392682_signature_28.png | Harris F. Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/5878482649218392682/5878482649218392682_signature_31.png
| 2024-04-02 | 2024-04-02 13:15 | Cabadbaran City Health Office | S. Curato, Brgy. 5 Cabadbaran City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27788 | DOH | samstercem@gmail.com | 27 |  |  | GX-4 | 838550 | None | Reported to Cepheid | XpertCheck Calibration of all modules | All modules are passed in XPC WORKING AND FUNCTIONAL | None |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5878444426978749263/5878444426978749263_signature_28.png | Sammy Cempron | https://www.jotform.com/uploads/macarewecare/222927978372067/5878444426978749263/5878444426978749263_signature_31.png
| 2024-04-02 | 2024-04-02 09:49 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | MOLECULAR LAB | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27787 | COMMERCIAL | cduhcovidpoct@gmail.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 65 ASSAY: Xpert CT/NG, COUNT: 139 ASSAY: Xpert HIV-1 Viral Load, COUNT: 144 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1085 ASSAY: Xpert HCV VL, COUNT: 114 |  | GX-IV | 110008524 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5878329071113515317/5878329071113515317_signature_28.png | Jennifer Marie Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5878329071113515317/5878329071113515317_signature_31.png
| 2024-04-02 | 2024-04-02 09:26 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | REPAIR | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27786 | COMMERCIAL | cduhlabresults2@yahoo.com | 23 |  |  | MCS+ | 17H639SPG | UPS DOES NOT HOLD A CHARGE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5878308294466347809/5878308294466347809_signature_28.png | CHRISTINE JADE T. LAGAT | https://www.jotform.com/uploads/macarewecare/222927978372067/5878308294466347809/5878308294466347809_signature_31.png
| 2024-04-01 | 2023-08-18 09:16 | Itogon Rural Health Unit | Itogon, Benguet | CAR | Benguet | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27785 | DOH | ggdepayso@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 45 ASSAY: Xpert MTB/RIF Ultra, COUNT: 321 | 2023-06-13 | GX4 | 838282 | Due for Calibration and Maintenance | #NAME? | Performed XpertCheck and Maintenance | Unit is operational | Module B3 failed, for reporting | XpertCheck Kits | DR 5336 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5877803805918073705/5877803805918073705_signature_28.png | Gloria Depayso | https://www.jotform.com/uploads/macarewecare/222927978372067/5877803805918073705/5877803805918073705_signature_31.png
| 2024-04-01 | 2024-03-27 09:34 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27782 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 350 ASSAY: Xpert SARS-CoV-2, COUNT: 3214 |  | GX16-D-10C | 110013856 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, Icore Optics Checked Filter if Clogged Performed Self-test Performed XpertCheck Archived Test Results Generate Reports | All Modules are Operational | 14 Passed, 2 Newly Replaced |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5877629543611641331/5877629543611641331_signature_28.png | Ervin Emmett P Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/5877629543611641331/5877629543611641331_signature_31.png
| 2024-04-01 | 2024-03-27 09:18 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27781 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 517 ASSAY: Xpert SARS-CoV-2, COUNT: 243 |  | GXIV-D-10C | 110014099 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, Icore Optics Checked Filter if Clogged Performed Self-test Performed XpertCheck Archived Test Results Generate Reports | All Modules are Operational | All Passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5877629543615851944/5877629543615851944_signature_28.png | Ervin Emmett P Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/5877629543615851944/5877629543615851944_signature_31.png
| 2024-04-01 | 2024-03-26 14:23 | Research Institute for Tropical Medicine | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Training Center | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27784 | MACARE-HBDC | nathanpabilando@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1061 | 2024-02-14 | GXIV-D-10C | 110017870 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod and Icore Optics Checked Filter if Clogged Performed Self-test Performed XpertCheck Generate Reports | All Modules are Operational | All Passed | XpertCheck Kit | DR 11219 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5877629553611071602/5877629553611071602_signature_28.png | Nathaniel Pabilando | https://www.jotform.com/uploads/macarewecare/222927978372067/5877629553611071602/5877629553611071602_signature_31.png
| 2024-04-01 | 2024-03-26 14:37 | Research Institute for Tropical Medicine | Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Training Center | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27783 | MACARE-HBDC | nathanpabilando@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 965 | 2024-02-14 | GXIV-D-10C | 110017869 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod and Icore Optics Checked Filter if Clogged Performed Self-test Performed XpertCheck Generate Reports | All Modules are Operational | All Passed | XpertCheck Kit | DR 11220 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5877629553618563123/5877629553618563123_signature_28.png | Nathaniel Pabilando | https://www.jotform.com/uploads/macarewecare/222927978372067/5877629553618563123/5877629553618563123_signature_31.png
| 2024-04-01 | 2024-04-01 14:32 | Corazon Locsin Montelibano M.R.H | Lacson St., Bacolod City | 6 | Negros Occidental | Laboratory Dept.(Micro) | NO | CALIBRATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27780 | DOH | jennybellegallego@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1049 ASSAY: Xpert HIV-1 Viral Load, COUNT: 490 |  | GX4 | 821219 | N/A | All modules passed the calibration today. | Performed Xpert Check calibration for 4 Modules and conduct Preventive Maintenance. | All modules are in good working conditions as well as the CPU and Monitor. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5877628911518648742/5877628911518648742_signature_28.png | Jennybelle Gallego | https://www.jotform.com/uploads/macarewecare/222927978372067/5877628911518648742/5877628911518648742_signature_31.png
| 2024-04-01 | 2024-04-01 11:15 | Talisay District Hospital | San Isidro, Talisay City Cebu | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27779 | DOH | csmcmicro@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1246 ASSAY: Xpert HPV, COUNT: 49 |  | GX4 | 838559 | Due for Preventive Maintenance Error 2008 and 2036 | Performed Preventive Maintenance, Clean Up Icore and Plunger, clean filter Generate IQR, GXX, Sys log, and Back up database | Performed Preventive Maintenance, Clean Up Icore and Plunger, clean filter Generate IQR, GXX, Sys log, and Back up database | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5877509271114812328/5877509271114812328_signature_28.png | Joseph Lorenzo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5877509271114812328/5877509271114812328_signature_31.png
| 2024-04-01 | 2024-03-27 09:38 | Pozorrubio Rural Health Unit | Pozorrubio, Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27778 | DOH | rhupozorrubio@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 607 | 2024-02-05 | GXIV-4-D-6C | 838467 | #NAME? | #NAME? | #NAME? | Unit is operational | Module B1 Failed XpertCheck. Remaining Modules passed XpertCheck | XpertCheck Kits | DR 11189 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5877439955911425743/5877439955911425743_signature_28.png | Jinky Grace Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/5877439955911425743/5877439955911425743_signature_31.png
| 2024-03-28 | 2024-03-28 12:45 | Bicol Medical Center | Naga | 5 | Camarines Sur | Blood Bank | YES | TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27777 | MACARE-HBDC | collaoleo17@gmail.com | 25 |  |  | MCS+ | 15H031G | MCS+no power Possible blown/busted fuse | Replaced 2 Fuse | Checked UPS Power cable Checked Fuse | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5874107039398477574/5874107039398477574_signature_28.png | Leo C. Collao, RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5874107039398477574/5874107039398477574_signature_31.png
| 2024-03-27 | 2024-03-27 11:09 | Cardinal Santos Medical Center | Wilson St. Greenhills | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27775 | COMMERCIAL | colomakath14@gmail.com | 20 |  |  | MCS+ | 08M461 | Red valve not correctly loaded. SPM error. | Checked valves and SPM for leak. | Simulation Run Performed maintenance Checked valve Checked rotor pumps Checked SPM Verified calibration parameters Functionality test | Unit is operational. No leak found. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306923612516906/5873306923612516906_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306923612516906/5873306923612516906_signature_31.png
| 2024-03-27 | 2024-03-27 11:09 | Cardinal Santos Medical Center | Wilson St. Greenhills | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27774 | COMMERCIAL | colomakath14@gmail.com | 20 |  |  | MCS+ | 08M461 | Red valve not correctly loaded. SPM error. | Checked valves and SPM for leak. | Simulation Run Performed maintenance Checked valve Checked rotor pumps Checked SPM Verified calibration parameters Functionality test | Unit is operational. No leak found. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306933612051958/5873306933612051958_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306933612051958/5873306933612051958_signature_31.png
| 2024-03-27 | 2024-03-27 11:09 | Cardinal Santos Medical Center | Wilson St. Greenhills | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27776 | COMMERCIAL | colomakath14@gmail.com | 20 |  |  | MCS+ | 08M461 | Red valve not correctly loaded. SPM error. | Checked valves and SPM for leak. | Simulation Run Performed maintenance Checked valve Checked rotor pumps Checked SPM Verified calibration parameters Functionality test | Unit is operational. No leak found. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306933614449062/5873306933614449062_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306933614449062/5873306933614449062_signature_31.png
| 2024-03-27 | 2024-03-27 11:09 | Cardinal Santos Medical Center | Wilson St. Greenhills | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27772 | COMMERCIAL | colomakath14@gmail.com | 20 |  |  | MCS+ | 08M461 | Red valve not correctly loaded. SPM error. | Checked valves and SPM for leak. | Simulation Run Performed maintenance Checked valve Checked rotor pumps Checked SPM Verified calibration parameters Functionality test | Unit is operational. No leak found. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306843614982066/5873306843614982066_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306843614982066/5873306843614982066_signature_31.png
| 2024-03-27 | 2024-03-27 11:09 | Cardinal Santos Medical Center | Wilson St. Greenhills | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27771 | COMMERCIAL | colomakath14@gmail.com | 20 |  |  | MCS+ | 08M461 | Red valve not correctly loaded. SPM error. | Checked valves and SPM for leak. | Simulation Run Performed maintenance Checked valve Checked rotor pumps Checked SPM Verified calibration parameters Functionality test | Unit is operational. No leak found. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306843616617580/5873306843616617580_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306843616617580/5873306843616617580_signature_31.png
| 2024-03-27 | 2024-03-27 11:09 | Cardinal Santos Medical Center | Wilson St. Greenhills | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27773 | COMMERCIAL | colomakath14@gmail.com | 20 |  |  | MCS+ | 08M461 | Red valve not correctly loaded. SPM error. | Checked valves and SPM for leak. | Simulation Run Performed maintenance Checked valve Checked rotor pumps Checked SPM Verified calibration parameters Functionality test | Unit is operational. No leak found. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306853614645850/5873306853614645850_signature_28.png | Kathleen Shane Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/5873306853614645850/5873306853614645850_signature_31.png
| 2024-03-27 | 2024-03-27 14:14 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27770 | DOH | nacosta@loveyourself.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 510 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1074 |  | GXIV | 831895 | Unit for maintenance and troubleshooting | Performed maintenance and troubleshooting | #NAME? | Unit is operational Modules A2, A3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873300760116021999/5873300760116021999_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5873300760116021999/5873300760116021999_signature_31.png
| 2024-03-27 | 2024-03-27 13:22 | Philippine Red Cross | Manila Service Road, Calamba Laguna | 4A | Laguna | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27769 | COMMERCIAL | prclaguna.calambadro@gmail.com | 25 |  |  | MCS+ | 02J046 | Unit for Preventive maintenance. Unit for calibration. | Preventive maintenance and calibration. | Performed Preventive Maintenance. Performed calibration. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873280852606273294/5873280852606273294_signature_28.png | Rommel Branco | https://www.jotform.com/uploads/macarewecare/222927978372067/5873280852606273294/5873280852606273294_signature_31.png
| 2024-03-27 | 2024-03-27 13:26 | Sta Rosa Community District Hospital | Sta Rosa, Olango, Lapulapu City | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27768 | DOH | ariestumz@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 312 |  | GX4 | 838534 | Due for Preventive Maintenance. | Conduct Preventive Maintenance, Clean Filter, Clean Plunger and ICore Generate IQR, Gxx, System Log, Assay Statistics | Conduct Preventive Maintenance, Clean Filter, Clean Plunger and ICore Generate IQR, Gxx, System Log, Assay Statistics | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873267129709771092/5873267129709771092_signature_28.png | Ariston Tumulak, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5873267129709771092/5873267129709771092_signature_31.png
| 2024-03-27 | 2024-03-27 09:15 | La Castellana Rural Health UnitMi | La Castellana, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | REPAIR ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 27767 | DOH | milouepascual@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 805 |  | GX4 | 838483 | Defective LCD monitor, HP Pro Display 19A, SN CNK9430B94 | Preventive Maintenance | -Install backup LCD monitor (temporary unit only), Philips LCD monitor 15". | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873124902317774980/5873124902317774980_signature_28.png | Mira Luna Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/5873124902317774980/5873124902317774980_signature_31.png
| 2024-03-27 | 2024-03-26 15:13 | Philippine Red Cross QC | Malakas St,QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27766 | MACARE-HBDC | prcqc.bloodservices@gmail.com | 28 |  |  | MCS+ | 02A227 | Encounter self test error non valital during self test. | Performed CPU reset Checked voltage output of cpu battery Reset cables. Checked keypad functionality. Checked connection | Performed troubleshooting | After performed the troubleshooting, error self test non valital still occur. Unit is unoperational for further troubleshooting suspect defective parts CPU board |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873016296547765763/5873016296547765763_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/5873016296547765763/5873016296547765763_signature_31.png
| 2024-03-27 | 2024-03-26 10:55 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomed@nch.doh.gov.ph |  | 27765 | MACARE-HBDC | zarabellehilario@yahoo.com | 28 |  |  | MCS + | 17H632SPG | CRC bios failed Low voltage output CPU battery | Checked unit Reset CPU board Checked CPU board battery output voltage .834volts Reset cables. | Troubleshooting | After performed the troubleshooting CRC Bios failed still occur Unit unoperational need to replace CPU board battery. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5873015266544249047/5873015266544249047_signature_28.png | zarabelle hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5873015266544249047/5873015266544249047_signature_31.png
| 2024-03-26 | 2024-03-26 13:21 | Davao Reproductive Health and Wellness Center | Jacinto St. Davao City | 11 | Davao del Sur | TB-HIV | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27764 | MACARE-HBDC | rhwclab@gmail.com | 25.8 | ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 125 ASSAY: Xpert HIV-1 Viral Load, COUNT: 790 |  | GX-IV 4d | 839236 | Error At Module B2 and B3 | Perform Calibration and Maintenance, Clean optics of the Module | Perform Xpc Check and Maintenance | Module B3 for Service Required | Data will be sent to Cepheid for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5872508289721985013/5872508289721985013_signature_28.png | Kent ferolin | https://www.jotform.com/uploads/macarewecare/222927978372067/5872508289721985013/5872508289721985013_signature_31.png
| 2024-03-26 | 2024-03-26 13:19 | Davao Reproductive Health and Wellness Center | Jacinto St. Davao City | 11 | Davao del Sur | TB-HIV | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27763 | DOH | rhwclab@gmail.com | 25.6 | ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 125 ASSAY: Xpert HIV-1 Viral Load, COUNT: 790 |  | GX-IV 4d | 838262 | For Calibration of Modules B1 and B3 | Calibrate the Modules C1 and C3 | Perform Calibration | Modules C1 and C3 Passed the Calirbation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5872505639722311866/5872505639722311866_signature_28.png | Kent ferolin | https://www.jotform.com/uploads/macarewecare/222927978372067/5872505639722311866/5872505639722311866_signature_31.png
| 2024-03-26 | 2024-03-26 13:21 | Davao Reproductive Health and Wellness Center | Jacinto St. Davao City | 11 | Davao del Sur | TB-HIV | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27762 | MACARE-HBDC | rhwclab@gmail.com | 25.8 | ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 125 ASSAY: Xpert HIV-1 Viral Load, COUNT: 790 |  | GX-IV 4d | 839236 | Error At Module B2 and B3 | Perform Calibration and Maintenance, Clean optics of the Module | Perform Xpc Check and Maintenance | Module B3 for Service Required | Data will be sent to Cepheid for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5872505269728828589/5872505269728828589_signature_28.png | Kent ferolin | https://www.jotform.com/uploads/macarewecare/222927978372067/5872505269728828589/5872505269728828589_signature_31.png
| 2024-03-26 | 2024-03-26 15:28 | Medellin Rural Health Unit | Poblacion Medellin Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27761 | DOH | rhumedellin@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1249 |  | GX4 | 210009597 | Stucked up cartridge in Module 4 SN# 210009597 | Removal of stucked cartridge , Generate GXX files, IQR, Sys Log, and back up database Conduct preventive maintenance of the genexpert machine | Removal of stucked cartridge , Generate GXX files, IQR, Sys Log, and back up database Conduct preventive maintenance of the genexpert machine | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5872475431114800615/5872475431114800615_signature_28.png | Raul Galano | https://www.jotform.com/uploads/macarewecare/222927978372067/5872475431114800615/5872475431114800615_signature_31.png
| 2024-03-26 | 2024-03-26 12:23 | Central Luzon Doctor's Hospital | Hospital drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27760 | COMMERCIAL | cldhlaboratory@cldh.com.ph | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 ASSAY: Xpert SARS-CoV-2, COUNT: 665 | 2024-03-14 | GXII-D | 830841 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Installation of xpertcheck application Archived test results Performed xpertcheck Generate reports Performed functionality test | Unit is operational C1 passed calibration | For calibration Module C2 | 1 xpertcheck kit | 97385 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5872396733415224708/5872396733415224708_signature_28.png | Floryvy Antonio | https://www.jotform.com/uploads/macarewecare/222927978372067/5872396733415224708/5872396733415224708_signature_31.png
| 2024-03-26 | 2024-03-26 11:51 | Central Luzon Doctor's Hospital | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27759 | COMMERCIAL | cldhlaboratory@cldh.com.ph | 22 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 ASSAY: Xpert SARS-CoV-2, COUNT: 665 | 2024-03-14 | GXIV-D | 839232 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Archived test results Installation of xpertcheck application Performed xpertcheck Generate reports Performed functionality test | Unit is operational B1,B2 and B3 passed calibration B4: 738203 failed xpertcheck | B4: 738203 for replacement | 4 xpertcheck kits | 97335 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5872396033419894566/5872396033419894566_signature_28.png | Floryvy Antonio | https://www.jotform.com/uploads/macarewecare/222927978372067/5872396033419894566/5872396033419894566_signature_31.png
| 2024-03-26 | 2024-03-26 11:53 | Philippine General Hospital | Taft. Ave, Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27758 | COMMERCIAL | Pghmolecularlab@gmail.com | 25 |  |  | GX16 | 844975 | Unit for Preventive Maintenance. | Preventive Maintenance. | Performed Preventive Maintenance. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5872350711462359093/5872350711462359093_signature_28.png | Kleante bautro | https://www.jotform.com/uploads/macarewecare/222927978372067/5872350711462359093/5872350711462359093_signature_31.png
| 2024-03-26 | 2024-03-26 10:36 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27757 | DOH | martertingzon20@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1244 | 2024-03-26 | GXIV | 838421 | Unit for Calibration | #NAME? | #NAME? | #NAME? |  | 1 xpc cartridge | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5872301574635698285/5872301574635698285_signature_28.png | Marter Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5872301574635698285/5872301574635698285_signature_31.png
| 2024-03-26 | 2024-03-26 10:16 | Gerona Rural Health Unit 2 | Gerona Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27756 | DOH | huwakie15@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 815 |  | GXIV-D | 831693 | Due for maintenance Re-installation of printer | Preventive maintenance Installation | Performed surfce cleaning, filter, icore, plunder rod and catridge bay Performed printer installation Performed functionality test Gathered reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5872289624191932045/5872289624191932045_signature_28.png | Artemio I. Nool III | https://www.jotform.com/uploads/macarewecare/222927978372067/5872289624191932045/5872289624191932045_signature_31.png
| 2024-03-25 | 2024-03-25 12:39 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27755 | DOH | rhucalbiga2018@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 782 | 2024-03-25 | GXIV | 838409 | n/a | #NAME? | #NAME? | - Unit is operational - Module B4 for monitoring (4012 error) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5871512245131481955/5871512245131481955_signature_28.png | vanessa mabag | https://www.jotform.com/uploads/macarewecare/222927978372067/5871512245131481955/5871512245131481955_signature_31.png
| 2024-03-25 | 2024-03-21 09:08 | Earnshaw Health Center | Earnshaw St., Sampaloc, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27753 | MACARE-HBDC | johndarrellturato@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 234 ASSAY: Xpert HIV-1 Viral Load, COUNT: 697 |  | GXIV-D-10C | 848262 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay and Plunger Rod Replaced Filter Performed Self-test Generate Reports | 3 Modules are Operational, 1 for Replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5871401923611886297/5871401923611886297_signature_28.png | John Darrell Turato | https://www.jotform.com/uploads/macarewecare/222927978372067/5871401923611886297/5871401923611886297_signature_31.png
| 2024-03-25 | 2024-03-20 14:16 | Calaca Rural Health Unit | Calaca, Batangas | 4A | Batangas | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27754 | DOH | loidamarasigan123@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1926 |  | GXIV-D-6C | 821300 | Modules 2 and 3 with Temperature Errors | XpertCheck | Cleaned Cartridge Bay and Plunger Rod Checked Filter if Clogged Performed XpertCheck on Modules 2 and 3 Generate Reports | 3 Modules are Operational, 1 Defective | Service Required: A2/ 730076 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5871401923618396159/5871401923618396159_signature_28.png | Loida C. marasigan | https://www.jotform.com/uploads/macarewecare/222927978372067/5871401923618396159/5871401923618396159_signature_31.png
| 2024-03-25 | 2024-03-22 14:55 | Detoxicare Molecular Diagnostics Inc. | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27752 | CONSORTIUM | imeehabaluyas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 131 ASSAY: Xpert MTB/RIF Ultra, COUNT: 131 | 2024-03-07 | GXIV 10C | 110018312 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed |  | XpertCheck Kit | 11269 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5871375743612082170/5871375743612082170_signature_28.png | Regene Mendezabal | https://www.jotform.com/uploads/macarewecare/222927978372067/5871375743612082170/5871375743612082170_signature_31.png
| 2024-03-25 | 2024-03-22 11:24 | Philippine Heart Center | East avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27751 | MACARE-HBDC | avabichuela3216@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 379 ASSAY: Xpert SARS-CoV-2, COUNT: 781 |  | GX16 D | 838655 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5871341693618933481/5871341693618933481_signature_28.png | Arjay Abichuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5871341693618933481/5871341693618933481_signature_31.png
| 2024-03-25 | 2024-03-22 10:14 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 27750 | MACARE-HBDC | roseannparnes@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 | 2024-03-22 | GX4 (10c) | 110014085 | Due for xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | B2 newly replaced,B1 ,B3 ,B4failed xpertcheck |  | Xpertcheck | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5871340863611362205/5871340863611362205_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/5871340863611362205/5871340863611362205_signature_31.png
| 2024-03-25 | 2024-03-22 09:27 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 27749 | MACARE-HBDC | cruzelle95@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 31 ASSAY: Xpert HPV, COUNT: 177 ASSAY: Xpert SARS-CoV-2, COUNT: 939 | 2024-03-22 | GX4 (10c) | 110016781 | Due for maintenance and xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance and xpertcheck | B1,B2,B3 0assed xpertcheck, B4 failed xpertcheck |  | Xpertcheck kits | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5871339223616838099/5871339223616838099_signature_28.png | Christelle Kimberly Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5871339223616838099/5871339223616838099_signature_31.png
| 2024-03-23 | 2024-03-22 13:22 | Camp Capinpin Army Station Hospital | Camp Capinpin | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27748 | DOH | cathialpha@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 139 |  | GX-IV-D-6C | 831896 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5870101564422767133/5870101564422767133_signature_28.png | Anne Catherine I Alpha | https://www.jotform.com/uploads/macarewecare/222927978372067/5870101564422767133/5870101564422767133_signature_31.png
| 2024-03-23 | 2024-03-21 13:19 | Sariaya Rural Health Unit | Sariaya | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27747 | DOH | marielinnetabril@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2098 |  | GX-IV-D-6C | 831890 | Stuck cartridge on module A1 | Removed stuck cartridge. | Manual removal of stuck cartridge. Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports. | Unit is operational. Module A4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5870101564424512405/5870101564424512405_signature_28.png | Marie Linnet Abril | https://www.jotform.com/uploads/macarewecare/222927978372067/5870101564424512405/5870101564424512405_signature_31.png
| 2024-03-22 | 2024-03-22 13:49 | Kalinga Health Hub | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27746 | CONSORTIUM | pearl.balderama@innovationsch.org | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 979 | 2024-03-05 | GXIV | 110008420 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed |  | XpertCheck Kit | 11259 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5868972288919444443/5868972288919444443_signature_28.png | Shad Famisaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5868972288919444443/5868972288919444443_signature_31.png
| 2024-03-22 | 2024-03-21 14:25 | University of The East - Ramon Magsaysay Memorial Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27745 | MACARE-HBDC | ljf7403@yahoo.com | 25 |  |  | MCS+ | 13C270 | Unit for calibration and maintenance | Performed calibration | #NAME? | Unit is operational All parameters are within range |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5868971708912363294/5868971708912363294_signature_28.png | Leonila Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5868971708912363294/5868971708912363294_signature_31.png
| 2024-03-22 | 2024-03-21 13:36 | University of The East - Ramon Magsaysay Memorial Hospital | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27744 | CONSORTIUM | lpcomendador@uerm.edu.ph | 25 | ASSAY: Xpert C. difficile BT, COUNT: 39 ASSAY: Xpert CT/NG, COUNT: 81 ASSAY: Xpert MTB/RIF Ultra, COUNT: 740 ASSAY: Others, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 25 | 2024-03-07 | GXIV | 847578 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed calibration |  | XpertCheck Kit Lot: 1000410214 | 11273 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5868971708919964565/5868971708919964565_signature_28.png | Osang Comendador | https://www.jotform.com/uploads/macarewecare/222927978372067/5868971708919964565/5868971708919964565_signature_31.png
| 2024-03-22 | 2024-03-22 12:07 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB - DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27743 | DOH | burauenlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 466 | 2024-03-22 | GXIV - 6C | 862127 | n/a | n/a | Preventive maintenance - Cleaned Cartridge bat, plunger rod, Fan Filter - Performed functionality test - Generate reports | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5868901937918879887/5868901937918879887_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5868901937918879887/5868901937918879887_signature_31.png
| 2024-03-21 | 2024-03-21 14:38 | University of Santo Tomas Hospital | España Street Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Rhea Anne Cañelas |  | rheaannecanelas65@gmail.com |  |  | 27742 | CONSORTIUM | ustimp@yahoo.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 911 ASSAY: Xpert BCR-ABL Ultra, COUNT: 8 | 2024-03-06 | GXIV-4D-6C | 110000711 | Unit is due for xpertcheck | Performed xpertcheck All modules passed. | Cleaned catridge bay Plunger maintenance Cleaned Icore Checked Filter Generate reports | Unit is operational | Work with Cienne Valerio | Xpertcheck-CE-5 | 11265 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5868130887403332512/5868130887403332512_signature_28.png | John Mark W. Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5868130887403332512/5868130887403332512_signature_31.png
| 2024-03-21 | 2024-03-21 11:48 | Baybay City Health Office | Baybay leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27741 | DOH | malitchelllanos@gmail.com | 18 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1207 | 2024-03-21 | GXIV - 6C | 838555 | n/a | n/a | #NAME? | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5868023960342015479/5868023960342015479_signature_28.png | Malitchel B. Llanos | https://www.jotform.com/uploads/macarewecare/222927978372067/5868023960342015479/5868023960342015479_signature_31.png
| 2024-03-21 | 2024-03-21 11:27 | The Medical City South Luzon | United Blvd., Sta Rosa Laguna | 4A | Laguna | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27740 | COMMERCIAL | tmcsl-molecularlab@themedicalcity.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 41 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 163 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GXIV | 839233 | Unit for PM. | PM | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5868011686268333363/5868011686268333363_signature_28.png | Ma. Veronica Agnas | https://www.jotform.com/uploads/macarewecare/222927978372067/5868011686268333363/5868011686268333363_signature_31.png
| 2024-03-21 | 2024-03-21 09:44 | Calamba Medical Center | National Highway, Crossing, Calamba Laguna | 4A | Laguna | Molecular Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27739 | COMMERCIAL | cmc.rtpcr@gmail.com | 20 | ASSAY: Xpert HBV VL, COUNT: 44 ASSAY: Xpert HPV, COUNT: 28 ASSAY: Xpert MTB/RIF Ultra, COUNT: 623 | 2024-03-11 | GXII | 812916 | Unit for PM. Unit for Xpertcheck. For module replacement of module B1. | PM, module replacement and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | GX module B1 (210096500) in replacement to defective B1 (685554). Xpertcheck Kit | SI# 97342 & 97324 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867962036262843977/5867962036262843977_signature_28.png | Andrea Gail B. Orgesvik | https://www.jotform.com/uploads/macarewecare/222927978372067/5867962036262843977/5867962036262843977_signature_31.png
| 2024-03-21 | 2024-03-21 09:19 | St. Elizabeth Hospital Inc. | General Santos City, South Cotabato | 12 | South Cotabato | Molecular lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27738 | COMMERCIAL | covidsupport@sehi.ph | 21.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 90 ASSAY: Xpert SARS-CoV-2, COUNT: 8 |  | Gx-IV | 110003713 | Module 1,2 and 3 H/W failed | Genexpert machine Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 1 operational modules | Module 2,3 and 4 requesting for module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867949442614530704/5867949442614530704_signature_28.png | Joana Simplicio | https://www.jotform.com/uploads/macarewecare/222927978372067/5867949442614530704/5867949442614530704_signature_31.png
| 2024-03-21 | 2024-03-21 09:32 | St. Elizabeth Hospital Inc. | General Santos City, South Cotabato | 12 | South Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27737 | COMMERCIAL | covidsupport@sehi.ph | 35.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 199 ASSAY: Xpert SARS-CoV-2, COUNT: 22 |  | Gx-IV | 831313 | Module 1 and 3 H/W failed due to temperature errors | Genexpert machine Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867945102611033813/5867945102611033813_signature_28.png | Joana simplicio | https://www.jotform.com/uploads/macarewecare/222927978372067/5867945102611033813/5867945102611033813_signature_31.png
| 2024-03-20 | 2024-03-20 13:33 | Victor R. Potenciano Medical Center | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27736 | CONSORTIUM | oro_janet@yahoo.com | 25 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 22 ASSAY: Xpert MTB/RIF Ultra, COUNT: 797 | 2024-03-07 | GXIV | 835600 | Unit for calibration and maintenance | Performed XpertCheck | #NAME? | Unit is operational B1: Service required B2: Insufficient cartridge B3: Passed B4: Insufficient cartridge No hardware error logs on B2. B2 is under observation. Use until hardware failure / temp. error Existing temp. errors on B4 B1 and B4 for reporting. Both modules are recommended for replacement |  | XpertCheck Kit Lot: 1000410214 | 11274 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867288288529496842/5867288288529496842_signature_28.png | Floriza Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5867288288529496842/5867288288529496842_signature_31.png
| 2024-03-20 | 2024-03-20 15:18 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27734 | MACARE-HBDC | sah.genexpertlaboratory@gmail.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1165 ASSAY: Xpert SARS-CoV-2, COUNT: 2167 |  | GXIV | 847728 | Unit for troubleshooting | Cleaned filter | #NAME? | - Unit is operational - Room temperature is recommended to be below 30 degrees Celsius (Current ave. temp. is 33 deg.) to avoid temperature errors on modules - Installation of additional air flow machine is recommended (i.e. electric fan pointed towards the back of GeneXpert) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867288268521175705/5867288268521175705_signature_28.png | Yollymel Mary Grejarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5867288268521175705/5867288268521175705_signature_31.png
| 2024-03-20 | 2024-03-20 12:06 | Our Lady of Lourdes Hospital | Sta. Mesa, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | jcmasaganda@ollh.ph |  | 27735 | CONSORTIUM | ljmalora@ollh.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 111 ASSAY: Xpert MTB/RIF Ultra, COUNT: 387 | 2024-03-07 | GXIV | 840974 | Unit for calibration and maintenance | Performed XpertCheck | #NAME? | Unit is operational All modules passed calibration |  | XpertCheck Kit Lot: 1000410214 | 11272 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867288268526083101/5867288268526083101_signature_28.png | Loureine Joy Obrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5867288268526083101/5867288268526083101_signature_31.png
| 2024-03-20 | 2024-03-20 14:09 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27733 | DOH | dy_rosemarie71@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1927 |  | GX-IV 4d | 838401 | For Relocation of GeneXpert | Relocation of GeneXpert | Relocation of GeneXpert | Successfully Relocated the Machine Frequently Occuring Error 5007 at Module B2 Monitor LCD Problem | Errors are sent to JotForm for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867287525299023448/5867287525299023448_signature_28.png | Rosemarie Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/5867287525299023448/5867287525299023448_signature_31.png
| 2024-03-20 | 2024-03-20 14:33 | Malapatan Rural Health Unit | Malapatan, Sarangani | 12 | Sarangani | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27732 | DOH | rhu.malapatan12@gmail.com | 27.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 |  | Gx-IV | 838419 | Genexpert Filter | Genexpert Cleaning Checking of errors Problem reporting Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules | Requesting for Genexpert filters |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867265726819707032/5867265726819707032_signature_28.png | Kamilah Nichol S. Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/5867265726819707032/5867265726819707032_signature_31.png
| 2024-03-20 | 2024-03-20 12:55 | Sarangani Health Care Facility | Alabel, Sarangani | 12 | Sarangani | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27731 | DOH | ronayadao@yahoo.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 722 |  | Gx-lV | 838542 | Module 3 under observation due to temperature errors | Genexpert cleaning Checking of errors Errors/module problem reporting Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867224426811479876/5867224426811479876_signature_28.png | Eunice Jurilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5867224426811479876/5867224426811479876_signature_31.png
| 2024-03-20 | 2024-03-15 11:41 | Cardona Rural Health Unit | Cardona | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27729 | DOH | rhucardona@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 631 |  | GX-IV-D-6C | 821215 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843611646155/5867125843611646155_signature_28.png | Khawla punzalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843611646155/5867125843611646155_signature_31.png
| 2024-03-20 | 2024-03-15 09:11 | Cainta City Health Office | Karangalan Dr. | NCR | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27726 | PBSP | ballesterrichard18@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 983 ASSAY: Xpert HIV-1 Viral Load, COUNT: 112 |  | GX-IV-D-6C | 838182 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | Unit is operational. Module B4/807137 failed. Modules B1, B2 and B3 passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843614080951/5867125843614080951_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843614080951/5867125843614080951_signature_31.png
| 2024-03-20 | 2024-03-20 09:38 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE Software Upgrade | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27727 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 23 | ASSAY: Xpert HPV, COUNT: 104 ASSAY: Xpert MTB/RIF Ultra, COUNT: 101 ASSAY: Xpert SARS-CoV-2, COUNT: 12704 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV-D-6C | 838627 | For software upgrade | Upgrade GeneXpert software | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Upgrade GeneXpert software from 5.1 to 6.4 Generated reports | Unit is operational. Modules A1, A2, B3, B4, C2, D2, and D4 not operational. Reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843614433144/5867125843614433144_signature_28.png | Jhon Patrick Ariola | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843614433144/5867125843614433144_signature_31.png
| 2024-03-20 | 2024-03-14 14:01 | Claro M. Recto District Hospital | Infanta | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27728 | PBSP | katelynaveryumerez@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1508 |  | GX-IV-D-6C | 807608 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. Modules C1 and C4 newly replaced. Modules C2 and C3 passed. | UPS has no power. Recommend to replace. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843614629031/5867125843614629031_signature_28.png | Lorelyn Miras | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843614629031/5867125843614629031_signature_31.png
| 2024-03-20 | 2024-03-19 14:38 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@themedicalcity.com |  | 27730 | COMMERCIAL | moleculardx@themedicalcity.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 32 ASSAY: Xpert CT/NG, COUNT: 44 ASSAY: Xpert HBV VL, COUNT: 434 ASSAY: Xpert HPV, COUNT: 334 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3731 ASSAY: Xpert HCV VL, COUNT: 67 ASSAY: Xpert HIV-1 Viral Load, COUNT: 572 ASSAY: Xpert SARS-CoV-2, COUNT: 474 | 2024-03-12 | GX-XVI-D-10C | 110015767 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. Module C2 failed. Other modules passed. Disabled module C2. |  | XpertCheck kit | DR# 11294 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843616745861/5867125843616745861_signature_28.png | Janica Katheryn Zamora | https://www.jotform.com/uploads/macarewecare/222927978372067/5867125843616745861/5867125843616745861_signature_31.png
| 2024-03-20 | 2024-03-20 10:02 | Philippine Red Cross | Port Area, Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27724 | COMMERCIAL | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for Preventive Maintenance. | Preventive Maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867116796265792221/5867116796265792221_signature_28.png | Lalaine ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5867116796265792221/5867116796265792221_signature_31.png
| 2024-03-20 | 2024-03-20 09:59 | Philippine Red Cross | Port Area, Intramuros Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27725 | COMMERCIAL | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 03J030 | Unit for Preventive Maintenance | Preventive Maintenance. | Performed Preventive Maintenance. Checked and cleaned rotor pumps. Verified calibration parameters. Performed simulation running. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867116796266211920/5867116796266211920_signature_28.png | Lalaine Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5867116796266211920/5867116796266211920_signature_31.png
| 2024-03-20 | 2024-03-20 10:02 | CHONG HUA HOSPITAL | FUENTE CEBU CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27723 | COMMERCIAL | chh_microbiology@chonghua.com.ph | 21 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 15 ASSAY: Xpert CT/NG, COUNT: 62 ASSAY: Xpert HIV-1 Viral Load, COUNT: 195 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1105 ASSAY: Xpert HCV VL, COUNT: 80 ASSAY: Xpert HIV-1 Viral Load, COUNT: 245 |  | GX-IV | 110020777 | ERROR 2008 ENCOUNTERED | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5867096714149456410/5867096714149456410_signature_28.png | Antonio Padernal | https://www.jotform.com/uploads/macarewecare/222927978372067/5867096714149456410/5867096714149456410_signature_31.png
| 2024-03-19 | 2024-03-19 10:46 | Morong Rural Health Unit | Morong Bataan | 3 | Bataan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27722 | DOH | lielyn.lim@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV-D | 831607 | CPU for replacement Defective CPU:2UA8461W0C | Repair Preventive maintenance | Performed CPU replacement Performed preventive maintenance Performed installation and setup of genexpert application Performed functionality test Gathered reports | Unit is operational New CPU relacement SN:2UA52027WB |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5866277075426808986/5866277075426808986_signature_28.png | Lielyn M. Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5866277075426808986/5866277075426808986_signature_31.png
| 2024-03-18 | 2024-03-18 15:46 | Sto Domingo District Hospital | Pulong Buli, Sto Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27721 | DOH | lhaydeesel@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 994 | 2024-03-11 | Genexpert IV | 831582 | Module A2 error 4006 | Performed Module Replacement A2(888301 - 210085270) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Module - 210085270 | 11290 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865575468717393110/5865575468717393110_signature_28.png | Cyndrell B. Samatra | https://www.jotform.com/uploads/macarewecare/222927978372067/5865575468717393110/5865575468717393110_signature_31.png
| 2024-03-18 | 2024-03-18 15:15 | Ciudad Medical Zamboanga | Mayor, Vitallino Agan Ave, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 27720 | COMMERCIAL | thairaphoebegrey13@gmail.com | 23 | ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 85 ASSAY: Xpert MTB/RIF Ultra, COUNT: 417 ASSAY: Xpert HCV VL, COUNT: 35 | 2023-03-18 | GX-4 6colors | 834128 | Module 2 excessive error 1002 | Complete Preventive Maintenence. And site assessment for Temperature related issues. | Complete Preventive Maintenence. And site assessment for Temperature related issues. Undergone test run for the said module. | All modules are not available. Module 1 error 4017 while module 2 error 1002. | All modules must be replaced. Temperature of the room and Machine are in ideal range. Genexpert Machine Must be xpertcheck Calibrated yearly | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865566934955345304/5865566934955345304_signature_28.png | Shereemer A. Cheong-Ynawat | https://www.jotform.com/uploads/macarewecare/222927978372067/5865566934955345304/5865566934955345304_signature_31.png
| 2024-03-18 | 2024-03-18 11:18 | Macare In-House | Quezon City | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Ronnel Belizar |  | rcb.macare@gmail.com |  | Computer Monitor | 27719 | MACARE-HBDC | rcb.macare@gmail.com | 25 |  |  | HP ProDisplay P19A | CNK8322445 | Unit for repair | Repaired unit | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865412273616728804/5865412273616728804_signature_28.png | Macare Medical, Inc. | https://www.jotform.com/uploads/macarewecare/222927978372067/5865412273616728804/5865412273616728804_signature_31.png
| 2024-03-18 | 2024-03-14 10:26 | Dingras District Hospital | Brgy Suyo, Dingras , Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27718 | PBSP | kenacudapas1022@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 725 ASSAY: Xpert MTB/RIF Ultra, COUNT: 725 |  | GXIV-D-6C | 807597 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics Performed Self-test Performed XpertCheck Generate Reports | 3 Modules are Operational, 1 Defective | 2 Newly Replaced, 1 Failed, 1 Passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865403963612991713/5865403963612991713_signature_28.png | Kenadine Cudapas | https://www.jotform.com/uploads/macarewecare/222927978372067/5865403963612991713/5865403963612991713_signature_31.png
| 2024-03-18 | 2024-03-14 08:04 | The Black Nazarene Hospital | Brgy 2, San Nicolas, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27717 | CONSORTIUM | tbnhi.lab@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 48 ASSAY: Xpert MTB/RIF Ultra, COUNT: 48 | 2024-03-05 | GXIV-D-6C | 835542 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod and Icore Checked if Filter is Clogged Performed Self-test Performed XpertCheck Generate Reports | All Modules are Operational | All Modules Passed | XpertCheck Kit | DR 11263 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865403063615108349/5865403063615108349_signature_28.png | Carmella A. Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5865403063615108349/5865403063615108349_signature_31.png
| 2024-03-18 | 2024-03-14 15:28 | Batanes General Hospital | National Road , Brgy. Kayhuvokan Basco, Batanes | 2 | Batanes | Laboratory | NO | MAINTENANCE | CEPHEID | Cienne Mae Valerio |  | ciennevalerio@gmail.com | ciennevalerio@gmail.com |  | 27716 | PBSP | lenyda@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 277 ASSAY: Xpert SARS-CoV-2, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 64 ASSAY: Xpert SARS-CoV-2, COUNT: 30 |  | GXIV-2D-6C | 834839 | Repack of old GX Unit 6C | Packed GX Unit for transferring. Defective UPS | Cleaned machine Cleaned cartridge bay Cleaned Filter Plunger Rod Maintenance Gather Reports Archived Tests Back up old files | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865392673617371055/5865392673617371055_signature_28.png | Eleonor D. Argonza | https://www.jotform.com/uploads/macarewecare/222927978372067/5865392673617371055/5865392673617371055_signature_31.png
| 2024-03-18 | 2024-03-14 11:54 | Batanes General Hospital | National Road , Brgy. Kayhuvokan Basco, Batanes | 2 | Batanes | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Cienne Mae Valerio |  | ciennevalerio@gmail.com |  |  | 27714 | MACARE-HBDC | lenyda@yahoo.com | 18 |  | 2024-03-14 | GXIV-4D-10C | 110001652 | CPU not powering on. | Install Network Switch between 110001652 and 110020685 to be able to use two machine. Defective CPU SN *CPH0621F22206* | Troubleshoot CPU still the same Performed maintenance Cleaned filter Cleaned cartridge bay Generate Reports | Unit is operational CPU for reporting |  | Network Switch | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865392673615729120/5865392673615729120_signature_28.png | Eleonor D. Argonza | https://www.jotform.com/uploads/macarewecare/222927978372067/5865392673615729120/5865392673615729120_signature_31.png
| 2024-03-18 | 2024-03-14 11:03 | Batanes General Hospital | National Road , Brgy. Kayhuvokan Basco, Batanes | 2 | Batanes | Laboratory | NO | INSTALLATION | CEPHEID | Cienne Mae Valerio |  | ciennevalerio@gmail.com | ciennevalerio@gmail.com |  | 27715 | PBSP | lenyda@yahoo.com | 20 |  |  | GXIV-4D-10C | 110020685 | Unit due for Installation | Defective UPS SN SKPHU2KV230910025 UPS Rating is 50Hz Temporarily using AVR Site is with Back up Generator in case of power outage UPS for Reporting | Install GX Unit Install CPU Install Monitor Install Barcode Scanner Install Printer Import Assay Create Admin Account for End user Defective UPS | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865392673618963071/5865392673618963071_signature_28.png | Eleonor D. Argonza | https://www.jotform.com/uploads/macarewecare/222927978372067/5865392673618963071/5865392673618963071_signature_31.png
| 2024-03-18 | 2024-03-14 08:04 | The Black Nazarene Hospital | Brgy 2, San Nicolas, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27712 | CONSORTIUM | tbnhi.lab@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 48 ASSAY: Xpert MTB/RIF Ultra, COUNT: 48 | 2024-03-05 | GXIV-D-6C | 835542 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod and Icore Checked if Filter is Clogged Performed Self-test Performed XpertCheck Generate Reports | All Modules are Operational | All Modules Passed | XpertCheck Kit | DR 11263 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865390883611307740/5865390883611307740_signature_28.png | Carmella A. Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5865390883611307740/5865390883611307740_signature_31.png
| 2024-03-18 | 2024-03-14 10:26 | Dingras District Hospital | Brgy Suyo, Dingras , Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27713 | PBSP | kenacudapas1022@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 725 ASSAY: Xpert MTB/RIF Ultra, COUNT: 725 |  | GXIV-D-6C | 807597 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, and Icore Optics Performed Self-test Performed XpertCheck Generate Reports | 3 Modules are Operational, 1 Defective | 2 Newly Replaced, 1 Failed, 1 Passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865390883616779468/5865390883616779468_signature_28.png | Kenadine Cudapas | https://www.jotform.com/uploads/macarewecare/222927978372067/5865390883616779468/5865390883616779468_signature_31.png
| 2024-03-18 | 2024-03-15 14:11 | Quezon City Jail | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | TB Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27711 | PBSP | carleecruz0@gmail.com | 25 |  |  | GXIV-D 6C | 838881 | None | None | Backed-up results Packed-up system and prepared for transfer | Unit is operational and packed. Ready for transfer |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5865288293611577241/5865288293611577241_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5865288293611577241/5865288293611577241_signature_31.png
| 2024-03-17 | 2024-03-12 14:11 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27710 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3050 ASSAY: Xpert SARS-CoV-2, COUNT: 4527 ASSAY: Others, COUNT: 10 | 2024-03-11 | GX16 D (10c) | 110014171 | For module replacement D1 210069284 - 614213 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | Genexpert 10 color module | 11291 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5864811084077156674/5864811084077156674_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5864811084077156674/5864811084077156674_signature_31.png
| 2024-03-17 | 2024-03-12 14:48 | National Childrens Hospital | E. Rodriguez QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27709 | MACARE-HBDC | zarabellehilario@yahoo.com | 26 |  |  | MCS + | 17H632SPG | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5864810564075522667/5864810564075522667_signature_28.png | Zarabelle Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5864810564075522667/5864810564075522667_signature_31.png
| 2024-03-17 | 2024-03-13 08:44 | Payatas B Health Center | Bulacan st.,Payatas QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27708 | MACARE-HBDC | racquelmrocero@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2287 |  | GX4 D | 831899 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5864810154071994500/5864810154071994500_signature_28.png | Racquel M.Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5864810154071994500/5864810154071994500_signature_31.png
| 2024-03-17 | 2024-03-13 14:21 | NTRL | Filinvest, Alabang Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27707 | PBSP | eradenice@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2512 | 2024-03-12 | GX4 D | 807606 | For module replacement B3 662385 - 713880 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | B1 and B3 operational B2 encounter 4009 and B4 encounter 2032 disable |  | Genexpert 6 color module | G6894 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5864809754073358261/5864809754073358261_signature_28.png | Era Denice R. Acueza | https://www.jotform.com/uploads/macarewecare/222927978372067/5864809754073358261/5864809754073358261_signature_31.png
| 2024-03-17 | 2024-03-14 11:16 | Rodriguez RHU 1 | J.P Rizal Rodriguez Rizal | 4A | Rizal | Laboratory | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27706 | PBSP | karlbanal25@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 |  | GX4 D (10c) | 120000503 | For installation and for re pack old unit s/n 807580 | Installation | Installation and repacking of old unit | Operational | Old unit complete repacking |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5864808334079430139/5864808334079430139_signature_28.png | Karl Michael Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5864808334079430139/5864808334079430139_signature_31.png
| 2024-03-17 | 2024-03-14 14:15 | San Mateo RHU 5 | Modesta, San Mateo | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27705 | PBSP | mmacahis@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 617 | 2024-03-12 | GX4 D | 838161 | For module replacement B3 807009 - 711149 B4 210091113 - 715597 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | Genexpert 6 color module | G6896 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5864807844076334363/5864807844076334363_signature_28.png | Monina Rica Macahis - De guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5864807844076334363/5864807844076334363_signature_31.png
| 2024-03-15 | 2024-03-13 14:26 | Bagong Silang Ph. 1 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27704 | PBSP | alcaydeirene07@gmail.com | 25 |  | 2024-03-13 | GXIV | 838856 | Module B3, B4 for replacement | Replaced B3, B4 | #NAME? | Unit is operational |  | GeneXpert Module SN: 719046, 699744 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5863002919501225254/5863002919501225254_signature_28.png | Airene Alcayde | https://www.jotform.com/uploads/macarewecare/222927978372067/5863002919501225254/5863002919501225254_signature_31.png
| 2024-03-15 | 2024-03-15 13:39 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27703 | PBSP | Cyrilcordova0515@gmail.com | 28 |  |  | GX 4 | 804838 | Defective Module C4 sn699776 | New module Installed C4 717815 Clean Fan And Fan Filter Data of genexpert | New module Installed C4 717815 Clean Fan And Fan Filter Data of genexpert | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862918771115858963/5862918771115858963_signature_28.png | Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5862918771115858963/5862918771115858963_signature_31.png
| 2024-03-15 | 2024-03-15 13:46 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27702 | PBSP | venvalizado@gmail.com | 25 |  |  | GXIV 10C | 110020277 | Unit for installation | Installed unit | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862913919502846555/5862913919502846555_signature_28.png | Venmarie Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5862913919502846555/5862913919502846555_signature_31.png
| 2024-03-15 | 2024-03-15 11:47 | Tagaytay City Health Office | Acle St. Tagaytay | 4A | Cavite | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27701 | PBSP | edith.mirando@gmail.com | 25 |  | 2024-03-11 | GXIV | 807583 | Unit for PM. For installation of backup laptop. | PM and installation. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed installation of backup laptop. Generated reports. | Unit is operational. |  | Lenovo ThinkPad Laptop (sn: CPH1019F03275). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862858284448306135/5862858284448306135_signature_28.png | Edith B.Mirando | https://www.jotform.com/uploads/macarewecare/222927978372067/5862858284448306135/5862858284448306135_signature_31.png
| 2024-03-15 | 2024-03-15 09:42 | Amai pak-pak Medical Center | Marawi City | 10 | Lanao del Norte | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 27700 | PBSP | rbarco2007@gmail.com | 24 |  | 2024-01-04 | GX4 | 807457 | Module C1 haraware fail. | Save files. | Save Installation Qualification Reports Save system Log | Module C1 not ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862768476324749745/5862768476324749745_signature_28.png | Raisah P. Barco | https://www.jotform.com/uploads/macarewecare/222927978372067/5862768476324749745/5862768476324749745_signature_31.png
| 2024-03-14 | 2024-03-14 21:19 | Amai Pak-pak Medical Center | Marawi City | 10 | Lanao del Norte | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 27699 | PBSP | rbarco2007@gmail.com | 23 |  | 2024-03-14 | GX4 10C | 110020384 | For Installation | Installation. | Install GeneXpert Machine Install Printer Install LCD Monitor Install CPU Install UPS | GeneXpert Machine ok to use. Printer ok to use UPS ok to use CPU ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862320555619397624/5862320555619397624_signature_28.png | Raisah P. Barco | https://www.jotform.com/uploads/macarewecare/222927978372067/5862320555619397624/5862320555619397624_signature_31.png
| 2024-03-14 | 2024-03-14 18:57 | RURAL HEALTH UNIT OF SAPANG DALAGA | SAPANG DALAGA, MISAMIS OCCIDENTAL | 10 | Misamis Occidental | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27698 | PBSP | cherryelogsong@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 492 |  | GX-IV | 847252 | FOR MODULE B3 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B3 REPLACEMENT: DEF SN:910152 NEW SN:698867 | EQUIPMENT IS OPERATIONAL | UPS FOR REPORTING |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862235165742387654/5862235165742387654_signature_28.png | Cherry Mae Elogsong | https://www.jotform.com/uploads/macarewecare/222927978372067/5862235165742387654/5862235165742387654_signature_31.png
| 2024-03-14 | 2024-03-14 17:10 | Bicol Medical Center | Naga | 5 | Camarines Sur | Bacteriology Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27697 | PBSP | disuanco3@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1384 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7436 ASSAY: Xpert SARS-CoV-2, COUNT: 2463 ASSAY: Xpert HIV-1 Viral Load, COUNT: 560 ASSAY: Others, COUNT: 10 |  | GX4-6C | 842225 | A2/758306 - temperature error. For module replacement | Module Replaced: A2 - 758306>>702996 Cleaned cartridge bays, plunger rod, icore Generate reports | Replaced module A2/758306 | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862179982327633897/5862179982327633897_signature_28.png | Jb disuanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5862179982327633897/5862179982327633897_signature_31.png
| 2024-03-14 | 2024-03-14 17:09 | Lanao Del Norte Provincial Hospital | Upper Sagadan, Baroy, Lanao Del Norte | 10 | Lanao del Norte | Tb laboratory | YES | Pull and packed the 6c machine for transfer | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27696 | PBSP | lnphlab2021@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1663 |  | GX4 | 807673 | Pull out and packed the genexpert machine, cpu unit, printer, monitor, bar code reader, and the UPS Back-up the machine, gxx files, generate sys log, IQR, | Pull out and packed the genexpert machine, cpu unit, printer, monitor, bar code reader, and the UPS Back-up the machine, gxx files, generate sys log, IQR, | Pull out and packed the genexpert machine, cpu unit, printer, monitor, bar code reader, and the UPS Back-up the machine, gxx files, generate sys log, IQR, | All units are already packed for transfer to Matungao RHU. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862167591114527816/5862167591114527816_signature_28.png | Anne Marie Tubio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5862167591114527816/5862167591114527816_signature_31.png
| 2024-03-14 | 2024-03-14 16:54 | Lanao Del Norte Provincial Hospital | Upper Sagadan, Baroy, Lanao Del Norte | 10 | Lanao del Norte | Tb laboratory | YES | INSTALLATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27695 | PBSP | lnphlab2021@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 |  | GX4 | 120000612 | Installation of 10C genexpert machine, printer, UPS, CPU, Monitor | IInstallation of 10C genexpert machine, printer, UPS, CPU, Monitor Pull and repack 6c genexpert for transfer. Installation of XDR assay, MTB-ultra, HIV-viral load back-up database | Installation of 10C genexpert machine, printer, UPS, CPU, Monitor Pull and repack 6c genexpert for transfer. Installation of XDR assay, MTB-ultra, HIV-viral load back-up database | All modules are operational. Except the UPS is defective during testing no power. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862164761118596076/5862164761118596076_signature_28.png | Anne Marie Tubio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5862164761118596076/5862164761118596076_signature_31.png
| 2024-03-14 | 2024-03-14 14:59 | Gabaldon Rural Health Unit | South Poblacion, Gabaldon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27694 | PBSP | helengrace.parungao@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 288 | 2024-03-14 | Genexpert II | 834804 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration (B1/787168 = Service Required) (B2/787169 = Service Required) Performed Functionality Test | Unit is Opetional | Module B1 and B2 for Replacement | Xpertcheck 2pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862095203729340986/5862095203729340986_signature_28.png | Helen Grace Parungao | https://www.jotform.com/uploads/macarewecare/222927978372067/5862095203729340986/5862095203729340986_signature_31.png
| 2024-03-14 | 2024-03-14 14:37 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27693 | PBSP | bmcgenexpertlab@yahoo.com | 25 |  | 2024-03-14 | GXIV | 834623 | Modules A3, A4 for replacement | Replaced modules A3, A4 | #NAME? | Unit is operational |  | GeneXpert Module SN: 702434, 702034 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862081115107725910/5862081115107725910_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5862081115107725910/5862081115107725910_signature_31.png
| 2024-03-14 | 2024-03-14 14:07 | San Fernando Rural Health Unit | San Fernando | 5 | Camarines Sur | Laboratory Gx4-6C | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27692 | DOH | josephus_garfin@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1088 |  | GX4-6C | 838496 | For Preventive Maintenance B1/722557 Temperature Error, for Investigation | Cleaned Cartridge Bays, Plunger rod, icore Generate report | Performed Preventive Maintenance | Unit is Operational | Insufficient error collected for B1/722557 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862061338015243989/5862061338015243989_signature_28.png | Josephus Garfin | https://www.jotform.com/uploads/macarewecare/222927978372067/5862061338015243989/5862061338015243989_signature_31.png
| 2024-03-14 | 2024-03-14 13:27 | Tambulig rural health unit | Tambulig zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | TROUBLESHOOTING Preventive Maintenence | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27691 | PBSP | tambuligmho@yahoo.com | 27 |  | 2023-03-14 | GX-2 6c | 834828 | System unit unable to power up. Module 2 failed in xpertcheck calibration last February 20, 2024 | Replaced module 787279 with module 714788 | Module Replacement, Preventive maintenence and system unit troubleshooting. | All modules are available but system unit must be replaced | System unit unable to start up and must be replaced | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862036890313435932/5862036890313435932_signature_28.png | Lucille Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5862036890313435932/5862036890313435932_signature_31.png
| 2024-03-14 | 2024-03-14 12:55 | Guiguinto Rural Health Unit | Guiguinto, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27690 | PBSP | joie_calica@yahoo.com | 25 |  |  | GXIV | 807466 | UPS for troubleshooting | Performed troubleshooting | #NAME? | Unit is operational UPS is recommended for replacement Unit is already packed for PBSP pull out (Rovering unit) UPS SN: GA64E45555 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5862015283105278640/5862015283105278640_signature_28.png | Ligaya Calica | https://www.jotform.com/uploads/macarewecare/222927978372067/5862015283105278640/5862015283105278640_signature_31.png
| 2024-03-14 | 2024-03-14 10:42 | Cotabato Regional and Medical Center - Molecular Laboratory | National Highway, Cotabato City | BARMM | Maguindanao | Molecular Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27689 | PBSP | crmcmolecularlab@gmail.com | 22.5 | ASSAY: Xpert SARS-CoV-2, COUNT: 12844 |  | GX-IV 4D | 834442 | Due for Preventive Maintenance and Module B2 encounter multiple temperature related errors | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Troubleshooting | Module B2 reported for further assessment. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861963537315020987/5861963537315020987_signature_28.png | Kevin Dave G. Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5861963537315020987/5861963537315020987_signature_31.png
| 2024-03-14 | 2024-03-14 10:48 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27688 | PBSP | crmcmolecularlab@gmail.comKevim | 23.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 12844 |  | Gx-IV | 834441 | N/A | Genexpert machine Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861962455475355496/5861962455475355496_signature_28.png | Kevin Dave G. Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5861962455475355496/5861962455475355496_signature_31.png
| 2024-03-14 | 2024-03-14 09:47 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27687 | DOH | crmcmolecularlab@gmail.comn | 23.5 | ASSAY: Xpert SARS-CoV-2, COUNT: 12844 |  | Gx-IV | 837662 | Module 2 for service required | Perform Xpertcheck at module 2 for troubleshooting Genexpert machine Cleaning Checking of errors Problem reporting Data collection | Genexpert machine Preventive maintenance Perform Genexpert module troubleshooting | 3 operational modules | Module 2 waiting for replacement module |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861960947424875446/5861960947424875446_signature_28.png | Kevin Dave G. Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5861960947424875446/5861960947424875446_signature_31.png
| 2024-03-14 | 2024-03-14 10:52 | Matalom Community Hospital | Matalom Leyte | 8 | Leyte | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27686 | WHO | fegorne@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 458 | 2024-03-14 | GXIV - 6C | 847227 | Module B3/910533 defective | #NAME? | - Replaced module B3/910533 into 885108 - Performed functionality test - Performed preventive maintenance (Clean Cartridge bay, Plunger rod, Fan Filter) - Generate reports - Pullout defective module SN/910533 | unit is operational |  | Genexpert module 6C SN: 885108 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861943516329314638/5861943516329314638_signature_28.png | FE D. GORNE | https://www.jotform.com/uploads/macarewecare/222927978372067/5861943516329314638/5861943516329314638_signature_31.png
| 2024-03-14 | 2024-03-13 22:32 | Cotabato Regional and Medical Center | National Highway, Cotabato City | BARMM | Maguindanao | TB-DOTS | YES | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27685 | PBSP | crmctblab@gmail.com | 21.5 | ASSAY: Xpert HBV VL, COUNT: 31 ASSAY: Xpert HPV, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 201 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-IV 4D | 807463 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861931767315226820/5861931767315226820_signature_28.png | Shaneen Ayn Meditar | https://www.jotform.com/uploads/macarewecare/222927978372067/5861931767315226820/5861931767315226820_signature_31.png
| 2024-03-14 | 2024-03-13 22:11 | Guimbal RHU | Guimbal, Iloilo | 7 | Aklan | TB laboratory | NO | INSTALLATION | CEPHEID | Harold Toco |  | htoco@macare-medicals.com | guimbalrhulaboratory@gmail.com |  | 27684 | PBSP | toto.davis.td@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 | 2024-03-14 | GX4-10C | 110013745 | None; For unit installation. | Perforned unit installation. Perforned functionality test. Conduct operator's basic training. | Performed unit installation. Disassemble of old 6C unit ready for pullout. | Unit is operational. | N/A | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861918551126754627/5861918551126754627_signature_28.png | Davis James Secapuri | https://www.jotform.com/uploads/macarewecare/222927978372067/5861918551126754627/5861918551126754627_signature_31.png
| 2024-03-14 | 2024-03-13 22:44 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27683 | PBSP | crmctblab@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 358 ASSAY: Xpert HPV, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3240 ASSAY: Xpert HIV-1 Viral Load, COUNT: 249 ASSAY: Xpert HIV-1 Viral Load, COUNT: 318 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3240 ASSAY: Xpert HPV, COUNT: 40 ASSAY: Xpert HBV VL, COUNT: 358 | 2024-03-13 | Gx-XVI | 839307 | Module B3 and B4 for Service Required Due for Calibration | Replace defective module B3 SN: 870481 with new module SN:717133 Replace defective module B4 SN: 868766 with new module SN:717192 Perform Xpertcheck/Calibration Perform Genexpert cleaning and checking of errors Data collection | Perform Genexpert machine module replacement/repair Perform Genexpert machine Xpertcheck/Calibration Perform Genexpert machine preventive maintenance | 14 operational modules 2 modules for service required (A1 & A2) 3 modules under observation due to Xpertcheck result (D2, D3, & D4 skipped then service required) |  | Genexpert module 6-color SN:717133, SN:717192 | Cepheid case: 1484259 & 1484261 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861853887426976100/5861853887426976100_signature_28.png | Shannon Ayn Meditar | https://www.jotform.com/uploads/macarewecare/222927978372067/5861853887426976100/5861853887426976100_signature_31.png
| 2024-03-13 | 2024-02-23 10:28 | Cardinal Santos Medical Center | Wilson St., | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | egatchalian@csmc.ph |  | 27682 | COMMERCIAL | egatchalian@csmc.ph | 25 | ASSAY: Xpert HBV VL, COUNT: 103 ASSAY: Xpert HPV, COUNT: 22 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1051 |  | GX-IV-D-10C | 110003714 | Module B4 has error 1002/4012 | Checked module B4 | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Checked module B4 Generated reports | Modules B2 and B3 operational. Module B1/210022152 for replacement failed in XpertCheck. Module B4/210048235 not operational, for replacement due to error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847158893618692814/5847158893618692814_signature_28.png | Elaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5847158893618692814/5847158893618692814_signature_31.png
| 2024-03-13 | 2024-03-13 16:25 | Buug rural health unit | Poblacion buug, Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27681 | PBSP | chingskie1980@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1120 | 2024-03-13 | GX-2 6c | 834832 | Module 2 temperature error | Replaced module 787228 with Module 807025 | Module Replacement B2 and Preventive Maintenence | All 2 modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861283737818941178/5861283737818941178_signature_28.png | Ratchel D. Guerrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5861283737818941178/5861283737818941178_signature_31.png
| 2024-03-13 | 2024-03-12 14:06 | Philippines Red Cross Dagupan | Dagupan City | 1 | Pangasinan | Laboratory | NO | Pullout | HAEMONETICS | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27680 | MACARE-HBDC | prcpangasinanbloodcenter@gmail.com | 25 |  |  | MCS+ | 03A043 | #NAME? | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861282095705371083/5861282095705371083_signature_28.png | Mae Dianne Alano | https://www.jotform.com/uploads/macarewecare/222927978372067/5861282095705371083/5861282095705371083_signature_31.png
| 2024-03-13 | 2024-03-12 13:00 | Sta. Barbara Rural Health Unit | Sta. Barbara, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27679 | DOH | dianeruiz147@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 386 | 2024-02-05 | GXIV-4-D-6C | 838438 | #NAME? | #NAME? | #NAME? | Unit is operational | Module B3 failed XpertCheck. Remaining Modules passed XpertCheck | XpertCheck Kits | DR 11187 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861279865703465452/5861279865703465452_signature_28.png | Diane Joy Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5861279865703465452/5861279865703465452_signature_31.png
| 2024-03-13 | 2024-03-12 09:42 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27678 | DOH | allan.vinluan2013@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1049 |  | GXIV-4-D-6C | 838472 | #NAME? | #NAME? | #NAME? | #NAME? | For observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861278325707819489/5861278325707819489_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5861278325707819489/5861278325707819489_signature_31.png
| 2024-03-13 | 2024-03-13 15:45 | Mariano Marcos Memorial Hospital and Medical Center | Manila North Road, Batac, Ilocos Norte | 1 | Ilocos Norte | Molecular | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27677 | PBSP | dennisoncatindig@gmail.com | 24 |  |  | GXIV-D-10C | 120000490 | Unit is due for Installation | Installation | Install GeneXpert System | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861266730249657414/5861266730249657414_signature_28.png | DENNISON CATINDIG | https://www.jotform.com/uploads/macarewecare/222927978372067/5861266730249657414/5861266730249657414_signature_31.png
| 2024-03-13 | 2024-03-13 15:45 | Mariano Marcos Memorial Hospital and Medical Center | Manila North Road, Batac, Ilocos Norte | 1 | Ilocos Norte | Molecular | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27676 | PBSP | dennisoncatindig@gmail.com | 24 |  |  | GXIV-D-10C | 120000490 | Unit is due for Installation | Installation | Install GeneXpert System | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861265430243879937/5861265430243879937_signature_28.png | DENNISON CATINDIG | https://www.jotform.com/uploads/macarewecare/222927978372067/5861265430243879937/5861265430243879937_signature_31.png
| 2024-03-13 | 2024-03-13 15:05 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | TB Dots | YES | INSTALLATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27675 | PBSP | sanjoseantiquegxsite@gmail.com | 24 |  |  | 10C GXIV-4D | 120000591 | None | None | -Installation of new 10 Colors GeneXpert unit -Setup GeneXpert workstation | All modules are functional. | Unit is daisy-chained with DOH unit SN: 838412. UPS and printer not installed; site used their own units. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861231534607687655/5861231534607687655_signature_28.png | Oscar Delano R. Salazar, IV | https://www.jotform.com/uploads/macarewecare/222927978372067/5861231534607687655/5861231534607687655_signature_31.png
| 2024-03-13 | 2024-03-13 14:00 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | TB-Dots | YES | REPAIR CALIBRATION MAINTENANCE Pull-out | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27674 | PBSP | sanjoseantiquegxsite@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2795 ASSAY: Xpert HIV-1 Viral Load, COUNT: 190 | 2024-03-13 | 6C GXIV-4D | 804714 | #NAME? | -Remove stuck cartridge in module C1 -Replace defective module C3/886646 with module SN: 715827 -Run XpertCheck on module C1, C2, and C4 -Pack up for pull-out | #NAME? | #NAME? |  | -900-0491RS GeneXpert Module 6-Colors -XpertCheck-CE-5 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861220381128525508/5861220381128525508_signature_28.png | Oscar Delano R. Salazar, IV | https://www.jotform.com/uploads/macarewecare/222927978372067/5861220381128525508/5861220381128525508_signature_31.png
| 2024-03-13 | 2024-03-07 14:38 | San Ildefonso Rural Health Unit | San Ildefonso, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27673 | WHO | marvinsanchez06@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 636 | 2024-03-07 | GXIV | 846041 | Module B4 for replacement | Replaced module B4 | #NAME? | Unit is operational |  | GeneXpert Module SN: 698949 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861217502328741064/5861217502328741064_signature_28.png | Marvin Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/5861217502328741064/5861217502328741064_signature_31.png
| 2024-03-13 | 2024-03-13 14:28 | Mariano Marcos Memorial Hospital and Medical Center | Manila North Road, Batac, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | Pull-out | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27672 | PBSP | dennisoncatindig@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 36 ASSAY: Xpert HPV, COUNT: 198 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1212 ASSAY: Xpert SARS-CoV-2, COUNT: 3806 |  | GXIV-D-6C | 838168 | Unit is for Pull-out | Pull-out | Backup Files Uninstall GX System Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861217270247211688/5861217270247211688_signature_28.png | DENNISON B. CATINDIG | https://www.jotform.com/uploads/macarewecare/222927978372067/5861217270247211688/5861217270247211688_signature_31.png
| 2024-03-13 | 2024-03-13 14:32 | Maysan 3S Health Station | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27671 | PBSP | neldimatera@gmail.com | 25 |  |  | GXIV | 120000727 | Unit for installation | Installed unit | #NAME? | Unit is operational Unit is ready to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861209982323017226/5861209982323017226_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/5861209982323017226/5861209982323017226_signature_31.png
| 2024-03-13 | 2024-03-13 14:27 | Maysan 3S Health Center | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27670 | PBSP | neldimatera@gmail.com | 25 |  | 2024-03-13 | GXIV | 814151 | Module C1 for replacement | Replaced module C1 | #NAME? | Unit is operational |  | GeneXpert Module SN: 716234 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861207112323562419/5861207112323562419_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/5861207112323562419/5861207112323562419_signature_31.png
| 2024-03-13 | 2024-03-13 13:55 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | Tb Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27669 | DOH | sanjoseantiquegxsite@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2795 ASSAY: Xpert HIV-1 Viral Load, COUNT: 190 |  | 6C GXIV-4D | 838412 | None | None | -Preventive Maintenance -Archive tests from September 2023 to recent | All modules are functional. | CPU and Monitor packed up. Unit is daisy-chained with 10C unit. |  |  |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5861193148837105282/5861193148837105282_signature_28.png | Oscar Delano R. Salazar, IV | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5861193148837105282/5861193148837105282_signature_31.png
| 2024-03-13 | 2024-03-13 10:43 | Dimasalang Health Center | Isagani St. Sta Cruz | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27668 | PBSP | joneevamarie@yahoo.com | 25 |  |  | GXIV | 707072 | Unit for PM. | PM. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861074893074690030/5861074893074690030_signature_28.png | Eva Marie C. Jone | https://www.jotform.com/uploads/macarewecare/222927978372067/5861074893074690030/5861074893074690030_signature_31.png
| 2024-03-13 | 2024-03-13 10:45 | Dimasalang Health Center | Isagani St. Sta Cruz | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27667 | PBSP | joneevamarie@yahoo.com | 25 |  |  | GXIV | 805111 | Unit for PM. | PM. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861073663076698145/5861073663076698145_signature_28.png | Eva Marie C. Jone | https://www.jotform.com/uploads/macarewecare/222927978372067/5861073663076698145/5861073663076698145_signature_31.png
| 2024-03-13 | 2024-03-13 09:29 | Hi-Precision Diagnostics | N. Amotanto St. Quezon City | NCR | NCR 2nd - Quezon City | Molecular laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | pcr@hi-precision.com.ph |  | 27666 | CONSORTIUM | jose@hi-precision.com.ph | 25 |  | 2024-03-07 | GXIV | 835602 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Modules A3 (904172) and A4 (210028982), service required. |  | Xpertcheck kit. | DR#11271 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861043741029961963/5861043741029961963_signature_28.png | Joselito Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5861043741029961963/5861043741029961963_signature_31.png
| 2024-03-13 | 2024-03-13 09:26 | Hi-Precision Diagnostics | N. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27665 | CONSORTIUM | jose@hi-precision.com.ph | 25 |  | 2024-03-07 | GXIV | 835601 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck kit. | DR#11270 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861042911022312077/5861042911022312077_signature_28.png | Joselito Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5861042911022312077/5861042911022312077_signature_31.png
| 2024-03-13 | 2024-03-12 15:49 | Fabella Memorial Hospital | Tayuman St. | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27664 | MACARE-HBDC | fabella.mollabgene@gmail.com | 24 |  |  | GX-IV-D-10C | 110015482 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned icore Cleaned plunger rod Generated reports Archive files | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5861031063615358717/5861031063615358717_signature_28.png | Christianne Bernadette Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5861031063615358717/5861031063615358717_signature_31.png
| 2024-03-13 | 2024-03-12 15:08 | South Cotabato Provincial Hospital | Koronadal City, South Cotabato | 12 | South Cotabato | Laboratory | YES | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27663 | PBSP | anamariermt22@gmail.com | 24.4 |  |  | Gx-IV 10C | 110013917 | Not functional UPS | Installed Gx machine SN:110013917 Installed Dell CPU SN CPH0523F42141 Installed lenovo Monitor V5GAG439 Installed hp printer CNB1RBK7RS | Genexpert machine 10C installation | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860981017611268185/5860981017611268185_signature_28.png | Ana Marie Grace C. Egasan | https://www.jotform.com/uploads/macarewecare/222927978372067/5860981017611268185/5860981017611268185_signature_31.png
| 2024-03-12 | 2024-03-12 20:26 | Gabriela Silang General Hospital | Brgy. Tamag, Vigan, Ilocos Sur | 1 | Ilocos Sur | Molecular | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27662 | PBSP | isphcovidlab21@gmail.com | 24 |  |  | GXIV-D-10C | 120000590 | Unit is due for Installation | Installation | Perfirmed Site Assessment Installed GeneXpert System Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860558530419305147/5860558530419305147_signature_28.png | Shaira Gelline Sorizo | https://www.jotform.com/uploads/macarewecare/222927978372067/5860558530419305147/5860558530419305147_signature_31.png
| 2024-03-12 | 2024-03-12 13:52 | Sta. Cruz Municipal Health Office | Poblacion Weste, Sta.Cruz | 1 | Ilocos Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27659 | PBSP | donna.rsantiago@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 763 |  | GXIV-D-6C | 808622 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger, Icore Checked if Filter is Clogged Performed XpertCheck Generate Reports | All Modules are Operational | Passed in XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860553240414794027/5860553240414794027_signature_28.png | Donna Fe Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5860553240414794027/5860553240414794027_signature_31.png
| 2024-03-12 | 2024-03-12 11:47 | Bacnotan District Hospital | Poblacion, La Union | 1 | La Union | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27660 | WHO | zharinaapolsnolasco@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 118 |  | GXIV-D-6C | 846036 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, Icore Checked if Filter is Clogged Performed XpertCheck Generate Reports | UPS is Defective | All Modules Passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860553240418951528/5860553240418951528_signature_28.png | Zharina Apols Nolasco | https://www.jotform.com/uploads/macarewecare/222927978372067/5860553240418951528/5860553240418951528_signature_31.png
| 2024-03-12 | 2024-03-12 19:09 | Gabriela Silang General Hospital | Brgy Tamag, Vigan | 1 | Ilocos Sur | Molecular | NO | Pull-out | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27661 | PBSP | isphcovidlab21@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1421 ASSAY: Xpert SARS-CoV-2, COUNT: 513 ASSAY: Xpert HIV-1 Viral Load, COUNT: 139 |  | GXIV-D-6C | 807578 | Unit is due for Packed-up | Pull-out | Uninstall GX System in Laboratory Checked Functionality | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860553250416119118/5860553250416119118_signature_28.png | Shaira Gelline Quiaot Sorizo | https://www.jotform.com/uploads/macarewecare/222927978372067/5860553250416119118/5860553250416119118_signature_31.png
| 2024-03-12 | 2024-03-12 16:07 | Bayugan City Health Office | Brgy. Poblacion, Bayugan City | 13 | Agusan del Sur | RTDL | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27658 | PBSP | lourdescayao@gmail.com | 25 |  |  | GX-4 | 815391 | Stuck Cartridge at Module B4 Error in Valve Motion persists after removing of the Cartridge at Module B4 H/W Failed at Module B2 due to Temperature related issues (Thermistor) | Reported to Jotform | Removing of Stuck Cartridge at Module B4 Troubleshooting of Module B2 and B4 | 2 modules are working. 2 modules for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860403755604200483/5860403755604200483_signature_28.png | Lourdes Diola | https://www.jotform.com/uploads/macarewecare/222927978372067/5860403755604200483/5860403755604200483_signature_31.png
| 2024-03-12 | 2024-03-12 11:26 | San Lazaro Hospital | Quiricada St. | NCR | NCR 1st - Manila | SACCL | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | egee_gtz@yahoo.com |  | 27657 | MACARE-HBDC | robethsonfp@gmail.com | 26 |  |  | GX-IV-L-6C | 821257 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned icore Run XpertCheck Generated reports Checked and cleaned printer | All modules operational. All modules passed. | Printer cannot print in black ink. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860389560029554025/5860389560029554025_signature_28.png | Robethson Perez, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5860389560029554025/5860389560029554025_signature_31.png
| 2024-03-12 | 2024-03-12 15:39 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | REPAIR | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27656 | COMMERCIAL | cduhlabresults2@yahoo.com | 23 |  |  | MCS+ | 17H639SPG | MCS+ KEYPAD DOES NOT FUNCTION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860385098029250270/5860385098029250270_signature_28.png | Ericka Jane Corro | https://www.jotform.com/uploads/macarewecare/222927978372067/5860385098029250270/5860385098029250270_signature_31.png
| 2024-03-12 | 2024-03-12 13:49 | Plaridel Rural Health Unit IV | Plaridel, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27655 | DOH | Yellei0908@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 854 | 2024-03-12 | GXIV | 831711 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational A1: Passed A2: Service required A3: Insufficient cartridge A4: Passed |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860336154617090877/5860336154617090877_signature_28.png | Ahlelei G. Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5860336154617090877/5860336154617090877_signature_31.png
| 2024-03-12 | 2024-03-12 14:07 | Chinese General Hospital | Blumentritt, Sta Cruz | NCR | NCR 1st - Manila | Bacteriology | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | biomed@cghmc.com.ph |  | 27654 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 |  | 2024-03-07 | GXIV | 835604 | Unit for PM. Unit for Xpertcheck. | PM.amd Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  | Xpertcheck kit | DR#11268 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860335321028994559/5860335321028994559_signature_28.png | Jerwin E. Hermosura | https://www.jotform.com/uploads/macarewecare/222927978372067/5860335321028994559/5860335321028994559_signature_31.png
| 2024-03-12 | 2024-03-12 13:47 | Cadiz City Health Office | Cadiz City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27653 | PBSP | zhevebacomo@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1617 |  | GX4 | 807614 | Module D2 with SN 808098 failed in XPC today. | Save Files(IQR, System Log, Assay Stats & GXC Files). Conduct also PM(Filter, Optics and Plunger Cleaning). | Performed Xpert Check calibration for 4 modules. | Module D1, D3 & D4 passed in XPC today. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860320809565858250/5860320809565858250_signature_28.png | Zherlyn Eve Bacomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5860320809565858250/5860320809565858250_signature_31.png
| 2024-03-12 | 2024-03-12 13:19 | Lamitan city health office | Lamitan city health office | BARMM | Basilan | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27652 | PBSP | lcholaboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 765 | 2024-03-12 | GX-4 6c | 807460 | System unit cmos battery already drained. | Replaced CMOS battery and conducted preventive Maintenance | Replaced CMOS battery and conducted preventive Maintenance | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860307975622501882/5860307975622501882_signature_28.png | Criselita A. Sojor | https://www.jotform.com/uploads/macarewecare/222927978372067/5860307975622501882/5860307975622501882_signature_31.png
| 2024-03-12 | 2024-03-12 13:02 | Lamitan city health office | Lamitan | BARMM | Basilan | Laboratory | YES | INSTALLATION | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27651 | PBSP | lcholaboratory@gmail.com | 26 |  | 2024-03-12 | GX 4 10c | 110020679 | None | Installation of GX-4 10c Genexpert Machine. | Installation of GX-4 10c Genexpert Machine. | All modules are available. UPS is defective and must be replaced emedietly. | UPS is defective and must be replaced ASAP | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860307975623281021/5860307975623281021_signature_28.png | Criselita A. Sojor | https://www.jotform.com/uploads/macarewecare/222927978372067/5860307975623281021/5860307975623281021_signature_31.png
| 2024-03-12 | 2024-01-30 13:49 | Bantayan RHU | Bantayan Island | 7 | Cebu | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27650 | DOH | mcdelle916bausin@gmail.com | 30 |  |  | Gx4 | 838509 | Defective Monitor | Install monitor for bantayan rhu, clean fan filter | Install monitor for bantayan rhu | Module B2 H/W fail |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824141020013347585/5824141020013347585_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/5824141020013347585/5824141020013347585_signature_31.png
| 2024-03-12 | 2024-03-12 10:16 | Dr. Rafael S. Tumbokon Memorial Hospital | Kalibo, Aklan | 6 | Aklan | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27649 | PBSP | delfinteodia@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1114 ASSAY: Xpert HIV-1 Viral Load, COUNT: 62 | 2024-03-12 | GXIV-4D | 804680 | Module B2 defective | -Replace defective module B2/210068729 with module SN: 896225 | #NAME? | All modules are functional. |  | 900-0491RS GeneXpert Module 6-Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860265929531694025/5860265929531694025_signature_28.png | Teodia P Delfin | https://www.jotform.com/uploads/macarewecare/222927978372067/5860265929531694025/5860265929531694025_signature_31.png
| 2024-03-12 | 2024-03-12 09:51 | Rafael Lazatin Memorial Medical Center | Ageles City Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27648 | PBSP | onabb_lab@yahoo.com.ph | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1519 |  | GXIV-D | 838874 | B1:863513 and B4:799806 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Performed functionality test Gathered reports | Unit is operational B1:701068 and B4: 715586 new modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860237988226565320/5860237988226565320_signature_28.png | Perla Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5860237988226565320/5860237988226565320_signature_31.png
| 2024-03-12 | 2024-03-12 11:13 | Northern Samar Provincial Hospital | Catarman Northern Samar | 8 | Northern Samar | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27647 | PBSP | dinahada70@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2337 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 | 2024-03-12 | GXIV - 6C | 807566 | Module B3/730410 and B4/657532 for replacement | #NAME? | - Install new module on B3 w/ SN: 674541 - Install new module on B4 w/ SN: 890880 - Performed preventive maintenance (Cleaned plunger rod, Cartridge bay, Fan Filter) - Generate reports - Pullout defective modules | unit is operational |  | 2 Genexpert module 6C SN: 674541 / 890880 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860228578857303869/5860228578857303869_signature_28.png | Dinah F. Ada | https://www.jotform.com/uploads/macarewecare/222927978372067/5860228578857303869/5860228578857303869_signature_31.png
| 2024-03-12 | 2024-03-12 10:39 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27646 | COMMERCIAL | covidlabjrrmmc@gmail.com | 25 |  | 2024-03-08 | GX16-10C | 110013856 | Unit for PM. For module replacement of module A4. | PM and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A4 (806481) in replacement to defective A4 (210078094). | DR#11285 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5860207441027513666/5860207441027513666_signature_28.png | Ervin Emmett Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/5860207441027513666/5860207441027513666_signature_31.png
| 2024-03-11 | 2024-03-11 09:51 | Jose Abad Santos Rural Health Unit | Jose Abad Santos | 11 | Davao Occidental | TB-DOTS | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27645 | DOH | alcalend47777@gmail.com | 33.9 |  |  | GX-IV 4d | 838460 | No power on CPU System Unit | Replaced the Defective Cpu System Unit with new System Unit SN: CPH1119F04296 | System Unit Replacement | System Unit was successfully installed but Airconditioning Unit is damaged/no power. Enduser is advised not to run specimen due to high ambient temperature. | Fix the Airconditioning Unit first before running specimen to avoid temperature related errors |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859706370225550765/5859706370225550765_signature_28.png | Dranreb John Alcalen | https://www.jotform.com/uploads/macarewecare/222927978372067/5859706370225550765/5859706370225550765_signature_31.png
| 2024-03-11 | 2024-03-11 15:05 | Chong Hua Hospital- Fuente | Capitol Site, Fuente, Cebu City | 7 | Cebu | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27644 | COMMERCIAL | antoniojrpadernal@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1018 ASSAY: Xpert HBV VL, COUNT: 188 ASSAY: Xpert CT/NG, COUNT: 54 ASSAY: Xpert HCV VL, COUNT: 74 ASSAY: Xpert HIV-1 Viral Load, COUNT: 234 |  | GX4 | 110020777 | - Repeated error 5007- samples processing problem for - Xpert BCR-ABL Ultra G2 assay - Error 2097: Assay-Specific Termination Error #2: 45, 28, 1, 0 | Conduct preventive maintenance for Module 2, Module 3, and Module 4 - Required end-user review sample procedure for Xpert BCR-ABL Ultra G2 - Generate IQR, Sys Log, Assay Statistics, Gxx files | - Conduct preventive maintenance for Module 2, Module 3, and Module 4 - Required end-user review sample procedure for Xpert BCR-ABL Ultra G2 - Generate IQR, Sys Log, Assay Statistics, Gxx files | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859508221113427047/5859508221113427047_signature_28.png | Antonio Padernal Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5859508221113427047/5859508221113427047_signature_31.png
| 2024-03-11 | 2024-03-06 09:10 | MIA Health Center | NAIA Road, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27643 | PBSP | mervheequiogue14@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1629 |  | GXIV-D-6C | 812833 | Module 3 with Temperature Error | Module Replacement | Replaced Module 3 Performed Self-test Generate Reports | 3 Modules are Operational, 1 for Replacement | 768349 -> 718714 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643612308916/5859438643612308916_signature_28.png | Mervhee Sorilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643612308916/5859438643612308916_signature_31.png
| 2024-03-11 | 2024-03-08 10:32 | De La Salle Health Sciences Institute | Special Health Sciences Institute, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angiepabruada@gmail.com |  | 27642 | CONSORTIUM | acpabruada@dlsmhsi.edu.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 169 | 2024-03-05 | GXIV-D-6C | 812380 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, Icore Checked if Filter is Clogged Performed Self-test Performed XpertCheck Generate Reports | 3 Modules Passed, 1 Failed | Failed Mod: 210005363 | XpertCheck Kit | DR 11257 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643615677822/5859438643615677822_signature_28.png | Angelita C. Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643615677822/5859438643615677822_signature_31.png
| 2024-03-11 | 2024-03-08 13:09 | Doña Marta Health Center | Don Carlos, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | Relocation | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27639 | DOH | tessa_0503@yahoo.com.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 881 |  | GXIV-D-6C | 831700 | Request for Relocation of GX System | Relocation | Relocate GX System to 2nd floor Laboratory Checked Outlet Voltage Checked Room if with A/C Generate Reports | Room is Good for GX Installation All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643611565890/5859438643611565890_signature_28.png | Tessa Rose Tabugo | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643611565890/5859438643611565890_signature_31.png
| 2024-03-11 | 2024-03-07 08:38 | Imus City Health Office III | Greengate Subd, Imus, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27641 | PBSP | tess19_chua@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1500 |  | GXIV-D-6C | 838166 | Module 3 with Temperature Error | Module Replacement | Replaced Module 3 Performed Self-test Generate Reports | All Modules are Operational | 702673-> 713513 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643612730530/5859438643612730530_signature_28.png | Marites M. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643612730530/5859438643612730530_signature_31.png
| 2024-03-11 | 2024-03-08 10:43 | De La Salle Health Sciences Institute | Special Health and Sciences Institute, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | angiepabruada@gmail.com |  | 27640 | CONSORTIUM | acpabruada@dlsmhsi.edu.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2411 | 2024-03-05 | GX16-D-6C | 842166 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, Icore Checked if Filter is Clogged Performed Self-test Performed XpertCheck Generate Reports | 15 Modules Passed, 1 Failed | Failed Mod: 880257 | XpertCheck Kit | DR 11258 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643614142741/5859438643614142741_signature_28.png | Angelita C. Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5859438643614142741/5859438643614142741_signature_31.png
| 2024-03-11 | 2024-03-08 09:02 | Kabacan Rural Health Unit | Kabacan, North Cotabato | 12 | Cotabato | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27638 | PBSP | yayo_kmsi@yahoo.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 631 ASSAY: Xpert MTB/RIF Ultra, COUNT: 633 | 2024-03-08 | Gx-IV 2modules | 835182 | Module 2 service required | Perform Genexpert Xpertcheck Genexpert Cleaning and checking of errors Data collection | Perform Genexpert machine preventive maintenance Perform Genexpert Xpertcheck/calibration | 2 operational modules | Module 1 passed and Module 2 for service required | Xpertcheck cartridge CE-5 | Lot #: 1000725070 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859389879423246598/5859389879423246598_signature_28.png | Joan A Sangines | https://www.jotform.com/uploads/macarewecare/222927978372067/5859389879423246598/5859389879423246598_signature_31.png
| 2024-03-11 | 2024-03-08 13:33 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27637 | PBSP | benguetphortdl@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2172 | 2024-03-08 | GXIV-4-D-6C | 807621 | #NAME? | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 656905 >>> 717616 - Generated Reports | #NAME? | Unit is operational |  | GeneXpert Module 717616 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859349115471744000/5859349115471744000_signature_28.png | Cirila O. Sabiano | https://www.jotform.com/uploads/macarewecare/222927978372067/5859349115471744000/5859349115471744000_signature_31.png
| 2024-03-11 | 2024-03-08 09:42 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE Temporary Relocation | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27636 | PBSP | keithtagudar@gmail.com | 21 | ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 625 ASSAY: Others, COUNT: 24 |  | GXIV-4-D-10C | 120000171 | #NAME? | #NAME? | #NAME? | Unit is Operational | For Temporary Relocation to Baguio Health Services Office |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859347745475261067/5859347745475261067_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5859347745475261067/5859347745475261067_signature_31.png
| 2024-03-11 | 2024-03-08 09:38 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27635 | PBSP | keithtagudar@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 423 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1670 | 2024-03-08 | GXIV-4-D-6C | 834685 | #NAME? | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 786825 >>> 718969 - Generated Reports | #NAME? | Unit is operational |  | GeneXpert Module 718969 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859345155472586446/5859345155472586446_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5859345155472586446/5859345155472586446_signature_31.png
| 2024-03-11 | 2024-03-06 11:07 | Rosario District Hospital | Rosario, La Union | 1 | La Union | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27634 | PBSP | rosariodhlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 551 | 2024-03-06 | GXIV-4-D-6C | 815396 | #NAME? | #NAME? | #NAME? | #NAME? | All Modules passed | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859343765479008900/5859343765479008900_signature_28.png | Juanita Orine | https://www.jotform.com/uploads/macarewecare/222927978372067/5859343765479008900/5859343765479008900_signature_31.png
| 2024-03-11 | 2024-03-06 09:40 | Sison Rural Health Unit | Sison, Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27633 | PBSP | sisonrhu@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 286 | 2024-03-06 | GXIV-4-D-6C | 834684 | #NAME? | #NAME? | #NAME? | Unit is Operational | Module B3 failed XpertCheck. Remaining Modules passed XpertCheck | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859337415478020306/5859337415478020306_signature_28.png | Jeannette Alcaide | https://www.jotform.com/uploads/macarewecare/222927978372067/5859337415478020306/5859337415478020306_signature_31.png
| 2024-03-11 | 2024-03-05 08:26 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27632 | PBSP | thisisformyfamily@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1491 | 2024-02-05 | GXIV-4-D-6C | 838286 | #NAME? | #NAME? | #NAME? | Unit is operational | Did not perform XpertCheck on Newly replaced Module B3. All remaining Modules passed XpertCheck | XpertCheck Kits | DR 11186 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859334824753710589/5859334824753710589_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5859334824753710589/5859334824753710589_signature_31.png
| 2024-03-11 | 2024-03-07 12:08 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27631 | DOH | gemmponce11@gmail.com | 24 |  | 2024-03-05 | GX-IV-D-6C | 821292 | CPU does not turn on | Installed backup CPU | Replaced CPU Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports Installed printer driver Import assay | Unit is operational. |  | Old CPU SN: 2UA8471VWM Backup CPU SN: 2UA70329CV | DR# 11264 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859256943611248713/5859256943611248713_signature_28.png | Gemma M. PONCE | https://www.jotform.com/uploads/macarewecare/222927978372067/5859256943611248713/5859256943611248713_signature_31.png
| 2024-03-11 | 2024-03-07 13:45 | Sta. Maria Rural Health Unit | Cailes St., Sta. Maria | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27630 | PBSP | maidauridorado@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 546 | 2024-03-07 | GX-IV-D-6C | 838133 | Module B3 for replacement | Replaced module B3 | Performed maintenance Replaced module B3 Generated reports | All modules operational. |  | Old: B3/806702 New: B3/891643 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859256203613446822/5859256203613446822_signature_28.png | Maida DORADO | https://www.jotform.com/uploads/macarewecare/222927978372067/5859256203613446822/5859256203613446822_signature_31.png
| 2024-03-11 | 2024-03-08 11:02 | Ospital ng Makati | Sampaguita St. | NCR | NCR 4th - Makati | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakpatho_labs@yahoo.com.ph |  | 27629 | MACARE-HBDC | osmakproperty@yahoo.com | 20 |  |  | GX-IV-D-10C | 110023460 | For Installation | Installed the GeneXpert system | Installed the GeneXpert system -UPS -GeneXpert unit -Desktop -Printer Functionality test Generated IQR | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859254583616145113/5859254583616145113_signature_28.png | Kim adrine federizo | https://www.jotform.com/uploads/macarewecare/222927978372067/5859254583616145113/5859254583616145113_signature_31.png
| 2024-03-11 | 2024-03-08 13:41 | Kalinga Health Hub - Marikina | JP Rizal St. | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27628 | CONSORTIUM | pearl.balderama@innovationsch.org | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 870 |  | GX-IV-D-6C | 110008418 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253783613791482/5859253783613791482_signature_28.png | Krisna Morante | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253783613791482/5859253783613791482_signature_31.png
| 2024-03-11 | 2024-03-07 13:45 | Sta. Maria Rural Health Unit | Cailes St., Sta. Maria | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27626 | PBSP | maidauridorado@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 546 | 2024-03-07 | GX-IV-D-6C | 838133 | Module B3 for replacement | Replaced module B3 | Performed maintenance Replaced module B3 Generated reports | All modules operational. |  | Old: B3/806702 New: B3/891643 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253033613098453/5859253033613098453_signature_28.png | Maida DORADO | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253033613098453/5859253033613098453_signature_31.png
| 2024-03-11 | 2024-03-08 13:41 | Kalinga Health Hub - Marikina | JP Rizal St. | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27627 | CONSORTIUM | pearl.balderama@innovationsch.org | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 870 |  | GX-IV-D-6C | 110008418 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253033613557925/5859253033613557925_signature_28.png | Krisna Morante | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253033613557925/5859253033613557925_signature_31.png
| 2024-03-11 | 2024-03-07 12:08 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27625 | DOH | gemmponce11@gmail.com | 24 |  | 2024-03-05 | GX-IV-D-6C | 821292 | CPU does not turn on | Installed backup CPU | Replaced CPU Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Generated reports Installed printer driver Import assay | Unit is operational. |  | Old CPU SN: 2UA8471VWM Backup CPU SN: 2UA70329CV | DR# 11264 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253033614858096/5859253033614858096_signature_28.png | Gemma M. PONCE | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253033614858096/5859253033614858096_signature_31.png
| 2024-03-11 | 2024-03-08 11:02 | Ospital ng Makati | Sampaguita St. | NCR | NCR 4th - Makati | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | osmakpatho_labs@yahoo.com.ph |  | 27624 | MACARE-HBDC | osmakproperty@yahoo.com | 20 |  |  | GX-IV-D-10C | 110023460 | For Installation | Installed the GeneXpert system | Installed the GeneXpert system -UPS -GeneXpert unit -Desktop -Printer Functionality test Generated IQR | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253033615118276/5859253033615118276_signature_28.png | Kim adrine federizo | https://www.jotform.com/uploads/macarewecare/222927978372067/5859253033615118276/5859253033615118276_signature_31.png
| 2024-03-10 | 2024-03-07 15:18 | Bernadette bautista medical clinic | 7 BF Rd QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27623 | CONSORTIUM | richeclevia@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 31 | 2024-03-04 | GX4 D | 835603 | Due for xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | B3 failed xpertcheck B1,B2,B4 passed |  | Xpertcheck kits | 11251 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5858791821719920437/5858791821719920437_signature_28.png | Richard Eclevia | https://www.jotform.com/uploads/macarewecare/222927978372067/5858791821719920437/5858791821719920437_signature_31.png
| 2024-03-10 | 2024-03-08 10:53 | Casimiro Ynares Sr.Memorial Hospital | Becco st., Burgos, Rodriguez Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27622 | DOH | caysmhlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 574 |  | GX4 | 831642 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | A1 encounter intemittent temp error 4012. For reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5858791371714692335/5858791371714692335_signature_28.png | Patricia Mae Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5858791371714692335/5858791371714692335_signature_31.png
| 2024-03-10 | 2024-03-08 14:07 | De los Santos Medical Center | E.rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomedtech@dlsmc.ph |  | 27621 | CONSORTIUM | jdvelarde@dlsmc.ph | 27 |  | 2024-03-04 | GX4 D (10C) | 110020700 | Due for maintenance and xpert check | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance and xpertcheck | All module passed in xpertcheck. |  | Xpertcheck kits | 11253 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5858791021718898349/5858791021718898349_signature_28.png | Jeanela Velarde | https://www.jotform.com/uploads/macarewecare/222927978372067/5858791021718898349/5858791021718898349_signature_31.png
| 2024-03-08 | 2024-03-08 15:49 | BAIS DISTRICT HOSPITAL | Bais, Negros Oriental | 7 | Negros Oriental | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27620 | PBSP | laboratorybdh@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 745 |  | GX-IV | 838857 | END USER REQUEST FOR PREVENTIVE MAINTENANCE AND CALIBRATION | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM XPERTCHECKC CALIBRATION: MODULE B1, B2, AND B4 PASSED MODULE B3 SN 878813 FOR SERVICE | EQUIPMENT OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856937740044798059/5856937740044798059_signature_28.png | Evanjun Paccurig | https://www.jotform.com/uploads/macarewecare/222927978372067/5856937740044798059/5856937740044798059_signature_31.png
| 2024-03-08 | 2024-03-08 15:33 | Corazon Locsin Montelibano M.R.H | Lacson St., Bacolod City | 6 | Negros Occidental | Immuno/Sero Dept. | NO | MAINTENANCE | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27619 | MACARE-HBDC | lorenzrhudy@gmail.com | 4 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 632 |  | GX4 | 110020743 | None | Save Files(IQR, System Log & Assay Stats). | Conduct Preventive Maintenance( Filter, Bay Area & Optics Cleaning) for GeneXpert Machine 10 C. | All modules are in good working conditions. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856927447215365051/5856927447215365051_signature_28.png | LORENZ RUDY | https://www.jotform.com/uploads/macarewecare/222927978372067/5856927447215365051/5856927447215365051_signature_31.png
| 2024-03-08 | 2024-03-06 13:19 | Morong Rural Health Unit | Morong Bataan | 3 | Bataan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27618 | DOH | lielyn.lim@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV-D | 831607 | CPU not booting up ( Error 3FO) | Troubleshooting | Performed BIOS battery replacement Performed hard drive checkick Performed RAM and connection reset | Unit is not operational CPU for replacement error 3F0 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856921104247901796/5856921104247901796_signature_28.png | Lielyn M. Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5856921104247901796/5856921104247901796_signature_31.png
| 2024-03-08 | 2024-03-08 15:00 | Aurora Memorial Hosital | Baler-Casiguran Rd. Brgy Reserva Baler Aurora | 3 | Aurora | Molecular Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27617 | PBSP | raymondrmt2010@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV-D | 807565 | Installation of machine B1:661317 B2:636286 | Installation Preventive maintenance Module repair | Performed installation and setup of 6C machine Performed module replacement Performed enduser training Gathered reports Performed functionality test | Unit is operational B1:702999 new module B2:701072 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856920744241384346/5856920744241384346_signature_28.png | Raymond DL. Alviar | https://www.jotform.com/uploads/macarewecare/222927978372067/5856920744241384346/5856920744241384346_signature_31.png
| 2024-03-08 | 2024-03-08 14:29 | Medical Center Manila | 850 United Nations Avenue, Brgy. 674 Zone 73 Dist. V Paco Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Rhea Anne Cañelas |  | rheaannecanelas65@gmail.com |  |  | 27616 | CONSORTIUM | fbvelarde.manilamed@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1054 | 2024-03-06 | GXIV-4D-6C | 835540 | Unit due for xpertcheck | Performed xpertcheck and maintenance B1 failed, B2-B4 passed | Collect reports Cleaned filter Cleaned cartridge bay Plunger rod maintenance I-core cleaning | Unit is operational | Work with Cienne Valerio | Xpertcheck-CE-5 | 11266 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856905385018284287/5856905385018284287_signature_28.png | Frances Levina Velarde | https://www.jotform.com/uploads/macarewecare/222927978372067/5856905385018284287/5856905385018284287_signature_31.png
| 2024-03-08 | 2024-03-08 14:50 | Rogaciano M. Mercado Memorial Hospital | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27615 | PBSP | alona0276@gmail.com | 25 |  |  | GXIV | 805252 | Unit for packing | Packed unit | #NAME? | Unit is operational Unit ready for pull out |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856903662029874360/5856903662029874360_signature_28.png | Alona Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5856903662029874360/5856903662029874360_signature_31.png
| 2024-03-08 | 2024-03-08 13:54 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27614 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1948 |  | GXIV-4D | 838561 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856868338475637360/5856868338475637360_signature_28.png | Farimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5856868338475637360/5856868338475637360_signature_31.png
| 2024-03-08 | 2024-03-08 11:08 | Cebu City Health Office | Carreta, Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27613 | PBSP | laureanjo@yahoo.com | 28 |  |  | Gx4 | 804840 | Defective Module B2 sn 210092147 | New Module B2 installed Clean Fan and Filter | New module installed B2 705575 Clean Fan and Filter Gathering data of genexpert machine | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856820971112156951/5856820971112156951_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5856820971112156951/5856820971112156951_signature_31.png
| 2024-03-08 | 2024-03-05 09:44 | Tubod Rural Health Unit | Tubod, lanao del norte | 10 | Lanao del Norte | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27612 | DOH | jasminerusselcuevas@gmail.com | 27 |  |  | Gx 4 | 847251 | Defective module B3 910220 | New module installed B3 784156 Clean fan filter Gathering data of gx machine | New module installed B3 784156 Clean fan filter Gathering data of gx machine | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856818281113473041/5856818281113473041_signature_28.png | Jasmine Russel Cuevas | https://www.jotform.com/uploads/macarewecare/222927978372067/5856818281113473041/5856818281113473041_signature_31.png
| 2024-03-08 | 2024-03-08 11:56 | Dona Remedios Trinidad Health Center | Dona Remedios Trinidad, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27611 | PBSP | maricel.rmt24@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 433 |  | GXIV | 838872 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: Passed B2: Passed B3: Passed B4: Service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856811822023033112/5856811822023033112_signature_28.png | Cecille Angeles | https://www.jotform.com/uploads/macarewecare/222927978372067/5856811822023033112/5856811822023033112_signature_31.png
| 2024-03-08 | 2024-03-08 12:09 | Cagayan Valley Center for Health Development | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27610 | PBSP | mctorio.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1922 ASSAY: Xpert SARS-CoV-2, COUNT: 3962 | 2024-03-08 | Genexpert IV | 808628 | Module C3 Error | Performed Module Replacement C3(686934 - 632425) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Module - 632425 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856805447689617931/5856805447689617931_signature_28.png | Mikeala Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5856805447689617931/5856805447689617931_signature_31.png
| 2024-03-08 | 2024-03-08 10:31 | Rizal Rural Health Unit | Rizal Nueva Ecija | 3 | Nueva Ecija | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27609 | PBSP | abigailestoque@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 569 |  | GXIV-D | 834691 | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Performed xpertcheck Generate reports Archived test results Performed functionality test | Xpertcheck Preventive maintenance | Due for calibration | Unit is operational All modules passed calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856759124813574753/5856759124813574753_signature_28.png | Abigail Estoque | https://www.jotform.com/uploads/macarewecare/222927978372067/5856759124813574753/5856759124813574753_signature_31.png
| 2024-03-08 | 2024-03-08 10:48 | Cagayan Valley Medical Center | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | TB-DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27608 | PBSP | zionclydematote@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3072 ASSAY: Xpert SARS-CoV-2, COUNT: 6296 | 2024-03-08 | Genexpert XVI | 842894 | Module B3 error 4008 | Reset Gatewayboard | Performed Maintenance Gathered Reports Performed Module Replacement A1(888955 - 627228) A3(210061145 - 635995) A4(655356 - 793031) Performed Functionality Test | Unit is Operational | Module B3 for Observation | Genexpert Module = 627228 & 635995 & 793031 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856757520211224213/5856757520211224213_signature_28.png | Sunshine Matote | https://www.jotform.com/uploads/macarewecare/222927978372067/5856757520211224213/5856757520211224213_signature_31.png
| 2024-03-08 | 2024-03-08 09:14 | Cagayan Valley Medical Center | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Serology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27607 | MACARE-HBDC | cvmcserology@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 216 | 2024-03-08 | GXIV-6C | 803870 | Assay problem | Archived Results | Performed Maintenance -Cleaned Filter -Cleaned Icore -Cleaned Cartridge bay -Cleaned Plunger Rod Gathered Reports Performed Functionality Test | Unit is Operational | No Preventive Maintenance Certificate | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856711956444670881/5856711956444670881_signature_28.png | Dorothy Pajarillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5856711956444670881/5856711956444670881_signature_31.png
| 2024-03-08 | 2024-03-08 09:27 | NEGROS ORIENTAL PROVINCIAL HOSPITAL | Dumaguete City | 7 | Negros Oriental | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27606 | PBSP | rezelmarielagos@gmail.com | 25 | ASSAY: Others, COUNT: 0 |  | GX-IV 10C | 110019626 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | #NAME? |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856707863249441042/5856707863249441042_signature_28.png | rezelmarielagos@gmail | https://www.jotform.com/uploads/macarewecare/222927978372067/5856707863249441042/5856707863249441042_signature_31.png
| 2024-03-08 | 2024-03-07 10:40 | FEU NRMF | Regalado ave.Novaliches QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | gqbugayong@feu-nrmf.edu.ph |  | 27605 | CONSORTIUM | clinicallab@feu-nrmf.edu.ph | 27 | ASSAY: Xpert HPV, COUNT: 60 ASSAY: Xpert MTB/RIF Ultra, COUNT: 415 | 2024-03-04 | GX4 D | 840975 | Due for xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | Operational All module passed in xpertcheck. |  | Xpertcheck kits | 11252 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856699620308733141/5856699620308733141_signature_28.png | Ley Arranz | https://www.jotform.com/uploads/macarewecare/222927978372067/5856699620308733141/5856699620308733141_signature_31.png
| 2024-03-07 | 2024-03-07 13:26 | Pikit Rural Health Unit | Pikit, North Cotabato | 12 | Cotabato | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27604 | PBSP | litosepelagior@gmail.com | 27.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: 651 | 2024-03-07 | Gx-IV 2modules (6 color) | 835181 | N/A | Perform Xpertcheck Genexpert Cleaning and checking of errors Data collection | Perform Genexpert machine preventive maintenance Perform Genexpert Xpertcheck/calibration | 2 operational modules | All modules passed Xpertcheck | Xpertcheck cartridge CE-5 | Lot #: 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856297205122617736/5856297205122617736_signature_28.png | Lito R. Sepelagio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5856297205122617736/5856297205122617736_signature_31.png
| 2024-03-07 | 2024-03-07 13:26 | Pikit Rural Health Unit | Pikit, North Cotabato | 12 | Cotabato | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27603 | PBSP | litosepelagior@gmail.com | 27.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: 651 | 2024-03-07 | Gx-IV 2modules (6 color) | 835181 | N/A | Perform Xpertcheck Genexpert Cleaning and checking of errors Data collection | Perform Genexpert machine preventive maintenance Perform Genexpert Xpertcheck/calibration | 2 operational modules | All modules passed Xpertcheck | Xpertcheck cartridge CE-5 | Lot #: 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856296925129987566/5856296925129987566_signature_28.png | Lito R. Sepelagio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5856296925129987566/5856296925129987566_signature_31.png
| 2024-03-07 | 2024-03-07 15:23 | Macare Medicals, Inc. | Quezon City | NCR | NCR 2nd - Quezon City | Macare | NO | REPAIR | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 27602 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  |  | HP Z230 | 2UA52027VVB | Defective HDD | Replaced HDD with SSD Installed new software | Replaced defective HDD | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856056315704391664/5856056315704391664_signature_28.png | Carlos Origenes | https://www.jotform.com/uploads/macarewecare/222927978372067/5856056315704391664/5856056315704391664_signature_31.png
| 2024-03-07 | 2024-03-07 14:24 | Tarlac Provincial Hospital | Hosital Drive San Vicente Tarlac City | 3 | Tarlac | PCR Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27601 | PBSP | molecularlab.tph@gmail.com | 28 | ASSAY: Xpert SARS-CoV-2, COUNT: 2747 |  | GXIV-D | 838142 | Error 2120 Lost Communication | Troubleshooting Preventive maintenance | Performed Preventive maintenance Troubleshoot UPS if Backing up when power is interupted Checked power connection Check ad reset LAN cable and gateway board Performed reliability test Performed functionality test Gthered reports | Unit is operational Under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856044885716135050/5856044885716135050_signature_28.png | Mark Rahael Turaray | https://www.jotform.com/uploads/macarewecare/222927978372067/5856044885716135050/5856044885716135050_signature_31.png
| 2024-03-07 | 2024-03-07 14:08 | Davao Reproductive Health and Wellness Center | Jacinto Street Davao City | 11 | Davao del Sur | TB-HIV | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27600 | DOH | rhwclab@gmail.com | 26.5 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 117 ASSAY: Xpert HIV-1 Viral Load, COUNT: 680 |  | GX-IV 4D | 838262 | Multiple Errors on Module B2 and B4 GeneXpert Relocation and to be Networked | Reset connection for B2 Module, and reset the valve position to its home position Relocate the GeneXpert for networked connection | Module Troubleshooting and Relocation for network of GeneXpert | Module B2 still occuring Error 4004: related to Valve position Error, Module B4 occurring temperature related even after XPC. Errors documented to be sent at Cepheid for further Assessment. Successfully Networked the GeneXpert Machine | Temporarily disabled Modules B2 and B4 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856027298825146030/5856027298825146030_signature_28.png | Elaine Malano RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5856027298825146030/5856027298825146030_signature_31.png
| 2024-03-07 | 2024-03-07 13:54 | Pandan Rural Health Unit | Pandan | 5 | Catanduanes | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27599 | PBSP | mjnamata10@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1186 |  | GX4-6C | 838863 | For Preventive Maintenance Test results not archived | Cleaned cartridge bays, plunger rod, icore Archived test results Generate reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856003083024167603/5856003083024167603_signature_28.png | Mary Joan Namata | https://www.jotform.com/uploads/macarewecare/222927978372067/5856003083024167603/5856003083024167603_signature_31.png
| 2024-03-07 | 2024-03-07 13:41 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb laboratory | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27598 | DOH | chachaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1800 |  | GX4 | 838283 | All modules encountered termination error due to syringe pressure exceeds at protocol limits | - End-user advice to properly keep the reagent buffer at 20-28 degrees including cartridges - Room temperature during weekends rise at approximately 40 degrees and end-user advice to store the cartridges at required temperature from 20-28 degrees to maintain. - buffer turns yellowish during their operation | #NAME? | All units of modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5856001851112194057/5856001851112194057_signature_28.png | Lucille Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5856001851112194057/5856001851112194057_signature_31.png
| 2024-03-07 | 2024-03-04 16:53 | Estevez Memorial Hospital, Inc. | Legazpi | 5 | Albay | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | reylovendino83@gmail.com |  | 27597 | COMMERCIAL | melosept22@Gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 36 ASSAY: Xpert CT/NG, COUNT: 10 ASSAY: Xpert HBV VL, COUNT: 235 ASSAY: Xpert MTB/RIF Ultra, COUNT: 231 ASSAY: Others, COUNT: 31 ASSAY: Xpert SARS-CoV-2, COUNT: 12 |  | GX4-10C | 110017942 | Error/Code: 2008 for B2/220009328 & B4/210083433 | Cleaned icore, carridge bays Check gx cables and power cables Collect Gxx file Generate reports | Performed troubleshooting Collect Data | Unit is Operational | will request assistance from tech specialist for 2008 error if problem persists (insufficient error) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855998243028262984/5855998243028262984_signature_28.png | Estevez Memorial Hoeputwl,Inc. | https://www.jotform.com/uploads/macarewecare/222927978372067/5855998243028262984/5855998243028262984_signature_31.png
| 2024-03-07 | 2024-03-07 13:34 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microobiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27595 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 31 | ASSAY: Xpert HPV, COUNT: 96 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3495 ASSAY: Xpert HIV-1 Viral Load, COUNT: 340 |  | GXIV-4D | 820206 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855993765522946365/5855993765522946365_signature_28.png | Sarah Nicole Aurea Mireno | https://www.jotform.com/uploads/macarewecare/222927978372067/5855993765522946365/5855993765522946365_signature_31.png
| 2024-03-07 | 2024-03-07 13:27 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27596 | DOH | vsmmcdepatho.microbiology@gmail.com | 32 | ASSAY: Xpert HPV, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3168 ASSAY: Xpert HIV-1 Viral Load, COUNT: 105 |  | GXIV- 4D | 838285 | None | None | #NAME? | All modules are functional but error 2097 keeps occurring. | Modules under observation. Error 2097 may be cause my insufficient sample or defective cartridge. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855993765523120821/5855993765523120821_signature_28.png | Sarah Nicole Aurea Merino | https://www.jotform.com/uploads/macarewecare/222927978372067/5855993765523120821/5855993765523120821_signature_31.png
| 2024-03-07 | 2024-03-07 13:32 | Minglanilla Rural Health Unit | Pider 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27594 | DOH | brendagonzagaparedes@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 822 |  | GX4 | 838512 | Module 1 encountered a termination error due to temperature related problem | Fix the temperature of the area, at 20 degrees. Generate IQR, GXX, Assay Statistics Run sample specimen at Module 1 Clean Filter, clean Icore, and plunger | Fix the temperature of the area, at 20 degrees. Generate IQR, GXX, Assay Statistics Run sample specimen at Module 1 Clean Filter, clean Icore, and plunger | As of March 6, 2024, Module 1 run at normal condition and generate proper results without termination error, - All modules are operation -Under Observation within the week, -Medtech end-user advise to properly close the exit area like doors, window -If aircon are malfunction, advise to stop the process and fix the aircon immediately. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855992611114866272/5855992611114866272_signature_28.png | Brenda Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/5855992611114866272/5855992611114866272_signature_31.png
| 2024-03-07 | 2024-03-07 13:29 | Don Carlos Rural Health Unit | Brgy. Poblacion, Don Carlos | 10 | Bukidnon | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27593 | PBSP | gibsz.l002@gmail.com | 34 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 358 |  | GX-4 | 807625 | Module C2 found service required thru XPC | Module Replacement | Replacing Module C2 with SN:675868 to SN: 885298 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855990020321551816/5855990020321551816_signature_28.png | Gilbert E. Loro | https://www.jotform.com/uploads/macarewecare/222927978372067/5855990020321551816/5855990020321551816_signature_31.png
| 2024-03-07 | 2024-03-07 11:53 | Schistosomiasis Research and Training Center | Palo Leyte | 8 | Leyte | TB-DOTS | YES | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27592 | PBSP | mbesplago.pbsp@gmail.com | 25 |  | 2024-03-07 | GXIV - 10C | 110020683 | Unit for Installation | #NAME? | #NAME? | unit is operational |  | #NAME? | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855940216115334362/5855940216115334362_signature_28.png | Maria Chrissa E. Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5855940216115334362/5855940216115334362_signature_31.png
| 2024-03-07 | 2024-03-07 12:01 | Amma Jadsac District Hospital | Pudtol, Apayao | CAR | Apayao | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27591 | PBSP | dlgblacar.pbsp@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3419 | 2024-03-07 | GXIV-6C | 807609 | Module 2 error | Performed Module Replacement on Module A2 A2(210091324 - 861019) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Module - 861019 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855937495517789772/5855937495517789772_signature_28.png | Daisy Lynn Grace B. Lacar | https://www.jotform.com/uploads/macarewecare/222927978372067/5855937495517789772/5855937495517789772_signature_31.png
| 2024-03-07 | 2024-03-07 10:44 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | synch28@yahoo.com |  | 27590 | MACARE-HBDC | msff-tondo-labtech@paris.msf.org | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5296 | 2024-02-14 | GXIV | 110016343 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Modules A4 (210076794), B2 (210083060), B4 (210083059), C2 (210081539) and C3 (210070572) are Service Required. |  | Xpertcheck kit. | DR#11218 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855904209981604378/5855904209981604378_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/5855904209981604378/5855904209981604378_signature_31.png
| 2024-03-07 | 2024-03-07 09:43 | Midsayap Rural Health Unit | Midsayap, North Cotabato | 12 | Cotabato | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27589 | PBSP | rhumidsayap@yahoo.com | 28.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1122 | 2024-03-07 | Gx-IV/2 module 6 color | 835183 | Module 2 under observation due to temperature errors | Perform Genexpert annual Xpertcheck Genexpert Cleaning and checking of errors Data collection | Perform Genexpert machine preventive maintenance Perform Genexpert Xpertcheck/Calibration | 2 operational modules | All modules passed Xpertcheck | Xpertcheck cartridge CE-5 | Lot #: 1001075092 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855895929417699228/5855895929417699228_signature_28.png | Milcah sugcawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5855895929417699228/5855895929417699228_signature_31.png
| 2024-03-07 | 2024-03-07 09:00 | Dr. Amado Diaz Provincial Foundation Hospital | Midsayap, North Cotabato | 12 | Cotabato | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27588 | PBSP | dadphf@yahoo.com | 27.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1154 |  | Gx-lV | 807623 | Module 2 multiple temperature errors | Perform Xpertcheck at module 2 for module troubleshooting Genexpert cleaning and Checking of errors Reporting of errors Data collection | Perform Genexpert machine preventive maintenance Perform Genexpert module troubleshooting | 3 operation modules | Module 2 for service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855833599411340980/5855833599411340980_signature_28.png | Ermariz Jaducana | https://www.jotform.com/uploads/macarewecare/222927978372067/5855833599411340980/5855833599411340980_signature_31.png
| 2024-03-07 | 2024-03-05 13:05 | Batasan Hills Super Health Center | IBP RD Batasan QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27587 | PBSP | ashwinkevin@gmail.com | 40 | ASSAY: Xpert HPV, COUNT: 467 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2992 | 2024-03-05 | GX4 D | 805254 | For modeule replacement D2 640940 -700788, D3 210019124 - 716971 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | Genexpert module | G6855 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855695061714759716/5855695061714759716_signature_28.png | Kevin Orlando D. Cabuso | https://www.jotform.com/uploads/macarewecare/222927978372067/5855695061714759716/5855695061714759716_signature_31.png
| 2024-03-07 | 2024-03-05 10:35 | Lung Center of the Philippines | Quezon avenue,QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27586 | PBSP | lcpmicrolab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5830 ASSAY: Xpert SARS-CoV-2, COUNT: 2536 |  | GX16 L | 801080 | C2 hw failed failure A1 and A2 broken uniframe | Swap module C2 to C1 after swapping the error hw failed still occur Checked voltage of ups Checked gateway board | Trounleshooting | C2 hw failed, 13 operational A1 and A2 still can be use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855694211719944881/5855694211719944881_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5855694211719944881/5855694211719944881_signature_31.png
| 2024-03-06 | 2024-03-06 14:17 | Sasa Health Center | Sasa, Davao City | 11 | Davao del Sur | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27585 | PBSP | mariarosario_mapayo@yahoo.com.sg | 26.8 |  |  | GX-IV 4D | 838889 | For Calibration | Run XpertCheck Calibration on Four Module | Perform Xpertcheck Calibration | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855257575512709289/5855257575512709289_signature_28.png | Maria Rosario M. Pino RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5855257575512709289/5855257575512709289_signature_31.png
| 2024-03-06 | 2024-03-06 16:01 | Ipil rural health unit | Ipil, Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27584 | PBSP | rheasolana@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1688 ASSAY: Xpert HIV-1 Viral Load, COUNT: 29 | 2023-03-06 | GX 4 | 808615 | Module 2 failed during xpertcheck | Replaced module 637066 with Module 607974. And Preventive maintenance. | Replaced module 637066 with Module 607974. And Preventive maintenance. | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855234722346816830/5855234722346816830_signature_28.png | Rheamay E. Solana | https://www.jotform.com/uploads/macarewecare/222927978372067/5855234722346816830/5855234722346816830_signature_31.png
| 2024-03-06 | 2024-03-06 15:50 | INTEGRATED PROVINCIAL HEALTH OFFICE - SIQUIJOR | DUMANHOG, SIQUIJOR | 7 | Siquijor | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27583 | PBSP | glinniericamor73@gmail.com | 27 | ASSAY: Others, COUNT: 0 |  | GX-IV 10C | 120000724 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | UPS FOR OBSERVATION DUE TO ALARM |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855210426819604547/5855210426819604547_signature_28.png | GLINNIE RICAMOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5855210426819604547/5855210426819604547_signature_31.png
| 2024-03-06 | 2024-03-06 15:36 | Malvar Rural Health Unit | Malvar | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27582 | PBSP | lopezsheryl13@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 691 | 2024-03-06 | GX-IV-D-6C | 838191 | Module B4 for replacement | Replaced module B4 | Performed maintenance Replaced module B4 Generated reports | All modules operational. |  | Old: B4/806888 New: B4/893800 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855204395268168497/5855204395268168497_signature_28.png | Sheryl E. Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5855204395268168497/5855204395268168497_signature_31.png
| 2024-03-06 | 2024-03-06 15:06 | Manila Central University | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | lab_mcuhospital@yahoo.com |  | 27581 | CONSORTIUM | purchasing@mcufdtmf.org | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 285 | 2024-03-04 | GXIV | 110018008 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: Passed B2: Service required B3: Passed B4: Passed Module B2 for reporting |  | XpertCheck Kit Lot: 1000410214 | 11250 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855199776542319333/5855199776542319333_signature_28.png | Marcjoven Angelo Superales | https://www.jotform.com/uploads/macarewecare/222927978372067/5855199776542319333/5855199776542319333_signature_31.png
| 2024-03-06 | 2024-03-06 15:09 | Carmen Rural Health Unit | Carmen, North Cotabato | 12 | Cotabato | Tb-Lab | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27580 | PBSP | lumawagrosalyn1@gmail.com | 26.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 573 | 2024-03-06 | Gx-IV 6 color | 815284 | Module 3 H/W failed Request for genexpert fan filter | Replace defective module 3 SN:695711 with new module SN:862382 Perform Xpertcheck at module 1,2 and 4 Genexpert Cleaning and checking of errors Data collection | Perform Genexpert machine preventive maintenance Perform Genexpert machine module replacement/repair Perform Genexpert Xpertcheck/Calibration | 4 operational modules | Module 1,2 and 4 passed Xpertcheck | Genexpert module 6-color SN:695711 | Cepheid case:1467083 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855194121517866022/5855194121517866022_signature_28.png | Rosalyn L. Lumawag | https://www.jotform.com/uploads/macarewecare/222927978372067/5855194121517866022/5855194121517866022_signature_31.png
| 2024-03-06 | 2024-03-06 14:14 | Iloilo Provincial Hospital | Pototan, Iloilo | 6 | Negros Occidental | TB DOTS | NO | CALIBRATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27579 | DOH | djpragale92@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 181 |  | GX4 | 838289 | Module C2 with SN 808243 and module C3 with SN 740037 failed in Xpert check today. | Conduct PM and save files(IQR, System Log, Assay Stats and GXC File). | Performed Xpert check calibration for 4 modules. | Module C1 and C4 passed the XPC and in good working conditions. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855155806222300477/5855155806222300477_signature_28.png | Rachel Pragale | https://www.jotform.com/uploads/macarewecare/222927978372067/5855155806222300477/5855155806222300477_signature_31.png
| 2024-03-06 | 2024-03-06 14:01 | Mataas na Kahoy Rural Health Unit | Brgy. Dos, Dimayuga St. Mataas na Kahoy | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27578 | PBSP | dmlescano1219@gmail.com | 25 |  |  | GXIV | 838325 | Unit for PM. Unit for Xpertcheck. | PM and Xpercheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855155742511863484/5855155742511863484_signature_28.png | Debbie Lescano | https://www.jotform.com/uploads/macarewecare/222927978372067/5855155742511863484/5855155742511863484_signature_31.png
| 2024-03-06 | 2024-03-06 12:41 | Batangas Medical Center | Batangas City | 4A | Batangas | Molecular laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27577 | PBSP | batmc.tbculturelab@gmail.com | 25 |  | 2024-03-04 | GX16 | 844120 | Unit for PM. For module replacement of A2, C3, C4 and D4. | PM and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX modules A2 (862414), C3 (638691), C4 (717781) and D4 (886347) in replacement to defective A2 (607813), C3 (892213), C4 (892212) and D4 (680439). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855117852512685625/5855117852512685625_signature_28.png | Patricia Louisse A. Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5855117852512685625/5855117852512685625_signature_31.png
| 2024-03-06 | 2024-03-06 12:46 | Batangas Medical Center | Batangas City | 4A | Batangas | Molecular laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27576 | PBSP | batmc.tbculturelab@gmail.com | 25 |  | 2024-03-04 | GXIV | 814153 | Unit for PM. For module replacement of C1. | PM and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module C1 (717610) in replacement to defective C1 (619967). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855117852514643912/5855117852514643912_signature_28.png | Patricia Louisse A. Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5855117852514643912/5855117852514643912_signature_31.png
| 2024-03-06 | 2024-03-06 12:52 | Buhangin Health Center | Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27575 | DOH | buhanginlaboratory@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 909 |  | GX-IV 4D | 838232 | Due for Preventive Maintenance and UPS Problem | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log Archived Data test results Check UPS | Perform Preventive Maintenance UPS Troubleshooting | All Modules are Operational UPS no power on battery mode, documented to be sent for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855106912917894943/5855106912917894943_signature_28.png | Catherine De Ocampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5855106912917894943/5855106912917894943_signature_31.png
| 2024-03-06 | 2024-03-06 12:35 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27574 | PBSP | djnrmhsmolecularlab@gmail.com | 25 |  | 2024-03-06 | GXIV | 814159 | Modules B1, B2, B3 for replacement | #NAME? | #NAME? | Unit is operational Module B3's status is hardware failed upon installation B3 for reporting |  | GeneXpert Module SN: 700319, 717214, 713929 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855098798817209324/5855098798817209324_signature_28.png | Noreen Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/5855098798817209324/5855098798817209324_signature_31.png
| 2024-03-06 | 2024-03-06 11:29 | Don Jose S. Monfort Medical Center | Barotac Nuevo, Iloilo | 6 | Iloilo | TB DOTS | NO | CALIBRATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27573 | DOH | shirleymaepecto@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1926 |  | GX4 | 838501 | Module B1 with SN 760020 failed in XPC today. | Save files(GXX File, IQR, System Log and Assay Stat). Conduct also PM and Jotf | Performed Xpert check calibration for 4 modules. | Module B2, B3 and B4 passed the calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855063106229260466/5855063106229260466_signature_28.png | Shirley Mae G. Pecto | https://www.jotform.com/uploads/macarewecare/222927978372067/5855063106229260466/5855063106229260466_signature_31.png
| 2024-03-06 | 2024-03-06 10:27 | Matalom Community Hospital | Matalom Leyte | 8 | Leyte | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27572 | PBSP | fegorne@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 460 | 2024-03-06 | GXIV 6 color | 847227 | stuck cartridge on B2,B3,B4 | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855016442542844652/5855016442542844652_signature_28.png | FE D. GORNE | https://www.jotform.com/uploads/macarewecare/222927978372067/5855016442542844652/5855016442542844652_signature_31.png
| 2024-03-06 | 2024-03-06 08:39 | Bataan Genaral Hospital and Medical Center | Balang Bataan | 3 | Bataan | PCR Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27571 | PBSP | honeyeichel@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 17 ASSAY: Xpert SARS-CoV-2, COUNT: 11512 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GXIV-D | 834433 | Due for calibration | Preventive maintenance Xpertcheck | Performed cleaning of cartridge bay, plunger rod, filter, icore and surface area Archived test results Performed machine calibration Gathered reports Performed functionality test | Unit is operational 2 modules passed calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5855011049619654979/5855011049619654979_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5855011049619654979/5855011049619654979_signature_31.png
| 2024-03-05 | 2024-03-05 19:41 | Jose B. Lingad Memorial and General Hospital | San fernado papanga | 3 | Pampanga | Immunology/ Serology | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27570 | MACARE-HBDC | jblmgh.is@gmail.com | 25 |  |  | GCIV-D | 812571 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Generate reports Archived test results Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854482783228406305/5854482783228406305_signature_28.png | Michelle C. Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5854482783228406305/5854482783228406305_signature_31.png
| 2024-03-05 | 2024-03-05 16:30 | Jose B. Lingad Memorial and General Hospital | San Fernando Pampanga | 3 | Pampanga | Immunology/Serology | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27569 | MACARE-HBDC | jblmgh.is@gmail.com | 31 | ASSAY: Xpert HBV VL, COUNT: 46 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 315 |  | GXIV-D 10C | 110013748 | Due for preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Archived test results Gathered reports Performed functionality test | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854383743227134025/5854383743227134025_signature_28.png | Michelle C. Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5854383743227134025/5854383743227134025_signature_31.png
| 2024-03-05 | 2024-03-05 15:47 | Baras Medicare and Community Hospital | Bagong Sirang, Baras | 5 | Catanduanes | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27568 | PBSP | malustesorerormt@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 709 |  | GX4-10C | 835215 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854353817119532404/5854353817119532404_signature_28.png | Malu S. Tesorero | https://www.jotform.com/uploads/macarewecare/222927978372067/5854353817119532404/5854353817119532404_signature_31.png
| 2024-03-05 | 2024-03-05 16:03 | Catarman Municipal Health Office | Brgy. Poblacion, Catarman | 10 | Camiguin | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27567 | PBSP | airishcamacho1225@gmail.com | 28 |  |  | GX-4 | 835063 | Module B2 found service required thru XPC | Module Replacement | Replacing Module B2 with SN: 787848 to SN: 660885 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854352883699139704/5854352883699139704_signature_28.png | Airish Camacho | https://www.jotform.com/uploads/macarewecare/222927978372067/5854352883699139704/5854352883699139704_signature_31.png
| 2024-03-05 | 2024-03-05 15:51 | Baras Medicare and Community Hospital | Bagong Sirang, Baras | 5 | Catanduanes | Laboratory | NO | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27566 | PBSP | malustesorerormt@gmail.com | 22 |  |  | GX4-10C | 110020684 | GX4-10C for Installation | Installed Gx4-10C, CPU, Monitor, UPS, Printer, Barcode Scanner Generate Reports | GX4-10C installed | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854351797114490506/5854351797114490506_signature_28.png | Malu S. Tesorero | https://www.jotform.com/uploads/macarewecare/222927978372067/5854351797114490506/5854351797114490506_signature_31.png
| 2024-03-05 | 2024-03-05 14:38 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27565 | DOH | abigailgsula@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1391 | 2024-02-14 | GXIV-D | 831912 | A2:649409 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Performed functionality test Gathered reports | Unit is operational A2:720724 ( new module) |  | 1 genexpert module(720724) | 11215 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854305195412465024/5854305195412465024_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5854305195412465024/5854305195412465024_signature_31.png
| 2024-03-05 | 2024-03-05 13:27 | Pandi Rural Health Unit | Pandi, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27564 | WHO | pandirhulaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1064 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1064 |  | GXIV | 838883 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854286192521050762/5854286192521050762_signature_28.png | Jhoana Marie Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5854286192521050762/5854286192521050762_signature_31.png
| 2024-03-05 | 2024-03-05 13:47 | Toril Health Center | Toril, Davao City | 11 | Davao del Sur | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27563 | PBSP | leabuelis@yahoo.com | 23.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1705 |  | GX-IV 4d | 838865 | Module B1 service required last Xpc Calibration | Replaced the Defective Module SN: 788487 with new Module SN: 210058025 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854274265911569304/5854274265911569304_signature_28.png | Lea L. Buelis RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5854274265911569304/5854274265911569304_signature_31.png
| 2024-03-05 | 2024-03-04 15:02 | Palimbang Rural Health Unit | Palimbang, Sultan Kudarat | 12 | Sarangani | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27562 | PBSP | tinavillegas1910@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 412 | 2024-03-04 | Gx-IV 6-color | 847257 | Module 1 for Service Required | Replace defective module 1 SN:910135 with new module SN:883410 Genexpert Cleaning and checking of errors Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert module replacement/repair | 4 operational modules |  | Genexpert module 6-color | Cepheid case:1470649/1470759 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854232836644759323/5854232836644759323_signature_28.png | Cristina Nelda Villegas | https://www.jotform.com/uploads/macarewecare/222927978372067/5854232836644759323/5854232836644759323_signature_31.png
| 2024-03-05 | 2024-03-05 12:02 | Mabalacat Rural Health Unit | Mabalacat Pamanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27561 | PBSP | perezriyadh@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1155 ASSAY: Xpert HIV-1 Viral Load, COUNT: 52 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 |  | GXII-D | 835104 | Due for calibration | Xpertcheck | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Archived test results Performed xpertcheck Generate reports Performed functionality test | Unit is operational except B2 B1 passed calibration B2 Failed calibration | B2:715593 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854224185417207043/5854224185417207043_signature_28.png | Riyadh Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5854224185417207043/5854224185417207043_signature_31.png
| 2024-03-05 | 2024-03-05 12:21 | Mabalacat Rural Health Unit | Mabalacat Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27560 | PBSP | perezriyadh@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 314 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GCIV-D | 807564 | Due for calibration Error 2034( unable to performed xpertcheck) | Xpertcheck Preventive maintenance Troubleshooting | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Installation of xpertcheck application Performed troubleshooting Generate reports Archived test results Performed functionality test | Unit is operational | Error 2034 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854223695412892519/5854223695412892519_signature_28.png | Riyadh Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5854223695412892519/5854223695412892519_signature_31.png
| 2024-03-05 | 2024-03-05 10:57 | Maitum Rural Health Unit | Maitum, Sarangani | 12 | Sarangani | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27559 | PBSP | mikexxxballesteros@gmail.com | 35.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1030 |  | Gx-IV (2 modules) 6-color | 835076 | Module 2 H/W failed due to temperature errors Genexpert module temperature not within the acceptable range of 15°C to 30°C Modules average temperature - 35.2°C | Perform Genexpert Xpertcheck/calibration Genexpert Cleaning and checking of errors Reporting of problem Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert Xpertcheck/Calibration | 1 operational modules | Need to install/replace air conditioning unit for genexpert required temperature |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854221616641677205/5854221616641677205_signature_28.png | Michael F. Ballesteros | https://www.jotform.com/uploads/macarewecare/222927978372067/5854221616641677205/5854221616641677205_signature_31.png
| 2024-03-05 | 2024-03-05 09:49 | Macare In-House | 1005 Aurora Blvd. Proj 3, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | OTHERS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  | GXIV CPU | 27558 | DOH | plosbanos@macare-medicals.com | 24 |  |  | HP Z240 | 2UA70329CV | Defective HDD | Replaced HDD | Replaced HDD Installed GX software Test the CPU with GX machine Setup system configuration | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854133963619854921/5854133963619854921_signature_28.png | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5854133963619854921/5854133963619854921_signature_31.png
| 2024-03-05 | 2024-03-05 09:33 | Magalang Rural Health Unit | Sa Nicolas 1 Magalang Pampanga | 3 | Pampanga | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27557 | WHO | bevlozano29@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 |  | GXIV-D | 846044 | Due for calibration | Xpertcheck | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Archived test results Installation of xpertcheck application Performed machine calibration Generate reports Performed functionality test | Unit is operational All modules passed calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5854126475419264792/5854126475419264792_signature_28.png | Beverly Lozano | https://www.jotform.com/uploads/macarewecare/222927978372067/5854126475419264792/5854126475419264792_signature_31.png
| 2024-03-04 | 2024-03-04 21:04 | Toledo City Health Office | Toledo, Ilihan, CEbu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27556 | PBSP | loriemaediapen@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1529 |  | GX4 | 808640 | Module 2 SN 79154 replacement due to failed last calibration | Module 2 New Module SN# 693979 replacement | Module 2 New Module SN# 693979 replacement | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853670121116362502/5853670121116362502_signature_28.png | Lorelie Mae Diapen, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5853670121116362502/5853670121116362502_signature_31.png
| 2024-03-04 | 2024-03-04 20:57 | Moalboal Rural Health Unit | Poblacion, Moalboal Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27555 | PBSP | rhu_moalboal@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 642 |  | GX4 | 805033 | Module 1 863544 failed during calibration, | Module 1 replaced new module Serial number : 787969 Generate IQR, Sys Log, Assay Statistics, GXX file | Module 1 replaced new module Serial number : 787969 Generate IQR, Sys Log, Assay Statistics, GXX file | All modules are functioning. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853666111119779622/5853666111119779622_signature_28.png | Ava Anaciano Gako | https://www.jotform.com/uploads/macarewecare/222927978372067/5853666111119779622/5853666111119779622_signature_31.png
| 2024-03-04 | 2024-03-04 16:04 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27554 | COMMERCIAL | cduhlabresults2@yahoo.com | 25 |  |  | MCS+ | 14F014 | MCS+ KEYPAD DOES NOT FUNCTION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | FOR QUOTATION |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853495812391651361/5853495812391651361_signature_28.png | Ericka Jane Corro | https://www.jotform.com/uploads/macarewecare/222927978372067/5853495812391651361/5853495812391651361_signature_31.png
| 2024-03-04 | 2024-02-22 08:45 | Jose Reyes Memorial Medical Center | Rizal Ave, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27553 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 398 |  | GX16-D-10C | 110013856 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Filter and Plunger Rod Archived Test Results Performed Self-test Generate Reports | 1 Module for Replacement, 15 Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853412573616518609/5853412573616518609_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5853412573616518609/5853412573616518609_signature_31.png
| 2024-03-04 | 2024-02-21 14:57 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Road, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27552 | COMMERCIAL | uphdmc_lab@yahoo.com | 27 | ASSAY: Xpert CT/NG, COUNT: 8 ASSAY: Xpert HBV VL, COUNT: 34 ASSAY: Xpert HPV, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 361 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2024-02-16 | GXIV-D-6C | 831316 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, Filter and Icore Performed XpertCheck Performed Self-test Generate Reports | 2 Modules Passed, 2 Modules Failed | A1: 784655 and A4: 784666 Failed in XpertCheck | XpertCheck Kit | SI 97180 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853412013615444358/5853412013615444358_signature_28.png | Inger Lia Jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/5853412013615444358/5853412013615444358_signature_31.png
| 2024-03-04 | 2024-01-11 11:36 | Alfonso Rural Health Unit | Burgos St., Alfonso, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27551 | DOH | aim_1225@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 784 | 2024-01-08 | GXIV-D-6C | 821279 | Module 1 with Temperature Related Problem | Module Replacement | Replaced Module 1 Performed Selft-test Performed Maintenance Generate Reports | All Modules are Operational | B1: 738405 -> 756445 | GeneXpert Module 6C | 11062 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853411603616294526/5853411603616294526_signature_28.png | Aim Arroyo | https://www.jotform.com/uploads/macarewecare/222927978372067/5853411603616294526/5853411603616294526_signature_31.png
| 2024-03-04 | 2024-02-22 08:50 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27550 | MACARE-HBDC | covidlabjrrmmc@gamil.com | 24 | ASSAY: Xpert HPV, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 451 ASSAY: Xpert SARS-CoV-2, COUNT: 273 |  | GXIV-D-10C | 110014099 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Filter and Plunger Rod Archived Test Results Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853411123619979777/5853411123619979777_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5853411123619979777/5853411123619979777_signature_31.png
| 2024-03-04 | 2024-02-22 08:54 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta.Cruz, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27549 | DOH | covidlabjrrmmc@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF Ultra, COUNT: 215 ASSAY: Xpert SARS-CoV-2, COUNT: 96 |  | GXIV-D-6C | 831530 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Filter and Plunger Rod Archived Test Results Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853410683614406387/5853410683614406387_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5853410683614406387/5853410683614406387_signature_31.png
| 2024-03-04 | 2024-02-22 11:34 | National Tuberculosis Reference Laboratory | RITM, Filinvest | NCR | NCR 4th - Muntinlupa | MBL | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | eracueza.pbsp@gmail.com |  | 27548 | PBSP | ntrlphil_cs@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 48 ASSAY: Others, COUNT: 1893 |  | GXIV-D-10C | 848244 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Filter, Icore and Plunger Rod Performed XpertCheck Performed Self-test Generate Reports | All Modules are Operational | 3 Pass, 1 New Module |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853409743615771859/5853409743615771859_signature_28.png | Era Denice R. Acueza | https://www.jotform.com/uploads/macarewecare/222927978372067/5853409743615771859/5853409743615771859_signature_31.png
| 2024-03-04 | 2024-02-22 11:45 | National Tuberculosis Reference Laboratory | RITM, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | MBL | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | eracueza.pbsp@gmail.com |  | 27547 | PBSP | ntrlphil_cs@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2411 |  | GXIV-D-6C | 807606 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Filter, Icore and Plunger Rod Performed XpertCheck Performed Self-test Generate Reports | 3 Modules are Operational | B3/662385 Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853408863611226576/5853408863611226576_signature_28.png | Era Denice R. Acueza | https://www.jotform.com/uploads/macarewecare/222927978372067/5853408863611226576/5853408863611226576_signature_31.png
| 2024-03-04 | 2024-02-23 14:10 | Tanza Rural Health Unit | Bawis, Tanza,Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27546 | PBSP | wenarcala@gmail.com | 24 |  |  | GXIV-D-6C | 812835 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Icore, Filter and Plunger Rod Performed Self-test Performed XpertCheck Generate Reports | 3 Modules are Operational, 1 for Replacement | B2 Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853408513619442700/5853408513619442700_signature_28.png | Edwina P. Arcala | https://www.jotform.com/uploads/macarewecare/222927978372067/5853408513619442700/5853408513619442700_signature_31.png
| 2024-03-04 | 2024-02-23 14:30 | Tanza Rural Health Unit | Biwas, Tanza, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27545 | PBSP | wenarcala@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2549 |  | GXIV-D-6C | 838162 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Icore, Filter and Plunger Rod Performed Self-test Performed XpertCheck Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853408033619007999/5853408033619007999_signature_28.png | Edwina P. Arcala | https://www.jotform.com/uploads/macarewecare/222927978372067/5853408033619007999/5853408033619007999_signature_31.png
| 2024-03-04 | 2024-02-21 09:33 | San Lazaro Hospital | Quiricada St, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27544 | PBSP | slhtbclab@gmail.com | 24 |  |  | GXIV-D-10C | 120000475 | Unit is due for Installation | Installation | Installed GeneXpert System Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853407013617579681/5853407013617579681_signature_28.png | Ashley Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5853407013617579681/5853407013617579681_signature_31.png
| 2024-03-04 | 2024-03-04 13:29 | Sto Tomas City Health Office | Sto. Tomas, Batangas | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27543 | PBSP | vickybana_412@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 | 2024-02-19 | GXIV | 821786 | Unit for PM. Unit for Xpertcheck. For module replacement of B2. | PM, Xpertcheck and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Performed module replacement. Generated reports. | Unit is operational. Module B3 (686333), service required. | Actual service date Feb. 21, 2024. | GX module B2 (627677) in replacement to defective B2 (643194). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853398333615126514/5853398333615126514_signature_28.png | Victoria Menor | https://www.jotform.com/uploads/macarewecare/222927978372067/5853398333615126514/5853398333615126514_signature_31.png
| 2024-03-04 | 2024-03-04 11:00 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Mabini, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Molecular Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27542 | PBSP | bryanmichaelbalunes@gmail.com | 24 |  | 2024-03-04 | GXIV-10C | 110020311 | N/a | N/a | Installed Genexpert Installed Drivers Installed PC Installed Monitor Installed Printer Installed UPS Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853313354125573578/5853313354125573578_signature_28.png | Bryan Michael N. Balunes | https://www.jotform.com/uploads/macarewecare/222927978372067/5853313354125573578/5853313354125573578_signature_31.png
| 2024-03-04 | 2024-02-29 13:36 | Indang Rural Health Unit | Alfonso-Indang Rd., Indang, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27540 | PBSP | rhuindang@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1220 |  | GXIV-D-6C | 838137 | Unit is due for XpertCheck | XpertCheck | Cleaned Plunger Rod, Icore, and Cartridge Bay Checked Filter if Clogged Performed XpertCheck Generate Reports | All Modules are Operational | All Modules Passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853244533618726188/5853244533618726188_signature_28.png | Cherry rose digma | https://www.jotform.com/uploads/macarewecare/222927978372067/5853244533618726188/5853244533618726188_signature_31.png
| 2024-03-04 | 2024-03-01 14:28 | Dasmariñas City Health Office IV | Paliparan, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27541 | PBSP | krlanday25@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2522 ASSAY: Xpert HIV-1 Viral Load, COUNT: 102 |  | GXIV-D-6C | 838139 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, an Icore Checked if Filter is Clogged Performed XpertCheck Performed Self-test Generate Reports | All Modules are Operational | 3 Modules Passed, 1 Newly Replaced |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853244543611033985/5853244543611033985_signature_28.png | Khristofer Ralph L. Andaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5853244543611033985/5853244543611033985_signature_31.png
| 2024-03-04 | 2024-03-01 10:20 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27539 | PBSP | molecularlab@rmc.doh.gov.ph | 23 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4061 ASSAY: Xpert SARS-CoV-2, COUNT: 15 ASSAY: Xpert HIV-1 Viral Load, COUNT: 38 |  | GX-IV-D-6C | 814374 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | Module B2 newly replaced. Modules B1, B3 and B4 passed. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853237953619185342/5853237953619185342_signature_28.png | Justine Lorenz V. Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5853237953619185342/5853237953619185342_signature_31.png
| 2024-03-04 | 2024-02-29 11:54 | Lumban Rural Health Unit | Lumban | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27538 | PBSP | rizalynesacdalan@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1184 |  | GX-IV-D-6C | 838148 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | Unit is operational. Module B3 newly replaced. Module B1, B2 and B4 passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863612897770/5853235863612897770_signature_28.png | Rizalyne Sacdalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863612897770/5853235863612897770_signature_31.png
| 2024-03-04 | 2024-03-01 10:20 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27535 | PBSP | molecularlab@rmc.doh.gov.ph | 23 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4061 ASSAY: Xpert SARS-CoV-2, COUNT: 15 ASSAY: Xpert HIV-1 Viral Load, COUNT: 38 |  | GX-IV-D-6C | 814374 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | Module B2 newly replaced. Modules B1, B3 and B4 passed. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863614768743/5853235863614768743_signature_28.png | Justine Lorenz V. Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863614768743/5853235863614768743_signature_31.png
| 2024-03-04 | 2024-02-28 14:48 | Sumilang Super Health Center | Sumilang | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27534 | PBSP | maricor.obnamia05@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 631 | 2024-02-28 | GX-IV-D-6C | 814163 | Module B4 for replacement | Replaced module B4 | Performed maintenance Replaced module B4 Generated reports | All modules operational. |  | Old: B4/718085 New: B4/876093 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863616395184/5853235863616395184_signature_28.png | Maria Corazon Obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863616395184/5853235863616395184_signature_31.png
| 2024-03-04 | 2024-03-01 15:34 | St. Luke's Medical Center Global City | Rizal Dr. | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 27537 | COMMERCIAL | ci.bgc@stluke.com.ph | 24 |  |  | GX-IV-D-6C | 839231 | For maintenance, checking of date and time | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Checked date and time of workstation Generated reports | Module C1 not operational, for replacement. Modules C2, C3 and C4 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863617724496/5853235863617724496_signature_28.png | Leslie | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863617724496/5853235863617724496_signature_31.png
| 2024-03-04 | 2024-03-01 15:23 | St. Luke's Medical Center Global City | Rizal Dr. | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 27536 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 173 ASSAY: Xpert CT/NG, COUNT: 1192 ASSAY: Xpert HIV-1 Viral Load, COUNT: 54 ASSAY: Xpert HPV, COUNT: 988 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2485 ASSAY: Xpert HCV VL, COUNT: 29 |  | GX-IV-L-6C | 808411 | For maintenance, checking of date and time. | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Checked Workstation date and time Generated reports | Modules B1 and B4 not operational, for replacement. Modules B2 and B3 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863619272256/5853235863619272256_signature_28.png | Leslie | https://www.jotform.com/uploads/macarewecare/222927978372067/5853235863619272256/5853235863619272256_signature_31.png
| 2024-03-04 | 2024-02-29 11:07 | San Mateo RHU 5 | Sto niño,San mateo | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27533 | PBSP | mmacahis@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: |  | GX4 D | 838161 | Due for xpertcheck and maintenance | Cleaned unit Performed xpertcheck Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | B3 and B4 failed xpertcheck B1 and B2 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853231683616009294/5853231683616009294_signature_28.png | Monina Rica M. De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5853231683616009294/5853231683616009294_signature_31.png
| 2024-03-04 | 2024-02-29 08:45 | Rodriguez RHU | J.P Rizal,Rodriguez Rizal | 4A | Rizal | Tb dots | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27532 | PBSP | karlbanal25@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3014 ASSAY: Xpert HIV-1 Viral Load, COUNT: 37 | 2024-02-27 | GX4 | 807580 | For module replacement B2 737085 - 801853 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | 900-0491rs genexpert module | G6843 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853231103613355053/5853231103613355053_signature_28.png | Karl Muchael M. Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5853231103613355053/5853231103613355053_signature_31.png
| 2024-03-04 | 2024-02-28 14:23 | Lung Center of The Philippines | Quezon avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27531 | MACARE-HBDC | lcpmicrolab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 466 ASSAY: Xpert SARS-CoV-2, COUNT: 1458 ASSAY: Xpert HIV-1 Viral Load, COUNT: 47 ASSAY: Xpert HIV-1 Viral Load, COUNT: 56 |  | GX4 D | 838837 | B1 temp error,and B4 h/w failed | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | B1 and B4 for reporting B2 nad B3 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853230563618429351/5853230563618429351_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5853230563618429351/5853230563618429351_signature_31.png
| 2024-03-04 | 2024-02-28 14:18 | Lung Center of The Philippines | Quezon avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27530 | MACARE-HBDC | lcpmicrolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 135 ASSAY: Xpert SARS-CoV-2, COUNT: 40 |  | GX4 L | 804205 | D2,D3,D4 disabled loss comunication error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Troubleshooting. And maintenance | D1 operational D2,D3,D4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853230133615698427/5853230133615698427_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5853230133615698427/5853230133615698427_signature_31.png
| 2024-03-04 | 2024-02-28 14:13 | Lung Center of The Philippines | Quezon avenue,QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27529 | PBSP | lcpmicrolab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5810 ASSAY: Xpert SARS-CoV-2, COUNT: 2607 ASSAY: Others, COUNT: 2 |  | GX16 L | 801080 | C2 H/W failed | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | 15 module operational C2 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853229643618152655/5853229643618152655_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5853229643618152655/5853229643618152655_signature_31.png
| 2024-03-04 | 2024-02-28 11:18 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27528 | MACARE-HBDC | bloodbankphc@gmail.cim | 28 |  |  | MCS + | 11H204 | As per end user they encounter lid lock mechanism problem of centrifuge. | Checked unit Checked lidlock assy Replaced centrifuge assy taken from 01M201 Performed functionality test Test run unit | Troubleshooting | Operational ready to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853228963611409933/5853228963611409933_signature_28.png | Philippine Heart Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5853228963611409933/5853228963611409933_signature_31.png
| 2024-03-04 | 2024-02-27 13:32 | Tatalon Health Center | Victory avenue QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27527 | MACARE-HBDC | tatalonhc@gmail.com | 34 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1005 |  | GX4 | 831610 | As per end user they encounter unable to power on the workstation. | Performed troubleshooting Checked ups voltage output ok Checked cable connection Cleaned RAM. Performed system restart. | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853228513618306696/5853228513618306696_signature_28.png | Tatalon Health Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5853228513618306696/5853228513618306696_signature_31.png
| 2024-03-04 | 2024-02-26 15:28 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27526 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3046 ASSAY: Xpert SARS-CoV-2, COUNT: 4660 ASSAY: Xpert SARS-CoV-2, COUNT: 10 | 2024-02-20 | GX16 | 110014171 | For module replacement C3 210069272 - 898257 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | D1 reported already, 15 module operational |  | 900-0491rs | 11229 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853228013615504862/5853228013615504862_signature_28.png | Jessica Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5853228013615504862/5853228013615504862_signature_31.png
| 2024-03-04 | 2024-02-29 02:42 | Abra Provincial Hospital | Bangued, Abra | CAR | Abra | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27525 | PBSP | aphpathologydepartment@gmail.com | 24 |  | 2024-02-29 | GXIV-4-D-10C | 110019619 | #NAME? | - Installed GeneXpert 10C System - Setup Necessary Accounts - Installed Required Assays - Generated Reports | - Installed GeneXpert 10C system | Unit is Operational |  | GeneXpert System | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853221265471166992/5853221265471166992_signature_28.png | Riverlie Valera | https://www.jotform.com/uploads/macarewecare/222927978372067/5853221265471166992/5853221265471166992_signature_31.png
| 2024-03-04 | 2024-02-29 02:38 | Abra Provincial Hospital | Bangued, Abra | CAR | Abra | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27524 | PBSP | aphpathologydepartment@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2437 | 2024-02-29 | GXIV-4-D-6C | 807616 | #NAME? | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 662126 >>> 210068732, 662128 >>> 806812 - Generated Reports | #NAME? | Unit is operational |  | GeneXpert Modules | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853219075476779731/5853219075476779731_signature_28.png | Riverlie Vallera | https://www.jotform.com/uploads/macarewecare/222927978372067/5853219075476779731/5853219075476779731_signature_31.png
| 2024-03-04 | 2024-02-29 02:35 | Manabo Rural Health Unit | Manabo, Abra | CAR | Abra | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27523 | DOH | lenmayclairelumiwag15@gmail.com | 28 |  | 2024-02-21 | GXIV-4-D-6C | 838274 | #NAME? | #NAME? | #NAME? | Unit is operational |  | CPU | DR 11232 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853216935479903137/5853216935479903137_signature_28.png | Ellen May Claire S. Lumiwag | https://www.jotform.com/uploads/macarewecare/222927978372067/5853216935479903137/5853216935479903137_signature_31.png
| 2024-03-04 | 2024-02-27 10:03 | Umingan Community Hospital | Umingan Community Hospital | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27522 | DOH | jeffreyvrumbaoa@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1120 | 2024-02-20 | GXIV-4-D-6C | 838473 | #NAME? | #NAME? | #NAME? | Unit is operational | Module B2 failed ReXpertCheck, For reporting | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5853209315477431088/5853209315477431088_signature_28.png | Jeffrey Rumbaoa | https://www.jotform.com/uploads/macarewecare/222927978372067/5853209315477431088/5853209315477431088_signature_31.png
| 2024-03-03 | 2024-03-01 12:26 | Masbate City Health Office | Masbate City | 5 | Masbate | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27521 | PBSP | desrmt2017@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1263 |  | Gx4-6C | 808633 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icores Replaced fan filter Generate reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5852860156619889720/5852860156619889720_signature_28.png | Desiree Mae D Dela Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5852860156619889720/5852860156619889720_signature_31.png
| 2024-03-01 | 2024-02-28 16:55 | Rakkk Prophet Medical Center | Maharlika Highway, Brgy. Rosario, Gumaca, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27520 | PBSP | rpmcimoleclab20@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 119 ASSAY: Xpert SARS-CoV-2, COUNT: 18 |  | GXIV 6C | 834625 | CPU beeping sound | Cleaned RAM pins Cleaned SSD pins Checked fan Problem Solved Performed PM Generated report | Troubleshoot CPU | unit is operational all modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850912563619044043/5850912563619044043_signature_28.png | Katrina Rose F. Aureada | https://www.jotform.com/uploads/macarewecare/222927978372067/5850912563619044043/5850912563619044043_signature_31.png
| 2024-03-01 | 2024-02-28 12:44 | Guinayangan Municipal Health Office | Brgy. Calimpak, Guinayangan, Quezon Province | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27519 | WHO | lala0117r.kampitan@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 597 | 2024-01-26 | GXIV 6C - D | 847334 | Unit due for calibration | Checked and cleaned cartridge bays, plunger rod, icore slots and filter Performed XpertCheck on modules B1, B2 and B4 . B3 newly replaced | Performed XpertCheck | Unit is operational All modules working | Used 4 xpertcheck cartridges | XpertCheck kit | DR 46612 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850902093611282113/5850902093611282113_signature_28.png | Leonila Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5850902093611282113/5850902093611282113_signature_31.png
| 2024-03-01 | 2024-02-29 08:03 | Gumaca District Hospital | Maharlika Highway, Brgy. Rosario, Gumaca | 4A | Quezon | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27518 | PBSP | danecadanna@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 409 | 2024-02-19 | GXIV 6C L | 804684 | Modules C1 and C2 for replacement Unit due for calibration | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced modules C1 and C2 Performed XpertCheck on C3 and C4 C3 and C4 passed XpertCheck Generated report | Replaced Modules C1: 667447 -> 719925 C2: 642538 -> 890611 Performed XpertCheck | Unit is operational All modules working | Used 4 xpertcheck cartridges | 2 GX Modules 6C, XpertCheck kit | DR 47083, DR 47170, DR 46613 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850900843617891202/5850900843617891202_signature_28.png | danna edan | https://www.jotform.com/uploads/macarewecare/222927978372067/5850900843617891202/5850900843617891202_signature_31.png
| 2024-03-01 | 2024-03-01 12:27 | MIA Health Center | MIA HEALTH CENTER, CHAPEL ROAD NAIA AVENUE, PASAY CITY | NCR | NCR 4th - Pasay | Laboratory | NO | INSTALLATION REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27517 | PBSP | mervheequiogue14@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-02-27 | GX IV 6C-D | 812833 | GX for setup to new location Module B1 for replacement Unit due for XpertCheck | Installed GX and accessories Performed PM Replaced Module B1 Setup User admin Performed XpertCheck on Modules 2 and 4 B1: Newly replaced B2: Passed B3: Reported waiting for replacement B4: Failed xpertcheck | Setup GX to new location Duay HC -> MIA HC Replaced Module B1 635266 -> 210073917 Performed XpertCheck | unit is operational 2 modules working 2 modules fo replacement | used 3 XpertCheck cartridges | Gx modules 6C | G6845 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850900353616817280/5850900353616817280_signature_28.png | mervhee sorilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5850900353616817280/5850900353616817280_signature_31.png
| 2024-03-01 | 2024-03-01 15:07 | Salvacion Oppus-Yñiguez Memorial Hospital | Maasin City, Southern Leyte | 8 | Southern Leyte | Laboratory | YES | INSTALLATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27516 | PBSP | montillacherry2@gmail.com | 25 |  |  | 10C GXIV-4D | 110013959 | None | None | -New GeneXpert 10C Unit Installation -Setup GeneXpert workstation including printer and UPS | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850864260322443679/5850864260322443679_signature_28.png | Cherry Oja | https://www.jotform.com/uploads/macarewecare/222927978372067/5850864260322443679/5850864260322443679_signature_31.png
| 2024-03-01 | 2024-03-01 15:04 | Salvacion Oppus-Yñiguez Memorial Hospital | Maasin City, Leyte | 8 | Southern Leyte | Laboratory | YES | Pull-out | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27515 | PBSP | montillacherry2@gmail.com | 25 |  |  | 6C GXIV-4D | 804262 | For pull-out, to be transferred to another site. | Pack-up unit for pull-out. | Pack-up unit for pull-out. | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850860390321412396/5850860390321412396_signature_28.png | Cherry Oja | https://www.jotform.com/uploads/macarewecare/222927978372067/5850860390321412396/5850860390321412396_signature_31.png
| 2024-03-01 | 2024-03-01 14:32 | Sta. Maria Rural Health Unit | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27514 | PBSP | naomidecastro35@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1514 |  | GXIV | 838866 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850852435603044864/5850852435603044864_signature_28.png | Naomi De Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5850852435603044864/5850852435603044864_signature_31.png
| 2024-03-01 | 2024-03-01 12:33 | Bacolod City Health Office | Bacolod City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 27513 | PBSP | jacabiaao.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6570 |  | GX-4 | 812020 | Defective module 643461/D4. | - Preventive Maintenance (optic cleaning, module bay cleaning, plunger rod maintenance, filter cleaning, archive tests). | - Replaced module D4 with new module SN 807732. | 4 Modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850832514571860073/5850832514571860073_signature_28.png | Jenzel Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5850832514571860073/5850832514571860073_signature_31.png
| 2024-03-01 | 2024-03-01 12:36 | Bacolod City Health Office | Bacolod City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27512 | PBSP | jacabiaao.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GX-4 | 804267 | Defective module 210068557/E1 and 702067/E2. | - Preventive Maintenance (optic cleaning, module bay cleaning, plunger rod maintenance, filter cleaning, archive tests). | -Replaced module E1 with new modules SN 684571/E1 and 210074399/E2. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850830674575485464/5850830674575485464_signature_28.png | Jenzel Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5850830674575485464/5850830674575485464_signature_31.png
| 2024-03-01 | 2024-03-01 12:53 | Maguindanao Provincial Hospital | Datu Hoffer, Maguindanao | BARMM | Maguindanao | Tb-Lab | YES | INSTALLATION | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27511 | PBSP | maulananor03@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | Gx-IV (10 color) | 110020260 | N/A | Installed Genexpert machine SN:110020260 Installed CPU SN:*CPHO523F42120 Installed Monitor SN:V5FTG688 Installed UPS - IT - M2000A Installed Printer HP laser jet M107a | Genexpert machine Installation | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850821697277892998/5850821697277892998_signature_28.png | Norhata Maulana | https://www.jotform.com/uploads/macarewecare/222927978372067/5850821697277892998/5850821697277892998_signature_31.png
| 2024-03-01 | 2024-03-01 13:51 | Maguindanao Provincial Hospital | Datu Hoffer, Maguindanao | BARMM | Maguindanao | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27510 | PBSP | maulananor03@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1670 | 2024-03-01 | Gx-IV 6 color | 807980 | N/A | Perform annual Xpertcheck/Calibration Genexpert Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules |  | Xpertcheck cartridge CE-5 | Lot #:1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850819407277363614/5850819407277363614_signature_28.png | Norhata Maulana | https://www.jotform.com/uploads/macarewecare/222927978372067/5850819407277363614/5850819407277363614_signature_31.png
| 2024-03-01 | 2024-03-01 12:33 | Cataingan Rural Health Unit | Cataingan | 5 | Masbate | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27509 | PBSP | catainganrhu@gmail.com | 27 |  |  | GX4-6C | 812787 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod Generate reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850804361816964722/5850804361816964722_signature_28.png | Rowena Alicante | https://www.jotform.com/uploads/macarewecare/222927978372067/5850804361816964722/5850804361816964722_signature_31.png
| 2024-03-01 | 2024-03-01 12:28 | Cataingan Rural Health Unit | Cataingan | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27508 | PBSP | catainganrhu@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 643 |  | GX4-6C | 838164 | B3/807011 for Module Replacement | Replaced Module: B3 - 807011>>210047214 Cleaned cartridge bays, plunger rod, fan Generate reports | Replaced Module B3/807011 Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850802301814280939/5850802301814280939_signature_28.png | Rowena Cadungog | https://www.jotform.com/uploads/macarewecare/222927978372067/5850802301814280939/5850802301814280939_signature_31.png
| 2024-03-01 | 2024-03-01 12:57 | Cabuyao City Health Office 1 | Rosario Village, Brgy. Sala, Cabuyao | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27507 | PBSP | shirleypalcatara18@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1443 | 2024-03-01 | GXIV | 812794 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  | Xpertcheck kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850800138907126263/5850800138907126263_signature_28.png | Shirley Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5850800138907126263/5850800138907126263_signature_31.png
| 2024-03-01 | 2024-03-01 13:07 | Meycauayan City Health Center VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27506 | DOH | rocelann.caranto14@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1675 ASSAY: Xpert HIV-1 Viral Load, COUNT: 142 |  | GXIV | 831904 | Performed additional troubleshooting by Cepheid | Performed troubleshooting | #NAME? | Unit is operational A2 is under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850791086499409468/5850791086499409468_signature_28.png | Rocel Ann A. Caranto | https://www.jotform.com/uploads/macarewecare/222927978372067/5850791086499409468/5850791086499409468_signature_31.png
| 2024-03-01 | 2024-03-01 12:50 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | TB DOTS | NO | INSTALLATION | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 27505 | PBSP | honeyeichel@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV-D 10C | 120000555 | Installation of 10C machine | Installation | Performed installation and setup of new 10c Machine Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850780937116338065/5850780937116338065_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5850780937116338065/5850780937116338065_signature_31.png
| 2024-03-01 | 2024-03-01 11:31 | Ibaan Municipal Health Office | Brgy. Talaibon, Ibaan | 4A | Batangas | Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27504 | DOH | bagsicemieliza@yahoo.com | 25 |  |  | GXIV | 821299 | Unit for installation. | Installation. | Performed installation. Performed basic training on end-user. Generated reports. | Unit is operational. | Unit was from Tanauan CHO. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850736748901897119/5850736748901897119_signature_28.png | Emie Liza H. Bagsic | https://www.jotform.com/uploads/macarewecare/222927978372067/5850736748901897119/5850736748901897119_signature_31.png
| 2024-03-01 | 2024-03-01 10:22 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | PCR Labortory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 27503 | MACARE-HBDC | honeyeichel@gmail.com | 28 | ASSAY: Xpert SARS-CoV-2, COUNT: 11432 ASSAY: Xpert HBV VL, COUNT: 17 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GXIV-D | 829390 | B1: 662752 temperature error B2:623847 temperature error B3:792278 temperature error | Preventive maintenance | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Generate reports Archived test results Performed functionality test | Unit is partially operational B1,B2 and B3 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850698487114568103/5850698487114568103_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5850698487114568103/5850698487114568103_signature_31.png
| 2024-02-29 | 2024-02-29 16:27 | GOVERNOR CELESTINO GALLARES MEMORIAL MEDICAL CENTER | TAGBILARAN, BOHOL | 7 | Bohol | TB DOTS | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27502 | DOH | majieocmen@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1095 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 |  | GX-IV | 820228 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850046771111387068/5850046771111387068_signature_28.png | Ma. Jernai Sosoban, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5850046771111387068/5850046771111387068_signature_31.png
| 2024-02-29 | 2024-02-29 16:19 | GOVERNOR CELESTINO GALLARES MEMORIAL MEDICAL CENTER | TAGBILARAN, BOHOL | 7 | Bohol | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27501 | PBSP | majieocmen@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1095 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 |  | GX-IV | 838878 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE B2 REPLACEMENT: DEF SN : 860580 NEW SN: 794001 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5850044631111968940/5850044631111968940_signature_28.png | Ma. Jernai Sosoban, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5850044631111968940/5850044631111968940_signature_31.png
| 2024-02-29 | 2024-02-29 14:34 | Tomas Claudio Health Center | Camus Street. Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27500 | DOH | arleen_p_alconera@yahoo.com | 23.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 648 |  | GX-IV 4D | 838492 | No power on System Unit Due for Preventive Maintenance | Reset connection of the system unit Clean ram slots Replaced the CMOS Battery Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance Troubleshoot the System Unit | System unit is able to boot and All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849979831716603853/5849979831716603853_signature_28.png | Arleen P Alconera | https://www.jotform.com/uploads/macarewecare/222927978372067/5849979831716603853/5849979831716603853_signature_31.png
| 2024-02-29 | 2024-02-29 13:13 | Padre Burgos Rural Health Unit | Barangay Burgos, Padre Burgos, Quezon Province | 4A | Quezon | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27499 | PBSP | miralynmendoza30@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 181 | 2024-01-26 | GX IV 6C - D | 847329 | Unit due for calibration | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed XpertCheck B2 and B4 failed xpertcheck Generated report | Performed XpertCheck | unit is operational 2 modules for reporting |  | XpertCheck kit | DR 46625 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849951589015263363/5849951589015263363_signature_28.png | Miralyn Pionilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5849951589015263363/5849951589015263363_signature_31.png
| 2024-02-29 | 2024-02-29 13:30 | Milagros Rural Health Unit | Milagros | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27498 | PBSP | gracemarymesa1029@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1446 |  | GX4-6C | 838244 | B2/610287, B3/807742 for module replacement | Replaced Modules: B2 - 610287>>652302 B3 - 807742>>718960 Cleaned cartridge bays, plunger rods, icores, fan Generate reports | Replaced Module B2/610287, B3/807742 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849948061425624081/5849948061425624081_signature_28.png | Mary Grace C.Mesa | https://www.jotform.com/uploads/macarewecare/222927978372067/5849948061425624081/5849948061425624081_signature_31.png
| 2024-02-29 | 2024-02-29 13:30 | Milagros Rural Health Unit | Milagros | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27497 | PBSP | gracemarymesa1029@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1446 |  | GX4-6C | 838244 | B2/610287, B3/807742 for module replacement | Replaced Modules: B2 - 610287>>652302 B3 - 807742>>718960 Cleaned cartridge bays, plunger rods, icores, fan Generate reports | Replaced Module B2/610287, B3/807742 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849948021427631817/5849948021427631817_signature_28.png | Mary Grace C.Mesa | https://www.jotform.com/uploads/macarewecare/222927978372067/5849948021427631817/5849948021427631817_signature_31.png
| 2024-02-29 | 2024-02-29 13:27 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27496 | DOH | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2433 ASSAY: Xpert HIV-1 Viral Load, COUNT: 123 |  | GXIV | 831596 | CPU not booting | Reset RAM | #NAME? | Unit is operational Unit is under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849940666939787917/5849940666939787917_signature_28.png | Rosebi Jalova | https://www.jotform.com/uploads/macarewecare/222927978372067/5849940666939787917/5849940666939787917_signature_31.png
| 2024-02-29 | 2024-02-29 12:54 | Biliran Provincial Hospital | Naval, Biliran | 8 | Biliran | TB-Dots | YES | Pull-out | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27494 | PBSP | ghelmission13@yahoo.com | 23 |  |  | 6C GXIV-4D | 815273 | For pull-out and transfer to another site. | Pack-up machine for pull-out. | Pack-up machine for pull-out. | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849926648112839470/5849926648112839470_signature_28.png | Gheland Corto | https://www.jotform.com/uploads/macarewecare/222927978372067/5849926648112839470/5849926648112839470_signature_31.png
| 2024-02-29 | 2024-02-29 12:59 | Biliran Provincial Hospital | Naval, Biliran | 8 | Biliran | TB-Dots | YES | INSTALLATION | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27495 | PBSP | ghelmission13@yahoo.com | 25 |  |  | 10C GXIV-4D | 110020264 | None | None | -New 10C GeneXpert Unit Installation -Setup GeneXpert workstation including printer and UPS | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849926648116344195/5849926648116344195_signature_28.png | Gheland Corto | https://www.jotform.com/uploads/macarewecare/222927978372067/5849926648116344195/5849926648116344195_signature_31.png
| 2024-02-29 | 2024-02-29 12:27 | Davao Chest Center | Villa Abrille Street, Davao City | 11 | Davao del Sur | TB-DOTS | YES | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27493 | WHO | dcclabdept@gmail.com | 25.8 |  |  | GX-IV 4D | 804835 | No power on UPS No power on System Unit | Check and reset the connection making sure it is properly plug Reset and Change Cmos Battery | Troubleshoot the System Unit and UPS | Still no power on ups and System Unit Errors documented to be sent to Cepheid for further Assessment | Installed a GeneXpert Software on End User's Personal Laptop for temporarily use the GeneXpert System |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849907285634653597/5849907285634653597_signature_28.png | Maria Lisa Paras RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5849907285634653597/5849907285634653597_signature_31.png
| 2024-02-29 | 2024-02-29 11:41 | Potrero Health Center | Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27492 | PBSP | ailyn.santos19@yahoo.com | 25 |  |  | GXIV | 804715 | CPU for troubleshooting | Reset RAM | #NAME? | Unit is operational CPU under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849880166933606384/5849880166933606384_signature_28.png | Ailyn Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5849880166933606384/5849880166933606384_signature_31.png
| 2024-02-29 | 2024-02-29 11:09 | Southern Isabela Medical Center | Rosario, Santiago City, Isabela | 2 | Isabela | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27491 | PBSP | cassflotildesvillanueva@yahoo.com | 27 |  | 2024-02-29 | GXIV-10C | 120000738 | N/a | N/a | Installed PC Installed Genexpert Installed UPS Installed Printer Installed Monitor Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849861775274861278/5849861775274861278_signature_28.png | Cassandra Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5849861775274861278/5849861775274861278_signature_31.png
| 2024-02-29 | 2024-02-29 10:50 | La Castella Rural Health Unit | La Castellana, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | REPAIR ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 27490 | DOH | milouepascual@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 789 |  | GX-4 | 838483 | Defective LCD monitor | Preventive maintenance | #NAME? | Same problem after troubleshooting, recommend replacement of LCD monitor. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849844800712055441/5849844800712055441_signature_28.png | Mira Luna Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/5849844800712055441/5849844800712055441_signature_31.png
| 2024-02-29 | 2024-02-29 06:41 | Bago City Health Office | Bago City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@protonmail.com |  |  | 27489 | PBSP | swabteambago@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2682 |  | GX-4 | 815433 | Defective module 800544/C1 | - Preventive Maintenance (optic cleaning, module bay cleaning, plunger rod maintenance, filter cleaning, archive tests). | - Replaced module 800544/C1 with new module with SN 766466. | All 4 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849772866617814584/5849772866617814584_signature_28.png | Glenly Casona | https://www.jotform.com/uploads/macarewecare/222927978372067/5849772866617814584/5849772866617814584_signature_31.png
| 2024-02-28 | 2024-02-28 15:32 | Region II Trauma and Medical Center | Magsaysay, Bayombong, Nueva Vizcaya | 2 | Nueva Vizcaya | Molecular Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | cecillepabance@gmail.com |  | 27488 | PBSP | maricrisvillanueva96@gmail.com | 27 |  | 2024-02-28 | GXIV-10C | 120000537 | N/a | N/a | Instaled Genexpert Installed PC Installed Monitor Installed UPS Installed Drivers Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849167033614286120/5849167033614286120_signature_28.png | Maricris Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5849167033614286120/5849167033614286120_signature_31.png
| 2024-02-28 | 2024-02-28 15:22 | Aroroy Rural Health Unit | Brgy. Baga-Uma, Aroroy | 5 | Masbate | Laboratory | NO | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27487 | PBSP | morenorhenaluz@gmail.com | 24 |  |  | GX4-10C | 110019625 | GX4-10C for Installation | Installed Gx4-10C, UPS, Printer, CPU, Monitor Generate Report | Installed Gx4-10C set | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849153583522816497/5849153583522816497_signature_28.png | Rhena Luz M. Etcobanez | https://www.jotform.com/uploads/macarewecare/222927978372067/5849153583522816497/5849153583522816497_signature_31.png
| 2024-02-28 | 2024-02-28 15:31 | Aroroy Rural Health Unit | Brgy. Baga-Uma, Aroroy | 5 | Masbate | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27486 | DOH | morenorhenaluz@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1665 |  | GX4-6C | 838268 | For Packing For Preventive Maintenance | Cleaned Cartridge bays, plunger rod, Packed Gx4-6C, Monitor, CPU, UPS, Printer Generate reports | Performed Preventive Maintenance Packed Gx4-6s set | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849151623525535114/5849151623525535114_signature_28.png | Rhena Luz M. Etcobanez | https://www.jotform.com/uploads/macarewecare/222927978372067/5849151623525535114/5849151623525535114_signature_31.png
| 2024-02-28 | 2024-02-28 14:25 | Vitali Lying-In Centre | Vitali, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27485 | PBSP | erianejanecardenas@gmail.com | 27 |  | 2024-02-28 | GX 4 | 838140 | System unit won't turn on | System unit undergone troubleshooting. | System unit undergone troubleshooting. | System unit is unable to turn on. | System unit is for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849112971855389715/5849112971855389715_signature_28.png | Eriane Jane M. Cardenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5849112971855389715/5849112971855389715_signature_31.png
| 2024-02-28 | 2024-02-28 14:04 | Hinoba-an Rural Health Unit | Hinoba-an, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@protonmail.com |  |  | 27484 | PBSP | johnericvillamater@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 662 |  | GX-4 | 834843 | Stuck Cartridge in module 2. | #NAME? | - Removes stuck cartridge. - Run Xpertcheck on all 2 modules. | All modules passed Xpertcheck; all modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849100146714348479/5849100146714348479_signature_28.png | Loralyn Villamater | https://www.jotform.com/uploads/macarewecare/222927978372067/5849100146714348479/5849100146714348479_signature_31.png
| 2024-02-28 | 2024-02-28 13:28 | San Simon Rural Health Unit | San Simon Pampanga | 3 | Pampanga | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 27483 | PBSP | pingolseven@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 311 |  | GXIV-D | 838895 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Performed xpertcheck Generate reports Archived test results Performed functionality test | Unit is operational 3 modules passed calibration B1 failed calibration | B1 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849088220603162273/5849088220603162273_signature_28.png | Virmar Pingol | https://www.jotform.com/uploads/macarewecare/222927978372067/5849088220603162273/5849088220603162273_signature_31.png
| 2024-02-28 | 2024-02-28 12:25 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Molecular Laboratory | NO | MAINTENANCE Software Upgrade | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27482 | COMMERCIAL | cmc.rtpcr@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 38 ASSAY: Xpert HPV, COUNT: 28 ASSAY: Xpert MTB/RIF Ultra, COUNT: 627 |  | GXII | 812916 | Unit for PM. Unit for software upgrade. | PM and software upgrade. | Performed PM. Performed Genexpert Software Upgrade. Generated reports. | Unit is operational. GeneXpert software is now running on version 6.4. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849038868442604382/5849038868442604382_signature_28.png | Andrea Gail B. Orgesvik | https://www.jotform.com/uploads/macarewecare/222927978372067/5849038868442604382/5849038868442604382_signature_31.png
| 2024-02-28 | 2024-02-28 11:51 | Minalin Rural Health Unit | Minalin Pampanga | 3 | Pampanga | TB DOTS | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27481 | PBSP | bernadetteann0318@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 476 |  | GXIV-D | 834688 | Due for calibration Manual setup of date and time Hard boot up | Xpertcheck Preventive maintenance Bios battery replacement | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Performed xpertcheck Generate reports Archived test results Replaced BIOS battery Performed functionality test | Unitis is operational Under obesrvation 2 modules passed calibration (B1,B2 and B3 2 failed calibration (B4) | Used 6 Xpertcheck kits, pulled out 4 exyra Xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849036590608372349/5849036590608372349_signature_28.png | Bernadette Lazaro | https://www.jotform.com/uploads/macarewecare/222927978372067/5849036590608372349/5849036590608372349_signature_31.png
| 2024-02-28 | 2024-02-28 12:07 | East Avenue Medical Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 27480 | MACARE-HBDC | paulinobernaldo1419@gmail.com | 28 |  |  | MCS | 15H044 | No power of ups and no voltage output, | Checked voltage output main supply ok Checked ups fuse ok Checked power cord continuity ok Checked supply board defective. Checked transformer ok | Troubleshooting | UPS unoperational | For quotation and dor replacement of ups 1kva Model accupower s/n AT1000190402ST0399 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849028334911886804/5849028334911886804_signature_28.png | Christopher Paul Bernaldo | https://www.jotform.com/uploads/macarewecare/222927978372067/5849028334911886804/5849028334911886804_signature_31.png
| 2024-02-28 | 2024-02-28 11:54 | Davao Reproductive Health and Wellness Center | Jacinto St. Davao City | 11 | Davao del Sur | TB-HIV Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27479 | DOH | rhwclab@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 226 ASSAY: Xpert HIV-1 Viral Load, COUNT: 869 ASSAY: Xpert HBV VL, COUNT: 1 |  | GX-IV 4D | 838262 | Due for Preventive Maintenance Module B4 encounter multiple temperature related errors | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck at Module 4 | Perform Preventive Maintenance Perform Functionality Test | Module 4 passed the functionality test | Module 4 for under observation if error. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849025385016967116/5849025385016967116_signature_28.png | Elaine Malano RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5849025385016967116/5849025385016967116_signature_31.png
| 2024-02-28 | 2024-02-28 11:58 | Davao Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27478 | MACARE-HBDC | rhwclab@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 114 ASSAY: Xpert HIV-1 Viral Load, COUNT: 660 ASSAY: Xpert HBV VL, COUNT: 6 |  | GX-IV 4d | 839236 | Due for Preventive Maintenance Error on Modules B2, and B4 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Module B2, B4 encounter cartidge related errors Enduser was advised to run sample under new cartidge |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849024925011854013/5849024925011854013_signature_28.png | Elaine Malano RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5849024925011854013/5849024925011854013_signature_31.png
| 2024-02-28 | 2024-02-28 11:50 | Imus City Health Office I | Velarde Subd, Imus, Cavite | 4A | Cavite | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27477 | PBSP | armandlasq@gmail.com | 28 |  |  | GXIV-D-10C | 120000593 | Unit is due for Installation | Installation | Installed GeneXpert System Retrieved Test Generate Reports | Unit is with UPS. Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849018659025788799/5849018659025788799_signature_28.png | Armand G. Lasquete | https://www.jotform.com/uploads/macarewecare/222927978372067/5849018659025788799/5849018659025788799_signature_31.png
| 2024-02-28 | 2024-02-27 09:47 | Santolan Super Health Center | Evangelista Ave. | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27475 | PBSP | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3530 | 2024-02-27 | GX-IV-D-6C | 838898 | Module B3 for replacement | Replaced module B3 | Performed maintenance Replaced module B3 Generated reports | All modules operational. |  | Old: B3/882702 New: B3/807020 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849009183611955997/5849009183611955997_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5849009183611955997/5849009183611955997_signature_31.png
| 2024-02-28 | 2024-02-27 13:35 | Rosario Super Health Center | Sixto Antonio Ave. | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27476 | PBSP | eden.cawayan@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2388 | 2024-02-27 | GX-IV-D-6C | 843924 | Module B1 for replacement | Replaced module B1 | Performed maintenance Replaced module B1 Generated reports | All modules operational. |  | Old: B1/210101926 New: B1/677762 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5849009183616753322/5849009183616753322_signature_28.png | Eden R. Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5849009183616753322/5849009183616753322_signature_31.png
| 2024-02-28 | 2024-02-28 11:08 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27474 | DOH | burauenlab@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 466 | 2024-02-28 | GXIV | 838551 | n/a | n/a | #NAME? | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848993567772632540/5848993567772632540_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5848993567772632540/5848993567772632540_signature_31.png
| 2024-02-28 | 2024-02-28 10:01 | San Fernando Rural Health Unit 2 | Sindalan San Fernando Pampanga | 3 | Pampanga | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 27473 | PBSP | punoricahrd@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1275 |  | GXIV-D | 838879 | B4:863566 for module replacement | Repair(module replacement) Preventive maintenance | Performed preventive maintenance Performed module replacement Performed functionality test Gathered reports | Unit is operational B4:807350 new module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848971906198629670/5848971906198629670_signature_28.png | Richard Puno | https://www.jotform.com/uploads/macarewecare/222927978372067/5848971906198629670/5848971906198629670_signature_31.png
| 2024-02-28 | 2024-02-28 10:26 | Binan City Health Office 1 | Brgy. Sto. Domingo, Binan | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27472 | PBSP | annefriam.almazora@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 115 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2556 | 2024-02-26 | GXIV | 812780 | Unit for PM. For module replacement of module B1. | PM and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B1 (862476) in replacement to defective B1 (886541). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848968248441070571/5848968248441070571_signature_28.png | Alexandra Mae Evangelista | https://www.jotform.com/uploads/macarewecare/222927978372067/5848968248441070571/5848968248441070571_signature_31.png
| 2024-02-28 | 2024-02-28 08:51 | Mayor Hilarion A Ramiro Sr. Medical Center | Ozamiz City | 10 | Misamis Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 27471 | PBSP | mharsmc.covidlabgenex@gmail.com | 24 |  | 2024-02-28 | Gx4 | 834437 | Stuck-up Cartridge with Lot number. 51801 | Pull out cartridge | Pull out cartridge. | All modules ok to use | Module B1 door is broken. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848912968813992917/5848912968813992917_signature_28.png | John Carlo B. Bongales | https://www.jotform.com/uploads/macarewecare/222927978372067/5848912968813992917/5848912968813992917_signature_31.png
| 2024-02-27 | 2024-02-27 16:12 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | RHIVDA SECTION | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | nimezsheilamarie@gmail.com |  | 27470 | SHELL FOUNDATION | nimezsheilamarie@gmail.com | 23.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 35 ASSAY: Xpert HIV-1 Viral Load, COUNT: 397 ASSAY: Xpert SARS-CoV-2, COUNT: 7 |  | GX - IV 10 COLORS 4L | 120000127 | Due for Preventive Maintenance Module B2 encounter multiple temperature related errors | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform Xpertcheck Calibration at Modules | Perform Preventive Maintenance Perform Module Functionality | Module B2 for Service Required | Errors will be sent to Cepheid for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848322380526940083/5848322380526940083_signature_28.png | Sheila Marie Nimez | https://www.jotform.com/uploads/macarewecare/222927978372067/5848322380526940083/5848322380526940083_signature_31.png
| 2024-02-27 | 2024-02-27 15:07 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27469 | PBSP | vincentmilestone20@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2485 ASSAY: Xpert SARS-CoV-2, COUNT: 1886 ASSAY: Xpert HIV-1 Viral Load, COUNT: 155 |  | GX-IV 4D | 834440 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration and operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848298360521469260/5848298360521469260_signature_28.png | Jenry Ken Vincent Mibato RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5848298360521469260/5848298360521469260_signature_31.png
| 2024-02-27 | 2024-02-27 15:20 | Davao Regional Medical Center | Apokon, Tagum City, Davao Del Norte | 11 | Davao del Norte | Molecular laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27468 | PBSP | vincentmilestone20@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2485 ASSAY: Xpert SARS-CoV-2, COUNT: 1886 ASSAY: Xpert HIV-1 Viral Load, COUNT: 144 |  | Gx-IV 6 color | 805063 | Module 3 H/W failed | Replace defective module SN: 684273 with new module SN: 615093 Perform Genexpert Xpertcheck Genexpert Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine module replacement/repair Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules | All modules passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848297980022997706/5848297980022997706_signature_28.png | Jenry Kent Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5848297980022997706/5848297980022997706_signature_31.png
| 2024-02-27 | 2024-02-27 12:42 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27467 | PBSP | vincentmilestone20@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3762 ASSAY: Xpert SARS-CoV-2, COUNT: 3181 ASSAY: Xpert HIV-1 Viral Load, COUNT: 86 |  | GX-XV1 16D | 838756 | Module A2, B2 and C4 encounter multiple temperature related errors. Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Replaced the Defective Module SN: 750639 with new SN: 677313 Replaced the Defective Module SN: 210091151 with new Module SN: 888732 Replaced the Defective Module SN: 870474 with New Module SN: 866121 Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are operational | Module B4 encounter error, errors sent to JotForm for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848297820524274161/5848297820524274161_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5848297820524274161/5848297820524274161_signature_31.png
| 2024-02-27 | 2024-02-27 15:07 | Davao Regional Medical Center | Apokon, Tagum City, Davao Del Norte | 11 | Davao del Norte | Molecular laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27466 | PBSP | vincentmilestone20@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2485 ASSAY: Xpert SARS-CoV-2, COUNT: 1886 ASSAY: Xpert HIV-1 Viral Load, COUNT: 155 |  | Gx-IV (6-color) | 834626 | Module 3 h/w failed | Replace defective module 3 SN: 656018 with new module SN:210037029 Perform Xpertcheck Genexpert Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert machine module replacement/repair Perform Genexpert Xpertcheck/Calibration | 4 operational modules | All modules passed Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848296130026211303/5848296130026211303_signature_28.png | Jenny Ken Vincent mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5848296130026211303/5848296130026211303_signature_31.png
| 2024-02-27 | 2024-02-27 15:18 | Lorenzo D. Zayco District Hospital | Kabankalan City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | INSTALLATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 27465 | PBSP | beuagravante@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Others, COUNT: 0 |  | GX4- 10c | 110020360 | None | Preventive Maintenance | -Installation of new 10c unit and accessories; -System Configuration - Genexpert, Printer - Daisy chained Gx4 6c (804755) and new 10c unit (110020360); -Functionality testing. | 2 units operational in daisy chain configuration with all 8 modules operational. | Waiting for UPS delivery, system connected directly to line voltage. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848294818314122091/5848294818314122091_signature_28.png | Beulah Agravante, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5848294818314122091/5848294818314122091_signature_31.png
| 2024-02-27 | 2024-02-27 14:13 | Tiaong Rural Health Unit | Don V. Robles St., Poblacion 3, Tiaong, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27464 | PBSP | ilabrgutiaong@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 27 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1627 | 2024-01-26 | GXIV 6C D | 838899 | Module B3 and B4 for replacement | Performed PM Replaced Modules B3 and B4 Generated report | Replaced Modules B3: 863780 -> 877236 B4: 863787 -> 877308 | Unit is operational All modules working |  | 2 GX modules 6C | DR 46642 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848239879429013736/5848239879429013736_signature_28.png | Tiaong RTDL | https://www.jotform.com/uploads/macarewecare/222927978372067/5848239879429013736/5848239879429013736_signature_31.png
| 2024-02-27 | 2024-02-27 12:05 | Democrito O. Plaza Memorial Hospital | Brgy. Patin-ay, Prosperidad | 13 | Agusan del Sur | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27463 | PBSP | lermabpabillan@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1980 |  | GX-4 | 805335 | Module B4 found temperature error and reported thru jotform | None | Replacing Module B4 with SN: 686369 to SN: 675667 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848164280761804867/5848164280761804867_signature_28.png | Lerma B. Pabillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5848164280761804867/5848164280761804867_signature_31.png
| 2024-02-27 | 2024-02-27 09:48 | Tarlac Provincial Hospital | Hospital drive San Vicente Tarlac City | 3 | Tarlac | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 27462 | PBSP | molecularlab.tph@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2078 ASSAY: Xpert HIV-1 Viral Load, COUNT: 521 ASSAY: Xpert HBV VL, COUNT: 20 |  | GXIV-D | 805108 | C1:717934 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Performed functionality test Gathered reports | Unit is operational C1:888375 (new module) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5848094899431119164/5848094899431119164_signature_28.png | Danielle Santillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5848094899431119164/5848094899431119164_signature_31.png
| 2024-02-26 | 2024-02-14 10:26 | Klinika Project 7 Social Hygiene Clinic | Project 7, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27461 | MACARE-HBDC | michelenaval@gmail.com | 29 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 |  | GX IV 10C - D | 120000577 | Printer assistance Forgot login credentials | Assisted end user on how to use printer Created end user credentials Generated report | Printer assistance Setup new user admin | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847467173617997258/5847467173617997258_signature_28.png | Michele Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5847467173617997258/5847467173617997258_signature_31.png
| 2024-02-26 | 2024-02-22 11:45 | National Tuberculosis Reference Laboratory | RITM, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | MBL | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | eracueza.pbsp@gmail.com |  | 27459 | PBSP | ntrlphil_cs@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2411 |  | GXIV-D-6C | 807606 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Filter, Icore and Plunger Rod Performed XpertCheck Performed Self-test Generate Reports | 3 Modules are Operational | B3/662385 Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421403612258676/5847421403612258676_signature_28.png | Era Denice R. Acueza | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421403612258676/5847421403612258676_signature_31.png
| 2024-02-26 | 2024-02-23 14:30 | Tanza Rural Health Unit | Biwas, Tanza, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27460 | PBSP | wenarcala@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2549 |  | GXIV-D-6C | 838162 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Icore, Filter and Plunger Rod Performed Self-test Performed XpertCheck Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421403612585065/5847421403612585065_signature_28.png | Edwina P. Arcala | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421403612585065/5847421403612585065_signature_31.png
| 2024-02-26 | 2024-02-22 11:34 | National Tuberculosis Reference Laboratory | RITM, Filinvest | NCR | NCR 4th - Muntinlupa | MBL | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | eracueza.pbsp@gmail.com |  | 27458 | PBSP | ntrlphil_cs@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 48 ASSAY: Others, COUNT: 1893 |  | GXIV-D-10C | 848244 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Filter, Icore and Plunger Rod Performed XpertCheck Performed Self-test Generate Reports | All Modules are Operational | 3 Pass, 1 New Module |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421403615095628/5847421403615095628_signature_28.png | Era Denice R. Acueza | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421403615095628/5847421403615095628_signature_31.png
| 2024-02-26 | 2024-02-23 14:10 | Tanza Rural Health Unit | Bawis, Tanza,Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27457 | PBSP | wenarcala@gmail.com | 24 |  |  | GXIV-D-6C | 812835 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Icore, Filter and Plunger Rod Performed Self-test Performed XpertCheck Generate Reports | 3 Modules are Operational, 1 for Replacement | B2 Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421403619426053/5847421403619426053_signature_28.png | Edwina P. Arcala | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421403619426053/5847421403619426053_signature_31.png
| 2024-02-26 | 2024-02-22 08:54 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta.Cruz, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27456 | DOH | covidlabjrrmmc@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF Ultra, COUNT: 215 ASSAY: Xpert SARS-CoV-2, COUNT: 96 |  | GXIV-D-6C | 831530 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Filter and Plunger Rod Archived Test Results Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421343612215701/5847421343612215701_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421343612215701/5847421343612215701_signature_31.png
| 2024-02-26 | 2024-02-21 09:33 | San Lazaro Hospital | Quiricada St, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27453 | PBSP | slhtbclab@gmail.com | 24 |  |  | GXIV-D-10C | 120000475 | Unit is due for Installation | Installation | Installed GeneXpert System Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421293617608514/5847421293617608514_signature_28.png | Ashley Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421293617608514/5847421293617608514_signature_31.png
| 2024-02-26 | 2024-02-22 08:45 | Jose Reyes Memorial Medical Center | Rizal Ave, Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27452 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 398 |  | GX16-D-10C | 110013856 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Filter and Plunger Rod Archived Test Results Performed Self-test Generate Reports | 1 Module for Replacement, 15 Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421293619409916/5847421293619409916_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421293619409916/5847421293619409916_signature_31.png
| 2024-02-26 | 2024-02-21 14:57 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Road, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27455 | COMMERCIAL | uphdmc_lab@yahoo.com | 27 | ASSAY: Xpert CT/NG, COUNT: 8 ASSAY: Xpert HBV VL, COUNT: 34 ASSAY: Xpert HPV, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 361 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2024-02-16 | GXIV-D-6C | 831316 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay, Plunger Rod, Filter and Icore Performed XpertCheck Performed Self-test Generate Reports | 2 Modules Passed, 2 Modules Failed | A1: 784655 and A4: 784666 Failed in XpertCheck | XpertCheck Kit | SI 97180 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421303612935555/5847421303612935555_signature_28.png | Inger Lia Jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421303612935555/5847421303612935555_signature_31.png
| 2024-02-26 | 2024-02-22 08:50 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27454 | MACARE-HBDC | covidlabjrrmmc@gamil.com | 24 | ASSAY: Xpert HPV, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 451 ASSAY: Xpert SARS-CoV-2, COUNT: 273 |  | GXIV-D-10C | 110014099 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Filter and Plunger Rod Archived Test Results Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421303614162915/5847421303614162915_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5847421303614162915/5847421303614162915_signature_31.png
| 2024-02-26 | 2024-02-26 14:50 | Institute of Biology - University of the Philippines Diliman | IB National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27451 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | Primo Star | 3116008122 | None | None | Checked lenses Power-on check | Unit is operational but right eyepiece prism has some growth the results in yellow tinge | Storage room/areas should be clean and dry - minimal to no dust, and low humidity. Highly recommended to have a cover for each microscope. Recommended for Preventive Maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847397584436234488/5847397584436234488_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5847397584436234488/5847397584436234488_signature_31.png
| 2024-02-26 | 2024-02-26 14:42 | Institute of Biology - University of the Philippines Diliman | IB National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27450 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | Primo Star | 3116007598 | Knobs are heavy | None | Checked knobs Checked lenses Power-on test | Unit is operational but knobs need further checking and lubrication | Storage room/areas should be clean and dry - minimal to no dust, and low humidity. Highly recommended to have a cover for each microscope. Recommended for Preventive Maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847393524437654256/5847393524437654256_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5847393524437654256/5847393524437654256_signature_31.png
| 2024-02-26 | 2024-02-26 13:02 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27449 | DOH | dyrosemarie71@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1904 |  | GX-IV 4D | 838401 | Module B1 Encounter Multiple Temperature Related Error | Run XpertCheck for Error Verification | Perform Module Troubleshooting | Module B1 for service Required | Data Sent to Cepheid for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847390334259831019/5847390334259831019_signature_28.png | Rosemarie Dy RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5847390334259831019/5847390334259831019_signature_31.png
| 2024-02-26 | 2024-02-26 14:31 | Institute of Biology - University of the Philippines Diliman | IB National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27448 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | Primо Star | 3116008129 | Defective | None | Checked power supply, swapped good power supply Checked bulb, no bulb Checked lenses | Unit is not operational. Need to replace power supply and bulb | Storage room/areas should be clean and dry - minimal to no dust, and low humidity. Highly recommended to have a cover for each microscope. Recommended for Preventive Maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847389344435864085/5847389344435864085_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5847389344435864085/5847389344435864085_signature_31.png
| 2024-02-26 | 2024-02-26 14:22 | Institute of Biology - University of the Philippines Diliman | IB National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27447 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | Axiostar | 41091 | None | Aligned Kohler | Checked lenses Power-on check | Unit is operational but need cleaning | Storage room/areas should be clean and dry - minimal to no dust, and low humidity. Highly recommended to have a cover for each microscope. Recommended for Preventive Maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847381564436391221/5847381564436391221_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5847381564436391221/5847381564436391221_signature_31.png
| 2024-02-26 | 2024-02-26 14:18 | Institute of Biology - University of the Philippines Diliman | IB National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 260 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27446 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 28 |  |  | KF 2 | 462838 | None | None | Checked lenses Power-on check | Unit is operational | Storage room/areas should be clean and dry - minimal to no dust, and low humidity. Highly recommended to have a cover for each microscope. Recommended for Preventive Maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847377164436408707/5847377164436408707_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5847377164436408707/5847377164436408707_signature_31.png
| 2024-02-26 | 2024-02-26 14:13 | Institute of Biology - University of the Philippines Diliman | IB National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 260 | NO | ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27445 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 28 |  |  | KF 2 | 466147 | None | None | Checked lenses power-on check | Unit is operational | Storage room/areas should be clean and dry - minimal to no dust, and low humidity. Highly recommended to have a cover for each microscope. Recommended for Preventive Maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847374544431326887/5847374544431326887_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5847374544431326887/5847374544431326887_signature_31.png
| 2024-02-26 | 2024-02-26 13:57 | Institute of Biology - University of the Philippines Diliman | IB National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27444 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF 2 | 466210 | Does not turn on | None | Checked light bulb Checked power supply = 0V output | Unit is not operational. Further testing and isolation of power supply | Storage room/areas should be clean and dry - minimal to no dust, and low humidity. Highly recommended to have a cover for each microscope. Recommended for Preventive Maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847371684433323351/5847371684433323351_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5847371684433323351/5847371684433323351_signature_31.png
| 2024-02-26 | 2024-02-26 13:39 | Institute of Biology - University of the Philippines Diliman | IB National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Room 232 | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27443 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 26 |  |  | KF 2 | 466151 | Working but light suddenly turns off after 1 to 2 hrs. of use | None | Checked light bulb Checked power supply output voltage | Unit is operational but need to further test power supply | Storage room/areas should be clean and dry - minimal to no dust, and low humidity. Highly recommended to have a cover for each microscope. Recommended for Preventive Maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847361904432268523/5847361904432268523_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5847361904432268523/5847361904432268523_signature_31.png
| 2024-02-26 | 2024-02-26 13:33 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 27442 | DOH | vanessadelmonte140@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 827 | 2024-02-26 | GXIV | 838409 | n/a | n/a | Performed preventive maintenance - Cleaned cartridge bay, plunger rod, fan filter, icore - Backed up files | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847352534313637523/5847352534313637523_signature_28.png | Vanessa Mabag | https://www.jotform.com/uploads/macarewecare/222927978372067/5847352534313637523/5847352534313637523_signature_31.png
| 2024-02-26 | 2024-02-22 10:23 | Gov. Celestino Gallares Memorial Hospital | Tagbilaran, Bohol | 7 | Bohol | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27441 | DOH | cyrilxoxocaguco@gmail.com | 28 |  |  | Gx4 | 820228 | Stucked-up Carthridge | Repair Module Door and Remove stucked-up carthridge | Damage Module Door | Systems operatuonal |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847340421118650534/5847340421118650534_signature_28.png | Cyril Caguco | https://www.jotform.com/uploads/macarewecare/222927978372067/5847340421118650534/5847340421118650534_signature_31.png
| 2024-02-26 | 2024-02-23 11:14 | Quezon Medical Center 10C | Provincial Health Office, Barangay XI, Lucena City | 4A | Quezon | laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27440 | PBSP | qmcrtdl2015@gmail.com | 27 |  | 2024-02-16 | GXIV 10C - D | 110019627 | GeneXpertIV 10C for installation | Perform maintenance and check the condition of existing Gx 6C unit. Disassemble of the existing Gx 6C unit. Pack the unit using the 10C boxes to make sure it is safe and ready for transport to new facility. Perform unit set up/Installation of the new Gx 10C unit. Perform operator's basic training. Generated report | Packed GXIV 6C, Installed GXIV 10C | Unit is operational. Allodules working |  | GXIV 10C - D UPS Printer | DR47081 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847270983617251224/5847270983617251224_signature_28.png | Jo-Ann P. Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5847270983617251224/5847270983617251224_signature_31.png
| 2024-02-26 | 2024-02-20 10:40 | Marikina City Health Office | Shoe Ave., Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27439 | PBSP | amymitra22@gmail.com | 26 |  |  | GX-IV-D-10C | 110020361 | For Installation | Installed GeneXpert system | Installed the GeneXpert system -Installed the GeneXpert unit -Installed Workstation -Installed barcode scanner -Installed printer -Installed UPS Functionality test Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847170973612712953/5847170973612712953_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5847170973612712953/5847170973612712953_signature_31.png
| 2024-02-26 | 2024-02-20 11:24 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27438 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1315 ASSAY: Xpert HIV-1 Viral Load, COUNT: 257 | 2024-02-20 | GX-IV-D-6C | 805211 | Modules A3 and A4 for replacement | Replaced modules A3 and A4 | Performed maintenance Replaced modules A3 and A4 Generated reports | All modules operational. |  | Old: A3/753414, A4/691313 New: A3/891641, A4/636812 | G 6826 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847169873613494598/5847169873613494598_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5847169873613494598/5847169873613494598_signature_31.png
| 2024-02-26 | 2024-02-21 15:12 | Antipolo City Health Office | Antipolo | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27435 | PBSP | acho.laboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1504 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX-XVI-6C | 842226 | Stuck cartridge on module D3 | Manual removal of stuck cartridge | Manual removal of stuck cartridge Performed maintenance Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847168203612224836/5847168203612224836_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5847168203612224836/5847168203612224836_signature_31.png
| 2024-02-26 | 2024-02-21 15:12 | Antipolo City Health Office | Antipolo | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27437 | PBSP | acho.laboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1504 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX-XVI-6C | 842226 | Stuck cartridge on module D3 | Manual removal of stuck cartridge | Manual removal of stuck cartridge Performed maintenance Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847168203613888762/5847168203613888762_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5847168203613888762/5847168203613888762_signature_31.png
| 2024-02-26 | 2024-02-21 15:12 | Antipolo City Health Office | Antipolo | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27436 | PBSP | acho.laboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1504 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX-XVI-6C | 842226 | Stuck cartridge on module D3 | Manual removal of stuck cartridge | Manual removal of stuck cartridge Performed maintenance Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847168203619244853/5847168203619244853_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5847168203619244853/5847168203619244853_signature_31.png
| 2024-02-26 | 2024-02-21 15:12 | Antipolo City Health Office | Antipolo | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27434 | PBSP | acho.laboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1504 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX-XVI-6C | 842226 | Stuck cartridge on module D3 | Manual removal of stuck cartridge | Manual removal of stuck cartridge Performed maintenance Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847168173613856106/5847168173613856106_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5847168173613856106/5847168173613856106_signature_31.png
| 2024-02-26 | 2024-02-22 12:08 | Pila Rural Health Unit | Pila | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27432 | PBSP | gezza.basarte@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1173 | 2024-02-22 | GX-IV-D-6C | 834692 | Modules B1 and B3 for replacement. CPU troubleshooting. | Replaced modules B1 and B3 Checked CPU | Performed maintenance Replaced modules B1 and B3 Checked CPU Generated reports | All modules operational. | CPU does not turn on sometimes. Under observation. | Old: B1/782521, B3/786604 New: B1/786197, B3/719057 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847166923611228313/5847166923611228313_signature_28.png | Gezzamin Basarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5847166923611228313/5847166923611228313_signature_31.png
| 2024-02-26 | 2024-02-22 12:08 | Pila Rural Health Unit | Pila | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27433 | PBSP | gezza.basarte@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1173 | 2024-02-22 | GX-IV-D-6C | 834692 | Modules B1 and B3 for replacement. CPU troubleshooting. | Replaced modules B1 and B3 Checked CPU | Performed maintenance Replaced modules B1 and B3 Checked CPU Generated reports | All modules operational. | CPU does not turn on sometimes. Under observation. | Old: B1/782521, B3/786604 New: B1/786197, B3/719057 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847166923614902857/5847166923614902857_signature_28.png | Gezzamin Basarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5847166923614902857/5847166923614902857_signature_31.png
| 2024-02-26 | 2024-02-22 14:56 | Pangil Rural Health Unit | Pangil | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor.macare@gmail.com |  |  | 27431 | PBSP | tacile10@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 572 |  | GX-IV-D-6C | 807661 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned icore Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847166083611984599/5847166083611984599_signature_28.png | Cecile R. Sagritalo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5847166083611984599/5847166083611984599_signature_31.png
| 2024-02-26 | 2024-02-20 10:40 | Marikina City Health Office | Shoe Ave., Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27430 | PBSP | amymitra22@gmail.com | 26 |  |  | GX-IV-D-10C | 110020361 | For Installation | Installed GeneXpert system | Installed the GeneXpert system -Installed the GeneXpert unit -Installed Workstation -Installed barcode scanner -Installed printer -Installed UPS Functionality test Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847165433611954201/5847165433611954201_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5847165433611954201/5847165433611954201_signature_31.png
| 2024-02-26 | 2024-02-20 11:24 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27429 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1315 ASSAY: Xpert HIV-1 Viral Load, COUNT: 257 | 2024-02-20 | GX-IV-D-6C | 805211 | Modules A3 and A4 for replacement | Replaced modules A3 and A4 | Performed maintenance Replaced modules A3 and A4 Generated reports | All modules operational. |  | Old: A3/753414, A4/691313 New: A3/891641, A4/636812 | G 6826 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847164603617323471/5847164603617323471_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5847164603617323471/5847164603617323471_signature_31.png
| 2024-02-26 | 2024-02-20 14:05 | Santolan Super Health Center | Evangelista Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27428 | PBSP | samandlet17@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3554 | 2024-02-20 | GX-IV-D-6C | 838898 | Module B2 for replacement | Replaced module B2 | Performed maintenance Replaced module B2 Generated reports | All modules operational. |  | Old: B2/210059845 New:B2/601491 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847163833612468950/5847163833612468950_signature_28.png | LOLITA FORMALEJO | https://www.jotform.com/uploads/macarewecare/222927978372067/5847163833612468950/5847163833612468950_signature_31.png
| 2024-02-26 | 2024-02-21 13:13 | Baras Rural Health Unit | Baras | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | mitchz0323@gmail.com |  | 27427 | PBSP | mesu.baras@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 606 | 2024-02-20 | GX-IV-D-6C | 815389 | Module B4 for replacement | Replaced module B4 | Performed maintenance Replaced module B4 Generated reports | All modules operational. |  | Old: B4/696136 New: B4/653834 | G 6833 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847162893614063914/5847162893614063914_signature_28.png | Michelle Avela | https://www.jotform.com/uploads/macarewecare/222927978372067/5847162893614063914/5847162893614063914_signature_31.png
| 2024-02-26 | 2024-02-21 15:12 | Antipolo City Health Office | Antipolo | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27426 | PBSP | acho.laboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1504 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX-XVI-6C | 842226 | Stuck cartridge on module D3 | Manual removal of stuck cartridge | Manual removal of stuck cartridge Performed maintenance Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847161433616937289/5847161433616937289_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5847161433616937289/5847161433616937289_signature_31.png
| 2024-02-26 | 2024-02-22 12:08 | Pila Rural Health Unit | Pila | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27425 | PBSP | gezza.basarte@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1173 | 2024-02-22 | GX-IV-D-6C | 834692 | Modules B1 and B3 for replacement. CPU troubleshooting. | Replaced modules B1 and B3 Checked CPU | Performed maintenance Replaced modules B1 and B3 Checked CPU Generated reports | All modules operational. | CPU does not turn on sometimes. Under observation. | Old: B1/782521, B3/786604 New: B1/786197, B3/719057 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847160693614431886/5847160693614431886_signature_28.png | Gezzamin Basarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5847160693614431886/5847160693614431886_signature_31.png
| 2024-02-26 | 2024-02-22 14:56 | Pangil Rural Health Unit | Pangil | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor.macare@gmail.com |  |  | 27424 | PBSP | tacile10@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 572 |  | GX-IV-D-6C | 807661 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned icore Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847159653615171229/5847159653615171229_signature_28.png | Cecile R. Sagritalo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5847159653615171229/5847159653615171229_signature_31.png
| 2024-02-26 | 2024-02-23 10:28 | Cardinal Santos Medical Center | Wilson St., | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | egatchalian@csmc.ph |  | 27423 | COMMERCIAL | egatchalian@csmc.ph | 25 | ASSAY: Xpert HBV VL, COUNT: 103 ASSAY: Xpert HPV, COUNT: 22 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1051 |  | GX-IV-D-10C | 110003714 | Module B4 has error 1002/4012 | Checked module B4 | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Checked module B4 Generated reports | Modules B2 and B3 operational. Module B1/210022152 disabled, reported already. Module B4/210048235 not operational, for reporting to Cepheid. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847158893618692814/5847158893618692814_signature_28.png | Elaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5847158893618692814/5847158893618692814_signature_31.png
| 2024-02-26 | 2024-02-23 10:53 | Cardinal Santos Medical Center | Wilson St. | NCR | NCR 2nd - San Juan | Blood bank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 27422 | COMMERCIAL | ftejada@csmc.ph | 24 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance. | Performed maintenance Run simulation test Checked rotor pumps Verified calibration parameters | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847158293611196370/5847158293611196370_signature_28.png | Fidel Tejada | https://www.jotform.com/uploads/macarewecare/222927978372067/5847158293611196370/5847158293611196370_signature_31.png
| 2024-02-26 | 2024-02-23 13:21 | Batis Health Center | F. Manalo St, | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27421 | DOH | almacenaio16@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1938 ASSAY: Xpert HIV-1 Viral Load, COUNT: 91 |  | GX-IV-D-6C | 831598 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned icore Cleaned fan filter Generated reports | Modules A2 and A3 operational. Modules A1/210091071 and A4/879974 has error 1002, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847156313613020576/5847156313613020576_signature_28.png | Rosario T. Almacen | https://www.jotform.com/uploads/macarewecare/222927978372067/5847156313613020576/5847156313613020576_signature_31.png
| 2024-02-26 | 2024-02-20 10:40 | Marikina City Health Office | Shoe Ave., Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27420 | PBSP | amymitra22@gmail.com | 26 |  |  | GX-IV-D-10C | 110020361 | For Installation | Installed GeneXpert system | Installed the GeneXpert system -Installed the GeneXpert unit -Installed Workstation -Installed barcode scanner -Installed printer -Installed UPS Functionality test Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5847154693613661603/5847154693613661603_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5847154693613661603/5847154693613661603_signature_31.png
| 2024-02-26 | 2024-02-21 14:22 | Tinagacan Rural Health Unit | Brgy. Tinagacan, General Santos City | 12 | South Cotabato | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27419 | PBSP | kolleenarcega24@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 982 | 2024-02-21 | Gx-IV | 838196 | Module 1 Service Required | Replace/repair defective module 1 SN:807303 with new module SN:655741 Genexpert machine Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert module replacement/repair | 4 operational modules |  | Genexpert module 6 color SN:655741 | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5846895321215454000/5846895321215454000_signature_28.png | Mary Faith Kollen Arcega | https://www.jotform.com/uploads/macarewecare/222927978372067/5846895321215454000/5846895321215454000_signature_31.png
| 2024-02-23 | 2024-02-22 08:04 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27418 | PBSP | r1mcxpertlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1397 ASSAY: Xpert SARS-CoV-2, COUNT: 413 | 2024-02-22 | GXXVI-4-D-6C | 842230 | #NAME? | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module D2 617706 >>> 210074100 - Generated Reports | #NAME? | Unit is operational |  | GeneXpert Module 210074100 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844955480245270441/5844955480245270441_signature_28.png | Mark Noni Pimentel | https://www.jotform.com/uploads/macarewecare/222927978372067/5844955480245270441/5844955480245270441_signature_31.png
| 2024-02-23 | 2024-02-21 09:48 | Urdaneta District Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27417 | PBSP | warlyacosta@gmail.com | 27 |  | 2024-02-21 | GXIV-4-D-10C | 120000719 | #NAME? | - Installed GeneXpert 10C System - Setup accounts - Installed Assays - Generated Reports | - Installed GeneXpert 10C system | Unit is operational |  | GeneXpert System | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844953280243575767/5844953280243575767_signature_28.png | Warlito Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5844953280243575767/5844953280243575767_signature_31.png
| 2024-02-23 | 2024-02-21 09:46 | Urdaneta District Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27416 | PBSP | warlyacosta@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2057 | 2024-02-21 | GXIV-4-D-6C | 805109 | #NAME? | #NAME? | #NAME? | Unit is operational |  | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844951690245217913/5844951690245217913_signature_28.png | Warlito Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5844951690245217913/5844951690245217913_signature_31.png
| 2024-02-23 | 2024-02-21 09:08 | Urdaneta Sacred Heart Hospital | Urdaneta City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27415 | COMMERCIAL | labmedsci@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 191 |  | GXII-2D | 820303 | #NAME? | #NAME? | #NAME? | Unit is operational | XpertCheck Required. Last XpertCheck was on 06-18-2020 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844947580247030989/5844947580247030989_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5844947580247030989/5844947580247030989_signature_31.png
| 2024-02-23 | 2024-02-23 18:47 | BALASAN MUNICIPAL HEALTH OFFICE | BALASAN, ILOILO | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27414 | PBSP | rojtullo19@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1552 |  | GX-IV | 807555 | MODULE B1,B2, AND B4 STUCK CARTRIDGE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844947575111754667/5844947575111754667_signature_28.png | MARIA ROJEANETTE TUTOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5844947575111754667/5844947575111754667_signature_31.png
| 2024-02-23 | 2024-02-20 13:22 | Western Pangasinan District Hospital | Alaminos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27413 | PBSP | wpdhlab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1862 |  | GXIV-4-D-6C | 807598 | #NAME? | #NAME? | #NAME? | Unit is operational | Current Air conditioner is not sufficient or too weak to cool the room where GeneXpert Systems are installed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844945060243911823/5844945060243911823_signature_28.png | Czarina Catabay | https://www.jotform.com/uploads/macarewecare/222927978372067/5844945060243911823/5844945060243911823_signature_31.png
| 2024-02-23 | 2024-02-20 12:11 | Western Pangasinan District Hospital | Alaminos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27412 | PBSP | wpdhlab@gmail.com | 28 |  |  | GXIV-4-D-10C | 120000731 | For GeneXpert 10C Installation | - Installed GeneXpert 10C system - Installed required Assays - Setup Accounts - Generated Reports | - Installed GeneXpert 10C system | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844941210248560866/5844941210248560866_signature_28.png | Czarina Catabay | https://www.jotform.com/uploads/macarewecare/222927978372067/5844941210248560866/5844941210248560866_signature_31.png
| 2024-02-23 | 2024-02-19 15:08 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27411 | PBSP | gbmmbg@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 597 | 2024-02-05 | GXIV-4-D-6C | 838437 | #NAME? | #NAME? | #NAME? | Unit is operational | XpertCheck not runned on newly replaced Module B1. 3 modules passed | XpertCheck Kits | DR 11188 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844939290244684576/5844939290244684576_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5844939290244684576/5844939290244684576_signature_31.png
| 2024-02-23 | 2024-02-23 11:08 | Quezon Medical Center | Provincial Health Office, Barangay XI, Lucena City | 4A | Quezon | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27410 | PBSP | qmcrtdl2015@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 56 ASSAY: Xpert MTB/RIF Ultra, COUNT: 291 | 2024-01-26 | GXIV 6C -D | 804692 | Unit due for calibration | Checked and cleaned cartridge bays, plunger rod, icore slot, and filter Performed XpertCheck All modules passed Generated report | Performed XpertCheck | Unit is operational. All modules working | Used 5 XpertCheck cartridges | XpertCheck Kit | DR46630 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844901373849168717/5844901373849168717_signature_28.png | Jo-Ann P. Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5844901373849168717/5844901373849168717_signature_31.png
| 2024-02-23 | 2024-02-22 13:28 | San Narciso Rural Health Unit | Brgy. Pagkakaisa, San Narciso, Quezon | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27409 | PBSP | morena_veluz@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 748 | 2024-01-26 | GXIV 6C- D | 838163 | unit due for calibration | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed Xpert Check All modules passed Generated report | Performed XpertCheck | Unit is operational All modules working | used 5 XpertCheck cartridges | XpertCheck kit | DR 46634 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844897253845583789/5844897253845583789_signature_28.png | MA. MORENA V. DIMACULANGAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5844897253845583789/5844897253845583789_signature_31.png
| 2024-02-23 | 2024-02-23 14:27 | Ifugao General Hospital | Poblacion North, Lagawe, Ifugao | CAR | Ifugao | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27408 | PBSP | meannebahiwag@gmail.com | 25 |  | 2024-02-23 | GXIV-10C | 110019622 | N/a | N/a | Installed Genexpert Installed PC Installed Monitor Installed Printer Installed Drivers Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844794347396549478/5844794347396549478_signature_28.png | Mary Ann Bahiwag | https://www.jotform.com/uploads/macarewecare/222927978372067/5844794347396549478/5844794347396549478_signature_31.png
| 2024-02-23 | 2024-02-23 13:25 | Angeles Rural Health Unit 3 | Angeles City Pampanga | 3 | Pampanga | TB Dots Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 27407 | PBSP | femiepangilinan@yahoo.com | 32 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 127 |  | GXIV-D | 807569 | CMOS Battery for replacement Module 2 Error 5007 | BIOS battery replacement Repair Preventive maintenance | Performed preventive maintenance Performed BIOS battery replacement Performed functionality test Archived test results Gathered reports | Unit is operational Module B2 under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844771421818396593/5844771421818396593_signature_28.png | Femie Pangilinan | https://www.jotform.com/uploads/macarewecare/222927978372067/5844771421818396593/5844771421818396593_signature_31.png
| 2024-02-23 | 2024-02-23 12:25 | Davao del Norte Provincial Hospital-Kapalong Zone | Poblacion, Kapalong | 11 | Davao del Norte | TB-DOTS | YES | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | lot2salcedo@yahoo.com |  | 27406 | PBSP | ddnhkapalongzone.laboratory@gmail.com | 20.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4404 |  | GX-IV 4D | 815270 | Module C4 encountered multiple temperature related errors | Replaced the Defective Module SN: 731230 with new Module SN: 801354 Generate IQR, Assay Stat, and System Log | Perform Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844760976113674244/5844760976113674244_signature_28.png | Leona Olila RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5844760976113674244/5844760976113674244_signature_31.png
| 2024-02-23 | 2024-02-23 13:06 | Maysan 3S Health Center | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.comrcb.macare@gmail.com |  |  | 27404 | PBSP | neldimatera@gmail.com | 25 |  |  | GXIV | 814151 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational C1: Service required C2: passed C3: passed C4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844754101207117449/5844754101207117449_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/5844754101207117449/5844754101207117449_signature_31.png
| 2024-02-23 | 2024-02-23 13:10 | Maysan 3S Health Center | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27405 | PBSP | neldimatera@gmail.com | 25 |  |  | GXIV | 815381 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: passed B2: insufficient cartridge B3: passed B4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844754111207640870/5844754111207640870_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/5844754111207640870/5844754111207640870_signature_31.png
| 2024-02-23 | 2024-02-23 12:04 | Bicol Medical Center | Naga | 5 | Camarines Sur | Bacteriology Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27403 | PBSP | bicolmedcovidlab@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1424 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7464 ASSAY: Xpert SARS-CoV-2, COUNT: 2465 ASSAY: Xpert HIV-1 Viral Load, COUNT: 497 ASSAY: Others, COUNT: 10 |  | GX16_6C | 842225 | D3/210092889 for Module Replacement | Replaced Module: D3 210092889>>210006124 Cleaned cartridge bays, plunger rod, icore, fan, fan filter Generate report | Replaced D3 Defective module Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844713532328205678/5844713532328205678_signature_28.png | Rosalie B. Nacario | https://www.jotform.com/uploads/macarewecare/222927978372067/5844713532328205678/5844713532328205678_signature_31.png
| 2024-02-23 | 2024-02-23 11:45 | Bagong Barrio Health Center | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27402 | DOH | cruzernitamay@gmail.com | 25 |  | 2024-01-16 | GXIV | 831906 | Monitor for replacement | Replaced monitor | #NAME? | Unit is operational |  | GeneXpert Monitor SN: AVI215CMQY05B01060472 | 11123 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844694922345031341/5844694922345031341_signature_28.png | Ernita May Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5844694922345031341/5844694922345031341_signature_31.png
| 2024-02-23 | 2024-02-22 13:27 | Laak Rural Health Unit | Poblacion, Laak | 11 | Davao de Oro | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27401 | PBSP | alonaangco@yahoo.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 741 ASSAY: Xpert MTB/RIF Ultra, COUNT: 741 |  | GX-IV 4D | 835218 | Module B1 and B2 encounter multiple temperature related errors | Replace defective module SN: 789109 with new Module SN:883222 Replaced the Defective Module SN:789161 with new Module SN:890331 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844538503714402816/5844538503714402816_signature_28.png | Alona A. ANGCO | https://www.jotform.com/uploads/macarewecare/222927978372067/5844538503714402816/5844538503714402816_signature_31.png
| 2024-02-22 | 2024-02-21 08:52 | East Avenue Medical Center | East avenue,QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27400 | PBSP | roseannparnes@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 320 | 2024-02-20 | GX4 | 803087 | For module replacement C2 682748 - 715604,C3 718171- 888973,C4 661016 - 714498 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | 900-0491rs genexpert module | G6828 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844135123012346507/5844135123012346507_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/5844135123012346507/5844135123012346507_signature_31.png
| 2024-02-22 | 2024-02-21 09:59 | Lung Center of the Philippines | Quezon avenue,QC | NCR | NCR 2nd - Quezon City | Tbdots | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27399 | PBSP | nilabis.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 573 | 2024-02-20 | GX4 R1 | 704584 | For module replacement B3 684304- 715711 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | 900-0491rs genexpert module | G6829 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844134743014291278/5844134743014291278_signature_28.png | Norman Charles Labis | https://www.jotform.com/uploads/macarewecare/222927978372067/5844134743014291278/5844134743014291278_signature_31.png
| 2024-02-22 | 2024-02-20 13:34 | Tropical Disease Foundation | Amorsolo st,Makati City | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27398 | CONSORTIUM | krcueno@tdf.org.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1637 | 2024-02-20 | GX4 | 812560 | For module replacement C3 684528 - 885095 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | Operational |  | 900-0491rs genexpert module | G6827 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844134033019383565/5844134033019383565_signature_28.png | Karl Francis Cueno | https://www.jotform.com/uploads/macarewecare/222927978372067/5844134033019383565/5844134033019383565_signature_31.png
| 2024-02-22 | 2024-02-22 18:50 | Bicol Medical Center | Naga City | 5 | Camarines Sur | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | bmc.bloodbank@gmail.com |  | 27397 | MACARE-HBDC | judebuenaventurarn@gmail.com | 28 |  | 2024-01-30 | MCS+ | 11J221 | SPM out of tolerance error | Replaced and installed spare spm parts. After replacing the spm the error still occur. Try swapping cpu board,motor control board | Troubleshooting | Unit unoperational For pull out for inhouse repair |  | SPM Sensor assy | 11173 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844133613019341928/5844133613019341928_signature_28.png | Jude Celso P. Buenaventura III | https://www.jotform.com/uploads/macarewecare/222927978372067/5844133613019341928/5844133613019341928_signature_31.png
| 2024-02-22 | 2024-02-22 18:59 | Bicol Medical Center | Naga city | 5 | Camarines Sur | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | bmc.bloodbank@gmail.com |  | 27396 | MACARE-HBDC | judebuenaventurarn@gmail.com | 26 |  | 2024-02-22 | MCS+ | 15H027G | Malfunction of display keypad | Performed calibration Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running Replaced temporary display cable ribbon to use the machine. | Troubleshooting | Unit ready to use | For quotation of display cable ribbon | Cable ribbon taken from 11J221. | 11173 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844133273011061621/5844133273011061621_signature_28.png | Jude Celso P. Buenaventura III | https://www.jotform.com/uploads/macarewecare/222927978372067/5844133273011061621/5844133273011061621_signature_31.png
| 2024-02-22 | 2024-02-22 19:32 | Bicol Medical Center | Naga city | 5 | Camarines Sur | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | bmc.bloodbank@gmail.com |  | 27395 | MACARE-HBDC | judebuenaventurarn@gmail.com | 28 |  |  | MCS+ | 15H031G | Due for calibration and maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Calibration and maintenance | Operational ready to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5844130403018490109/5844130403018490109_signature_28.png | Jude Celso P. Buenaventura III | https://www.jotform.com/uploads/macarewecare/222927978372067/5844130403018490109/5844130403018490109_signature_31.png
| 2024-02-22 | 2024-02-22 15:12 | PHILIPPINE RED CROSS - ILOILO CHAPTER | ILOILO CITY | 6 | Iloilo | BLOOD BANK | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27394 | COMMERCIAL | rlcerbas@gmail.com | 24 |  |  | MCS+ | 98G049 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE, REPLACE RUBBER ROLLERS(4 PCS.) FOR BLOOD PUMP AND TRANSFER PUMP -PERFORM CALIBRATION PROCEDURE -PERFORM CYCLE TEST AND DIAGNOSTICS,PASSED | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843954225118246808/5843954225118246808_signature_28.png | Ryan L Cerbas | https://www.jotform.com/uploads/macarewecare/222927978372067/5843954225118246808/5843954225118246808_signature_31.png
| 2024-02-22 | 2024-02-22 14:38 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27393 | PBSP | bmcgenexpertlab@yahoo.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 890 |  | GXIV | 834623 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational A1: Newly replaced A2: Passed A3: Service required A4: Service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843950255611846596/5843950255611846596_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5843950255611846596/5843950255611846596_signature_31.png
| 2024-02-22 | 2024-02-22 14:59 | Quirino Province Medical Center | Cabarroguis, Quirino | 2 | Quirino | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27392 | PBSP | ylbumidang0817@gmail.com | 23 |  | 2024-02-29 | GXIV-10C | 120000609 | N/a | N/a | Installed Genexpert Installed Monitor Installed PC Installed Printer Installed Drivers Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843949911832057868/5843949911832057868_signature_28.png | Yvone Bumidang | https://www.jotform.com/uploads/macarewecare/222927978372067/5843949911832057868/5843949911832057868_signature_31.png
| 2024-02-22 | 2024-02-22 15:04 | Quirino Province Medical Center | Cabarroguis, Quirino | 2 | Quirino | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | roxasmatthew659@yahoo.com.ph |  | 27391 | PBSP | ylbumidang0817@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1633 | 2024-02-22 | Genexpert IV | 814166 | Module 1 error | Gathered Results | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843949781833695425/5843949781833695425_signature_28.png | Yvone Bumidang | https://www.jotform.com/uploads/macarewecare/222927978372067/5843949781833695425/5843949781833695425_signature_31.png
| 2024-02-22 | 2024-02-22 15:03 | WESTERN VISAYAS SANITARIUM | SANTA BARBARA, ILOILO | 6 | Iloilo | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27390 | DOH | ruthiedequilla@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1177 |  | GX-IV | 838393 | MODULE B3 STUCK CARTRIDGE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843948945111859238/5843948945111859238_signature_28.png | Ruth Dequilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5843948945111859238/5843948945111859238_signature_31.png
| 2024-02-22 | 2024-02-22 13:39 | San Carlos City Health Office | Poblacion, San Carlos City, Negros Occidental | 7 | Negros Occidental | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27389 | PBSP | minozac450@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2444 |  | GX4 | 805033 | Module 1 Serial Number 713029 failed during calibration. | Replacement of Module 1 Serial Number 713029 failed during calibration. Generate Data IQR, Sys log, GXX files, and Assay statistics | Replacement of Module 1 Serial Number 713029 failed during calibration. | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843905351153109591/5843905351153109591_signature_28.png | Christine Minoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5843905351153109591/5843905351153109591_signature_31.png
| 2024-02-22 | 2024-02-22 10:23 | Gov. Celestino Gallares Memorial Hospital | Tagbilaran, Bohol | 7 | Bohol | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27388 | DOH | cyrilxoxocaguco@gmail.com | 28 |  |  | Gx4 | 820228 | Stucked-up Carthridge | Repair Module Door and Remove stucked-up carthridge | Damage Module Door | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843836546918701264/5843836546918701264_signature_28.png | Cyril Caguco | https://www.jotform.com/uploads/macarewecare/222927978372067/5843836546918701264/5843836546918701264_signature_31.png
| 2024-02-22 | 2024-02-22 11:18 | Samar Provincial Hospital | Catbalogan Samar | 8 | Samar | TB - Dots | YES | MAINTENANCE Pullout | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27387 | PBSP | evelynvdomingo@yahoo.com | 25 |  | 2024-02-22 | GXIV - 6C | 808291 | unit for transfer | #NAME? | - Archive data - Back up data (Genexpert folder) - Performed preventive maintenance (Cleaned surface , plunger rod, Cartridge bay, Fan filter) - Packed 6C unit and peripherals using boxes of new 10c system. - Monitor SN: V5GTL380 - CPU SN: MJOC9TWS - UPS SN: 2101784060193700026 - Printer - VNF8R45696 | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843820641524372022/5843820641524372022_signature_28.png | Evelyn Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5843820641524372022/5843820641524372022_signature_31.png
| 2024-02-22 | 2024-02-22 10:56 | Dr.Jorge P. Royeca Hospital | General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27386 | PBSP | fatchoi12001@yahoo.com | 29.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1122 ASSAY: Xpert HIV-1 Viral Load, COUNT: 224 |  | Gx-IV (6-colors) | 838130 | N/A | Re arrange Genexpert set up Genexpert Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843809385218114626/5843809385218114626_signature_28.png | Maria Ivy A. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5843809385218114626/5843809385218114626_signature_31.png
| 2024-02-22 | 2024-02-22 10:58 | Samar Provincial Hospital | Catbalogan Samar | 8 | Samar | TB-DOTS | YES | INSTALLATION | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27385 | PBSP | evelynvdomingo@yahoo.com | 24 |  | 2024-02-22 | GXIV - 10C | 120000737 | N/a | n/a | - Installed Genexpert 10C SN/120000737 - Installed CPU SN: CPH0523F42097 - Intalled Monitor SN: V5GAG428 - Installed UPS SN: SKPHU2KV231210098 - Intalled printer SN: CNBIRBK82W - import assay (MTB Ultra, XDR, HIV-1) - Retrieved data - Generate reports | Unit it Operational |  | Genexpert Machine 10C SN/120000737 CPU SN: CPH0523F42097 Monitor SN: V5GAG428 UPS SN: SKPHU2KV231210098 0rinter SN: CNBIRBK82W | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843808461523336944/5843808461523336944_signature_28.png | evelyn v.domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5843808461523336944/5843808461523336944_signature_31.png
| 2024-02-22 | 2024-02-22 10:50 | Dr. Jorge P. Royeca Hospital | General Santos City | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27384 | PBSP | fatchoi12001@yahoo.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1355 ASSAY: Xpert HIV-1 Viral Load, COUNT: 54 | 2024-02-22 | Gx-IV (6 color) | 805007 | Module 3 service Required | Replace/repair defective module 3 SN: 603701 with new module SN: 893019 Genexpert machine Cleaning Checking of errors Data collection | Perform Genexpert machine Preventive maintenance Perform Genexpert module repair | 4 operational modules |  | Genexpert module 6-color SN: 893019 | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843807855211037719/5843807855211037719_signature_28.png | Maria Ivy A. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5843807855211037719/5843807855211037719_signature_31.png
| 2024-02-22 | 2024-02-22 10:18 | WESTERN VISAYAS MEDICAL CENTER | ILOILO CITY | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27383 | PBSP | mitorilla.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4832 |  | GX-XVI | 839316 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE REPLACEMENT: MODULE B4 DEF SN: 870437 - NEW SN:635869 MODULE D1 DEF SN: 870322 - NEW SN:784716 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843780845111292228/5843780845111292228_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5843780845111292228/5843780845111292228_signature_31.png
| 2024-02-22 | 2024-02-22 09:55 | Navotas Medical and Wellness Center | M. Naval St. Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27382 | PBSP | alejandrinoknowell@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 81 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1710 | 2024-02-20 | GXIV | 807592 | Unit for PM. For module replacement of module B1. | PM and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B1 (758216) in replacement to defective B1 (908485 | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843770689617590854/5843770689617590854_signature_28.png | Noel Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5843770689617590854/5843770689617590854_signature_31.png
| 2024-02-22 | 2024-02-22 08:42 | Tondo Medical Center | H. Lopez Blvd., Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27381 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 10 ASSAY: Xpert HPV, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2832 ASSAY: Xpert SARS-CoV-2, COUNT: 6821 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 | 2024-02-20 | GX16 | 844121 | Unit for PM. For module replacement of module C3. | PM and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module C3 (786433) in replacement to defective C3 (210070826). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843728459611060891/5843728459611060891_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5843728459611060891/5843728459611060891_signature_31.png
| 2024-02-22 | 2024-02-22 08:47 | Tondo Medical Center | H. Lopez Blvd., Tondo Manila | NCR | NCR 1st - Manila | Molecular laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27380 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 88 ASSAY: Xpert SARS-CoV-2, COUNT: 13 | 2024-02-20 | GXIV | 834430 | Unit for PM. For module replacement of modules B1 and B2. | PM and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX modules B1 (783209) and B2 (878345) in replacement to defective B1 (677430) and B2 (210070789). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843727679617756296/5843727679617756296_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5843727679617756296/5843727679617756296_signature_31.png
| 2024-02-21 | 2024-02-21 16:23 | Adela Serra Ty Memorial Medical Center | Capitol Hills, Brgy. Telaje, Tandag City | 13 | Surigao del Sur | PMDT | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27379 | PBSP | astmmclaboratory17@gmail.com | 25 |  |  | GX-4 | 804255 | Module A1 found Error in Temperature | None | Replacing Module A1 with SN: 210018751 to SN; 910202 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843273925317179895/5843273925317179895_signature_28.png | Editha Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5843273925317179895/5843273925317179895_signature_31.png
| 2024-02-21 | 2024-02-21 14:47 | Nabunturan Rural Health Unit | Poblacion Nabunturan | 11 | Davao de Oro | TB-DOTS | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27378 | WHO | mbacerraiii@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1647 |  | GX-IV 4d | 847253 | Due to for Preventive Maintenance, and XpertCheck Error at Module B4 | Replace the defective Module SN: 910233 with new Module SN:628666 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck Calibration | Perform Module Replacement, Preventive Maintenance, and XpertCheck Calibration | Module B3 For Service Required Module B1 and B2 pass the Calibration | Module B3 for under observation if error persist, enduser was advised to disable the Module |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843273551612040686/5843273551612040686_signature_28.png | Manueal Bacerra III | https://www.jotform.com/uploads/macarewecare/222927978372067/5843273551612040686/5843273551612040686_signature_31.png
| 2024-02-21 | 2024-02-21 15:47 | Sorsogon Provincial Hospital | Sorsogon | 5 | Sorsogon | Laboratory | NO | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27377 | PBSP | wizard682000@yahoo.com | 26 |  |  | GX4-10C | 110020680 | Gx4-10 for Installation | Installed Gx4-10C, CPU, Monitor, UPS, Printer Generate Reports | Installed Gx4-10C set | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843121780115844347/5843121780115844347_signature_28.png | Reynaldo B. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5843121780115844347/5843121780115844347_signature_31.png
| 2024-02-21 | 2024-02-21 15:38 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27376 | PBSP | rsjalova1974@yahoo.com | 25 |  |  | GXIV | 805062 | CPU not booting up | Replaced BIOS battery Reset internal cables | #NAME? | Unit is |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843120642713476503/5843120642713476503_signature_28.png | Rosevi Jalova | https://www.jotform.com/uploads/macarewecare/222927978372067/5843120642713476503/5843120642713476503_signature_31.png
| 2024-02-21 | 2024-02-21 15:43 | Sorsogon Provincial Hospital | Sorsogon | 5 | Sorsogon | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27375 | PBSP | wizard682000@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1155 |  | Gx4-6C | 835213 | For Preventive Maintenance Gx for Packing, for relocation | Cleaned cartridge bays, plunger rod, icore Packed Gx4-6X, CPU, Monitor, UPS, Printer Generate reports | Performed Preventive Maintenance Packed GX4 set | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843118250111515793/5843118250111515793_signature_28.png | Reynaldo B. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5843118250111515793/5843118250111515793_signature_31.png
| 2024-02-21 | 2024-02-21 14:39 | Banlat Health Center | Daisy St cor Dafoodial St Banlat Tandang Sora, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Rhea Ann Cañelas |  | rheaannecanelas65@gmail.com |  |  | 27374 | PBSP | reyzelmariesantos@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 114 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1759 | 2024-02-20 | GXIV-6C-D | 804688 | Module A1 for replacement | Checked and cleaned cartridge bays, plunger rod, Icore slot and filter. Replaced module A1 Generated Report | Replaced module A1 608353->210001095 | Unit is operational. All modules are working. |  | GX Module 6C | G 6836 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843111375398061433/5843111375398061433_signature_28.png | Reyzel marie Dulay | https://www.jotform.com/uploads/macarewecare/222927978372067/5843111375398061433/5843111375398061433_signature_31.png
| 2024-02-21 | 2024-02-21 15:15 | University of Perpetual Help Dalta Medical Center | Alabang zapote rd las piñas | NCR | NCR 4th - Las Piñas | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27373 | COMMERCIAL | uphdmc_lab@yahoo.com | 28 |  | 2024-02-21 | MCs + | 11J220 | Safety system error | Performed troubleshooting replaced motor control board,error still occur Replaced AC pump motor error still occur Replaced Transfer pump motor Replaced blood pump motor error disappear, After replacing blood pump motor simulation running complete | Troubleshooting | Unit unoperational Need to replace blood pump motor | For quotation of blood pump motor | Motor Motor control board | 11233 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843101684317791503/5843101684317791503_signature_28.png | Prichi Tonido | https://www.jotform.com/uploads/macarewecare/222927978372067/5843101684317791503/5843101684317791503_signature_31.png
| 2024-02-21 | 2024-02-21 13:42 | Veterans Memorial Medical Center | North Ave. Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Rhea Ann Cañelas |  | rheaannecanelas65@gmail.com |  |  | 27372 | MACARE-HBDC | vmmc.moleclab@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 125 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1590 ASSAY: Xpert HCV VL, COUNT: 15 |  | GXIV-6C-D | 839150 | GX Unit for relocation | Disassembled GX and all accessories Transferred and setup to new location Performed PM Installed SARS-CoV-2 assay Generated Report | Relocate GX from Bacteriology to Molecular Laboratory | Unit is operational. All modules working. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843060731781969091/5843060731781969091_signature_28.png | Cyril Barros | https://www.jotform.com/uploads/macarewecare/222927978372067/5843060731781969091/5843060731781969091_signature_31.png
| 2024-02-21 | 2024-02-20 13:28 | Tampakan Rural Health Unit | Poblacion Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | INSTALLATION TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27371 | DOH | jmedtech_lab@yahoo.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 997 |  | Gx-IV | 838544 | Module 4 multiple temperature errors Module 4 for Service Required | Perform Genexpert Xpertcheck at module 4 for troubleshooting Checking of errors Reporting of problem Cleaning Data collections | Perform Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5843056752213335543/5843056752213335543_signature_28.png | Jovi Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/5843056752213335543/5843056752213335543_signature_31.png
| 2024-02-21 | 2024-02-21 12:06 | Davao de Oro Provincial Hospital - Montevista | Bankerohan Sur, National Highway Montevista | 11 | Davao de Oro | TB-DOTS | YES | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27370 | PBSP | cecillejagurilles1122@gmail.com | 24.8 |  |  | GX-IV 4D | 808624 | Module C3 for replacement | Replaced the Defective Module SN: 210068322 with new Module SN: 210061224 Generate IQR | Perform Module Replacement | Not operational due to cpu system unit problem |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842981869711538261/5842981869711538261_signature_28.png | Cecille A. Jagurilles | https://www.jotform.com/uploads/macarewecare/222927978372067/5842981869711538261/5842981869711538261_signature_31.png
| 2024-02-21 | 2024-02-21 11:28 | Malilipot Rural Health Unit | Malilipot | 5 | Albay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27369 | PBSP | faith.buban2124@gmail.com | 29 |  |  | Gx4-6C | 808280 | Gx4 for Calibration | XpertCheck Check Result: B1/870447 - Passed B2/210102150 - Passed B3/870950 - Passed B4/713836 - Passed Cleaned cartridge bays, plunger rod, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842961165594079508/5842961165594079508_signature_28.png | Fatima ana cabiles | https://www.jotform.com/uploads/macarewecare/222927978372067/5842961165594079508/5842961165594079508_signature_31.png
| 2024-02-21 | 2024-02-21 10:41 | Klinika Bernardo | Quezon City Health Department, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Rhea Ann Cañelas |  | rheaannecanelas65@gmail.com |  |  | 27368 | PBSP | klinikabernardolaboratory@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 27 ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 605 | 2024-02-20 | GXIV-6C-D | 808618 | Module B3 for replacement | Checked and Cleaned cartridge bays, plunger rod, Icore slot and filter Replace module B3 Generated report | Replace module B3 683694-> 806718 | Unit is operational All modules working |  | GX Module 6C | G 6838 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842931968063893458/5842931968063893458_signature_28.png | Mylen Sibal | https://www.jotform.com/uploads/macarewecare/222927978372067/5842931968063893458/5842931968063893458_signature_31.png
| 2024-02-21 | 2024-02-21 10:11 | University of the Philippines - Diliman | Institute of Biology, National Science Complex, UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Mammalian Cell Culture Laboratory | NO | ON-SITE ASSISTANCE | OTHERS | Carlos Leniel Origenes |  | carlos.peniel.origenes@gmail.com |  | AIRINSPACE | 27367 | COMMERCIAL | palim1@up.edu.ph | 24 |  |  | Plasmair Sentinel | 2006-TS-0015 | The room smells sour when the Sentinel is in use | None | Check filters | Unit is operational but the VOC Adsorber R34 (CP01027) needs to be replaced as it is already saturated leading to the sour smell. Pre Filter (CP11200) also needs to be replaced to prevent blockages as this filters out particles coming into the Sentinel | Also recommended for replacement are the Catalyst R33 (CP01026), Collector R32 (CP01066), and Plasma GeN. R31 (CP01024) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842923035496060138/5842923035496060138_signature_28.png | Patricia Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5842923035496060138/5842923035496060138_signature_31.png
| 2024-02-21 | 2024-02-21 10:33 | Malvar Rural Health Unit | Malvar Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27366 | PBSP | lopezsheryl13@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 663 | 2024-02-20 | GXIV | 838191 | Unit for PM. For module replacement of B3. | PM and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module B3 (883330) in replacement to defectivr B3 (806738). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842922971516428581/5842922971516428581_signature_28.png | Sheryl E. Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5842922971516428581/5842922971516428581_signature_31.png
| 2024-02-21 | 2024-02-21 09:28 | San Lazaro Hospital | Quiricada St, Sta.cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27365 | PBSP | slhtbclab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1633 ASSAY: Xpert HIV-1 Viral Load, COUNT: 614 |  | GXIV-D-6C | 707843 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 2 Modules are Operational. 2 Modules are Defective |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842884224642713931/5842884224642713931_signature_28.png | Ashley Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5842884224642713931/5842884224642713931_signature_31.png
| 2024-02-21 | 2024-02-20 10:13 | Alabang Health Center | Liwasan ng mga Bayani, Alabang, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27364 | DOH | cathycarasig@yahoo.com | 27 | ASSAY: Xpert HPV, COUNT: 146 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1980 |  | GXIV-D-6C | 831913 | CPU of GX System is Not Functioning Module 2 with Temperature Error | Troubleshooting | Checked Voltage Output and Cables Cleaned RAM XpertCheck Module 2 Generate Reports | CPU is Operational but for Observation 3 Modules are Operational | Module 2 Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842867864641815625/5842867864641815625_signature_28.png | catheeine carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/5842867864641815625/5842867864641815625_signature_31.png
| 2024-02-21 | 2024-02-20 12:17 | Almanza Uno Health Center | Alabang-Zapote Rd, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27363 | DOH | milcerpl@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1508 ASSAY: Xpert HIV-1 Viral Load, COUNT: 276 |  | GXIV-D-6C | 831918 | Modules 2,3 and 4 with Temperature Errors UPS and Monitor Not Functioning | Troubleshooting | Performed XpertCheck on Modules 2,3 and 4 Checked UPS Output Voltage and Backup Function Checked Monitor Power Cable Generate Reports | 1 Module is Operational. UPS and Monitor are Defective | Modules 2,3 and 4 Failed in XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842867864647960524/5842867864647960524_signature_28.png | Milcer P. Lladoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5842867864647960524/5842867864647960524_signature_31.png
| 2024-02-21 | 2024-02-20 16:05 | Tabaco City Health Unit | Tabaco | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | elyjoy17rmt@gmail.com |  | 27362 | PBSP | chutabacontp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3825 |  | GX4-6C | 804257 | For Preventive Maintenance GX set for packing | Cleaned cartridge bays, plunger rod, icore Packed Gx4, UPS, CPU, Monitor, Printer Generate reports | Performed Preventive Maintenance Packed Gx4 set, for relocation | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842559792334782114/5842559792334782114_signature_28.png | Ely Joy Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5842559792334782114/5842559792334782114_signature_31.png
| 2024-02-21 | 2024-02-20 18:09 | Tabaco City Health Unit | Tabaco | 5 | Albay | Laboratory | NO | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | elyjoy17rmt@gmail.com |  | 27361 | PBSP | chutabacontp@gmail.com | 25 |  |  | GX4-10C | 110019623 | For Installation | Installed Gx4-10C, CPU, Monitor, UPS, Printer Generate Reports | Installed Gx4-10C | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842558302332038672/5842558302332038672_signature_28.png | Ely Joy Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5842558302332038672/5842558302332038672_signature_31.png
| 2024-02-20 | 2024-02-20 15:14 | Agusan del Norte Provincial Health Office | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27360 | PBSP | ronniebalansag74@gmail.com | 30 |  |  | GX4 | 807461 | Module A2 reported thru Jotform (Temperature Error) | None | Replacing Module A2 with SN: 663320 to SN: 877141 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842229092418433771/5842229092418433771_signature_28.png | Ronnie Balansag | https://www.jotform.com/uploads/macarewecare/222927978372067/5842229092418433771/5842229092418433771_signature_31.png
| 2024-02-20 | 2024-02-20 14:47 | Malolos Rural Health Unit | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27359 | WHO | arneilquilao17@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1999 |  | GXIV | 838862 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed calibration Module B3 status is H/W failure Module B3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842225722717721360/5842225722717721360_signature_28.png | Arneil Quilao | https://www.jotform.com/uploads/macarewecare/222927978372067/5842225722717721360/5842225722717721360_signature_31.png
| 2024-02-20 | 2024-02-20 15:09 | Bayugan City Health Office | Brgy. Poblacion, Bayugan City | 13 | Agusan del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27358 | PBSP | lourdescayao@gmail.com | 26 |  |  | GX4 | 815391 | B1 and B3 modules reported in Jotform (Temperature Errors) | None | Replacing Module B1 with SN: 634834 to SN: 685791 Module B3 with SN: 694385 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842224722417331764/5842224722417331764_signature_28.png | Lourdes Diola | https://www.jotform.com/uploads/macarewecare/222927978372067/5842224722417331764/5842224722417331764_signature_31.png
| 2024-02-20 | 2024-02-20 13:59 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27357 | COMMERCIAL | pshclinlab@gmail.com | 22 |  |  | MCS+ | 11H307 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842180836637933244/5842180836637933244_signature_28.png | Roda Fatima Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5842180836637933244/5842180836637933244_signature_31.png
| 2024-02-20 | 2024-02-20 13:56 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27356 | COMMERCIAL | pshclinlab@gmail.com | 22 |  |  | MCS+ | 02D167 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842179726633903940/5842179726633903940_signature_28.png | Roda Fatima Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5842179726633903940/5842179726633903940_signature_31.png
| 2024-02-20 | 2024-02-20 13:08 | Sta. Ignacia Rural Health Unit | Sta. Ignacia Tarla | 3 | Tarlac | TB Dots Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | marj10venzon@gmail.com |  |  | 27355 | DOH | jdi0973@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 508 | 2024-01-30 | GXIV-D | 831909 | UPS for replacement Defective UPS SN:AT1000181128ST0157 | Repair/ Parts replacement | Performed UPS replacement Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Gathered reports Performed functionality test | Unit is operational New UPS SN:AT1K230830ST0016 |  | UPS | 11172 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842164742358058723/5842164742358058723_signature_28.png | Jesus D. Ibarra Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5842164742358058723/5842164742358058723_signature_31.png
| 2024-02-20 | 2024-02-20 11:28 | Davao del Sur Provincial Hospital | Poblacion, Digos City | 11 | Davao del Sur | TB-DOTS | YES | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27354 | PBSP | shammyj0308@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3174 |  | GX-IV 4D | 807464 | Module A4 for replacement | Replace the defective Module SN: 676407 with new Module SN: 861278 Generate IQR, Assay Stat, and System Log | Perform Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842108526215528884/5842108526215528884_signature_28.png | SHARMELLA JOY D. PINOTE | https://www.jotform.com/uploads/macarewecare/222927978372067/5842108526215528884/5842108526215528884_signature_31.png
| 2024-02-20 | 2024-02-20 11:13 | Moncada Rural Health Unit | Moncada Tarlac | 3 | Tarlac | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27353 | PBSP | moncadarhulaboratory.2023@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 654 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 |  | GXIV-D | 808292 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational All modules passed calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5842092052358767349/5842092052358767349_signature_28.png | Jessa Pabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5842092052358767349/5842092052358767349_signature_31.png
| 2024-02-20 | 2024-02-15 09:57 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27352 | PBSP | dmgonzales.pbsp@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 484 |  | GXIV-4-D-6C | 834685 | #NAME? | #NAME? | #NAME? | Unit is operational | Module B4 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841864370241029356/5841864370241029356_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5841864370241029356/5841864370241029356_signature_31.png
| 2024-02-19 | 2024-02-19 16:25 | Legazpi City Health Office | Legazpi | 5 | Albay | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27351 | WHO | menchiecorpuz@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1330 |  | GX4-6C | 846354 | B4/907516 frequent Invalid result | Performed Reliability test Cleaned Icore Generate reports | Performed troubleshooting | Unit is Operational | B4/907516 passed reliability test |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841408469224394956/5841408469224394956_signature_28.png | Menchie corpuz | https://www.jotform.com/uploads/macarewecare/222927978372067/5841408469224394956/5841408469224394956_signature_31.png
| 2024-02-19 | 2024-02-19 15:19 | Bicol Region Diagnostics and Reference Laboratory | Legazpi | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27350 | PBSP | rbalzaga.pbsp@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 532 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1859 |  | GX4-6C | 808636 | B1/210068731 for Module replacement | Replaced Module: B1 - 210068731>>617185 Cleaned cartridge bays, plunger rod, icore, fan Generate reports | Replaced Module B1/210068731 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841378079228447613/5841378079228447613_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5841378079228447613/5841378079228447613_signature_31.png
| 2024-02-19 | 2024-02-19 15:25 | Bicol South Luzon Sub National Reference Laboratory | Legazpi | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27349 | PBSP | rbalzaga.pbsp@gmail.com | 23 | ASSAY: Others, COUNT: 585 |  | GX4-10C | 110016954 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841377189225568724/5841377189225568724_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5841377189225568724/5841377189225568724_signature_31.png
| 2024-02-19 | 2024-02-19 15:31 | Bicol Regional Diagnostics and Reference Laboratory | Legazpi | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27348 | PBSP | rbalzaga.pbsp@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 352 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1794 ASSAY: Xpert HIV-1 Viral Load, COUNT: 328 |  | GX4-6C | 807579 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841376079228808422/5841376079228808422_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5841376079228808422/5841376079228808422_signature_31.png
| 2024-02-19 | 2024-02-19 15:37 | Chong Hua Hospital- Mandaue | Subangdaku , Mandaue City Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27347 | COMMERCIAL | irl.maderazo@gmail.com | 24 |  |  | MCS+ | 11J217 | Red Valve error | Check compressor unit, check valve , conduct valve oculation | Check compressor unit, check valve , conduct valve oculation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841375921115669737/5841375921115669737_signature_28.png | Inna Maderazo,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5841375921115669737/5841375921115669737_signature_31.png
| 2024-02-19 | 2024-02-19 14:25 | Zamboanga City Health Office | Pettit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27346 | PBSP | cityhealthlaboratoryzamboanga@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 773 | 2024-02-19 | GX 4 | 838128 | Requested for xpertcheck calibration assistance for all four modules. | All four modules undergone Xpertcheck calibration and Preventive maintenance. | All four modules undergone Xpertcheck calibration and Preventive maintenance. | All modules passed Xpertcheck calibration | None | N/A | N_A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841353991417214948/5841353991417214948_signature_28.png | CHARLANE DELOS REYES | https://www.jotform.com/uploads/macarewecare/222927978372067/5841353991417214948/5841353991417214948_signature_31.png
| 2024-02-19 | 2024-02-19 14:23 | Danao City health office | Danao ciry, cebu | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27345 | PBSP | chockenriquez@gmail.com | 26 |  |  | Gx 4 | 808286 | Defective module 3 sn 210091171 | New module 3 installed sn 758585 | New module 3 installed sn 758585 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841332569683064418/5841332569683064418_signature_28.png | Nicanor Enriquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5841332569683064418/5841332569683064418_signature_31.png
| 2024-02-19 | 2024-02-16 10:46 | Bontoc General Hospital | Bontoc | CAR | Mountain Province | TB DOTS | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27344 | PBSP | rpdelim.pbsp@gmail.com | 26 |  | 2024-02-16 | GXIV-4-D-10C | 110020314 | For installation | Installed 10 C and required Assays | Installed GeneXpert 10 C | Unit is operational |  | GeneXpert 10C system | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841243950248712334/5841243950248712334_signature_28.png | Roxanne P. Delim | https://www.jotform.com/uploads/macarewecare/222927978372067/5841243950248712334/5841243950248712334_signature_31.png
| 2024-02-19 | 2024-02-16 11:01 | Bontoc General Hospital | Bontoc | CAR | Mountain Province | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27343 | PBSP | rpdelim.pbsp@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1687 |  | GXIV-4-D-6C | 807582 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841241280244849036/5841241280244849036_signature_28.png | Roxanne P. Delim | https://www.jotform.com/uploads/macarewecare/222927978372067/5841241280244849036/5841241280244849036_signature_31.png
| 2024-02-19 | 2024-02-15 12:16 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | TB DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27342 | PBSP | benguetphortdl@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1968 | 2024-02-15 | GXIV-4-D-6C | 807621 | #NAME? | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 714131 >>> 719050 - Run Reliability test on Module B3 and B4 - Generated Reports | #NAME? | Unit is Operational | Modules B3 and B4 for Reporting | GeneXpert Module 719050 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841231200247478927/5841231200247478927_signature_28.png | Cirila O, Sabiano | https://www.jotform.com/uploads/macarewecare/222927978372067/5841231200247478927/5841231200247478927_signature_31.png
| 2024-02-19 | 2024-01-26 09:41 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27341 | PBSP | keithtagudar@gmail.com | 21 | ASSAY: Xpert HPV, COUNT: 552 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1692 |  | GXIV-4-D-6C | 834685 | For Troubleshooting and Maintenance | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841226520247538438/5841226520247538438_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5841226520247538438/5841226520247538438_signature_31.png
| 2024-02-19 | 2024-02-15 09:54 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27340 | DOH | sherryannemarrero@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1353 |  | GXIV-4-D-6C | 838281 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841225040249608378/5841225040249608378_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5841225040249608378/5841225040249608378_signature_31.png
| 2024-02-19 | 2024-02-17 18:12 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Rd, Las Piñas | NCR | NCR 4th - Las Piñas | Blood Bank | NO | TROUBLESHOOTING | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27339 | MACARE-HBDC | uphdmc_lab@yahoo.com | 24 |  |  | MCS+ | 11J220 | Autoloading Safety Error | Troubleshooting | Reset Boards Connection Reset Motor Pump Connections Checked Pump Rollers Performed Functionality Test | Unit is Not Operational | For Additional/Further Troubleshooting. Test kit for Pull-out |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841147413612260578/5841147413612260578_signature_28.png | Prichi Tonido | https://www.jotform.com/uploads/macarewecare/222927978372067/5841147413612260578/5841147413612260578_signature_31.png
| 2024-02-19 | 2024-02-14 09:07 | Imus City Health Office III | Greengate Subd, Imus, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27338 | PBSP | tess19_chua@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1467 |  | GXIV-D-6C | 838166 | Unit is due for XpertCheck | XpertCheck | Performed XpertCheck Checked Filter Cleaned Icore Performed Self-test Generate Reports | 3 Modules are Operational, 1 Defective | 2 Passed, 1 Failed, 1 Newly Replaced |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841147413612081102/5841147413612081102_signature_28.png | Marites M. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5841147413612081102/5841147413612081102_signature_31.png
| 2024-02-19 | 2024-02-15 11:02 | Dasmariñas City Health Office I | D. Mangubat Ave., Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27336 | PBSP | cho1lab@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3020 ASSAY: Xpert HIV-1 Viral Load, COUNT: 127 |  | GXIV-D-6C | 838893 | Module 3 Failed in XpertCheck | Module Replacement | Replaced Module 3 Performed Self-test Generte Reports | All Modules are Operational | B3: 863195 -> 634942 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841147413612990489/5841147413612990489_signature_28.png | Elma R.Belostrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5841147413612990489/5841147413612990489_signature_31.png
| 2024-02-19 | 2024-02-13 15:01 | E. Aldana Health Center | E. Aldana, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratort | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27337 | PBSP | egaguilar10@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1351 | 2024-02-13 | GXIV-D-6C | 804261 | Modules 1,2 and 3 Service Required | Module Replacement | Replaced Modules 1,2 and 3. Cleaned Filter and Fan Performed Self-test Generate Reports | All Modules are Operational | 639339, 725156, 688549 -> 662522, 894805, 884275 | GX Modules | G 6810 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841147413617356456/5841147413617356456_signature_28.png | Edita D.Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5841147413617356456/5841147413617356456_signature_31.png
| 2024-02-19 | 2024-02-19 08:24 | NORTHERN MINDANAO MEDICAL CENTER | CAGAYAN DE ORO, MISAMIS ORIENTAL | 10 | Misamis Oriental | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27335 | PBSP | nmmc.gxmtbrif@gmail.com | 26 | ASSAY: Others, COUNT: 0 |  | GX-IV 10C | 110020363 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5841117344096890449/5841117344096890449_signature_28.png | Rione Frances Lomongo | https://www.jotform.com/uploads/macarewecare/222927978372067/5841117344096890449/5841117344096890449_signature_31.png
| 2024-02-18 | 2024-02-15 13:03 | Bagumbayan Rural Health Unit | Poblacion, Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27334 | DOH | marizalenoid@yahoo.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1288 |  | Gx-IV | 838261 | N/A | Genexpert machine Cleaning Checking of errors Data collection: IQR, System log, Assay statistics report and Archive test | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5840787443489544549/5840787443489544549_signature_28.png | MA. RIZA R. DIONELA | https://www.jotform.com/uploads/macarewecare/222927978372067/5840787443489544549/5840787443489544549_signature_31.png
| 2024-02-16 | 2024-02-16 13:55 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27333 | DOH | dyrosemarie71@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1874 |  | GX-IV 4D | 838401 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838881248425388502/5838881248425388502_signature_28.png | Rosemarie B. Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/5838881248425388502/5838881248425388502_signature_31.png
| 2024-02-16 | 2024-02-16 12:10 | Toril Health Center | Toril, Davao City | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | leabuelis@yahoo.com |  | 27332 | PBSP | torilhealth@yahoo.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1664 |  | GX-IV 4D | 838865 | For XpertCheck Calibration and Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XpertCheck Calibration | Perform Preventive Maintenance and XpertCheck Calibration | Module B1 for service required, other modules passed the Calibration | Module B1 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838880698422193842/5838880698422193842_signature_28.png | Lea L. Buelis RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5838880698422193842/5838880698422193842_signature_31.png
| 2024-02-16 | 2024-02-16 16:44 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27331 | DOH | chaska.lp14@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1997 |  | GX4 | 838283 | Module 2 encounter running problem, hang up module 2 software side | Forced stop module 2, restart GXDX , restart pc Generate IQR, Sys log, Assay Stats, Clean up filter | Forced stop module 2, restart GXDX , restart pc Generate IQR, Sys log, Assay Stats, Clean up filter | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838858711111286499/5838858711111286499_signature_28.png | Lucille Mae A. Potoy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5838858711111286499/5838858711111286499_signature_31.png
| 2024-02-16 | 2024-02-16 16:36 | Moalboal Rural Health Unit | Poblacion, Moalboal Cebu | 7 | Cebu | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27330 | PBSP | rhu_moalboal@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 620 |  | GX4 | 838892 | Operator ask for assistance for annual calibration of genexpert machines. | Performed maintenance , clean Icore, and plunger, generate IQR, GXX and Assay statistics Conduct genexpert check Clean filter | Performed maintenance , clean Icore, and plunger, generate IQR, GXX and Assay statistics Conduct genexpert check Clean filter | B1- Failed during calibration B2, B3 and B4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838821771116086779/5838821771116086779_signature_28.png | Ava Anaciano Gako | https://www.jotform.com/uploads/macarewecare/222927978372067/5838821771116086779/5838821771116086779_signature_31.png
| 2024-02-16 | 2024-02-16 16:07 | Wenceslao Dela Paz Health Center | Culiat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27329 | DOH | ameliafclemente2001@yahoo.com | 29 | ASSAY: Xpert HPV, COUNT: 411 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2908 |  | GX IV 6C - D | 831876 | Unit to return to original location | Renovation done on Wenceslao DLPHC Unit to return to original location Packed GX and all accessories Setup GX and accessories to original location Tested GX Generated report | Unit transferred from SB Park -> Culiat | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838805113847257153/5838805113847257153_signature_28.png | Camille Joy Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5838805113847257153/5838805113847257153_signature_31.png
| 2024-02-16 | 2024-02-16 13:03 | Bicol Region General Hospital and Geriatric Medical Center | San Pedro, Cabusao | 5 | Camarines Sur | Laboratory | NO | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27328 | PBSP | jlvdvera_rmt@yahoo.com | 28 |  |  | GX4-10C | 110019624 | GX4-10C installation | Installed Gx4-10C, Printer, UPS, CPU, Monitor Generate reports | Installed Gx4-10C | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838802949418457421/5838802949418457421_signature_28.png | Jessie lowell de vera | https://www.jotform.com/uploads/macarewecare/222927978372067/5838802949418457421/5838802949418457421_signature_31.png
| 2024-02-16 | 2024-02-16 12:56 | Bicol Region General Hospital and Geriatric Medical Center | San Pedro, Cabusoa | 5 | Camarines Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27327 | PBSP | jlvdvera_rmt@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1851 |  | GX4-6C | 807601 | For Preventive Maintenance | Cleaned Cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838802519417118770/5838802519417118770_signature_28.png | JESSIE LOwell de vera | https://www.jotform.com/uploads/macarewecare/222927978372067/5838802519417118770/5838802519417118770_signature_31.png
| 2024-02-16 | 2024-02-16 15:01 | CEBU DOCTORS UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27326 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 57 ASSAY: Xpert CT/NG, COUNT: 111 ASSAY: Xpert HIV-1 Viral Load, COUNT: 135 ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1055 ASSAY: Xpert HCV VL, COUNT: 103 |  | GX-IV | 110008524 | MODULE B1,B2, AND B3 STUCK CARTRIDGE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838764761115730553/5838764761115730553_signature_28.png | Xavier C. Corraje,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5838764761115730553/5838764761115730553_signature_31.png
| 2024-02-16 | 2024-02-16 12:34 | Batasan Super Health Center | Batasan Hills, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27325 | PBSP | deandomingo12@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 502 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2950 | 2024-01-29 | GX IV 6C-D | 805254 | Unit due for XpertCheck | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed XpertCheck on module D2 Module 1 and 4 passed XpertCheck on 1/29/24 Maodule 3 service required on 01/2924 Module D2 failed xpertcheck Generated report | Performed XpertCheck | unit is operational. Modules 2 and 3 service required | Used 2 cartridges | XpertCheck Kit | Lot 1000907498 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838709074861224920/5838709074861224920_signature_28.png | Dean Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5838709074861224920/5838709074861224920_signature_31.png
| 2024-02-16 | 2024-02-16 09:08 | Fairview Clinical Laboratory | Brgy. Compound Dahlia Ave., Fairview, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27324 | PBSP | fairviewclinicallab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 283 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2017 | 2024-01-26 | GX IV 6C-D | 838876 | Unit due for XpertCheck | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed XpertCheck All modules passed Generated report | Performed XpertCheck | Unit is operational All modules working |  | XpertCheck kit | DR 46570 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838605249299128222/5838605249299128222_signature_28.png | Rinna May Bureros | https://www.jotform.com/uploads/macarewecare/222927978372067/5838605249299128222/5838605249299128222_signature_31.png
| 2024-02-16 | 2024-02-16 08:56 | Governor Celestestino Gallares Memorial Hospital | Tagbilaran, Bohol | 7 | Bohol | Laboratory | NO | REPAIR ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27323 | DOH | cyrilxoxocaguco@gmail.com | 28 |  |  | Gx 4 | 820228 | Stucked up carrthridge module 2 and 3 | Removing of stucked up carthridge | Removing of stucked up carthridge | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838546730751629618/5838546730751629618_signature_28.png | Cyril Caguco | https://www.jotform.com/uploads/macarewecare/222927978372067/5838546730751629618/5838546730751629618_signature_31.png
| 2024-02-15 | 2024-02-15 11:38 | Don Emilio Del Valle Memorial Hospital | Bood, Ubay, Bohol | 7 | Bohol | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27322 | PBSP | janarkid@gmail.com | 29 |  |  | Gx 4 | 807603 | Dective CPU software, SQL error | Intall genexpert software of thier existing CPU unit.. | Intall genexpert software of thier existing CPU unit.. | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838025880133564824/5838025880133564824_signature_28.png | JAN ARKID DIACOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5838025880133564824/5838025880133564824_signature_31.png
| 2024-02-15 | 2024-02-15 14:43 | National Children Hospital | E. Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27321 | PBSP | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 25 ASSAY: Xpert SARS-CoV-2, COUNT: 28 | 2024-02-15 | GX4 D | 803109 | Due for xpertcheck and maintenance | Perfromed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | Module all passed |  | Xpertcheck | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838005584984845253/5838005584984845253_signature_28.png | Jessica Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5838005584984845253/5838005584984845253_signature_31.png
| 2024-02-15 | 2024-02-15 14:19 | National Childrens Hospital | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27320 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3076 ASSAY: Xpert SARS-CoV-2, COUNT: 4744 ASSAY: Others, COUNT: 10 |  | GX 16D 10C | 110014171 | D2 stuck cartridge | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Fixed stuck cartridge,Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838004964985706557/5838004964985706557_signature_28.png | Jessica Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5838004964985706557/5838004964985706557_signature_31.png
| 2024-02-15 | 2024-02-15 11:33 | St Lukes Medical Center Global City | Global City, Taguig | NCR | NCR 4th - Taguig | Cellular Immunology | YES | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomedreports@stlukes.com.ph |  | 27319 | COMMERCIAL | ci.bgc@stlukes.com.ph | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 175 ASSAY: Xpert CT/NG, COUNT: 1236 ASSAY: Xpert HBV VL, COUNT: 52 ASSAY: Xpert HPV, COUNT: 988 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2462 ASSAY: Xpert HCV VL, COUNT: 26 ASSAY: Xpert HIV-1 Viral Load, COUNT: 49 |  | GX-4L | 839231 | C1 reported already for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | C1 diabled C2,C3,C4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838004534986962144/5838004534986962144_signature_28.png | Stephanie Infante | https://www.jotform.com/uploads/macarewecare/222927978372067/5838004534986962144/5838004534986962144_signature_31.png
| 2024-02-15 | 2024-02-14 13:50 | Lung Center of the Philippines ( Koica) | Quezon avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27318 | PBSP | ibtee.pbsp@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 505 | 2024-02-13 | GX4 L R1 | 707450 | For module replacement C1 685469 - 210051258 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | Operational |  | 900-0491rs | G6813 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838003854986536824/5838003854986536824_signature_28.png | Ivan Rhyan B. Tee | https://www.jotform.com/uploads/macarewecare/222927978372067/5838003854986536824/5838003854986536824_signature_31.png
| 2024-02-15 | 2024-02-15 11:50 | St Lukes Medical Center Global City | Global City, Taguig City | NCR | NCR 4th - Taguig | Cellular Immunology | YES | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomedreports@stlukes.com.ph |  | 27317 | COMMERCIAL | ci.bgc@stlukes.com.ph | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 175 ASSAY: Xpert CT/NG, COUNT: 1236 ASSAY: Xpert HBV VL, COUNT: 52 ASSAY: Xpert HPV, COUNT: 988 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2462 ASSAY: Xpert HCV VL, COUNT: 26 ASSAY: Xpert HIV-1 Viral Load, COUNT: 49 |  | GX4 L | 808411 | B1 encounter h/w failed,B4 ecounter temp error. 1002 | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Troubleshooting | B1 and B4 disabled for reporting B2 and B3 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5838003314986816639/5838003314986816639_signature_28.png | Sapla, Grace | https://www.jotform.com/uploads/macarewecare/222927978372067/5838003314986816639/5838003314986816639_signature_31.png
| 2024-02-15 | 2024-02-15 16:40 | Mary Johnston Hospital | 1221 J. Nolasco St., Tondo, Manila | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27316 | PBSP | mjhlaboratory@gmail.com | 28 |  |  | GX IV 6C - D | 804702 | Machine relocated from North Daang Hari Health Center -> Mary Johnston Hospital | Perfomed PM Perform unit set up/Installation of the Gx 6C. Perform operator's basic training. Generated report | Setup Gx and accessories to new location | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837974514722340177/5837974514722340177_signature_28.png | Clarice Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/5837974514722340177/5837974514722340177_signature_31.png
| 2024-02-15 | 2024-02-15 15:20 | Maria Aurora Community Hospital | Brgy. 1, Maria Aurora, Aurora | 3 | Aurora | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27315 | PBSP | mariabelenth@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3303 ASSAY: Xpert HIV-1 Viral Load, COUNT: 35 | 2024-02-15 | Genexpert IV | 807565 | Module 2 error | Module 2 for reporting | Performed Miantenance Gathered Reports Performed Functionality Test Disassembled Genexpert Packed Genexpert | Unit is Operational | Module 2 for replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837920891516440335/5837920891516440335_signature_28.png | Maria Belem T. Teh | https://www.jotform.com/uploads/macarewecare/222927978372067/5837920891516440335/5837920891516440335_signature_31.png
| 2024-02-15 | 2024-02-15 11:50 | Sta. Cruz Rural Health Unit | Poblacion Sta Cruz | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27314 | DOH | chendalagan72@gmail.com | 27.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 564 |  | GX-IV 4d | 838505 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration | Perform Preventive Maintenance and XpertCheck Calibration | All Modules are operational and Passed the Calibration | Damaged filters, will be reported to JotForm |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837911322826308369/5837911322826308369_signature_28.png | Dorothy A. Dalagan RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5837911322826308369/5837911322826308369_signature_31.png
| 2024-02-15 | 2024-02-15 14:17 | Maria Aurora Community Hospital | Brgy. 1, Maria Aurora, Aurora | 3 | Aurora | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27313 | PBSP | mariabelenth@gmail.com | 27 |  | 2024-02-15 | GXIV-10C | 110020362 | N/a | N/a | Installed Genexpert Installed Genexpert PC Installed Monitor Installed Driver Installed UPS Installed Printer Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837883771511505505/5837883771511505505_signature_28.png | Maria Belen T. Teh | https://www.jotform.com/uploads/macarewecare/222927978372067/5837883771511505505/5837883771511505505_signature_31.png
| 2024-02-15 | 2024-02-15 13:18 | San Ildefonso Rural Health Unit | San Ildefonso, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27312 | WHO | marvinisanchez06@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 629 | 2024-02-15 | GXIV | 846041 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: Passed B2: Insufficient cartridge B3: Insufficient cartridge B4: Service required |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837850689313856925/5837850689313856925_signature_28.png | Marvin Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/5837850689313856925/5837850689313856925_signature_31.png
| 2024-02-15 | 2024-02-15 13:16 | Basud Municipal Health Office | Poblacion Dos, Basud | 5 | Camarines Norte | Laboratory | NO | INSTALLATION | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27311 | PBSP | kenedypomposo1925@gmail.com | 25 |  |  | GX4-10C | 110019621 | For Installation of GX4-10C | Installed Gx4-10C, Monitor, UPS, CPU, Printer Generate Reports | Installed Gx4-10C, Monitor, UPS, CPU, Printer 1 | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837850426289444117/5837850426289444117_signature_28.png | Kenedy M. Pomposo | https://www.jotform.com/uploads/macarewecare/222927978372067/5837850426289444117/5837850426289444117_signature_31.png
| 2024-02-15 | 2024-02-15 13:12 | Basud Municipal Health Office | Poblacioon Dos, Basud | 5 | Camarines Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27310 | PBSP | kenedypomposo1925@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1646 |  | GX4-6C | 807600 | For Preventive Maintenance For Packing Gx, Preparation for Relocation | Cleaned cartridge bays, plunger rod, icore Packed Gx4, Monitor, CPU, UPS, Printer Generate reports | Performed Preventive Maintenance Packed Gx4 | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837849716285330970/5837849716285330970_signature_28.png | Kenedy M. Pomposo | https://www.jotform.com/uploads/macarewecare/222927978372067/5837849716285330970/5837849716285330970_signature_31.png
| 2024-02-15 | 2024-02-15 11:23 | Rafael Lazatin Memorial Medical Center | Visitation St. Angeles City Pampanga | 3 | Tarlac | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 27309 | PBSP | onabb_lab@yahoo.com.ph | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1482 |  | GXIV-D | 838874 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Archived test results Performed xpertcheck Generate reports Performed functionality test | Unit is fully operational 2 modules passed calibration 2 modules failed calibration B1 and B4 Used 6 Xpertcheck kits | B1 and B4 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837784249512975372/5837784249512975372_signature_28.png | Perla Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5837784249512975372/5837784249512975372_signature_31.png
| 2024-02-15 | 2024-02-15 11:06 | Vinzons RHU | Brgy. 2, Vinzons | 5 | Camarines Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27308 | DOH | bardonamabel@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 708 |  | GX4-6C | 838576 | UPS for Replacement | Replaced UPS: AT1000 Cleaned cartridge bays. Plunger rod, icore, fan Generate Reports | Replaced UPS Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837760886289932300/5837760886289932300_signature_28.png | Amabel A. Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/5837760886289932300/5837760886289932300_signature_31.png
| 2024-02-15 | 2024-02-15 10:25 | Taguig City Health Office | Bay Breeze, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27307 | PBSP | ginabuendumawal@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3451 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2024-02-13 | GX IV 6C - D | 814372 | Module B4 for replacement | Performed PM Replaced Module B4 Generated report | Replaced module B4 210059802 -> 892293 | Unit is operational All modules working |  | GX module 6C | G6812 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837737264728411352/5837737264728411352_signature_28.png | Gina B.Dumawal | https://www.jotform.com/uploads/macarewecare/222927978372067/5837737264728411352/5837737264728411352_signature_31.png
| 2024-02-15 | 2024-02-15 09:42 | Talisay Municipal Health Office | Maharlika Highway, Poblacion, Yalisay | 5 | Camarines Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27306 | PBSP | savillaeditha@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 935 |  | GX4-6C 2Placer | 835188 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate Reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837708496282537227/5837708496282537227_signature_28.png | Editha V.Savilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5837708496282537227/5837708496282537227_signature_31.png
| 2024-02-15 | 2024-02-14 12:11 | Panabo City Health Office | Poblacion Panabo | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27305 | DOH | nadineeust67@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2501 |  | GX-IV 4D | 838295 | Due for Preventive Maintenance and New Printer Installation | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Install new printer | Perform Preventive Maintenance and Printer Installation | All Modules are operational | New Printer from LGU - Panabo (Epson L5290) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837684479015331244/5837684479015331244_signature_28.png | Nadine Gemma S. Eustaqiou | https://www.jotform.com/uploads/macarewecare/222927978372067/5837684479015331244/5837684479015331244_signature_31.png
| 2024-02-15 | 2024-02-15 08:02 | Tacloban City Health Office | Tacloban City | 8 | Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27304 | DOH | marteetingzon20@gmail.con | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1294 | 2024-02-15 | GXIV | 838421 | n/a | n/a | #NAME? | Unit is operational - Module B2 disabled (Defective) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837656607113056626/5837656607113056626_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5837656607113056626/5837656607113056626_signature_31.png
| 2024-02-14 | 2024-02-14 18:17 | Valencia City Health Office | Brgy. Hagkol, Valencia City | 10 | Bukidnon | RTDL | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27303 | DOH | caluyajopay@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1945 | 2024-02-14 | GX-4 | 838271 | B3 Found Service Required Thru XPC | NONE | Error Checking Using XpertCheck Calibration of Module B3 | 3 MODULES WORKING, 1 MODULE FOR REPLACEMENT (SERVICE REQUIRED) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837156916226123416/5837156916226123416_signature_28.png | Joefelyn A. Caluya | https://www.jotform.com/uploads/macarewecare/222927978372067/5837156916226123416/5837156916226123416_signature_31.png
| 2024-02-14 | 2024-02-14 15:17 | Commonwealth Health Center | 20 Katuparan St., Commonwealth, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27302 | DOH | commonwealthclinical@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 104 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3000 | 2024-02-14 | GXIV 6C - D | 831615 | Printer for replacement | Performed PM Removed old printer Setup new printer Generated report | Replaced printer K10347 -> KLWR16647 | Unit is operational All modules working |  | FC-IT-PRINTER | DR11214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5837056464729992899/5837056464729992899_signature_28.png | Iron Capistrano | https://www.jotform.com/uploads/macarewecare/222927978372067/5837056464729992899/5837056464729992899_signature_31.png
| 2024-02-14 | 2024-02-14 13:45 | Lung Center of the Philippines | Quezon ave QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27301 | PBSP | ibtee.pbsp@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 573 | 2024-02-13 | GX4L R1 | 704584 | For module replacement B2 650425 - 903572 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | Operational |  | 900-0491rs | G6813 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836996530321073731/5836996530321073731_signature_28.png | Ivan Rhyan B. Tee | https://www.jotform.com/uploads/macarewecare/222927978372067/5836996530321073731/5836996530321073731_signature_31.png
| 2024-02-14 | 2024-02-14 12:53 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27300 | PBSP | venvalizado@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 29 ASSAY: Xpert HBV VL, COUNT: 37 ASSAY: Others, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2430 ASSAY: Xpert SARS-CoV-2, COUNT: 10587 ASSAY: Xpert HIV-1 Viral Load, COUNT: 163 | 2024-02-14 | GXIV | 814159 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational B1: Service required B2: Service required B3: Service required B4: Passed |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836992871915160084/5836992871915160084_signature_28.png | Venmarie Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5836992871915160084/5836992871915160084_signature_31.png
| 2024-02-14 | 2024-02-14 12:51 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27299 | PBSP | venvalizado@gmail.com | 25 |  | 2024-02-14 | GXIV | 707523 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational C1: Passed C2: Insufficient cartridge C3: Insufficient cartridge C4: Insufficient cartridge |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836985801914258636/5836985801914258636_signature_28.png | Venmarie Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5836985801914258636/5836985801914258636_signature_31.png
| 2024-02-14 | 2024-02-14 11:46 | Bongabong Municipal Health Office | Bongabong | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27298 | PBSP | cortezmariel30@gmail.com | 25 |  |  | GX-IV-D-6C | 834811 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned icore Cleaned filter Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836963356591949224/5836963356591949224_signature_28.png | Mariel A. Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/5836963356591949224/5836963356591949224_signature_31.png
| 2024-02-14 | 2024-02-14 11:41 | Bongabong Municipal Health Office | Bongabong | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27297 | PBSP | cortezmariel30@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 761 |  | GX-IV-D-6C | 834687 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned icore Cleaned filter Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836963356592852203/5836963356592852203_signature_28.png | Mariel A. Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/5836963356592852203/5836963356592852203_signature_31.png
| 2024-02-14 | 2024-02-14 | South Cotabato Provincial Hospital | Koronadal City, South Cotabato | 12 | South Cotabato | Molecular laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | dappmc.tblab@gmail.com |  | 27296 | COMMERCIAL | angelicaparba132@gmail.com | 28.1 | ASSAY: Xpert SARS-CoV-2, COUNT: 2244 |  | Gx-XVI | 839314 | Module C2,C4,D1,And D2 problem due multiple Temperature errors UPS problem - non functional power-packed up | Perform Genexpert machine Cleaning Checking of errors Data collection: Installation Qualification Report, System log, Assay statistics report and Archive test | Genexpert machine Preventive maintenance | 12 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836928505024663917/5836928505024663917_signature_28.png | ANGELICA PARBA | https://www.jotform.com/uploads/macarewecare/222927978372067/5836928505024663917/5836928505024663917_signature_31.png
| 2024-02-14 | 2024-02-14 11:39 | CAGAYAN DE ORO CITY HEALTH OFFICE | Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27295 | DOH | chocdotbdots@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3085 | 2024-01-15 | GX-IV | 838435 | FOR MODULE C1 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE C1 REPLACEMENT: DEF SN:861422 NEW SN:891040 | EQUIPMENT IS OPERATIONAL |  | GENEXPERT PRE-OWNED GENEXPERT MODULE | 11109 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836914060488824346/5836914060488824346_signature_28.png | Charissa Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5836914060488824346/5836914060488824346_signature_31.png
| 2024-02-14 | 2024-02-14 10:20 | Oriental Mindoro Southern District Hospital | Roxas, Oriental Mindoro | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27294 | PBSP | markedisonmartinez@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2508 | 2024-01-31 | GXIV | 807590 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  | Xpertcheck Kit. | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836882257015847676/5836882257015847676_signature_28.png | Mark Edison Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5836882257015847676/5836882257015847676_signature_31.png
| 2024-02-14 | 2024-02-14 10:31 | San Isidro Rural Health Unit | Poblacion, San Isidro, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27293 | PBSP | urian_marlyn@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1030 | 2024-02-14 | Genexpert IV | 805107 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration (All Modules Passed) Performed Functionality Test | Unit is Operational | N/a | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836873086712768509/5836873086712768509_signature_28.png | Marlyn Urian | https://www.jotform.com/uploads/macarewecare/222927978372067/5836873086712768509/5836873086712768509_signature_31.png
| 2024-02-14 | 2024-02-13 11:14 | North Daang Hari Super Health | 14 road 1,North Daang Hari,Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | CALIBRATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27292 | PBSP | onlychumz@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3043 | 2024-02-13 | GX4 D | 804702 | For packing to be transfer to another facility,and for xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck,Maintenance, | A1,A2 newly replaced, A3 and A4 passed | Packed unit ready for transfer | Xpertcheck -ce-5 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836869193612749018/5836869193612749018_signature_28.png | Charmaine D. Trinidad RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5836869193612749018/5836869193612749018_signature_31.png
| 2024-02-14 | 2024-02-13 11:08 | North Daang Hari Super Health Center | 14 road 1,North daang hari,Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27291 | PBSP | onlychumz@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX4D 10c | 110013753 | For installation | Performed installation Complete accesories Test run unit | Installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836868773612218367/5836868773612218367_signature_28.png | Charmaine D. Trinidad RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5836868773612218367/5836868773612218367_signature_31.png
| 2024-02-14 | 2024-02-13 14:46 | Makati Medical Center (Medilinx) | Amorsolo st, Makati City | NCR | NCR 4th - Makati | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | John.Manzano@sodexo.com |  | 27290 | COMMERCIAL | reinaluz.bactol@medilinx.com.ph | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 161 ASSAY: Xpert CT/NG, COUNT: 682 ASSAY: Xpert HPV, COUNT: 776 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2820 ASSAY: Others, COUNT: 14 ASSAY: Xpert HCV VL, COUNT: 70 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1841 |  | GX4 D (10C) | 110001647 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836868423619570636/5836868423619570636_signature_28.png | Maria Reinaluz Bactol | https://www.jotform.com/uploads/macarewecare/222927978372067/5836868423619570636/5836868423619570636_signature_31.png
| 2024-02-14 | 2024-02-14 08:25 | Bulalacao Municipal Health Office | Bulalacao | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27289 | PBSP | queence.macaraig@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 75 |  | GX-IV-D-6C (2 modules) | 835185 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned icore Cleaned filter Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836865063243225837/5836865063243225837_signature_28.png | Queencelyn festin | https://www.jotform.com/uploads/macarewecare/222927978372067/5836865063243225837/5836865063243225837_signature_31.png
| 2024-02-14 | 2024-02-13 08:41 | Occidental Mindoro Provincial Hospital | Mamburao | 4B | Occidental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27288 | PBSP | omphtbGxlab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1522 ASSAY: Xpert SARS-CoV-2, COUNT: 37 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 |  | GX-IV-D-6C | 807987 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned icore Cleaned filter Run XpertCheck Generated reports | Unit is operational. Mouse for replacement. Unresponsive right click. All modules passed. All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836865063245044613/5836865063245044613_signature_28.png | Larry Amita | https://www.jotform.com/uploads/macarewecare/222927978372067/5836865063245044613/5836865063245044613_signature_31.png
| 2024-02-14 | 2024-02-14 09:40 | Eastern Visayas Medical Center | Bagacay, Tacloban City | 8 | Leyte | laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27287 | PBSP | merl_agner62@yahoo.com.sg | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4014 ASSAY: Xpert HIV-1 Viral Load, COUNT: 202 | 2024-02-14 | GXIV | 808631 | - Module B1/662778 defective - Due for Calibration | #NAME? | - Install new module on B1 w/ SN: 717118 - Performed Functionality test - Performed preventive maintenance (Cleaning surface ,Cartridge bay, plunger rod and Fan filter) - Generate reports - Performed Xpertcheck Calibration - Pullout defective module SN/662778 | Unit is Operational |  | - Genexpert module 6 Color SN/717118 - xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836864872367058800/5836864872367058800_signature_28.png | MA.MERLINA A. VISTAL | https://www.jotform.com/uploads/macarewecare/222927978372067/5836864872367058800/5836864872367058800_signature_31.png
| 2024-02-14 | 2024-02-14 10:00 | East Avenue Medical Center | East avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 27286 | MACARE-HBDC | florescanoel77@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 174 |  | GX4 D 10C | 110014085 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836856303614060091/5836856303614060091_signature_28.png | Rose Ann Parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5836856303614060091/5836856303614060091_signature_31.png
| 2024-02-14 | 2024-02-14 09:57 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27285 | PBSP | bmcgenexpertlab@yahoo.com | 20 |  | 2024-02-08 | GXIV | 834623 | Module A1 (789085) for replacement | Replaced module A1 | #NAME? | Unit is operational No hardware problems found on modules |  | GeneXpert Module | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836854043912990133/5836854043912990133_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5836854043912990133/5836854043912990133_signature_31.png
| 2024-02-14 | 2024-02-14 09:15 | Lagonoy Rural Health Unit | San Rafael, Lagonoy | 5 | Camarines Sur | RTDL & TML Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27284 | WHO | barbacenaml28@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1563 |  | GX4-6C | 846355 | GeneXpert for Calibration | XpertCheck Result: B1/905884 - Passed B2/906159 - Passed B3/906400 - Passed B4/906502 - Passed Cleaned module bays, plunger rod, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836830853114093285/5836830853114093285_signature_28.png | Marietta L. Barbacena | https://www.jotform.com/uploads/macarewecare/222927978372067/5836830853114093285/5836830853114093285_signature_31.png
| 2024-02-14 | 2024-02-14 08:10 | Minglanilla Rural Health Unit | Pide 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27283 | DOH | brendagonzagaparedes@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1468 |  | GX4 | 838512 | Error 1002 - temperature related error in Module 1 SN: 862340 | Conduct calibration Module 1 SN: 862340 Relocate genexpert Generate data IQR, System Log, Gxx files, Assay Statistics | Calibration in 862340 Module 1, Relocate genexpert machine for temperature stabilization, | Module 1 Module 1 SN: 862340 - Pass during calibration All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836791730199227886/5836791730199227886_signature_28.png | Brenda Paredes , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5836791730199227886/5836791730199227886_signature_31.png
| 2024-02-14 | 2024-02-14 08:10 | Minglanilla Rural Health Unit | Pide 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27282 | DOH | brendagonzagaparedes@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1468 |  | GX4 | 838512 | Error 1002 - temperature related error in Module 1 SN: 862340 | Conduct calibration Module 1 SN: 862340 Relocate genexpert Generate data IQR, System Log, Gxx files, Assay Statistics | Calibration in 862340 Module 1, Relocate genexpert machine for temperature stabilization, | Module 1 Module 1 SN: 862340 - Pass during calibration All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836791640193392253/5836791640193392253_signature_28.png | Brenda Paredes , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5836791640193392253/5836791640193392253_signature_31.png
| 2024-02-13 | 2024-02-13 19:36 | NORTHERN MINDANAO MEDICAL CENTER | Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | For Pull-Out | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27281 | PBSP | ronnel_mabasa@yahoo.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 164 |  | GX IV | 804682 | FOR PULL OUT(TRANSFER GX EQUIPMENT) TO ANOTHER FACILITY | COLLECT IQR,SYSTEM LOG, AND ARCHIVE | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836336326315557073/5836336326315557073_signature_28.png | Ronnel Mabasa | https://www.jotform.com/uploads/macarewecare/222927978372067/5836336326315557073/5836336326315557073_signature_31.png
| 2024-02-13 | 2024-02-13 19:10 | NORTHERN MINDANAO MEDICAL CENTER | Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27280 | PBSP | nmmc.gxmtbrif@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 578 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7842 ASSAY: Xpert SARS-CoV-2, COUNT: 9648 ASSAY: Xpert HIV-1 Viral Load, COUNT: 332 |  | GX XVI | 844118 | FOR MODULE A1 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE A1 REPLACEMENT: DEF SN: 210067859 NEW SN:672777 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836322965111064737/5836322965111064737_signature_28.png | Rione Frances Lomongo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5836322965111064737/5836322965111064737_signature_31.png
| 2024-02-13 | 2024-02-13 19:22 | NORHTERN MINDANAO MEDICAL CENTER | Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27279 | PBSP | nmmc.gxmtbrif@gmail.com | 26 | ASSAY: Others, COUNT: 0 |  | GX-IV 10C | 110020363 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836328396315563585/5836328396315563585_signature_28.png | Rione Frances Lomongo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5836328396315563585/5836328396315563585_signature_31.png
| 2024-02-13 | 2024-02-13 19:10 | NORTHERN MINDANAO MEDICAL CENTER | Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27278 | PBSP | nmmc.gxmtbrif@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 578 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7842 ASSAY: Xpert SARS-CoV-2, COUNT: 9648 ASSAY: Xpert HIV-1 Viral Load, COUNT: 332 |  | GX XVI | 844118 | FOR MODULE A1 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE A1 REPLACEMENT: DEF SN: 210067859 NEW SN:672777 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836322965111064737/5836322965111064737_signature_28.png | Rione Frances Lomongo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5836322965111064737/5836322965111064737_signature_31.png
| 2024-02-13 | 2024-02-13 18:57 | Northern Mindanao Medical Center | Cagayan de Oro Ciry | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27277 | PBSP | nmmcgenexpress@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 164 |  | GX4 | 804682 | B4 Found Temperature Errors | None | Replacing Module B4 with SN: 671936 to SN: 627039 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836315995468114355/5836315995468114355_signature_28.png | Rione Frances Lomongo | https://www.jotform.com/uploads/macarewecare/222927978372067/5836315995468114355/5836315995468114355_signature_31.png
| 2024-02-13 | 2024-02-13 16:18 | Presentacion Rural Health | Brgy. Pili, Presentacion | 5 | Camarines Sur | RTDL Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27276 | PBSP | tabo.juliet@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 577 |  | GX4-6C 2placer | 834808 | Module B2/699326 for module replacement | Replaced Module: B2 - 699326>>889837 Cleaned module bays, plunger rod, icore, fan Generate Report | Replaced Module B2 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836229795222152535/5836229795222152535_signature_28.png | Julieta G Tabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5836229795222152535/5836229795222152535_signature_31.png
| 2024-02-13 | 2024-02-13 14:45 | San Jose Public Health and Diagnostic Center | San Jose, Occidental Mindoro | 4B | Occidental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27275 | PBSP | mhosanjose2018@gmail.com | 25 |  |  | GXIV | 838197 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module B4 (807346), service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836186894616920645/5836186894616920645_signature_28.png | Reina M. Marasigan | https://www.jotform.com/uploads/macarewecare/222927978372067/5836186894616920645/5836186894616920645_signature_31.png
| 2024-02-13 | 2024-02-13 14:40 | Zaragoza Rural Health Unit I | Del Pilar, Zaragoza, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27274 | PBSP | palapalzarah2425@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 669 | 2024-02-13 | Genexpert IV | 838888 | Genexpert for Calibration | Performed Calibration (All Modules Passed) | Performed Maintenance Gathered Reports Archive Tests Performed Functionality Test | Unit is Operational | N/a | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836158820275321364/5836158820275321364_signature_28.png | Zarah A. Palapal | https://www.jotform.com/uploads/macarewecare/222927978372067/5836158820275321364/5836158820275321364_signature_31.png
| 2024-02-13 | 2024-02-13 14:21 | Ipil rural health unit | Ipil,Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | YES | CALIBRATION MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27273 | PBSP | rheasolana@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1731 ASSAY: Xpert HIV-1 Viral Load, COUNT: 21 | 2024-02-13 | GX 4 | 808615 | Module 2 hardware failed | Replaced module 688497 with Module 640264. Module 1 passed Module 3 service required Module 4 retry. | Module Replacement, Preventive Maintenance and xpertcheck calibration | All 4 module are available | Module 3 failed during xpertcheck | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836150411177582609/5836150411177582609_signature_28.png | RHEAMAY SOLANA | https://www.jotform.com/uploads/macarewecare/222927978372067/5836150411177582609/5836150411177582609_signature_31.png
| 2024-02-13 | 2024-02-13 13:26 | Tomas Claudio Health Center | Camus Street, Corner E. Quirino Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27272 | DOH | arleen_p_alconera@yahoo.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 693 |  | GX-IV 4D | 838492 | AVR not functioning | Replaced the fuse of the AVR Visually check the avr the are burn marks and damaged components | Troubleshoot the AVR | Still no power on the AVR, upon visual inspection. There are black/burned spots on wirings | Pictures will be documented and sent for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836119264427893930/5836119264427893930_signature_28.png | Arleen Alconera RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5836119264427893930/5836119264427893930_signature_31.png
| 2024-02-13 | 2024-02-07 14:15 | Caloocan City Health Office | A. Mabini St., Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27271 | PBSP | caloocancentrallaboratory@gmail.com | 29 |  |  | GX IV 6C - D | 814154 | PC not booting up | Checked CPU Checked and cleaned RAM, ok Checked Bios battery, ok Checked power and data cables, ok Problem persists Network switched 2 GX machine in 1 CPU temporarily | Troubleshoot CPU | CPU for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836117983618645859/5836117983618645859_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5836117983618645859/5836117983618645859_signature_31.png
| 2024-02-13 | 2024-02-13 11:12 | San Sebastian District Hospital | Sablayan, Occidental Mindoro | 4B | Occidental Mindoro | Labotatory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27270 | PBSP | merj3101996@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2494 | 2024-01-31 | GXIV | 812791 | Unit for PM. Unit for Xpertcheck. For module replacement of module B3. | PM, Xpertcheck and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Performed module replacement. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | GX module B3 (210015095) in replacement to defective B3 (210091102). Xpertcheck kit. | Monarch DR 0046883 & 0046784. |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836048783319678947/5836048783319678947_signature_28.png | Merjorie Ann R. Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5836048783319678947/5836048783319678947_signature_31.png
| 2024-02-13 | 2024-02-13 11:16 | Salvacion Oppus-Yñiguez Memorial Provincial Hospital | Maasin City, Southern Leyte | 8 | Southern Leyte | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27269 | PBSP | montillacherry2@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1680 | 2024-02-13 | GXIV | 804262 | Module B1 and B2 defective. | -Replace defective module B1/640300 with module SN: 210004146 -Replace defective module B2/757944 with module SN: 860932 | #NAME? | All modules are functional. |  | 2 GeneXpert module 6-colors 900-0491RS | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836036900324491200/5836036900324491200_signature_28.png | Cherry M Oja | https://www.jotform.com/uploads/macarewecare/222927978372067/5836036900324491200/5836036900324491200_signature_31.png
| 2024-02-13 | 2024-02-13 09:58 | Manila Doctors Hospital | 667 U.N. Ave., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27268 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 21 | ASSAY: Xpert CT/NG, COUNT: 103 ASSAY: Xpert HBV VL, COUNT: 79 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1237 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 93 |  | GX IV 10C-D | 110001648 | Printer not working | Troubleshoot printer Replaced printer HP Deskjet 2335 -> Canon LBP2900 | Checked Printer Data and power cables working Resume Light, power light, and alert lights blinking Restart the printer; problem persists Reset the printer; problem persists Printer beyond repair | Unit is operational. Printer replaced |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836029934317690422/5836029934317690422_signature_28.png | Ryamin F. Buentiempo | https://www.jotform.com/uploads/macarewecare/222927978372067/5836029934317690422/5836029934317690422_signature_31.png
| 2024-02-13 | 2024-02-13 10:12 | Toledo City Health Office | Toledo city | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 27267 | PBSP | loriemaediapen@yahoo.com | 23 |  |  | Gx4 | 808640 | Due for Xpertcheck | Perform Xpertcheck | Perform Xpertcheck | Module B2 H/W Serial Number 7915 Module B1, 3,4 ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5836000575711534399/5836000575711534399_signature_28.png | LORIE MAE DIAPEN | https://www.jotform.com/uploads/macarewecare/222927978372067/5836000575711534399/5836000575711534399_signature_31.png
| 2024-02-13 | 2024-02-13 10:06 | Giporlos Rural Health Unit | Giporlos Eastern Samar | 8 | Eastern Samar | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27266 | PBSP | ningrediang@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 918 | 2024-02-13 | GXIV | 835060 | Module B2/210100869 Defective | #NAME? | - Installed new module on B2 w/ SN 713811 - Performed Functionality test - Performed preventive maintenance ( Cleaning plunger rod, Cartridge bay, Fan Filter) - Generate reports - pullout defective module SN/210100869 | Unit is Operational |  | 1 Genexpert module 6 Color SN/713811 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835999102513349631/5835999102513349631_signature_28.png | Luningning Abuda | https://www.jotform.com/uploads/macarewecare/222927978372067/5835999102513349631/5835999102513349631_signature_31.png
| 2024-02-13 | 2024-02-06 14:36 | Lingap Kalusugan CHU I - Alima | Gen. Evangelista St., Alima, Bacoor, Cavitre | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27265 | PBSP | eng112803@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1732 | 2024-02-05 | GXIV-D-6C | 838185 | Module 3 with Temperature Errors Unit is due for XpertCheck | Module Replacement XpertCheck | Cleaned Cartridge Bay Cleaned Plunger Rod Checked Filter Cleaned Icore Optics Replaced Module 3 Performed Self-test Generate Reports | All Modules are Operational 1 Module Passed, 3 Modules Newly Replaced | B3: 807167 -> 877178 | GeneXpert Module 6-Colors | G 6791 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835988143612680863/5835988143612680863_signature_28.png | Aide E. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5835988143612680863/5835988143612680863_signature_31.png
| 2024-02-13 | 2024-02-06 14:36 | Lingap Kalusugan CHU I - Alima | Gen. Evangelista St., Alima, Bacoor, Cavitre | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27264 | PBSP | eng112803@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1732 | 2024-02-05 | GXIV-D-6C | 838185 | Module 3 with Temperature Errors Unit is due for XpertCheck | Module Replacement XpertCheck | Cleaned Cartridge Bay Cleaned Plunger Rod Checked Filter Cleaned Icore Optics Replaced Module 3 Performed Self-test Generate Reports | All Modules are Operational 1 Module Passed, 3 Modules Newly Replaced | B3: 807167 -> 877178 | GeneXpert Module 6-Colors | G 6791 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835987453616188981/5835987453616188981_signature_28.png | Aide E. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5835987453616188981/5835987453616188981_signature_31.png
| 2024-02-12 | 2024-02-12 17:27 | Amai Pakpak Medical Center RTDL | Marawi City | BARMM | Lanao del Sur | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27263 | PBSP | rbarco2007@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 435 |  | GX-IV | 807457 | FOR MODULE C4 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE C4 REPLACEMENT: DEF SN: 771991 NEW SN:210030509 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835398604426182063/5835398604426182063_signature_28.png | Raisah P. Barco | https://www.jotform.com/uploads/macarewecare/222927978372067/5835398604426182063/5835398604426182063_signature_31.png
| 2024-02-12 | 2024-02-12 16:12 | Zamboanga City Medical Centre | Dr. Evangelista Street, Snta Catalina, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27262 | PBSP | zcmctblab@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 4 ASSAY: Xpert HPV, COUNT: 884 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5547 ASSAY: Xpert SARS-CoV-2, COUNT: 9184 ASSAY: Xpert HCV VL, COUNT: 1 | 2024-02-12 | GX 16 | 839309 | Module B2, B2 and D3 temperature error | Replaced Modules 210068680 , 210068674 and 657350 with modules 877546, 878234 and 880307 respectively. Preventive Maintenance | Replaced Modules 210068680 , 210068674 and 657350 with modules 877546, 878234 and 880307 respectively. Preventive Maintenance | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835364799518451367/5835364799518451367_signature_28.png | Aira Faith Runas | https://www.jotform.com/uploads/macarewecare/222927978372067/5835364799518451367/5835364799518451367_signature_31.png
| 2024-02-12 | 2024-02-12 16:07 | Bacolod City Health Office | Burgos St., Bacolod City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27261 | PBSP | jacabiao.pbsp@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6671 |  | GX4 | 804267 | Module E1 with SN 210068557 and E2 with SN 702067 failed in XPC and need service. | Saved GXX Files, IQR, System Log and Assay Stats. Conduct also PM. | Performed Xpert check calibration for 3 modules module E1, E2 & E4 only. | Modules E3 and E4 passed the calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835348863262269819/5835348863262269819_signature_28.png | Jenzel Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5835348863262269819/5835348863262269819_signature_31.png
| 2024-02-12 | 2024-02-12 15:58 | Bacolod City Health Office | Burgos St., Bacolod City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27260 | PBSP | jacabiao.pbsp@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6671 |  | GX4 | 812020 | Module D4 failed in XPC and need service. | Save GXX File, IQR, System Log and Assay Stats. Conduct also PM. | Run xpert check calibration for 3 modules. Module D1, D2 and D4. | Module D1 & D2 passed in XPC while module D3 is already calibrated and new installed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835343323265477414/5835343323265477414_signature_28.png | Jenzel Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5835343323265477414/5835343323265477414_signature_31.png
| 2024-02-12 | 2024-02-12 13:37 | Ormoc City Health Office | Ormoc City, Leyte | 8 | Leyte | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27259 | PBSP | theresesaladaga@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3141 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3141 | 2024-02-12 | GXIV | 804253 | Module B1 defective | -Replace defective module B1/644830 with module SN: 674245 | #NAME? | All modules are functional. | Waiting for filter allocation. | 900-0491RS | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835259617429234255/5835259617429234255_signature_28.png | Ma. Therese A. Saladaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5835259617429234255/5835259617429234255_signature_31.png
| 2024-02-12 | 2024-02-12 13:37 | Ormoc City Health Office | Ormoc City, Leyte | 8 | Leyte | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27258 | PBSP | theresesaladaga@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3141 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3141 | 2024-02-12 | GXIV | 804253 | Module B1 defective | -Replace defective module B1/644830 with module SN: 674245 | #NAME? | All modules are functional. | Waiting for filter allocation. | 900-0491RS | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835259407422203413/5835259407422203413_signature_28.png | Ma. Therese A. Saladaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5835259407422203413/5835259407422203413_signature_31.png
| 2024-02-12 | 2024-02-07 14:03 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27257 | PBSP | venvalizado@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1252 ASSAY: Xpert HIV-1 Viral Load, COUNT: 134 | 2024-02-07 | GXIV | 834431 | Module D3 for replacement | Replaced D3 | #NAME? | Unit is operational |  | GeneXpert Module SN: 705580 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835135723614215049/5835135723614215049_signature_28.png | Ma. Janeal Mizal | https://www.jotform.com/uploads/macarewecare/222927978372067/5835135723614215049/5835135723614215049_signature_31.png
| 2024-02-12 | 2024-01-30 13:58 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27256 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1293 ASSAY: Xpert HIV-1 Viral Load, COUNT: 221 |  | GX-IV-D-6C | 805211 | Module A3 and A4 has temperature related errors | Performed maintenance | Performed maintenance Checked module A3 and A4 Generated reports | Modules A3 and A4 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835086623614797276/5835086623614797276_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5835086623614797276/5835086623614797276_signature_31.png
| 2024-02-12 | 2024-01-30 14:08 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27255 | MACARE-HBDC | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 318 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835086113619877150/5835086113619877150_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5835086113619877150/5835086113619877150_signature_31.png
| 2024-02-12 | 2024-01-31 10:56 | Bangkal Health Center | Rodriguez Ave, Bangkal | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27254 | PBSP | mhdbangkalsatlab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3349 | 2024-01-29 | GX-IV-D-6C | 838875 | Module B2 for replacement | Replaced module B2 | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Replaced module B2 Generated reports | All modules operational. |  | Old: B2/863410 New: B2/916341 | G 6776 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835085783618512898/5835085783618512898_signature_28.png | Eduardo O. Sepe V | https://www.jotform.com/uploads/macarewecare/222927978372067/5835085783618512898/5835085783618512898_signature_31.png
| 2024-02-12 | 2024-02-07 08:59 | Aguilar Rural Health Unit | Aguilar, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27253 | PBSP | riverabrendalyn@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 507 | 2024-02-07 | GXIV-4-D-6C | 834369 | #NAME? | #NAME? | #NAME? | Unit is operational | All modules passed | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835085103611152418/5835085103611152418_signature_28.png | Brendalyn Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5835085103611152418/5835085103611152418_signature_31.png
| 2024-02-12 | 2024-02-01 11:05 | Pila Rural Health | Rizal St., Pila | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27252 | PBSP | gezza.basarte@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1164 |  | GX-IV-D-6C | 834692 | For XpertCheck, module B1 and B2 has error 1001 | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | Unit is operational. Modules B1 and B3 failed in XpertCheck. CPU for bios battery replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835084773614469353/5835084773614469353_signature_28.png | Gezzamin Basarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5835084773614469353/5835084773614469353_signature_31.png
| 2024-02-12 | 2024-02-01 13:30 | Sta. Maria Rural Health Unit | Sta. Maria | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27251 | PBSP | maidauridorado@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 518 |  | GX-IV-D-6C | 838133 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835084003617164843/5835084003617164843_signature_28.png | Maida Dorado | https://www.jotform.com/uploads/macarewecare/222927978372067/5835084003617164843/5835084003617164843_signature_31.png
| 2024-02-12 | 2024-02-02 13:24 | Nagpayong Super Health Center | Nagpayong | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27250 | DOH | mnoreend@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1374 |  | GX-IV-D-6C | 831905 | Modules B2, B3 and B4 has error 2008 | Run reliability test on module B2 B3 and B4 | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run reliability test on modules B2 B3 and B4 Generated reports | Module B1 not operational. Modules B2 B3 and B4 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835082453614149112/5835082453614149112_signature_28.png | Ma.Noreen L DeLeon | https://www.jotform.com/uploads/macarewecare/222927978372067/5835082453614149112/5835082453614149112_signature_31.png
| 2024-02-12 | 2024-01-30 09:22 | Metro Antipolo Hospital and Medical Center | Marcos Highway, Antipolo | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27249 | COMMERCIAL | mahmci_laboratory@yahoo.com | 24 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 178 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 | 2024-01-29 | GX-IV-D-6C | 110005265 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Run XpertCheck Generated reports | All modules operational. |  | XpertCheck | DR# 11160 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835081723615928639/5835081723615928639_signature_28.png | Sarah Almeria | https://www.jotform.com/uploads/macarewecare/222927978372067/5835081723615928639/5835081723615928639_signature_31.png
| 2024-02-12 | 2024-01-30 13:58 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27248 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1293 ASSAY: Xpert HIV-1 Viral Load, COUNT: 221 |  | GX-IV-D-6C | 805211 | Module A3 and A4 has temperature related errors | Performed maintenance | Performed maintenance Checked module A3 and A4 Generated reports | Modules A3 and A4 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835080923619761214/5835080923619761214_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5835080923619761214/5835080923619761214_signature_31.png
| 2024-02-12 | 2024-01-30 14:03 | Marikina City Health Office | Shoe Ave. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27246 | DOH | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1293 |  | GX-IV-D-6C | 831588 | Module B2 has error 1001 | Performed maintenance | Performed maintenance Checked module B2 Generated reports | Module B2 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835080123614214215/5835080123614214215_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5835080123614214215/5835080123614214215_signature_31.png
| 2024-02-12 | 2024-01-30 14:03 | Marikina City Health Office | Shoe Ave. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27247 | DOH | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1293 |  | GX-IV-D-6C | 831588 | Module B2 has error 1001 | Performed maintenance | Performed maintenance Checked module B2 Generated reports | Module B2 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835080123614257750/5835080123614257750_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5835080123614257750/5835080123614257750_signature_31.png
| 2024-02-12 | 2024-02-08 11:08 | Rodriguez RHU 3 | San isdro,Rodriguez | 4A | Rizal | Tbdots | NO | CALIBRATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27245 | PBSP | medinawenona@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-02-08 | GX4 | 838894 | Due for xpertcheck and maintenance, filter worn out | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Performed xpertcheck Test run unit | Xpertcheck and maintenance | All module passed in xpertcheck |  | Xpertcheck-ce-5 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835079733617492115/5835079733617492115_signature_28.png | Wenona Beers | https://www.jotform.com/uploads/macarewecare/222927978372067/5835079733617492115/5835079733617492115_signature_31.png
| 2024-02-12 | 2024-01-30 14:08 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27244 | MACARE-HBDC | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 318 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835079533611985005/5835079533611985005_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5835079533611985005/5835079533611985005_signature_31.png
| 2024-02-12 | 2024-01-30 14:40 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27243 | MACARE-HBDC | amymitra22@gmail.com | 24 |  |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835079063614246710/5835079063614246710_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5835079063614246710/5835079063614246710_signature_31.png
| 2024-02-12 | 2024-01-31 10:56 | Bangkal Health Center | Rodriguez Ave, Bangkal | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27242 | PBSP | mhdbangkalsatlab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3349 | 2024-01-29 | GX-IV-D-6C | 838875 | Module B2 for replacement | Replaced module B2 | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Replaced module B2 Generated reports | All modules operational. |  | Old: B2/863410 New: B2/916341 | G 6776 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835075943612683626/5835075943612683626_signature_28.png | Eduardo O. Sepe V | https://www.jotform.com/uploads/macarewecare/222927978372067/5835075943612683626/5835075943612683626_signature_31.png
| 2024-02-12 | 2024-01-31 13:18 | MR Flores Health Center | 321 P. Herrera St. | NCR | NCR 4th - Pateros | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27241 | PBSP | paterosclinicallaboratory@gmail.com | 23 |  |  | GX-IV-D-6C | 804254 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck | All modules operational. Modules D1 and D4 passed. Module D2 and D3 newly replaced. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835075373613380604/5835075373613380604_signature_28.png | Jose Mari Zoilo B. De Borja | https://www.jotform.com/uploads/macarewecare/222927978372067/5835075373613380604/5835075373613380604_signature_31.png
| 2024-02-12 | 2024-02-01 11:05 | Pila Rural Health | Rizal St., Pila | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27240 | PBSP | gezza.basarte@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1164 |  | GX-IV-D-6C | 834692 | For XpertCheck, module B1 and B2 has error 1001 | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | Unit is operational. Modules B1 and B3 failed in XpertCheck. CPU for bios battery replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835073633619180061/5835073633619180061_signature_28.png | Gezzamin Basarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5835073633619180061/5835073633619180061_signature_31.png
| 2024-02-12 | 2024-02-01 13:30 | Sta. Maria Rural Health Unit | Sta. Maria | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27239 | PBSP | maidauridorado@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 518 |  | GX-IV-D-6C | 838133 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835072933613516807/5835072933613516807_signature_28.png | Maida Dorado | https://www.jotform.com/uploads/macarewecare/222927978372067/5835072933613516807/5835072933613516807_signature_31.png
| 2024-02-12 | 2024-02-01 15:34 | Siniloan Rural Health Unit | Siniloan | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27238 | PBSP | siniloanruralhealthunit@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 816 |  | GX-IV-D-6C | 838165 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. Module B3 newly replaced. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835072413619129119/5835072413619129119_signature_28.png | Ethel May Merana | https://www.jotform.com/uploads/macarewecare/222927978372067/5835072413619129119/5835072413619129119_signature_31.png
| 2024-02-12 | 2024-02-02 13:24 | Nagpayong Super Health Center | Nagpayong | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27237 | DOH | mnoreend@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1374 |  | GX-IV-D-6C | 831905 | Modules B2, B3 and B4 has error 2008 | Run reliability test on module B2 B3 and B4 | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run reliability test on modules B2 B3 and B4 Generated reports | Module B1 not operational. Modules B2 B3 and B4 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835070953612693737/5835070953612693737_signature_28.png | Ma.Noreen L DeLeon | https://www.jotform.com/uploads/macarewecare/222927978372067/5835070953612693737/5835070953612693737_signature_31.png
| 2024-02-12 | 2024-02-08 13:45 | Pagsanjan Rural Health Unit | Pagsanjan | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27236 | PBSP | dotspagsanjan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 423 ASSAY: Xpert HIV-1 Viral Load, COUNT: 141 |  | GX-IV-D-6C | 807568 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835044963614977759/5835044963614977759_signature_28.png | Lei Arnie Abrigo | https://www.jotform.com/uploads/macarewecare/222927978372067/5835044963614977759/5835044963614977759_signature_31.png
| 2024-02-12 | 2024-02-08 14:43 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27235 | DOH | gemmponce11@gmail.com | 24 |  |  | GX-IV-D-6C | 821292 | CPU has no power. | Troubleshoot CPU | Checked CPU Checked UPS Checked connections Reseat internal connections Checked voltage of UPS Swap power cable | Unit is not operational. Possible defective power supply of CPU. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835043583614446360/5835043583614446360_signature_28.png | Gemma Ponce | https://www.jotform.com/uploads/macarewecare/222927978372067/5835043583614446360/5835043583614446360_signature_31.png
| 2024-02-12 | 2024-02-08 13:45 | Pagsanjan Rural Health Unit | Pagsanjan | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27233 | PBSP | dotspagsanjan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 423 ASSAY: Xpert HIV-1 Viral Load, COUNT: 141 |  | GX-IV-D-6C | 807568 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835041633612999527/5835041633612999527_signature_28.png | Lei Arnie Abrigo | https://www.jotform.com/uploads/macarewecare/222927978372067/5835041633612999527/5835041633612999527_signature_31.png
| 2024-02-12 | 2024-02-08 14:43 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27234 | DOH | gemmponce11@gmail.com | 24 |  |  | GX-IV-D-6C | 821292 | CPU has no power. | Troubleshoot CPU | Checked CPU Checked UPS Checked connections Reseat internal connections Checked voltage of UPS Swap power cable | Unit is not operational. Possible defective power supply of CPU. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5835041633619553867/5835041633619553867_signature_28.png | Gemma Ponce | https://www.jotform.com/uploads/macarewecare/222927978372067/5835041633619553867/5835041633619553867_signature_31.png
| 2024-02-12 | 2024-02-07 10:43 | Dr. Cornelio T Martinez Sr. Memorial Hospital | Kiamba, Sarangani | 12 | Sarangani | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27232 | PBSP | mamartinez.pbsp@gmail.com | 28.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: 975 | 2024-02-07 | Gx-IV | 812017 | N/A | Perform Genexpert Xpertcheck/Calibration Genexpert Cleaning Checking of errors Data collection: IQR, System log, Assay statistics report, Archive test and Calibration data | Genexpert machine Preventive maintenance Genexpert machine Xpertcheck/Calibration | 4 operational modules | All modules passed annual Xpertcheck/Calibration | Xpertcheck cartridge CE-5 | Lot 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5834764533488074527/5834764533488074527_signature_28.png | Madelene Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5834764533488074527/5834764533488074527_signature_31.png
| 2024-02-11 | 2024-02-08 16:55 | Estevez Memorial Hospital, Inc. | Legazpi | 5 | Albay | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | reylovendino83@gmail.com |  | 27231 | COMMERCIAL | yer_lovendino@yahoo.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 32 ASSAY: Xpert CT/NG, COUNT: 10 ASSAY: Xpert HBV VL, COUNT: 212 ASSAY: Xpert MTB/RIF Ultra, COUNT: 206 ASSAY: Others, COUNT: 30 ASSAY: Xpert SARS-CoV-2, COUNT: 12 |  | GX4-10C | 110017942 | GX for Calibration | XpertCheck Result: B1/220009435 - Passed B2/220009328 - Passed B3/210090192 - Passed B4/210083433 - Passed Cleaned cartridge bays, plunger rod Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5834610296413679369/5834610296413679369_signature_28.png | Rey Lovendino | https://www.jotform.com/uploads/macarewecare/222927978372067/5834610296413679369/5834610296413679369_signature_31.png
| 2024-02-09 | 2024-02-09 11:09 | Southern Isabela Medical Center | Rosario, Santiago City, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27230 | PBSP | cassflotildesvillanueva@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4493 ASSAY: Xpert HIV-1 Viral Load, COUNT: 192 | 2024-02-09 | Genexpert IV | 805104 | Module A3 error 4008 | Checked IP address Swapped Module A1 and A3 Returned Module A1 and A3 Reset Gatewayboard | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module A3 for Observation | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5832576966811580963/5832576966811580963_signature_28.png | Soneleth Abenoja | https://www.jotform.com/uploads/macarewecare/222927978372067/5832576966811580963/5832576966811580963_signature_31.png
| 2024-02-08 | 2024-02-08 13:25 | Wenceslao Dela Paz Health Center | Apollo St., Tandang Sora, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27229 | DOH | ameliafclemente2001@yahoo.com | 29 |  |  | GX IV 6C -D | 831876 | Unit due for maintenance | Checked and cleaned GX Parts *cartridge bays *plunger rod *icore slot *filter | Performed preventive maintenance Generated report | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831911030715705838/5831911030715705838_signature_28.png | Camille Joy Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5831911030715705838/5831911030715705838_signature_31.png
| 2024-02-08 | 2024-02-08 13:30 | Balagtas Rural Health Unit | Balagtas, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27228 | WHO | jessaleisantiago@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1056 ASSAY: Xpert HIV-1 Viral Load, COUNT: 28 | 2024-02-08 | GXIV | 846037 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational All modules passed |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831817223338472482/5831817223338472482_signature_28.png | Jessa Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5831817223338472482/5831817223338472482_signature_31.png
| 2024-02-08 | 2024-02-08 13:46 | Alicia Rural Health Unit | Magsaysay, Alicia, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27227 | DOH | jerayvaldez@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 781 | 2024-02-08 | Genexpert IV | 838521 | Module B3 error 1002 | Performed Xpertcheck on module B3 = Service Required | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module B3 for Replacement | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831808173329757611/5831808173329757611_signature_28.png | Jeslie Raymunda Valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/5831808173329757611/5831808173329757611_signature_31.png
| 2024-02-08 | 2024-02-08 13:16 | Pagsanjan Rural Health Unit | Pagsanjan | 4A | Laguna | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27226 | PBSP | dotspagsanjan@gmail.com | 25 |  |  | GX-IV-D-10C | 110013962 | For Installation | Installed the GeneXpert system | Installed the GeneXpert system -GeneXpert machine -Workstation -Printer -UPS Setup system configuration Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831795634161779804/5831795634161779804_signature_28.png | Lei Arnie Abrigo | https://www.jotform.com/uploads/macarewecare/222927978372067/5831795634161779804/5831795634161779804_signature_31.png
| 2024-02-08 | 2024-02-07 14:48 | Taytay Rural Health Unit 2A | Kadalagahan St., Taytay | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27223 | PBSP | dannelli11@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1449 |  | GX-IV-D-6C | 812836 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831792164164072589/5831792164164072589_signature_28.png | Marissa M Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5831792164164072589/5831792164164072589_signature_31.png
| 2024-02-08 | 2024-02-06 14:00 | Santolan Super Health Center | Evangelista Ave., Santolan | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27224 | PBSP | samandlet17@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3639 |  | GX-IV-D-6C | 838898 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | Unit is operational. Modules B1, B2 and B4 passed. Module B3 failed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831792164164640953/5831792164164640953_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5831792164164640953/5831792164164640953_signature_31.png
| 2024-02-08 | 2024-02-07 10:27 | Sumilang Super Health Center | Dr. Garcia, Sumilang | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27225 | PBSP | sumilangshc@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 516 |  | GX-IV-D-6C | 814163 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. Modules B1, B2 and B3 passed. Module B4 failed. Monitor has dead pixels on the lower part of the screen, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831792164169147257/5831792164169147257_signature_28.png | Myrna Lising | https://www.jotform.com/uploads/macarewecare/222927978372067/5831792164169147257/5831792164169147257_signature_31.png
| 2024-02-08 | 2024-02-06 10:34 | Rosario Super Health Center | Sixto Avenue, Rosario | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27222 | WHO | eden.cawayan@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2410 |  | GX-IV-D-6C | 843924 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay Cleaned plunger rod Cleaned instrument surface Cleaned icore Cleaned fan filter Run XpertCheck Generated reports | All modules operational. Modules B2, B3 and b4 passed. Module B1 failed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831792154164502960/5831792154164502960_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5831792154164502960/5831792154164502960_signature_31.png
| 2024-02-08 | 2024-02-07 08:59 | Aguilar Rural Health Unit | Aguilar, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27221 | PBSP | riverabrendalyn@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 507 | 2024-02-07 | GXIV-2-D-6C | 834369 | #NAME? | #NAME? | #NAME? | Unit is operational | All modules passed | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831770962047817059/5831770962047817059_signature_28.png | Brendalyn Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5831770962047817059/5831770962047817059_signature_31.png
| 2024-02-08 | 2024-02-08 11:52 | Sta Rosa City Health Office 1 | City Health Office 1 Bldg., Sta. Rosa Laguna | 4A | Laguna | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27219 | COMMERCIAL | csrlcho1covidlab2@gmail.com | 25 |  |  | GXIV | 838519 | Unit for PM. | PM. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831754058323340177/5831754058323340177_signature_28.png | Mariza B. Garing | https://www.jotform.com/uploads/macarewecare/222927978372067/5831754058323340177/5831754058323340177_signature_31.png
| 2024-02-08 | 2024-02-08 11:54 | Sta Rosa City Health Office 1 | City Health Office 1 Bldg., Sta. Rosa Laguna | 4A | Laguna | Molecular laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27218 | COMMERCIAL | csrlcho1covidlab2@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 468 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 | 2024-01-30 | GXIV | 838522 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Modules B1 (644155) and B4 (769312), service required. |  | Xpertcheck kit. | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831754058325121706/5831754058325121706_signature_28.png | Mariza B. Garing | https://www.jotform.com/uploads/macarewecare/222927978372067/5831754058325121706/5831754058325121706_signature_31.png
| 2024-02-08 | 2024-02-08 11:49 | Sta Rosa City Health Office 1 | City Health Office 1 Bldg., Sta. Rosa Laguna | 4A | Laguna | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27220 | COMMERCIAL | csrlcho1covidlab2@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 64 ASSAY: Xpert SARS-CoV-2, COUNT: 2572 ASSAY: Xpert HIV-1 Viral Load, COUNT: 48 |  | GX16 | 844981 | Unit for PM. | PM. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831754068325202121/5831754068325202121_signature_28.png | Mariza B. Garing | https://www.jotform.com/uploads/macarewecare/222927978372067/5831754068325202121/5831754068325202121_signature_31.png
| 2024-02-08 | 2024-02-08 10:53 | Sta Rosa City Health Office 1 | City Health Office 1 Bldg B., Sta Rosa Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27217 | PBSP | santarosacho1clinlab@gmail.com | 25 |  | 2024-01-31 | GXIV | 804839 | Unit for PM. Unit for Xpertcheck. For module replacement on B2. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Performed Xpertcheck. Generated reports. | Unit is operational. |  | Xpertcheck kkt. GX module B2 (802611) in replacement to defective B2 (600006). | G 6806 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831711121224333911/5831711121224333911_signature_28.png | ALONA NINA L. ALVARO | https://www.jotform.com/uploads/macarewecare/222927978372067/5831711121224333911/5831711121224333911_signature_31.png
| 2024-02-08 | 2024-02-08 10:40 | Manila Doctors Hospital | 667 UN Avenue, Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | xcaguas@maniladoctorscomph.onmicrosoft.com |  | 27216 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 20 | ASSAY: Xpert CT/NG, COUNT: 105 ASSAY: Xpert HPV, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1229 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 94 | 2024-02-07 | GX IV 10C - D | 110001648 | Module B3 for replacement | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Replaced module B3 Pulled out wrong module (6C) Generated report | Replaced Module B3 210010361 -> 210111218 | Unit is operational All modules working |  | GX Module 10C | DR 11192 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831709666137856176/5831709666137856176_signature_28.png | Ryamin F. Buentiempo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5831709666137856176/5831709666137856176_signature_31.png
| 2024-02-08 | 2024-02-08 09:04 | Binan City Health Office 1 | Brgy. Sto. Domingo, Binan Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27215 | PBSP | annefriam.almazora@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 149 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2527 | 2024-01-31 | GXIV | 812780 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module B1 (886541), service required. |  | Xpertcheck kit. | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831653849129556515/5831653849129556515_signature_28.png | Annefriam Almazora | https://www.jotform.com/uploads/macarewecare/222927978372067/5831653849129556515/5831653849129556515_signature_31.png
| 2024-02-08 | 2024-02-08 08:54 | Appari Provincial Hospital | Appari, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27214 | PBSP | glorivimaruquin16@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3121 | 2024-02-08 | Genexpert IV | 808632 | Filter broken | N/a | Performed Maintenance Gathered Reports Performed Module Replacement A2(666402 - 806898) A3(666395 - 861254) Performed Functionality Test | Unit is Operational | Filter for replacement | Genexpert Modules - 806898 & 861254 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831653150527683465/5831653150527683465_signature_28.png | Grace De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5831653150527683465/5831653150527683465_signature_31.png
| 2024-02-08 | 2024-02-08 08:37 | Rodriguez rhu (main) | J.P rizal rodriguez rizal | 4A | Rizal | Laboratory | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27213 | PBSP | karlbanal25@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3025 ASSAY: Xpert HIV-1 Viral Load, COUNT: 37 | 2024-02-08 | GX4 | 807580 | For rexpertcheck B2 and B3, barcode scanner malfunction of trigger during press the button | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Performed rexpertcheck Performed barcode troubleshooting Test run unit | Rexpertcheck and troubleshooting | B2 failed xpertcheck B3 passed xpertcheck |  | Xpertcheck-ce-5 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831623364223013004/5831623364223013004_signature_28.png | Karl Michael M. Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5831623364223013004/5831623364223013004_signature_31.png
| 2024-02-07 | 2024-02-07 19:58 | Davao de Oro Provincial Hospital - Montevista | Poblacion Montevista | 11 | Davao de Oro | TB-DOTS | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27212 | PBSP | cecillejagurilles1122@gmail.com | 25.9 |  |  | GX IV 4D | 808624 | No Power on CPU System Unit Module C2 reported multiple temperature related errors | Reset CPU System Connections Reseat and clean RAM Change CMOS Battery Replaced the Defective Module SN: 792061 with new Module SN: 661274 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform System Unit troubleshooting, Module Repair and Preventive Maintenance | CPU System Unit still no power, all troubleshooting made was document to be sent to Cepheid for further assessment. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831167231822575333/5831167231822575333_signature_28.png | Cecille Jagurilles RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5831167231822575333/5831167231822575333_signature_31.png
| 2024-02-07 | 2024-02-07 15:22 | East Avenue Medical Center | East Avenue,QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27211 | PBSP | roseannparnes@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 174 | 2024-02-07 | GX4 D | 803087 | Due for xpertcheck and maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Performed xpertcheck Test run unit | Xpertcheck and maintenance | C1 passed in xpertcheck C2, C3,C4 failed in xpertcheck |  | Xpertcheck-ce-5 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831147175421322694/5831147175421322694_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/5831147175421322694/5831147175421322694_signature_31.png
| 2024-02-07 | 2024-02-07 09:22 | Best Diagnostic Corporation | Masikap ext. Central QC | NCR | NCR 2nd - Quezon City | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27210 | COMMERCIAL | laboratory@bestdiagnostic.com.ph | 29 | ASSAY: Xpert CT/NG, COUNT: 149 ASSAY: Xpert HBV VL, COUNT: 67 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 | 2024-01-30 | GX4 | 110008529 | Due for xpertcheck and maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Performed xpertcheck Test run unit | Maintenance and xpertcheck | Operational All module passed in xpertcheck |  | Xpertcheck-ce-5 | 11171 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831145599919056605/5831145599919056605_signature_28.png | Best Diagnostic Corp | https://www.jotform.com/uploads/macarewecare/222927978372067/5831145599919056605/5831145599919056605_signature_31.png
| 2024-02-07 | 2024-02-06 13:19 | Tropical Disease Foundation | Amorsolo st., Makati | NCR | NCR 4th - Makati | Laboratory | NO | Software upgrade | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | aageronimo@tdf.org.ph |  | 27209 | CONSORTIUM | krcueno@tdf.org.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1625 |  | GX4 D | 812560 | For software upgrade unable to link LIS, C3 h/w failed. | Performed software 5.1 to 6.4 Gathered reports. Test run unit | For upgrade | Operational | C3 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831145029919463291/5831145029919463291_signature_28.png | Karl Francis Cueno | https://www.jotform.com/uploads/macarewecare/222927978372067/5831145029919463291/5831145029919463291_signature_31.png
| 2024-02-07 | 2024-02-05 15:59 | Philippine Heart Center | East avenue ,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27208 | PBSP | molecularlab@phc.gov.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1238 ASSAY: Others, COUNT: 16 ASSAY: Xpert SARS-CoV-2, COUNT: 3581 |  | GX4 | 834443 | Workstation unable to power,malfunction of on/off switch corroded terminal | Checked and measure voltage input. Cleaned RAM Unplug and Plug connector inside workstation. Checked cables Cleaned on/off switch terminal Test run and observe unit | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831144359911955091/5831144359911955091_signature_28.png | Casidsid, Kenneth Melvin | https://www.jotform.com/uploads/macarewecare/222927978372067/5831144359911955091/5831144359911955091_signature_31.png
| 2024-02-07 | 2024-02-07 15:47 | Philippine Red Cross - Ormoc | Ormoc City | 6 | Leyte | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 27207 | COMMERCIAL | ormoc@redcross.org.ph | 28 |  |  | Mcs+ | 15H043G | CPU board using was spare parts of cebu office | Re-installing Original CPU board of PRC ormoc haemonitics machine/update UPP settings configuration/PM and calibration | Re-installing Original CPU board of PRC ormoc haemonitics machine/update UPP settings configuration/PM and calibration | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831022979845882964/5831022979845882964_signature_28.png | Phil. Red Cross Ormoc City | https://www.jotform.com/uploads/macarewecare/222927978372067/5831022979845882964/5831022979845882964_signature_31.png
| 2024-02-07 | 2024-02-07 15:47 | Philippine Red Cross - Ormoc | Ormoc City | 6 | Leyte | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 27206 | COMMERCIAL | ormoc@redcross.org.ph | 28 |  |  | Mcs+ | 15H043G | CPU board using was spare parts of cebu office | Re-installing Original CPU board of PRC ormoc haemonitics machine/update UPP settings configuration/PM and calibration | Re-installing Original CPU board of PRC ormoc haemonitics machine/update UPP settings configuration/PM and calibration | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831021879844584744/5831021879844584744_signature_28.png | Phil. Red Cross Ormoc City | https://www.jotform.com/uploads/macarewecare/222927978372067/5831021879844584744/5831021879844584744_signature_31.png
| 2024-02-07 | 2024-02-07 15:47 | Philippine Red Cross - Ormoc | Ormoc City | 6 | Leyte | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 27205 | COMMERCIAL | ormoc@redcross.org.ph | 28 |  |  | Mcs+ | 15H043G | CPU board using was spare parts of cebu office | Re-installing Original CPU board of PRC ormoc haemonitics machine/update UPP settings configuration/PM and calibration | Re-installing Original CPU board of PRC ormoc haemonitics machine/update UPP settings configuration/PM and calibration | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831021309849851520/5831021309849851520_signature_28.png | Phil. Red Cross Ormoc City | https://www.jotform.com/uploads/macarewecare/222927978372067/5831021309849851520/5831021309849851520_signature_31.png
| 2024-02-07 | 2024-02-07 15:07 | Pilar Rural Health Unit | Calongay, Pilar | 5 | Sorsogon | RTDL Laboratory | YES | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27204 | DOH | manalangmark88@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 255 |  | GX4-6C | 838266 | B1/807683 | Re-XpertCheck Result: B1/807683 - passed | Performed Re-XpertCheck | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5831002806318661467/5831002806318661467_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5831002806318661467/5831002806318661467_signature_31.png
| 2024-02-07 | 2024-02-07 14:50 | Julian Felipe Health Center | Julian Felipe St., Brgy. 8, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27203 | PBSP | maoi.lloyd@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1506 | 2024-02-05 | GX IV 6C - D | 838884 | Module B3 for replacement | Performed PM Replaced module b3 generated report | Replaced Module B3 863557 -> 612274 | Unit is operational All modules working |  | GX module 6C | G6802 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830983261023333538/5830983261023333538_signature_28.png | Lloyd Maoi De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5830983261023333538/5830983261023333538_signature_31.png
| 2024-02-07 | 2024-02-07 14:19 | Batangas Medical Center | Kumintang Ibaba, Batangas City | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27202 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 100 | 2024-01-31 | GXIV | 804683 | Unit for PM. Unit for Xpertcheck. For checking of CPU. | PM, Xpertcheck and troubleshooting. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Troubleshooting of CPU. Generated reports. | Unit is operational. All modules passed Xpertcheck using different CPU. CPU unable to poweron and for replacement. |  | Xpertcheck kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830972888617646967/5830972888617646967_signature_28.png | Patricia Louisse A. Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5830972888617646967/5830972888617646967_signature_31.png
| 2024-02-07 | 2024-02-07 14:16 | Batangas Medical Center | Kumintang Ibaba, Batangas City | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27201 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 267 ASSAY: Xpert SARS-CoV-2, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 60 | 2024-01-31 | GXIV | 834627 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830972398614879468/5830972398614879468_signature_28.png | Patricia Louisse A. Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5830972398614879468/5830972398614879468_signature_31.png
| 2024-02-07 | 2024-02-07 14:11 | Batangas Medical Center | Kumintang Ibaba, Batangas City | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27200 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 67 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2440 ASSAY: Xpert SARS-CoV-2, COUNT: 330 ASSAY: Xpert HIV-1 Viral Load, COUNT: 611 | 2024-01-31 | GXI6 | 844120 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Modules A2 (607813), C3 (892213), C4 (892212) and D4 (680439) are Service Required. |  | Xpertcheck kit. | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830971768615698076/5830971768615698076_signature_28.png | Patricia Louisse A. Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5830971768615698076/5830971768615698076_signature_31.png
| 2024-02-07 | 2024-02-07 14:23 | Batangas Medical Center | Kumintang Ibaba, Batangas City | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27199 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 42 ASSAY: Others, COUNT: XDR 356 | 2024-01-31 | GXIV-10C | 110016955 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck kit. | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830970888615633017/5830970888615633017_signature_28.png | Patricia Louisse A. Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5830970888615633017/5830970888615633017_signature_31.png
| 2024-02-07 | 2024-02-07 14:28 | Batangas Medical Center | Kumintang Ibaba, Batangas City | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27198 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 129 ASSAY: Xpert SARS-CoV-2, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 32 | 2024-01-31 | GXIV | 814153 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module C1 (619967) is service required. |  | Xpertcheck kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830969018616894491/5830969018616894491_signature_28.png | Patricia Louisse A. Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5830969018616894491/5830969018616894491_signature_31.png
| 2024-02-07 | 2024-02-06 09:58 | San Fernando Rural Health Unit 2 | Sindalan San Fernando Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27197 | PBSP | punoricahrd@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1236 |  | GXIV-D | 838879 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational 3 modules passed calibration 1 failed calibration | A4 service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830951433906253190/5830951433906253190_signature_28.png | Richard M. Puno | https://www.jotform.com/uploads/macarewecare/222927978372067/5830951433906253190/5830951433906253190_signature_31.png
| 2024-02-07 | 2024-02-06 09:58 | San Fernando Rural Health Unit 2 | Sindalan San Fernando Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27196 | PBSP | punoricahrd@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1236 |  | GXIV-D | 838879 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational 3 modules passed calibration 1 failed calibration | A4 service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830949913904343133/5830949913904343133_signature_28.png | Richard M. Puno | https://www.jotform.com/uploads/macarewecare/222927978372067/5830949913904343133/5830949913904343133_signature_31.png
| 2024-02-07 | 2024-02-07 11:30 | Olongapo City Health Offic | West Bajac-bajac Olongapo City Zambales | 3 | Zambales | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27195 | PBSP | olongapochotmlandrtdl@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1703 |  | GXIB-D | 804700 | D1:210089680 and D4:721209 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational New modules D1:629603 D4:210053177 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830948673908699919/5830948673908699919_signature_28.png | Roselyn Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5830948673908699919/5830948673908699919_signature_31.png
| 2024-02-07 | 2024-02-07 12:17 | Olongapo City Health Office | West bajac-bajac Olongapo City, Zambales | 3 | Zambales | Laboratory | NO | INSTALLATION | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27194 | PBSP | olongapochotmlandrtdl@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV-D 10C | 120000606 | Genexpert 10C installation | Installation | Installation and setup of Genexpert Machine Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830947053907477855/5830947053907477855_signature_28.png | Roselyn M. Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5830947053907477855/5830947053907477855_signature_31.png
| 2024-02-07 | 2024-02-07 13:07 | Amma Jadsac District Hospital | Poblacion, Pudtol, Apayao | CAR | Apayao | Laboratory | NO | INSTALLATION | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27193 | PBSP | dlgblacar.pbsp@gmail.com | 24 |  | 2024-02-07 | Genexpert IV-10c | 120000598 | N/a | N/a | Installed Genexpert Installed Computer Installed Barcode Scanner Install Printer Installed Monitor Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830918718455069368/5830918718455069368_signature_28.png | Daisy Lynn Grace | https://www.jotform.com/uploads/macarewecare/222927978372067/5830918718455069368/5830918718455069368_signature_31.png
| 2024-02-07 | 2024-02-07 12:51 | Amma Jadsac District Hospital | Poblacion, Pudtol, Apayao | CAR | Apayao | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27192 | PBSP | dlgblacar.pbsp@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3415 | 2024-02-07 | Genexpert IV | 807609 | Module 4 error 2008 | N/a | Performed Maintenance Performed Module Replacement A1(210102203 - 784265) Archive Results Performed Functionality Test | Unit is Operational | N/a | Genexpert Module - 784265 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830913038459844616/5830913038459844616_signature_28.png | Daisy Lynn Grace B. Lacar | https://www.jotform.com/uploads/macarewecare/222927978372067/5830913038459844616/5830913038459844616_signature_31.png
| 2024-02-07 | 2024-02-07 12:18 | Naic Rural Health Unit | Capt. Nazareno St., Naic, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27190 | PBSP | jengcasamar2007@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2243 |  | GXIV-D-6C | 807567 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay Cleaned Icore Optics Cleaned Plunger Rod Checked Filter Performed XpertCheck Performed Self-test Generate Reports | All Modules Passed and Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830894779142624616/5830894779142624616_signature_28.png | Jennifer S Casamar RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5830894779142624616/5830894779142624616_signature_31.png
| 2024-02-07 | 2024-02-07 12:22 | Naic Rural Health Unit | Capt. Nazareno St, Naic, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27191 | PBSP | jengcasamar2007@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2243 |  | GXIV-D-6C | 812006 | Unit is due for XpertCheck | XpertCheck | Cleaned Cartridge Bay Cleaned Icore Optics Cleaned Plunger Rod Checked Filter Performed XpertCheck Performed Self-test Generate Reports | All Modules Passed and Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830894779149501679/5830894779149501679_signature_28.png | Jennifer Casamar | https://www.jotform.com/uploads/macarewecare/222927978372067/5830894779149501679/5830894779149501679_signature_31.png
| 2024-02-07 | 2024-02-07 11:56 | Quezon City Jail | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27189 | PBSP | carleecruz0@gmail.com | 25 |  | 2024-02-07 | GXIV | 803089 | Module A2 for replacement | Replaced B2 | #NAME? | Unit is operational |  | GeneXpert Module SN: 760346 | na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830875581379983626/5830875581379983626_signature_28.png | Carlee Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5830875581379983626/5830875581379983626_signature_31.png
| 2024-02-07 | 2024-02-06 15:10 | Ospital ng Lungsod ng San Jose Del Monte | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27188 | PBSP | krizetteamorio@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1847 ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-02-05 | GXIV | 807623 | Module B1 for replacement Unit for XpertCheck | Replaced module B1 Performed XpertCheck | #NAME? | Unit is operational B1: newly replaced B2: service required B3: passed B4: passed |  | GeneXpert Module SN: 764899 | G 6804 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830873931371685499/5830873931371685499_signature_28.png | Krizette Marie Amorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5830873931371685499/5830873931371685499_signature_31.png
| 2024-02-07 | 2024-02-07 11:39 | Western Visayan Medical Center | Iloilo City | 6 | Iloilo | TB Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27187 | PBSP | mitorilla.pbsp@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 988 ASSAY: Xpert MTB/RIF Ultra, COUNT: 818 ASSAY: Xpert HIV-1 Viral Load, COUNT: 485 |  | GX4 R1 | 706910 | Defective module A1/639353 | Preventive maintenance | Replaced module A1 with mew module SN 895555 | All 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830868701348619551/5830868701348619551_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5830868701348619551/5830868701348619551_signature_31.png
| 2024-02-07 | 2024-02-07 11:32 | Western Visayan Medical Center | Iloilo City | 6 | Iloilo | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 27186 | PBSP | mitorilla.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4637 |  | GX 16 | 839316 | Defective module D3/871127. | Preventive Maintenance | Replaced module D3 with new module SN 839377. | All 16 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830862081343892138/5830862081343892138_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5830862081343892138/5830862081343892138_signature_31.png
| 2024-02-07 | 2024-02-07 10:25 | Mabini Rural Health Unit | Poblacion Mabini Batangas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27185 | PBSP | glee.maravilla@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1024 | 2024-01-31 | GXIV | 838867 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. |  | Xpertcheck Kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830828527011893608/5830828527011893608_signature_28.png | Gliceria C. Maravilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5830828527011893608/5830828527011893608_signature_31.png
| 2024-02-07 | 2024-02-06 16:07 | Palimbang Rural Health Unit | Poblacion, Palimbang, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27184 | WHO | tinavillegas1910@gmail.com | 23.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 430 ASSAY: Xpert MTB/RIF Ultra, COUNT: 430 ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-02-06 | Gx-IV | 847257 | Module 3 H/W failed Module 1 for Service Required | Replace/repair defective module 3 SN:910253 with new module SN:896614 Genexpert Cleaning Perform Calibration at module 1,2 and 4 Data collection: Pre and Post IQR, System log, Assay statistics report and Archive test | Perform Genexpert machine Preventive maintenance Perform Genexpert machine module replacement/repair Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules |  | Genexpert module 6 color Xpertcheck CE-5 | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830820415456962417/5830820415456962417_signature_28.png | Cristina Nelda Villegas | https://www.jotform.com/uploads/macarewecare/222927978372067/5830820415456962417/5830820415456962417_signature_31.png
| 2024-02-07 | 2024-02-07 08:43 | GMA Rural Health Unit II | Congressional Road, General Mariano Alvarez, Cavite | 4A | Cavite | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27183 | PBSP | ldlconti@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1719 |  | GXIV-D-6C | 814152 | UPS Problem | Troubleshoot | Checked Outlet Voltage Checked Output Voltage of UPS Performed Self-test Generate Reports | UPS is Defective. No Power and Backup Function. UPS for Replacement | UPS S/N: 20002103290032 .All Modules are Operational |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772903388266069/5830772903388266069_signature_28.png | Lorena L.Conti | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772903388266069/5830772903388266069_signature_31.png
| 2024-02-07 | 2024-02-01 08:13 | Tablas Island District Hospital | TIDH Rd., San Agustin, Romblon | 4B | Romblon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27182 | PBSP | nelda.grefiel@yahoo.com.ph | 24 |  |  | GXIV-2-D-6C | 835088 | CPU is Not Booting | CPU Replacement | Performed Maintenance Replaced CPU Setup GeneXpert Software Generate Reports | All Modules are Operational. New CPU is Good | CPH1119F04197 -> CPH1023F45887 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713383463796/5830772713383463796_signature_28.png | Nelda S. Grefiel | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713383463796/5830772713383463796_signature_31.png
| 2024-02-07 | 2024-01-30 13:30 | Pulang Lupa Dos Health Center | St. Joseph Subd, Pulang Lupa Dos, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27181 | PBSP | elvimir08@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1775 | 2024-01-29 | GXIV-D-6C | 838183 | Modukes 1 and 2 with Temperature Errors | Module Replacement | Performed Maintenance Replaced Modules 1 and 2 Performed Self-test Generate Reports | All Modules are Operational | 807075&807118 -> 866511&210067976 | GX Module | G 6777 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713384679361/5830772713384679361_signature_28.png | Elvira alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713384679361/5830772713384679361_signature_31.png
| 2024-02-07 | 2024-01-25 16:24 | San Luis Rural Health Unit | Poblacion, San Luis, Batangas | 4A | Batangas | Laboratory | NO | MAINTENANCE Installation of Backup Laptop | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27180 | PBSP | sanluisbatangas.rhu@gmail.com | 24 |  | 2024-01-24 | GXIV-D-6C | 838131 | CPU of GX System Not Booting Up | Installation of Backup Laptop | Checked CPU Problem Installed Back-up Laptop Generate Reports | All Modules are Operational |  | GX Laptop L480 | G 6762 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713385201560/5830772713385201560_signature_28.png | Rhoda Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713385201560/5830772713385201560_signature_31.png
| 2024-02-07 | 2024-01-23 09:53 | UST Hospital | Dapitan St, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27178 | MACARE-HBDC | usth_bloodbank@yahoo.com | 24 |  |  | MCS+ | 02F069 | Unit is due for Maintenance | Preventive Maintenance | Checked and Ckeaned Sensors, Rollers, and Centrifuge Checked A/D Values Performed Simulation Test Cleaned Surface Area of the Unit | Unit is in Good Condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713385390171/5830772713385390171_signature_28.png | Mylene Balboa | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713385390171/5830772713385390171_signature_31.png
| 2024-02-07 | 2024-02-05 14:01 | UST Hospital | Dapitan St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27179 | CONSORTIUM | ustimp@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 869 |  | GXIV-D-6C | 1100000711 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartirdge Bay Cleaned Plunger Rod Cleaned Icore Optics Checked Filter if Clogged Archived Test Results Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713385804587/5830772713385804587_signature_28.png | Maria Liova Sherleen D. Calaunan,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5830772713385804587/5830772713385804587_signature_31.png
| 2024-02-06 | 2024-02-05 14:01 | UST Hospital | Dapitan St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27176 | CONSORTIUM | ustimp@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 869 |  | GXIV-D-6C | 1100000711 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartirdge Bay Cleaned Plunger Rod Cleaned Icore Optics Checked Filter if Clogged Archived Test Results Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830227442157459422/5830227442157459422_signature_28.png | Maria Liova Sherleen D. Calaunan,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5830227442157459422/5830227442157459422_signature_31.png
| 2024-02-06 | 2024-01-23 09:53 | UST Hospital | Dapitan St, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27175 | MACARE-HBDC | usth_bloodbank@yahoo.com | 24 |  |  | MCS+ | 02F069 | Unit is due for Maintenance | Preventive Maintenance | Checked and Ckeaned Sensors, Rollers, and Centrifuge Checked A/D Values Performed Simulation Test Cleaned Surface Area of the Unit | Unit is in Good Condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830227432155495886/5830227432155495886_signature_28.png | Mylene Balboa | https://www.jotform.com/uploads/macarewecare/222927978372067/5830227432155495886/5830227432155495886_signature_31.png
| 2024-02-06 | 2024-02-01 08:13 | Tablas Island District Hospital | TIDH Rd., San Agustin, Romblon | 4B | Romblon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27177 | PBSP | nelda.grefiel@yahoo.com.ph | 24 |  |  | GXIV-2-D-6C | 835088 | CPU is Not Booting | CPU Replacement | Performed Maintenance Replaced CPU Setup GeneXpert Software Generate Reports | All Modules are Operational. New CPU is Good | CPH1119F04197 -> CPH1023F45887 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830227442154796256/5830227442154796256_signature_28.png | Nelda S. Grefiel | https://www.jotform.com/uploads/macarewecare/222927978372067/5830227442154796256/5830227442154796256_signature_31.png
| 2024-02-06 | 2024-02-06 16:34 | Rapu-Rapu RHU | Rapu-Rapu | 5 | Albay | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27174 | DOH | suzetteoguan1@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 351 |  | GX4-6C | 838565 | B2/737128, B3/862633 for module replacement | Replaced Module: B2 - 737128>>730958 B3 - 862633>>757493 Cleaned fan, module bays, plunger rod Generate reports | Replaced defective module B2, B3 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830181420914863234/5830181420914863234_signature_28.png | Suzette C. Oguan | https://www.jotform.com/uploads/macarewecare/222927978372067/5830181420914863234/5830181420914863234_signature_31.png
| 2024-02-06 | 2024-02-06 09:09 | Davao del Sur Provincial Hospital | Poblacion, Digos City | 11 | Davao del Sur | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27173 | PBSP | shammyj0308@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: 3141 |  | GX-IV 4D | 807464 | Module A3 encounter multiple temperature related errors | Replaced the defective Module SN: 672526 with new Module SN: 632623 Generate IQR, Assay Stat, and System Log Clean Surface area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Archive data test result | Perform Preventive Maintenance and Module Replacement | All Modules are operational. | Module A4 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830175380124237796/5830175380124237796_signature_28.png | Sharmella Joy D. Pinote | https://www.jotform.com/uploads/macarewecare/222927978372067/5830175380124237796/5830175380124237796_signature_31.png
| 2024-02-06 | 2024-02-06 09:58 | San Fernando Rural Health Unit 2 | Sindalan San Fernando Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27172 | PBSP | punoricahrd@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1236 |  | GXIV-D | 838879 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational 3 modules passed calibration 1 failed calibration | A4 service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830140821413860694/5830140821413860694_signature_28.png | Richard M. Puno | https://www.jotform.com/uploads/macarewecare/222927978372067/5830140821413860694/5830140821413860694_signature_31.png
| 2024-02-06 | 2024-02-06 15:10 | Klinika Bernardo Health Center | Quezon City Health Department, Quezon City Hall | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27171 | PBSP | klinikabernardolaboratory@gmail.com | 22 |  | 2024-01-30 | GX IV 6C - D | 808618 | Unit for XpertCheck Module B3 and B4 for XpertCheck retry Module B1 and B2 passed January 25, 2024 | Performed PM Performed XpertCheck Module 3 - Service Required Module 4 - Passed | Performed Xpert Check on Modules 3 and 4 Module 3 - Service Required Module 4 - Passed Modules 1 and 2 passed on January 25, 2024 | Unit is operational Module B3 for reporting | Used 3 XpertCheck cartridges | XpertCheck Kit | PL 014900 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830135564623953129/5830135564623953129_signature_28.png | Mylen Sibal | https://www.jotform.com/uploads/macarewecare/222927978372067/5830135564623953129/5830135564623953129_signature_31.png
| 2024-02-06 | 2024-02-06 14:42 | Masinloc Rural Health Unit | Masinloc Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27170 | DOH | rodenne@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2210 |  | GXIV-D | 831608 | Temperature error on module A3 | Xoertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | A3 passed calibration under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830127296221976446/5830127296221976446_signature_28.png | Rodenn Eduvala | https://www.jotform.com/uploads/macarewecare/222927978372067/5830127296221976446/5830127296221976446_signature_31.png
| 2024-02-06 | 2024-02-06 13:11 | Amparo Health Center | Amparo, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27169 | PBSP | jjpsjasperjoyce@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1775 | 2024-02-05 | GXIV | 838873 | Modules B1, B2 for replacement | Replaced modules B1, B2 | #NAME? | Unit is operational |  | GeneXpert Module SN: 706384, 210035367 | G 6793 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830068413417013165/5830068413417013165_signature_28.png | Jasper Joyce Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/5830068413417013165/5830068413417013165_signature_31.png
| 2024-02-06 | 2024-02-06 13:00 | Cagayan Valley Medical Center | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27168 | PBSP | zionclydematote@gmail.com | 27 |  | 2024-02-06 | Genexpert XVI | 842894 | Module B3 Hardware Failure | N/a | Performed Maintenance Gathered Reports Performed Calibration (A1, A3, A4 = Service Required) (B1, B2, B4, C2, C3, D1, D2, D3, D4 = Passed) Performed Functionality Test | Unit is Operational | Module A1, A3, A4, B3 for Replacement | Xpertcheck 15pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830066072147651555/5830066072147651555_signature_28.png | Olive Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5830066072147651555/5830066072147651555_signature_31.png
| 2024-02-06 | 2024-02-06 13:03 | Belmonte Health Center | Don Quijote St. Sampaloc Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27167 | PBSP | spsurell@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1217 | 2024-02-05 | GXIV | 814376 | Unit for PM. For module replacement of module A3. | PM and module replacement. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed module replacement. Generated reports. | Unit is operational. |  | GX module A3 (210063191) in replacement to defective A3 (684928). | G 6794 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830055270169527459/5830055270169527459_signature_28.png | Clarinda Sharon P.Surell | https://www.jotform.com/uploads/macarewecare/222927978372067/5830055270169527459/5830055270169527459_signature_31.png
| 2024-02-06 | 2024-02-06 11:46 | Cagayan Valley Center for Health Development | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27166 | PBSP | mctorio.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 457 ASSAY: Others, COUNT: 96 | 2024-02-06 | Genexpert IV - 10c | 120000178 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement B3(220011645 - 210074806) Performed Functionality Test | Unit is Operational | N/a | Genexpert Module - 210074806 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830009793623214648/5830009793623214648_signature_28.png | Mikeala Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5830009793623214648/5830009793623214648_signature_31.png
| 2024-02-06 | 2024-02-06 11:43 | Bago City Health Office | Bago City, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 27165 | PBSP | swabteambago@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2386 |  | GX4 | 815433 | Defective module C2/210100867. | Preventive Maintenance - optics cleaning, module bay cleaning, filter cleaning. | Replaced module C2 with new module SN 877174. | 3 modules operational, 1 module (C1) for repoeting to Cepheid due to temperature-related errors. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830008242033047489/5830008242033047489_signature_28.png | GLenly CAsona | https://www.jotform.com/uploads/macarewecare/222927978372067/5830008242033047489/5830008242033047489_signature_31.png
| 2024-02-06 | 2024-02-06 11:16 | Bureau of Quarantine | 25th St., Intramuros Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27164 | COMMERCIAL | dklo.lab@quarantine.doh.gov.ph | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 279 | 2024-01-26 | GX16 | 844418 | Unit for PM. Unit for Xpertcheck. | PM and Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | Unit is operational. Module D2 (893186), Service Required. |  | Xpertcheck Kit | - |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5830007290169078777/5830007290169078777_signature_28.png | Dulce Karmela Lo | https://www.jotform.com/uploads/macarewecare/222927978372067/5830007290169078777/5830007290169078777_signature_31.png
| 2024-02-06 | 2024-02-06 09:47 | Banlat Health Center | Daisy St., cor Daffodil St., Banlat, Tandang Sora, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27163 | PBSP | reyzelmariesantos@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 113 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1701 | 2024-01-26 | GX IV 6C - D | 804688 | Unit due for XpertCheck | Performed PM Performed XpertCheck Module A1 service required Generated report | Performed XpertCheck Module A1 failed Module A2, 3,4 passed | Unit is operational Module A1 for reporting | Used 5 cartridges | XpertCheck Kit | DR 0046564 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829970666886952335/5829970666886952335_signature_28.png | Reyzel Marie S.Dulay | https://www.jotform.com/uploads/macarewecare/222927978372067/5829970666886952335/5829970666886952335_signature_31.png
| 2024-02-05 | 2024-02-05 17:02 | Nuestra Sra. De Piat District Hospital | Maguilling, Piat, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27162 | PBSP | mariagenevaligas@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 356 | 2024-02-05 | Genexpert IV | 835091 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement B2(210091136 - 210059766) Performed Functionality Test | Unit is Operational | N/a | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829331719982213772/5829331719982213772_signature_28.png | Maria Geneva M. Ligas | https://www.jotform.com/uploads/macarewecare/222927978372067/5829331719982213772/5829331719982213772_signature_31.png
| 2024-02-05 | 2024-02-05 15:53 | Paniqui Super Health Center | Paniqui Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27161 | PBSP | forylab76@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 164 |  | GXIV-D | 812792 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational 4 modules passed calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829304142911641972/5829304142911641972_signature_28.png | Fortune Abutin | https://www.jotform.com/uploads/macarewecare/222927978372067/5829304142911641972/5829304142911641972_signature_31.png
| 2024-02-05 | 2024-02-05 13:06 | Talomo North Health Center | Sandawa, S.I.R Matina, Davao City | 11 | Davao del Sur | TB-DOTS | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27160 | PBSP | alfredojrsantos@gmail.com | 27.5 | ASSAY: , COUNT: 2091 |  | GX-IV 4D | 815385 | Module B3 encounter temperature error | Replaced the defective Module SN: 625094 with new Module SN: 806700 Clean Surface area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run Xpertcheck on Modules 1,2 and 4 | Perform Preventive Maintenance, Module Replacement and XpertCheck Calibration | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829284669229944753/5829284669229944753_signature_28.png | Alfredo Santos Jr. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5829284669229944753/5829284669229944753_signature_31.png
| 2024-02-05 | 2024-02-05 13:37 | Tarlac Provincial Hospital | Hospital Drive Brgy. San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27159 | PBSP | molecularlab.tph@gmail.com | 26 | ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2164 ASSAY: Xpert HIV-1 Viral Load, COUNT: 546 |  | GXIV-D | 805108 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except module C1(service required) | 3 modules passed calibration 1 failed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829231953319348292/5829231953319348292_signature_28.png | Ethel Mae D. Avellanosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5829231953319348292/5829231953319348292_signature_31.png
| 2024-02-05 | 2024-02-05 13:42 | Tarlac Provincial Hospital | Hospital drive Brgy. San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27158 | PBSP | molecularlab.tph@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 2377 |  | GXIV-D | 838142 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational 2 modules passed calibration 2 newly replaced module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829231043312835705/5829231043312835705_signature_28.png | Ethel Mae D. Avellanosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5829231043312835705/5829231043312835705_signature_31.png
| 2024-02-05 | 2023-12-06 13:27 | Tarlac Provincial Hospital | Hospital Drive Brgy. San Vicente Tarlac City | 3 | Tarlac | PCR | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27157 | PBSP | genextphmpl@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 1651 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1196 |  | GXIV-D | 838142 | B3:909438 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replavement Gathered reports Performed functionality test | Unit is operational | New B3 SN:672264 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829201335128254788/5829201335128254788_signature_28.png | Ethel Mae D. Avellanosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5829201335128254788/5829201335128254788_signature_31.png
| 2024-02-05 | 2024-02-05 11:27 | Love Yourself | Cebu City | 7 | Cebu | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 27156 | COMMERCIAL | preforzado@loveyourself.ph | 24 |  | 2024-01-01 | Gx4 | 814150 | Error code 2096 all modules | Perform expert check | Clean all modules and perform Expert check for all modules. | Module 1,2,3 ok to use and passed for expert check. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829132221119593311/5829132221119593311_signature_28.png | Paulo Crispo Reforzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5829132221119593311/5829132221119593311_signature_31.png
| 2024-02-05 | 2024-02-02 08:41 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27155 | PBSP | romarmtamayo@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2202 ASSAY: Xpert SARS-CoV-2, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 122 | 2024-02-02 | GXIV-4-D-6C | 804726 | For XpertCheck and Preventive Maintenance | #NAME? | Performed Preventive Maintenance and XpertCheck | Unit is operational | All Modules passed XpertCheck | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829053652041321102/5829053652041321102_signature_28.png | Romar M. Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5829053652041321102/5829053652041321102_signature_31.png
| 2024-02-05 | 2024-02-02 10:07 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27154 | PBSP | r1mcxpertlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1417 ASSAY: Xpert SARS-CoV-2, COUNT: 476 | 2024-02-02 | GX 16 | 842230 | For Preventive Maintenance and XpertCheck | #NAME? | Performed Preventive Maintenance and XpertCheck | Unit is operational | Module B3 recently installed, 15 modules for XpertCheck. 14 modules passed Module; D2 for reporting | XpertCheck Kits | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5829051252044731864/5829051252044731864_signature_28.png | Efren Notario Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5829051252044731864/5829051252044731864_signature_31.png
| 2024-02-04 | 2024-02-02 08:41 | Philippine Heart Center | East avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | Pull out tie up unit | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27153 | MACARE-HBDC | phcmolecularlab@gmail.com | 28 |  |  | GX4 L | 804260 | N/A | N/A | Pull out genexpert system | Pull out tie up unit | 1kva ups accupower # AT1000190402ST0439, barcode scanner # 1324600508420,HP elite book 8470p # CNU330BSTR |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5828450561214768217/5828450561214768217_signature_28.png | Joyce Jasmin Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5828450561214768217/5828450561214768217_signature_31.png
| 2024-02-02 | 2024-02-02 15:35 | Rodriguez RHU | J.P Rizal.,Rodriguez Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27152 | PBSP | karlbanal25@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3014 ASSAY: Xpert HIV-1 Viral Load, COUNT: 37 | 2024-02-02 | GX4 | 807580 | Due for xpertcheck | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | B2 and B3 failed xpertcheck, |  | Xpertcheck | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826695241222239384/5826695241222239384_signature_28.png | Karl Michael M. Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5826695241222239384/5826695241222239384_signature_31.png
| 2024-02-02 | 2024-02-02 15:22 | Ragay Rural Health Unit | Ragay | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27151 | DOH | phylgomez0115@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 639 |  | GX4-6C | 838560 | B2/801695, B4/862868 for module replacement | Replaced Module: B2 - 801695>>757700 B4 - 862868>>757448 Cleaned cartridge bays, plunger rod, fan Generate reports | Replaced Module B2/801695, B4/862868 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826693533218313541/5826693533218313541_signature_28.png | Phyl Frances C. Gomez | https://www.jotform.com/uploads/macarewecare/222927978372067/5826693533218313541/5826693533218313541_signature_31.png
| 2024-02-02 | 2024-02-02 15:16 | Graman Medical And Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27150 | COMMERCIAL | gramanmolecularlab@gmail.com | 19 | ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 505 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GXIV | 110008530 | Module B4 for check up | - checked and scanned system log (temperature errors 1002 found) - generated reports - checked and cleaned cartrdige bay, plunger rod, filter | Scanned system log for possible hardware errors | Unit is operational Module B4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826682972122296190/5826682972122296190_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5826682972122296190/5826682972122296190_signature_31.png
| 2024-02-02 | 2024-02-02 14:00 | Julian Felipe Health Center | Jualin Felipe St., Sangandaan, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27149 | PBSP | maoi.lloyd@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1505 | 2024-01-26 | Gx IV 6C - D | 838884 | Unit for XpertCheck | Performed PM Performed Xpertcheck on modules B1, B2 and B4 All three modules passed Generated report | Performed XpertCheck on modules B1, B2, and B4 Module B3 reported and waiting for replacement | unit is operational 3 modules passed Module B3 waiting for replacement | Used 3 Xpert Cartridges | XpertCheck Kit | DR 46573 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826640311119345674/5826640311119345674_signature_28.png | lloyd maoi de guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5826640311119345674/5826640311119345674_signature_31.png
| 2024-02-02 | 2024-02-02 13:36 | Rogaciano M. Mercado Memorial General Hospital | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27148 | PBSP | alona0276@gmail.com | 23 |  |  | GXIV 10C | 120000608 | Unit for installation | #NAME? | Installed unit | Unit is operational and ready to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826615292122918322/5826615292122918322_signature_28.png | Alona Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5826615292122918322/5826615292122918322_signature_31.png
| 2024-02-02 | 2024-02-02 13:11 | Bay Rural Health Unit | Bay, Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com | princessanncatungos@gmail.com |  | 27147 | PBSP | rhubaylaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 561 |  | GXIV | 838132 | Unit for PM. Unit for Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | PM and Xpertcheck | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826613556013089714/5826613556013089714_signature_28.png | Princess Ann D. Catungos | https://www.jotform.com/uploads/macarewecare/222927978372067/5826613556013089714/5826613556013089714_signature_31.png
| 2024-02-02 | 2024-02-02 12:43 | Libmanan Rural Health Unit | Libmanan | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27146 | PBSP | amyquiapo1972@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1420 |  | Gx4-6C | 805260 | B2/210092909 for module replacement | Replaced Module: B2 - 210092909>>916366 Cleaned fan filter, fan, module bay Generate Report | Replaced B2/210092909 Performed Module Replacement | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826586555812746285/5826586555812746285_signature_28.png | Amy Quiapo | https://www.jotform.com/uploads/macarewecare/222927978372067/5826586555812746285/5826586555812746285_signature_31.png
| 2024-02-02 | 2024-02-02 12:08 | Nagcarlan Rural Health Unit | Nagcarlan Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27145 | DOH | jayongtangco@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1091 |  | GXIV | 821163 | Unit for PM. For Re-Xpertcheck of modules B1, B2 and B3. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Re-Xpertcheck. Generated reports. | PM and ReXpertcheck. | Unit is operational. Modules B1 (892641), B2 (627110) and B3 (862106), Service Required. | Used 4 spare Xpertcheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826568616225536725/5826568616225536725_signature_28.png | Jay M. Ongtangco | https://www.jotform.com/uploads/macarewecare/222927978372067/5826568616225536725/5826568616225536725_signature_31.png
| 2024-02-02 | 2024-02-02 11:35 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27144 | DOH | riazarizzamae@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1071 | 2024-02-02 | Genexpert IV | 831637 | Module A1 Error 4009 | Performed Xpertcheck on Module A1 (A1 = Service Required) | Performed Maintenance Gathered Reports Performed Module Replacement on A2 (784358 - 210012191) Performed Functionality Test | Unit is Operational | Module A1 for Replacement | Genexpert Module and Xpertcheck 1pc | 11117 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826552678987797850/5826552678987797850_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5826552678987797850/5826552678987797850_signature_31.png
| 2024-02-02 | 2024-02-02 11:18 | Malibay Health Center | Malibay, Pasay City | NCR | NCR 4th - Pasay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27143 | PBSP | ag_abad@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 933 ASSAY: Xpert CT/NG, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 138 | 2024-01-26 | GXIV 6C D | 838886 | Unit for Calibration | Performed PM Performed XpertCheck All modules passed Generated report | Performed XpertCheck | Unit is operational All modules working | Used 4 xpertcheck cartridges | XpertCheck Kit | DR 46578 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826534381115880937/5826534381115880937_signature_28.png | Aljandro G. Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5826534381115880937/5826534381115880937_signature_31.png
| 2024-02-02 | 2024-02-02 09:32 | Malvar Rural Health Unit | J. P. Laurel Highway, Malvar | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27142 | PBSP | malvarrhu@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 600 | 2024-01-29 | GXIV | 838191 | Unit for PM. Unit for Xpertcheck. For module replacement of module B1. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Performed module replacement. Generated reports. | PM, Xpertcheck and module replacement. | Unit is operational. |  | GX module B1 (877137) in replacement to defective B1 (806219) | G6778 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826491006326040436/5826491006326040436_signature_28.png | Baby Bhell A. Abraham | https://www.jotform.com/uploads/macarewecare/222927978372067/5826491006326040436/5826491006326040436_signature_31.png
| 2024-02-02 | 2024-01-25 16:24 | San Luis Rural Health Unit | Poblacion, San Luis, Batangas | 4A | Batangas | Laboratory | NO | MAINTENANCE Installation of Backup Laptop | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27138 | PBSP | sanluisbatangas.rhu@gmail.com | 24 |  | 2024-01-24 | GXIV-D-6C | 838131 | CPU of GX System Not Booting Up | Installation of Backup Laptop | Checked CPU Problem Installed Back-up Laptop Generate Reports | All Modules are Operational |  | GX Laptop L480 | G 6762 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826458025025248633/5826458025025248633_signature_28.png | Rhoda Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5826458025025248633/5826458025025248633_signature_31.png
| 2024-02-02 | 2024-01-23 09:53 | UST Hospital | Dapitan St, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27139 | MACARE-HBDC | usth_bloodbank@yahoo.com | 24 |  |  | MCS+ | 02F069 | Unit is due for Maintenance | Preventive Maintenance | Checked and Ckeaned Sensors, Rollers, and Centrifuge Checked A/D Values Performed Simulation Test Cleaned Surface Area of the Unit | Unit is in Good Condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826458035026726072/5826458035026726072_signature_28.png | Mylene Balboa | https://www.jotform.com/uploads/macarewecare/222927978372067/5826458035026726072/5826458035026726072_signature_31.png
| 2024-02-02 | 2024-01-30 13:30 | Pulang Lupa Dos Health Center | St. Joseph Subd, Pulang Lupa Dos, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27141 | PBSP | elvimir08@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1775 | 2024-01-29 | GXIV-D-6C | 838183 | Modukes 1 and 2 with Temperature Errors | Module Replacement | Performed Maintenance Replaced Modules 1 and 2 Performed Self-test Generate Reports | All Modules are Operational | 807075&807118 -> 866511&210067976 | GX Module | G 6777 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826458035026943656/5826458035026943656_signature_28.png | Elvira alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5826458035026943656/5826458035026943656_signature_31.png
| 2024-02-02 | 2024-02-01 08:13 | Tablas Island District Hospital | TIDH Rd., San Agustin, Romblon | 4B | Romblon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27140 | PBSP | nelda.grefiel@yahoo.com.ph | 24 |  |  | GXIV-2-D-6C | 835088 | CPU is Not Booting | CPU Replacement | Performed Maintenance Replaced CPU Setup GeneXpert Software Generate Reports | All Modules are Operational. New CPU is Good | CPH1119F04197 -> CPH1023F45887 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5826458035029353582/5826458035029353582_signature_28.png | Nelda S. Grefiel | https://www.jotform.com/uploads/macarewecare/222927978372067/5826458035029353582/5826458035029353582_signature_31.png
| 2024-02-01 | 2024-01-31 13:58 | Potrero Health Center | Potrero, Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | CALIBRATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27137 | PBSP | ailyn.santos19@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: 2223 ASSAY: Xpert HIV-1 Viral Load, COUNT: 62 | 2024-01-31 | GXIV | 804715 | Unit for calibration Printer for installation | #NAME? | Performed XpertCheck Installed Printer | Unit is operational All modules passed |  | XpertCheck Kit | na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5825776947125092192/5825776947125092192_signature_28.png | Ailyn Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5825776947125092192/5825776947125092192_signature_31.png
| 2024-02-01 | 2024-02-01 14:09 | Dinas Rural Health Unit | Dinas, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27136 | DOH | bernantereynacres@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 655 | 2023-12-20 | GX 4 | 838533 | Module 1 and 4 temperature error | Replaced module 862293 with module 866128 | Preventive Maintenence and Module Replacement for module 1 | Module 4 is h/w failed | Module 4 will be xpertchecked in the following weeks | 900-0491RS | 10992 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5825773948118458019/5825773948118458019_signature_28.png | Reyna Cres Bernante | https://www.jotform.com/uploads/macarewecare/222927978372067/5825773948118458019/5825773948118458019_signature_31.png
| 2024-02-01 | 2024-02-01 13:55 | OFW Hospital | Brgy. Sindalan, San Fernando City, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27135 | MACARE-HBDC | michellehalili17@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 7 | 2024-02-01 | Genexpert GXIV - 10c | 110018556 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration (All Modules Passed) Performed Functionality Test | Unit is Opetational | N/a | Xpertcheck kit | 11079 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5825770750017802178/5825770750017802178_signature_28.png | Michelle Angelica Halili | https://www.jotform.com/uploads/macarewecare/222927978372067/5825770750017802178/5825770750017802178_signature_31.png
| 2024-02-01 | 2024-02-01 10:11 | Commonwealth Health Center | Katuparan St., Commonwealth, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27134 | DOH | danicalaisa.025@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 104 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2960 |  | GX IV 6C | 831615 | Printer doesn't print | Performed PM Checked printer Data and power cable working Reset printer, problem persists Printer for replacement Generated report | Troubleshoot printer | GX is operational All modules working Printer for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5825757183619663574/5825757183619663574_signature_28.png | Danica Laisa Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5825757183619663574/5825757183619663574_signature_31.png
| 2024-02-01 | 2024-02-01 12:44 | Tondo Health Center | Juan Luna St. Gagalangi Tondo | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27133 | DOH | ma.lorinamanera@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1686 |  | GXIV | 831620 | Unit for PM. For Re-Xpertcheck of module A4. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Re-Xpertcheck. Generated reports. | PM and Re-Xpertcheck | Unit is operational. Module A4 (644057), service required. | Used 2 spare Xpertchecks. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5825741620756308233/5825741620756308233_signature_28.png | Lontoc, Ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5825741620756308233/5825741620756308233_signature_31.png
| 2024-02-01 | 2024-02-01 11:15 | Legazpi City Health Office | Ilawod, Legazpi City | 5 | Albay | RTDL Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27132 | WHO | janzchel_neneh10@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1262 |  | GX4-6C | 846354 | GX for Calibration | XpertCheck Result: B1/882523 - Passed B2/210101409 - Passed B3/210101426 - Passed B4/907516 - Passed Cleaned cartridge bays, plunger rods, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5825683793012427921/5825683793012427921_signature_28.png | Janzhel marie chan | https://www.jotform.com/uploads/macarewecare/222927978372067/5825683793012427921/5825683793012427921_signature_31.png
| 2024-02-01 | 2024-01-31 12:12 | Philippine Red Cross dagupan chapter | Dagupan City | 1 | Pangasinan | Bloodbank | NO | REPAIR MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27131 | MACARE-HBDC | prcpangasinanbloodcenter@gmail.com | 26 |  | 2024-01-24 | MCS + | 03A043 | No display,Defective lcd | Delivered and replaced Membrane panel and display cable Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Troubleshooting and repair | Operational | Work with Yasser | Display cable and Membrane panel | 11148 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5825456931896282116/5825456931896282116_signature_28.png | Emily Quiros | https://www.jotform.com/uploads/macarewecare/222927978372067/5825456931896282116/5825456931896282116_signature_31.png
| 2024-01-31 | 2024-01-31 15:36 | Southern Philippines Medical Center | JP Laurel Bajada Davao City | 11 | Davao del Sur | TB-HIV | YES | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27130 | PBSP | spmcrhivda@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2332 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1038 |  | GX-IV 4D | 808298 | Module B4 for replacement | Replaced the defective Module SN: 804865 with new Module SN: 870522 Generate IQR, Assay Stat and System Log | Perform Module Replacement | 3 Modules are operational | Module 3 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5825007134423015320/5825007134423015320_signature_28.png | Flordeliza Go RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5825007134423015320/5825007134423015320_signature_31.png
| 2024-01-31 | 2024-01-31 14:04 | Southern Philippines Medical Center | JP Laurel, Bajada Davao City | 11 | Davao del Sur | TB-HIV | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27129 | MACARE-HBDC | spmcrhivda@gmail.com | 18.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1671 ASSAY: Xpert HIV-1 Viral Load, COUNT: 813 |  | GX-IV 4D | 839152 | Stuck Cartridge at Module B2 | Remove the stuck cartridge | Perform Removal of Stuck Cartridge | Only Module B3 is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5825005574424702595/5825005574424702595_signature_28.png | Flordeliza Go RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5825005574424702595/5825005574424702595_signature_31.png
| 2024-01-31 | 2024-01-25 15:07 | Tigaon RHU | Tigaon | 5 | Camarines Sur | Tigaon RTDL | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27128 | DOH | jingimperial24@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1282 |  | GX-4 | 838558 | B1/862620, B3/862622, B4/695719 - temperature Error | Performed Re-XpertCheck for B1, B3, B4 Cleaned cartridge bays, plunger rod, icore Generate reports | Re-XpertCheck B1, B3, B4 Performed Preventive Maintenance | Unit is Operational Re-XC result: B1/862620 - service required B3/862622 - service required B4/695719 - passed | B1/862620, B3/862622 for module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824943793101121927/5824943793101121927_signature_28.png | Maria Jesusa P. Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/5824943793101121927/5824943793101121927_signature_31.png
| 2024-01-31 | 2024-01-25 16:56 | San Jose Rural Health Unit | San Jose | 5 | Camarines Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27127 | DOH | jhuneric2008@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1866 |  | GX-4 | 838389 | B2/607968, B4/718686 temperature error | XpertCheck Result: B2/607968 - service required B4/718686 - passed Generate Reports/Result | Performed Re-XpertCheck for B2, B4 | Unit is Operational B2/607968 for module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824943293109240403/5824943293109240403_signature_28.png | Rosmar Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5824943293109240403/5824943293109240403_signature_31.png
| 2024-01-31 | 2024-01-31 11:30 | Tondo Medical Center | H. Lopez Blvd., Balut Tondo | NCR | NCR 1st - Manila | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27126 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 10 ASSAY: Xpert HPV, COUNT: 69 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2728 ASSAY: Xpert SARS-CoV-2, COUNT: 7255 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 |  | GX16 | 844121 | Unit for PM. Unit for Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Module C3 (210070826), service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824830165619681055/5824830165619681055_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5824830165619681055/5824830165619681055_signature_31.png
| 2024-01-31 | 2024-01-31 11:27 | Tondo Medical Center | H. Lopez Blvd., Balut Tondo | NCR | NCR 1st - Manila | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27125 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 150 ASSAY: Xpert SARS-CoV-2, COUNT: 11 |  | GXIV | 834430 | Unit for PM. Unit for Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Module B1 (677430) & B2 (210070789), service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824829515615448628/5824829515615448628_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5824829515615448628/5824829515615448628_signature_31.png
| 2024-01-31 | 2024-01-31 10:48 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27124 | DOH | laboratoryadh@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 522 | 2024-01-31 | GXIV | 838548 | N/a | N/a | Performed preventive maintenance (Cartridge bay cleaning, plunger rod cleaning, changed fan filter, archrive data) | Unit is in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824789521919611757/5824789521919611757_signature_28.png | Mary Grace Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5824789521919611757/5824789521919611757_signature_31.png
| 2024-01-31 | 2024-01-31 08:10 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27123 | DOH | sajecopooh20@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1283 | 2024-01-16 | GXIV | 838555 | Module A2 defective | -Replace defective module A2/862903 with module SN: 878222 | #NAME? | All modules are functional. |  | 900-0491RS | DR no. 11116 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824785901868513765/5824785901868513765_signature_28.png | Sarah Jean R. Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5824785901868513765/5824785901868513765_signature_31.png
| 2024-01-31 | 2024-01-31 10:31 | Molave Rural Health unit | Molave Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27122 | DOH | aldrinacain@gmail.com | 16 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 778 | 2024-12-20 | GX 4 | 838233 | Module 3 temperature error | Replaced module 807260 with module 758220 | Module Replacement and Preventive Maintenence | All modules are available | None | 900-0491RS | 10991 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824777802404638758/5824777802404638758_signature_28.png | Aldrin L. Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/5824777802404638758/5824777802404638758_signature_31.png
| 2024-01-31 | 2024-01-31 09:22 | Navotas Medical and Wellness Center | M. Naval St., Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27121 | DOH | alejandrinoknowell@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 70 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1415 |  | GXIV | 831626 | Unit for PM. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Generated reports. | PM | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824741539743726928/5824741539743726928_signature_28.png | Noel Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5824741539743726928/5824741539743726928_signature_31.png
| 2024-01-31 | 2024-01-31 09:13 | Navotas Medical and Wellness Center | M. Naval St., Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27120 | PBSP | alejandrinoknowell@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 70 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1715 |  | GXIV | 807592 | Unit for PM. Unit for Xpertcheck. | Performed PM. -Cleaned and disinfect plunger -Cleaned and disinfect cartridge bay -Cleaned and disinfect instrument surface -Cleaned icore sensors -Cleaned filter fan -Archived test results Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Module B1 (908485), service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824741039746680217/5824741039746680217_signature_28.png | Noel Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5824741039746680217/5824741039746680217_signature_31.png
| 2024-01-30 | 2024-01-30 16:29 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27119 | PBSP | ibtee.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 548 | 2024-01-30 | GX4 L R1 | 704584 | Due for xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | B3 failed xpert check ,b4 newly replaced, 2 module passed xpertcheck |  | Xpertcheck | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824147330524414886/5824147330524414886_signature_28.png | Ivan Tee | https://www.jotform.com/uploads/macarewecare/222927978372067/5824147330524414886/5824147330524414886_signature_31.png
| 2024-01-30 | 2024-01-30 16:35 | Lung Center of the Philippines | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27118 | PBSP | ibtee.pbsp@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 592 | 2024-01-30 | GX4 L R1 | 707450 | Due for xoertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | C1 failed xpertcheck, 3 modules passed |  | Xpertcheck | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824146520525743041/5824146520525743041_signature_28.png | Ivan Tee | https://www.jotform.com/uploads/macarewecare/222927978372067/5824146520525743041/5824146520525743041_signature_31.png
| 2024-01-30 | 2024-01-30 13:49 | Bantayan RHU | Bantayan Island | 7 | Cebu | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27117 | DOH | mcdelle916bausin@gmail.com | 30 |  |  | Gx4 | 838509 | Defective Monitor | Install monitor for bantayan rhu, clean fan filter | Install monitor for bantayan rhu | Module B2 H/W fail |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824141020013347585/5824141020013347585_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/5824141020013347585/5824141020013347585_signature_31.png
| 2024-01-30 | 2024-01-30 13:32 | Juan Posadas Health Center | Rodriguez St., Balut Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27116 | PBSP | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 173 |  | GXIV | 805060 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5824029741241240944/5824029741241240944_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/5824029741241240944/5824029741241240944_signature_31.png
| 2024-01-30 | 2024-01-30 12:03 | Research Institute for Tropical Medicine | Filinvest Corporate City, Alabang Muntinlupa City | NCR | NCR 4th - Muntinlupa | Clinical Laboratory (Training Center) | YES | INSTALLATION Gx operation and assay training | CEPHEID | Noel Vicente |  | nvicente@macare-medicals.com | ritmclinlab@gmail.com |  | 27115 | PBSP | ritm.covid19reception@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-01-30 | GeneXpert IV | 707527 | Installation | None | Installation, gx operation, assay training | Good and operational |  | None | None |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5823973570428006570/5823973570428006570_signature_28.png | Maria Nyrissa B. Ubaldo | https://www.jotform.com/uploads/macarewecare/222927978372067/5823973570428006570/5823973570428006570_signature_31.png
| 2024-01-30 | 2024-01-30 10:31 | Liloy rural health unit | Liloy, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com |  |  | 27114 | DOH | gisellegipulan75@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1427 | 2023-12-20 | GX 4 | 838536 | Module 3 error 2014, 4012 | Replaced module 654516 with module 786645 | Module Replacement and Preventive Maintenence | All modules are available | None | 900-0491RS | 10990 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5823924211529424483/5823924211529424483_signature_28.png | GISELLE M.GIPULAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5823924211529424483/5823924211529424483_signature_31.png
| 2024-01-29 | 2024-01-23 12:57 | Malungon Rural Health Unit | Poblacion, Malungon, Sarangani | 12 | Sarangani | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27113 | DOH | ulbenarioe@gmail.com | 25.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 816 | 2024-01-23 | Gx-IV | 838546 | Module 4 for Service Required | Replace defective module 4 SN:862858 with new module SN:786232 Genexpert machine Cleaning Checking of errors Data collection: Pre and Post IQR, System log, Assay statistics report and Archive test | Perform Genexpert machine module replacement/repair Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 color SN:786232 | DR number: 10989 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5823232574314363567/5823232574314363567_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5823232574314363567/5823232574314363567_signature_31.png
| 2024-01-29 | 2024-01-24 11:15 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27113 | DOH | jan19.cam@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1555 | 2024-01-24 | Gx-IV | 838504 | Module 2 and 3 H/W failed | Replace defective module 2 SN:701378 with new module SN:786418 Replace defective module 3 SN:862410 with new module SN:807256 Genexpert machine Cleaning Checking of errors Data collection: Pre and Post IQR, System log, Assay statistics report, and archive test | Perform Genexpert machine module replacement/repair Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert 6 color module SN:807256 and 786418 | DR number: 10988 and 10980 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5823232574314919686/5823232574314919686_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5823232574314919686/5823232574314919686_signature_31.png
| 2024-01-29 | 2024-01-25 10:23 | Bicol Medical Center | Naga | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27112 | PBSP | bmc.bacteriology@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1414 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7558 ASSAY: Xpert SARS-CoV-2, COUNT: 2458 ASSAY: Xpert HIV-1 Viral Load, COUNT: 352 ASSAY: Others, COUNT: 10 |  | GX4 - 6C | 842225 | A1/885370 - broken module door C2/718898 - for module replacement C4/896403 - for module replacement | A1/885370 - replaced module door C2/718898>>789699 C4/896403>>790018 | A1 - replaced module door C2, C4 - replaced module | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5823115118214523480/5823115118214523480_signature_28.png | Gianne Pornillos | https://www.jotform.com/uploads/macarewecare/222927978372067/5823115118214523480/5823115118214523480_signature_31.png
| 2024-01-29 | 2024-01-25 13:18 | Bugasong Rural Health Unit | Bugasong, antique | 6 | Antique | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27111 | PBSP | jeanelyngabuyogimalay@gmail.com | 30 |  |  | Gx 4 | 815427 | Stucked up carthridge module 2 | Removal of Stucked up carthridge module 2 | Removal of Stucked up carthridge module 2 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5823061651112980232/5823061651112980232_signature_28.png | Jeanelyn G. Imalay | https://www.jotform.com/uploads/macarewecare/222927978372067/5823061651112980232/5823061651112980232_signature_31.png
| 2024-01-29 | 2024-01-25 13:50 | Morong Rural Health Unit | Morong | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27110 | DOH | morongrhu@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 513 | 2024-01-08 | GX-IV-D-6C | 821213 | Module A1 for replacement | Replaced module A1 | Performed maintenance Replaced module A1 Generated reports | All modules operational |  | Old: A1/717338 New: A1/662285 | DR# 11070 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5823015513617131122/5823015513617131122_signature_28.png | Necita Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5823015513617131122/5823015513617131122_signature_31.png
| 2024-01-29 | 2024-01-26 11:05 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27109 | COMMERCIAL | moleculardx@themedicalcity.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 32 ASSAY: Xpert CT/NG, COUNT: 44 ASSAY: Xpert HBV VL, COUNT: 434 ASSAY: Xpert HPV, COUNT: 334 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3731 ASSAY: Xpert HCV VL, COUNT: 67 ASSAY: Xpert HIV-1 Viral Load, COUNT: 572 ASSAY: Xpert SARS-CoV-2, COUNT: 474 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5823014643611257454/5823014643611257454_signature_28.png | Fritzie Mae T. Perante | https://www.jotform.com/uploads/macarewecare/222927978372067/5823014643611257454/5823014643611257454_signature_31.png
| 2024-01-29 | 2024-01-26 13:32 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27108 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Simulation run Checked pump rotors Verified calibration parameters Functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822997313615118847/5822997313615118847_signature_28.png | Analyn Ponferrada | https://www.jotform.com/uploads/macarewecare/222927978372067/5822997313615118847/5822997313615118847_signature_31.png
| 2024-01-29 | 2024-01-26 15:35 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27107 | COMMERCIAL | rmcbloodbank@gmail.com | 24 |  |  | MCS+ | 15H045G | For Maintenance | Performed maintenance | Performed maintenance Simulation run Verified calibration parameters Checked pump rotors Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822997183613882253/5822997183613882253_signature_28.png | Jennifer reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5822997183613882253/5822997183613882253_signature_31.png
| 2024-01-29 | 2024-01-26 11:34 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27106 | COMMERCIAL | pedelacruz@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For maintenance | Performed maintenance | Performed maintenance Simulation run Checked rotor pump Verified calibration parameters Functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822997173616352236/5822997173616352236_signature_28.png | Perlita De La Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5822997173616352236/5822997173616352236_signature_31.png
| 2024-01-29 | 2024-01-26 11:05 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27104 | COMMERCIAL | moleculardx@themedicalcity.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 32 ASSAY: Xpert CT/NG, COUNT: 44 ASSAY: Xpert HBV VL, COUNT: 434 ASSAY: Xpert HPV, COUNT: 334 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3731 ASSAY: Xpert HCV VL, COUNT: 67 ASSAY: Xpert HIV-1 Viral Load, COUNT: 572 ASSAY: Xpert SARS-CoV-2, COUNT: 474 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913611192076/5822996913611192076_signature_28.png | Fritzie Mae T. Perante | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913611192076/5822996913611192076_signature_31.png
| 2024-01-29 | 2024-01-25 13:50 | Morong Rural Health Unit | Morong | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27105 | DOH | morongrhu@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 513 | 2024-01-08 | GX-IV-D-6C | 821213 | Module A1 for replacement | Replaced module A1 | Performed maintenance Replaced module A1 Generated reports | All modules operational |  | Old: A1/717338 New: A1/662285 | DR# 11070 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913613869307/5822996913613869307_signature_28.png | Necita Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913613869307/5822996913613869307_signature_31.png
| 2024-01-29 | 2024-01-23 10:55 | Department oh Health Treatment and Rehabilitation Center - Bicutan | Camp Bagong Diwa, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27103 | PBSP | dohtrcbicutan.lab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1896 |  | GX-IV-D-6C | 807452 | CPU does not continue to boot and blue screen error appears | Troubleshooting of CPU | Reseat internal connection of CPU Cleaned RAM Performed disk checking Performed maintenance Generated reports | Unit is operational. CPU under observation. | CPU SN: 2UA5212D9Y |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913614786015/5822996913614786015_signature_28.png | John Jay O. Madlangbayan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913614786015/5822996913614786015_signature_31.png
| 2024-01-29 | 2024-01-25 09:44 | Cainta City Health Office | Karangalan Dr. | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27102 | DOH | ballesrerichard18@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 869 ASSAY: Xpert HIV-1 Viral Load, COUNT: 68 | 2024-01-16 | GX-IV-D-6C | 831623 | Module A2 for replacement | Replaced module A2 | Performed maintenance Replaced module A2 Generated reports | Unit is operational. Module A1 not operational. Reported already. |  | Old: A2/754519 New: A2/220005534 | DR# 11119 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913617982494/5822996913617982494_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913617982494/5822996913617982494_signature_31.png
| 2024-01-29 | 2024-01-23 14:12 | Taguig Pateros District Hospital | C5 Road | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27104 | DOH | kristinejane.kjd25@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3182 ASSAY: Xpert HIV-1 Viral Load, COUNT: 791 |  | GX-IV-D-6C | 831638 | Module A1 and A4 has error 4009 | Run XpertCheck on modules A1 and A4 | Performed maintenance Run XpertCheck on modules A1 and A4 Generated reports | Unit is Operational. Modules A1 and A4 passed the Xpertcheck but under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913618297628/5822996913618297628_signature_28.png | Kristine Jane Base | https://www.jotform.com/uploads/macarewecare/222927978372067/5822996913618297628/5822996913618297628_signature_31.png
| 2024-01-28 | 2024-01-26 12:06 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27101 | PBSP | keithtagudar@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 80 ASSAY: Xpert MTB/RIF Ultra, COUNT: 625 ASSAY: Others, COUNT: 24 | 2024-01-26 | GXIV-4-D-10C | 120000171 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 220011692 >>> 210056927 - Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210056927 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822648140947050824/5822648140947050824_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5822648140947050824/5822648140947050824_signature_31.png
| 2024-01-28 | 2024-01-26 09:41 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27100 | PBSP | keithtagudar@gmail.com | 21 | ASSAY: Xpert HPV, COUNT: 552 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1692 |  | GXIV-4-D-6C | 834685 | For Troubleshooting and Maintenance | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822646580947669003/5822646580947669003_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5822646580947669003/5822646580947669003_signature_31.png
| 2024-01-28 | 2024-01-26 09:30 | Baguio Health Services Office | Baguio City | 1 | Benguet | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27099 | DOH | sherryannemarrero@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1287 |  | GXIV-4-D-6C | 838281 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822644880947698556/5822644880947698556_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5822644880947698556/5822644880947698556_signature_31.png
| 2024-01-28 | 2024-01-24 08:51 | Mariano Marcos State University | Batac, Ilocos Norte | 1 | Ilocos Norte | College of Medicine | NO | INSTALLATION | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27098 | MACARE-HBDC | jpkperalta@mmsu.edu.ph | 23 |  | 2024-01-10 | GX4 10C | 110022474 | For Installation | #NAME? | Performed Installation | Unit is Operational |  | GXIV-4-D-10C | DR 11134 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5822640540941872016/5822640540941872016_signature_28.png | Dr. Patrick James Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/5822640540941872016/5822640540941872016_signature_31.png
| 2024-01-26 | 2024-01-26 15:08 | Hi Precision Diagnostics | 674 NS Amoranto Sr. ST. QC | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  | htoco@macare-medicals.com | jose@hi-precision.com.ph |  | 27097 | CONSORTIUM | pcr@hi-precision.com.ph | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 58 ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5602 | 2024-01-26 | GX4 | 835601 | Module B1:677244 defective and for module replacement | Replaced module B1 with new moduke sn: 687400 | Performed module replacement. Performed maintenance. Geberate reports. | All 4 modules operational | NA | Good mod SN: 687400 | DR number: 11154 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820630040395107642/5820630040395107642_signature_28.png | Glenn Christian A. Banlaygas | https://www.jotform.com/uploads/macarewecare/222927978372067/5820630040395107642/5820630040395107642_signature_31.png
| 2024-01-26 | 2024-01-26 13:28 | Caloocan City Health Office | A. Mabini St., Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27096 | PBSP | caloocancentrallaboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1204 ASSAY: Xpert HIV-1 Viral Load, COUNT: 440 | 2024-01-24 | GX IV 6C - D | 805271 | Module A4 for replacement | Performed PM Replaced module A4 Generated report | Replaced module A4 210059851 -> 877032 | Unit is operational All modules working |  | GX Module 6C | G 6764 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820617773721516927/5820617773721516927_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5820617773721516927/5820617773721516927_signature_31.png
| 2024-01-26 | 2024-01-26 13:20 | Cabanatuan City Health Office | Kapt. Pepe, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27095 | DOH | cholab20@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2387 | 2024-01-26 | Genexpert IV | 831592 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820558860316943868/5820558860316943868_signature_28.png | Maricris So | https://www.jotform.com/uploads/macarewecare/222927978372067/5820558860316943868/5820558860316943868_signature_31.png
| 2024-01-26 | 2024-01-26 13:02 | Manila Central University | Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27094 | CONSORTIUM | mikaelamagno02@gmail.com | 21 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 271 |  | GXIV | 110018008 | Unit for check up | #NAME? | Performed maintenance | Unit is operational Unit is due for calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820553838513356994/5820553838513356994_signature_28.png | Michelle Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5820553838513356994/5820553838513356994_signature_31.png
| 2024-01-26 | 2024-01-26 12:32 | Philippine Red Cross Zamboanga | Pettit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION | HAEMONETICS | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27093 | COMMERCIAL | zamboangacity@redcross.org.ph | 0 |  | 2024-01-26 | MCS + | 01M256 | None | Installed new purchased UPS (Uninterrupted power supply | Installed new purchased UPS (Uninterrupted power supply | Unit is available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820534057905703337/5820534057905703337_signature_28.png | Philippine Red Cross Zamboanga City Chapter | https://www.jotform.com/uploads/macarewecare/222927978372067/5820534057905703337/5820534057905703337_signature_31.png
| 2024-01-26 | 2024-01-25 10:13 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Stem Cell | YES | Checked up | OTHERS | Noel Floresca |  | florescanoel77@gmail.com |  | Plasmaire | 27092 | MACARE-HBDC | lcp.mdctl2024@gmail.com | 25 |  |  | 2006T | 1188 | For quotation of filters | Checked filters | Checked up | Operational | Need to replace pre filter and filter 2,3,4 for quotation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820516058997735677/5820516058997735677_signature_28.png | Mary Suzette A. De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5820516058997735677/5820516058997735677_signature_31.png
| 2024-01-26 | 2024-01-25 14:59 | National Kidney and Transplant Institute | East avenuv.,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27091 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1208 ASSAY: Others, COUNT: 43 ASSAY: Xpert HBV VL, COUNT: 866 ASSAY: Xpert HPV, COUNT: 94 ASSAY: Xpert MTB/RIF, COUNT: 524 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1939 ASSAY: Xpert SARS-CoV-2, COUNT: 15094 ASSAY: Xpert HCV VL, COUNT: 88 | 2024-01-24 | GX16D | 110005587 | For module replacement C2 210009197- 646651. C4 890075 and A4 882420 encounter continous temp error 4017,4012 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | A4 and C4 for replacement and for quotation, A2 for reporting 13 module operational |  | 900-0491rs | 11149 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820514938991752688/5820514938991752688_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5820514938991752688/5820514938991752688_signature_31.png
| 2024-01-26 | 2024-01-26 12:04 | Tanauan City Health Office | Tanauan City | 4A | Batangas | Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27090 | PBSP | edithramilo04@gmail.com | 25 |  |  | GXIV 10C | 120000605 | Unit for installation. | Performed installation. Generated reports. | Installation | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820514469113567424/5820514469113567424_signature_28.png | EDITA D. RAMILO | https://www.jotform.com/uploads/macarewecare/222927978372067/5820514469113567424/5820514469113567424_signature_31.png
| 2024-01-26 | 2024-01-26 11:40 | Talavera General Hospital | Poblacion, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27089 | MACARE-HBDC | pjgteh@gmail.com | 28 | ASSAY: Xpert HBV VL, COUNT: 79 ASSAY: Xpert MTB/RIF Ultra, COUNT: 642 ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 184 | 2024-01-26 | Genexpert IV | 110013774 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | For Calibration this July 2024 | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820501560315406803/5820501560315406803_signature_28.png | John Carlo P. Cuanan | https://www.jotform.com/uploads/macarewecare/222927978372067/5820501560315406803/5820501560315406803_signature_31.png
| 2024-01-26 | 2024-01-26 10:31 | Sto Domingo District Hospital | Pulong buli, Sto. Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27088 | DOH | lhaydeesel@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1033 | 2024-01-26 | Genexpert IV | 831582 | Module 2 error 4006 | Module 2 for reporting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 2 for replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820457360319355478/5820457360319355478_signature_28.png | Haydee Alejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5820457360319355478/5820457360319355478_signature_31.png
| 2024-01-26 | 2024-01-26 09:09 | Dipolog City Health Office | Dipolog City, Zamboanga del Norte | 9 | Zamboanga del Norte | TB-Dots | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27087 | PBSP | jz.ebron@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1310 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 |  | GXIV-II | 834826 | Module B2 temp error | None | #NAME? | Module B2 disabled but no recent error. Module B1 operational. | Module B2 under observation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820408953742278309/5820408953742278309_signature_28.png | James Zebedee C. Ebron | https://www.jotform.com/uploads/macarewecare/222927978372067/5820408953742278309/5820408953742278309_signature_31.png
| 2024-01-26 | 2024-01-26 08:59 | Arsenia J. Maximo Health Center | 925 Quirino Highway, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27086 | PBSP | elyvettemacalalagbasa@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 198 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1767 ASSAY: Xpert HIV-1 Viral Load, COUNT: 265 | 2024-01-24 | GX IV 6C - D | 814162 | Module B3 for replacement | Performed PM Replaced Module B3 Generated report | Replaced module B3 688602 -> 874837 | U it is operational. All modules working |  | GX IV 6C module | G 6763 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5820405535795782170/5820405535795782170_signature_28.png | ma, Jocelyn c utanes | https://www.jotform.com/uploads/macarewecare/222927978372067/5820405535795782170/5820405535795782170_signature_31.png
| 2024-01-25 | 2024-01-25 16:09 | Nueva Vizcaya Provincial Hospital | Almaguer North, Bambang, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | dumlaojoy4@gmail.com |  | 27085 | DOH | dumlaojoy4@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1529 | 2024-01-25 | Genexpert IV | 838272 | Modules B1 and B2 error Printer not working | Troubleshoot Printer Performed Xpertcheck on Modules B1 and B2 B1 = Service Required B2 = Service Required | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Printer and Modules B1 & B2 for Replacement | Xpertcheck 2pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819799522213517915/5819799522213517915_signature_28.png | Bernadette Dumlao | https://www.jotform.com/uploads/macarewecare/222927978372067/5819799522213517915/5819799522213517915_signature_31.png
| 2024-01-25 | 2024-01-25 15:48 | Doña Marta Health Center | Pasay City, Metro Manila | NCR | NCR 4th - Pasay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27084 | DOH | cerissalynsabado@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 846 |  | GX IV 6C- D | 831700 | Module A3 and A4 frequent lost communication error | Performed PM Checked system log 1000+ lost comm error Replaced Gateway board Tightened screws GX for monitoring if lost comm problem solved Generated report | Replaced Gateway board | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819786111819783833/5819786111819783833_signature_28.png | jouana marie sumando | https://www.jotform.com/uploads/macarewecare/222927978372067/5819786111819783833/5819786111819783833_signature_31.png
| 2024-01-25 | 2024-01-25 14:57 | Dr. Jose Rizal Memorial Hospital | Dapitan City, Zamboanga del Norte | 9 | Zamboanga del Norte | TB-Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 27083 | PBSP | calumbamendel@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1477 | 2024-01-25 | GXIV | 804698 | Module C2 defective | -Replace defective module C2/888548 with module SN: 754396 | #NAME? | All modules are functional. |  | 900-0491RS | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819756664425509802/5819756664425509802_signature_28.png | Mendel Q. Calumba | https://www.jotform.com/uploads/macarewecare/222927978372067/5819756664425509802/5819756664425509802_signature_31.png
| 2024-01-25 | 2024-01-25 14:16 | Compostela Rural Health Unit | Poblacion Compostela, Davao De Oro | 11 | Davao de Oro | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | comp.clin.lab@gmail.com |  | 27082 | PBSP | nielmaquiling20@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1570 |  | GX-IV 4D | 807360 | System Unit for Replacement | Replaced the Defective System Unit SN: 2UA52027TS with new System Unit SN: CPH1023F45886 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Migrate the datas to new system unit | Perform System Unit replacement | Module B3 encounter temperature related errors | Data gathered to be sent at Cepheid for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819748350912740601/5819748350912740601_signature_28.png | Neil C. Maquiling RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5819748350912740601/5819748350912740601_signature_31.png
| 2024-01-25 | 2024-01-25 14:19 | Cebu TB Reference Laboratory | Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 27081 | PBSP | asmantiquilla.pbsp@gmail.com | 23 |  | 2024-01-25 | Gx4 | 838882 | Module B3 hardware fail serial number 863551 B3 | Replace module | Replace module B3 with new serial number 877045 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819731353693606337/5819731353693606337_signature_28.png | Alan Montiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5819731353693606337/5819731353693606337_signature_31.png
| 2024-01-25 | 2024-01-25 14:04 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27080 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1282 | 2024-01-16 | GXIV | 831691 | Monitor for replacement | #NAME? | Replaced monitor | Unit is operational |  | GeneXpert Monitor SN: AV1215CMQY05B01060568 | 11124 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819727975191010278/5819727975191010278_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5819727975191010278/5819727975191010278_signature_31.png
| 2024-01-25 | 2024-01-23 14:31 | Catmon Health Center | Catmon, Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27077 | DOH | mariathes0218@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2024-01-23 | GXIV | 831903 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational C1: passed C3: service required |  | XpertCheck Kit | na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267321417938/5819681267321417938_signature_28.png | Maria Theresa Nunez | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267321417938/5819681267321417938_signature_31.png
| 2024-01-25 | 2024-01-25 12:14 | Norzagaray Rural Health Unit | Norzagaray, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27079 | PBSP | chelseajempolicarpio@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1245 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1245 | 2024-01-25 | GXIV | 838167 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational B1, B3, B4 passed calibration |  | XpertCheck Kit | na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267322604956/5819681267322604956_signature_28.png | Rowena Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267322604956/5819681267322604956_signature_31.png
| 2024-01-25 | 2024-01-24 15:45 | Mandaluyong City Health Office | Mandaluyong City | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27075 | PBSP | rsjalova1974@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2489 ASSAY: Xpert HIV-1 Viral Load, COUNT: 132 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2489 ASSAY: Xpert HIV-1 Viral Load, COUNT: 132 | 2024-01-24 | GXIV | 805062 | Unit for calibration UPS for troubleshooting | #NAME? | Performed XpertCheck | Unit is operational All modules passed UPS is recommended for replacement (no back up power) |  | XpertCheck Kit | na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267323647298/5819681267323647298_signature_28.png | Teresita Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267323647298/5819681267323647298_signature_31.png
| 2024-01-25 | 2024-01-23 14:27 | Catmon Health Center | Catmon, Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27076 | PBSP | mariathes0218@gmail.com | 25 |  | 2024-01-23 | GXIV | 838880 | Unit for calibration | #NAME? | Performed calibration | Unit is operational B4 passed the calibration |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267329657673/5819681267329657673_signature_28.png | Maria Theresa Nunez | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267329657673/5819681267329657673_signature_31.png
| 2024-01-25 | 2024-01-24 15:49 | Mandaluyong City Health Office | Mandaluyong City | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27078 | DOH | rsjalova1974@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2548 ASSAY: Xpert HIV-1 Viral Load, COUNT: 151 |  | GXIV | 831596 | Module C3 ultrasonic error | #NAME? | Performed XpertCheck on C3 | Unit is operational C3 requires service |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267329904139/5819681267329904139_signature_28.png | Rosevi Jalova | https://www.jotform.com/uploads/macarewecare/222927978372067/5819681267329904139/5819681267329904139_signature_31.png
| 2024-01-25 | 2024-01-25 12:27 | Southern Isabela Medical Center | Rosario, Santiago City, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27074 | PBSP | cassflotildesvillanueva@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4504 ASSAY: Xpert HIV-1 Viral Load, COUNT: 196 | 2024-01-25 | Genexpert IV | 805104 | Module 3 error 4008 | Troubleshoot module 3 Check the fuse of module 3 | Performed Maintenance Gathered Reports Performed Module Replacement (874222 - 861608) Performed Functionality Test | Unit is Operational | Fuse Broken | Genexpert Module - 861608 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819665492215547879/5819665492215547879_signature_28.png | Cassandra Cates Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5819665492215547879/5819665492215547879_signature_31.png
| 2024-01-25 | 2024-01-25 12:03 | Belmonte Health Center | Sampaloc Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27073 | PBSP | spsurell@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1250 | 2024-01-22 | GXIV | 814376 | Unit for PM. For module replacement of modules A2 and A4. | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module A2 (877065) ans A4 (877133) in replacement to defective A2 (671861) and A4 (882082). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819654361266122205/5819654361266122205_signature_28.png | Clarinda Sharon P.Surell | https://www.jotform.com/uploads/macarewecare/222927978372067/5819654361266122205/5819654361266122205_signature_31.png
| 2024-01-25 | 2024-01-25 11:55 | Belmonte Health Center | Sampaloc Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27072 | PBSP | spsurell@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1250 | 2024-01-22 | GXIV | 808619 | Unit for PM. For module replacement of module B1. | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module B1 (869596) in replacement to defective B1 (210101692). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819653661262142721/5819653661262142721_signature_28.png | Clarinda Sharon P, Surell | https://www.jotform.com/uploads/macarewecare/222927978372067/5819653661262142721/5819653661262142721_signature_31.png
| 2024-01-25 | 2024-01-25 11:47 | Klinika Bernardo Health Center | Quezon City Health Department Building, Gate 4, Quezon City Hall Compound | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27071 | PBSP | klinikabernardolaboratory@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 27 ASSAY: Xpert MTB/RIF Ultra, COUNT: 720 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1458 |  | GX IV 6C | 808618 | Unit for setup in new location Unit due for Calibration UPS Checking | Performed PM Setup GX and Hardware Checked UPS UPS does not turn on Recommended for replacement Performed XpertCheck Generated report | Setup GX machine and hardware to new location Klinika Bernardo, Magsaysay Ave. Ermin Garcia St, Brgy. Pinagkaisagan transferred to Quezon City Health Department Building, Gate 4, Quezon City Hall Compound Module 1&2 passed XpertCheck Module 3&4 for retry, cartridge lacking | Unit is operational All modules working UPS for replacement To retry XpertCheck when kits arrive | Used 5 XpertCheck kit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819649667699580655/5819649667699580655_signature_28.png | Mylen Sibal, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5819649667699580655/5819649667699580655_signature_31.png
| 2024-01-24 | 2024-01-24 22:25 | Medellin Rural Health Unit | Poblacion Medellin Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27070 | DOH | rhumedellin@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1780 |  | GX4 | 838506 | Stucked up cartridge in Module 3 | Removed cartridge in module 3 | Removed cartridge in module 3 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819156821915650603/5819156821915650603_signature_28.png | Raul Galano, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5819156821915650603/5819156821915650603_signature_31.png
| 2024-01-24 | 2024-01-23 15:24 | National Childrens Hospital | E.Rodriguez, QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27069 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 28 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3019 ASSAY: Xpert SARS-CoV-2, COUNT: 4999 ASSAY: Others, COUNT: 10 |  | GX16 D 10C | 110014171 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819127479963477420/5819127479963477420_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5819127479963477420/5819127479963477420_signature_31.png
| 2024-01-24 | 2024-01-23 09:48 | Tatalon Health Center | Victory avenue QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27068 | MACARE-HBDC | jel472020@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 968 |  | GX4 D | 831610 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819126969964815617/5819126969964815617_signature_28.png | Jasmin A. Magtalas | https://www.jotform.com/uploads/macarewecare/222927978372067/5819126969964815617/5819126969964815617_signature_31.png
| 2024-01-24 | 2024-01-22 16:08 | Philippine Heart Center | East Avenue, QC | NCR | NCR 2nd - Quezon City | Molecular | YES | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27067 | MACARE-HBDC | avabichuela3216@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 572 ASSAY: Others, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 1062 |  | GX16 | 838655 | For module replacement D2 653968 - 220015088, C4 628454-22001525, C3 645559 - 220015090, C1 668429 -220015305, B3 662646 - 22005301, A3 644007 -220015266, A4 634060 - 220015133, B4 804727 -220015270 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819125699964558870/5819125699964558870_signature_28.png | Arjay V. Abichuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5819125699964558870/5819125699964558870_signature_31.png
| 2024-01-24 | 2024-01-24 15:45 | East Avenue Medical Center | East ave. QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | Relocation of genexpert | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27066 | MACARE-HBDC | roseannparnes@yahoo.com | 28 | ASSAY: Xpert CT/NG, COUNT: 291 ASSAY: Xpert HPV, COUNT: 159 ASSAY: Xpert SARS-CoV-2, COUNT: 858 |  | GX4D 10C | 110014085 | For transfer from molecular to tb dots | Re installed gx, Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Relocation and installation | Operational | Network 2 gx machine pbsp and hbdc |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5819125149963333902/5819125149963333902_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/5819125149963333902/5819125149963333902_signature_31.png
| 2024-01-24 | 2024-01-24 15:22 | National Government Center Super Health Center | NGC Sentrong Sigla Health Center, MRB compound, Pilot Drive, Commonwealth | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27065 | DOH | tmcolisao@gmail.com | 24 |  |  | GX IV 6C-D | 831586 | Filter for replacement Bios time not synced | Performed PM Changed filter Bios battery changed Monitor if time syncs automatically Generated report | Replaced Filter Replaced bios battery | unit is operational all modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818911221817878637/5818911221817878637_signature_28.png | Mila N. Baldovizo | https://www.jotform.com/uploads/macarewecare/222927978372067/5818911221817878637/5818911221817878637_signature_31.png
| 2024-01-24 | 2024-01-24 15:16 | NAZARETH HEALTH CENTER | Brgy. Nazareth, Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27064 | DOH | nanettebtolentino@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1324 |  | GX-IV | 838269 | MODULE B3 TESTS ARE ABORTED | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM XPERTCHECK CALIBRATION ON MODULE B3 SN 688652: RETRY THEN FOR SERVICE | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818900375514795652/5818900375514795652_signature_28.png | Nanette Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/5818900375514795652/5818900375514795652_signature_31.png
| 2024-01-24 | 2024-01-23 13:09 | Cebu TB reference lab - SRP | Cebu City | 7 | Cebu | Laboratory | NO | REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 27063 | PBSP | tivalerio.pbsp@gmail.com | 24 |  |  | Gx 4 | 120000180 | Stuck up cartridge module B1 | Removing stuck up carthride module B1 | Removing stuck up carthride module B1 | Module B2 and B4 H/W fail |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818850761111721638/5818850761111721638_signature_28.png | Titus Ian Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5818850761111721638/5818850761111721638_signature_31.png
| 2024-01-24 | 2024-01-24 12:54 | Bagong Pag-Asa Health Center | 35 Road 9, cor. Road 10, Brgy. Bagong Pag-asa | NCR | NCR 2nd - Quezon City | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27062 | DOH | erikatagorda28@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 752 | 2024-01-24 | GX IV 6C-D | 831597 | Module A2 for replacement | Performed PM Replaced Module A2 Generated report | Replaced module A2 914754 -> 210106611 | unit is operational all modules working |  | GX module 6C | DR 11118 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818848981812465935/5818848981812465935_signature_28.png | erika lucas | https://www.jotform.com/uploads/macarewecare/222927978372067/5818848981812465935/5818848981812465935_signature_31.png
| 2024-01-24 | 2024-01-24 10:51 | Quezon Rural Health Unit | Poblacion, Quezon | 10 | Bukidnon | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27061 | DOH | quenxlopez@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 782 | 2024-01-05 | GX-4 | 838440 | Module B2 found Service Required thru XPC | None | Replacing module B2 with SN: 806408 to SN: 629512 Preventive Maintenance | All modules and units are functional |  | Certified pre-owned GeneXpert module | 11035 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818746558324756448/5818746558324756448_signature_28.png | Quenie Rose Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5818746558324756448/5818746558324756448_signature_31.png
| 2024-01-24 | 2024-01-17 10:11 | Cainta City Health Office | Karangalan Dr | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27060 | PBSP | ballesterrichard18@gmail.com | 24 |  |  | GX-IV-D-6C | 838182 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Archive test Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818686476232370766/5818686476232370766_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5818686476232370766/5818686476232370766_signature_31.png
| 2024-01-24 | 2024-01-17 10:11 | Cainta City Health Office | Karangalan Dr | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27059 | PBSP | ballesterrichard18@gmail.com | 24 |  |  | GX-IV-D-6C | 838182 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Archive test Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818686053615096297/5818686053615096297_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5818686053615096297/5818686053615096297_signature_31.png
| 2024-01-24 | 2024-01-17 10:11 | Cainta City Health Office | Karangalan Dr | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27058 | PBSP | ballesterrichard18@gmail.com | 24 |  |  | GX-IV-D-6C | 838182 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Archive test Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818685673619277672/5818685673619277672_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5818685673619277672/5818685673619277672_signature_31.png
| 2024-01-24 | 2024-01-17 10:03 | Cainta City Health Office | Karangalan Dr | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27057 | PBSP | ballesterrichard18@gmail.com | 24 |  |  | GX-IV-D-6C | 804687 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Archive test Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818685463613715658/5818685463613715658_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5818685463613715658/5818685463613715658_signature_31.png
| 2024-01-23 | 2024-01-23 17:47 | NORTHERN MINDANAO MEDICAL CENTER | Cagayan de Oro City | 10 | Misamis Oriental | HIV LAB | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27056 | PBSP | ronnel_mabasa@yahoo.com | 26 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 118 |  | GX-IV | 804682 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE B4 SN 671936 SELF TEST, H/W FAILED | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818127116224087378/5818127116224087378_signature_28.png | Ronnel Mabasa | https://www.jotform.com/uploads/macarewecare/222927978372067/5818127116224087378/5818127116224087378_signature_31.png
| 2024-01-23 | 2024-01-23 17:34 | NORTHERN MINDANAO MEDICAL CENTER | Cagayan de Oro City | 10 | Misamis Oriental | LABORATORY | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27055 | PBSP | nmmc.gxmtbrif@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 670 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7768 ASSAY: Xpert SARS-CoV-2, COUNT: 10136 ASSAY: Xpert HIV-1 Viral Load, COUNT: 334 |  | GX-XVI | 844118 | #NAME? | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE REPLACEMENT: MODULE A3 DEF SN: 210059879 , NEW SN: 791986 MODULE B2 DEF SN: 678540 , NEW SN 798296 | EQUIPMENT IS OPERATIONAL | MODULE C2 AND C3 FOR OBSERVATION DUE TO INTERMITTENT ERRORS |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818122736229827959/5818122736229827959_signature_28.png | Rione Frances Lomongo | https://www.jotform.com/uploads/macarewecare/222927978372067/5818122736229827959/5818122736229827959_signature_31.png
| 2024-01-23 | 2024-01-23 17:26 | LAGUINDINGAN RURAL HEALTH UNIT | POBLACION LAGUINDINGAN | 10 | Misamis Oriental | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27054 | DOH | chellesambaan@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 965 |  | GX-IV | 838575 | -MODULE B3 ERROR 1001 | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B3 SN 863072 XPERTCHECK CALIBRATION, PASSED | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818115336222247902/5818115336222247902_signature_28.png | Rochelle M. Samban | https://www.jotform.com/uploads/macarewecare/222927978372067/5818115336222247902/5818115336222247902_signature_31.png
| 2024-01-23 | 2024-01-23 16:14 | Bulan Rural Health Unit | Bulan | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27053 | PBSP | bing.gonzales33@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2688 |  | GX-4 | 807983 | B3/210070808 for Module Replacement | Replaced Module: B3 - 210070808>>757241 Cleaned Gx fan, cartridge bays, plunger rod Generate reports | Replaced Module B3 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818073519212583259/5818073519212583259_signature_28.png | Claribel G. Nacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5818073519212583259/5818073519212583259_signature_31.png
| 2024-01-23 | 2024-01-23 14:28 | Kairos Diagnostic Laboratory - Ermita | 1123 A. Mabini St., Malate, Manila | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27052 | CONSORTIUM | kairoslaboperations@gmail.com | 27 |  | 2023-12-22 | GX IV 10C - D | 110022622 | Unit for installation | Performed PM Installed GX unit Setup hardwares Generated report | Installed GX Machine | Unit is operational All modules working (2 Mod) |  | GX IV 10C 2 modules GX Barcode Scanner GX Desktop GX Monitor | SI 9688 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5818029136626084276/5818029136626084276_signature_28.png | Christine Joy Feliciano | https://www.jotform.com/uploads/macarewecare/222927978372067/5818029136626084276/5818029136626084276_signature_31.png
| 2024-01-23 | 2024-01-23 14:00 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 27051 | DOH | ammegzedd@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1309 | 2024-01-23 | GXIV | 838421 | Module B2/617811 H/W Failed | #NAME? | - Performed xpertcheck on B2 - Preventive maintenance - Performed Functionality test - Generate reports - | Unit is Operational | Module B2 for service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817991484526331208/5817991484526331208_signature_28.png | gemma malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5817991484526331208/5817991484526331208_signature_31.png
| 2024-01-23 | 2024-01-23 13:11 | Baganga Rural Health Unit | Poblacion Baganga, Davao Oriental | 11 | Davao Oriental | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27050 | DOH | rhubaganga@gmail.com | 30.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 565 | 2023-12-18 | GX-IV 4D | 838395 | Module B2 for Replacement and due for Preventive Maintenance | Replace the defective Module SN: 860970 with new Module SN:888380 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Replacement | All Modules are operational | High Room Temperature, Enduser | 900-0491rs | DR NO. 10971 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817982212513010927/5817982212513010927_signature_28.png | Liza Narciso RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5817982212513010927/5817982212513010927_signature_31.png
| 2024-01-23 | 2024-01-22 13:06 | Davao del Norte Provincial Hospital-Kapalong Zone | Poblacion Kapalong, Davao Del Norte | 11 | Davao del Norte | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27049 | PBSP | ddnhkapalongzone.laboratory@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4431 |  | GX-IV 4D | 815270 | Module C1 encounter multiple temp related errors and for Preventive Maintenance | Replaced the Defective Module SN: 690965 with new Module SN: 749076 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Relocate the GeneXpert Machine to achieve optimal temperature for operation | Perform Module Replacement and Preventive Maintenance | Module C3 for under observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817965252514588662/5817965252514588662_signature_28.png | Leona Olila RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5817965252514588662/5817965252514588662_signature_31.png
| 2024-01-23 | 2024-01-23 10:40 | Josefina Belmonte Duran Albay Provincial Hospital | Tuburan, Ligao | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 27048 | PBSP | salvadorlyndonrempillo@gmail.com | 25 |  |  | GX-4 | 814371 | For CPU Replacement | Replaced CPU: CPH0723F43470 | Replaced CPU Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817880903696960569/5817880903696960569_signature_28.png | Salvador Lyndon P. Rempillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5817880903696960569/5817880903696960569_signature_31.png
| 2024-01-22 | 2024-01-19 11:48 | Dasmariñas (Paliparan) City Health Office IV | N. Dela Cruz, Paliparan, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27047 | PBSP | krlandaya25@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2314 ASSAY: Xpert HIV-1 Viral Load, COUNT: 91 | 2024-01-15 | GXIV-D-6C | 838139 | Module 2 with Temperature Error | Module Replacement | Replaced Module 2 Performed Maintenance Performed Self-test Generate Reports | All Moduled are Operational | B2:210091142-> 782599 | GX Module | G 6741 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175903612738774/5817175903612738774_signature_28.png | Khristofer Ralph L Andaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175903612738774/5817175903612738774_signature_31.png
| 2024-01-22 | 2024-01-18 09:04 | E. Aldana Health Center | St. Thomas St., Pulang Lupa Dos, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | MAINTENANCE ReXpertCheck | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27044 | PBSP | elvimir08@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1233 |  | GXIV-D-6C | 804261 | Unit is for ReXpertCheck | ReXpertCheck | Performed XpertCheck Performed Maintenance Performed Self-test Generate Reports | 3 Modules Failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175783612831300/5817175783612831300_signature_28.png | Elvira alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175783612831300/5817175783612831300_signature_31.png
| 2024-01-22 | 2024-01-18 13:05 | Imus City Health Office 3 | Greengate Subd, Imus, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27046 | PBSP | tess19_chua@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1434 | 2024-01-15 | GXIV-D-6C | 838166 | Module 4 with Temperature Erorr | Module Replacement | Performed Maintenance Replaced Module 4 Performed Self-test Generate Reports | 3 Modules are Operational, 1 Module for Reporting | B4: 807102-> 785038 | GeneXpert Module | G 6740 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175783614579493/5817175783614579493_signature_28.png | Marites M. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175783614579493/5817175783614579493_signature_31.png
| 2024-01-22 | 2024-01-18 10:47 | Kawit Rural Health Unit | Tirona Highway, Kawit, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27045 | DOH | kawitrhu@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 983 | 2024-01-08 | GXIV-D-6C | 820705 | Module 1 Failed in XpertCheck | Module Replacement | Replaced Module 1 Performed Maintenance Performed Self-test Generate Reports | All Modules are Operational | A1: 790930 -> 757992 | GX Module | 11061 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175783615154046/5817175783615154046_signature_28.png | Aljhon G. Alambra | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175783615154046/5817175783615154046_signature_31.png
| 2024-01-22 | 2024-01-19 09:23 | Dasmariñas City Health Office I | D. Mangubat Ave , Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27043 | PBSP | cho1lab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2907 ASSAY: Xpert HIV-1 Viral Load, COUNT: 124 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2565 ASSAY: Xpert HIV-1 Viral Load, COUNT: 122 |  | GXIV-D-6C | 838893 | Unit is due for XpertCheck | XpertCheck Preventive Maintenance | Performed Maintenance Performed XpertCheck Performed Self-test Generate Reports | 3 Modules are Operational | Modules B1,B4 Passed. B3 Failed. B2 Newly Replaced Module |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175783615260192/5817175783615260192_signature_28.png | Elma R.Belostrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175783615260192/5817175783615260192_signature_31.png
| 2024-01-22 | 2024-01-18 08:59 | Pulang Lupa Dos Health Center | St. Thomas St., Pulang Lupa Dos, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 27042 | PBSP | elvimir08@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1757 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1575 |  | GXIV-D-6C | 838183 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports Performed Self-test | 2 Modules are Operational, 2 Reported | 2 Reported, 1 Passed, 1 Newly Replaced |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175773616383608/5817175773616383608_signature_28.png | Elvira alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175773616383608/5817175773616383608_signature_31.png
| 2024-01-22 | 2024-01-17 09:58 | Cainta City Health Office | Karangalan Dr. | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27041 | DOH | ballesterrichard18@gmail.com | 24 |  |  | GX-IV-D-6C | 831623 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Archive test Generated reports | Unit is operational. Modules A1 and A2 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175333618916714/5817175333618916714_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5817175333618916714/5817175333618916714_signature_31.png
| 2024-01-22 | 2024-01-17 13:45 | Atang dela Rama Health Center | Gagalangin Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 27040 | DOH | monskie_22@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1301 |  | GXIV | 831644 | Unit for PM. For installation of new battery to UPS. | Performed PM. Performed repair of UPS. Generated reports | PM and repair. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817171307422720118/5817171307422720118_signature_28.png | Monaliza Balcos | https://www.jotform.com/uploads/macarewecare/222927978372067/5817171307422720118/5817171307422720118_signature_31.png
| 2024-01-22 | 2024-01-22 14:58 | Mexico Rural Health Unit IV | Pandacaqui, Mexico, Pampanga | 2 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27039 | DOH | annbathan79@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 482 | 2024-01-22 | Genexpert IV | 831902 | Filter broken | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Filter needs to be replaced | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817165761118476076/5817165761118476076_signature_28.png | Mary Ann D. Carreon | https://www.jotform.com/uploads/macarewecare/222927978372067/5817165761118476076/5817165761118476076_signature_31.png
| 2024-01-22 | 2024-01-22 12:58 | Magsaysay Rural Health Unit | Kakilgan St. Poblacion, Magsaysay | 13 | Misamis Oriental | RTDL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 27038 | DOH | jkakilgan12@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 326 |  | GX-4 | 838557 | No signal on Monitor reported thru Jotform Ticket | None | Troubleshooting CPU and Monitor Preventive Maintenance | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817090526326391748/5817090526326391748_signature_28.png | Jesusa N. Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/5817090526326391748/5817090526326391748_signature_31.png
| 2024-01-22 | 2024-01-15 17:38 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27037 | DOH | elma.bragat@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5591 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 148 |  | GX-IV | 838508 | MODULE B1 ERROR 2014 | #NAME? | -PERFORM MAINTENANCE PROCEDURE -PERFORM MODULE B1 SN 629957 CALIBRATION, PASSED | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5817021700479049241/5817021700479049241_signature_28.png | Elma Gonzales, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5817021700479049241/5817021700479049241_signature_31.png
| 2024-01-21 | 2024-01-19 15:06 | Rakkk Prophet Medical Center | Gumaca, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27036 | PBSP | rpmcimoleclab20@gmail.com | 29 |  | 2024-01-08 | GX IV 6C - D | 808268 | CPU for replacement | Perfomed PM Replaced CPU Setup new CPU Generated report | Replaced CPU SGH530PD9W -> CPH0923F45088 | unit is operational All modules working |  | GX desktop | DR 45920 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816553236903912719/5816553236903912719_signature_28.png | Katrina Rose F. Aureada | https://www.jotform.com/uploads/macarewecare/222927978372067/5816553236903912719/5816553236903912719_signature_31.png
| 2024-01-21 | 2024-01-18 12:13 | San Francisco Rural Health Unit | San Francisco, Quezon Province | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27035 | DOH | micovelarde13@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 440 | 2023-10-25 | GX IV 6C -D | 831698 | Unit for calibration | Performed PM Performed XpertCheck Module A2 service required Generated report | Performed XpertCheck | Unit is operational 3 modules working 1 module for reporting | Used 6 XpertCheck cartridges | XpertCheck Kit | DR10779 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816547196907578625/5816547196907578625_signature_28.png | Michael Francis Velarde, RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5816547196907578625/5816547196907578625_signature_31.png
| 2024-01-21 | 2024-01-19 18:36 | Quezon Medical Center | Lucena, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27034 | PBSP | qmcrtdl2015@gmail.com | 29 |  | 2024-01-11 | GX Iv 6C - D | 804692 | CPU for replacement | Performed PM Removed GX from network switch Replaced CPU Setup system admin Setup system configuration Generated report | Replaced CPU CPH0619F01767 -> CPH1023F45888 | Unit is operational All modules working |  | GX Desktop | DR 46016 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816542606908070158/5816542606908070158_signature_28.png | Ariel carmelo R. Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5816542606908070158/5816542606908070158_signature_31.png
| 2024-01-21 | 2024-01-21 21:19 | Quezon Medical Center | Lucena, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27033 | PBSP | qmcrtdl2015@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 373 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2735 ASSAY: Xpert HIV-1 Viral Load, COUNT: 83 | 2024-01-08 | GX IV 6c - D | 807462 | Module D2 for replacement | Performed PM Replaced Mdoule D2 Generated report | Replaced Module D2 861624 -> 776975 | Unit is operational. All modules working | Jotform Error. Module Replacement Performed on January 17, 2024 | GX Module 6C | DR 45919 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816527376906848565/5816527376906848565_signature_28.png | JoAnn P. Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5816527376906848565/5816527376906848565_signature_31.png
| 2024-01-21 | 2024-01-17 13:40 | Candelaria Rural Health Unit | Poblacion, Candelaria, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27032 | PBSP | shirleyzrubz@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1692 | 2024-01-09 | GX IV 6C -D | 814155 | Monitor for replacement | Performed PM Replaced Monitor Generated report | Replaced monitor CNK6290NRK -> SV5FTH182 | Unit is operational All modules working |  | GX Monitor | DR 45912 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816507756906182819/5816507756906182819_signature_28.png | Shirley R. Dollente | https://www.jotform.com/uploads/macarewecare/222927978372067/5816507756906182819/5816507756906182819_signature_31.png
| 2024-01-21 | 2024-01-19 11:00 | Eastern Pangasinan District Hospital | Tayug, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 27031 | PBSP | epdhlaboratory@gmail.com | 29 |  | 2024-01-19 | GX4 | 815394 | For CPU Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective CPU CPH1221F29401 >>> CPH0923F44932 - Generated Reports | Replaced Defective CPU | Unit is operational |  | CPU CPH0923F44932 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816411140942778834/5816411140942778834_signature_28.png | Fergyl Escaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5816411140942778834/5816411140942778834_signature_31.png
| 2024-01-21 | 2024-01-17 14:50 | Antipolo City Health Office | M. Santos St. | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27030 | PBSP | acho.laboratory@gmail.com | 24 |  | 2024-01-15 | GX-XVI-D-6C | 842226 | Module B3 for replacement For CPU Replacement | Replaced module B3 and CPU | Performed maintenance Replaced module B3 Replaced CPU Generated reports | All modules operational. |  | Old: B3/885762,Backup laptop/CPH1019F03472 New: B3/766856 New CPU: CPH1023F45878 | G 6736 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816106104129987173/5816106104129987173_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5816106104129987173/5816106104129987173_signature_31.png
| 2024-01-21 | 2024-01-18 12:43 | Claro M. Recto Memorial District Hospital | Infanta | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27028 | PBSP | katelynaveryumerez@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1451 | 2024-01-18 | GX-IV-D-6C | 807608 | Module C4 for replacement | Replaced module C4 | Performed maintenance Replaced module C4 Generated reports | All modules operational. |  | Old: C4/210068710 New: C4/717884 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816105424127244026/5816105424127244026_signature_28.png | Lorelyn Miras | https://www.jotform.com/uploads/macarewecare/222927978372067/5816105424127244026/5816105424127244026_signature_31.png
| 2024-01-21 | 2024-01-19 13:50 | Pililla Rural Health Unit | M.A. Roxas St. Pililla | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27029 | DOH | tbdotspililla@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 708 | 2024-01-08 | GX-IV-D-6C | 821293 | Module B4 for replacement | Replaced module B4 | Performed maintenance Replaced module B4 Generated reports | All modules operational. |  | Old: B4/670429 New: B4/210044515 | DR# 11069 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816105424125656730/5816105424125656730_signature_28.png | Pililla RTDL | https://www.jotform.com/uploads/macarewecare/222927978372067/5816105424125656730/5816105424125656730_signature_31.png
| 2024-01-21 | 2024-01-19 11:26 | Jala-Jala Rural Health Unit | Jala-Jala | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 27028 | DOH | loisitalibertad@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 353 | 2024-01-08 | GX-IV-D-6C | 820184 | For monitor replacement | Replaced defective monitor | Performed maintenance Replaced defective monitor Generated reports | All modules operational |  | Old monitor: CNK831071T New monitor: AVI215CMQY05B01060027 | DR# 11052 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5816105424128825217/5816105424128825217_signature_28.png | Loisita Libertad | https://www.jotform.com/uploads/macarewecare/222927978372067/5816105424128825217/5816105424128825217_signature_31.png
| 2024-01-20 | 2024-01-19 14:48 | San Mateo Super Health Center | Guitnang bayan,San Mateo | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 27027 | PBSP | wilhelminanartatez@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1556 | 2024-01-15 | GX4D | 808616 | For module replacement A3 667542 - 754442 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | 900-0491RS | G6742 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5815337150647932792/5815337150647932792_signature_28.png | Wilhelmina R. Nartatez | https://www.jotform.com/uploads/macarewecare/222927978372067/5815337150647932792/5815337150647932792_signature_31.png
| 2024-01-19 | 2024-01-19 12:01 | Lake Sebu Rural Health Unit | Poblacion, Lake Sebu, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27026 | DOH | jennypingkian23@yahoo.com | 28.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 683 | 2024-01-19 | Gx-IV | 838569 | Module 4 h/w failed | Replace defective module 4 SN:862119 with new module SN:756995 Genexpert Cleaning Checking of errors Data collection: Pre and Post Installation qualification report, System log, Assay statistics report and Archive test | Perform Genexpert machine Preventive maintenance Perform Genexpert machine module replacement/repair | 4 operational modules |  | Genexpert 6 color module SN: 756995 | DR number: 11026 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814761901212361258/5814761901212361258_signature_28.png | Jenny A. Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/5814761901212361258/5814761901212361258_signature_31.png
| 2024-01-19 | 2024-01-19 19:07 | Marantao Rural Health Unit | Marawi City | 8 | Lanao del Norte | Laboratory | YES | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 27025 | DOH | sorayajrm19@yahoo.com | 23 |  | 2024-01-18 | Gx4 | 838549 | Module D1 with serial number 869706 calibration fail. | Replace module D1 | Replace module D1 with new serial number 739641 | All modules are functional. | CPU with serial number CPH0320F05903 are not functioning schedule for repair. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814720102468047451/5814720102468047451_signature_28.png | Soraya JM. Muti | https://www.jotform.com/uploads/macarewecare/222927978372067/5814720102468047451/5814720102468047451_signature_31.png
| 2024-01-19 | 2024-01-19 18:52 | Amai Pakpak Medical Center RTDL | Marawi City | 8 | Lanao del Norte | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 27024 | PBSP | rbarco2007@gmail.com | 23 |  | 2024-01-18 | Gx4 | 807457 | Module C4 with serial number 771991 hardware fail. | Save Files | Save Installation Qualification, SystemLog | Module C1, C2, C3 ok to use | Cpu running to slow while opening files and scaner also not functioning the trigger sometimes. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814711692465573687/5814711692465573687_signature_28.png | Raisah P. Barco | https://www.jotform.com/uploads/macarewecare/222927978372067/5814711692465573687/5814711692465573687_signature_31.png
| 2024-01-19 | 2024-01-19 18:45 | Amai Pakpak Medical Center RTDL | Marawi city | 8 | Lanao del Norte | Laboratory | YES | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 27023 | DOH | rbarco2007@gmail.com | 23 |  | 2024-01-18 | Gx4 | 838293 | Module B1 serial number 781987 hardware fail. Module B4 calibration fail last year 2023 | Replace module B1 | Replace module B1 with new serial number 210078829 | Module B1, B2, B3 ok to use. Module B4 not ok to use always error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814707522468783255/5814707522468783255_signature_28.png | Raisah P. Barco | https://www.jotform.com/uploads/macarewecare/222927978372067/5814707522468783255/5814707522468783255_signature_31.png
| 2024-01-19 | 2024-01-19 13:18 | Southern Philippines Medical Center | J.P Laurel , Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27022 | MACARE-HBDC | spmcrhivda@gmail.com | 23.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1672 ASSAY: Xpert HIV-1 Viral Load, COUNT: 790 |  | GX-IV 4D | 839152 | Error at Modules B1, B3 and B4 | Run functionality test for error Modules | Perform Functionality Test | B1 and B4 for service required only B3 Passed the functionality test | Modules B1, B2, B4 excluded from Test |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814699352515345676/5814699352515345676_signature_28.png | Maria Ronna Grace B. Seprado RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5814699352515345676/5814699352515345676_signature_31.png
| 2024-01-19 | 2024-01-19 13:29 | Southern Philippines Medical Center | JP Laurel Bajada Davao City | 11 | Davao del Sur | TB-HIV | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 27021 | PBSP | spmcrhivda@gmail.com | 28.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2267 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1003 |  | GX-IV 4D | 808298 | Module B3 for error | Run Functionality Test for Module B3 | Perform Functionality Test on Module B3 | Module B3 passed the functionality test. Module B3 for under observation | Open a new MTB Ultra Cartidge for error verification |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814698272519019592/5814698272519019592_signature_28.png | Maria Ronna Grace B. Seprado | https://www.jotform.com/uploads/macarewecare/222927978372067/5814698272519019592/5814698272519019592_signature_31.png
| 2024-01-19 | 2024-01-19 14:21 | Labuan General hospital | Labuan, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27020 | DOH | jenne060213@gmail.com | 27 |  | 2023-12-20 | GX 4 | 838568 | Module 3 temperature error | Replaced module 862916 with Module 768535 | Module Replacement and Preventive | All module are available | None | 900-0491RS | 10993 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814603007674290098/5814603007674290098_signature_28.png | Jennerose Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5814603007674290098/5814603007674290098_signature_31.png
| 2024-01-19 | 2024-01-19 14:21 | Labuan General hospital | Labuan, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27019 | DOH | jenne060213@gmail.com | 27 |  | 2023-12-20 | GX 4 | 838568 | Module 3 temperature error | Replaced module 862916 with Module 768535 | Module Replacement and Preventive | All module are available | None | 900-0491RS | 10993 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814602807675910667/5814602807675910667_signature_28.png | Jennerose Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5814602807675910667/5814602807675910667_signature_31.png
| 2024-01-19 | 2024-01-16 12:53 | Cuyapo Rural Health Unit | Cuyapo Nueva Ecija | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27018 | DOH | jolenerodico09171998@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1156 | 2024-01-05 | GXIV-D | 831672 | A3 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performwd functionality test | Unit is operational |  | 1 genexpert moduke | 11048 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814573353021225754/5814573353021225754_signature_28.png | Jolene Rodico | https://www.jotform.com/uploads/macarewecare/222927978372067/5814573353021225754/5814573353021225754_signature_31.png
| 2024-01-19 | 2024-01-16 12:53 | Cuyapo Rural Health Unit | Cuyapo Nueva Ecija | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 27017 | DOH | jolenerodico09171998@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1156 | 2024-01-05 | GXIV-D | 831672 | A3 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performwd functionality test | Unit is operational |  | 1 genexpert moduke | 11048 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814572943022935731/5814572943022935731_signature_28.png | Jolene Rodico | https://www.jotform.com/uploads/macarewecare/222927978372067/5814572943022935731/5814572943022935731_signature_31.png
| 2024-01-19 | 2024-01-19 12:57 | Sta Rosa Community District Hospital | Sta Rosa, Olango, Lapulapu City | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27016 | DOH | ariestumz@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 250 |  | GX4 | 838534 | Due for monthly maintenance Icore clean, Plunger clean up, and filter clean up | performed maintenance in GeneXpert clean Icore , Plunger clean up, and filter clean-up | Conduct for monthly maintenance Icore clean, Plunger clean up, and filter clean up | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814497201115510932/5814497201115510932_signature_28.png | Ariston Tumulak, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5814497201115510932/5814497201115510932_signature_31.png
| 2024-01-19 | 2024-01-19 12:52 | PHILIPPINE RED CROSS - ILOILO CHAPTER | ILOILO CITY | 6 | Iloilo | BLOOD BANK | NO | REPAIR | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27015 | COMMERCIAL | rlcerbas@gmail.com | 22 |  |  | MCS+ | 98G049 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814495337326332679/5814495337326332679_signature_28.png | Ryan L Cerbas | https://www.jotform.com/uploads/macarewecare/222927978372067/5814495337326332679/5814495337326332679_signature_31.png
| 2024-01-19 | 2024-01-19 12:42 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 27014 | DOH | chaska.lp14@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1911 |  | GX4 | 838283 | Repeated occurrence of error 2008 in all modules | Extend time in specimen mixing process for sticky samples and review the sample preparation. | Troubleshoot and clean up Icore, plunger, clean filter, and generate GXX, IQR, and Sys Log | All units are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814489901116607714/5814489901116607714_signature_28.png | Lucille Mae A. Potoy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5814489901116607714/5814489901116607714_signature_31.png
| 2024-01-19 | 2024-01-19 09:18 | Lopez Rural Health Unit | Lopez, Quezon Province | 4A | Quezon | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 27013 | DOH | janette_villasanta@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 822 | 2024-01-12 | GX IV 6C -D | 831889 | Module C1 for replacement | Performed PM Replaced Module C1 Performed reliability test Generated report | Replaced Module C1 783221 -> 635746 | Unit is operational All modules working |  | GX Module 6C | DR 11094 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814368088023806031/5814368088023806031_signature_28.png | Janette Villasanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5814368088023806031/5814368088023806031_signature_31.png
| 2024-01-19 | 2024-01-18 11:57 | Makati Medical Center | Amorsolo st Makati City | NCR | NCR 4th - Makati | Bloodbank | YES | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | troxoctubre@gmail.com |  | 27012 | COMMERCIAL | Blood.Bank@makatimed.net.ph | 27 |  | 2024-01-16 | MCS + | 97H330 | Blue pinch valve encounter error, malfunction of open and close of pinch valve due to wear and tear. | Checked pinch valve Checked tubings Checked leakage Installed spare pinch valve for troubleshooting of defective parts.After installed the spare pinch valve the units is operational.no error encountered Performed functionality test | Troubleshooting | Unit unoperational, need to replace blue pinch valve | For quotation of blue pinch valve | Pinch valve assy | 11125 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5814226442291184792/5814226442291184792_signature_28.png | Blood Bank | https://www.jotform.com/uploads/macarewecare/222927978372067/5814226442291184792/5814226442291184792_signature_31.png
| 2024-01-18 | 2024-01-18 15:51 | Bureau of Animal Industry | Visayas Ave., Quezon City | NCR | NCR 2nd - Quezon City | Bacteriology | YES | MAINTENANCE | CARL ZEISS | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 27011 | COMMERCIAL | addrl.bacte@yahoo.com | 23 |  |  | Axioscope Standard 20 | 8589 | Unit for preventive maintenance and cleaning | None | Performed functionality test Cleaned optics and microscope body Performed optical alignment | Unit is operational for brightfield use | Camera, camera controller, and mercury light source are obsolete and unusable |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813743557515448354/5813743557515448354_signature_28.png | Revelyn R. Suyat | https://www.jotform.com/uploads/macarewecare/222927978372067/5813743557515448354/5813743557515448354_signature_31.png
| 2024-01-18 | 2024-01-18 15:09 | Quezon City Jail | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27010 | PBSP | carleecruz0@gmal.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: na |  | GXIV | 838881 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813714882324430296/5813714882324430296_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5813714882324430296/5813714882324430296_signature_31.png
| 2024-01-18 | 2024-01-18 15:11 | Quezon City Jail | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 27010 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: na |  | GXIV | 803089 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational Module 2 requires service |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813714882324655501/5813714882324655501_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5813714882324655501/5813714882324655501_signature_31.png
| 2024-01-18 | 2024-01-18 15:05 | Bureau of Animal Industry | Visayas Ave., Quezon City | NCR | NCR 2nd - Quezon City | Parasitology | NO | TROUBLESHOOTING | CARL ZEISS | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 27009 | COMMERCIAL | michelle.barcelona366@gmail.com | 24 |  |  | Primo Star | 3144017776 | Light source is flickering | Verified functionality of light source | Check-up and troubleshooting | Recommend replacement of LED light source |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813710157511135287/5813710157511135287_signature_28.png | Michelle Ann Barcelona | https://www.jotform.com/uploads/macarewecare/222927978372067/5813710157511135287/5813710157511135287_signature_31.png
| 2024-01-18 | 2024-01-18 13:41 | UP Diliman - Institute of Biology | Commonwealth Ave, Quezon City | NCR | NCR 2nd - Quezon City | MCCL | NO | REPAIR | CYTIVA | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 27008 | COMMERCIAL | aamanua@up.edu.ph | 22 |  |  | Plasmair Sentinel | 2006-TS-0015 | Unable to start operation due to error | Updated machine firmware | Updated machine firmware Performed functionality test | Unit is operational | Recommend replacing filters as they have exceeded manufacturer-recommended lifetime |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813670361492208384/5813670361492208384_signature_28.png | Al James A. Manua | https://www.jotform.com/uploads/macarewecare/222927978372067/5813670361492208384/5813670361492208384_signature_31.png
| 2024-01-18 | 2024-01-18 13:30 | Casiguran Rural Health Unit | Poblacion, Casiguran, Aurora | 3 | Aurora | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 27007 | DOH | cezalen@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 951 | 2023-12-12 | Genexpert IV | 831844 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement (699723 - 807311) Performed Functionality Test | Unit is Operational | N/a | Genexpert Module - 807311 | 10902 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813653117612272334/5813653117612272334_signature_28.png | Princess Alen Cabunoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5813653117612272334/5813653117612272334_signature_31.png
| 2024-01-18 | 2024-01-18 13:19 | Zamboanga City Health Office | Pettit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 27006 | PBSP | floresteresita25@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 271 | 2024-01-18 | GX 4 | 807599 | Module 2 h/w failed | Replaced module 892103 with module 868378 | Module Replacement and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813652698738014121/5813652698738014121_signature_28.png | Teresita Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5813652698738014121/5813652698738014121_signature_31.png
| 2024-01-18 | 2024-01-18 11:14 | Bureau of Animal Industry | Visayas Ave., Quezon City | NCR | NCR 2nd - Quezon City | Chemical and Feeds Analysis Section | YES | MAINTENANCE | CARL ZEISS | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 27005 | COMMERCIAL | leo.yambot@gmail.com | 28 |  |  | Primo Star | 3116006792 | None, for preventive maintenance and cleaning | None | Performed functionality test Cleaned optics and microscope body Performed optical alignment | Unit is operational | Recommended adding a blue light filter for more comfortable viewing |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813576789217994577/5813576789217994577_signature_28.png | Leonardo Dante P. Yambot | https://www.jotform.com/uploads/macarewecare/222927978372067/5813576789217994577/5813576789217994577_signature_31.png
| 2024-01-18 | 2024-01-18 11:01 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | JARO, ILOILO CITY | 6 | Iloilo | TB DOTS | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27004 | PBSP | tvanrlujan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3438 |  | GX-IV | 804699 | -PREVENTIVE MAINTENANCE -MODULE D4 SN 210068572 FAILED : ERROR 1002,4009 | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | MODULE D4 FOR REPORTING |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813563241747743180/5813563241747743180_signature_28.png | Van Ryan Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5813563241747743180/5813563241747743180_signature_31.png
| 2024-01-18 | 2024-01-18 10:57 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | JARO, ILOILO CITY | 6 | Iloilo | LABORATORY | NO | TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27003 | PBSP | tvanrlujan@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3438 |  | GX-IV | 707531 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL, FOR SERVICE OF FAN/PSU COMPONENT |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813560611748875294/5813560611748875294_signature_28.png | Van Ryan Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5813560611748875294/5813560611748875294_signature_31.png
| 2024-01-18 | 2024-01-18 09:48 | Bureau of Animal Industry | DA Compound, Visayas Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Veterenary Laboratory Division | YES | MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 27002 | COMMERCIAL | kpmalabas@gmail.com | 25 |  |  | Axio Scope.A1 | 3335000463 | None | None | Performed cleaning of external surfaces Performed cleaning of lenses, objectives, eyepieces and condenser Performed functionality check of the microscope Performed functionality check of the cameras with the Labscope software Performed kohler illumination alignment Performed overall preventive maintenance Performed overall system functionality checks | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813534554225926713/5813534554225926713_signature_28.png | Krista Mae Malabad | https://www.jotform.com/uploads/macarewecare/222927978372067/5813534554225926713/5813534554225926713_signature_31.png
| 2024-01-18 | 2024-01-18 08:27 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 27001 | COMMERCIAL | crmcbloodcenter2016@gmail.com | 23 |  |  | MCS+ Heamonetics | 15HO42G | N/A | MCS+ Heamonetics machine Cleaning Checking of errors and parts functionality | Perform MCS+ Heamonetics Preventive maintenance | In good condition/functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813512320613235207/5813512320613235207_signature_28.png | CRMC- Blood Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5813512320613235207/5813512320613235207_signature_31.png
| 2024-01-18 | 2024-01-18 09:04 | WESTERN VISAYAS UNIVERSITY MEDICAL CENTER | JARO, ILOILO CITY | 6 | Iloilo | MOLECULAR LABORATORY | NO | TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 27000 | COMMERCIAL | zielglet_09@yahoo.com | 26 | ASSAY: , COUNT: N/A |  | GX-IV | 110021716 | MODULES NOT DETECTED | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813493320008748839/5813493320008748839_signature_28.png | Glezie T. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5813493320008748839/5813493320008748839_signature_31.png
| 2024-01-18 | 2024-01-18 08:51 | Oriental Mindoro Southern District Hospital | Brgy. Odiong, Roxas Oriental Mindoro | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26999 | PBSP | markdisonmartinez@yhoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2469 | 2024-01-18 | GXIV | 807590 | Unit for PM. For module replacement of module B3. For installation of new gateway board. | Performed PM. Performed module replacement. Performed installation of gateway board. Generated reports. | PM, installation and module replacement. | Unit is operational. |  | GX module B3 (701116) in replacement to defective B3 (905814). | Monarch DR#0045749 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813489343917937861/5813489343917937861_signature_28.png | Mark Edison G. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5813489343917937861/5813489343917937861_signature_31.png
| 2024-01-18 | 2024-01-17 09:23 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26998 | DOH | marisatuansi123@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 633 | 2024-01-17 | Gx-IV | 838528 | Module 2 and 4 h/w failed | Replace/repair defective module 2 SN:862010 with new module SN: 749865 Replace/repair defective module 4 SN:673141 with new module SN: 750083 Genexpert machine Cleaning Data collection: Pre and Post IQR, System log, Assay statistics report, and archive test | Perform Genexpert machine Preventive maintenance Perform Genexpert machine module replacement/repair | 4 operational modules |  | Genexpert module 6 color SN:749865 and SN:750083 | DR number: 10987 and 10978 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813468350611990509/5813468350611990509_signature_28.png | Marisa Tuansi | https://www.jotform.com/uploads/macarewecare/222927978372067/5813468350611990509/5813468350611990509_signature_31.png
| 2024-01-18 | 2024-01-17 13:57 | Cotabato Sanitarium | Sultan Kudarat, Maguindanao | BARMM | Maguindanao | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26998 | PBSP | cotsanlaboratory@gmail.com | 29.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1059 |  | Gx-IV | 815277 | Defective Genexpert machine monitor | Replace defective monitor SN:611NTYT6B386 with new monitor SN:V5GAH836 Genexpert machine Cleaning Checking of errors Data collection: Installation qualification report, System log, Assay statistics report and Archive test | Perform Genexpert machine Preventive maintenance Perform Genexpert machine monitor Replacement/repair | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5813468350616969005/5813468350616969005_signature_28.png | Carolyn R, Paggao | https://www.jotform.com/uploads/macarewecare/222927978372067/5813468350616969005/5813468350616969005_signature_31.png
| 2024-01-17 | 2024-01-17 12:16 | TEODORO B. GALAGAR DISTRICT HOSPITAL | Jagna Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26997 | DOH | applecpalma09@yahoo.com | 29 |  |  | Gx 4 | 838292 | Defective module b2 735121 | New module installed B2 786200 | New module installed module b2 786200 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812939447424073263/5812939447424073263_signature_28.png | Therese Arjelin Palma | https://www.jotform.com/uploads/macarewecare/222927978372067/5812939447424073263/5812939447424073263_signature_31.png
| 2024-01-17 | 2024-01-17 10:13 | Candijay Community Hospital | Candijay, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26996 | DOH | angelinaredulla@yahoo.com.ph | 28 |  |  | Gx4 | 838396 | Defective module B1 860992 | New module installed B1 786121 | New module installed B1 786121 | Module B2 HW fail |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812934547429086194/5812934547429086194_signature_28.png | ANGELINA O. REDULLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5812934547429086194/5812934547429086194_signature_31.png
| 2024-01-17 | 2024-01-17 15:41 | Sariaya Rural Health Unit | Gomez St., Sariaya, Quezon | NCR | Quezon | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26995 | DOH | smcl_2012@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1947 | 2024-01-12 | GX IV 6C -D | 831890 | Module A2 and A3 for replacement | Performed PM Replaced modules A2 and A3 Performed reliability test Generated report | Replaced modules A2: 784340 -> 210005873 A3: 784372 -> 210040526 | Unit is operational 3 modules working 1 module for reporting |  | GX modules 6C | DR 11092, 11093 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812879050811306056/5812879050811306056_signature_28.png | rhea parreño | https://www.jotform.com/uploads/macarewecare/222927978372067/5812879050811306056/5812879050811306056_signature_31.png
| 2024-01-17 | 2024-01-17 11:38 | Capitol Medical Center | Scout Magbanua, Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26994 | COMMERCIAL | gfalcantara.rmt@gmail.com | 24 |  |  | MCS+ | 07C002 | Unit is due for Preventive Maintenance | Preventive Maintenance | Checked Sensors, Valves, Rollers, Pressure and Centrifuge Checked A/D Values Performed Simulation Run Cleaned Surface Area | Unit is Operational | Defective UPS Backup Function. For Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812846553613902468/5812846553613902468_signature_28.png | Gretha Faye B. Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5812846553613902468/5812846553613902468_signature_31.png
| 2024-01-17 | 2024-01-16 12:39 | Asian Hospital and Medical Center | Civic Dr.,Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | rpescador@asianhospital.com |  | 26993 | COMMERCIAL | angelgimang17@gmail.com | 24 |  |  | MCS+ | 10B129 | Unit is due for Preventive Maintenance | Preventive Maintenance | Checked Pumps,Valves, Pressure, Sensors and Centrifuge Checked A/D Values Cleaned Surface Unit Performed Simulation Run | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812846553618031749/5812846553618031749_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5812846553618031749/5812846553618031749_signature_31.png
| 2024-01-17 | 2024-01-16 12:31 | Asian Hospital and Medical Center | Civic Dr., Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | rpescasdor@asianhospital.com |  | 26992 | COMMERCIAL | angelgimang17@gmail.com | 24 |  |  | MCS+ | 01F116 | Unit is due for Preventive Maintenance | Preventive Maintenance | Checked Pumps, Valves, Sensors, Pressure and Centrifuge Checked A/D Values if Pass Performed Simulation Test Cleaned Surface Area | Unit is Operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812846543613801412/5812846543613801412_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5812846543613801412/5812846543613801412_signature_31.png
| 2024-01-17 | 2024-01-16 10:17 | Asian Hospital And Medical Center | Civic Drive, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular Laboratort | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com | rpescador@asianhospital.com |  | 26991 | COMMERCIAL | agalang@asianhospital.com | 24 | ASSAY: Others, COUNT: 43 ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert HPV, COUNT: 81 ASSAY: Xpert MTB/RIF Ultra, COUNT: 477 ASSAY: Xpert HIV-1 Viral Load, COUNT: 5 |  | GXIV-D-10C | 110014996 | Unit is due for Preventive Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod, Icore, and Filter Checked System Log for Errors Performed Self-test Generate Reports | All Modules are Operationsl. Unit is Good |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812846543615063593/5812846543615063593_signature_28.png | Alyssa Galang | https://www.jotform.com/uploads/macarewecare/222927978372067/5812846543615063593/5812846543615063593_signature_31.png
| 2024-01-17 | 2024-01-17 14:49 | Mabalacat Rural Health Unit | Mabalacat Pampanga | 3 | Pampanga | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26990 | PBSP | perezriyadh@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1302 ASSAY: Xpert HIV-1 Viral Load, COUNT: 56 | 2024-01-08 | GXII-D | 835104 | Monitor for replacement Defective newly delivered monitor | For reporting | Performed preventive maintenance Gathered reports Performed functionality test | Unit is operational, monitor for replacement |  | Lenovo monitor | 45916 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812833617512576457/5812833617512576457_signature_28.png | Riyadh L. Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5812833617512576457/5812833617512576457_signature_31.png
| 2024-01-17 | 2024-01-17 14:44 | Mabalacat Rural Health Unit | Mabalacat Pampanga | 3 | Pampanga | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26989 | PBSP | perezriyadh@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 33 | 2024-01-11 | GXIV-D | 807564 | CPU for replacement CPU SN:2UA52027WB | Installation and setup | Performed preventive ,aintenance Installation and setup of new CPE Gathered reports Performed functionality test | Unit is operational |  | CPU: COH1023F45889 | 46019 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812830807517306265/5812830807517306265_signature_28.png | Riyadh L. Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5812830807517306265/5812830807517306265_signature_31.png
| 2024-01-17 | 2024-01-17 14:40 | National Center For Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26988 | PBSP | iancristrono@gmail.com | 33 | ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1626 | 2024-01-15 | GXIV | 808625 | Module B3 for replacement | #NAME? | Replaced B3 | Unit is operational Fan filter for request to PBSP |  |  | G 6739 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812829695327547677/5812829695327547677_signature_28.png | Ian Cris Trono | https://www.jotform.com/uploads/macarewecare/222927978372067/5812829695327547677/5812829695327547677_signature_31.png
| 2024-01-17 | 2024-01-17 13:45 | Atang dela Rama Health Center | Gagalangin Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26987 | DOH | monskie_22@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1301 |  | GXIV | 831644 | Unit for PM. For installation of new battery to UPS. | Performed PM. Performed repair of UPS. Generated reports | PM and repair. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812798423128556301/5812798423128556301_signature_28.png | Monaliza Balcos | https://www.jotform.com/uploads/macarewecare/222927978372067/5812798423128556301/5812798423128556301_signature_31.png
| 2024-01-17 | 2024-01-17 13:45 | Atang dela Rama Health Center | Gagalangin Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26986 | DOH | monskie_22@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1301 |  | GXIV | 831644 | Unit for PM. For installation of new battery to UPS. | Performed PM. Performed repair of UPS. Generated reports | PM and repair. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812797173127524739/5812797173127524739_signature_28.png | Monaliza Balcos | https://www.jotform.com/uploads/macarewecare/222927978372067/5812797173127524739/5812797173127524739_signature_31.png
| 2024-01-17 | 2024-01-17 12:40 | Victoria Luna Medical Center - Armed Forces of The Philippines Health Service Command | V. Luna, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26985 | MACARE-HBDC | mtvalderama@gmail.com | 30 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 707 ASSAY: Xpert SARS-CoV-2, COUNT: 2008 ASSAY: Xpert HCV VL, COUNT: 60 ASSAY: Others, COUNT: 49 ASSAY: Xpert HIV-1 Viral Load, COUNT: 80 |  | GXIV | 845462 | Stuck cartridge on B1, B2, B3 | #NAME? | Removed stuck cartridges | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812784325324806294/5812784325324806294_signature_28.png | Maria Teressa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/5812784325324806294/5812784325324806294_signature_31.png
| 2024-01-17 | 2024-01-17 11:18 | Pilar Rural Health Unit | Pilar Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26984 | DOH | aimna_24@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 691 |  | GXIV-D | 831613 | A2: 754415 error 1001 temperature error | Preventive maintenance Xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational, A2 under observation | A2 passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812711343743136465/5812711343743136465_signature_28.png | Marciana Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5812711343743136465/5812711343743136465_signature_31.png
| 2024-01-17 | 2023-12-06 13:27 | Tarlac Provincial Hospital | Hospital Drive Brgy. San Vicente Tarlac City | 3 | Tarlac | PCR | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26983 | PBSP | genextphmpl@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 1651 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1196 |  | GXIV-D | 838142 | B3:909438 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replavement Gathered reports Performed functionality test | Unit is operational | New B3 SN:672264 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812701703749725207/5812701703749725207_signature_28.png | Ethel Mae D. Avellanosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5812701703749725207/5812701703749725207_signature_31.png
| 2024-01-17 | 2024-01-17 10:32 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | YES | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 26982 | DOH | jonnifercodoyceniza@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 96 ASSAY: Xpert HCV VL, COUNT: 4 | 2024-01-17 | Gx4 | 829419 | Error code 5007,5015,5002 | Perform preventive maintenance | Clean all modules and plunger, change filter | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812686906944736653/5812686906944736653_signature_28.png | Jonnifer ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5812686906944736653/5812686906944736653_signature_31.png
| 2024-01-17 | 2024-01-17 10:16 | WESTERN VISAYAS MEDICAL CENTER | MANDURRIAO, ILOILO CITY | 6 | Iloilo | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26981 | PBSP | mitorilla.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4745 |  | GX-XVI | 839316 | -MODULE D3 ERROR 1001 AND 2014 | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -DATA COLLECTION FOR MODULE D3 SN:871127 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812672536189138869/5812672536189138869_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5812672536189138869/5812672536189138869_signature_31.png
| 2024-01-17 | 2024-01-17 10:10 | WESTERN VISAYAS MEDICAL CENTER | MANDURRIAO, ILOILO CITY | 6 | Iloilo | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26980 | PBSP | mitorilla.pbsp@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 1113 ASSAY: Xpert MTB/RIF Ultra, COUNT: 824 ASSAY: Xpert HIV-1 Viral Load, COUNT: 409 |  | GX-IV | 706910 | -MODULE A1 ERROR 1002 | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812669646181049905/5812669646181049905_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5812669646181049905/5812669646181049905_signature_31.png
| 2024-01-17 | 2024-01-17 10:02 | WESTERN VISAYAS MEDICAL CENTER | MANDURRIAO, ILOILO CITY | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26979 | WHO | mitorilla.pbsp@gmail.com | 26 | ASSAY: Others, COUNT: XDR - 910 |  | GX-IV | 848241 | #NAME? | #NAME? | -PERFORM PREVENTIVE MAINTENANCE PROCEDURE -PERFORM MODULE B4 REPLACEMENT : DEF SN:210002436 NEW SN:210068892 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812665776183522267/5812665776183522267_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5812665776183522267/5812665776183522267_signature_31.png
| 2024-01-17 | 2024-01-16 15:12 | Mauban Municipal Health Office | Signal St. Mauban | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26978 | MACARE-HBDC | jayveepasoot3@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HPV, COUNT: 143 ASSAY: Xpert MTB/RIF Ultra, COUNT: 602 |  | GX-IV-D-6C | 110015487 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812659153557288058/5812659153557288058_signature_28.png | Jayvee Pasoot | https://www.jotform.com/uploads/macarewecare/222927978372067/5812659153557288058/5812659153557288058_signature_31.png
| 2024-01-17 | 2024-01-17 07:22 | Cong. Toribio Semion Memorial Hospital | Carmen, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26977 | DOH | Himaya.vale@yahoo.com | 28 |  |  | Gx 4 | 838394 | Defective module 702990 | New module installed 784515 | New module installed 784515 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812657211121640706/5812657211121640706_signature_28.png | Himaya vale | https://www.jotform.com/uploads/macarewecare/222927978372067/5812657211121640706/5812657211121640706_signature_31.png
| 2024-01-17 | 2024-01-17 09:06 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 26976 | DOH | materesaaliman78@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1063 | 2024-01-17 | GXIV | 838406 | n/a | n/a | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812630809612407299/5812630809612407299_signature_28.png | ma.teresa aliman | https://www.jotform.com/uploads/macarewecare/222927978372067/5812630809612407299/5812630809612407299_signature_31.png
| 2024-01-17 | 2024-01-16 12:23 | Laoang Rural Health Unit | Laoang Northern Samar | 8 | Northern Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 26975 | DOH | aliceiytan@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1319 | 2024-01-16 | GXIV | 838202 | N/a | n/@ | Preventive Maintenance - Cleaned plunger rod, cartridge bay - Cleaned Fan filter - archive data | unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812627319612976608/5812627319612976608_signature_28.png | dr alice iy tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5812627319612976608/5812627319612976608_signature_31.png
| 2024-01-16 | 2024-01-16 08:14 | Dagupan City Health Office | Dagupan City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26974 | DOH | thisisformyfamily@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1464 | 2024-01-04 | GX4 | 838286 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 808254 >>> 750392 - Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 750392 | DR 11045 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812242330206046594/5812242330206046594_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5812242330206046594/5812242330206046594_signature_31.png
| 2024-01-16 | 2024-01-16 08:09 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26973 | DOH | arleneabad32@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 633 | 2024-01-04 | GX4 | 838288 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 808257 >>> 739420 - Generated reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module | DR 11046 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812237950201401711/5812237950201401711_signature_28.png | Arlene U. Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5812237950201401711/5812237950201401711_signature_31.png
| 2024-01-16 | 2024-01-16 15:47 | Governor Generoso Rural Health Unit | Poblacion, Governor Generoso | 11 | Davao Oriental | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26972 | DOH | josephmagallanes72@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 497 | 2024-01-05 | GX-IV 4D | 838490 | Module B1 and B2 for repair and due for Preventive Maintenance | Replaced the Defective Module SN: 862212 with new Module SN: 616039 Replaced the Defective Module SN: 862216 with new Module SN: 615764 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | 900-0491rs | DR 11033 & DR 11034 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812047095012482147/5812047095012482147_signature_28.png | Charlene V. Magallanes RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5812047095012482147/5812047095012482147_signature_31.png
| 2024-01-16 | 2024-01-16 15:39 | National Government Center Super Health Center | MRB Compound Pilot Drive, Brgy. Commonwealth, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26971 | DOH | tmcolisao@gmail.com | 25 |  | 2024-01-12 | GX IV 6C - D | 831586 | Module A2 for replacement CPU for troubleshooting; not booting up | Performed PM Checked cables Tightened data cable and RAM Problem persists Replaced HDD Replaced module A2 Generated report | Replaced module A2 754226 -> 784178 Troubleshoot CPU | unit is operational all modules working |  | GX module 6C | DR 11095 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812040060814990513/5812040060814990513_signature_28.png | Trecia Marie D. Colisao | https://www.jotform.com/uploads/macarewecare/222927978372067/5812040060814990513/5812040060814990513_signature_31.png
| 2024-01-16 | 2024-01-16 16:30 | Adela Serra Ty Memorial Medical Center | Capitol Hills, Brgy. Telaje, Tandag City | 13 | Surigao del Sur | RTDL | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26970 | PBSP | astmmclaboratory17@gmail.com | 28 |  |  | GX-4 | 804255 | Module B1 found Service Required Defective CPU | None | Replacing Defective CPU with New CPU (SN. CPH0923F45093) | 3 modules functional, 1 module service required CPU is now functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812035251612747607/5812035251612747607_signature_28.png | Editha M. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5812035251612747607/5812035251612747607_signature_31.png
| 2024-01-16 | 2024-01-16 16:24 | SARA RURAL HEALTH UNIT | SARA, ILOILO | 6 | Iloilo | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26969 | PBSP | dohro6hrh.sara@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 182 |  | GX-IV | 838200 | FOR MODULE B3 REPLACEMENT | #NAME? | -PERFORM REPLACEMENT OF MODULE B3 DEF SN:807217 WITH NEW MODULE SN:758436 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812029997327885743/5812029997327885743_signature_28.png | Candelaria Parrenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5812029997327885743/5812029997327885743_signature_31.png
| 2024-01-16 | 2024-01-16 14:50 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26968 | PBSP | bmcgenexpertlab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA | 2024-01-16 | GXIV | 834623 | Replacement CPU for installation | #NAME? | Installed CPU | Unit is operational Pulled out back up laptop |  | GeneXpert Desktop SN: CPH1023F45879 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5812029439014231809/5812029439014231809_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5812029439014231809/5812029439014231809_signature_31.png
| 2024-01-16 | 2024-01-16 14:32 | Mauban Rural Health Unit | Mauban | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26967 | PBSP | rhumauban.laboratory@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 152 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1177 | 2024-01-16 | GX-IV-D-6C | 814375 | For CPU Replacement | Replaced CPU | Performed maintenance Replaced CPU Archive test Import assay Generated reports | Unit is operational. |  | Old CPU: 2UA70329D8 New CPU: CPH0723F43757 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811983109351445985/5811983109351445985_signature_28.png | Rhandy Buraga,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5811983109351445985/5811983109351445985_signature_31.png
| 2024-01-16 | 2024-01-16 14:40 | Dingalan Community Hospital | Caragsacan, Dingalan, Aurora | 3 | Aurora | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26966 | DOH | bnjlvrz@9gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 571 | 2024-01-05 | Genexpert IV | 831621 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement (754456 - 757233) (754457 - 754439) Performed Functionality Test | Unit is Operational | N/a | Genexpert Module | 11044 & 11043 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811967375123936947/5811967375123936947_signature_28.png | Benjo A Alvarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5811967375123936947/5811967375123936947_signature_31.png
| 2024-01-16 | 2024-01-16 14:24 | Nabua Rural Health Unit | Nabua | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26965 | PBSP | fevalenzuela1965@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 764 |  | GX-4 | 834620 | GX relocated. Assistance for re installation of GX | Setup Gx: connected cables opened Gx, CPU, UPS, Monitor Cleaned cartridge bays, plunger rod, icore Generate reports | Setup GX Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811958242225133130/5811958242225133130_signature_28.png | Valenzuela, Fe L. | https://www.jotform.com/uploads/macarewecare/222927978372067/5811958242225133130/5811958242225133130_signature_31.png
| 2024-01-16 | 2024-01-16 13:02 | Calaanan Health Center | Nadurata St. Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26964 | DOH | asheilacatlirn@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1249 | 2024-01-15 | GXIV | 831892 | Unit for PM. For installation of new UPS and monitor. | Performed PM. Performed installation of UPS and monitor. Generated reports. | PM and installation. | Unit is operational. |  | Focus monitor (AVI215CMQY05B01060156) in replacement to defective monitor (CNK83224MH). Accupower 1KVA UPS (AT1K230830ST0014) in replacement to defective UPS (AT1000180910ST0289). | DR#11104 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811915963359402865/5811915963359402865_signature_28.png | Arlene Sheila T. catli | https://www.jotform.com/uploads/macarewecare/222927978372067/5811915963359402865/5811915963359402865_signature_31.png
| 2024-01-16 | 2024-01-16 12:35 | St Luke's Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | Laptop replacement | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | biomed.qc@stlukes.com.ph |  | 26963 | COMMERCIAL | cellularimmunology.qc@stlukes.com.ph | 28 |  | 2024-01-15 | GX4 L | 807155 | For laptop replacement unable to upload adf files because of o.s version | Replaced and installed laptop replacement s.n CNU1321BXY1. Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Laptop replacement | Operational | For pull out for repair dell laptop s.n GCN2SM1 | Hp laptop | 1111 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811899312716040623/5811899312716040623_signature_28.png | Roiland A. Baybayon | https://www.jotform.com/uploads/macarewecare/222927978372067/5811899312716040623/5811899312716040623_signature_31.png
| 2024-01-16 | 2024-01-16 11:53 | Bagong Pag-asa Health Center | 35 Road 9 cor. Road 10. Brgy. Bagong Pag-Asa | NCR | NCR 2nd - Quezon City | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26962 | DOH | erikatagorda18@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 664 | 2024-01-12 | GX IV 6C -D | 831597 | Module A1 for replacement | Performed PM Replaced module A1 Generated report | Replaced module A1 698524 -> 722419 | Unit is operational 3 modules working |  | GX module 6C | DR 11091 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811889980819115928/5811889980819115928_signature_28.png | erika lucas | https://www.jotform.com/uploads/macarewecare/222927978372067/5811889980819115928/5811889980819115928_signature_31.png
| 2024-01-16 | 2024-01-16 12:00 | Belmonte Health Center | Sampaloc Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26961 | PBSP | spsurell@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1250 | 2024-01-15 | GXIV | 808619 | Unit for PM. For module replacement of module B3. | Performed PM. Performed module replacement. Generated reports | PM and module replacement. | Unit is operational. |  | GX module B3 (791983) in replacement to defective B3 (653817). | G6738 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811870563352668392/5811870563352668392_signature_28.png | Clarinda Sharon P. Surell | https://www.jotform.com/uploads/macarewecare/222927978372067/5811870563352668392/5811870563352668392_signature_31.png
| 2024-01-16 | 2024-01-16 11:40 | Guiguinto Rural Health Unit | Guguinto, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26960 | PBSP | joie_calica@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA | 2024-01-16 | GXIV | 814165 | Workstation for replacement | #NAME? | Replaced workstation | Unit is operational |  | GeneXpert Workstation CPU SN: CPH0723F43759 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811861257715259948/5811861257715259948_signature_28.png | Ligaya Calica | https://www.jotform.com/uploads/macarewecare/222927978372067/5811861257715259948/5811861257715259948_signature_31.png
| 2024-01-16 | 2024-01-16 10:58 | Guimba Community Hospital | Guimba Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | marquezandrelyn18@gmail.com |  | 26959 | PBSP | laboratorygch@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 956 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 |  | GXIV-D | 807560 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational, all modules passed calibration Genexpert machine/ referigirator needs to change location due to high ambient temperature. CPU BIOS battery for replacement | Used 4 xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811851224522356530/5811851224522356530_signature_28.png | Andrelyn Nazar Marquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5811851224522356530/5811851224522356530_signature_31.png
| 2024-01-16 | 2024-01-16 10:46 | San Roque Rural Health Unit | San Roque Northern Samar | 8 | Northern Samar | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 26958 | DOH | lynnjarito@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 707 | 2024-01-16 | GXIV | 838498 | Module B4/862238 Defective | #NAME? | #NAME? | Unit is Operational |  | 1 Genexpert module 6 Color SN/905745 | DOh |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811830954022291749/5811830954022291749_signature_28.png | LYNN N JARITO | https://www.jotform.com/uploads/macarewecare/222927978372067/5811830954022291749/5811830954022291749_signature_31.png
| 2024-01-16 | 2024-01-16 10:09 | Victoria Rural Health unit | Brgy Bulo Victoria Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26957 | DOH | diazrmt@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 961 |  | GXIV-D | 831669 | Stuck cartridge | Preventive maintenance Stuck cartridge removal | Performed cartridge removal Performed preventive maintenance Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811803874528153243/5811803874528153243_signature_28.png | Alma Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5811803874528153243/5811803874528153243_signature_31.png
| 2024-01-16 | 2024-01-11 16:20 | Dr. Elvira Lagrosa Health Center (Duay) | Dolores Hidago St., Pasay Metro Manila | NCR | NCR 4th - Pasay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26956 | PBSP | rhemejetbotaquia@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1522 |  | GX IV 6c - D | 812833 | Module B2 invalid results | Performed PM Checked system log Archived reports Generated report | Troubleshoot module B2 Archived files | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811770073616301180/5811770073616301180_signature_28.png | Rheme Tañeza | https://www.jotform.com/uploads/macarewecare/222927978372067/5811770073616301180/5811770073616301180_signature_31.png
| 2024-01-16 | 2024-01-11 13:36 | Philippine Tuberculosis Society, Inc. | E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26955 | PBSP | sannymisorena@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1406 |  | GX IV 6C - D | 707527 | New machine to arrive Files for back up | Performed PM Backed up all reports Store files in CD Generated report | Backed up files stored files in CD | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811760963618594261/5811760963618594261_signature_28.png | sanny m isorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5811760963618594261/5811760963618594261_signature_31.png
| 2024-01-15 | 2024-01-15 14:58 | Sta. Ignacia Rural Health Unit | Sta ignacia Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 26954 | DOH | jdi0973@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 499 | 2024-01-05 | GXIV-D | 831909 | A2:758401 for module replacement A3 for re-calibration UPS for troubleshooting | Repair Xpertcheck Troubleshooting | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Performed xpertcheck Performed module replacement Performed UPS troubleshooting Generate reports Archived test results Performed functionality test | Unit is operational | Defective UPS not charging for replacement. | 1 genexpert module 1 xpertcheck kit | 11047 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811327963023232650/5811327963023232650_signature_28.png | Jesus D. Ibarra Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5811327963023232650/5811327963023232650_signature_31.png
| 2024-01-15 | 2024-01-15 17:58 | Medellin Rural Health Unit | Poblacion Medellin Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26953 | DOH | rhumedellin@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1307 |  | GX4 | 838506 | Defective Modules in B1 862252 and B3 662546 | Module replacement of defective modules | Module replacement of New Module SR 786303 and 786385 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811222823919677738/5811222823919677738_signature_28.png | Raul Galano, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5811222823919677738/5811222823919677738_signature_31.png
| 2024-01-15 | 2024-01-15 16:29 | Tomas Claudio (District A) Health Center | Camus Street, Corner Quirino Avenue. Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26952 | DOH | arleen_p_alconera@yahoo.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 688 | 2023-12-18 | GX-IV 4D | 838492 | Module B2 for repair and Due for Maintenance | Replaced the defective Module SN: 8662215 with new Module SN:791145 Clean Surface Area, Cartridge Bay, Fan Filters, Plunger Rod and Icore Lense Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | 900-0491RS | DR 10972 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811173179913653937/5811173179913653937_signature_28.png | Arleen Alconera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5811173179913653937/5811173179913653937_signature_31.png
| 2024-01-15 | 2024-01-15 13:04 | Tarlac Provincial Hospital | Hospital Drive San Vicente Tarlac City | 3 | Tarlac | PCR Labortory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | mark10venzon@gmail.com |  | 26951 | PBSP | molecularlab.tph@gmail.com | 26 | ASSAY: Xpert SARS-CoV-2, COUNT: 2069 |  | GXIV-D | 838142 | B2:644413 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational | B2:862237 new installed module |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5811058030518125090/5811058030518125090_signature_28.png | Camille Ann Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5811058030518125090/5811058030518125090_signature_31.png
| 2024-01-15 | 2024-01-11 11:36 | Alfonso Rural Health Unit | Burgos St., Alfonso, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26950 | DOH | aim_1225@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 784 | 2024-01-08 | GXIV-D-6C | 821279 | Module 1 with Temperature Related Problem | Module Replacement | Replaced Module 1 Performed Selft-test Performed Maintenance Generate Reports | All Modules are Operational | B1: 738405 -> 756445 | GeneXpert Module 6C | 11062 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810968073612145846/5810968073612145846_signature_28.png | Aim Arroyo | https://www.jotform.com/uploads/macarewecare/222927978372067/5810968073612145846/5810968073612145846_signature_31.png
| 2024-01-15 | 2024-01-12 09:59 | General Trias City Health Office | Hospital Dr., General Trias, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26949 | DOH | loremontoya1169@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3152 ASSAY: Xpert HIV-1 Viral Load, COUNT: 116 | 2024-01-08 | GXIV-D-6C | 821340 | Module 2 Failed in XpertCheck UPS with No Back-up | Module Replacement | Performed Maintenance Replaced Module 2 Checked UPS Output Voltage and Backup Function Replaced UPS Battery Performed Self-test Generate Reports | All Modules are Operational | A2: 626065 -> 671136 | GeneXpert Module | 11063 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810968073612543561/5810968073612543561_signature_28.png | Lorena R.Montoya | https://www.jotform.com/uploads/macarewecare/222927978372067/5810968073612543561/5810968073612543561_signature_31.png
| 2024-01-15 | 2024-01-09 13:09 | Molino 3 City Health Unit | Malinis Compound, Molino, Bacoor | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26947 | DOH | joanrnd@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1867 | 2024-01-08 | GXIV-D-6C | 821221 | Module 2 with Temperature Error | Module Replacement | Replaced Module 2 Performed Self-test Performed Maintenance Generate Reports | All Modules are Operational | B2: 784350 -> 724167 | GeneXpert Module | 11060 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810968073613009438/5810968073613009438_signature_28.png | Jonalyn Aranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5810968073613009438/5810968073613009438_signature_31.png
| 2024-01-15 | 2024-01-04 13:51 | Silang Rural Health Unit | Brgy. Poblacion, Silang, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26948 | DOH | bencq78@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1929 | 2023-12-20 | GXIV-D-6C | 820708 | Module 4 Failed in XpertCheck | Module Replacement | Replaced Module 4 Performed Maintenance Performed Self-test Generate Reports | All Modules are Operational | A4: 629212 -> 786530 | GX Module 6C | 10994 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810968073616863557/5810968073616863557_signature_28.png | Benessie madlangsakay | https://www.jotform.com/uploads/macarewecare/222927978372067/5810968073616863557/5810968073616863557_signature_31.png
| 2024-01-15 | 2024-01-15 09:34 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26946 | DOH | luzvimindanavidad2@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1092 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1092 | 2023-12-14 | GXIV | 838297 | Module B4 XpertCheck service required. | -Replace defective module B4/808286 with new module SN: 785115 | #NAME? | All modules are functional. |  | 900-0491RS | DR no. 10946 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810921988984911358/5810921988984911358_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5810921988984911358/5810921988984911358_signature_31.png
| 2024-01-15 | 2024-01-15 09:34 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26945 | DOH | luzvimindanavidad2@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1092 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1092 | 2023-12-14 | GXIV | 838297 | Module B4 XpertCheck service required. | -Replace defective module B4/808286 with new module SN: 785115 | #NAME? | All modules are functional. |  | 900-0491RS | DR no. 10946 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810921758987743117/5810921758987743117_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5810921758987743117/5810921758987743117_signature_31.png
| 2024-01-15 | 2024-01-10 13:55 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26944 | DOH | laureanjo@yahoo.com | 29 |  |  | Gx 4 | 838527 | Defective Module D3 791668 | New module D3 758219, clean fan filter | New module D3 758219 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810909433919271577/5810909433919271577_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5810909433919271577/5810909433919271577_signature_31.png
| 2024-01-15 | 2024-01-10 13:55 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26943 | DOH | laureanjo@yahoo.com | 29 |  |  | Gx 4 | 838527 | Defective Module D3 791668 | New module D3 758219, clean fan filter | New module D3 758219 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810908373914585508/5810908373914585508_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5810908373914585508/5810908373914585508_signature_31.png
| 2024-01-15 | 2024-01-15 09:12 | Cebu TB reference lab | Cebu TB Referende lab-SRP | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26942 | DOH | asmantiquilla.pbsp@gmail.com | 25 |  |  | Gx4 | 838263 | Defective module B1 693532 and B2 807729 | New module installed B1 786431 and B2 210014805 | New module installed B1 786431 and B2 210014805 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810907483914364901/5810907483914364901_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5810907483914364901/5810907483914364901_signature_31.png
| 2024-01-14 | 2024-01-11 14:42 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26941 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 20 | ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert SARS-CoV-2, COUNT: 9 ASSAY: Xpert HCV VL, COUNT: 2 |  | GX4 10 color | 110016112 | For XpertCheck and Preventive Maintenance | #NAME? | Performed XpertCheck and Preventive Maintenance | Unit is operational | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810339630203880893/5810339630203880893_signature_28.png | Ma. Victoria Teresa R. Mamuri | https://www.jotform.com/uploads/macarewecare/222927978372067/5810339630203880893/5810339630203880893_signature_31.png
| 2024-01-14 | 2024-01-11 12:12 | Itogon Rural Health Unit | Itogon, Benguet | CAR | Benguet | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26940 | DOH | ggdepayso@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 385 | 2023-12-14 | GX4 | 838282 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 807798 >>> 863306 - Generated Reports | #NAME? | Unit is operational 1 | Work with Engineer Venzon | GeneXpert Module 863306 | DR 10916 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810337420208193664/5810337420208193664_signature_28.png | Gloria Depayso | https://www.jotform.com/uploads/macarewecare/222927978372067/5810337420208193664/5810337420208193664_signature_31.png
| 2024-01-14 | 2024-01-10 15:13 | Bontoc General Hospital | Bontoc, Mountain Province | CAR | Mountain Province | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26939 | PBSP | mffanglayan@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1220 | 2024-01-10 | GX4 | 807582 | For Module Replacement and Troubleshooting | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 688593 >>> 714782 - Swapped module B3 and B2 and performed Reliability Test - Generated Reports | Performed Module Replacement and Troubleshooting | Unit is operational | Work with with Engineer Venzon. Module B3 for observation. | GeneXpert Module 714782 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810334660209319504/5810334660209319504_signature_28.png | Mary Lourdes Fanglayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5810334660209319504/5810334660209319504_signature_31.png
| 2024-01-14 | 2024-01-12 14:31 | Angono Rural Health Unit | Angono | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26938 | DOH | siennablanco@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1040 | 2024-01-08 | GX-IV-D-6C | 819864 | Module A2 for replacement | Replaced module A2 | Performed maintenance Replaced module A2 Generated reports | All modules operational. |  | Old: A2/784784 New: A2/754306 | DR# 11068 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810189385706850016/5810189385706850016_signature_28.png | Sienna Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5810189385706850016/5810189385706850016_signature_31.png
| 2024-01-14 | 2024-01-09 15:59 | Laoag City Health Office | Laoag, Ilocos Norte | 1 | Ilocos Norte | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26937 | DOH | denver.mahor@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 559 | 2024-01-05 | GX4 | 838416 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods, I-CORE - Replaced Defective Module 861111 >>> 753904 - Generated Reports | Performed Module Replacement | Unit is operational | Work with Engineer Venzon | GeneXpert Module 753904 | DR 11040 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810163030201641108/5810163030201641108_signature_28.png | Rodenver D. Mahor | https://www.jotform.com/uploads/macarewecare/222927978372067/5810163030201641108/5810163030201641108_signature_31.png
| 2024-01-14 | 2024-01-09 14:59 | Dingras District Hospital | Dingras, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26936 | PBSP | kenacudapas1022@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 667 | 2024-01-09 | GX4 | 807597 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Modules 658000>>>705668, 612388>>>705964 - Generated Reports | Performed Module Replacement | Unit is operational | Work with Engineer Venzon | GeneXpert Modules 705668, 705964 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810159130209964180/5810159130209964180_signature_28.png | KENADINE Cudapas | https://www.jotform.com/uploads/macarewecare/222927978372067/5810159130209964180/5810159130209964180_signature_31.png
| 2024-01-14 | 2024-01-09 12:48 | Mariano Marcos Memorial Hospital and Medical Center | Batac, Ilocos Norte | 1 | Ilocos Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26935 | PBSP | mmmhmc.molecularlab@gmail.com | 28 | ASSAY: Xpert C. difficile BT, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 26 ASSAY: Xpert HPV, COUNT: 175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1137 ASSAY: Xpert HBV VL, COUNT: 3598 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 11 |  | GX4 | 838168 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is operational | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810155430202451874/5810155430202451874_signature_28.png | Allan Fellizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5810155430202451874/5810155430202451874_signature_31.png
| 2024-01-14 | 2024-01-09 12:48 | Mariano Marcos Memorial Hospital and Medical Center | Batac, Ilocos Norte | 1 | Ilocos Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26934 | PBSP | mmmhmc.molecularlab@gmail.com | 28 | ASSAY: Xpert C. difficile BT, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 26 ASSAY: Xpert HPV, COUNT: 175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1137 ASSAY: Xpert HBV VL, COUNT: 3598 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 11 |  | GX4 | 838168 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is operational | Work with Engineer Venzon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810154380202239904/5810154380202239904_signature_28.png | Allan Fellizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5810154380202239904/5810154380202239904_signature_31.png
| 2024-01-14 | 2024-01-09 12:38 | Mariano Marcos Memorial Hospital and Medical Center | Batac, Ilocos Norte | 1 | Ilocos Norte | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26933 | DOH | mmmhmc.molecularlab@gmail.com | 29 | ASSAY: Xpert C. difficile BT, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 431 ASSAY: Xpert SARS-CoV-2, COUNT: 1489 ASSAY: Xpert HCV VL, COUNT: 3 | 2023-12-12 | GX4 | 831908 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Modules 645827 >>> 754288, 713828 >>> 701383, 654653 >>> 699719 - Generated Reports | Performed Module Replacement | Unit is operational | Work with Engr. Venzon | GeneXpert Modules 754288, 701383, 699719q | DR 11041, 10900, 10904 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810148300204317207/5810148300204317207_signature_28.png | Allan Fellizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5810148300204317207/5810148300204317207_signature_31.png
| 2024-01-14 | 2024-01-09 10:13 | Gabriela Silang General Hospital | Vigan, Ilocos Sur | 1 | Ilocos Sur | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26932 | PBSP | isphcovidlab21@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1277 ASSAY: Xpert SARS-CoV-2, COUNT: 342 ASSAY: Xpert HIV-1 Viral Load, COUNT: 103 | 2024-01-09 | GX4 | 807578 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 865209 >>> 705557 - Generated reports | Replaced Defective module | Unit is operational | Work with Engineer Venzon | GeneXpert Module | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810062840203375344/5810062840203375344_signature_28.png | Mary Alexis F Tahil | https://www.jotform.com/uploads/macarewecare/222927978372067/5810062840203375344/5810062840203375344_signature_31.png
| 2024-01-14 | 2024-01-09 08:37 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | La Union | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26931 | DOH | piaramuj@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 698 | 2023-12-14 | GX4 | 838270 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 795994 >>> 210068699 and 808128 >>> 210039876 - Generated Reports | Replaced Defective Module | Unit is operational | Work with Engineer Venzon | GeneXpert Modules | DR 10913 , DR 10917 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810058090201081761/5810058090201081761_signature_28.png | Pia Jumar T. Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/5810058090201081761/5810058090201081761_signature_31.png
| 2024-01-14 | 2024-01-09 08:37 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | La Union | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26930 | DOH | piaramuj@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 698 | 2023-12-14 | GX4 | 838270 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 795994 >>> 210068699 and 808128 >>> 210039876 - Generated Reports | Replaced Defective Module | Unit is operational | Work with Engineer Venzon | GeneXpert Modules | DR 10913 , DR 10917 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5810057020207649092/5810057020207649092_signature_28.png | Pia Jumar T. Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/5810057020207649092/5810057020207649092_signature_31.png
| 2024-01-12 | 2024-01-12 15:51 | Klinika Project 7 | Project 7, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26929 | PBSP | freyaoriana@yahoo.com | 25 | ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1739 | 2024-01-12 | GX IV 10C - D | 845446 | Module B2 and B4 for replacement | Performed PM Replaced modules b2 and b4 Generated report | Replaced modules B2: 899269 -> 210024871 B4: 899271 -> 210078213 | Unit is operational All modules working |  | GX Module 10C | G6714 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808590993975948911/5808590993975948911_signature_28.png | alyanna faye zerrudo | https://www.jotform.com/uploads/macarewecare/222927978372067/5808590993975948911/5808590993975948911_signature_31.png
| 2024-01-12 | 2024-01-12 14:57 | Graman Medical And Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | MAINTENANCE On site manual removal of stucked cartridge. | CEPHEID | Harold Toco |  | htoco@macare-medicals.com |  |  | 26928 | COMMERCIAL | gramanmolecularlab@gmail.com | 20 | ASSAY: Xpert HBV VL, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 473 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 | 2024-01-12 | GX4 6C | 110008530 | Stucked cartridge | Performed manual removal of stucked cartridge | Removed stucked cartridge Performed maintenance. | All 4 modules operational. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808525950416480988/5808525950416480988_signature_28.png | Maria Elizabeth de Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5808525950416480988/5808525950416480988_signature_31.png
| 2024-01-12 | 2024-01-12 15:02 | Bagong Barrio Health Center | Bagong Barrio, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26927 | DOH | cruzernitamay@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1292 | 2024-01-08 | GXIV | 831906 | Module A1, A3 for replacement | - replaced A1, A3 - checked and cleaned cartridge bay, plunger rod, - generated reports | Replaced A1, A3 | Unit is operational |  | GeneXpert Module SN: 677426, 684092 | 11085 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808523822611170895/5808523822611170895_signature_28.png | Ernita May Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5808523822611170895/5808523822611170895_signature_31.png
| 2024-01-12 | 2024-01-12 13:58 | Caloocan City Health Office | A. Mabini St., Caloocan City | NCR | NCR 3rd - Caloocan | laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26926 | PBSP | caloocancentrallaboratory@gmail.com | 29 |  | 2024-01-12 | GX IV 6C-D | 814154 | CPU for replacement | Performed PM Replaced CPU Checked Aircon Aircon error E8 after 30 minutes of running Aircon plug burnt Aircon needs troubleshooting asap GX machine can't function properly without AC Generated report | Replaced CPU 2UA7021MS6 -> CPH0923F44892 | Unit is operational Aircon for troubleshooting |  | GeneXpert Desktop | G6715 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808491403437444405/5808491403437444405_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5808491403437444405/5808491403437444405_signature_31.png
| 2024-01-12 | 2024-01-12 13:49 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26925 | DOH | renlizros11@gmail.com | 27.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1213 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 | 2024-01-12 | Gx-IV | 838570 | Module 2 and 3 H/W failed | Replace module 2 SN:679154 with new module SN:80759 Replace module 3 SN:708473 with new module SN:807731 Genexpert machine Cleaning Data collection: Pre and Post IQR, System log , Assay statistics report, and archive test | Perform Genexpert machine Preventive maintenance Perform Genexpert module replacement/repair | 4 operational modules |  | Genexpert module 6-color SN:807594 and 87731 | DR number: 10979 and 10985 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808486441916064145/5808486441916064145_signature_28.png | Renee Liz Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5808486441916064145/5808486441916064145_signature_31.png
| 2024-01-12 | 2024-01-12 13:31 | Catmon Health Center | Catmon, Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26924 | DOH | mariathes0218@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3284 ASSAY: Xpert HIV-1 Viral Load, COUNT: 107 | 2024-01-08 | GXIV | 831903 | Module C2, C4 for replacement | #NAME? | Replaced C2, C4 | Unit is operational |  | GeneXpert Module SN: 742879, 753902 | 11072 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808486182613506449/5808486182613506449_signature_28.png | Maria Theresa Nunez | https://www.jotform.com/uploads/macarewecare/222927978372067/5808486182613506449/5808486182613506449_signature_31.png
| 2024-01-12 | 2024-01-12 13:56 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26923 | PBSP | renlizros11@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1213 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 |  | Gx-IV | 815428 | N/A | Genexpert machine Cleaning Data collection: IQR, System log, Assay statistics report, and Archive test | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808484841915788948/5808484841915788948_signature_28.png | Renee Liz Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5808484841915788948/5808484841915788948_signature_31.png
| 2024-01-12 | 2024-01-12 13:51 | San Pedro City Health Office 1 | San Pedro Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26922 | DOH | venersantos0425@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 102 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1863 | 2024-01-08 | GXIV | 821224 | Unit for PM. For module replacement of module A4. | Performed PM. Performed module replacement. Generated reports | PM and module replacement. | Unit is operational. |  | GX module A4 (743364) in replacement to defective A4 (730041). | DR#11056 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808481545116893670/5808481545116893670_signature_28.png | Vener R. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5808481545116893670/5808481545116893670_signature_31.png
| 2024-01-12 | 2024-01-12 12:44 | Bago City Health Office | Bago City, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 26921 | DOH | swabteambago@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2487 |  | GX4 | 838486 | Defective modules B2/711021 and B3/862205. | #NAME? | - Replaced B2 and B3 with new modules with SN 745913 and 758221. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808479531193254030/5808479531193254030_signature_28.png | Grace L. Soria | https://www.jotform.com/uploads/macarewecare/222927978372067/5808479531193254030/5808479531193254030_signature_31.png
| 2024-01-12 | 2024-01-12 13:40 | Davao Reproductive Health and Wellness Center | Jacinto Street. Poblacion Davao | 11 | Davao del Sur | TB-HIV | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26920 | MACARE-HBDC | rhwclab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 102 ASSAY: Xpert HIV-1 Viral Load, COUNT: 577 | 2023-12-18 | GX-IV 4D | 839236 | Module B1 and B2 for repair | Replaced the Defective Module B1 SN:869012 with new Module SN: 210003270 Replaced the Defective Module SN:676723 with new Module SN: 210014848 Generate IQR, Assay Stat and System Log | Perform Module Replacement | All Modules are operational |  | 900-0491rs | DR 10973 & DR 10970 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808477072257837119/5808477072257837119_signature_28.png | Charlene Mae Dueňas | https://www.jotform.com/uploads/macarewecare/222927978372067/5808477072257837119/5808477072257837119_signature_31.png
| 2024-01-12 | 2024-01-12 12:58 | Binan City Health Office 2 | Binan Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26919 | DOH | eloicyparra@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1790 ASSAY: Xpert HPV, COUNT: 148 ASSAY: Xpert HIV-1 Viral Load, COUNT: 331 | 2024-01-08 | GXIV | 821278 | Unit for PM. For module replacement of A3. | Performed PM. Performed module replacement. Generated reports | PM and module replacement. | Unit is operational. |  | GX module A3 (757236) in replacement to defective A3 (730170). | DR#11057 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808449140417183853/5808449140417183853_signature_28.png | Eloicy Parra | https://www.jotform.com/uploads/macarewecare/222927978372067/5808449140417183853/5808449140417183853_signature_31.png
| 2024-01-12 | 2024-01-12 11:32 | Talacogon District Hospital | San Isidro, Talacogon | 13 | Agusan del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26918 | DOH | tdh.labdept@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1453 | 2023-12-14 | GX-4 | 838275 | C1, C2 and C4 found Service Required thru XPC | None | Replacing C1 with SN: 807851 to SN: 874275 C2 with SN: 807837 to SN: 870802 C4 with SN: 807799 to SN: 863529 | All modules and units are functional |  | Certified Pre-Owned GeneXpert Modules | 10933, 10932 and 10931 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808407100785339989/5808407100785339989_signature_28.png | Shyra Ericka T. Calamba | https://www.jotform.com/uploads/macarewecare/222927978372067/5808407100785339989/5808407100785339989_signature_31.png
| 2024-01-12 | 2024-01-12 11:42 | Santiago City Health Office | San Andres, Santiago City, Isabela | 2 | Isabela | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | cho.santiago@gmail.com |  | 26917 | DOH | cho.santiago@gmail.com | 25 |  | 2024-01-12 | Genexpert IV | 831718 | N/a | N/a | Installed Genexpert System Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808403584813778411/5808403584813778411_signature_28.png | Marichelle Zapanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5808403584813778411/5808403584813778411_signature_31.png
| 2024-01-12 | 2024-01-12 10:53 | Bangkal Health Center | Rodriguez St. Bangkal | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26916 | PBSP | kapid5sepe@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3135 |  | GX-IV-D-6C | 838875 | Module B2 has error 1002 | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Checked module B2 Generated reports | Unit is operational. Module B2 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808383019104434782/5808383019104434782_signature_28.png | Eduardo Sepe | https://www.jotform.com/uploads/macarewecare/222927978372067/5808383019104434782/5808383019104434782_signature_31.png
| 2024-01-12 | 2024-01-12 10:35 | Bangkal Health Center | Rodriguez St., Bangkal | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26915 | DOH | kapid5sepe@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3135 | 2024-01-08 | GX-IV-D-6C | 831875 | Modules A1 and A2 for replacement | Replaced module A1 and A2 | Performed maintenance Replaced module A1 and A2 Generated reports | Modules A1 and A2 operational. Module A3 not operational, for reporting. Module A4 not operational, reported already. |  | Old: A1/210020089, A2/862667 New: A1/757959, A2/757999 | DR#11066, 11067 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808382699108939910/5808382699108939910_signature_28.png | Eduardo Sepe | https://www.jotform.com/uploads/macarewecare/222927978372067/5808382699108939910/5808382699108939910_signature_31.png
| 2024-01-12 | 2024-01-03 11:14 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | YES | MAINTENANCE For software upgrade | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26914 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 24 | ASSAY: Xpert HPV, COUNT: 25 ASSAY: Xpert MTB/RIF Ultra, COUNT: 103 ASSAY: Xpert SARS-CoV-2, COUNT: 13335 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GX-XVI-6C-D | 838627 | For software upgrade | Performed maintenance, archive and purge tests | Performed maintenance Tried to upgrade GeneXpert software Generated reports Archive and purge tests | Unit is operational. Modules A1,A2,B3,B4,C2,D2,D4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808376669101522061/5808376669101522061_signature_28.png | Justine Lorenz V. Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5808376669101522061/5808376669101522061_signature_31.png
| 2024-01-12 | 2024-01-03 09:59 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26913 | PBSP | molecularlab@rmc.doh.gov.ph | 24 |  |  | GX-IV-D-6C | 814374 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808376399108144695/5808376399108144695_signature_28.png | Justine Lorenz V. Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5808376399108144695/5808376399108144695_signature_31.png
| 2024-01-12 | 2024-01-12 09:43 | Teresita L. Jalandoni Provincial Hospital | Silay City | 6 | Negros Occidental | Laboratory Dept. | YES | REPAIR | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com | 06emerson23@gmail.com |  | 26912 | DOH | mrvillan.pbsp@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3644 |  | GX4 | 838464 | Module B1 with SN 802411 failed during last XPC. | Conduct PM and saves files. | Installed the new module for B1 with SN 775357. | All modules are now in good working conditions. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808340964822616543/5808340964822616543_signature_28.png | Ma. Arna Villan | https://www.jotform.com/uploads/macarewecare/222927978372067/5808340964822616543/5808340964822616543_signature_31.png
| 2024-01-12 | 2024-01-12 09:56 | Balamban Rural Health Unit | Poblacion, Balamban Cebu | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com |  |  | 26911 | DOH | climaco.mildred@yahoo.com | 23 |  | 2024-01-12 | Gx4 | 838564 | Module A2 calibration fail. | Replace module | Replace module A2 with new module. | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808340872712297163/5808340872712297163_signature_28.png | Jonnel Barnayha | https://www.jotform.com/uploads/macarewecare/222927978372067/5808340872712297163/5808340872712297163_signature_31.png
| 2024-01-12 | 2024-01-12 09:12 | Ormoc City Health Office | Ormoc City, Leyte | 8 | Leyte | Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26910 | PBSP | theresesaladaga@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3363 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3363 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GXIV | 804253 | Module B3 stuck cartridge | #NAME? | #NAME? | All modules are functional. Module B1 has temp error history, for observation. | -ADF HIV-VL 2.0 Error 2097 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808324550753602024/5808324550753602024_signature_28.png | Ma. Therese A. Saladaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5808324550753602024/5808324550753602024_signature_31.png
| 2024-01-12 | 2024-01-11 11:00 | Tampakan Rural Health Unit | Poblacion Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26909 | DOH | jmedtech_lab@yahoo.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1045 | 2024-01-11 | Gx-IV | 838544 | Module 1 H/W failed | Replace defective module 1 SN:801766 with new module SN: 861471 Genexpert machine Cleaning Data collection: Pre and Post IQR, S.Log, Assay statistics report and Archive test | Perform Genexpert machine Preventive maintenance Perform module repair | 4 operational modules |  | Genexpert module 6 color SN:861471 | DR number: 10986 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5808302751217439093/5808302751217439093_signature_28.png | Jovi G Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/5808302751217439093/5808302751217439093_signature_31.png
| 2024-01-11 | 2024-01-11 17:01 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26908 | DOH | elma.bragat@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5547 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 | 2023-12-14 | GX-IV | 838510 | #NAME? | #NAME? | -PERFORM CLEANING OF GX EQUIPMENT AND WORKSTATION -PERFORM MODULE REPLACEMENT: C1 DEF SN: 862195 WITH NEW SN: 758210 C4 DEF SN: 861023 WITH NEW SN: 757488 | #NAME? |  | GENEXPERT MODULE | 1094810949 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807732548226938926/5807732548226938926_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5807732548226938926/5807732548226938926_signature_31.png
| 2024-01-11 | 2024-01-11 16:47 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26907 | DOH | elma.bragat@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5547 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 | 2023-12-14 | GX-IV | 838508 | #NAME? | #NAME? | -PERFORM CLEANING OF GX EQUIPMENT AND WORKSTATION -PERFORM MODULE REPLACEMENT: B2 DEF SN: 802814 WITH NEW SN: 754298 B4 DEF SN: 636240 WITH NEW SN: 757225 | EQUIPMENT IS OPERATIONAL | MODULE B1 DISABLED | GENEXPERT MODULE | 1093510944 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807728488226367601/5807728488226367601_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5807728488226367601/5807728488226367601_signature_31.png
| 2024-01-11 | 2024-01-10 17:00 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26906 | MACARE-HBDC | spmcrhivda@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1642 ASSAY: Xpert HIV-1 Viral Load, COUNT: 773 |  | GX-IV 4D | 839152 | For Preventive Maintenance and reported Stuck Cartidge in Module B4 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Removed the stuck cartidge | Perform Preventive Maintenance and Module Troubleshooting | Module B2 and B4 h/w failure | All datas are sent and reported to Jotform for further assessments |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807725776423036243/5807725776423036243_signature_28.png | Flordeliza M. Go RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5807725776423036243/5807725776423036243_signature_31.png
| 2024-01-11 | 2024-01-11 15:59 | Agusan del Norte Provincial Health Office | Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26905 | PBSP | ronniebalansag74@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1990 |  | GX-4 | 807461 | B2 found Service Required thru XPC Temperature Error E2014 | None | XPC Error Checking at Module B2 | 3 modules working, 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807694430288997028/5807694430288997028_signature_28.png | Ronnie Balansag | https://www.jotform.com/uploads/macarewecare/222927978372067/5807694430288997028/5807694430288997028_signature_31.png
| 2024-01-11 | 2024-01-11 15:50 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26904 | DOH | sajecopooh20@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 45 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1282 | 2023-12-14 | GXIV | 838555 | Module A1 and A3 defective | -Replace defective module A1/862891 with new module SN: 784638 -Replace defective module A3/862904 with new module SN: 738209 | #NAME? | Waiting for module A2 allocation. Other modules are functional. |  | 2 modules 900-0491RS | 10952 10969 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807689523681188174/5807689523681188174_signature_28.png | Sarah Jean R. Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5807689523681188174/5807689523681188174_signature_31.png
| 2024-01-11 | 2024-01-11 14:29 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-DOTS | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26903 | PBSP | spmcrhivda@gmail.com | 29.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2200 ASSAY: Xpert HIV-1 Viral Load, COUNT: 218 |  | GX-IV 4D | 808298 | Reported Module Errors at B3 and B4 | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log | Performed Module Maintenance and perform functionality test | Module B3 and B4 h/w failure | Additional Data reported to Jotform and Cepheid for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807683588612066502/5807683588612066502_signature_28.png | Flordeliza M. Go,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5807683588612066502/5807683588612066502_signature_31.png
| 2024-01-11 | 2024-01-11 15:26 | Minglanilla Rural Health Unit | Pide 1, Minglanilla ,Cebu | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26902 | DOH | brendagonzagaparedes@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1417 |  | GX4 | 838512 | Module 3 SN 862243 defective | Module replacement of module 3 | Module replacement of Module with new Module Serial Number 787229 | All module are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807674563919759894/5807674563919759894_signature_28.png | Brenda Paredes , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5807674563919759894/5807674563919759894_signature_31.png
| 2024-01-11 | 2024-01-11 15:22 | LAPULAPU CITY HEALTH OFFICE | PAJO, LAPULAPU CITY, CEBU | 7 | Cebu | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | tanya_pax@yahoo.com |  | 26901 | DOH | tanya_pax@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2126 |  | GX4 | 837661 | Module 4 SN 753535 defective | Module replacement of Module 4 | Module replacement of Module 4 with new module 210074454 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807671473916179550/5807671473916179550_signature_28.png | Maria Lucille Arong, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5807671473916179550/5807671473916179550_signature_31.png
| 2024-01-11 | 2024-01-11 14:55 | Punturin 3S Health Center | Punturin, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26900 | DOH | lukegonzaga09@gmail.comuxh | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1153 | 2024-01-08 | GXIV | 831914 | Module B3 for replacement | #NAME? | Replaced B3 | Unit is operational |  | GeneXpert Module SN: 756767 | 11071 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807660064535511819/5807660064535511819_signature_28.png | Judelyn Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5807660064535511819/5807660064535511819_signature_31.png
| 2024-01-11 | 2024-01-11 14:37 | Libertad Main Health Center | Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26899 | DOH | suzzettekho34@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1349 | 2023-12-14 | GX-4 | 838554 | Module B3 Found Defective | None | Replacing Module B3 with SN: 862879 to SN:791131 | All modules and units are functional |  | Certified Pre-Owned Genexpert Module | 10934 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807650273029127736/5807650273029127736_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/5807650273029127736/5807650273029127736_signature_31.png
| 2024-01-11 | 2024-01-11 13:46 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 26898 | DOH | ajbtrryes@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1247 |  | GX4 | 838477 | Defective module 862003/B2. | #NAME? | - Replaced module B2 with new module SN 749944. | 4 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807615916428165176/5807615916428165176_signature_28.png | jose ali c treyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5807615916428165176/5807615916428165176_signature_31.png
| 2024-01-11 | 2024-01-11 13:27 | Nuestra Señora De Piat District Hospital | Maguilling, Piat, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26897 | PBSP | mariagenevaligas@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 345 | 2024-01-11 | Genexpert II | 835091 | Module B2 error 2008 | Performed Xpertcheck | Performed Maintenance Gathered Reports Performed Calibration (B1 = Passed), (B2 = Service Required) Performed Functionality Test | Unit is Operational | Module B2 for Replacement | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807615514519697267/5807615514519697267_signature_28.png | Maria Geneva M. Ligas | https://www.jotform.com/uploads/macarewecare/222927978372067/5807615514519697267/5807615514519697267_signature_31.png
| 2024-01-11 | 2024-01-11 12:53 | Aurora Rural Health Unit | Aurora, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com |  |  | 26896 | PBSP | joann8850@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 371 | 2024-01-11 | GX 4 | 834435 | Module 2 hardware failed | Replaced module 776329 with module 680310 | Module Replacement and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807607497612133650/5807607497612133650_signature_28.png | Jo Ann Duran | https://www.jotform.com/uploads/macarewecare/222927978372067/5807607497612133650/5807607497612133650_signature_31.png
| 2024-01-11 | 2024-01-11 12:00 | Cadiz CHO | Cadiz City | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com | 06emerson23@gmail.com |  | 26895 | DOH | zhevebacomo@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1698 |  | GX4 | 838499 | Module B2, B3 and B4 failed in XPC during last calibration. | Conduct PM and saves files. | Replaced the old module for B2 with SN 606933 to new module SN 723150. Replaced old module for B3 with SN 71560 to new module SN745137. Replaced the old module for B4 with SN 777925 to new module SN 210008412. | All modules are now operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807593158276817790/5807593158276817790_signature_28.png | Zherlyn Bacomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5807593158276817790/5807593158276817790_signature_31.png
| 2024-01-11 | 2024-01-11 12:47 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 26894 | DOH | lornsabaigar@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 | 2024-01-11 | GXIV | 838409 | Module B1/861109 Defective | #NAME? | - Module replacement - Install new module on B1 w/ SN 862407 - Performed Functionality test - Performed preventive maintenance - Generate reports | Unit is Operational |  | Genexpert module 6 Color SN/862407 | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807579806398018340/5807579806398018340_signature_28.png | Vanessa Mabag | https://www.jotform.com/uploads/macarewecare/222927978372067/5807579806398018340/5807579806398018340_signature_31.png
| 2024-01-11 | 2024-01-11 11:16 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan | NCR | Bulacan | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26893 | DOH | laguitannataniel@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 784 | 2024-01-10 | GXIV | 831945 | Module B1, B2 for replacement | #NAME? | Replaced B1, B2 | Unit is operational |  | GeneXpert Module SN: 617067, 618478 | 11082 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5807526065143143589/5807526065143143589_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5807526065143143589/5807526065143143589_signature_31.png
| 2024-01-10 | 2024-01-10 16:20 | Bacolod City Health Office | Bacolod City, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | ON-SITE ASSISTANCE Installation of UPS and AVR | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 26892 | PBSP | cyrilcordova0515@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6673 |  | GX4 | 804267 | Defective UPS and AVR | Preventive Maintenance | - Installation of UPS (2 KVA) and AVR (5 KVA) - Functionality testing of UPS and AVR. | - New UPS and AVR functional - All 4 modules functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806845833513377875/5806845833513377875_signature_28.png | Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5806845833513377875/5806845833513377875_signature_31.png
| 2024-01-10 | 2024-01-10 16:21 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | TB DOTS | NO | REPAIR | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 26891 | DOH | cyrilcordova0515@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6673 |  | GX4 | 837612 | Module B1 w/ SN 806654 and B2 with SN 806655 failed during last XPC. | Conduct PM and save files. | Installed the new module for B1 w/ SN758240 and module B2 w/ SN 758602. | All modules are now operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806844300917480514/5806844300917480514_signature_28.png | Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5806844300917480514/5806844300917480514_signature_31.png
| 2024-01-10 | 2024-01-10 15:47 | Marik | Marikina Health City Center, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26890 | MACARE-HBDC | amymitra22@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 139 | 2024-01-10 | GX IV 10C-D | 110016107 | Unit due for XpertCheck | Performed PM Performed XpertCheck all modules passed Generated report | Performed XpertCheck | unit is operational all modules working | used5 XpertCheck cartridges | XpertCheck Kit | DR 11078 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806829827423283159/5806829827423283159_signature_28.png | May Verona | https://www.jotform.com/uploads/macarewecare/222927978372067/5806829827423283159/5806829827423283159_signature_31.png
| 2024-01-10 | 2024-01-10 15:39 | Parola Health Center | Parola Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26889 | DOH | friaslynette@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1521 | 2024-01-08 | GXIV | 831871 | Unit for PM. For installation of new UPS. | Performed PM. Performed installation of UPS. Generated reports. | PM and installation. | Unit is operational. |  | Accupower UPS (SN: AT1K230830ST0012) in replacement to defective UPS (SN: AT1000180910ST0672). | DR#11053 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806818730317832824/5806818730317832824_signature_28.png | Lynette Olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5806818730317832824/5806818730317832824_signature_31.png
| 2024-01-10 | 2024-01-10 14:21 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-Dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 26888 | DOH | marteetingzon20@gmail.con | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1315 | 2024-01-10 | GXIV | 838421 | Module B1/861073 and B4/861213 for module replacement | #NAME? | - Install new module on B1 w/ SN:717426 - Install new module on B4 w/ SN:738410 - Performed Functionality test - Performed preventive maintenance - Generate reports | unit is Operational |  | Genexpert module 6 Color (SN: 717426, 738410) | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806775264525656030/5806775264525656030_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5806775264525656030/5806775264525656030_signature_31.png
| 2024-01-10 | 2024-01-10 13:58 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | REPAIR | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26887 | DOH | vsmmsdeptpatho.microbiology@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3289 ASSAY: Xpert HIV-1 Viral Load, COUNT: 218 | 2023-12-14 | GXIV | 820206 | Module A2 defective | -Replace defective module A2/716103 with new module SN: 784455 | #NAME? | All modules are functional. |  | 900-0491RS | 10939 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806760004956840291/5806760004956840291_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5806760004956840291/5806760004956840291_signature_31.png
| 2024-01-10 | 2024-01-10 13:15 | Ramon magsaysay rural health unit | Ramon magsaysay, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com |  |  | 26886 | DOH | rhuramonmagsaysay@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 637 | 2023-11-21 | GX 4 | 838443 | Module 2 and 3 hardware failed | Replaced module 806383 and 806382 with modules 623329 and 210059989 respectively. Preventive Maintenance | Replaced module 806383 and 806382 with modules 623329 and 210059989 respectively. Preventive Maintenance | All modules are available | None | 900-0491RS | G 6460 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806753175416273512/5806753175416273512_signature_28.png | Gina B Borinaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5806753175416273512/5806753175416273512_signature_31.png
| 2024-01-10 | 2024-01-10 13:46 | Solana Municipal Health Office | Centro South East, Solana, Cagayan | 2 | Cagayan | Laboratory | NO | INSTALLATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26885 | PBSP | daisyjane2094@gmail.com | 21 |  | 2024-01-10 | Genexpert II | 835090 | UPS not powering on | Troubleshoot UPS | Installed Genexpert System Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806750925214917833/5806750925214917833_signature_28.png | Daisy Jane Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5806750925214917833/5806750925214917833_signature_31.png
| 2024-01-10 | 2024-01-10 12:37 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26884 | PBSP | venvalizado@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1000 |  | GXIV | 834431 | Unit for relocation | #NAME? | Relocated unit from Molecular Lab to Main Lab | Unit is operational Module D3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806714389825845213/5806714389825845213_signature_28.png | Venmarie Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5806714389825845213/5806714389825845213_signature_31.png
| 2024-01-10 | 2024-01-10 11:14 | Siargao Island Medical Center | Dapa | 13 | Surigao del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26883 | DOH | carlynvirtudazo10@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 761 | 2023-08-01 | GX-4 | 838444 | None | None | XpertCheck Calibration Preventive Maintenance | All modules passed in XPC Modules and Units are functional |  | XpertCheck, CE-IVD 5 Cartridges/Kit | 10391 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806685662106193045/5806685662106193045_signature_28.png | Carlyn Virtudazo | https://www.jotform.com/uploads/macarewecare/222927978372067/5806685662106193045/5806685662106193045_signature_31.png
| 2024-01-10 | 2024-01-10 11:00 | Vinzons RHU | Vinzons, Camarines Norte | 5 | Camarines Norte | Lab | NO | UPS in house repair | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | florescanoel77@gmail.com |  | 26882 | DOH | florescanoel77@gmail.com | 26 |  |  | GX4 | 838576 | UPS unable to power on,defective power board, shorted | Checked power output, checked power board continuity,checked fuse continuity | Troubleshooting | UPS unoperational | For replacement of UPS |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806652823619852717/5806652823619852717_signature_28.png | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5806652823619852717/5806652823619852717_signature_31.png
| 2024-01-10 | 2024-01-10 10:43 | Isabela Municipal Health Office | Borromeo St., Isabela, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 26881 | PBSP | laurelmarjorie1@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1305 |  | GX4 | 838855 | Defective CPU | #NAME? | #NAME? | - All 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806642845918015041/5806642845918015041_signature_28.png | Marjorie Laurel | https://www.jotform.com/uploads/macarewecare/222927978372067/5806642845918015041/5806642845918015041_signature_31.png
| 2024-01-10 | 2024-01-10 08:56 | Belmonte Health Center | Sampaloc Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26880 | PBSP | maricar.lira@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1300 |  | GXIV | 808619 | Unit for PM. For generating of GXX report. For troubleshooting of monitor. | Performed PM. Performed troubleshooting of monitor. Generated reports. Generated GXX files as per request of Cepheid. | PM and troubleshooting | Unit is operational. Monitor for replacement to due pixel problem. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806577961329737461/5806577961329737461_signature_28.png | Maricar P. Lira | https://www.jotform.com/uploads/macarewecare/222927978372067/5806577961329737461/5806577961329737461_signature_31.png
| 2024-01-09 | 2024-01-09 15:33 | National Kidney and Transplant Institute | East Ave. QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26879 | MACARE-HBDC | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1175 ASSAY: Others, COUNT: 40 ASSAY: Xpert HBV VL, COUNT: 799 ASSAY: Xpert HPV, COUNT: 143 ASSAY: Xpert MTB/RIF, COUNT: 627 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1838 ASSAY: Xpert SARS-CoV-2, COUNT: 14961 ASSAY: Xpert HCV VL, COUNT: 82 | 2024-01-08 | Gx16 | 110005587 | For module replacement b4 686979 - 655306 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | 13 modules operational a4,c2,c4 disable encounter temp error |  | 900 0491rs | 11050 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5806093441893036738/5806093441893036738_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5806093441893036738/5806093441893036738_signature_31.png
| 2024-01-09 | 2024-01-09 14:48 | Caraga TB Reference Laboratory | Rizal Street, Brgy. Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26878 | PBSP | abpabellan.pbsp@gmail.com | 23.95 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5651 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GX-4 | 807602 | None | None | XpertCheck Calibration and Preventive Maintenance of SN: 807602 | All modules and units are functional | None |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805992681911146364/5805992681911146364_signature_28.png | Andrew B. Pabellan | https://www.jotform.com/uploads/macarewecare/222927978372067/5805992681911146364/5805992681911146364_signature_31.png
| 2024-01-09 | 2024-01-09 15:03 | San Mateo Integrated Community Hospital | Brgy 4, San Mateo, Isabela | 2 | Isabela | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26877 | PBSP | smichlab@yahoo.com | 26 |  | 2024-01-09 | Genexpert II | 835092 | N/a | N/a | Installed Genexpert System Performed Maintenance Gathered Reports Performed Functionality | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805934730613681618/5805934730613681618_signature_28.png | Regina Ann Ramento | https://www.jotform.com/uploads/macarewecare/222927978372067/5805934730613681618/5805934730613681618_signature_31.png
| 2024-01-09 | 2024-01-09 14:34 | Hinigaran Rural Health Unit | Hinigaran, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 26876 | PBSP | mfeliprada@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1013 |  | GX4 | 815283 | New Printer for installation | #NAME? | #NAME? | New printer operational, 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805916412828831543/5805916412828831543_signature_28.png | May Flo M. Feliprada | https://www.jotform.com/uploads/macarewecare/222927978372067/5805916412828831543/5805916412828831543_signature_31.png
| 2024-01-09 | 2024-01-09 13:56 | Bulakan Rural Health Unit | Bulakan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26875 | DOH | l_gemino@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1038 ASSAY: Xpert HIV-1 Viral Load, COUNT: 97 | 2024-01-08 | GXIV | 831631 | Module A4 for replacement | #NAME? | Replaced A4 | Unit is operational |  | GeneXpert Module SN: 717981 | 11074 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805891761214197030/5805891761214197030_signature_28.png | Beth Gemin9 | https://www.jotform.com/uploads/macarewecare/222927978372067/5805891761214197030/5805891761214197030_signature_31.png
| 2024-01-09 | 2024-01-09 12:23 | Tagum City Health Office | Mabini Street, Tagum City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26874 | DOH | tagumcholab.mabini@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2469 |  | GX-IV 4D | 838553 | Module Error at B1 and B3 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPC Calibration for error verification and Run Dummy Test for Functionality Test | Perform Preventive Maintenance and Functionality Test for error Modules | Module B3 reported to Cepheid for further Assessment, Only 3 Modules are functional | Module B1 for under observation, High Room Temp, end user is advised to have additional cooling system to achieve the GeneXpert prescribed room operating temperature |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805883301422740808/5805883301422740808_signature_28.png | Cristina Tiara A. Fernandez,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5805883301422740808/5805883301422740808_signature_31.png
| 2024-01-09 | 2024-01-09 11:43 | Lung Center of the Philippines | Quezon ave , QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26873 | MACARE-HBDC | lcpbloodbank@gmail.com | 28 |  |  | MCS + | 02D166 | Due for maintenance and calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Maintenance and calibration | Operational | Next due of pms 4/8/24 due of calibration 1/8/25 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805869814813489448/5805869814813489448_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/5805869814813489448/5805869814813489448_signature_31.png
| 2024-01-09 | 2024-01-09 12:31 | Lung Center of the Philippines | Quezon ave., QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26872 | MACARE-HBDC | lcpbloodbank@gmail.com | 28 |  |  | MCS + | 08H188 | Due for maintenance and calibration | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed simulation running | Maintenance and calibration | Operational | Need to replace rollers, Next due of pms 4/8/24 and due of calibratin 1/8/25 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805869304814609225/5805869304814609225_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/5805869304814609225/5805869304814609225_signature_31.png
| 2024-01-09 | 2024-01-09 12:49 | Virac Rural Health Unit | Virac | 5 | Catanduanes | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | danemilbodino23@gmail.com |  | 26871 | DOH | rogerericfacundo@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1717 |  | GX-4 | 838563 | B1/616796 for module replacement | Replaced Module: B1 - 616796>>730980 Cleaned cartridge bays, plunger rod, icores Generate reports | Replaced module B1/616796 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805861658113350820/5805861658113350820_signature_28.png | Roger Eric Facundo | https://www.jotform.com/uploads/macarewecare/222927978372067/5805861658113350820/5805861658113350820_signature_31.png
| 2024-01-09 | 2024-01-09 11:10 | Zamboanga del sur medical center | Zamboanga del sur Medicalcenter, Pagadian City | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com |  |  | 26870 | DOH | limboemillie@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1686 ASSAY: Xpert HIV-1 Viral Load, COUNT: 91 | 2023-11-21 | GX 4 | 838204 | Module 1, 2 and 3 hardware failed | Replaced module 807030 with module 210012372 and conducted xpertcheck calibration for module 2 and 4. | Replaced module 807030 with module 210012372 and conducted xpertcheck calibration for module 2 and 4. | Module 2 and 4 failed the xpertcheck calibration | Module 2 and 4 are for replacement | 900-0491RS | G 6461 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805809794657094331/5805809794657094331_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/5805809794657094331/5805809794657094331_signature_31.png
| 2024-01-09 | 2024-01-04 11:55 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26869 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1246 | 2023-12-14 | GXIV | 831691 | B3 for replacement | #NAME? | Replaced B3 | Unit is operational. |  | GeneXpert Module SN: 890579 | 10124 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805796961211437947/5805796961211437947_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5805796961211437947/5805796961211437947_signature_31.png
| 2024-01-09 | 2024-01-09 10:41 | Meycauayan City Health Center VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26868 | DOH | ronamarieiris_25@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1574 ASSAY: Xpert HIV-1 Viral Load, COUNT: 109 |  | GXIV | 831904 | Module A2 for re XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. A2 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805796051213847492/5805796051213847492_signature_28.png | Earl Zuniega | https://www.jotform.com/uploads/macarewecare/222927978372067/5805796051213847492/5805796051213847492_signature_31.png
| 2024-01-09 | 2024-01-09 10:51 | Himamaylan City Health Office | Himamaylan City, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 26867 | PBSP | fayemariejuguan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1206 |  | GX4 | 838138 | Defective module 806683/B1 due to temperature-related errors. | #NAME? | Replaced module B1 with new module SN 700632. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805783421262196883/5805783421262196883_signature_28.png | Faye Marie Juguan | https://www.jotform.com/uploads/macarewecare/222927978372067/5805783421262196883/5805783421262196883_signature_31.png
| 2024-01-08 | 2024-01-08 16:57 | Masinloc Rural Health Unit | Masinloc Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26866 | DOH | rodenne@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2191 | 2024-01-05 | GXIV-D | 831608 | Monitor for replacement A2:754352 A4: 784650 for module replacement | Repair Preventive maintenance | Performed preventive maintenance Performed monitor and module replacement Gathered reports Performed functionality test | Unit is operational | A2:732005 and A4:670638 | 1 monitor 2 modules | 11038/11037/10959 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805142433618518252/5805142433618518252_signature_28.png | Rodenn Eduvala | https://www.jotform.com/uploads/macarewecare/222927978372067/5805142433618518252/5805142433618518252_signature_31.png
| 2024-01-08 | 2024-01-08 16:16 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26865 | DOH | chaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1804 |  | GX4 | 838283 | Due for Preventive Maintenance | Performed preventive maintenance , check ups and voltage regulator voltage | Performed preventive maintenance in Icore, clean filter, archive data gxx, generate Sys log, IQR , Assay Statistics | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805113373918599489/5805113373918599489_signature_28.png | Lucille Mae A. Potoy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5805113373918599489/5805113373918599489_signature_31.png
| 2024-01-08 | 2024-01-08 13:03 | Olongapo Ciyt Health Office | West Bajac Bajac Olongapo City Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26864 | PBSP | olongapochotmlandrtdl@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1491 ASSAY: Xpert HIV-1 Viral Load, COUNT: 134 | 2023-04-27 | GX2-D | 834806 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modukes passed calibration | 4 xpertcheck kits | 40656 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805003843611712193/5805003843611712193_signature_28.png | Roselyn Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5805003843611712193/5805003843611712193_signature_31.png
| 2024-01-08 | 2024-01-08 13:07 | Olongapo City Health Office | West Bajac Bajac Olongapo City Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26863 | PBSP | olongapochotmlandrtdl@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1728 | 2023-03-10 | GXIV-D | 804700 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except D1 D2 unable to calibrate D3 and D4 Passed calibration | D1 failed calibration | 3 xpertcheck kits | 39470 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5805001803619512618/5805001803619512618_signature_28.png | Roselyn Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5805001803619512618/5805001803619512618_signature_31.png
| 2024-01-08 | 2024-01-08 13:12 | Olongapo City Health Office | West Bajac Bajac Olongapo City Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26862 | PBSP | olongapochotmlandrtdl@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1491 ASSAY: Xpert HIV-1 Viral Load, COUNT: 134 | 2023-04-27 | GXII-D | 834807 | Due for calibration | Xoertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration | 2 xpertcheck kits | 40656 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804999593619888765/5804999593619888765_signature_28.png | Roselyn Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5804999593619888765/5804999593619888765_signature_31.png
| 2024-01-08 | 2024-01-08 10:26 | Orani Rural Health Unit | Mulawin Orani Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26861 | DOH | oranirhu@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 614 | 2024-01-05 | GXIV-D | 831668 | A1:757000 for module replacement A2 and A3 for Re-calibration | Troubleshooting Repair Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Performed module replacement Gathered reports Performed functionality test | Unit is operational except A2 | A1:758120 | Genexpert module 3 xpertcheck kits | 11042 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804904803617968928/5804904803617968928_signature_28.png | Belinda C. Guinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5804904803617968928/5804904803617968928_signature_31.png
| 2024-01-08 | 2024-01-05 10:39 | Nagpayong Super Health Center | Nagpayong | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26860 | DOH | mnoreend@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1103 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1103 |  | GX-IV-D-6C | 831905 | Module B1 has error 1001/2014 | Run XpertCheck on module B1 | Performed maintenance Run XpertCheck on module B1 Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804882613615075577/5804882613615075577_signature_28.png | Ma.Noreen De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5804882613615075577/5804882613615075577_signature_31.png
| 2024-01-08 | 2024-01-05 14:23 | Angono Rural Health Unit | Angono | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26859 | DOH | siennablanco@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1035 | 2023-12-20 | GX-IV-D-6C | 819864 | Module A4 for replacement | Replaced module A4 | Performed maintenance Replaced module A4 Generated reports | Unit is operational. Module A2 not operational, reported already. |  | Old: A4/719514 New: A4/863009 | DR# 10998 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804882613617648999/5804882613617648999_signature_28.png | Rose Sienna Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5804882613617648999/5804882613617648999_signature_31.png
| 2024-01-08 | 2024-01-08 07:43 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26858 | DOH | abigailgsula@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1367 | 2024-01-05 | GXIV-D | 831912 | Monitor replacement: defective monitor SN:CNK83224PD Module replacement: A2 SN:792321 | Repair Preventive maintenance | Performed preventive maintenance Performed module and monitor replacement Gathered reports Performed functionality test | Unit is operational | A2:749409 and new monitor:AVI215CMQY05B01060577 | Genexpert module and monitor | 11039 and11049 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804809013611300393/5804809013611300393_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5804809013611300393/5804809013611300393_signature_31.png
| 2024-01-08 | 2024-01-05 13:27 | Iriga City Health Office | Irigq, Camarines Sur | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26857 | DOH | mongealyzza@yahoo.com.ph | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1718 |  | GX-4 | 838391 | B1/780882 for Module Replacement | Replaced Module: B1 - 780882>>758212 Cleaned fan, cartridge bays, plunger rod, icore Generate reports | Replaced module B1 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804529692121466345/5804529692121466345_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5804529692121466345/5804529692121466345_signature_31.png
| 2024-01-08 | 2024-01-04 13:24 | Pilar RHU | Pilar | 5 | Sorsogon | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26856 | DOH | manalangmark88@gmail.com | 25 |  |  | GX-4 | 838266 | CPU for replacement | Cleaned cartridge bays, plunger rod Replaced CPU Generate reports | Replaced CPU | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804528262129579970/5804528262129579970_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5804528262129579970/5804528262129579970_signature_31.png
| 2024-01-07 | 2024-01-03 14:16 | Juban RHU | Juban, Sorsogon | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26855 | PBSP | gaytarean@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 800 |  | GX-4 | 835214 | B2/210091002 for module replacement | Replaced Module: B2 - 210091002>>701213 | Replaced module B2/210091002 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804521092125746906/5804521092125746906_signature_28.png | Rean G. Gracilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5804521092125746906/5804521092125746906_signature_31.png
| 2024-01-07 | 2023-12-29 15:06 | Calabanga Rural Health Unit | Calabanga | 5 | Camarines Sur | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26854 | DOH | martyarc11@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 710 |  | GX4 | 838398 | B2/860997, B3/860998, B4/860999 for module replacement B1/807565 temperature error | Replaced module: B2-860997>>737294, B3-860998>>731543, B4-860999>>807261 Performed Re-XpertCheck for B1 Generate reports | Replaced module B2, B3, B4 Performed re-XpertCheck for B1 | Unit is operational | Re-XpertCheck: B1/807565 - service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804519192122266460/5804519192122266460_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/5804519192122266460/5804519192122266460_signature_31.png
| 2024-01-07 | 2023-12-27 20:50 | Bicol Medical Center | Naga | 5 | Camarines Sur | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26853 | PBSP | bmc.bacteriology@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1417 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7422 ASSAY: Xpert SARS-CoV-2, COUNT: 2319 ASSAY: Xpert HIV-1 Viral Load, COUNT: 273 ASSAY: Others, COUNT: 10 |  | GX-16 | 842225 | B2/612544 for module replacement | Replaced Module: B2 - 612544>>669505 Cleaned cartridge bays, plunger rod, fan Generate reports | Replaced module B2/612544 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804515472123627539/5804515472123627539_signature_28.png | Angelika Mae Bichara | https://www.jotform.com/uploads/macarewecare/222927978372067/5804515472123627539/5804515472123627539_signature_31.png
| 2024-01-07 | 2023-12-27 20:50 | Bicol Medical Center | Naga | 5 | Camarines Sur | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26852 | PBSP | bmc.bacteriology@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1417 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7422 ASSAY: Xpert SARS-CoV-2, COUNT: 2319 ASSAY: Xpert HIV-1 Viral Load, COUNT: 273 ASSAY: Others, COUNT: 10 |  | GX-16 | 842225 | B2/612544 for module replacement | Replaced Module: B2 - 612544>>669505 Cleaned cartridge bays, plunger rod, fan Generate reports | Replaced module B2/612544 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804513052121071283/5804513052121071283_signature_28.png | Angelika Mae Bichara | https://www.jotform.com/uploads/macarewecare/222927978372067/5804513052121071283/5804513052121071283_signature_31.png
| 2024-01-07 | 2023-12-27 14:55 | San Fernando Rural Health Unit | San Fernando, Camarines 3 | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26851 | DOH | josephus_garfin@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1102 |  | GX4 | 838496 | B1/861697, B2/861662, B3/862246 for module replacement | Module replaced: B1-861697>>722557, B2-861662>>723641, B3-862246>>725516 Cleaned cartridge bay, plunger rod, icore, fan Generate reports | Replaced module: B1/861697, B2/861662, B3/862246 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804511462122490273/5804511462122490273_signature_28.png | Josephus Garfin | https://www.jotform.com/uploads/macarewecare/222927978372067/5804511462122490273/5804511462122490273_signature_31.png
| 2024-01-07 | 2023-12-27 10:40 | Bicol Medical Center- Blood Bank | Naga | 5 | Camarines Sur | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com | charitypreston88@gmail.com |  | 26850 | MACARE-HBDC | ma.charitydelosreyes@yahoo.com | 25 |  |  | MCS+ | 15H027G | For Preventive Maintenance | Cleaned MCS+ Surface Cleaned centrifudge, rollers, checked diagnostic parameters | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804509282127138003/5804509282127138003_signature_28.png | Ma. CHARITY P. DE LOS REYES | https://www.jotform.com/uploads/macarewecare/222927978372067/5804509282127138003/5804509282127138003_signature_31.png
| 2024-01-07 | 2023-12-27 10:37 | Bicol Medical Center - Blood Bank | Naga | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  | danemilbodino23@gmail.com | charitypreston88@gmail.com |  | 26849 | MACARE-HBDC | ma.charitydelosreyes@yahoo.com | 25 |  |  | MCS+ | 11J221 | SPM out of tolerance | Followed Diagnostic Parameters Check/Cleaned SPM Cleaned MCS+ Surface Checked calibration data | Performed troubleshooting Performed Preventive Maintenance | Unit is not operational | For further troubleshooting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804504352127187774/5804504352127187774_signature_28.png | MA.CHARITY P. DE LOS REYES | https://www.jotform.com/uploads/macarewecare/222927978372067/5804504352127187774/5804504352127187774_signature_31.png
| 2024-01-07 | 2024-01-07 22:36 | Mini Forest Health Center | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26848 | DOH | miniforestlaboratory@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 519 |  | GX-IV 4d | 838507 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804475541422538903/5804475541422538903_signature_28.png | Joy L. Nio,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5804475541422538903/5804475541422538903_signature_31.png
| 2024-01-07 | 2024-01-07 15:59 | Quezon Medical Center Wellness Hub | Quezon Medical Center Compound, Brgy. 11, Lucena | 4A | Quezon | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26847 | SHELL FOUNDATION | villanuevazenjee18@gmail.com | 26 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 79 |  | GX IV 10C - D | 120000146 | Error 5016 and Jotform Assistance | Performed PM Archived files for technical troubleshooting assisted in jotform reporting Gathered report | Checked system log and reports Walkthrough jotform assistance | Unit is operational. | Serviced on January 4, 2023. Jotform error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804248071818432361/5804248071818432361_signature_28.png | Zenjee V. Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5804248071818432361/5804248071818432361_signature_31.png
| 2024-01-07 | 2024-01-07 09:16 | Quezon Medical Center | Provincial Health Office, Barangay XI, Lucena, Quezon City | 4A | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26846 | PBSP | qmcrtdl2015@gmail.com | 28 |  |  | GX IV 6C - D | 807462 | UPS not working Module D2 frequent error | Performed PM UPS not backing up, for reporting Module D2 for reporting Generated report | Troubleshoot UPS Checked module D2 | unit is operational UPS for replacement Module D2 for replacement | Serviced on Dec. 21, 2023. Jotform Error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5804000911814913672/5804000911814913672_signature_28.png | JoAnn P. Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5804000911814913672/5804000911814913672_signature_31.png
| 2024-01-07 | 2024-01-03 13:52 | Domingo Santiago Health Center | 844 Domingo Santiago, Sampaloc, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26845 | DOH | fei_paas@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1115 |  | GX IV 6C - D | 831537 | Module 2 stuck cartridge | Perfomed PM Removed stuck cartridge Generated report | Removed stuck cartridge | unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5803973941817118688/5803973941817118688_signature_28.png | Flory Fe Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5803973941817118688/5803973941817118688_signature_31.png
| 2024-01-07 | 2024-01-04 15:12 | Quezon Medical Center | Provincial Helath Office, Barangay XI, Lucena City, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26844 | PBSP | qmcrtdl2015@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 340 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2521 ASSAY: Xpert HIV-1 Viral Load, COUNT: 83 | 2023-12-22 | GX IV 6C -D | 807462 | Module D1 for replacement | Performed PM Replaced Module D1 Generated report | Replaced Module D1 210003676 -> 717400 | unit is operational 3 modules working. 1 module waiting for replacement |  | GX Module 6C | DR 45839 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5803967311814715011/5803967311814715011_signature_28.png | JoAnn P. Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5803967311814715011/5803967311814715011_signature_31.png
| 2024-01-07 | 2024-01-05 10:42 | Guinayangan Municipal Health Office | Brgy. Calimpak, Guinayangan, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | lala0117r.kampitan@yahoo.com |  | 26843 | PBSP | reamoongab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 598 | 2023-12-06 | GX IV 6C - D | 847334 | Module B3 for replacement | Performed PM Replaced Module B3 Generated report | Replaced module B3 210101928 -> 685248 | Unit is operational. all modules working |  | GX Module 6C | DR 45632 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5803964291812878328/5803964291812878328_signature_28.png | Rea Joan Formanes | https://www.jotform.com/uploads/macarewecare/222927978372067/5803964291812878328/5803964291812878328_signature_31.png
| 2024-01-06 | 2024-01-04 14:45 | Davao Reproductive Health and Wellness Center | Jacinto St., Poblacion Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26842 | DOH | rhwclab@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 211 ASSAY: Xpert HIV-1 Viral Load, COUNT: 841 |  | GX-IV 4D | 838262 | Module B2, having errors For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Run dummy cartridge to check functionality | Perform Preventive Maintenance and Module Troubleshooting | All Modules are operational | Module B2 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5803412508816164669/5803412508816164669_signature_28.png | Ellaine B. Mallano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5803412508816164669/5803412508816164669_signature_31.png
| 2024-01-05 | 2024-01-05 17:01 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26841 | DOH | nacosta@loveyourself.ph | 25 |  |  | GXIV | 831895 | Stuck cartridge on A1 A1, A3 for reXpertCheck | #NAME? | Removed stuck cartridge Performed XpertCheck on A1, A3 | Unit is operational A1, A3 requires service |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802563436912784615/5802563436912784615_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5802563436912784615/5802563436912784615_signature_31.png
| 2024-01-05 | 2024-01-05 17:06 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26840 | PBSP | nacosta@loveyourself.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 415 ASSAY: Xpert MTB/RIF Ultra, COUNT: 365 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2529 |  | GXIV | 808649 | Unit for maintenance Bios battery problem Laptop SN: 5CG5511V52 | #NAME? | Checked unit | Unit is operational Laptop for repair / bios battery for replacement Modules for further analysis |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802561476913865061/5802561476913865061_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5802561476913865061/5802561476913865061_signature_31.png
| 2024-01-05 | 2024-01-05 15:21 | Mexico RHU I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26839 | DOH | jasperdizon20@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 520 | 2023-12-12 | Genexpert IV | 831877 | Monitor Problem | Replaced Monitor | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Monitor - AV125CMQY05B01060049 | 10891 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802488197614931280/5802488197614931280_signature_28.png | Jasper P. Dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5802488197614931280/5802488197614931280_signature_31.png
| 2024-01-05 | 2024-01-04 14:35 | Bulakan Rural Health Unit | Bulakan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26838 | DOH | l_gemino@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1026 ASSAY: Xpert HIV-1 Viral Load, COUNT: 97 | 2023-12-15 | GXIV | 831641 | Modules A2, A3 for replacement | #NAME? | Replaced A2, A3 | Unit is operational. |  | GeneXpert Module SN: 717280, 715736 | 10956 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802463746914020348/5802463746914020348_signature_28.png | Lilibeth Gemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5802463746914020348/5802463746914020348_signature_31.png
| 2024-01-05 | 2024-01-04 14:35 | Bulakan Rural Health Unit | Bulakan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26837 | DOH | l_gemino@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1026 ASSAY: Xpert HIV-1 Viral Load, COUNT: 97 | 2023-12-15 | GXIV | 831641 | Modules A2, A3 for replacement | #NAME? | Replaced A2, A3 | Unit is operational. |  | GeneXpert Module SN: 717280, 715736 | 10956 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802463746917209458/5802463746917209458_signature_28.png | Lilibeth Gemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5802463746917209458/5802463746917209458_signature_31.png
| 2024-01-05 | 2024-01-05 13:58 | Bagumbong Health Center | Bagumbong, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26836 | DOH | jenilyn_duenas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2657 | 2023-12-20 | GXIV | 831907 | Module A4 for replacement UPS for check up | - replaced A4 - checked and cleaned cartridge bay, plunger rod, filter - archived tests from January to September 2023 - checked UPS | Replaced A4 Checked UPS | Unit is operational |  | GeneXpert Module SN: 722350 | 10997 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802463266918133779/5802463266918133779_signature_28.png | Jenilyn Duenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5802463266918133779/5802463266918133779_signature_31.png
| 2024-01-05 | 2024-01-05 10:37 | Cardinal Santos Medical Center | Wilson st., Greenhills | NCR | NCR 2nd - San Juan | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | biomed@csmc.ph |  | 26835 | COMMERCIAL | ftejada@csmc.ph | 26 |  |  | MCS + | 08M461 | As per end user they encounter red valve error during process | Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Checked and performed simulation running of valves Performed functionality test Performed isolation running | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802341412645568702/5802341412645568702_signature_28.png | Fidel Tejada | https://www.jotform.com/uploads/macarewecare/222927978372067/5802341412645568702/5802341412645568702_signature_31.png
| 2024-01-05 | 2024-01-03 14:20 | Pasig Treatment Hub | Champ Bldg. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26832 | MACARE-HBDC | pathmedtech@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 18 ASSAY: Xpert HIV-1 Viral Load, COUNT: 685 |  | GX-IV-L-6C | 821648 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. Modules A1, A2 and A4 not operational, reported. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802335161916863452/5802335161916863452_signature_28.png | Jhozza Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5802335161916863452/5802335161916863452_signature_31.png
| 2024-01-05 | 2024-01-04 10:56 | Pililla Rural Health Unit | M.A Roxas St., Pililla | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26834 | DOH | tbdotspililla@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 660 | 2023-12-20 | GX-IV-D-6C | 821293 | Module B2 for replacement | Replaced module B2 | Performed maintenance Replaced module B2 Generated reports | All modules operational. |  | Old: B2/687232 New: B2/784488 | DR# 10995 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802335161917023825/5802335161917023825_signature_28.png | Pililla RHU | https://www.jotform.com/uploads/macarewecare/222927978372067/5802335161917023825/5802335161917023825_signature_31.png
| 2024-01-05 | 2024-01-04 13:42 | Teresa Rural Health Unit | Teresa | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26833 | DOH | rosebenildatamayo@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 894 |  | GX-IV-D-6C | 820710 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802335161918117510/5802335161918117510_signature_28.png | Rosbenilda A.Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5802335161918117510/5802335161918117510_signature_31.png
| 2024-01-05 | 2024-01-04 14:48 | Davao Reproductive Health and Wellness Center | Jacinto St. Poblacion Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26831 | MACARE-HBDC | rhwclab@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 102 ASSAY: Xpert HIV-1 Viral Load, COUNT: 570 |  | GX-IV 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Only 2 Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802320720712059506/5802320720712059506_signature_28.png | Ellaine B. Mallano,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5802320720712059506/5802320720712059506_signature_31.png
| 2024-01-05 | 2024-01-04 14:45 | Davao Reproductive Health and Wellness Center | Jacinto St., Poblacion Davao City | 11 | Davao del Sur | RHIVDA | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26830 | DOH | rhwclab@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 211 ASSAY: Xpert HIV-1 Viral Load, COUNT: 841 |  | GX-IV 4D | 838262 | Module B2, having errors For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Run dummy cartridge to check functionality | Perform Preventive Maintenance and Module Troubleshooting | All Modules are operational | Module B2 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802319680715050022/5802319680715050022_signature_28.png | Ellaine Mallano RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5802319680715050022/5802319680715050022_signature_31.png
| 2024-01-05 | 2024-01-05 10:03 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION | CEPHEID | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26829 | COMMERCIAL | biopathqc@biopath.ph | 23 | ASSAY: Xpert CT/NG, COUNT: 49 ASSAY: Xpert SARS-CoV-2, COUNT: 86 ASSAY: Others, COUNT: 16 | 2024-01-04 | GXXVID | 839306 | A3 Hardware Fail | None | Performed XpertCheck Perform Maintenace, cleaning I-cores, cartridge bays, optics, and plunger rods. Generated reports Disabled service required module Performed functionality test | Unit is operational with 14 modules, A3 h/w fail, and C2 service required. |  | XpertCheck Kits | SI 96919 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5802315627435853385/5802315627435853385_signature_28.png | Amanda Inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/5802315627435853385/5802315627435853385_signature_31.png
| 2024-01-04 | 2024-01-04 16:22 | Zamboanga City Health Office | Pettit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 26828 | PBSP | floresteresitaj25@gmail.com | 25 |  | 2024-01-04 | GX 4 | 807599 | Module 3 error 2014. | Undergone Xpertcheck calibration and Preventive Maintenence | Undergone Xpertcheck calibration and Preventive Maintenence | Modules 1, 2 and 4 passed while module 3 service required | Module 3 needs to be replaced | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801661684517070624/5801661684517070624_signature_28.png | Teresita Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5801661684517070624/5801661684517070624_signature_31.png
| 2024-01-04 | 2024-01-04 15:06 | Zamboanga City Health Office | Pettit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 26827 | PBSP | floresteresitaj25@gmail.com | 23 |  | 2024-01-04 | GX 4 | 818902 | Uninterrupted power supply (UPS) is unable to work properly. Error F06 | Perform Troubleshooting on site | Troubleshoot on site with enduser | All modules are available | (Uninterrupted power supply) UPS Must be replaced emedietly | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801614552519163204/5801614552519163204_signature_28.png | Teresita Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5801614552519163204/5801614552519163204_signature_31.png
| 2024-01-04 | 2024-01-04 12:00 | Makati Medical Center | Amorsolo st., Makati | NCR | NCR 4th - Makati | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | troxoctubre@gmail.com |  | 26826 | COMMERCIAL | rovijaelle.lare@makatimed.net.ph | 28 |  |  | MCS + | 97H330 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed isolation running | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801535718812235010/5801535718812235010_signature_28.png | . | https://www.jotform.com/uploads/macarewecare/222927978372067/5801535718812235010/5801535718812235010_signature_31.png
| 2024-01-04 | 2024-01-04 12:48 | INTEGRATED PROVINCIAL HEALTH OFFICE - SIQUIJOR | DUMANHOG, SIQUIJOR | 7 | Siquijor | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26825 | PBSP | glinniericamor73@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 95 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1963 |  | GX-IV | 808638 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801531062619202723/5801531062619202723_signature_28.png | GLINNIE RICAMOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5801531062619202723/5801531062619202723_signature_31.png
| 2024-01-04 | 2024-01-04 12:35 | INTEGRATED PROVINCIAL HEALTH OFFICE - SIQUIJOR | DUMANHOG, SIQUIJOR | 7 | Siquijor | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26824 | PBSP | glinniericamor73@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 89 ASSAY: Xpert MTB/RIF Ultra, COUNT: 961 |  | GX-II | 835056 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801528692611458793/5801528692611458793_signature_28.png | GLINNIE RICAMOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5801528692611458793/5801528692611458793_signature_31.png
| 2024-01-04 | 2024-01-04 11:55 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26823 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1246 | 2023-12-14 | GXIV | 831691 | Module B3 (793226) for replacement | #NAME? | Replaced B3 | Unit is operational. |  | GeneXpert Module SN: 890579 | 10912 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801514008066036432/5801514008066036432_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5801514008066036432/5801514008066036432_signature_31.png
| 2024-01-04 | 2024-01-04 11:55 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26822 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1246 | 2023-12-14 | GXIV | 831691 | Module B3 (793226) for replacement | #NAME? | Replaced B3 | Unit is operational. |  | GeneXpert Module SN: 890579 | 10912 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801512948063906203/5801512948063906203_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5801512948063906203/5801512948063906203_signature_31.png
| 2024-01-04 | 2024-01-04 11:38 | Soccsksargen General Hospital | Surallah, South Cotabato | 12 | South Cotabato | Tb-Lab | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26821 | PBSP | sghlab21@gmail.com | 18.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1887 |  | Gx-IV | 815432 | N/A | Genexpert machine Cleaning Data collection: IQR, System log, Assay statistics report and Archive test Report and request for fan filter | Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801497954104344819/5801497954104344819_signature_28.png | Soccsksargen General Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/5801497954104344819/5801497954104344819_signature_31.png
| 2024-01-04 | 2024-01-04 11:43 | Soccsksargen General Hospital | Surallah, South Cotabato | 12 | South Cotabato | Tb-Lab | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26820 | PBSP | sghlab21@gmail.com | 20.1 | ASSAY: Others, COUNT: |  | Gx-IV (10color) | 120000172 | N/A | Perform Genexpert machine Cleaning Data collection: IQR, System log, Assay statistics report and Archive test | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801496614108372415/5801496614108372415_signature_28.png | Soccsksargen General Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/5801496614108372415/5801496614108372415_signature_31.png
| 2024-01-04 | 2024-01-04 10:32 | Northern Mindanao Medical Center | Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | HIV Lab | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26819 | PBSP | ronnel_mabasa@yahoo.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 69 | 2024-01-04 | GXIV | 804682 | Module 1 H/W failed | -Replace defective module B1/210068712 with new module SN: 715713 | #NAME? | All modules are functional. |  | GeneXpert Module 6C | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801463179522165752/5801463179522165752_signature_28.png | Ronnel Mabasa | https://www.jotform.com/uploads/macarewecare/222927978372067/5801463179522165752/5801463179522165752_signature_31.png
| 2024-01-04 | 2024-01-04 08:50 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26818 | DOH | riazarizzamae@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1023 | 2024-01-04 | Genexpert IV | 831637 | Module 1 error | Performed Xpertcheck on Module 1 (Passed) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 1 for Observation | Xpertcheck 2pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5801388597612722495/5801388597612722495_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5801388597612722495/5801388597612722495_signature_31.png
| 2024-01-03 | 2024-01-03 14:02 | Batasan Social Hygiene Clinic | IBP Road, Batasan, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26817 | MACARE-HBDC | russelldelavega030210@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GXIVD 10C | 110020697 | None | None | Installed new GXIV 10C unit Installed printer Performed basic end-user training on software and maintenance Installed HIV VL ADF Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800800836357591666/5800800836357591666_signature_28.png | RUSSELL C. DELA VEGA, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5800800836357591666/5800800836357591666_signature_31.png
| 2024-01-03 | 2024-01-03 16:26 | Davao Chest Center | Villa Abrille St, Poblacion Davao City | 11 | Davao del Sur | TB-DOTS | YES | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26816 | PBSP | dcclabdept@gmail.com | 22.5 |  |  | GX-IV 4d | 804835 | Modules not detected when launching the GeneXpert Software | Reconfigure Ip Address intended for GeneXpert Tried reseating cable connection Tried to use usb to rj45 cable adaptor Tried to reinstall Gx4.8 software and upgrade to version 5.1. Changing cables | Perfrom Onsite troubleshooting | Still Modules are not detected | Installed temporarily a GeneXpert Software on the personal laptop of the enduser. Documented videos gathered will be sent to Cepheid foe Further Evaluation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800799341428465985/5800799341428465985_signature_28.png | Maria Lisa Paras, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5800799341428465985/5800799341428465985_signature_31.png
| 2024-01-03 | 2024-01-03 12:01 | Koronadal City Health Office | Koronadal City, South Cotabato | 12 | South Cotabato | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26815 | PBSP | calvokenneth12@gmail.com | 24.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3052 | 2024-01-03 | Gx-IV | 707529 | Module 1 H/W failed | Repair/replace defective module 1 SN:866833 with new module SN: 730998 Genexpert machine Cleaning Data collection: Pre and Post IQR, System log, Assay statistics report, and Archive test | Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6- color SN: 730998 | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800790754993806586/5800790754993806586_signature_28.png | Kenneth Paolo T. Calvo | https://www.jotform.com/uploads/macarewecare/222927978372067/5800790754993806586/5800790754993806586_signature_31.png
| 2024-01-03 | 2024-01-03 15:47 | Sto Domingo District Hospital | Sto Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | roxasmatthew659@yahoo.com.ph |  | 26814 | DOH | lhaydeesel@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 984 | 2023-12-12 | Genexpert IV | 831582 | N/a | N/a | Performed Maintenance Gatherex Reports Performed Module Replacement (754239 - 888301) Performed Functionality Test | Unit is Operational | Module 2 for observation | Genexpert Module | 10906 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800774849219274367/5800774849219274367_signature_28.png | Haydee Alejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5800774849219274367/5800774849219274367_signature_31.png
| 2024-01-03 | 2024-01-03 11:14 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | YES | MAINTENANCE For software upgrade | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26813 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 24 | ASSAY: Xpert HPV, COUNT: 25 ASSAY: Xpert MTB/RIF Ultra, COUNT: 103 ASSAY: Xpert SARS-CoV-2, COUNT: 13335 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GX-XVI-6C-D | 838627 | For software upgrade | Performed maintenance, archive and purge tests | Performed maintenance Tried to upgrade GeneXpert software Generated reports Archive and purge tests | Unit is operational. Modules A1,A2,B3,B4,C2,D2,D4 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800720282956926794/5800720282956926794_signature_28.png | Justine Lorenz V. Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5800720282956926794/5800720282956926794_signature_31.png
| 2024-01-03 | 2024-01-03 09:59 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26812 | PBSP | molecularlab@rmc.doh.gov.ph | 24 |  |  | GX-IV-D-6C | 814374 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800720282954399298/5800720282954399298_signature_28.png | Justine Lorenz V. Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5800720282954399298/5800720282954399298_signature_31.png
| 2024-01-03 | 2024-01-03 12:36 | GOVERNOR CELESTINO GALLARES MEMORIAL MEDICAL CENTER | TAGBILARAN, BOHOL | 7 | Bohol | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26811 | DOH | majieocmen@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2743 ASSAY: Xpert HIV-1 Viral Load, COUNT: 56 |  | GX-IV | 820228 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800659655525019622/5800659655525019622_signature_28.png | Ma. Jernai Sosoban, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5800659655525019622/5800659655525019622_signature_31.png
| 2024-01-03 | 2023-12-07 16:17 | University of The Philippines Diliman | Quezon City | NCR | NCR 2nd - Quezon City | Marine Science | NO | INSTALLATION | CARL ZEISS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26810 | COMMERCIAL | scgonzaga@up.edu.ph | 25 |  |  | Carl Zeiss Primovert | 415510-1101-000 | Unit for installation | #NAME? | Installed unit for demo | Unit is operational. | with Engineers Bagbaguen, Daroya, Doctor |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800655894655446989/5800655894655446989_signature_28.png | Shienna Mae Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5800655894655446989/5800655894655446989_signature_31.png
| 2024-01-03 | 2024-01-03 12:23 | GOVERNOR CELESTINO GALLARES MEMORIAL MEDICAL CENTER | TAGBILARAN, BOHOL | 7 | Bohol | LABORATORY | NO | MAINTENANCE Update GeneXpert Software Version | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26809 | PBSP | majieocmen@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2743 ASSAY: Xpert HIV-1 Viral Load, COUNT: 56 |  | GX-IV | 838878 | #NAME? | #NAME? | -PERFORM GENEXPERT SOFTWARE UPDATE FROM VERSION 5.1 TO VERSION 6.1 -PERFORM PREVENTIVE MAINTENANCE AND CLEANING OF GENEXPERT WORKSTATION | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800652515523383438/5800652515523383438_signature_28.png | Ma. Jernai Sosoban, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5800652515523383438/5800652515523383438_signature_31.png
| 2024-01-03 | 2024-01-03 11:09 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26808 | PBSP | bmcgenexpertlab@yahoo.com | 25 |  | 2024-01-03 | GXIV | 834623 | Back up laptop for replacement (SN: CPH1019F03266) | #NAME? | Replaced back up laptop with new back up laptop | Unit is operational. Laptop is suspected to have defective charger/charging port. |  | Lenovo ThinkPad Laptop SN: CPH1019F03275 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800644214655435231/5800644214655435231_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5800644214655435231/5800644214655435231_signature_31.png
| 2024-01-03 | 2024-01-03 10:54 | Project 7 Social Hygiene Clinic | 39 Bansalangin St., Veterans Village, Project 7, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26807 | MACARE-HBDC | michelenaval@gmail.com | 28 |  |  | GXIV 10C-D | 120000577 | GX unit for installation | Performed PM Installed necessary drivers Setup User admin Generated report | Setup GX and accessories | unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800600110636653229/5800600110636653229_signature_28.png | Michele Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5800600110636653229/5800600110636653229_signature_31.png
| 2024-01-03 | 2024-01-03 08:00 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOT | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 26806 | DOH | marteetingzon20@gmail.con | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1312 | 2024-01-03 | GXIV | 838421 | n/a | n/a | -Monthly Preventive maintenance - Archive data - | Unit is Operational (module B1 & B4 disabled - waiting for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5800496334529462598/5800496334529462598_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5800496334529462598/5800496334529462598_signature_31.png
| 2024-01-02 | 2023-12-22 14:10 | Dasmariñas City Health Office I | D. Mangubat Ave., Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26805 | PBSP | cho1lab@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2892 ASSAY: Xpert HIV-1 Viral Load, COUNT: 143 | 2023-12-18 | GXIV-D | 838893 | Module 2 with Temperature Error | Module Replacement | Replaced Module 2 Performed Self-test Generate Reports | All Modules are Operational | B2: 863327 -> 662467 | GX Module | G 6608 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844523617120192/5799844523617120192_signature_28.png | Elma R.Belostrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844523617120192/5799844523617120192_signature_31.png
| 2024-01-02 | 2023-12-20 09:19 | Pulang Lupa Dos Health Center | St. Thomas St.,Pulang Lupa Dos, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26803 | PBSP | elvimir08@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1764 | 2023-12-18 | GXIV-D-6C | 838183 | Module 4 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 4 Generate Reports Performed Self-test | All Modules are Operational | B4: 692293 -> 696802 | GX Module 6-C | G 6598 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393612215598/5799844393612215598_signature_28.png | Elvira Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393612215598/5799844393612215598_signature_31.png
| 2024-01-02 | 2023-12-20 09:03 | E. Aldana Health Center | Interior Tramo St, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26802 | PBSP | egaguilar10@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1226 |  | GXIV-D-6C | 804261 | Unit is due fro XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 3 Modules Failed, 1 Module Passed | Modules 1,2 and 3 for Replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393616349642/5799844393616349642_signature_28.png | Edita D.Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393616349642/5799844393616349642_signature_31.png
| 2024-01-02 | 2023-12-21 11:10 | Molino III City Health Unit | Molino Road, Molino, Bacoor, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26804 | DOH | joanrnd@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1861 | 2023-12-14 | GXIV-D-6C | 821221 | Module 4 with Temperature Error | Module Replacement | Replaced Module 4 Performed Maintenance Generate Reports Performed Self-test | 3 Modules are Operational, 1 Module is Defective | B4: 886282 -> 862002 | GX Module | 10924 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393616635743/5799844393616635743_signature_28.png | Jonalyn N. Aranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393616635743/5799844393616635743_signature_31.png
| 2024-01-02 | 2023-12-20 11:14 | Caridad Health Center | P. Paterno St., Caridad, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26801 | DOH | seyerlanie85@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 615 | 2023-12-18 | GXIV-D-6C | 821227 | 2 Modules for Replacement | Module Replacement | Performed Maintenance Replaced Modules 1 and 3 Generate Reports Performed Self-test | All Modules are Operational | B1: 729891 -> 714076 & B3: 894009 -> 706171 | GX Modules | 10962 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393615110628/5799844393615110628_signature_28.png | Melanie Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393615110628/5799844393615110628_signature_31.png
| 2024-01-02 | 2023-12-21 09:52 | Putatan Health Center | MSR, Putatan, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26800 | PBSP | guimo.demesa@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 102 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1571 ASSAY: Xpert HIV-1 Viral Load, COUNT: 58 | 2023-12-18 | GXIV-D-6C | 834690 | Module 3 Failed in XpertCheck | Module Replacement | Replaced Module 3 Performed Maintenance Generate Reports Performed Self-test | All Modules are Operational | B3: 786785-> 696025 | GX Module | G 6600 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393617963938/5799844393617963938_signature_28.png | Guillermo De Mesa II | https://www.jotform.com/uploads/macarewecare/222927978372067/5799844393617963938/5799844393617963938_signature_31.png
| 2024-01-02 | 2023-12-27 13:56 | Veterans Memorial Medical Center | VMMC North Ave, Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26799 | MACARE-HBDC | evitaleona@yahoo.com | 29 | ASSAY: Xpert HBV VL, COUNT: 79 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1559 ASSAY: Xpert HCV VL, COUNT: 10 |  | GX IV 6C - D | 839150 | Unit due for preventive maintenance | Checked and cleaned cartridge bay, plunger rod, icore slot and filter | Performed preventive maintenance | Unit is operational. All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799667513611244797/5799667513611244797_signature_28.png | Evita Leona Buenavista | https://www.jotform.com/uploads/macarewecare/222927978372067/5799667513611244797/5799667513611244797_signature_31.png
| 2024-01-02 | 2023-12-28 11:07 | Cardinal Santos Medical Center | 10 Wilson street Greenhills West San Juan City | NCR | NCR 2nd - San Juan | Molecular laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | gpdmagnaye30@gmail.com |  | 26798 | COMMERCIAL | egatchalian@csmc.ph | 29 | ASSAY: Xpert HBV VL, COUNT: 93 ASSAY: Xpert HPV, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1013 |  | GX IV 10C-D | 110003714 | Run HBV on modules 2, 3, 4 Same sample all error Sample is fine | Performed PM Checked error. All error 2097 - Assay-Specific termination error Archived report for further checking by technical team. End user wants to know if it's still safe to run HBV on GX to avoid wasting cartridges Wants to know asap to avoid back logs Generated report | Checked report Archived report for technical to check | Unit is operational. Assay HBV for checking |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799663173612788302/5799663173612788302_signature_28.png | Ellaine Marie Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5799663173612788302/5799663173612788302_signature_31.png
| 2024-01-02 | 2023-12-27 10:04 | Caloocan City Health Office | Sangandaan, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26797 | PBSP | caloocancentrallaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1051 ASSAY: Xpert HIV-1 Viral Load, COUNT: 298 |  | GX IV 6C -D | 805271 | Frequent error on module A4 | Performed PM Checked module A4 system log Informed end user to get reports right away when module turns red Recommended to try using the module again Generated report | Checked module A4 | unit is operational. 2 modules under monitoring 6 modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799639923616030375/5799639923616030375_signature_28.png | Anna Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5799639923616030375/5799639923616030375_signature_31.png
| 2024-01-02 | 2023-12-27 11:16 | Tanay RHU | Municipal cmpd.Tanay Rizal | 4A | Rizal | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26796 | PBSP | amcruz1970@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1593 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 | 2023-12-18 | GX4 | 807593 | For module replacement A3 683099 - 684492 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | 900-0491rs | G6610 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799609330813524025/5799609330813524025_signature_28.png | Aireen m. Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5799609330813524025/5799609330813524025_signature_31.png
| 2024-01-02 | 2023-12-27 13:24 | Taytay rhu 5 | Sitio tayapan,taytay rizal | 4A | Rizal | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26795 | PBSP | zenresurreccion@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1420 | 2023-12-18 | GX4 | 838136 | For module replacement b1 806711 - 684235 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  |  | G6609 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5799608900816546449/5799608900816546449_signature_28.png | Zenilynn bobiles | https://www.jotform.com/uploads/macarewecare/222927978372067/5799608900816546449/5799608900816546449_signature_31.png
| 2023-12-30 | 2023-12-29 08:07 | Region 1 Medical Center | Dagupan City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26794 | PBSP | romarmtamayo@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2171 ASSAY: Xpert SARS-CoV-2, COUNT: 28 ASSAY: Xpert HIV-1 Viral Load, COUNT: 166 |  | GX4 | 804726 | For Preventive Maintenance | #NAME? | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5797189877534744209/5797189877534744209_signature_28.png | Romar M. Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5797189877534744209/5797189877534744209_signature_31.png
| 2023-12-30 | 2023-12-29 08:43 | Region 1 Medical Center | Dagupan City, Pangasinan | 1 | Pangasinan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26793 | PBSP | r1mc.molecularlab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1435 ASSAY: Xpert SARS-CoV-2, COUNT: 955 | 2023-12-29 | GX 16 | 842230 | For Preventive Maintenance and Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 897569 >>> 694832 - Generated reports | Performed Preventive Maintenance and Module Replacement | Unit is operational |  | GeneXpert Module 694832 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5797188177538792088/5797188177538792088_signature_28.png | Mark Noni Pimentel | https://www.jotform.com/uploads/macarewecare/222927978372067/5797188177538792088/5797188177538792088_signature_31.png
| 2023-12-30 | 2023-12-29 10:24 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | Tb-Dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 26792 | DOH | burauenlab@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 433 | 2023-12-29 | GXIV | 838551 | n/a | n/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5796967131915529205/5796967131915529205_signature_28.png | Mary Ann C. Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5796967131915529205/5796967131915529205_signature_31.png
| 2023-12-27 | 2023-12-27 10:47 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | REPAIR ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26791 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 195 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3410 ASSAY: Xpert HIV-1 Viral Load, COUNT: 181 |  | GXIV | 820206 | CPU is not booting | #NAME? | #NAME? | CPU is functional. Waiting for module A2 replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5794547556317445432/5794547556317445432_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5794547556317445432/5794547556317445432_signature_31.png
| 2023-12-27 | 2023-12-18 11:55 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26790 | DOH | sherryannemarrero@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1096 | 2023-11-21 | GX4 | 838281 | For Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 807811 >>> 651808 - Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 651808 | G 6455 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5794509367537021821/5794509367537021821_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5794509367537021821/5794509367537021821_signature_31.png
| 2023-12-27 | 2023-12-18 11:58 | Baguio General Hospital and Medical Center | Baguio City | 1 | Benguet | TB DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26789 | PBSP | lmsacyafen.pbsp@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 80 ASSAY: Xpert MTB/RIF Ultra, COUNT: 625 ASSAY: Others, COUNT: 24 | 2023-12-18 | GX4 10C | 120000171 | Due for Module Replacement Module B2 for Troubleshooting | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 220011687 >>>210002907 - Checked Filter - Performed Troubleshooting - Generated reports | Performed Module Replacement and Troubleshooting | Unit is operational | Self Test triggered Error 4012 | GeneXpert Module 210002907 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5794507247537409594/5794507247537409594_signature_28.png | La Femmia M. Sacyafen | https://www.jotform.com/uploads/macarewecare/222927978372067/5794507247537409594/5794507247537409594_signature_31.png
| 2023-12-27 | 2023-12-19 10:01 | Manabo Rural Health Unit | Manabo, Abra | CAR | Abra | TB DOTS | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26788 | DOH | lenmayclairelumiwag15@gmail.com | 29 |  | 2023-06-13 | GX4 | 838274 | For XpertCheck and Troubleshooting | #NAME? | Performed XpertCheck and Troubleshooting | Unit is not operational | Both fuses are blown. CPU no power. For Reporting | XpertCheck Kits - 4 | DR 5337 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5794503227535809868/5794503227535809868_signature_28.png | Ellen May Claire S. Lumiwag | https://www.jotform.com/uploads/macarewecare/222927978372067/5794503227535809868/5794503227535809868_signature_31.png
| 2023-12-24 | 2023-12-19 16:11 | Rodriguez RHU 3 | San isidro,Montalban | 4A | Rizal | Tb dots | NO | INSTALLATION | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26787 | PBSP | melodie_medina@yahoo.com | 27 |  |  | GX4 | 838894 | For installation unit transfer frm montalban rhu to rodriguez rhu 3 | Installation | Installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5792351233256187749/5792351233256187749_signature_28.png | Wenona M. Beers | https://www.jotform.com/uploads/macarewecare/222927978372067/5792351233256187749/5792351233256187749_signature_31.png
| 2023-12-23 | 2023-12-19 11:43 | North Daang Hari Super Health Center | Seagull Ave. Brgy. San Miguel, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26786 | PBSP | onlychumz@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2986 | 2023-12-18 | GX IV 6C-D | 804702 | Module A1 and A2 for replacement | Performed PM Replaced Modules Generated report | Replaced Modules A1: 661743 -> 706111 A2: 713660 -> 708486 | unit is operational All modules working |  | GX modules 6C | G 6594 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5791256613914356108/5791256613914356108_signature_28.png | Charmaine Trinidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5791256613914356108/5791256613914356108_signature_31.png
| 2023-12-23 | 2023-12-22 16:02 | Calauag Municipal Hospital | Barangay 3, Calauag, Quezon Province | 4A | Quezon | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26785 | DOH | jamila17carmel@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 834 | 2023-12-14 | GX IV 6C - D | 831910 | Module 3 for replacement | Performed PM Replaced Module 3 Generated report | Replaced Module 3 909151 -> 905825 | unit is operational All modules working |  | GX module 6C | DR 10926 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5791247293914948517/5791247293914948517_signature_28.png | Mark Errol M. Liwanag | https://www.jotform.com/uploads/macarewecare/222927978372067/5791247293914948517/5791247293914948517_signature_31.png
| 2023-12-22 | 2023-12-21 13:21 | Philippine General Hospital | Taft Ave. Manila | NCR | NCR 1st - Manila | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26784 | MACARE-HBDC | pghmolecularlab@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 161 ASSAY: Others, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 7454 |  | GX16 | 844975 | Monitor display appear continous movement of cursor in display settings | Checked and reset vga cable Checked monitor button Performed reset in monitor Test run unit | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790609544224088736/5790609544224088736_signature_28.png | Jesa Pedron | https://www.jotform.com/uploads/macarewecare/222927978372067/5790609544224088736/5790609544224088736_signature_31.png
| 2023-12-22 | 2023-12-22 15:14 | Antipolo CHO | M.Santos Antipolo | 4A | Rizal | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26783 | PBSP | acho.laboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1515 ASSAY: Xpert HIV-1 Viral Load, COUNT: 81 | 2023-12-18 | GX16 | 842226 | For module replacement B1 886028 -688971 B4 881319 -685962 C1 886306 - 686052 D4 884386 - 686588 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | B3 for reporting 15 module operational |  | 900 0491rs | G6611 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790609054228883285/5790609054228883285_signature_28.png | Gwendolyn C. Batingan | https://www.jotform.com/uploads/macarewecare/222927978372067/5790609054228883285/5790609054228883285_signature_31.png
| 2023-12-22 | 2023-12-22 16:38 | Philippine Heart Center | East Ave. QC | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR ON-SITE ASSISTANCE | HAEMONETICS | Harold Toco |  | htoco@macare-medicals.com |  |  | 26782 | COMMERCIAL | phc_bb@yahoo.com | 30 |  | 2023-12-22 | MCS+ | 11H204 | Centrufuge lock issue on opening cover. | Perfuge check up and functionality test. | Performed check up. Apply WD40 and exercise centrifuge lock mrchanism. Run actual procedure. | Unit is operational. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790439365321687516/5790439365321687516_signature_28.png | Sheila decena | https://www.jotform.com/uploads/macarewecare/222927978372067/5790439365321687516/5790439365321687516_signature_31.png
| 2023-12-22 | 2023-12-22 13:28 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | Site Visit, Check-up Machine | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26781 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1931 |  | GXIV-4D | 838561 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790326945038963844/5790326945038963844_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5790326945038963844/5790326945038963844_signature_31.png
| 2023-12-22 | 2023-12-22 13:01 | Delfin Albano Memorial Hospital | Ragan Sur, Delfin Albano, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | maryrosebacanirmt@gmail.com |  | 26780 | DOH | maryrosebacanirmt@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 904 | 2023-12-14 | Genexpert IV | 838208 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement B1(807314 - 880325) B2(807356 - 880365) Performed Functionality Test | Unit is Operational | N/a | Genexpert Modules | 10915 & 10914 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790309040311268874/5790309040311268874_signature_28.png | Mary Rose G. Bacani | https://www.jotform.com/uploads/macarewecare/222927978372067/5790309040311268874/5790309040311268874_signature_31.png
| 2023-12-22 | 2023-12-22 11:18 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26779 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 464 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2023-12-20 | GXIV | 110008530 | Unit for XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. All modules passed the calibration. |  | XpertCheck Kit Lot: 1001075092 | 11000 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790261644316836215/5790261644316836215_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5790261644316836215/5790261644316836215_signature_31.png
| 2023-12-22 | 2023-12-22 11:18 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26778 | COMMERCIAL | gramanmolecularlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-12-20 | GXIV | 110008530 | Unit for XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. All modules passed the calibration. |  | XpertCheck Kit Lot: 1001075092 | 11000 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790257994317660093/5790257994317660093_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5790257994317660093/5790257994317660093_signature_31.png
| 2023-12-22 | 2023-12-22 10:56 | Kalinga Provincial Hospital | Bulanao Centro, Tabuk City, Kalinga | CAR | Kalinga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING Xpertcheck | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26777 | DOH | leombelgica139@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1240 | 2023-11-21 | Genexpert IV | 838466 | Monitor bottom display problem | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test Performed Module Replacement B2(861495 - 604122) Performed Xpertcheck (B1 = Service Required) | Unit is Operational | Monitor for Replacement | Genexpert Module and Xpertcheck 1pc | G 6469 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790234455023225664/5790234455023225664_signature_28.png | Eden V. Asan | https://www.jotform.com/uploads/macarewecare/222927978372067/5790234455023225664/5790234455023225664_signature_31.png
| 2023-12-22 | 2023-12-22 10:49 | Kalinga Provincial Hospital | Bulanao Centro, Tabuk City, Kalinga | CAR | Kalinga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26776 | PBSP | leombelgica139@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2409 | 2023-12-22 | Genexpert IV | 807622 | Error 2008 | Performed plunger rod maintenance | Peformed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | For observation | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790229285022556654/5790229285022556654_signature_28.png | Eden V. Asan | https://www.jotform.com/uploads/macarewecare/222927978372067/5790229285022556654/5790229285022556654_signature_31.png
| 2023-12-22 | 2023-12-20 16:27 | Gubat Rural Health Unit | Gubat, Sorsogon | 5 | Sorsogon | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26775 | PBSP | concepcionbrin08@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1254 |  | GX4 - 6C | 814370 | B3/691587 for module replacement GX for Calibration | Replaced module: B3 - 691587>>682130 XpertCheck Result: B1 - passed B2 - passed B3 - new module B4 - passed Cleaned cartridge bays, plunger rod, fan, icore Generate reports | Replaced B3/691587 module Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790226558117275852/5790226558117275852_signature_28.png | Concepcion E. Brin | https://www.jotform.com/uploads/macarewecare/222927978372067/5790226558117275852/5790226558117275852_signature_31.png
| 2023-12-22 | 2023-12-22 10:28 | Tondo Medical Center | H. Lopez Blvd., Tondo Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26774 | DOH | microscan@tmc.doh.gov.ph | 25 |  | 2023-12-18 | GXIV | 831872 | Unit for PM. For installation of backup CPU. For module replacement of module A1. | Performed PM. Performed installation of CPU. Performed module replacement. Generated reports. | PM, installation and module replacement. | Unit is operational. |  | GX module A1 (879806) in replacement to defective A1 (655225). Backup Gx Desktop HP Z240 (2UA6511RLL) in replacement to defective HP Z240 (2UA8471W4X). | DR# 10964 & 11007 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790219933264216912/5790219933264216912_signature_28.png | Irychrist Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5790219933264216912/5790219933264216912_signature_31.png
| 2023-12-22 | 2023-12-21 10:17 | Philippine Children and Medical Center | Quezon Ave.QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | ppenasian1105@gamil.com |  | 26773 | PBSP | kate_july27@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2448 | 2023-12-20 | GX4 | 807672 | For module replacement b2 714190 - 699194 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | Operational |  | 900 0491 RS | G6596 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790181814223879593/5790181814223879593_signature_28.png | Katherine Tabaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5790181814223879593/5790181814223879593_signature_31.png
| 2023-12-22 | 2023-12-22 09:25 | San Jose Del Monte City Health Center VI | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26772 | PBSP | krizetteamorio@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3108 ASSAY: Xpert HIV-1 Viral Load, COUNT: 131 | 2023-12-22 | GXIV | 807620 | Module C4(798804) for replacement | #NAME? | Replacec C4 | Unit is operational. |  | GeneXpert Module SN: 693697 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790177684311441634/5790177684311441634_signature_28.png | Krizette Marie Amorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5790177684311441634/5790177684311441634_signature_31.png
| 2023-12-22 | 2023-12-22 09:12 | Atang dela Rama Health Center | Gagalangin, Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26771 | DOH | monskie_22@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2382 | 2023-12-18 | GXIV | 831644 | Unit for PM. For module replacement of B1 and B3. | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module B1 (892303) and B3 (903079) in replacement to defective B1 (756758) and B3 (614658). | DR# 10968 & 10967 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790169683261398228/5790169683261398228_signature_28.png | Monaliza Balcos | https://www.jotform.com/uploads/macarewecare/222927978372067/5790169683261398228/5790169683261398228_signature_31.png
| 2023-12-22 | 2023-12-22 07:49 | Negros Oriented Provincial Hospital | Dumaguete City, Negros Oriental | 7 | Negros Oriental | TB Laboratory | YES | REPAIR TROUBLESHOOTING | CEPHEID | Gerald Fritz Selga |  | engineering@macare-medicals.com |  |  | 26770 | PBSP | rezelmarielagos@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 130 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3965 ASSAY: Xpert HIV-1 Viral Load, COUNT: 113 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 |  | GX-4 | 805275 | Defective CPU, Boot problem | CPU cleaning, preventive maintenance. | #NAME? | System Operational with 8 modules functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5790152591445409775/5790152591445409775_signature_28.png | Rezel Marie Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5790152591445409775/5790152591445409775_signature_31.png
| 2023-12-21 | 2023-12-21 15:55 | Cagayan Vallel Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | TB Lab | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26769 | PBSP | olivesalvador10@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3366 ASSAY: Xpert SARS-CoV-2, COUNT: 6630 | 2023-12-21 | Genexpert XVI | 842894 | N/a | N/a | Installed GENEXPERT Installed UPS Installed Barcode Scanner Installed PC Installed Monitor Performed Maintenance Gathered Reports Performed Functionality Test Installed HPV assay | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789575900113593851/5789575900113593851_signature_28.png | Olive A. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5789575900113593851/5789575900113593851_signature_31.png
| 2023-12-21 | 2023-12-21 16:18 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | jaemariemacaballug1111@gmail.com |  | 26768 | MACARE-HBDC | ruellamorena@gmail.com | 27 | ASSAY: Xpert SARS-CoV-2, COUNT: 1433 | 2023-12-13 | Genexpert IV | 110008457 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test Performed Calibration | Unit is Operational | Xpertcheck Result (B1, B2, B4 = Passed) | Xpertcheck 3pc | 10894 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789564599412560409/5789564599412560409_signature_28.png | Tallion, Mary Angel Faye | https://www.jotform.com/uploads/macarewecare/222927978372067/5789564599412560409/5789564599412560409_signature_31.png
| 2023-12-21 | 2023-12-21 15:40 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE Disassemble | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26767 | DOH | zionclydematote@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert HPV, COUNT: 622 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3349 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 242 | 2023-11-21 | Genexpert IV | 831718 | Module 3 error | Performed Module Replacement (757492 - 908549) | Performed Maintenance Gathered Reports Performed Functionality Test Disassemble GENEXPERT | Unit is Operational | Genexpert ready for pickup | Genexpert Module | G 6447 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789543989412774944/5789543989412774944_signature_28.png | Sunshine Matote | https://www.jotform.com/uploads/macarewecare/222927978372067/5789543989412774944/5789543989412774944_signature_31.png
| 2023-12-21 | 2023-12-21 14:00 | Sariaya Rural Health Unit | Sariaya, Quezon Province | 4A | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26766 | DOH | marielinnetabril@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1978 |  | GXIv 6C - D | 831890 | Module A4 frequent error | Performed PM Re-XpertCheck module A4 Module A4 passed Module A4 under monitoring | Performed Re- XpertCheck on module A4 | Unit is operational. 2 modules reported 2 modules working | used 1 XpertCheck cartridge |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789483630369707258/5789483630369707258_signature_28.png | rhea parreno | https://www.jotform.com/uploads/macarewecare/222927978372067/5789483630369707258/5789483630369707258_signature_31.png
| 2023-12-21 | 2023-12-20 13:24 | Wenceslao Dela Paz Health Center | Felicia St., San Antonio Homes, Culiat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | gpdmagnaye30@gmail.com |  | 26765 | DOH | ameliafclemente2001@yahoo.com | 29 | ASSAY: Xpert HPV, COUNT: 362 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2792 | 2023-12-14 | GX IV 6C-D | 831876 | Module b4 for replacement | Performed PM Replaced module b4 Generated report | Replaced module b4 783192 -> 807309 | unit is operational all modules working |  | Gx Module 6C | DR 10910 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789475750365391599/5789475750365391599_signature_28.png | amelia clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5789475750365391599/5789475750365391599_signature_31.png
| 2023-12-21 | 2023-12-21 13:35 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26764 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 15 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 986 ASSAY: Xpert HCV VL, COUNT: 80 ASSAY: Xpert HBV VL, COUNT: 109 ASSAY: Xpert CT/NG, COUNT: 101 ASSAY: Xpert HPV, COUNT: 6 ASSAY: Xpert BCR-ABL Ultra, COUNT: 48 |  | GX-IV | 110008524 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789469055059340703/5789469055059340703_signature_28.png | Alyssa Sayson | https://www.jotform.com/uploads/macarewecare/222927978372067/5789469055059340703/5789469055059340703_signature_31.png
| 2023-12-21 | 2023-12-21 13:19 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26763 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 158 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3093 ASSAY: Xpert HIV-1 Viral Load, COUNT: 202 |  | GXIV-4D | 838285 | Module A4 frequent temperature error. | None | #NAME? | Module A4 is disabled. Other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789468369382078266/5789468369382078266_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5789468369382078266/5789468369382078266_signature_31.png
| 2023-12-21 | 2023-12-21 12:47 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26762 | DOH | nacosta@loveyourself.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 595 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1067 |  | GXIV | 831895 | Unit for troubleshooting | #NAME? | Checked unit | Unit is operational. A1, A3 for reXpertCheck A2, A4 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789435940999836014/5789435940999836014_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5789435940999836014/5789435940999836014_signature_31.png
| 2023-12-21 | 2023-12-21 12:12 | Baggao Rural Health Unit | San Jose, Baggao, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE Disassemble | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26761 | PBSP | eslabrajan02@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 838 | 2023-12-21 | Genexpert II | 835092 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test Disassemble Genexpert System | Unit is Operational | Genexpert Ready for pickup | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789414058859724239/5789414058859724239_signature_28.png | Jan Anthony Eslabra | https://www.jotform.com/uploads/macarewecare/222927978372067/5789414058859724239/5789414058859724239_signature_31.png
| 2023-12-21 | 2023-12-21 11:43 | Baggao Rural Health Unit | San Jose, Baggao, Cagayan | 2 | Cagayan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26760 | DOH | eslabrajan02@gmail.com | 27 |  | 2023-12-21 | Genexpert IV | 838531 | N/a | N/a | Installed Genexpert Installed PC Installed Monitor Installed Printer Installed UPS Installed Barcode scanner Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Genexpert from Alcala MHO | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5789410878857126636/5789410878857126636_signature_28.png | Jan Anthony Eslabra | https://www.jotform.com/uploads/macarewecare/222927978372067/5789410878857126636/5789410878857126636_signature_31.png
| 2023-12-20 | 2023-12-20 14:19 | Sanchez Mira Rural Health Unit | Sanchez Mira, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE Disassemble | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26759 | PBSP | sherylcaoagasralleca@yahoo.com | 27 |  | 2023-12-20 | Genexpert II | 835090 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test Disassemble Genexpert Disassemble PC Disassemble Monitor Disassemble Printer | Unit is Operational | Genexpert Ready for pickup | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788627800222931949/5788627800222931949_signature_28.png | Sheryl Ralleca | https://www.jotform.com/uploads/macarewecare/222927978372067/5788627800222931949/5788627800222931949_signature_31.png
| 2023-12-20 | 2023-12-20 13:34 | Padada Rural Health Unit | Rizal St. Poblacion Padada | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26758 | DOH | zhakie.marvz@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 303 | 2023-11-21 | GX-IV 4D | 838511 | Module B1 for replacement and for Preventive Maintenance | Replace the defective Module SN: 861798 with new Module SN: 738840 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Replacement | All Modules are operational | Enduser is advised to achieve the optimal room temperature before using the GeneXpert | 900-0491rs | G-6459 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788604442373296396/5788604442373296396_signature_28.png | Marvielyn C. Lusica, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5788604442373296396/5788604442373296396_signature_31.png
| 2023-12-20 | 2023-12-20 12:24 | Meycauayan City Health Center VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26757 | DOH | ronamarieiris_25@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1563 ASSAY: Xpert HIV-1 Viral Load, COUNT: 109 |  | GXIV | 831904 | Module A4 for re XpertCheck | #NAME? | Performed XpertCheck on A4 | Unit is operational. A4: passed A4 is under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788581357759132556/5788581357759132556_signature_28.png | Rona Marie Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5788581357759132556/5788581357759132556_signature_31.png
| 2023-12-20 | 2023-12-20 12:14 | Talugtug Rural Health Unit | Talugtug Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26756 | PBSP | christinejoyaliado@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 350 |  | GXIV-D | 834840 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | 2 modules passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788552690634949387/5788552690634949387_signature_28.png | Christine Joy Aliado | https://www.jotform.com/uploads/macarewecare/222927978372067/5788552690634949387/5788552690634949387_signature_31.png
| 2023-12-20 | 2023-12-20 10:40 | Tanauan City Health Office | Tanauan City | 4A | Batangas | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26755 | DOH | edithramilo04@gmail.com | 25 |  | 2023-12-18 | GXIV | 821299 | Unit for PM. For installation of assigned CPU. | Performed PM. Performed installation of CPU. Generated reports. | PM and installation | Unit is operational |  | GX desktop HPZ240 (2UA8471RRX). Pulled out backup unit HPZ240 (2UA6511RLL). | DR#10966 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788525809927471349/5788525809927471349_signature_28.png | EDITA D. RAMILO | https://www.jotform.com/uploads/macarewecare/222927978372067/5788525809927471349/5788525809927471349_signature_31.png
| 2023-12-20 | 2023-12-20 10:33 | Rogaciano M. Mercado Memorial General Hospital | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26754 | PBSP | alona0276@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2315 | 2023-12-20 | GXIV | 805252 | Module A1 (622369) for replacement | #NAME? | Replaced module A1 | Unit is operational. |  | GeneXpert Module SN: 691305 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788510737218819306/5788510737218819306_signature_28.png | Alona Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5788510737218819306/5788510737218819306_signature_31.png
| 2023-12-20 | 2023-12-20 10:40 | Marikina City Health Office | Marikina Health City Center, Marikina City | NCR | NCR 2nd - Marikina | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26753 | DOH | francescaherrera415@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1111 ASSAY: Xpert HIV-1 Viral Load, COUNT: 173 | 2023-12-18 | GX IV 6C-D | 831588 | Module B3 for replacement | Performed PM Replaced Module B3 Generated report | Replaced Module B3 657408 -> 880368 | unit is operational All modules working |  | GX module 6c | DR 10965 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788499864598445946/5788499864598445946_signature_28.png | Francesca Herrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5788499864598445946/5788499864598445946_signature_31.png
| 2023-12-20 | 2023-12-20 10:40 | Appari Provincial Hospital | Toran, Appari, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26752 | PBSP | glorivimaruquin16@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2848 | 2023-12-20 | Laboratory | 808632 | Module 4 error | Performed Module Replacement (666487 - 674284) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 2 & 3 failed in xpertcheck last 06/22/23 | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788495810221862137/5788495810221862137_signature_28.png | Glorivi R. Maruquin | https://www.jotform.com/uploads/macarewecare/222927978372067/5788495810221862137/5788495810221862137_signature_31.png
| 2023-12-20 | 2023-12-20 10:14 | Sta. Maria Rural Health Unit | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26751 | PBSP | naomidecastro35@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1443 |  | GXIV | 838886 | Stuck cartridge on module B3 | #NAME? | Removed stuck cartridge | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788477979461726962/5788477979461726962_signature_28.png | Naomi De Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5788477979461726962/5788477979461726962_signature_31.png
| 2023-12-20 | 2023-12-20 09:51 | Cuyapo rural Health Unit | Cuyapo Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26750 | DOH | jolenerodico09171998@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1114 | 2023-12-15 | GXIV-D | 831672 | A2:757048 A4:757050 for module replacemen | Module replacement Preventive maintenance | Oerformed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational A2 new sn:885042 A4 new sn: 884790 |  | 2 modulea | 10907 and 10909 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788468630788084896/5788468630788084896_signature_28.png | Jolene Rodico | https://www.jotform.com/uploads/macarewecare/222927978372067/5788468630788084896/5788468630788084896_signature_31.png
| 2023-12-20 | 2023-12-20 09:48 | Tagoloan Rural Health Unit | Poblacion, Tagoloan | 10 | Misamis Oriental | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26749 | DOH | macalisangglezeal@gmail.com | 25.925 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1258 |  | GX-4 | 838574 | Module B3 Defective | None | XpertCheck Calibration for Module B3. (Error Checking) | 3 modules working, 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788468158416484164/5788468158416484164_signature_28.png | Glezeal J. Macalisang | https://www.jotform.com/uploads/macarewecare/222927978372067/5788468158416484164/5788468158416484164_signature_31.png
| 2023-12-20 | 2023-12-20 09:08 | Puerto Princesa City Health Office | Puerto Princesa, Palawan | 4B | Palawan | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 26748 | PBSP | jocal0324@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1202 | 2023-12-20 | GXIV | 804272 | Module B4/683832 Defective | #NAME? | - Install new module on B4 w/ SN:669283 - Performed Functionality test - Performed Preventive maintenance - Cleaned Cartrige bay, Fan Filter, plunger rod - Generate reports - pullout defective module SN/683832 | unit is Operational |  | 1 Genexpert module 6 color SN/669283 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788445871817693703/5788445871817693703_signature_28.png | Jocelyn Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5788445871817693703/5788445871817693703_signature_31.png
| 2023-12-20 | 2023-12-20 08:25 | Victoria Rural Health Unit | Brgy. Bulo Victoria Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26747 | DOH | diazrmt@yahoo.com | 27 |  |  | GXIV-D | 831660 | A3 for re-xpertcheck | Re-xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except module A3 service required | A3 failed calibration, for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5788411790783451552/5788411790783451552_signature_28.png | Alma Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5788411790783451552/5788411790783451552_signature_31.png
| 2023-12-19 | 2023-12-19 16:59 | Cagayan Valley Center for Health Development | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26746 | PBSP | mctorio.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1601 ASSAY: Xpert SARS-CoV-2, COUNT: 3500 | 2023-12-19 | Genexpert IV | 834434 | Module B3 error | Performed Module Replacement 210102183 - 675756 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787858924916823820/5787858924916823820_signature_28.png | Mikeala Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5787858924916823820/5787858924916823820_signature_31.png
| 2023-12-19 | 2023-12-19 16:28 | Cagayan Valley Center for Health Development | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26745 | PBSP | mctorio.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3368 ASSAY: Xpert SARS-CoV-2, COUNT: 6630 | 2023-12-19 | Genexpert IV | 842894 | Module C1 and C4 error | Performed Module Replacement C1(620327 - 676462) C4(888629 - 679808) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Ready for pickup for delivery | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787843784919260073/5787843784919260073_signature_28.png | Mikeala Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5787843784919260073/5787843784919260073_signature_31.png
| 2023-12-19 | 2023-12-19 16:19 | Cordlife Medical Philippines Inc. | U.P.-Ayalaland Technohub, Commonwealth Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | REPAIR CALIBRATION | CYTIVA | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 26744 | COMMERCIAL | katrina.castro@cordlife.com | 25 |  | 2023-12-18 | Sepax 2 | 6220.4122 | Defective power supply | Performed parts replacement, performed functionality test, performed calibration procedure | Replaced defective power supply, calibrated unit | Unit is operational |  | Power Supply Assembly for Sepax 2 | DR# 10960 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787841763026826559/5787841763026826559_signature_28.png | Katrina Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5787841763026826559/5787841763026826559_signature_31.png
| 2023-12-19 | 2023-12-19 16:05 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26743 | PBSP | crmcmolecularlab@gmail.com | 25.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 10787 |  | Gx-IV | 834441 | Module 2 temperature error | Perform module 2 troubleshooting Genexpert machine Cleaning Generate data, IQR, System log, Assay statistics report, and archive test | Genexpert machine Preventive maintenance | 4 operational modules | Module 2 under observation due to temperature error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787841295414654346/5787841295414654346_signature_28.png | Cotabato Regional Medical Center Molecular Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/5787841295414654346/5787841295414654346_signature_31.png
| 2023-12-19 | 2023-12-19 16:09 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26742 | PBSP | crmcmolecularlab@gmail.com | 25.3 | ASSAY: Xpert SARS-CoV-2, COUNT: 10787 |  | Gx-IV | 834442 | N/A | Genexpert machine Cleaning Generate data IQR, System log, Assay statistics report | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787839265411448751/5787839265411448751_signature_28.png | Cotabato Regional Medical Center Molecular laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/5787839265411448751/5787839265411448751_signature_31.png
| 2023-12-19 | 2023-12-19 15:56 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26741 | DOH | crmcmolecularlab@gmail.com | 22.7 | ASSAY: Xpert SARS-CoV-2, COUNT: 10787 | 2023-12-19 | Gx-IV | 837662 | Module 4 for Service Required/ h/w failed | Replace/repair defect module 4 SN:902440 with new module SN:210005670 Genexpert machine Cleaning Generate data, IQR, System log, Assay statistics report and Archive test | Perform Genexpert Preventive maintenance | 4 operational modules |  | Genexpert module 6 color SN 210005670 | G-6463 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787837365416409974/5787837365416409974_signature_28.png | Cotabato Regional Medical Center Molecular Laboratory | https://www.jotform.com/uploads/macarewecare/222927978372067/5787837365416409974/5787837365416409974_signature_31.png
| 2023-12-19 | 2023-12-06 12:42 | Norala District Hospital | Poblacion, Norala, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26740 | DOH | rezielrds0709@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1226 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1226 | 2023-12-06 | Gx-IV | 838543 | Module 2 and 3 for Service Required | Replace defective module 2 SN:807699 with new module SN: 210068700 Replace defective module 3 SN:861377 with new module SN: 666557 Genexpert machine Cleaning Data collection: Pre and Post IQR, System log, Assay statistics report and Archive test | Genexpert machine Preventive maintenance Genexpert machine module repair | 4 operational modules |  | Genexpert module 6 color SN:210068700 and SN:666557 | G-6449 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787835385412241989/5787835385412241989_signature_28.png | Reziel Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/5787835385412241989/5787835385412241989_signature_31.png
| 2023-12-19 | 2023-12-19 15:06 | RIVERSIDE MEDICAL CENTER, INC. | B.S AQUINO DRIVE, BACOLOD CITY | 6 | Negros Occidental | LABORATORY DEPT. | NO | CALIBRATION | CEPHEID | Emerson Bejemino |  | 06emerson23@gmail.com |  |  | 26739 | COMMERCIAL | liza.paltriguera@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 270 ASSAY: Xpert BCR-ABL Ultra, COUNT: 19 ASSAY: Xpert HBV VL, COUNT: 132 ASSAY: Xpert HPV, COUNT: 63 ASSAY: Xpert HCV VL, COUNT: 100 |  | GX4 6 color | 110005268 | None | Conduct Preventive Maintenance and save files(IQR, System Log & Assay Stats) | Performed Xpert check calibration for 4 modules. | All 4 modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787807247215759910/5787807247215759910_signature_28.png | Liza Mae B. Paltriguera | https://www.jotform.com/uploads/macarewecare/222927978372067/5787807247215759910/5787807247215759910_signature_31.png
| 2023-12-19 | 2023-12-19 14:11 | Matanao Rural Health Unit | Poblacion, Matanao | 11 | Davao del Sur | TB-DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26738 | DOH | pamsustento@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 509 ASSAY: Xpert MTB/RIF, COUNT: 11 | 2023-11-23 | GX-IV 4D | 838456 | For Module Replacement, Preventive Maintenance and retry Calibration of Module B2 | Replace the defective Module SN: 806379 with new Module SN: 636114 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration at Module B2 | Perfrom Module Repair, Preventive and XPC Calibration | Module B2 for Service Required, Module is under observation, data will be sent to Cepheid for further Assessment | Enduser is advised to makesure achieved the optimum room temp for GeneXpert Operation | 900-0491rs | G-6473 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787776034514463821/5787776034514463821_signature_28.png | Pamela Joyce P. Sustento | https://www.jotform.com/uploads/macarewecare/222927978372067/5787776034514463821/5787776034514463821_signature_31.png
| 2023-12-19 | 2023-12-19 13:57 | Lower Bicutan Health Center | Bicutan, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26737 | PBSP | ellainemagtoto35@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3045 | 2023-12-18 | GX IV 6C-D | 838858 | Module 4 for replacement | Performed PM Replaced module 4 Generated report | Replaced module 4 863341-> 709910 | unit is operational. All modules working |  | GX module 6C | G 6592 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787753052129103546/5787753052129103546_signature_28.png | Ellaine D. Magtoto | https://www.jotform.com/uploads/macarewecare/222927978372067/5787753052129103546/5787753052129103546_signature_31.png
| 2023-12-19 | 2023-12-19 12:06 | Kalilangan Rural Health Unit | West Poblacion, Kalilangan | 10 | Bukidnon | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26736 | DOH | myrnasuyao17@gmail.com | 25.55 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 476 | 2023-11-21 | GX-4 | 838571 | B3 module defective | None | Replacing module b3 with SN: 862919 to SN: 674281 | All modules and units are functional |  | Certified pre-owned Genexpert Module | G-6451 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787692856415050769/5787692856415050769_signature_28.png | Myrna T. Suyao | https://www.jotform.com/uploads/macarewecare/222927978372067/5787692856415050769/5787692856415050769_signature_31.png
| 2023-12-19 | 2023-12-19 09:53 | Rosario Reyes Health Center | 627 San Andres St., Malate, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26735 | DOH | moninabatil16@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1596 |  | GX IV 6C-D | 831843 | Module A3 frequent error | Performed PM Re XpertCheck Module A3 Module A3 failed XpertCheck Generated report | Troubleshoot module A3 | Unit is operational 3 modules working 1 module for reporting | used 1 XpertCheck cartridge |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787613892125993552/5787613892125993552_signature_28.png | Monina Batil | https://www.jotform.com/uploads/macarewecare/222927978372067/5787613892125993552/5787613892125993552_signature_31.png
| 2023-12-19 | 2023-12-19 09:57 | Pilar Rural Health Unit | Pilar Bataan | 3 | Bataan | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26734 | DOH | aimna_24@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 620 | 2023-12-15 | GXIV-D | 831613 | A3:754416 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational | A3 new module sn:210028938 | 1 module | 10958 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787602194455611415/5787602194455611415_signature_28.png | Marciana Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5787602194455611415/5787602194455611415_signature_31.png
| 2023-12-19 | 2023-12-19 08:19 | Western Visayas Medical Center | Iloilo city | 6 | Iloilo | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26733 | PBSP | mitorilla.pbsp@gmail.com | 28 |  |  | Gx16 | 839316 | Defective module B2 sn 653541 | New module installed B2 sn 691154 | New module installed B2 sn 691154 | Systems opeartional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787584212729587126/5787584212729587126_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5787584212729587126/5787584212729587126_signature_31.png
| 2023-12-19 | 2023-12-19 08:05 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26732 | DOH | abigailgsula@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1357 | 2023-12-15 | GXIV-D | 831912 | UPS and Module A3 for replacement | Module replacement UPS replacement Preventive maintenance | Performed preventive maintenance Performed UPS and module replacement Gathered reports Performed functionality test | Unit is operational |  | Ups and 1 module | 10953 and 10905 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787536154453089650/5787536154453089650_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5787536154453089650/5787536154453089650_signature_31.png
| 2023-12-18 | 2023-12-18 19:12 | West Visayas State University Medical Center | Jaro, Iloilo | 6 | Iloilo | Labiratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26731 | PBSP | tvanrlujan@gmail.com | 34 |  |  | Gx 4 | 707531 | Damaged Fan | Open board panel and fan / clean Fan | Open board panel and fan | Defective Fan |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787079661969587926/5787079661969587926_signature_28.png | Van Ryan Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5787079661969587926/5787079661969587926_signature_31.png
| 2023-12-18 | 2023-12-18 19:03 | Davao Chest Center | Villa Abrille St,. Poblacion District Davao City | 11 | Davao del Sur | TB-DOTS | YES | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26730 | PBSP | dcclabdept@gmail.com | 21.5 |  |  | GX-IV 4D | 804835 | Modules Not Detected by the GeneXpert Software | Reset Cable Connection, From Gateway Boards connection to PSU connection Reset Network IP Connection Reinstall the GX Software | System Unit and Software troubleshooting | Still Error persist, Modules not detected by the System Unit | Tried to connect the GeneXpert to other unit (Laptop) GX Modules are detected when the GX Software Launches. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5787069693618757078/5787069693618757078_signature_28.png | Maria Lisa Paras RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5787069693618757078/5787069693618757078_signature_31.png
| 2023-12-18 | 2023-12-18 09:14 | Bangkal Health Center | Rodriguez St. | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26729 | PBSP | kapid5sepe@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2790 |  | GX-IV-D-6C | 838875 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5786768401016891835/5786768401016891835_signature_28.png | Eduardo Sepe | https://www.jotform.com/uploads/macarewecare/222927978372067/5786768401016891835/5786768401016891835_signature_31.png
| 2023-12-18 | 2023-12-18 09:11 | Bangkal Health Center | Rodriguez St. | NCR | NCR 4th - Makati | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26728 | DOH | kapid5sepe@yahoo.com | 54 |  |  | GX-IV-D-6C | 831875 | Stuck cartridge on module A3. Error 4017 on module A4 | Removed stuck cartridge | Manual removal of stuck cartridge Run XpertCheck on module A4 Performed maintenance Generated reports | Unit is operational. Module A1, A2 not operational. Module A4 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5786767781016692431/5786767781016692431_signature_28.png | Eduardo Sepe | https://www.jotform.com/uploads/macarewecare/222927978372067/5786767781016692431/5786767781016692431_signature_31.png
| 2023-12-18 | 2023-12-18 10:30 | Macare In-house | Aurora Grand QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26727 | DOH | htoco@macare-medicals.com | 25 |  |  | GXIV | 821299 | For inhouse repair of CPU | Performed repair of CPU HP Z240 (2UA8471RRX) c/o Sir Noel. | Repair | CPU is operational and ready to deploy. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5786760234912505693/5786760234912505693_signature_28.png | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5786760234912505693/5786760234912505693_signature_31.png
| 2023-12-18 | 2023-12-18 07:31 | Cauyan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | gfmselga@gmail.com |  |  | 26726 | DOH | cauayanmho@gmail.com | 27 |  |  | GX-IV | 838477 | Defective module 862945/B1. | #NAME? | Replaced module B1 with new module SN 695314. | 3 modules operational; waiting module replacement for defective module 862003/B2. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5786722834892579189/5786722834892579189_signature_28.png | Doriza Joy A. Tizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5786722834892579189/5786722834892579189_signature_31.png
| 2023-12-16 | 2023-12-16 14:44 | BILIRAN PROVINCIAL HOSPITAL | NAVAL, BILIRAN | 8 | Biliran | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26725 | PBSP | gheland01@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1780 |  | GX4 | 815273 | SN 220003268 B1 DEFECTIVE MODULE | REPLACED NEW MODULE, GENERATE IQR, SYS LOG , GXX FILES | REPLACED NEW MODULE 695537 | ALL MODULES ARE OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5785185413916685386/5785185413916685386_signature_28.png | GHELAND CORTO, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5785185413916685386/5785185413916685386_signature_31.png
| 2023-12-15 | 2023-12-15 10:20 | Makati Medical Center | Amorsolo st., Makati | NCR | NCR 4th - Makati | CRM | YES | MAINTENANCE | OTHERS | Noel Floresca |  | florescanoel77@gmail.com | biomedsupervisor@makatimed.net.ph | Miltenyi biotech | 26724 | COMMERCIAL | ma.portia.ronquillo@makatimed.net.ph | 26 |  |  | CLINIMACS | 797 | Due for maintenance | Cleaned unit Checked program parameters Checked and test liquid sensors Checked valves Checked and test display and buttons Performed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5784167069323672065/5784167069323672065_signature_28.png | Marie Portia Vanessa P. Ronquillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5784167069323672065/5784167069323672065_signature_31.png
| 2023-12-13 | 2023-12-13 11:22 | Sta Rosa Community District Hospital | Sta Rosa, Olango Island, Lapulapu City | 7 | Cebu | Tb laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26723 | DOH | ariestumz@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 199 |  | GX4 | 838534 | Due for Monthly Maintenance , Bar Code Scanner erratic | Performed Maintenance and bar code scanner calibrate | Peformed bar code scanner calibration , Clean Icore, Cartridge Bay , Filter Generate IQR, Gxx , Sys Log | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5782472843919494962/5782472843919494962_signature_28.png | Ariston Tumulak, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5782472843919494962/5782472843919494962_signature_31.png
| 2023-12-12 | 2023-12-07 09:49 | Klinika Project 7 | 29 Bansalangin, Project 7, Quezon City, Metro Manila | NCR | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26722 | MACARE-HBDC | freyaoriana@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1567 |  | GX IV 10C - D | 845446 | Modules B2 and B4 frequent error | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Module 2 and 4 for reporting Generated report | Troubleshoot module B2 and B4 | Unit is operational B1 and B3 working B2 , B4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5781710223616046316/5781710223616046316_signature_28.png | Faye Oriana | https://www.jotform.com/uploads/macarewecare/222927978372067/5781710223616046316/5781710223616046316_signature_31.png
| 2023-12-12 | 2023-12-12 09:29 | Lapulapu City Health Office | Lapulapu City Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26721 | PBSP | tanya_pax@yahoo.com | 24 |  |  | Gx4 | 808637 | Genxpert Machine unit no power | Powerswitch defective | Performed troubleshooting check fuse, check power switch | No power input Ac 220 Volts defective power switch |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5781540670102270684/5781540670102270684_signature_28.png | Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5781540670102270684/5781540670102270684_signature_31.png
| 2023-12-08 | 2023-12-08 13:52 | Vicente Sotto Memorial Medical Center | B Rodriguez St, Cebu City | 7 | Cebu | Tb laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26720 | DOH | aureamerino@gmail.com | 32 |  |  | gx4 | 820206 | 3187 | power supply of cpu need to be replaced immediately intermitent voltage supply and hard to power start up | check Cpu power supply check connection and voltage | Cpu unit as of the moment are working but need to replaced the power supply of the cpu |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5778263603913108089/5778263603913108089_signature_28.png | Sarah Nicole Aurea Merino, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5778263603913108089/5778263603913108089_signature_31.png
| 2023-12-07 | 2023-12-07 16:15 | University of The Philippines - Diliman | Diliman | NCR | NCR 2nd - Quezon City | Marine Science Institute | YES | INSTALLATION | CARL ZEISS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26719 | COMMERCIAL | scgonzaga@up.edu.ph | 24 |  |  | Primostar | 3153001559 | For installation | Installed the microscope system | Installed the microscope unit Installed router connection Checked allignment Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777495591614421754/5777495591614421754_signature_28.png | Shienna Mae Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5777495591614421754/5777495591614421754_signature_31.png
| 2023-12-07 | 2023-12-07 11:39 | Institute of Biological Science UPLB | University of the Philippines Los Baños, Los Baños, Laguna | 4A | Laguna | Phycology | NO | TROUBLESHOOTING | CARL ZEISS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26718 | COMMERCIAL | mmartinezgoss@gmail.com | 23 |  |  | Axio Scope.A1 | 3331000862 | No light | None | Checked connections Minor cleaning Performed Functionality check for 15 minutes | Unit is operational. Recommended for Preventive Maintenance. | Spotted some molds on the lens. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777304091811228795/5777304091811228795_signature_28.png | Dr. Milagrosa Goss | https://www.jotform.com/uploads/macarewecare/222927978372067/5777304091811228795/5777304091811228795_signature_31.png
| 2023-12-07 | 2023-12-07 11:36 | Davao Chest Center | Villa Abrille St. Poblacion Davao City | 11 | Davao del Sur | TB-DOTS | YES | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26717 | PBSP | dcclabdept@gmail.com | 23.5 |  |  | GX-IV 4D | 804835 | Modules are not Detected | Change Lan Cable Reset Connection for the GeneXpert Reset IP configuration at the network control panel Change and reconfigure lan ports available at System Unit | Onsite troubleshooting for Network Connection | Still error persist, no modules Detected after troubleshooting | Documented serial number of gateway board to be sent at Cepheid |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777300731426179647/5777300731426179647_signature_28.png | Maria lisa paras | https://www.jotform.com/uploads/macarewecare/222927978372067/5777300731426179647/5777300731426179647_signature_31.png
| 2023-12-07 | 2023-12-07 11:34 | Claver Rural Health Unit | Poblacion, Claver | 13 | Surigao del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26716 | DOH | rhodaruthdalogdog@yahoo.com.ph | 25 |  | 2023-11-21 | GX-4 | 838512 | Module B3 Defective | None | Replacing Module B3 with SN: 862375 to SN: 645516 | All modules and units are functional |  | Certified pre-owned Genexpert Module serial no. 645516 | G-6478 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777295231316542468/5777295231316542468_signature_28.png | Joyce E. Andoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5777295231316542468/5777295231316542468_signature_31.png
| 2023-12-07 | 2023-12-06 12:42 | Norala District Hospital | Poblacion, Norala, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26715 | DOH | rezielrds0709@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1226 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1226 | 2023-12-06 | Gx-IV | 838543 | Module 2 and 3 for Service Required | Replace defective module 2 SN:807699 with new module SN: 210068700 Replace defective module 3 SN:861377 with new module SN: 666557 Genexpert machine Cleaning Data collection: Pre and Post IQR, System log, Assay statistics report and Archive test | Genexpert machine Preventive maintenance Genexpert machine module repair | 4 operational modules |  | Genexpert module 6 color SN:210068700 and SN:666557 | G-6449 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777288974315808683/5777288974315808683_signature_28.png | Reziel Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/5777288974315808683/5777288974315808683_signature_31.png
| 2023-12-07 | 2023-12-07 09:07 | San Lazaro Hospital | Quiricada St., Sta.Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26713 | PBSP | acong.pbsp@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1504 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1155 |  | GXIV-D-6C | 804766 | Module 2 Failed in XpertCheck | Module Replacement | Replaced Module 2 Performed Maintenance Performed Self-test Generate Reports | All Modules are Operational | B2: 682956 -> 672525 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777285089021510564/5777285089021510564_signature_28.png | Ashley Cynna Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5777285089021510564/5777285089021510564_signature_31.png
| 2023-12-07 | 2023-12-07 09:03 | San Lazaro Hospital | Quiricada St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26714 | PBSP | acong.pbsp@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1504 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1155 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1504 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1155 |  | GXIV-D-6C | 707843 | Module 4 with Temperature Error | Module Replacement | Replaced Module 4 Performed Self-test Performed Maintenance Generate Reports | All Modules are Operational | C4: 210100884 -> 672497 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777285089021694613/5777285089021694613_signature_28.png | Ashley Cynna Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5777285089021694613/5777285089021694613_signature_31.png
| 2023-12-07 | 2023-12-05 12:27 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26712 | PBSP | laureanjo@yahoo.com | 28 |  |  | GX4 | 804840 | Defective module A4 717390 | New module Installed A4 662316 | New module Installed A4 662316 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777263954016578236/5777263954016578236_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5777263954016578236/5777263954016578236_signature_31.png
| 2023-12-07 | 2023-12-06 12:27 | Claro M. Recto Memorial District Hospital | Infanta | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26711 | PBSP | katelynaveryumerez@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1445 | 2023-12-06 | GX-IV-D-6C | 807608 | Module C1 for replacement | Replaced module C1 | Performed maintenance Replaced module C1 Generated reports | Unit is operational. Module C4 has error 1002, for reporting. |  | Old: C1/743403 New: C1/653351 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777257573618139469/5777257573618139469_signature_28.png | Lorelyn U. Miras | https://www.jotform.com/uploads/macarewecare/222927978372067/5777257573618139469/5777257573618139469_signature_31.png
| 2023-12-07 | 2023-12-07 10:23 | Lapulapu City Health Office | Poblacion Lapu lapu City | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26710 | PBSP | ceren_sy@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2580 |  | Gx4 | 808637 | Module 2 Defective 626040 | Replaced new module 2 | Replaced Module 2 New Module 666460 Generate IQR, Sys Log, GXX files | All.modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777252840312644252/5777252840312644252_signature_28.png | jovy villariasa | https://www.jotform.com/uploads/macarewecare/222927978372067/5777252840312644252/5777252840312644252_signature_31.png
| 2023-12-07 | 2023-12-07 09:58 | BUREAU OF PLANT INDUSTRY | BPI COMPOUND, LOS BAÑOS, LAGUNA | 4A | Laguna | ENTOM LAB | YES | REPAIR | CARL ZEISS | Carlos Leniel Origenes |  | carloa.leniel.origenes@gmail.com |  |  | 26709 | COMMERCIAL | jestongaylo@gmail.com | 23 |  | 2023-12-06 | STEMI 305 | 3947019527 | Defective Spot Illuminator, works internittently | Replaced Spot Illuminator | Replaced Spot Illuminator | Unit is operational |  | Spot illuminator | DR 10880, SI 96817 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777237891714887566/5777237891714887566_signature_28.png | Jestoni O. Gaylo | https://www.jotform.com/uploads/macarewecare/222927978372067/5777237891714887566/5777237891714887566_signature_31.png
| 2023-12-07 | 2023-12-07 09:32 | Northern Mindanao Medical Center | Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | HIV Area | YES | ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26708 | PBSP | ronnel_mabasa@yahoo.com | 26 |  |  | GXIV | 804682 | Monitor no display/signal. | #NAME? | None | Unit is functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777220452148480168/5777220452148480168_signature_28.png | Ronnel Mabasa | https://www.jotform.com/uploads/macarewecare/222927978372067/5777220452148480168/5777220452148480168_signature_31.png
| 2023-12-07 | 2023-12-07 09:03 | Bulan Rural Health Unit | Bulan | 5 | Sorsogon | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | danemilbodino23@gmail.com |  | 26707 | PBSP | gemmaalbello1970@gmail.com | 23 |  |  | GX4 | 807983 | Temperature error, 1001 | Checked fan filter, distance of machine from wall, ambient temp Generate and review reports | Performed troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777210014341590138/5777210014341590138_signature_28.png | Gemma M. Albello | https://www.jotform.com/uploads/macarewecare/222927978372067/5777210014341590138/5777210014341590138_signature_31.png
| 2023-12-07 | 2023-12-06 15:05 | National Kidney and Transplant Institute | East avenue,QC | NCR | NCR 2nd - Quezon City | Molecular | YES | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26706 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1053 ASSAY: Others, COUNT: 35 ASSAY: Xpert HBV VL, COUNT: 712 ASSAY: Xpert HPV, COUNT: 132 ASSAY: Xpert MTB/RIF, COUNT: 834 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1623 ASSAY: Xpert SARS-CoV-2, COUNT: 15336 ASSAY: Xpert HCV VL, COUNT: 75 | 2023-12-04 | GX16 | 110005587 | B4 encounter temp error 1002, C4 encounter temp error 4009,4012 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | A4 reported for replacement, B4 and C4 for reporting. | C2 module replacement done 801983 (old) - 210009197 (new) | 900-0491rs | 10875 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5777183983618836461/5777183983618836461_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5777183983618836461/5777183983618836461_signature_31.png
| 2023-12-06 | 2023-12-06 16:54 | Northern Mindanao Medical Center | Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | HIV Laboratory | YES | INSTALLATION TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26705 | PBSP | ronnel_mabasa@yahoo.com | 27 |  |  | GXIV | 804682 | For reinstallation | #NAME? | Transfer and Installation of GeneXpert Unit | Module 1 is defective; for replacement. Other modules are functional. | HDMI to VGA converter is defective, no display to monitor. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776624650221355154/5776624650221355154_signature_28.png | Ronnel Mabasa | https://www.jotform.com/uploads/macarewecare/222927978372067/5776624650221355154/5776624650221355154_signature_31.png
| 2023-12-06 | 2023-12-06 16:30 | Madrid District Hospital | Poblacion, Madrid | 13 | Surigao del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26704 | DOH | madrid_dh@yahoo.com | 29 |  | 2023-11-21 | GX-4 | 838439 | None | None | Replacing Module B3 with Sn: 806426 to SN: 633480 | All modules and units are functional |  | Certified pre-owned genexpert module Serial no. 633480 | G-6472 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776609852129502490/5776609852129502490_signature_28.png | Nicole Correos | https://www.jotform.com/uploads/macarewecare/222927978372067/5776609852129502490/5776609852129502490_signature_31.png
| 2023-12-06 | 2023-12-06 15:34 | National Center for Mental Health | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26703 | PBSP | iancristrono@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1591 |  | GXIV | 808625 | Unit for relocation | #NAME? | Relocated unit | Unit is operational. Temp. errors on B3 B3 is under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776578935421878814/5776578935421878814_signature_28.png | Ian Cris Trono | https://www.jotform.com/uploads/macarewecare/222927978372067/5776578935421878814/5776578935421878814_signature_31.png
| 2023-12-06 | 2023-12-06 13:55 | Panabo City Health Office | Poblacion, Panabo City | 11 | Davao del Norte | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26702 | DOH | nadineeust67@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2530 | 2023-11-21 | GX-IV 4d | 838295 | Module B4 for replacement due to Calibration Failure For Preventive Maintenance | Replaced the Defective Module SN:808298 with new Module SN: 608521 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Replacement | All Modules are Operational |  | 900-0491rs | G-6452 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776519901424014931/5776519901424014931_signature_28.png | Louella Jane G. Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5776519901424014931/5776519901424014931_signature_31.png
| 2023-12-06 | 2023-12-06 13:19 | Philippine Red Cross - Ormoc | Ormoc City | 6 | Leyte | Laboratory | YES | REPAIR TROUBLESHOOTING | HAEMONETICS | Raymond Laurito |  | daymon899@gmail.com |  |  | 26701 | COMMERCIAL | ormoc@redcross.org.ph | 28 |  |  | MCS + | 15H043G | CRC BIOS ERROR | Take out CPU board and bring to cebu office to replace the battery | Temporary Replacement of CPU Board from Cebu Office spareparts | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776506792472601225/5776506792472601225_signature_28.png | Phil. Red Cross Ormoc City | https://www.jotform.com/uploads/macarewecare/222927978372067/5776506792472601225/5776506792472601225_signature_31.png
| 2023-12-06 | 2023-12-06 13:27 | Tarlac Provincial Hospital | Hospital Drive Brgy. San Vicente Tarlac City | 3 | Tarlac | PCR | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26700 | PBSP | genextphmpl@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 1651 |  | GXIV-D | 838142 | B3:909438 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replavement Gathered reports Performed functionality test | Unit is operational | New B3 SN:672264 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776502644914578727/5776502644914578727_signature_28.png | Ethel Mae D. Avellanosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5776502644914578727/5776502644914578727_signature_31.png
| 2023-12-06 | 2023-12-06 11:10 | Muñoz City Health Office | Science City of Muñoz, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26699 | DOH | scomrtdl@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 747 | 2023-12-06 | Genexpert IV | 831673 | CPU Problem | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | CPU for monitoring | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776419445326754384/5776419445326754384_signature_28.png | Catherine Celestino | https://www.jotform.com/uploads/macarewecare/222927978372067/5776419445326754384/5776419445326754384_signature_31.png
| 2023-12-06 | 2023-12-06 09:47 | Lung Center of the Philippines | Quezon Avenue, Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Microbiology and Bacteriology | NO | ON-SITE ASSISTANCE | CEPHEID | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26698 | MACARE-HBDC | lcppathologyresults@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 378 ASSAY: Xpert SARS-CoV-2, COUNT: 1714 ASSAY: Xpert HIV-1 Viral Load, COUNT: 93 ASSAY: Others, COUNT: 12 |  | GXIV-D | 838837 | Modules 1 and 4 are disabled | None, modules are for purchase. | Upgrade software from 5.1 to 6.4 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776376275029936554/5776376275029936554_signature_28.png | Lucrecia S Bongato | https://www.jotform.com/uploads/macarewecare/222927978372067/5776376275029936554/5776376275029936554_signature_31.png
| 2023-12-06 | 2023-12-06 10:01 | NUEVA VALENCIA RURAL HEALTH UNIT | NUEVA VALENCIA, GUIMARAS | 6 | Guimaras | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26697 | DOH | nuevarhu@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 133 | 2023-11-21 | GX-IV | 838291 | FOR MODULE B2 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE B2 REPLACEMENT DEF SN:806869 NEW SN:649042 | EQUIPMENT IS OPERATIONAL |  | GX MODULE | G 6479 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776375905221005923/5776375905221005923_signature_28.png | Rhea April Q Jaleco | https://www.jotform.com/uploads/macarewecare/222927978372067/5776375905221005923/5776375905221005923_signature_31.png
| 2023-12-06 | 2023-12-06 09:19 | Nazareth Health Center | Brgy. Nazareth, Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26696 | DOH | nanettebtolentino@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1302 |  | GXIV | 838269 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776353218399530691/5776353218399530691_signature_28.png | Nanette Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/5776353218399530691/5776353218399530691_signature_31.png
| 2023-12-06 | 2023-12-05 10:06 | Datu halun sakilan Memorial Hospital | Datu halun | BARMM | Tawi-Tawi | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 26695 | DOH | joel_kamlian2003@yahoo.com | 21 |  | 2023-11-21 | GX 4 | 838834 | Module 1 hardware failed | Replaced module 861191 with Module 636514 | Module Replacement and Preventive Maintenence | All modules are available | None | 900-0491RS | G 6475 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776242954224949089/5776242954224949089_signature_28.png | Joel A Kamlian | https://www.jotform.com/uploads/macarewecare/222927978372067/5776242954224949089/5776242954224949089_signature_31.png
| 2023-12-06 | 2023-12-04 15:39 | Makati Medical Center | Amorsolo st., | NCR | NCR 4th - Makati | Bloodbank | NO | ON-SITE ASSISTANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | majarocon022@gmail.com |  | 26694 | COMMERCIAL | rovijaelle.lare@makatimed.net.ph | 28 |  |  | MCS + | 97H330 | Barcode scanner open wire, for removal from the unit. Ever since the barcode scanner is not in use | Removed barcode scanner,checked unit and program parameters, Test run unit | For removal of barcode scanner | Unit operational ready to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5776191097494910002/5776191097494910002_signature_28.png | Elle Lare | https://www.jotform.com/uploads/macarewecare/222927978372067/5776191097494910002/5776191097494910002_signature_31.png
| 2023-12-05 | 2023-12-05 13:32 | Bagumbayan Rural Health Unit | Poblacion Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26693 | DOH | bambi.vic6@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1403 | 2023-12-05 | Gx-IV | 838261 | Module 2 for Service Required | Perform Genexpert machine Preventive maintenance | Replace defective module 2 SN:765336 with new module SN: 624012 Genexpert machine Cleaning Wash/clean filter Data collection: Pre and Post IQR, System log, Assay statistics report and archive test | 4 operational modules |  | Genexpert module 6 color SN: 624012 | G 6471 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775946735723920949/5775946735723920949_signature_28.png | Videlyne Carlet | https://www.jotform.com/uploads/macarewecare/222927978372067/5775946735723920949/5775946735723920949_signature_31.png
| 2023-12-05 | 2023-12-05 15:57 | Juban Rural Health Unit | North Poblacion, Juban | 5 | Sorsogon | RTDL/TML Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | danemilbodino23@gmail.com |  | 26692 | PBSP | gaytarean@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 828 |  | GX4 | 835214 | B1/789036 for Module Replacement GX for Calibration | Replaced Module: B1 - 789036>>649810 XpertCheck Result: B1/649810 - new module B2/210091002 - service required(after 2 tries) | Replaced B2 Module Performed XpertCheck | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775737223114118922/5775737223114118922_signature_28.png | Rean G. Gracilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5775737223114118922/5775737223114118922_signature_31.png
| 2023-12-05 | 2023-12-05 15:41 | Nasipit Rural Health Unit | Brgy. Talisay, Nasipit | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26691 | DOH | peterdolotina@gmail.com | 25.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1359 | 2023-11-21 | GX-4 | 837611 | None | None | Replacing Module B4 with SN: 678339 to SN: 210060886 | All modules and units are functional |  | Certified pre-owned GeneXpert Module Serial No. 210060886 | G-6468 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775722517756588627/5775722517756588627_signature_28.png | Peter Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/5775722517756588627/5775722517756588627_signature_31.png
| 2023-12-05 | 2023-12-05 13:04 | Quirino Memorial Medical Center | J.P. Rizal, Project 4, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26690 | PBSP | carlopalomera58@gmail.com | 29 | ASSAY: Xpert CT/NG, COUNT: 26 ASSAY: Others, COUNT: 21 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3254 ASSAY: Xpert HIV-1 Viral Load, COUNT: 58 | 2023-12-04 | GX IV 6C-D | 815484 | Module B1 for replacement Stuck cartridge module B4 | Performed PM Replaced module B1 Removed stuck cartridge on B4 Generated report | Replaced module B1 750224 -> 674840 Troubleshoot module b4 | unit is operational. all modules working |  | GX module 6C | DR 45620 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775656233619645819/5775656233619645819_signature_28.png | Carlo Palomera | https://www.jotform.com/uploads/macarewecare/222927978372067/5775656233619645819/5775656233619645819_signature_31.png
| 2023-12-05 | 2023-12-05 13:41 | Siniloan Rural Health Unit | Siniloan | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26689 | PBSP | siniloanruralhealthunit@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 779 | 2023-12-04 | GX-IV-D-6C | 838165 | Module B3 for replacement | Replaced module B3 | Performed maintenance Replaced module B3 Generated reports | All modules operational. |  | Old: B3/806778 New: B3/667078 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775648618639725628/5775648618639725628_signature_28.png | Ethel May G. Meraña | https://www.jotform.com/uploads/macarewecare/222927978372067/5775648618639725628/5775648618639725628_signature_31.png
| 2023-12-05 | 2023-12-05 13:37 | Monkayo Rural Health Unit | Poblacion Monkayo, Davao de Oro | 11 | Davao de Oro | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26688 | DOH | kendralois730@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1691 | 2023-11-21 | GX-IV 4d | 838487 | Module B1 for replacement due to Calibration Error For Preventive Maintenance | Replaced the Defective Module SN:862177 with New Module SN: 643605 Clean Surface Area, Cartridge Bay, Plunger Rod fan filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | 900-0491rs | G6477 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775643731429537782/5775643731429537782_signature_28.png | Karen S. Anuta, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5775643731429537782/5775643731429537782_signature_31.png
| 2023-12-05 | 2023-12-05 13:19 | Magsaysay Rural Health Unit | Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26687 | DOH | jkakilgan12@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 317 | 2023-11-21 | GX-4 | 838557 | Non | Nonw | Replacing Module B1 with SN: 862581 to SN: 623297 Replacing B4 with SN: 862613 to SN: 623346 | All modules and units are functional |  | Certified Pre-Owned Genexpert Module Serial No. 623297 and 623346 | G-6470 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775632745719110993/5775632745719110993_signature_28.png | Jesusa Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/5775632745719110993/5775632745719110993_signature_31.png
| 2023-12-05 | 2023-12-05 12:39 | Carmen Health Center | Brgy. Carmen, Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26686 | DOH | carmenhealthcenterlab@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 | 2023-11-21 | GXIV | 838442 | Module B1, B3, & B4 - xpertcheck service required. | -Replace module B1/806450 with new module SN: 878085 -Replace module B3/806452 with new module SN: 790874 -Replace module B4/731631 with new module SN: 623431 | #NAME? | All modules are functional. |  | 3 GeneXpert Module 6-Colors | G6445 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775611638395462139/5775611638395462139_signature_28.png | Juvy T. Madarang | https://www.jotform.com/uploads/macarewecare/222927978372067/5775611638395462139/5775611638395462139_signature_31.png
| 2023-12-05 | 2023-12-05 12:27 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26685 | PBSP | laureanjo@yahoo.com | 28 |  |  | GX4 | 804840 | Defective module A4 717390 | New module Installed A4 662316 | New module Installed A4 662316 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775610553643000791/5775610553643000791_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5775610553643000791/5775610553643000791_signature_31.png
| 2023-12-05 | 2023-12-05 11:10 | Cagayan de Oro City Health Office | Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26684 | DOH | chocdotbdots@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3418 | 2023-12-05 | GXIV | 838435 | Module C1 H/W failed | None | #NAME? | Module C1 xpertcheck service required; for replacement. Other modules are functional. |  | 1 XpertCheck | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775551978395969975/5775551978395969975_signature_28.png | Charissa Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5775551978395969975/5775551978395969975_signature_31.png
| 2023-12-05 | 2023-12-05 11:01 | Cagayan de Oro City Health Office | Cagayan de Oro City, Misamis Oriental | 10 | Misamis Oriental | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26683 | DOH | chocdotbdots@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3418 | 2023-11-21 | GXIV | 838441 | Module B2 and B3 Xpert Check service required. | -Replace defective module B2/731414 with new module SN: 883448 -Replace defective module B3/806414 with new module SN: 699450 | #NAME? | All modules are functional. |  | 2 GeneXpert Module 6-Colors | G6448 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775548928397164695/5775548928397164695_signature_28.png | Charissa Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5775548928397164695/5775548928397164695_signature_31.png
| 2023-12-05 | 2023-12-05 10:40 | Vira Rural Health Unit | Virac | 5 | Catanduanes | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26682 | DOH | limarmariano50@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1734 |  | GX4 | 838563 | B2/862611 for module replacement | Replaced module: B2 - 862611>>210030593 Cleaned cartridge bays, plunger rod, icores,fan Generate reports | Replaced module B2 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775535862228700681/5775535862228700681_signature_28.png | Ramil A. Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/5775535862228700681/5775535862228700681_signature_31.png
| 2023-12-05 | 2023-12-04 12:33 | Iriga City Health Office | Iriga | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26681 | DOH | mongealyzza@yahoo.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1834 |  | GX4-6C | 838391 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icores Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775531642227224129/5775531642227224129_signature_28.png | Alyzza Rosabelle Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5775531642227224129/5775531642227224129_signature_31.png
| 2023-12-05 | 2023-12-05 09:22 | Datu halun sakilan Memorial Hospital | Lamion, Datu halun | BARMM | Tawi-Tawi | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 26680 | DOH | rappaiballaho@gmail.com | 27 |  | 2023-12-05 | GX 4 | 808643 | Module 3 hardware failed | Replaced module 732750 with Module 624779 | Module Replacement and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775531216014522916/5775531216014522916_signature_28.png | Rappai o. Ballaho | https://www.jotform.com/uploads/macarewecare/222927978372067/5775531216014522916/5775531216014522916_signature_31.png
| 2023-12-05 | 2023-12-05 10:22 | Eastern Samar Provincial Hospital | Borongan City | 8 | Eastern Samar | TB-DOTS | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | lermafernandez@yahoo.com |  | 26679 | PBSP | nathanycarmel@gmail.com | 34 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1281 ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 | 2023-12-05 | GXIV | 804256 | - CPU system error (not booting) - SN/2UA3450YRR | #NAME? | #NAME? | Unit is Operational | CPU is now back to normal condition |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775530720811675718/5775530720811675718_signature_28.png | Glynnie Operario | https://www.jotform.com/uploads/macarewecare/222927978372067/5775530720811675718/5775530720811675718_signature_31.png
| 2023-12-05 | 2023-12-05 09:33 | Bagong Pag-Asa Health Center | 35 Rd. 9 cor Rd. 10, Bagong Pag-Asa, Quezon City | 4A | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26678 | DOH | erikatagorda18@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 471 |  | GX IV 6C-D | 831597 | module A2 error 2014 | performed pm performed ReXpertCheck on module A2 Module A2 failed ReXpertCheck Generated report | troubleshoot module A2 | unit is operational A1 reported A2 failed reXC A3, A4 working | used 1 XpertCheck cartridge |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775501761456634421/5775501761456634421_signature_28.png | erika lucas | https://www.jotform.com/uploads/macarewecare/222927978372067/5775501761456634421/5775501761456634421_signature_31.png
| 2023-12-05 | 2023-12-04 14:16 | Medilinx (MMC) | Amorsolo st., Makati | NCR | NCR 4th - Makati | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | John.Manzano@sodexo.com |  | 26677 | COMMERCIAL | reinaluz.bactol@medilinx.com.ph | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 145 ASSAY: Xpert CT/NG, COUNT: 657 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 686 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2555 ASSAY: Others, COUNT: 2 ASSAY: Others, COUNT: 12 ASSAY: Xpert HCV VL, COUNT: 63 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1827 | 2023-12-04 | GX4- D(10C) | 110001647 | For module replacement b1 21008154- 210076418 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement | Operational |  | 900-0650rs | 10869 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5775325610917076140/5775325610917076140_signature_28.png | Reinaluz Bactol | https://www.jotform.com/uploads/macarewecare/222927978372067/5775325610917076140/5775325610917076140_signature_31.png
| 2023-12-04 | 2023-12-04 19:38 | CEBU CITY MEDICAL CENTER | CEBU CITY | 7 | Cebu | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26676 | DOH | ccmclaboratory01@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1729 |  | GX-IV | 838561 | FOR MODULE A4 REPLACEMENT | #NAME? | -PERFORM CLEANING AND PREVENTIVE MAINTENANCE OF GENEXPERT EQUIPMENT -PERFORM MODULE A4 REPLACEMENT: DEFECTIVE SN: 802146 , NEW SN:636313 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774993399117157103/5774993399117157103_signature_28.png | Fatimah Maraviles, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5774993399117157103/5774993399117157103_signature_31.png
| 2023-12-04 | 2023-12-04 16:28 | Calauag Municipal Hospital | New Municipal Govermment Site, Barangay Calauag, Quezon | 4A | Quezon | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26675 | DOH | jamila17carmel@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 827 | 2023-11-21 | GX IV 6C-D | 831910 | Module A1 for replacement | Performed PM Replaced Module A1 GGenerated report | Replaced Module A1 758305 -> 210009608 | Unit is operational 3 modules working 1 module reported | Jotform Sync Error. Original MR date 12/01/23 | Gx 4 6c | G6466 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774882523614790660/5774882523614790660_signature_28.png | Jamila Pedrezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5774882523614790660/5774882523614790660_signature_31.png
| 2023-12-04 | 2023-11-30 15:17 | Quezon Medical Center | Quezon Medical Center, Brgy. 11, Lucena, Quezon City | 4A | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26674 | PBSP | mariannenaynes081774@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 295 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2191 ASSAY: Xpert HIV-1 Viral Load, COUNT: 105 |  | GX IV 6C | 807462, 804692 | Modules frequent error 2008 | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test on modules All modules passed Advised end user to continue using the modules and perform maintenance regularly Generated report | Troubleshoot modules | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774873963616338701/5774873963616338701_signature_28.png | Marianne Naynes | https://www.jotform.com/uploads/macarewecare/222927978372067/5774873963616338701/5774873963616338701_signature_31.png
| 2023-12-04 | 2023-12-04 14:37 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | INHAIL | 26673 | COMMERCIAL | biopathqc@biopath.ph | 23 |  | 2023-07-19 | HEPA Filter | FA-EQ-HF | NONE | SEALED HOLES IN FLEXIBLE DUCTING | REPLACED FILTERS SEALED REMAINING OPENINGS | VENTILLATION UNITS ARE OPERATIONAL WITH NEW FILTERS |  | 30cm x 30cm x 2cm HEPA Filters (FA-EQ-HF) x4 units | 10106 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774815711136130491/5774815711136130491_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5774815711136130491/5774815711136130491_signature_31.png
| 2023-12-04 | 2023-12-01 12:07 | Lingap Kalusugan CHU I - Alima | Alima, Bacoor, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26671 | PBSP | lyzz.jimenez@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1637 | 2023-11-28 | GXIV-D | 838185 | Module 1 with Temperature Error | Module Replacement | Replaced Module 1 Performed Self-test Generate Reports | All Modules are Operational |  | 900-0491RS | G 6498 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774776803613313998/5774776803613313998_signature_28.png | Alyzza janelle jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/5774776803613313998/5774776803613313998_signature_31.png
| 2023-12-04 | 2023-11-29 09:56 | Balayan Rural Health Unit | Brgy. Gumamela, Balayan, Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26671 | PBSP | balayantbdots@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1092 | 2023-11-28 | GXIV-4 | 838451 | Modules 1 and 2 with Temperature Errors | Module Replacement | Replaced Modules 1 and 2 Performed Maintenance Performed Self-test Generate Reports | All Modules are Operational |  | GX Module | G 6497 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774776803614857921/5774776803614857921_signature_28.png | Veronica Dalida | https://www.jotform.com/uploads/macarewecare/222927978372067/5774776803614857921/5774776803614857921_signature_31.png
| 2023-12-04 | 2023-11-30 14:57 | National Tuberculosis Reference Laboratory | Filinvest Corporate City, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26672 | PBSP | mspogosa.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 30 ASSAY: Others, COUNT: 2525 | 2023-11-24 | GXIV-D-10C | 848244 | Module 1 with Temperature Error | Module Replacement | Replaced Module 1 Performed Maintenance Performed Self-test Generate Reports | All Modules are Operational |  | 900-0650RS | G 6496 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774776803619983902/5774776803619983902_signature_28.png | Maria Guileane Pogosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5774776803619983902/5774776803619983902_signature_31.png
| 2023-12-04 | 2023-12-04 13:03 | Gingoog City Health Office | Rizal Street, Poblacion, Gingoog City | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26670 | PBSP | lagbasmabel@gmail.com | 27.45 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1630 |  | GX-4 | 807572 | None | None | Replacing Module B2 with SN: 674371 to SN: 210057600 | All modules and units are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774771328318900440/5774771328318900440_signature_28.png | Mabel Lagbas | https://www.jotform.com/uploads/macarewecare/222927978372067/5774771328318900440/5774771328318900440_signature_31.png
| 2023-12-04 | 2023-12-04 13:13 | Lala Rural Health Unit | Lala, Lanipao Lanao del Norte | 12 | Lanao del Norte | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26669 | DOH | villagonzalojoyb@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 916 |  | GX4 | 838229 | Module 2 Serial Number 806826 failed during calibration | Replaced new module , clean Icore , clean fan, Required end user to transfer the unit to other area due high temperature might cause irreparable damage to any modules in genexpert | Replaced new module 2 serial number 210005638 Generate data System Log, Installation Qualification report, Assay Statistics | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774765022517009408/5774765022517009408_signature_28.png | Maria Glenda Villagonzalo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5774765022517009408/5774765022517009408_signature_31.png
| 2023-12-04 | 2023-12-04 11:32 | Kapatagan Rural Health Unit | Kapatagan , Lanao del Norte | 12 | Lanao del Norte | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26668 | DOH | jhayrajaculbe@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 730 |  | GX4 | 838224 | Module 2 Serial Number 659378 failed during calibration | Replaced new module , clean filter and fan | Replaced new module in Module 2 Serial Number 210070828 Generate Gxx files, IQR , System Log | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774758802517239095/5774758802517239095_signature_28.png | Jhayra Jaculbe, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5774758802517239095/5774758802517239095_signature_31.png
| 2023-12-04 | 2023-12-01 11:54 | Tipo-Tipo Rural Health Unit | Tipo-Tipo, Basilan | BARMM | Basilan | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia@gmail.com | niagnozia54@gmail.com |  | 26667 | PBSP | fatmaaunal84@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 673 | 2023-12-01 | GX 4 | 812773 | Module 2 hardware failed | Replaced module 686161 with Module 686363 | Module Replacement and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774757802438528868/5774757802438528868_signature_28.png | Fatma Aunal | https://www.jotform.com/uploads/macarewecare/222927978372067/5774757802438528868/5774757802438528868_signature_31.png
| 2023-12-04 | 2023-12-04 11:05 | Hinaplanon RTDL | Hinaplanon, Iligan City | 12 | Lanao del Norte | Tb laboratory | NO | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com |  |  | 26666 | DOH | Villacortaamy50@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 971 |  | GX4 | 838227 | Module 2 and Module 4 failed during calibration 807236 and 807440 respectively. | Replacement of Defective module 2 and module 4 | Replaced new Module 2- 667515 and Module 4 667655 Generate data Sys log, Installation Qualification Report, Gxx files Clean Filter and fan motor | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774692982518927068/5774692982518927068_signature_28.png | Amy S. Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/5774692982518927068/5774692982518927068_signature_31.png
| 2023-12-04 | 2023-11-30 13:18 | DRSTMH | Kalibo, Aklan | 6 | Aklan | Laboratory | NO | Installing AVR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26665 | MACARE-HBDC | hact.drstmh@gmail.com | 28 |  |  | Gx 4 | 110020745 | No AVR unit during installation | Installing AVR back up unit from Cebu Office | Installing AVR back up unit from Cebu Office | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774684393445574459/5774684393445574459_signature_28.png | Carlo franz bontogon | https://www.jotform.com/uploads/macarewecare/222927978372067/5774684393445574459/5774684393445574459_signature_31.png
| 2023-12-04 | 2023-11-30 13:18 | DRSTMH | Kalibo, Aklan | 6 | Aklan | Laboratory | NO | Installing AVR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26664 | MACARE-HBDC | hact.drstmh@gmail.com | 28 |  |  | Gx 4 | 110020745 | No AVR unit during installation | Installing AVR back up unit from Cebu Office | Installing AVR back up unit from Cebu Office | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774681123447155939/5774681123447155939_signature_28.png | Carlo franz bontogon | https://www.jotform.com/uploads/macarewecare/222927978372067/5774681123447155939/5774681123447155939_signature_31.png
| 2023-12-04 | 2023-11-22 11:49 | Nagpayong Super Health Center | Nagpayong | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26663 | DOH | michelindg23@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 925 |  | GX-IV-D-6C | 831905 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765171052694774066/5765171052694774066_signature_28.png | Michelle De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5765171052694774066/5765171052694774066_signature_31.png
| 2023-12-04 | 2023-11-29 13:21 | Tampakan Rural Health Unit | Poblacion Tampakan, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26662 | DOH | jmedtech_lab@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 968 |  | Gx4 | 838544 | Module 1 H/W failed waiting for replacement module | Perform Genexpert machine Preventive maintenance | Genexpert machine Cleaning Data collection, IQR, Assay statistics report, S. Log and archive test | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774559944317242307/5774559944317242307_signature_28.png | Jovi G. Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/5774559944317242307/5774559944317242307_signature_31.png
| 2023-12-04 | 2023-12-01 10:52 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26661 | DOH | jan19.cam@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1545 |  | Gx-IV | 838504 | Module 3 H/w failed waiting for replacement module | Perform Genexpert machine Preventive maintenance | Genexpert machine Cleaning Wash/Clean Filter Data collection: IQR, S.log, Assay statistics repot and Archive test | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5774559944318347330/5774559944318347330_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5774559944318347330/5774559944318347330_signature_31.png
| 2023-12-01 | 2023-12-01 17:34 | St Lukes Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | Patrick Los Baños |  | plosbanos@macare-medicals.com |  |  | 26660 | MACARE-HBDC | laboratory@slec.ph | 22 |  | 2023-11-28 | GX Infinity 48s | 809643 | Modules A1, A3, B1, B2, B3, C2, C3, D1, E1, E2, E4, F2, F3, F4, G2, G3 | Performed maintenance cleaning Replaced defective modules: 749978 704797 804179 702072 676644 710723 713581 702345 677939 718717 752696 701127 692430 701126 749224 719051 682673 695487 699651 694626 615706 695334 636265 692341 699635 710088 631346 702129 714776 695066 696264 694263 Perfomed Xpertcheck on Modules H1, H2, H3, H4: H2/696244, H4/672198 needs servicing | Replaced Defective Module Performed Xpertcheck on Modules H1, H2, H3, H4 | Unit is operational | With John Monsour Doctor, Cienne Mae Valerio, Rhea Anne A. Cañelas, Mike Pritos | GX Module 6-color: 749978 804179 676644 713581 677939 752696 692430 749224 682673 699651 615706 636265 699635 631346 714776 696264 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772331450419752520/5772331450419752520_signature_28.png | MA. SARAH R. SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5772331450419752520/5772331450419752520_signature_31.png
| 2023-12-01 | 2023-11-29 10:38 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26659 | DOH | agnesabalos220@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1018 | 2023-10-24 | GX4 | 838473 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Modules 724823 >>>706187 and 862048 >>> 627680 | Performed Module Replacement | Unit is operational | Work with Engr. Venzon | GeneXpert Modules 706187 and 627680 | DR 10774 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772259762087680542/5772259762087680542_signature_28.png | Agnes Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/5772259762087680542/5772259762087680542_signature_31.png
| 2023-12-01 | 2023-11-29 13:27 | Philippine Red Cross Dagupan | Dagupan City, Pangasinan | 1 | Pangasinan | Blood bank | NO | MAINTENANCE | HAEMONETICS | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26658 | MACARE-HBDC | prcpangasinanbloodcenter@gmail.com | 23 |  |  | MCS + | 03A043 | Due for Inspection | #NAME? | Performed Inspection and Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772258092087781995/5772258092087781995_signature_28.png | Mae Dianne Alano | https://www.jotform.com/uploads/macarewecare/222927978372067/5772258092087781995/5772258092087781995_signature_31.png
| 2023-12-01 | 2023-11-30 08:11 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26657 | PBSP | r1mcxpertlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1430 ASSAY: Xpert SARS-CoV-2, COUNT: 1422 |  | GX 16 | 842230 | Due for Troubleshooting | #NAME? | Performed Troubleshooting | Unit is operational | Module B3 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772254102087191332/5772254102087191332_signature_28.png | Kirstin Deinielle S. Banaag | https://www.jotform.com/uploads/macarewecare/222927978372067/5772254102087191332/5772254102087191332_signature_31.png
| 2023-12-01 | 2023-12-01 10:18 | La Union Medical Center | Agoo, La Union | 1 | La Union | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26656 | PBSP | lumcrtdl@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2004 |  | GX4 | 707400 | For Troubleshooting of UPS | #NAME? | Performed Troubleshooting on UPS | Unit is operational | UPS not functioning. No power is measured when UPS is on or in backup mode. Site is to request UPS to PBSP. GeneXpert is currently plugged directly to the wall outlet. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772249772089876961/5772249772089876961_signature_28.png | Josephine Paglingayen | https://www.jotform.com/uploads/macarewecare/222927978372067/5772249772089876961/5772249772089876961_signature_31.png
| 2023-12-01 | 2023-12-01 13:41 | Sto Domingo District Hospital | Sto Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26655 | DOH | lhaydeesel@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 948 ASSAY: Xpert MTB/RIF, COUNT: 48 ASSAY: Xpert MTB/RIF Ultra, COUNT: 948 ASSAY: Xpert MTB/RIF, COUNT: 49 | 2023-12-01 | Genexpert IV | 831582 | Module 4 error | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 4 for observation | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772187972187512015/5772187972187512015_signature_28.png | Myrish Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/5772187972187512015/5772187972187512015_signature_31.png
| 2023-12-01 | 2023-12-01 11:25 | Lower Bicutan Health Center | Lower Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26654 | PBSP | ellainemagtoto35@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3067 |  | GX-IV-D-6C | 838858 | Module B4 has error 4009/1002 | Checked module B4 | Performed maintenance Checked module B4 Disabled module B4 Generated reports | Unit is operational. Module B4 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772106751153633453/5772106751153633453_signature_28.png | Ellaine Magtoto | https://www.jotform.com/uploads/macarewecare/222927978372067/5772106751153633453/5772106751153633453_signature_31.png
| 2023-12-01 | 2023-12-01 10:56 | Laguindingan Rural Health Unit | Poblacion | 9 | Misamis Oriental | Laboratory | YES | INSTALLATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 26653 | DOH | chellesambaan@gmail.com | 23 |  | 2023-12-12 | GX4 | 838575 | Module B1 hardware fail | Replace module B1 | Replace module B1 with new serial number 210041816 | Module B1, 2,4 ok to use module B3 not ok to use due to error code 1002 20x from October to December 2023 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772101773119470715/5772101773119470715_signature_28.png | Rochelle M. Samban | https://www.jotform.com/uploads/macarewecare/222927978372067/5772101773119470715/5772101773119470715_signature_31.png
| 2023-12-01 | 2023-11-30 14:46 | Amai Pakpak Medical Center | Marawi City | 12 | Basilan | laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 26652 | PBSP | rbarco2007@gmail.com | 23 |  | 2023-11-30 | Gx4 | 807457 | Module C1, C3 calibration fail. | Replace module | Replace module C1, C3 with new serial number 655108,655962 | All modules are functional ok to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772085973117389651/5772085973117389651_signature_28.png | Raisah P. Barco | https://www.jotform.com/uploads/macarewecare/222927978372067/5772085973117389651/5772085973117389651_signature_31.png
| 2023-12-01 | 2023-12-01 10:27 | North Daang Hari Super Health Center | Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26651 | PBSP | onlychumz@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3055 |  | GX-IV-D-6C | 804702 | Module A1 has error 4009. Module A2 is resetting | Checked modules A1 and A2 | Performed maintenance Checked module A1 and module A2 Checked internal connections Checked internal voltage and gateway board Generated reports Checked UPS | Unit is operational. Module A1 and A2 not operational. UPS not operational. | Module A2's ultrasonic horn cable is cut. UPS SN: 40020332002 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772075461159275951/5772075461159275951_signature_28.png | Chummy | https://www.jotform.com/uploads/macarewecare/222927978372067/5772075461159275951/5772075461159275951_signature_31.png
| 2023-12-01 | 2023-11-30 14:14 | Kalinga Health Hub - Marikina | JP Rizal St. | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26650 | CONSORTIUM | jasmin.catral@innovationsch.org | 24 |  | 2023-11-29 | GX-IV-D-6C | 110008418 | Module Be for replacement | Replaced module B3 | Performed maintenance Replaced module B3 Generated reports | All modules operational. |  | Old: B3/210035666 New: B3/616087 | DR# 10859 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772068271152539401/5772068271152539401_signature_28.png | Jasmin Catral | https://www.jotform.com/uploads/macarewecare/222927978372067/5772068271152539401/5772068271152539401_signature_31.png
| 2023-12-01 | 2023-12-01 09:40 | WVSUMC | Jaro, iloilo city | 6 | Iloilo | Laboratory | NO | INSTALLATION | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26649 | MACARE-HBDC | zielglet_09@yahoo.com | 27 |  |  | Gx 4 | 110021716 | New installation | New installation | New installation | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772044154111216564/5772044154111216564_signature_28.png | Gleziel T. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5772044154111216564/5772044154111216564_signature_31.png
| 2023-12-01 | 2023-12-01 08:50 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26648 | DOH | sajecopooh20@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1282 | 2023-12-01 | GXIV | 838555 | Module A2 frequent temperature error. | None | #NAME? | Module A2 failed XPC; for replacement. Waiting for module A1 and A3 replacements. |  | XpertCheck | None |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772019111701953318/5772019111701953318_signature_28.png | Sarah Jean Raga Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5772019111701953318/5772019111701953318_signature_31.png
| 2023-12-01 | 2023-12-01 08:55 | WVSU | Iloilo City | 7 | Iloilo | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26647 | PBSP | tvanrlujan@gmail.com | 34 |  |  | Gx 4 | 707531 | Internal Temp High | Clean Fan | Clean Fan | Fan does not turning on |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772016934112668169/5772016934112668169_signature_28.png | Van Ryan R . Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5772016934112668169/5772016934112668169_signature_31.png
| 2023-12-01 | 2023-11-30 14:29 | Philippine Children Medical Center | Quezon Avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com | ernaandreaescobar013@gmail.com |  | 26646 | PBSP | sdtl.pcmc@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2341 |  | GX4 | 807672 | Due for maintenance, B2 module encounter temp error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | B2 for reporting,B1,B3,B4 operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5772001423614436226/5772001423614436226_signature_28.png | Marie Zarina Mae D. Orbito | https://www.jotform.com/uploads/macarewecare/222927978372067/5772001423614436226/5772001423614436226_signature_31.png
| 2023-11-30 | 2023-11-30 13:18 | DRSTMH | Kalibo, Aklan | 6 | Aklan | Laboratory | NO | Installing AVR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26645 | MACARE-HBDC | hact.drstmh@gmail.com | 28 |  |  | Gx 4 | 110020745 | No AVR unit during installation | Installing AVR back up unit from Cebu Office | Installing AVR back up unit from Cebu Office | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771498711965000753/5771498711965000753_signature_28.png | Carlo franz bontogon | https://www.jotform.com/uploads/macarewecare/222927978372067/5771498711965000753/5771498711965000753_signature_31.png
| 2023-11-30 | 2023-11-30 13:18 | DRSTMH | Kalibo, Aklan | 6 | Aklan | Laboratory | NO | Installing AVR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26644 | MACARE-HBDC | hact.drstmh@gmail.com | 28 |  |  | Gx 4 | 110020745 | No AVR unit during installation | Installing AVR back up unit from Cebu Office | Installing AVR back up unit from Cebu Office | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771498069939779360/5771498069939779360_signature_28.png | Carlo franz bontogon | https://www.jotform.com/uploads/macarewecare/222927978372067/5771498069939779360/5771498069939779360_signature_31.png
| 2023-11-30 | 2023-11-30 10:42 | Baao Rural Health Unit | Baao | 5 | Camarines Sur | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26643 | PBSP | naschw.4u@gmail.com | 28 |  |  | GX4 | 808639 | CPU for replacement | Replaced CPU - CPH0723F43458 Setup GXDx settings Generate reports | Replaced CPU | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771456603426949897/5771456603426949897_signature_28.png | Ignacio B. Britanico | https://www.jotform.com/uploads/macarewecare/222927978372067/5771456603426949897/5771456603426949897_signature_31.png
| 2023-11-30 | 2023-11-29 20:36 | Bicol Medical Center | Naga | 5 | Camarines Sur | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26642 | PBSP | bmc.bacteriology@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1387 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7237 ASSAY: Xpert SARS-CoV-2, COUNT: 2368 ASSAY: Xpert HIV-1 Viral Load, COUNT: 342 ASSAY: Others, COUNT: 8 |  | GX16 | 84222 | GX16 for additional troubleshooting and data collection | Checked Gateway board of C4. Generate reports | Collected sufficient Data as requested by Cepheid | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771454613425713881/5771454613425713881_signature_28.png | Angelika Mae Bichara | https://www.jotform.com/uploads/macarewecare/222927978372067/5771454613425713881/5771454613425713881_signature_31.png
| 2023-11-30 | 2023-11-29 12:05 | Bicol South Luzon Sub National Reference Laboratory | BRTTH Compound, Rizal St., Legazpi City | 5 | Albay | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26641 | PBSP | bsl-snrl@bicol.doh.gov.ph | 28 | ASSAY: Xpert HPV, COUNT: 444 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1754 ASSAY: Xpert HIV-1 Viral Load, COUNT: 285 |  | GX4 | 807579 | Error 2126 for all modules | Checked power cords, plugged Gx to UPS Generate reports | Performed troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771453013422043457/5771453013422043457_signature_28.png | Verson Olicia | https://www.jotform.com/uploads/macarewecare/222927978372067/5771453013422043457/5771453013422043457_signature_31.png
| 2023-11-30 | 2023-11-30 16:44 | BIOPATH CAP | CEBU CITY | 7 | Cebu | MOLECULAR LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26640 | COMMERCIAL | jkee@biopath.ph | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 1679 |  | GX-XVI | 838657 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771431899115266731/5771431899115266731_signature_28.png | Mary Joy Kee | https://www.jotform.com/uploads/macarewecare/222927978372067/5771431899115266731/5771431899115266731_signature_31.png
| 2023-11-30 | 2023-11-30 16:33 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26639 | DOH | riazarizzamae@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 986 | 2023-11-30 | Genexpert IV | 831637 | Module 2 Hardware Failure | Performed Xpertcheck on Module 2 = (Service Required) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771428468104298077/5771428468104298077_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5771428468104298077/5771428468104298077_signature_31.png
| 2023-11-30 | 2023-11-30 14:47 | Bongabon District Hospital | Curva, Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26638 | DOH | tsantuajr0128@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 534 | 2023-11-30 | Genexpert IV | 831562 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 1 for observation | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771364813644366189/5771364813644366189_signature_28.png | Liezl J. Adriano | https://www.jotform.com/uploads/macarewecare/222927978372067/5771364813644366189/5771364813644366189_signature_31.png
| 2023-11-30 | 2023-11-30 13:09 | Nabua Rural Health Unit | Nabua | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | danemilbodino23@gmail.com |  | 26637 | PBSP | mongealyzza@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 307 |  | GX4 | 834620 | B1/648078 & B3/674962 for module replacement | Replaced Module: B1/648078>>734479, B3/674962>>793128 Cleaned cartridge bays, plunger rod, fan Generate reports | Replaced B1 & B3 module | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771342324013790126/5771342324013790126_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5771342324013790126/5771342324013790126_signature_31.png
| 2023-11-30 | 2023-11-30 09:54 | Biliran Provincial Hospital | Naval, Biliran | 8 | Biliran | TB Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26636 | PBSP | ghelmission13@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1461 ASSAY: Xpert HIV-1 Viral Load, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1476 ASSAY: Xpert HIV-1 Viral Load, COUNT: 17 | 2023-11-30 | GXIV | 815273 | Module B4 defective. | -Replace defective module B4/877768 with new module SN: 657173 | #NAME? | Module B1 for replacement. Other modules are functional. |  | GeneXpert Module 6 Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771318029511161948/5771318029511161948_signature_28.png | Gheland Corto | https://www.jotform.com/uploads/macarewecare/222927978372067/5771318029511161948/5771318029511161948_signature_31.png
| 2023-11-30 | 2023-11-30 13:03 | Calubian Rural Health Unit | Calubian Leyte | 8 | Leyte | TB dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 26635 | DOH | melitaaringoy1@gmail.com | 27 |  | 2023-11-30 | GXIV | 838399 | Module B1/662277 and B4/860988 for replacement | #NAME? | - Install new module on B1 w/ SN 642566 - Install new module on B4 w/ SN 642770 - Performed Functionality test - Cleaned Cartridge bay, Plunger rod and Fan Filter - Generate reports - Pullout defective modules (662277, 860988) | Unit is Operational |  | 2 GeneXpert Module SN (642566,642770) | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771305325393326628/5771305325393326628_signature_28.png | Melita E. Aringoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5771305325393326628/5771305325393326628_signature_31.png
| 2023-11-30 | 2023-11-30 12:05 | Muñoz City Health Office | Science City of Muñoz, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26634 | DOH | scomrtdl@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 746 | 2023-11-21 | Genexpert IV | 831673 | Genexpert PC auto restart | Performed Genexpert PC troubleshooting | Performed Maintenance Gathered Reports Performed Module Replacement (754609 - 210005937) Performed Functionality Test | Unit is Operational | Genexpert PC for observation | Genexpert Module - 210005937 | G 6464 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771266426507055432/5771266426507055432_signature_28.png | Catherine Celestino | https://www.jotform.com/uploads/macarewecare/222927978372067/5771266426507055432/5771266426507055432_signature_31.png
| 2023-11-30 | 2023-11-30 11:31 | Oriental Mindoro Souther District Hospital | Brgy. Odiong, Oriental Mindoro | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26633 | PBSP | markedisonmartinez@yahoo.com | 25 |  | 2023-11-07 | GXIV | 807590 | Unit for PM. For module replacement of modules B2 and B4. | Performed PM. Performed module replacement on B2 and B4. Generated reports. | PM and module replacement. | Unit is operational. Module B3 for reporting. |  | GX module B2 (655759) and B4 (686738) in replacement to defective B2 (712980) and B4 (718766) | 45422 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771247899223164921/5771247899223164921_signature_28.png | Mark Edison G. Martinez, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5771247899223164921/5771247899223164921_signature_31.png
| 2023-11-30 | 2023-11-30 10:55 | Quezon Medical Center Wellness Hub | Quezon Medical Center Compound, Brgy. 11, Lucena, Quezon Province | 4A | Quezon | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26632 | SHELL FOUNDATION | qmclivepositivewellnesshub@gmail.com | 25 |  | 2023-07-28 | GX IV 10C | 120000146 | GX machine for installation | Checked installation site Unpacked GX and accessories Setup GX and accessories Added users Generated report | Installed GX and all its accessories | Unit is operational. All modules working |  | GX Machine | Bill Landing No. 1551 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771225331225994750/5771225331225994750_signature_28.png | Zenjee V. Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5771225331225994750/5771225331225994750_signature_31.png
| 2023-11-30 | 2023-11-30 08:42 | Cainta City Health Office | Karangalan Dr | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26631 | DOH | ballesterrichard18@gmail.com | 24 |  |  | GX-IV-D-6c | 831623 | Module A1 and A2 has error 1002 | Run XpertCheck on module A1 and A2 | Performed maintenance Run XpertCheck on module A1 and A2 Generated reports | Unit is operational. Module A1 and A2 passed in XpertCheck but under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771206486723886957/5771206486723886957_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5771206486723886957/5771206486723886957_signature_31.png
| 2023-11-30 | 2023-11-30 10:24 | Cainta City Health Office | Karangalan Dr. | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26630 | PBSP | ballesterrichard18@gmail.com | 24 |  |  | GX-IV-D-6C | 838182 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771205076721012459/5771205076721012459_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5771205076721012459/5771205076721012459_signature_31.png
| 2023-11-30 | 2023-11-30 08:27 | Passi City Health Office | Passi City | 7 | Iloilo | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26629 | PBSP | passicholab@gmail.com | 28 |  |  | Gx 4 | 804701 | Defective module SN 646647 | New module installed SN 654986, clean fan fliter | New module installed SN 654986 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5771139996712226906/5771139996712226906_signature_28.png | John Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5771139996712226906/5771139996712226906_signature_31.png
| 2023-11-29 | 2023-11-29 17:44 | West Visayas Medical Center | Iloilo City | 6 | Iloilo | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26628 | PBSP | mitorilla.pbsp@gmail.com | 27 |  |  | Gx 4 | 706910 | Defective Module A3 683885 | New module installed A3 660775 | New module installed A3 660775 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5770607222284496772/5770607222284496772_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5770607222284496772/5770607222284496772_signature_31.png
| 2023-11-29 | 2023-11-29 15:45 | San Jose Del Monte City Health Center VI | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26627 | PBSP | theadelrosario1@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2038 ASSAY: Xpert HIV-1 Viral Load, COUNT: 121 | 2023-11-28 | GXIV | 814161 | Monitor for replacement | #NAME? | Replaced monitor | Unit is operational. |  | GeneXpert Monitor SN: V5GAH819 | G 6499 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5770533759126068484/5770533759126068484_signature_28.png | Thea Del Rosario | https://www.jotform.com/uploads/macarewecare/222927978372067/5770533759126068484/5770533759126068484_signature_31.png
| 2023-11-29 | 2023-11-29 15:05 | Candelaria Rural Health Unit | Candelaria, Quezon Province | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26626 | PBSP | shirleyzrubz@gmail.com | 32 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1680 |  | GX IV 6C | 814155 | Modules frequent 2008 error | Checked and cleaned cartridge bays, plunger rod, icore slot and filters Checked system log, temp too high Performed reliability test, all passed Advised to use fan to lower the temp Generated report | Checked modules | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5770516150614605800/5770516150614605800_signature_28.png | Shirley R. Dollente | https://www.jotform.com/uploads/macarewecare/222927978372067/5770516150614605800/5770516150614605800_signature_31.png
| 2023-11-29 | 2023-11-29 14:22 | Mexico Rural Health Unit IV | Pandacaqui, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26625 | DOH | annbathan79@gmail.com | 27 | ASSAY: , COUNT: 457 | 2023-11-29 | Genexpert IV | 831902 | N/a | N/a | Performed Maintenance Gathered Reports Installed Printer Driver Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5770501968572727925/5770501968572727925_signature_28.png | Mary Ann Carreon | https://www.jotform.com/uploads/macarewecare/222927978372067/5770501968572727925/5770501968572727925_signature_31.png
| 2023-11-29 | 2023-11-29 14:25 | VICENTE SOTTO MEMORIAL MEDICAL CENTER | CEBU CITY | 7 | Cebu | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26624 | DOH | vsmmcdeptpatho.genexpert@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 180 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3166 ASSAY: Xpert HIV-1 Viral Load, COUNT: 145 |  | GX-IV | 820206 | SYSTEM UNIT ABLE TO BOOT BUT DISPLAY IS NOT SHOWN | #NAME? | -INSPECT EQUIPMENT POWER AND CABLE CONNECTIONS -RESEAT DISPLAY PORT CONNECTION FROM SYSTEM UNIT SN:2UA8471RRQ TO DISPLAY CONNECTION | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5770486273913704058/5770486273913704058_signature_28.png | Myra Riccil Estose | https://www.jotform.com/uploads/macarewecare/222927978372067/5770486273913704058/5770486273913704058_signature_31.png
| 2023-11-29 | 2023-11-29 10:31 | Western Visayas Sanitarium | Santa barbara, iloilo | 6 | Iloilo | Laboratiry | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26623 | DOH | migzko@yahoo.com | 28 |  |  | Gx 4 | 838393 | Defective D4 module Sn 692166 | New module installed D4 Sn 883440, clean fan filter | New module installed D4 Sn 883440 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5770348565964035246/5770348565964035246_signature_28.png | Ruth T. Dequilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5770348565964035246/5770348565964035246_signature_31.png
| 2023-11-29 | 2023-11-29 08:49 | Hinabangan Rural Health Unit | Hinabangn Samar | 8 | Samar | TB-DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 26622 | PBSP | nabloruth58@gmail.com | 25 |  | 2023-11-29 | GXIV-2D | 835099 | Due for Calibration/Preventive maintenance | #NAME? | #NAME? | Unit is Operational | All module passed (B1,B2) | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5770305798418438818/5770305798418438818_signature_28.png | Ruth c nablo | https://www.jotform.com/uploads/macarewecare/222927978372067/5770305798418438818/5770305798418438818_signature_31.png
| 2023-11-29 | 2023-11-28 11:11 | Eastern Visayas Medical Center | Bagacay, Tacloban City | 8 | Leyte | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 26621 | PBSP | shairamethystcantos@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4338 | 2023-11-28 | GXIV | 808631 | Module B4/681450 defective | #NAME? | - Install new module on B4 w/ SN: 657253 - Performed Functionality test - Performed Preventive maintenance - Generate reports - Pullout defective module SN/681450 | Unit is Operational |  | 1 GeneXpert Module 6 Color SN/657253 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5770160958328779689/5770160958328779689_signature_28.png | Shaira Amethyst C. Nicolas | https://www.jotform.com/uploads/macarewecare/222927978372067/5770160958328779689/5770160958328779689_signature_31.png
| 2023-11-28 | 2023-11-28 18:59 | Chong Hua Hospital- Mandaue | Subangdaku , Mandaue City Cebu | 7 | Cebu | Bloodbank | YES | REPAIR | HAEMONETICS | Mark Efren Ramas |  | ramasmarkefren@gmail.co |  |  | 26620 | COMMERCIAL | irl.maderazo@gmail.com | 24 |  |  | Mcs+ | 11J217 | Red Pump no motion | Replaced pump | Replaced Red Pump and White Pump | Units are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769786135114289757/5769786135114289757_signature_28.png | Inna Maderazo,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5769786135114289757/5769786135114289757_signature_31.png
| 2023-11-28 | 2023-11-28 14:41 | Bugasong Rural Health Unit | Bugasong, Antique | 6 | Antique | TB Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26619 | PBSP | jeanelyngabuyogimalay@gmail.com | 24.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 609 | 2023-11-28 | GXIV | 815427 | Module B4 defective | -Replace defective module B4/801807 with new module SN: 871535 | #NAME? | All modules are functional. |  | GeneXpert Module 6 Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769632414817136345/5769632414817136345_signature_28.png | Jeanelyn Gabuyog Imalay | https://www.jotform.com/uploads/macarewecare/222927978372067/5769632414817136345/5769632414817136345_signature_31.png
| 2023-11-28 | 2023-11-28 13:28 | Minqlin Rural Health Unit | Minalin Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26618 | PBSP | bernadetteann0318@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 472 |  | GXIV-D | 834688 | CPU hard boot up | Troubleshooting Preventive maintenance | Performed CPU Troubleshooting Gathered reports Performed functionality test | Unit is operational | CPU under observation(power button error) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769587326615356199/5769587326615356199_signature_28.png | Bernadette Lazaro | https://www.jotform.com/uploads/macarewecare/222927978372067/5769587326615356199/5769587326615356199_signature_31.png
| 2023-11-28 | 2023-11-28 11:58 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26617 | DOH | abigailgsula@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1339 |  | GXIV-D | 831912 | A4 for module replacement UPS not backing-up | Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational except A2 and A3 | UPS fir replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769551046613451661/5769551046613451661_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5769551046613451661/5769551046613451661_signature_31.png
| 2023-11-28 | 2023-11-24 14:56 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26616 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | VORTEX MIXER | K3CT82010446 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 200RPM AND 3000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769483549321461637/5769483549321461637_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769483549321461637/5769483549321461637_signature_31.png
| 2023-11-28 | 2023-11-24 15:02 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26615 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | VORTEX MIXER | K3CT82010445 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 200RPM AND 3000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769482969324343929/5769482969324343929_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769482969324343929/5769482969324343929_signature_31.png
| 2023-11-28 | 2023-11-24 15:09 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26614 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | VORTEX MIXER | K3CT82010021 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 200RPM AND 3000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769482469324838970/5769482469324838970_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769482469324838970/5769482469324838970_signature_31.png
| 2023-11-28 | 2023-11-24 15:24 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | IRIS | 26613 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | STATSPIN EXPRESS 2 | 731M50106146 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 1000RPM AND 3000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769481889327822928/5769481889327822928_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769481889327822928/5769481889327822928_signature_31.png
| 2023-11-28 | 2023-11-24 15:33 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26612 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | THERMAL MIXER | 1014319110746 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 500RPM AND 2000RPM VERIFIED TEMPERATURES AT 55°C AND 60°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769481039325308560/5769481039325308560_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769481039325308560/5769481039325308560_signature_31.png
| 2023-11-28 | 2023-11-24 15:18 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26611 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MYSPIN 12 | HSH12116 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 500RPM AND 12500RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769479829325728189/5769479829325728189_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769479829325728189/5769479829325728189_signature_31.png
| 2023-11-28 | 2023-11-24 15:21 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26610 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MYSPIN 12 | HSH12119 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 500RPM AND 12500RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769478929326484687/5769478929326484687_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769478929326484687/5769478929326484687_signature_31.png
| 2023-11-28 | 2023-11-24 15:22 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26609 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MYSPIN 12 | HSH12117 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 500RPM AND 12500RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769477869328110155/5769477869328110155_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769477869328110155/5769477869328110155_signature_31.png
| 2023-11-28 | 2023-11-24 15:22 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26608 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MYSPIN 12 | HSH12117 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 500RPM AND 12500RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769477269326873816/5769477269326873816_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769477269326873816/5769477269326873816_signature_31.png
| 2023-11-28 | 2023-11-24 15:28 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26607 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MEDIFUGE | 42358139 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 1000RPM AND 5000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769476419323963647/5769476419323963647_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769476419323963647/5769476419323963647_signature_31.png
| 2023-11-28 | 2023-11-24 15:31 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | LW SCIENTIFIC | 26606 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | EZ SWING | 1101273 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 1000RPM AND 5000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769475839324074131/5769475839324074131_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769475839324074131/5769475839324074131_signature_31.png
| 2023-11-28 | 2023-11-24 15:37 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | ESCO | 26605 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | IFA-32-8 | 2020-T05163 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 56°C AND 60°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769475129324001373/5769475129324001373_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769475129324001373/5769475129324001373_signature_31.png
| 2023-11-28 | 2023-11-24 15:39 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | ESCO | 26604 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | IFA-32-8 | 2020-T05006 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 56°C AND 60°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769467549325187881/5769467549325187881_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769467549325187881/5769467549325187881_signature_31.png
| 2023-11-28 | 2023-11-24 15:41 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | CONDURA | 26603 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | CTD311MNI | 18956716056 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 2°C, 8°C AND -20°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769465979324312899/5769465979324312899_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769465979324312899/5769465979324312899_signature_31.png
| 2023-11-28 | 2023-11-24 15:45 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | CONDURA | 26602 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | CBC283-NL | 688103101W | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 2°C AND 8°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769465279326051419/5769465279326051419_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769465279326051419/5769465279326051419_signature_31.png
| 2023-11-28 | 2023-11-24 15:47 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | CONDURA | 26601 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | CBC283-NL | 688103125W | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 2°C AND 8°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769464629322865476/5769464629322865476_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769464629322865476/5769464629322865476_signature_31.png
| 2023-11-28 | 2023-11-24 15:47 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | CONDURA | 26600 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | CBC283-NL | 688103125W | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 2°C AND 8°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769464039329108084/5769464039329108084_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769464039329108084/5769464039329108084_signature_31.png
| 2023-11-28 | 2023-11-24 15:49 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | BIOBASE | 26599 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | BKQ-B50 II | BKQ50II2007541 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 134°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769463179326512610/5769463179326512610_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769463179326512610/5769463179326512610_signature_31.png
| 2023-11-28 | 2023-11-24 15:49 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | BIOBASE | 26598 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | BKQ-B50 II | BKQ50II2007541 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 134°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769462639329584039/5769462639329584039_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769462639329584039/5769462639329584039_signature_31.png
| 2023-11-28 | 2023-11-24 15:51 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | MEDX | 26597 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | LX-B50L | F5020078 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 129°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769462029325381250/5769462029325381250_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769462029325381250/5769462029325381250_signature_31.png
| 2023-11-28 | 2023-11-24 15:51 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | MEDX | 26596 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | LX-B50L | F5020078 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 129°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769461469322919146/5769461469322919146_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769461469322919146/5769461469322919146_signature_31.png
| 2023-11-28 | 2023-11-24 15:51 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | MEDX | 26595 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | LX-B50L | F5020078 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 129°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769460949324657075/5769460949324657075_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769460949324657075/5769460949324657075_signature_31.png
| 2023-11-28 | 2023-11-24 15:53 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | PERKIN ELMER | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26594 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | PRE NAT II | 3310-2141 | NONE | NONE | PERFORMED SURFACE CLEANING CHECKED AND VERIFIED ALIGNMENTS PERFORMED WATER RUN PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769460239329183442/5769460239329183442_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769460239329183442/5769460239329183442_signature_31.png
| 2023-11-28 | 2023-11-24 15:53 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | PERKIN ELMER | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26593 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | PRE NAT II | 3310-2141 | NONE | NONE | PERFORMED SURFACE CLEANING CHECKED AND VERIFIED ALIGNMENTS PERFORMED WATER RUN PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769459579326898557/5769459579326898557_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769459579326898557/5769459579326898557_signature_31.png
| 2023-11-28 | 2023-11-24 15:59 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | RBC BIOSCIENCE | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26592 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MAGCORE PLUS II | P26R022020135 | NONE | NONE | PERFORMED SURFACE CLEANING CHECKED AND VERIFIED ALIGNMENTS PERFORMED WATER RUN PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769458969326434279/5769458969326434279_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769458969326434279/5769458969326434279_signature_31.png
| 2023-11-28 | 2023-11-24 16:02 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26591 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | TSX2330FZ | 1161322301200520 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURE AT -20°C PERFORMED DEFROSTING PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769458509326346099/5769458509326346099_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769458509326346099/5769458509326346099_signature_31.png
| 2023-11-28 | 2023-11-24 16:04 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26590 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | 88400D | 160414501161220 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURE AT -80°C PERFORMED DEFROSTING PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769458069327860322/5769458069327860322_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769458069327860322/5769458069327860322_signature_31.png
| 2023-11-28 | 2023-11-24 15:57 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. ROD. SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR LAB | NO | CALIBRATION MAINTENANCE | RBC BIOSCIENCE | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26589 | COMMERCIAL | biopathqc@biopath.ph | 23 |  |  | MAGCORE PLUS II | P26R021720135 | NONE | NONE | PERFORMED SURFACE CLEANING CHECKED AND VERIFIED ALIGNMENTS PERFORMED WATER RUN PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769457539325151626/5769457539325151626_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5769457539325151626/5769457539325151626_signature_31.png
| 2023-11-28 | 2023-11-24 09:56 | Hi-Precision Diagnostics | 674 N.S. Amoranto Sr. St, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26588 | CONSORTIUM | jose@hi-precision.com.ph | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 40 ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5201 |  | GX IV 6C | 835601 | Module B2 frequent error 2008 | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed functionality test Archived results Generated report | Checked modules Performed maintenance | Unut ia operational Archived results to be checked by specialist |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769456653619524450/5769456653619524450_signature_28.png | Rachel Anne Alhambra | https://www.jotform.com/uploads/macarewecare/222927978372067/5769456653619524450/5769456653619524450_signature_31.png
| 2023-11-28 | 2023-11-24 14:50 | Domingo Santiago Health Center | D. Santiago St., Sampaloc, Manila | NCR | NCR 2nd - Quezon City | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26587 | DOH | fei_paas@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1089 | 2023-11-21 | GX IV 6C | 831527 | Module A4 for replacement | Performed PM Replaced module A4 Generated report | Replaced module A4 716610 -> 644056 | Unit is operational. All modules working |  | Gx Module 6C | G6484 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769453133615985173/5769453133615985173_signature_28.png | Flory Fe A. Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5769453133615985173/5769453133615985173_signature_31.png
| 2023-11-28 | 2023-11-23 11:00 | General Trias Social Hygiene Clinic | Hospital Rd, General Trias, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26586 | MACARE-HBDC | chogentri.shc@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 203 |  | GXIV-D-10C | 110016785 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Filter Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769432763617784171/5769432763617784171_signature_28.png | Maria Christina Camille Saddul | https://www.jotform.com/uploads/macarewecare/222927978372067/5769432763617784171/5769432763617784171_signature_31.png
| 2023-11-28 | 2023-11-23 10:27 | General Trias City Health Office | Hospital Rd., General Trias, Cavite | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26585 | DOH | loremontoya1169@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2987 ASSAY: Xpert HIV-1 Viral Load, COUNT: 175 | 2023-09-19 | GXIV-D | 821340 | Module 1 with Temperature Error Unit is due for XpertCheck | XpertCheck Module Replacement Preventive Maintenance | Replaced Module 1 Performed Maintenance Performed XpertCheck Generate Reports Performed Self-test | 3 Modules are Operational, 1 For Reporting | A2 Failed in XpertCheck | XpertCheck Kit | 10673 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769432283614147679/5769432283614147679_signature_28.png | Lorena R.Montoya | https://www.jotform.com/uploads/macarewecare/222927978372067/5769432283614147679/5769432283614147679_signature_31.png
| 2023-11-28 | 2023-11-21 08:42 | Alabang Health Center | Liwasan ng mga Bayani, Alabang, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26584 | DOH | cathycarasig@yahoo.com | 26 | ASSAY: Xpert HPV, COUNT: 138 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1993 |  | GXIV-D | 831913 | Modules 3 and 4 with Temperature Error | Module Replacement | Performed Maintenance Replaced Modules 3 and 4 Generate Reports | 3 Modules are Operational, 1 Module for Replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769431863614142157/5769431863614142157_signature_28.png | Catherine carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/5769431863614142157/5769431863614142157_signature_31.png
| 2023-11-28 | 2023-11-24 12:47 | General Emilio Aguinaldo Memorial Hospital | Indang-Trece Rd, Trece Martires,Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26581 | PBSP | djaemiel@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1486 ASSAY: Xpert HIV-1 Viral Load, COUNT: 206 |  | GXIV-L | 802860 | Module 2 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 2 Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633611882401/5769430633611882401_signature_28.png | JAEMIEL JUVIRENS A. DIMARANAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633611882401/5769430633611882401_signature_31.png
| 2023-11-28 | 2023-11-23 11:00 | General Trias Social Hygiene Clinic | Hospital Rd, General Trias, Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26582 | MACARE-HBDC | chogentri.shc@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 203 |  | GXIV-D-10C | 110016785 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Filter Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633612958646/5769430633612958646_signature_28.png | Maria Christina Camille Saddul | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633612958646/5769430633612958646_signature_31.png
| 2023-11-28 | 2023-11-21 09:18 | Putatan Health Center | National Road, Putatan, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26583 | PBSP | apmolera@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 123 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1638 ASSAY: Xpert HIV-1 Viral Load, COUNT: 89 |  | GXIV-D | 834690 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Performed Selft-test Backup Tests Generate Reports | 3 Modules Passed, 1 Failed | B3 Failed in XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633614742806/5769430633614742806_signature_28.png | Aleli P. Molera | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633614742806/5769430633614742806_signature_31.png
| 2023-11-28 | 2023-11-23 10:27 | General Trias City Health Office | Hospital Rd., General Trias, Cavite | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26579 | DOH | loremontoya1169@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2987 ASSAY: Xpert HIV-1 Viral Load, COUNT: 175 | 2023-09-19 | GXIV-D | 821340 | Module 1 with Temperature Error Unit is due for XpertCheck | XpertCheck Module Replacement Preventive Maintenance | Replaced Module 1 Performed Maintenance Performed XpertCheck Generate Reports Performed Self-test | 3 Modules are Operational, 1 For Reporting | A2 Failed in XpertCheck | XpertCheck Kit | 10673 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633618646958/5769430633618646958_signature_28.png | Lorena R.Montoya | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633618646958/5769430633618646958_signature_31.png
| 2023-11-28 | 2023-11-21 08:42 | Alabang Health Center | Liwasan ng mga Bayani, Alabang, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26580 | DOH | cathycarasig@yahoo.com | 26 | ASSAY: Xpert HPV, COUNT: 138 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1993 |  | GXIV-D | 831913 | Modules 3 and 4 with Temperature Error | Module Replacement | Performed Maintenance Replaced Modules 3 and 4 Generate Reports | 3 Modules are Operational, 1 Module for Replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633616571045/5769430633616571045_signature_28.png | Catherine carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/5769430633616571045/5769430633616571045_signature_31.png
| 2023-11-28 | 2023-11-23 13:30 | Bantay Municipal Health Office | Bantay, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26578 | PBSP | lornarabara_23@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 569 | 2023-11-23 | GX4 | 807585 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 609462 >>> 863141 - Generated reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 863141 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769414472082164060/5769414472082164060_signature_28.png | Lorna rabara | https://www.jotform.com/uploads/macarewecare/222927978372067/5769414472082164060/5769414472082164060_signature_31.png
| 2023-11-28 | 2023-11-23 08:37 | Dingras District Hospital | Dingras, Ilocos Norte | 1 | Ilocos Norte | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26577 | PBSP | kenacudapas1022@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 619 |  | GX4 | 807597 | Due for Troubleshooting | #NAME? | Performed Troubleshooting and Maintenance | Unit is operational | Modules B1, B3 and B4 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769409542084176149/5769409542084176149_signature_28.png | Kenadine Cudapas | https://www.jotform.com/uploads/macarewecare/222927978372067/5769409542084176149/5769409542084176149_signature_31.png
| 2023-11-28 | 2023-11-24 14:23 | East Avenue Medical Center | East Avenue,QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com | eamc.biomed@gmail.com |  | 26576 | MACARE-HBDC | jocelynrote.1970@gmail.com | 28 |  |  | MCS + | 15H044 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed isolation running | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769408633617502103/5769408633617502103_signature_28.png | Jocelyn Rote | https://www.jotform.com/uploads/macarewecare/222927978372067/5769408633617502103/5769408633617502103_signature_31.png
| 2023-11-28 | 2023-11-21 09:01 | Dagupan City Health Office | Dagupan City | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26575 | DOH | thisisformyfamily@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 1408 |  | GX4 | 838286 | For Troubleshooting | #NAME? | Performed Troubleshooting and Maintenance | Unit is operational | Module B1 and B4 for observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5769406392086540799/5769406392086540799_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5769406392086540799/5769406392086540799_signature_31.png
| 2023-11-27 | 2023-11-25 09:55 | Camalig Rural Health Unit | Camalig, Albay | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | danemilbodino23@gmail.com |  | 26574 | PBSP | danemilbodino23@gmail.com | 20 | ASSAY: , COUNT: 798 |  | GX4 | 835187 | B1/788971 for module replacement | Replaced module: B1-788971>>680782 Cleaned cartridge bays, plunger rod, icore, fan Generate reports | Replaced Module B1 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5768962085318458696/5768962085318458696_signature_28.png | Leilani M. Grageda | https://www.jotform.com/uploads/macarewecare/222927978372067/5768962085318458696/5768962085318458696_signature_31.png
| 2023-11-27 | 2023-11-24 14:45 | Matnog Medicare Community Hospital | Matnog, Sorsogon | 5 | Sorsogon | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | macaretablet.engineering@gmail.com |  | 26573 | PBSP | macaretablet.engineering@gmail.com | 25 |  |  | GX4 | 835200 | CPU not powering on | Check power cord, cpu, power button, cpu and ups ports, RAM. Cleaned internal CPU. | Performed troubleshooting | Unit is not operational. |  |  |  |  | https://www.jotform.com/uploads/data:, | Mernadith Garcera | https://www.jotform.com/uploads/macarewecare/222927978372067/5768951755319072810/5768951755319072810_signature_31.png
| 2023-11-27 | 2023-11-24 12:20 | Bulan Rural Health Unit | Bulan, Sorsogon | 5 | Sorsogon | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | danemilbodino23@gmail.com |  | 26572 | PBSP | gemmaalbello1970@g.mail.com | 23 | ASSAY: , COUNT: 2574 |  | GX4 | 807983 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5768944355314939727/5768944355314939727_signature_28.png | Gemma G. Albello | https://www.jotform.com/uploads/macarewecare/222927978372067/5768944355314939727/5768944355314939727_signature_31.png
| 2023-11-27 | 2023-11-23 11:58 | San Jacinto Rural Health Unit | San Jacinto, Masbate | 5 | Masbate | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | danemilbodino23@gmail.com |  | 26571 | PBSP | danemilbodino23@gmail.com | 28 | ASSAY: , COUNT: 1285 |  | GX4 | 812793 | For Anual Calibration | XpertCheck Result: All modules passed Cleaned cartridge bays, plunger rod, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/data:, | Janet A. Bragais | https://www.jotform.com/uploads/macarewecare/222927978372067/5768933075312741640/5768933075312741640_signature_31.png
| 2023-11-27 | 2023-11-22 15:36 | Bulusan RHU | Bulusan, Sorsogon | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | danemilbodino23@gmail.com |  | 26570 | PBSP | danemilbodino23@gmail.com | 23 | ASSAY: , COUNT: 718 |  | GX4 | 835197 | B2/788991 for Module Replacement | Replaced Module: B2-788991>>621480 Cleaned cartridge bays, plunger rod, icore, fan Generate reports | Replaced Module B2 Performed Preventive Maintenance | Unjt is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5768923405316557890/5768923405316557890_signature_28.png | Maria Arlene G. Ragasa | https://www.jotform.com/uploads/macarewecare/222927978372067/5768923405316557890/5768923405316557890_signature_31.png
| 2023-11-24 | 2023-11-24 11:00 | PASSI CITY HEALTH OFFICE | PASSI CITY , ILOILO | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26569 | PBSP | passicholab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4421 ASSAY: Xpert HIV-1 Viral Load, COUNT: 264 |  | GX-IV | 807610 | FOR MODULE A1 AND MODULE A4 REPLACEMENT | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE REPLACEMENT: A1 DEF SN:210068726 WITH NEW SN:892448 A4 DEF SN:684889 WITH NEW SN:862309 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766437592275827487/5766437592275827487_signature_28.png | John Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5766437592275827487/5766437592275827487_signature_31.png
| 2023-11-24 | 2023-11-24 16:00 | Antipolo City Health Office | M. Santos St. | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26568 | PBSP | acho.laboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 |  | GX-XVI-L-6C | 842226 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Run XpertCheck Generated reports | Unit is operational. Module A2 not calibrated, module does not appear in XpertCheck. Module C1, D4 and B4 service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766244726222353849/5766244726222353849_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5766244726222353849/5766244726222353849_signature_31.png
| 2023-11-24 | 2023-11-24 14:54 | BAGO CITY HEALTH OFFICE | Bago City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  | engineering@macare-medicals.com |  |  | 26567 | PBSP | swabteambago@gmail.com | 25 |  |  | GX-4 | 815433 | None | Module bay cleaning, optics cleaning, plunger rod maintenance. | Installation of backup Genexpert unit. | 4 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766184264566810615/5766184264566810615_signature_28.png | GLENLY C. CASONA | https://www.jotform.com/uploads/macarewecare/222927978372067/5766184264566810615/5766184264566810615_signature_31.png
| 2023-11-24 | 2023-11-24 14:48 | BAGO CITY HEALTH OFFICE | Bago City, Negros Occidental | 6 | Negros Occidental | TB LABORATORY | NO | REPAIR ON-SITE ASSISTANCE | CEPHEID | Emerson Bejemino |  | engineering@macare-medicals.com |  |  | 26566 | DOH | swabteambago@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2482 |  | GX-4 | 838486 | Defective Module B4/862207 | Module bay cleaning, optics cleaning, plunger rod maintenance. | Replace module B4 with new module with SN 671816. | 2 modules operational, B2 and B3 waiting for availability of replacement modules. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766180914565108621/5766180914565108621_signature_28.png | Glenly Casona | https://www.jotform.com/uploads/macarewecare/222927978372067/5766180914565108621/5766180914565108621_signature_31.png
| 2023-11-24 | 2023-11-24 14:48 | BAGO CITY HEALTH OFFICE | Bago City, Negros Occidental | 6 | Negros Occidental | TB LABORATORY | NO | REPAIR ON-SITE ASSISTANCE | CEPHEID | Emerson Bejemino |  | engineering@macare-medicals.com |  |  | 26565 | DOH | swabteambago@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2482 |  | GX-4 | 838486 | Defective Module B4/862207 | Module bay cleaning, optics cleaning, plunger rod maintenance. | Replace module B4 with new module with SN 671816. | 2 modules operational, B2 and B3 waiting for availability of replacement modules. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766180564562901081/5766180564562901081_signature_28.png | Glenly Casona | https://www.jotform.com/uploads/macarewecare/222927978372067/5766180564562901081/5766180564562901081_signature_31.png
| 2023-11-24 | 2023-11-24 13:45 | Bagong Barrio Health Center | Bagong Barrio, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26564 | DOH | cruzernitamay@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1231 |  | GXIV | 831906 | Modules A1, A3 for re XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. A1: service required A3: service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766167132619300556/5766167132619300556_signature_28.png | Ernita May Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5766167132619300556/5766167132619300556_signature_31.png
| 2023-11-24 | 2023-11-24 13:59 | Don Juan C. Cojuwngco Memorial Health Center | Cuyapo Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26563 | DOH | jolenerodico09171998@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1112 |  | GXIV-D | 831672 | HW Failed on module A3 | Xoertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is partially operational | A3 failed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766150662212475336/5766150662212475336_signature_28.png | John M. Esteban | https://www.jotform.com/uploads/macarewecare/222927978372067/5766150662212475336/5766150662212475336_signature_31.png
| 2023-11-24 | 2023-11-24 13:38 | Balingasag Rural Health Unit | Brgy. Waterfall, Balingasag | 10 | Misamis Oriental | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26562 | PBSP | joelquino79@gmail.com | 28 |  |  | GX-4 | 815382 | None | None | Replacing Module B1 with SN: 615094 to 696057 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766148594178112692/5766148594178112692_signature_28.png | Joel Quiño | https://www.jotform.com/uploads/macarewecare/222927978372067/5766148594178112692/5766148594178112692_signature_31.png
| 2023-11-24 | 2023-11-24 12:32 | BACOLOD CITY HEALTH OFFICE | Bacolod City, Negros Occidental | 6 | Negros Occidental | TB LABORATORY | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  | engineering@macare-medicals.com |  |  | 26561 | PBSP | cyrilcordova0515@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6537 |  | GX4 | 804267 | Defective module E3/634474. | Module bay cleaning, optics cleaning, plunger rod maintenance. | Replaced module E3 with new module with SN 754045. | 4 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766099664129380333/5766099664129380333_signature_28.png | CYRIL CORDOVA, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5766099664129380333/5766099664129380333_signature_31.png
| 2023-11-24 | 2023-11-24 11:33 | Paniqui Super Health Center | Basio Poblacion Norte, Paniqui Tarlac | 3 | Tarlac | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26560 | PBSP | forylab76@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2023-11-23 | GXIV-D | 812792 | CPU and Monitor for replacement | Installation and setup Preventive maintenance | Performed CPU and Monitor Replace,ent Installation and user setup Gathered reports Performed functionality test | Unit is operational |  | CPU and Moniyor | 45561 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766072192216253915/5766072192216253915_signature_28.png | Fortune Abutin | https://www.jotform.com/uploads/macarewecare/222927978372067/5766072192216253915/5766072192216253915_signature_31.png
| 2023-11-24 | 2023-11-24 11:36 | Pascual Laboratories | McArthur Highway, Balagtas | 3 | Bulacan | QC | YES | CALIBRATION MAINTENANCE | CARL ZEISS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | qcventura@pascuallab.com |  | 26559 | COMMERCIAL | mbjimenez@pascuallab.com | 24 |  | 2023-11-21 | Axioscope A.1 | 3321001650 | For Calibration | Replaced bulb | Performed maintenance Replaced bulb Cleaned microscope body Cleaned lenses Cleaned filter Checked camera Performed Kohler Illumination Functionality test | Unit is operational. |  | Bulb 12V 50W Halogen Rectangular Filament | DR# 10831, PO# 7310003931 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766070101591689749/5766070101591689749_signature_28.png | Maureen Jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/5766070101591689749/5766070101591689749_signature_31.png
| 2023-11-24 | 2023-11-23 14:00 | Taytay Rural Health Unit 5B | Taytay, Arenda | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26558 | PBSP | zenresurreccion@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1410 |  | GX-IV-D-6C | 838136 | Module B1 has error 1002 | Checked module B1 | Performed maintenance Checked module B1 Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. Module B1 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766061931594515490/5766061931594515490_signature_28.png | Zenilynn Bobiles | https://www.jotform.com/uploads/macarewecare/222927978372067/5766061931594515490/5766061931594515490_signature_31.png
| 2023-11-24 | 2023-11-23 15:20 | Philippine Red Cross Q.C | Malakas st QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26557 | MACARE-HBDC | e.villamar@yahoo.com | 28 |  |  | MCS + | O2A227 | Due for maintenance | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed isolation running | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5766003483611051850/5766003483611051850_signature_28.png | Evangeline Villamar | https://www.jotform.com/uploads/macarewecare/222927978372067/5766003483611051850/5766003483611051850_signature_31.png
| 2023-11-23 | 2023-11-23 11:50 | Polomolok Rural Health Unit | Poblacion, Polomolok, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26556 | DOH | jmmirasolrmt@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2260 |  | Gx-IV | 838265 | N/A | Perform Genexpert machine Preventive maintenance | Genexpert machine Cleaning Data collection, IQR, S.Log, Assay statistics report, and archive test | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765318129128984550/5765318129128984550_signature_28.png | June Michele mirasol | https://www.jotform.com/uploads/macarewecare/222927978372067/5765318129128984550/5765318129128984550_signature_31.png
| 2023-11-23 | 2023-11-23 14:06 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26555 | DOH | vsmmcdeptpatho.genexpert@gmail.com | 32 | ASSAY: Xpert HPV, COUNT: 193 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2998 ASSAY: Xpert HIV-1 Viral Load, COUNT: 152 ASSAY: Xpert HPV, COUNT: 193 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2998 ASSAY: Xpert HIV-1 Viral Load, COUNT: 152 | 2023-11-23 | GXIV | 838285 | Module A4 frequent temp error | None | #NAME? | Module A4 disabled. Other modules are functional. | Module A4 are within calibrated parameters. Module temperature is too high (ranging from 32C-33C) | XpertCheck | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765304629911588809/5765304629911588809_signature_28.png | Myra Riccil Estose | https://www.jotform.com/uploads/macarewecare/222927978372067/5765304629911588809/5765304629911588809_signature_31.png
| 2023-11-23 | 2023-11-23 14:14 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com | laguitannataniel@gmail.com |  | 26554 | DOH | delacerna.sheryl22@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 848 | 2023-11-23 | GXIV | 831845 | Module B3 (786592) for replacement | #NAME? | Replaced B3 | Unit is operational. |  | GeneXpert Module SN: 644074 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765295342617523734/5765295342617523734_signature_28.png | Sheryl Dela Cerna | https://www.jotform.com/uploads/macarewecare/222927978372067/5765295342617523734/5765295342617523734_signature_31.png
| 2023-11-23 | 2023-11-23 13:23 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26553 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 195 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3160 ASSAY: Xpert HIV-1 Viral Load, COUNT: 133 ASSAY: Xpert HPV, COUNT: 195 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3160 ASSAY: Xpert HIV-1 Viral Load, COUNT: 133 |  | GXIV | 820206 | None | None | #NAME? | Module A2 disabled, waiting for module replacement. Other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765265019917648161/5765265019917648161_signature_28.png | Mariel Monilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5765265019917648161/5765265019917648161_signature_31.png
| 2023-11-23 | 2023-11-23 12:56 | Compostela Rural Health Unit | Poblacion, Compostela | 11 | Davao de Oro | TB-DOTS | YES | REPAIR TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26552 | PBSP | nielmaquiling20@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1348 |  | GX-IV 4D | 807360 | H/W Failure at Modules B1 and B4 UPS has no power at Battery Mode | Replaced the Defective Module SN :715684 with New Module SN:807990 Replaced the Defective Module SN:676635 with new Module SN:719912 Troubleshoot the UPS | Module Replacement UPS Troubleshooting | All Modules are operational UPS Battery Mode is Defective, Pictures and Videos are documented to be sent at Cepheid for Further Assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765251757221513612/5765251757221513612_signature_28.png | GIRNALYN D. ECONAR | https://www.jotform.com/uploads/macarewecare/222927978372067/5765251757221513612/5765251757221513612_signature_31.png
| 2023-11-23 | 2023-11-23 12:10 | Dipolog City Health Office | Dipolog City, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 26551 | PBSP | jz.ebron@yahoo.com.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-11-23 | GX 4 | 834826 | Module 2 excessive error 1002 | Evaluation of temperature errors, module 2 undergone some test run regarding error 1002, Preventive maintenence. | Evaluation of temperature errors, module 2 undergone some test run regarding error 1002 | Module 2 give excessive error of 1002. Module 1 is available | Module 2 need to be replaced | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765222178112556829/5765222178112556829_signature_28.png | James Zebedee C. Ebron | https://www.jotform.com/uploads/macarewecare/222927978372067/5765222178112556829/5765222178112556829_signature_31.png
| 2023-11-23 | 2023-11-22 15:49 | Commonwealth Health Center | 8 Katuparan St., Commonwealth, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26550 | DOH | commonwealthclinical@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2827 ASSAY: Xpert HPV, COUNT: 107 | 2023-11-21 | GX IV 6C | 831615 | Module A2 for replacement | Performed PM Replaced Module Generated report | Replaced Module A2 754436 -> 643802 | Unit is operational All modules working |  | GX Module 6C | G 6485 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765202476904584679/5765202476904584679_signature_28.png | Rizza Clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5765202476904584679/5765202476904584679_signature_31.png
| 2023-11-23 | 2023-11-23 11:14 | Kadingilan Rural Health Unit | Poblacion, Kadingilan | 10 | Bukidnon | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26549 | PBSP | graycey_mikay@yahoo.com | 26.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 217 |  | GX-2 | 835194 | Reported Defective Module B2 | None | Replacing Module B2 with SN: 731556 to SN: 696056 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765200868114543836/5765200868114543836_signature_28.png | Grace B. Remando | https://www.jotform.com/uploads/macarewecare/222927978372067/5765200868114543836/5765200868114543836_signature_31.png
| 2023-11-23 | 2023-11-23 11:05 | Marikina City Health Office | Shoe Ave. | NCR | NCR 2nd - Marikina | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26548 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1058 ASSAY: Xpert HIV-1 Viral Load, COUNT: 174 |  | GX-IV-D-6C | 805211 | Modules lost communication when running | Checked connections | Checked connections Checked LAN cable Checked network switch Run reliability test Run actual sample Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765191412694200263/5765191412694200263_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5765191412694200263/5765191412694200263_signature_31.png
| 2023-11-23 | 2023-11-22 11:49 | Nagpayong Super Health Center | Nagpayong | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26547 | DOH | michelindg23@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 925 |  | GX-IV-D-6C | 83195 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765171052694774066/5765171052694774066_signature_28.png | Michelle De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5765171052694774066/5765171052694774066_signature_31.png
| 2023-11-23 | 2023-11-23 09:54 | Sta. Lucia Health Center | Sta. Lucia Ave., Sta. Lucia, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26546 | DOH | josiedingle2004@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1877 | 2023-11-21 | GX IV 6C | 831901 | Module A4 for replacement | Performed PM Replaced Module Generated report | Replaced Module A4 643103 -> 702936 | unit is operational All modules working |  | GX module 6C | G6454 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765150417377959235/5765150417377959235_signature_28.png | Josephine L. Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5765150417377959235/5765150417377959235_signature_31.png
| 2023-11-23 | 2023-11-22 15:21 | Payatas B HC | Bulacan st,QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26545 | DOH | racquelmrocero@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2277 |  | GX4 | 831899 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765147123615950567/5765147123615950567_signature_28.png | Racquel M. Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5765147123615950567/5765147123615950567_signature_31.png
| 2023-11-23 | 2023-11-22 11:15 | Best Diagnostic | Matalino st,QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26544 | COMMERCIAL | valbuenashy@gmail.com | 29 | ASSAY: Xpert CT/NG, COUNT: 115 ASSAY: Xpert HBV VL, COUNT: 90 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GX4 | 110008529 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765146723617688985/5765146723617688985_signature_28.png | Best Diagnostic Corporation | https://www.jotform.com/uploads/macarewecare/222927978372067/5765146723617688985/5765146723617688985_signature_31.png
| 2023-11-23 | 2023-11-23 09:30 | AJUY RURAL HEALTH UNIT | Ajuy, Iloilo | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26543 | PBSP | ajuyrhu_11@yahoo.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1148 |  | GX IV | 815429 | FOR MODULE B4 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE B4 REPLACEMENT DEF SN: 695906 WITH NEW SN:702385 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765125415061455817/5765125415061455817_signature_28.png | Barbara Kylie Balleza | https://www.jotform.com/uploads/macarewecare/222927978372067/5765125415061455817/5765125415061455817_signature_31.png
| 2023-11-23 | 2023-11-23 09:02 | Polanco Rural Health Unit | Polanco, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | YES | MAINTENANCE Modules Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 26542 | PBSP | polanco.rhu@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 713 | 2023-11-23 | GX 4 | 838539 | None | Replaced modules 600015 and 862129 with modules 802387 and 210004150 respectively. | Preventive Maintenence and Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765112198116971749/5765112198116971749_signature_28.png | Winielyn F. Tapal | https://www.jotform.com/uploads/macarewecare/222927978372067/5765112198116971749/5765112198116971749_signature_31.png
| 2023-11-23 | 2023-11-23 08:44 | Bongabon District Hospital | Curva, Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26541 | DOH | tsantuajr0128@gmail.com | 25 |  | 2023-11-23 | Genexpert IV | 831562 | Module 1 error | Performed Xpertcheck on Module 1 = Passed | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 1 for observation | Xpertcheck 2pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765105244109444456/5765105244109444456_signature_28.png | Teodorico Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5765105244109444456/5765105244109444456_signature_31.png
| 2023-11-23 | 2023-11-23 08:08 | Macare In - House | 1005 Aurora Grand,Aurora blvd. | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26540 | COMMERCIAL | htoco@macare-medicals.com | 28 |  |  | HPZ220 | 2UA3520VNV | Unable to boot | Replaced SSD Installed GX DX Installed Operating System Performed functionality test | Repair CPU | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5765076293613484590/5765076293613484590_signature_28.png | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5765076293613484590/5765076293613484590_signature_31.png
| 2023-11-22 | 2023-11-22 16:48 | Macare In-House | Aurora Grand, QC | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Aspex | 26539 | COMMERCIAL | carlos.leniel.origenes@gmai.com | 25 |  |  | AT1000 | AT1000181128ST0151 | For checkup | None | Checked fuse Checked batteries Checked switch | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764523453612097414/5764523453612097414_signature_28.png | Carlos | https://www.jotform.com/uploads/macarewecare/222927978372067/5764523453612097414/5764523453612097414_signature_31.png
| 2023-11-22 | 2023-11-22 16:45 | Macare In-House | Aurora Grand, QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR TROUBLESHOOTING | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Aspex | 26538 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  |  | AT1000 | AT1000180910ST0568 | Doesnt turn on | Replaced fuse | Checked fuse | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764520363612621563/5764520363612621563_signature_28.png | Carlos | https://www.jotform.com/uploads/macarewecare/222927978372067/5764520363612621563/5764520363612621563_signature_31.png
| 2023-11-22 | 2023-11-22 16:42 | Macare In-House | Aurora Grand, QC | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Aspex | 26537 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  |  | AT1000 | AT1000180910ST0429 | Doesnt turn on | None | Checked fuse Checked battery Checked switch | Condemn unit. Cannot be repaired |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764518783618304196/5764518783618304196_signature_28.png | Carlos | https://www.jotform.com/uploads/macarewecare/222927978372067/5764518783618304196/5764518783618304196_signature_31.png
| 2023-11-22 | 2023-11-22 16:39 | Macare In-House | Aurora Grand, QC | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  | Aspex | 26536 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 25 |  |  | AT1000 | AT1000181128ST0219 | Doesnt turn on | None | Checked fuse Checked switch Checked batteries | Condemn unit. Cannot be repaired |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764517133618097905/5764517133618097905_signature_28.png | Carlos | https://www.jotform.com/uploads/macarewecare/222927978372067/5764517133618097905/5764517133618097905_signature_31.png
| 2023-11-22 | 2023-11-22 14:24 | Gov. Celestino Gallares Memorial Hospital | Tagbilaran, Bohol | 7 | Bohol | Laboratory | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26535 | PBSP | majieocmen@gmail.com | 28 |  |  | Gx 4 | 838878 | Stucked up cartridge | Removing Stucked up carthridge/switch hub connection 2 genexpert unit | Removing Stucked up carthridge/switch hub connection 2 genexpert unit | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764471407517037130/5764471407517037130_signature_28.png | Ma. Jernai Sosoban | https://www.jotform.com/uploads/macarewecare/222927978372067/5764471407517037130/5764471407517037130_signature_31.png
| 2023-11-22 | 2023-11-22 14:37 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando,Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26534 | MACARE-HBDC | jblmgh.is@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 239 ASSAY: Xpert HIV-1 Viral Load, COUNT: 573 ASSAY: Xpert HBV VL, COUNT: 88 ASSAY: Xpert HCV VL, COUNT: 22 | 2023-11-22 | Genexpert IV | 812571 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764451434219308061/5764451434219308061_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5764451434219308061/5764451434219308061_signature_31.png
| 2023-11-22 | 2023-11-22 13:23 | Mexico Rural Health Unit I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26533 | DOH | jasperdizon20@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 | 2023-11-22 | Genexpert IV | 831877 | Monitor no display on and off | Troubleshoot the wirings | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Monitor for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764403201577462846/5764403201577462846_signature_28.png | Jasper Dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5764403201577462846/5764403201577462846_signature_31.png
| 2023-11-22 | 2023-11-21 12:43 | Bulakan Rural Health Unit I | Bulakan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26532 | DOH | l_gemino@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 958 |  | GXIV | 831641 | Module A4 for reXpertCheck | #NAME? | Peformed XpertCheck on A4 | Unit is operational. A4 passed XpertCheck. A4 is under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764367088513957871/5764367088513957871_signature_28.png | Lilibeth Gemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5764367088513957871/5764367088513957871_signature_31.png
| 2023-11-22 | 2023-11-22 09:50 | St. Luke's Medical Center - Global City | Rizal Dr. | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 26531 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 |  |  | GX-IV-L-6C | 808411 | Module B1 has error 5007/2014 | Performed maintenance | Performed maintenance Checked module B1 Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. Module B1 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764344957571802395/5764344957571802395_signature_28.png | Stephanie Infante | https://www.jotform.com/uploads/macarewecare/222927978372067/5764344957571802395/5764344957571802395_signature_31.png
| 2023-11-22 | 2023-11-22 09:53 | St. Luke's Medical Center - Global City | Rizal Dr. | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomedreports@stlukes.com.ph |  | 26530 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 |  |  | GX-IV-L-6C | 839231 | Module C1 has error 4009 | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Checked module C1 Generated reports | Unit is operational. Module C1 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764344957576613136/5764344957576613136_signature_28.png | Stephanie Infante | https://www.jotform.com/uploads/macarewecare/222927978372067/5764344957576613136/5764344957576613136_signature_31.png
| 2023-11-22 | 2023-11-22 11:01 | Aurora rural health unit | Juan luna Street, poblacion aurora, zamboanaga del sur | 9 | Zamboanga del Sur | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | joann8850@yahoo.com |  | 26529 | PBSP | joann8850@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 304 | 2033-11-22 | GX 4 | 834435 | Module 2 h/w failed, Uninterrupted power supply malfunction | Module 2 undergone Xpertcheck | Module 2 undergone calibration, UPS battery in fault. | Module 2 needs to be replaced, and UPS needs to be replaced. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764317870359207890/5764317870359207890_signature_28.png | Jo Ann B. Duran | https://www.jotform.com/uploads/macarewecare/222927978372067/5764317870359207890/5764317870359207890_signature_31.png
| 2023-11-22 | 2023-11-22 09:05 | Klinika Novaliches Sundown Clinic | 2nd flr. Bautista Building, 8 Doña Isaura, Novaliches, QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com | klinikanovaliches2022@gmail.com |  | 26528 | MACARE-HBDC | laiza.lanuza@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GX IV 10C | 120000140 | GX and UPS needs socket adaptor | Performed PM Added socket adaptor and fixed cable Generated report | Added socket adaptor | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5764300977519889311/5764300977519889311_signature_28.png | ma. laiza mae lanuza | https://www.jotform.com/uploads/macarewecare/222927978372067/5764300977519889311/5764300977519889311_signature_31.png
| 2023-11-21 | 2023-11-21 16:09 | Calinan Health Center | Poblacion, Calinan Davao City | 11 | Davao del Sur | TB-DOTS | YES | REPAIR TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26527 | PBSP | mailfegelacio@gmail.com | 27.5 |  |  | GX-IV 4D | 815390 | Error at System Unit Erratic Lcd Display Fan Filter is damaged | Replace the defective System Unit SN:2UA7021M0R with New System Unit SN: CPH0273F43756 Replace the VGA chord, replace the vga to DP converter into VGA to HDMI | System Unit Replacement Troubleshoot monitor | After troubleshooting The Monitor, Erratic display still persist, Documented Videos and Pictures to be sent at Cepheid for further assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763725900715224387/5763725900715224387_signature_28.png | Mail Fe J. Gelacio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5763725900715224387/5763725900715224387_signature_31.png
| 2023-11-21 | 2023-11-21 15:31 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26526 | DOH | elma.bragat@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5037 ASSAY: Xpert HIV-1 Viral Load, COUNT: 133 |  | GX-IV | 838508 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763612352428401258/5763612352428401258_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5763612352428401258/5763612352428401258_signature_31.png
| 2023-11-21 | 2023-11-21 14:58 | Macare In-House | 1005 Aurora Blvd. Proj. 3 | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 26525 | MACARE-HBDC | htoco@macare-medicals.com | 24 |  |  | HP Z220 | 2UA3520VTR | Unable to boot | Replaced SSD Installed GX DX Installed Operating System Functionality test | Repair CPU | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763595003615898527/5763595003615898527_signature_28.png | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5763595003615898527/5763595003615898527_signature_31.png
| 2023-11-21 | 2023-11-21 14:55 | Macare In-House | 1005 Aurora Blvd. Proj. 3 | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 26524 | MACARE-HBDC | htoco@macare-medicals.com | 24 |  |  | HP Z240 | 2UA6511RLL | Unable to boot | Replaced SSD Installed GX DX Installed Operating System Functionality test | Repair CPU | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763591113619055711/5763591113619055711_signature_28.png | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5763591113619055711/5763591113619055711_signature_31.png
| 2023-11-21 | 2023-11-21 14:50 | Macare In-House | 1005 Auroa Blvd. Proj. 3 | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | CEPHEID | Noel Floresca |  | nfloresca@macare-medicals.com |  |  | 26523 | MACARE-HBDC | htoco@macare-medicals.com | 24 |  |  | HP Z220 | 2UA3520VNV | Unable to boot | Replaced SSD Installed GX DX Installed operating system Functionality test | Repair CPU | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763588703614066428/5763588703614066428_signature_28.png | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5763588703614066428/5763588703614066428_signature_31.png
| 2023-11-21 | 2023-11-21 14:27 | Minalin Rural Health Unit | Minalin Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26522 | PBSP | bernadetteann0318@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 510 |  | GXIV-D | 834688 | CPU hard startup | PC Troubleshooting Preventive maintenance | Performed CPU troubleshooting Performed preventive maintenance Gathered reports Performed functionality test | Unit is operational | Under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763576484132128330/5763576484132128330_signature_28.png | Lady Immaculate Magat | https://www.jotform.com/uploads/macarewecare/222927978372067/5763576484132128330/5763576484132128330_signature_31.png
| 2023-11-21 | 2023-11-21 12:29 | Lubao Rural Health Unit | Lubao Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26521 | PBSP | nathaliecoleencovar@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 918 |  | GXIV-D | 812834 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is fully operational | All modules passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763511314139340046/5763511314139340046_signature_28.png | Nathalie Coleen Covar | https://www.jotform.com/uploads/macarewecare/222927978372067/5763511314139340046/5763511314139340046_signature_31.png
| 2023-11-21 | 2023-11-21 10:33 | Pilar Rural Health Unit | Brgy. Pilaring, Pilar | 13 | Surigao del Norte | RTDL | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26520 | PBSP | DexterJohnTDotillos@gmail.com | 30 |  |  | GX-2 | 834833 | UPS unable to sustain in battery mode. Automatically shutdown after power interruption. | None | Troubleshooting | All modules are functional. CPU and Desktop are functional except for UPS. | UPS must be replaced with recommended 1KVA POWERCOM Brand Raptor Series. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763438001349408719/5763438001349408719_signature_28.png | Dexter John T. Dotillos | https://www.jotform.com/uploads/macarewecare/222927978372067/5763438001349408719/5763438001349408719_signature_31.png
| 2023-11-21 | 2023-11-21 10:17 | Pilar Rural Health Unit | National Rd. Panilao Pilar Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 26519 | DOH | aimna_24@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 617 |  | GXIV-D | 831613 | A2 for re-xpertcheck due to temperature error | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | A2 passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763427764137329655/5763427764137329655_signature_28.png | Marciana Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5763427764137329655/5763427764137329655_signature_31.png
| 2023-11-21 | 2023-11-21 08:40 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | Tb-dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | johnpaullabutap@gmail.com |  | 26518 | DOH | lornsabaigar@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 705 | 2023-11-21 | GXIV | 838409 | N/a | N/a | #NAME? | Unit is Operational - Module B1 disabled (waiting for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5763367109229770841/5763367109229770841_signature_28.png | Lorna C. Abaigqr | https://www.jotform.com/uploads/macarewecare/222927978372067/5763367109229770841/5763367109229770841_signature_31.png
| 2023-11-20 | 2023-11-20 11:36 | Eastern Samar Provincial Hospital | Borongan City | 8 | Eastern Samar | Tb-dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com |  |  | 26517 | PBSP | lermafernandez@yahoo.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1226 ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 | 2023-11-20 | GXIV | 804256 | Module B2/668765 defective (for replacement) | #NAME? | - Install new module on B2 w/ SN: 655081 - Performed Functionality test - Performed Preventive maintenance - Generate reports - Pullout defective module SN: 668765 | #NAME? |  | Genexpert Module 6 Color SN/655081 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762739796112758881/5762739796112758881_signature_28.png | Lerma Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5762739796112758881/5762739796112758881_signature_31.png
| 2023-11-20 | 2023-11-16 11:48 | Alabang Health Center | Liwasan ng Bayani, Alabang, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26516 | DOH | cathycarasig@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 138 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2000 |  | GXIV-D | 831913 | UPS is Defective | Repair | Performed Maintenance Replaced UPS Battery Generate Reports | 1 Module is Operational, 3 Modules are Defective |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762678483613665092/5762678483613665092_signature_28.png | Catherine Carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/5762678483613665092/5762678483613665092_signature_31.png
| 2023-11-20 | 2023-11-16 10:07 | Asian Hospital and Medical Center | Corporate Ave, Filinvest, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26515 | COMMERCIAL | molecularpathology@asianhospital.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 300 ASSAY: Xpert HIV-1 Viral Load, COUNT: 5 ASSAY: Xpert HPV, COUNT: 61 ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Others, COUNT: 16 |  | GXIV-D | 110014996 | Unit is due for Maintenance | Preventive Maintenance | Cleaned Cartridge Bay, Plunger Rod and Filter Performed Self-test Checked UPS Function Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762677173613314480/5762677173613314480_signature_28.png | Daphnie Barrato Sarmiento | https://www.jotform.com/uploads/macarewecare/222927978372067/5762677173613314480/5762677173613314480_signature_31.png
| 2023-11-20 | 2023-11-15 11:29 | Dasmariñas City Health Office II | DBB-E, Dasmariñas, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26514 | DOH | hazelmilan0777@gmail.com | 27 |  | 2023-11-06 | GXIV-D | 838535 | CPU Problem | Laptop Replacement | Pulled-out Defective CPU Installed Back-up Unit Performed Maintenance Generate Reports | All Modules are Operational |  | GX Laptop | 10796 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762674613611515039/5762674613611515039_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5762674613611515039/5762674613611515039_signature_31.png
| 2023-11-20 | 2023-11-15 09:13 | Bacoor Rural Health Unit (Panapaan Health Center) | Floraville Subd., Bacoor, Cavite | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26513 | PBSP | tobiasmariacecila@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2748 ASSAY: Xpert HIV-1 Viral Load, COUNT: 191 |  | GXIV-D | 807584 | Module 4 with Temperature Error Unit is due for XpertCheck | Module Replacement XpertCheck | Performed Maintenance Replaced Module 4 Performed XpertCheck Performed Selft-Test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762672813617293397/5762672813617293397_signature_28.png | Maria Cecila V. Tobias | https://www.jotform.com/uploads/macarewecare/222927978372067/5762672813617293397/5762672813617293397_signature_31.png
| 2023-11-20 | 2023-11-14 10:43 | San Nicolas Health Center | San Nicolas St, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26512 | DOH | guadex7@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1520 |  | GXIV-D | 831618 | UPS is Defective | Repair | Replaced UPS Batteries Performed Maintenance Performed Self-test Generate Reports | All Modules are Available |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762671643619908332/5762671643619908332_signature_28.png | Guadalupe L. Cube | https://www.jotform.com/uploads/macarewecare/222927978372067/5762671643619908332/5762671643619908332_signature_31.png
| 2023-11-20 | 2023-11-14 12:16 | San Lazaro Hospital | Quiricada st., Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26511 | PBSP | smportugues.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1415 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1177 |  | GXIV-D | 804766 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Performed Self-test Generate Reports | 3 Modules are Operational | Module B2 failed in XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762669043619559903/5762669043619559903_signature_28.png | Sheradyne M. Portugues, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5762669043619559903/5762669043619559903_signature_31.png
| 2023-11-20 | 2023-11-14 08:40 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26510 | DOH | eeps101487sambolledo@yahoo.com | 28 | ASSAY: Xpert HPV, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 173 ASSAY: Xpert SARS-CoV-2, COUNT: 110 |  | GXIV-D | 831530 | Unit is due for Maintenance | Preventive Maintenance | Performed Maintenance Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762667953613209919/5762667953613209919_signature_28.png | Ervin Emmett Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/5762667953613209919/5762667953613209919_signature_31.png
| 2023-11-20 | 2023-11-14 08:37 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26509 | MACARE-HBDC | eeps101487sambolledo@yahoo.com | 28 | ASSAY: Xpert HPV, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 280 ASSAY: Xpert SARS-CoV-2, COUNT: 331 ASSAY: Xpert HPV, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 280 ASSAY: Xpert SARS-CoV-2, COUNT: 331 |  | GXIV-D | 110014099 | Unit is due for Maintenance | Preventive Maintenance | Performed Maintenance Performed Self-test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762666623613333342/5762666623613333342_signature_28.png | Ervin Emmett Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/5762666623613333342/5762666623613333342_signature_31.png
| 2023-11-20 | 2023-11-14 08:33 | Jose Reyes Memorial Medical Center | Rizal Ave, Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26508 | MACARE-HBDC | eeps101487sambolledo@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 414 ASSAY: Xpert SARS-CoV-2, COUNT: 4380 ASSAY: Xpert HPV, COUNT: 414 ASSAY: Xpert SARS-CoV-2, COUNT: 4380 |  | GX16-D | 110013856 | Unit is due for Maintenance | Preventive Maintenance | Performed Maintenance Purged Test Results Performed Self-test Generate Reports | 14 Modules are Operational | Modules A2 and A4 for Reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762665243614615105/5762665243614615105_signature_28.png | Ervin Emmett Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/5762665243614615105/5762665243614615105_signature_31.png
| 2023-11-20 | 2023-11-17 11:12 | Lian Rural Health Unit | Lian-Calatagan Rd, Lian, Batangas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  | bagbaguenbrenguiler@gmail.com |  |  | 26507 | PBSP | malcrmt2009@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 314 |  | GXIV-D | 835179 | Unit is due for XpertCheck UPS Problem | XpertCheck | Performed Maintenance Performed XpertCheck Checked UPS Function Checked UPS Output Voltage Generate Reports | UPS is Defective. All Modules are Operational | UPS No Back-up Function. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762663633612607058/5762663633612607058_signature_28.png | Mary Anne C. Sison | https://www.jotform.com/uploads/macarewecare/222927978372067/5762663633612607058/5762663633612607058_signature_31.png
| 2023-11-20 | 2023-11-16 09:53 | Dr. Jose Rizal Memorial Hospital | Dipolog, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 26506 | PBSP | lovinoantonieta@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1396 | 2023-11-16 | GX 4 | 804698 | None | Replaced module 723386 with Module 888548. Preventive Maintenance | Replaced module 723386 with Module 888548. Preventive Maintenance | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762617491793558640/5762617491793558640_signature_28.png | Antonieta B. Lovino | https://www.jotform.com/uploads/macarewecare/222927978372067/5762617491793558640/5762617491793558640_signature_31.png
| 2023-11-20 | 2023-11-15 11:28 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26505 | PBSP | dmgonzales.pbsp@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 349 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1779 | 2023-11-15 | GX4 6C | 805212 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 801381 >>> 657281 - Generated reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 657281 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762539147067504694/5762539147067504694_signature_28.png | Donnel M. Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5762539147067504694/5762539147067504694_signature_31.png
| 2023-11-20 | 2023-11-15 11:23 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26504 | PBSP | dmgonzales.pbsp@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 75 ASSAY: Xpert MTB/RIF Ultra, COUNT: 574 ASSAY: Others, COUNT: 21 | 2023-11-15 | GX4 10C | 120000171 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 220011687 >>> 210048180 - Generated reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210048180 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762536897069976288/5762536897069976288_signature_28.png | Donnel M. Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5762536897069976288/5762536897069976288_signature_31.png
| 2023-11-20 | 2023-11-17 13:33 | Philippine Heart Center | East Avenue, QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26503 | MACARE-HBDC | phcmolecularlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 902 ASSAY: Others, COUNT: 14 ASSAY: Xpert SARS-CoV-2, COUNT: 3001 |  | GX4 | 834443 | Workstation unable to booth,loose power inlet terminal | Restart workstation Reset RAM Checked voltage output Fixed power inlet terminal Checked power supply | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762474223618133562/5762474223618133562_signature_28.png | ISON SAN JOSE | https://www.jotform.com/uploads/macarewecare/222927978372067/5762474223618133562/5762474223618133562_signature_31.png
| 2023-11-20 | 2023-11-17 16:33 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26502 | MACARE-HBDC | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert CT/NG, COUNT: 8 ASSAY: Others, COUNT: 39 ASSAY: Xpert HPV, COUNT: 69 ASSAY: Xpert MTB/RIF, COUNT: 101 ASSAY: Xpert MTB/RIF Ultra, COUNT: 117 ASSAY: Xpert HCV VL, COUNT: 35 |  | GX16 | 804476 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | 9 module disabled, 3module operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762473883611750160/5762473883611750160_signature_28.png | Ma Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5762473883611750160/5762473883611750160_signature_31.png
| 2023-11-20 | 2023-11-17 16:34 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.com |  |  | 26501 | MACARE-HBDC | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 363 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Others, COUNT: 20 ASSAY: Xpert HBV VL, COUNT: 281 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 28 ASSAY: Xpert MTB/RIF, COUNT: 169 ASSAY: Xpert MTB/RIF Ultra, COUNT: 26 ASSAY: Xpert HCV VL, COUNT: 25 |  | Gx4 | 838254 | Due for maintenance B1,B3,B4 disabled | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | B2 passed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762473613615185982/5762473613615185982_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5762473613615185982/5762473613615185982_signature_31.png
| 2023-11-20 | 2023-11-17 16:34 | National Kidney and Transplant Institute | East avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@gmail.con |  |  | 26500 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 978 ASSAY: Others, COUNT: 31 ASSAY: Xpert HBV VL, COUNT: 675 ASSAY: Xpert HPV, COUNT: 128 ASSAY: Xpert MTB/RIF, COUNT: 933 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1476 ASSAY: Xpert SARS-CoV-2, COUNT: 15088 ASSAY: Xpert HCV VL, COUNT: 71 |  | GX16 | 110005587 | Due for maintenance C2 and A4 hw failed | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | 14 module operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762473323614303344/5762473323614303344_signature_28.png | Ma. Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5762473323614303344/5762473323614303344_signature_31.png
| 2023-11-19 | 2023-11-14 16:10 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | NO | INSTALLATION MAINTENANCE | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com |  |  | 26499 | DOH | onehope834@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 943 |  | Gx-IV | 838528 | CPU problem | Perform Genexpert machine Preventive maintenance | Replace defective CPU SN: *CPH0520F08552* with refurbished CPU SN: *CPH1119FO4268* | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762183144319317567/5762183144319317567_signature_28.png | Evelyn A,Cadiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5762183144319317567/5762183144319317567_signature_31.png
| 2023-11-19 | 2023-11-17 16:18 | Ospital ng Palawan GX XVI | San Juan Rd. Cor. Malvar St., Puerto Princesa City, Palawan | 4B | Palawan | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26498 | PBSP | svsilva.pbsp@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3284 ASSAY: Xpert SARS-CoV-2, COUNT: 1747 ASSAY: Xpert HIV-1 Viral Load, COUNT: 173 | 2023-11-06 | GX IV 6C | 842893 | Modules A4 and B1 for replacment | Performed pm Replaced modules | Replaced 2 modules A4: 210046004 -> 880285 B1: 686440 -> 730690 | Unit is operational All 16 modules working |  | 2 GX modules 6C | DR 45256 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762107328323673363/5762107328323673363_signature_28.png | Shayne Marie Silva | https://www.jotform.com/uploads/macarewecare/222927978372067/5762107328323673363/5762107328323673363_signature_31.png
| 2023-11-19 | 2023-11-17 14:48 | Ospital ng Palawan GXIV | San Juan Rd. Cor. Malvar St., Puerto Princesa City, Palawan | 4B | Palawan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26497 | PBSP | svsilva.pbsp@gmail.com | 29 |  | 2023-10-30 | GX IV 6C | 804282 | Laptop for replacement | Performed PM Replaced Laptop Changed date & time Setup admin account Generated report | Replaced Laptop | Unit is operational All modules working |  | Lenovo Laptop | FedEx TRK# 7739 0338 7097 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762099978327895428/5762099978327895428_signature_28.png | Shayne Marie V. Silva | https://www.jotform.com/uploads/macarewecare/222927978372067/5762099978327895428/5762099978327895428_signature_31.png
| 2023-11-19 | 2023-11-16 15:25 | Puerto Princesa City Health Office | Valencia St. Brgy. Modal, Puerto Proncesa City, Palawan | 4B | Palawan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26496 | PBSP | jocal0324@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1239 | 2023-11-06 | GX IV 6C | 804272 | Modules D1, D2, D3 for replacement | Performed PM Replaced modules Generated report | Replaced 3 modules D1: 698691 -> 729707 D2: 643930 -> 709361 D3: 640245 -> 804652 | Unit is operational 3 modules working |  | 3 Gx modules 6C | DR 45259 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5762096408324243798/5762096408324243798_signature_28.png | Jocelyn L. CALALIN | https://www.jotform.com/uploads/macarewecare/222927978372067/5762096408324243798/5762096408324243798_signature_31.png
| 2023-11-18 | 2023-11-16 14:51 | Kalinga Health Hub - Marikina | JP Rizal St. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26495 | CONSORTIUM | krisna.morante@innovationsvh.org | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 371 |  | GX-IV-D-6C | 110008418 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Checked module B3 Generated reports | Unit is operational. Module B3/210035667 not operational. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760860340317295698/5760860340317295698_signature_28.png | Krisna Morante | https://www.jotform.com/uploads/macarewecare/222927978372067/5760860340317295698/5760860340317295698_signature_31.png
| 2023-11-18 | 2023-11-17 10:17 | Lumban Rural Health Unit | Lumban | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26494 | PBSP | ferreradelyn@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1202 | 2023-11-17 | GX-IV-D-6C | 838148 | Module B3 for replacement | Replaced module B3 | Performed maintenance Replaced module B3 Generated reports | All modules operational |  | Old: B3/630625 New: B3/787983 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760860340317911236/5760860340317911236_signature_28.png | Adelyn V. Ferrer | https://www.jotform.com/uploads/macarewecare/222927978372067/5760860340317911236/5760860340317911236_signature_31.png
| 2023-11-18 | 2023-11-17 11:50 | Siniloan Rural Health Unit | Siniloan | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26493 | PBSP | siniloanruralhealthunit@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 780 |  | GX-IV-D-6C | 838165 | Module B3 has error 1002 | Performed maintenance Checked module B3 | Performed maintenance Checked module B3 Functionality test Generated reports | Unit is operational. Module B3 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760860340318880720/5760860340318880720_signature_28.png | Jacobina A. Ybardolaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5760860340318880720/5760860340318880720_signature_31.png
| 2023-11-17 | 2023-11-17 13:45 | Rogaciano M. Mercado Memorial Hospital | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26492 | PBSP | alona0276@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2267 |  | GXIV | 805252 | Stuck cartridge on A1 | #NAME? | Performed troubleshooting on A1 | Unit is operational. Module A1 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760228071919866808/5760228071919866808_signature_28.png | Alona Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5760228071919866808/5760228071919866808_signature_31.png
| 2023-11-17 | 2023-11-17 15:03 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | TB Dots Laboratory | NO | MAINTENANCE Xpertcheck | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | roxasmatthew659@yahoo.com.ph |  | 26491 | DOH | zionclydematote@gmail.com | 26 | ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert HPV, COUNT: 581 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3355 | 2023-11-17 | Genexpert IV | 831718 | Module 1 and 4 error | Performed Xpertcheck on module 1, 4 | Performed Maintenance Gathered Reports Performed Functionality | Unit is Operational | A1, A4 = Passed | Xpertcheck 3pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760145813923773152/5760145813923773152_signature_28.png | Sunshine Matote | https://www.jotform.com/uploads/macarewecare/222927978372067/5760145813923773152/5760145813923773152_signature_31.png
| 2023-11-17 | 2023-11-17 15:06 | Lapu-Lapu City Health Office | Lapu-Lapu City, Cebu | 7 | Cebu | Tb-Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  | wbreoma.2020@gmail.com |  |  | 26490 | PBSP | tanya_pax@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 210 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2497 | 2023-11-17 | GXIV-4D | 808637 | Due for XPC | #NAME? | #NAME? | Module B2 failed XpertCheck. Module B4 frequent temp error. | Module B4 under observation. | Xpertcheck-CE-5 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760142981578537748/5760142981578537748_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5760142981578537748/5760142981578537748_signature_31.png
| 2023-11-17 | 2023-11-17 13:05 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26489 | MACARE-HBDC | ruellamorena@gmail.com | 26 | ASSAY: Xpert SARS-CoV-2, COUNT: 1238 | 2023-11-17 | Genexpert IV | 110008457 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760071264741740310/5760071264741740310_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5760071264741740310/5760071264741740310_signature_31.png
| 2023-11-17 | 2023-11-17 11:48 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Serology | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26488 | MACARE-HBDC | cvmcserology@gmail.com | 27 |  | 2023-11-13 | Genexpert IV | 803870 | N/a | N/a | Installed Genexpert Installed laptop Installed UPS Installed Barcode Scanner Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert IV Laptop UPS BARCODE | 10822 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760028003961218531/5760028003961218531_signature_28.png | Dorothy Pajarillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5760028003961218531/5760028003961218531_signature_31.png
| 2023-11-17 | 2023-11-17 11:48 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Serology | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26487 | MACARE-HBDC | cvmcserology@gmail.com | 27 |  | 2023-11-13 | Genexpert IV | 803870 | N/a | N/a | Installed Genexpert Installed laptop Installed UPS Installed Barcode Scanner Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert IV Laptop UPS BARCODE | 10822 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760027733964069903/5760027733964069903_signature_28.png | Dorothy Pajarillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5760027733964069903/5760027733964069903_signature_31.png
| 2023-11-17 | 2023-11-17 11:48 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Serology | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26486 | MACARE-HBDC | cvmcserology@gmail.com | 27 |  | 2023-11-13 | Genexpert IV | 803870 | N/a | N/a | Installed Genexpert Installed laptop Installed UPS Installed Barcode Scanner Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert IV Laptop UPS BARCODE | 10822 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760027093083555410/5760027093083555410_signature_28.png | Dorothy Pajarillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5760027093083555410/5760027093083555410_signature_31.png
| 2023-11-17 | 2023-11-17 11:09 | Balamban Rural Health Unit | Poblacion | 7 | Cebu | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 26485 | DOH | climaco.mildred@gmail.com | 23 |  | 2023-11-17 | GX4 | 838564 | Module A2 hardware fail | Module A2 with serial number 754214 fail for calibration. | Perform maintenance for all modules Calibrate module A2 only. | Module A2 not ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5760009805413644376/5760009805413644376_signature_28.png | Jonnel P. Barnayha | https://www.jotform.com/uploads/macarewecare/222927978372067/5760009805413644376/5760009805413644376_signature_31.png
| 2023-11-17 | 2023-11-17 09:54 | PHO MBL Palawan | PGP Compund, south National Hi-Way, Brgy. Irawan, Puerto Princesa | 4B | Palawan | Molecular Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26484 | PBSP | phomblpalawan@gmail.com | 29 |  |  | GC IV 10C | 120000182 | GX for installation | Installed GX and its accessories Performed PM Generated report | Installed GX and all accessories | Unit is operational UPS defective out of the box UPS RPT 2000AP for reporting s/n 40208832105 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759960309526260723/5759960309526260723_signature_28.png | Cyrus Cecilio A. Caabay | https://www.jotform.com/uploads/macarewecare/222927978372067/5759960309526260723/5759960309526260723_signature_31.png
| 2023-11-17 | 2023-11-17 07:47 | Baggao Rural Health Unit | San Jose, Baggao, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26483 | PBSP | eslabrajan02@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 785 | 2023-11-17 | Genexpert IV | 835092 | Module 2 Hardware Failure | Performed Module Replacement on Module 2 - (787950 - 713053) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759886492317907327/5759886492317907327_signature_28.png | Jan Anthony Eslabra | https://www.jotform.com/uploads/macarewecare/222927978372067/5759886492317907327/5759886492317907327_signature_31.png
| 2023-11-17 | 2023-11-16 14:32 | Medilinx | Amorsolo st,Makati CIty | NCR | NCR 4th - Makati | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  | florescanoel77@fmail.com | reinaluz.bactol@medilinx.com.ph |  | 26482 | COMMERCIAL | angelynlangreo.medilinx@outlook.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 137 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 619 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 659 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2427 ASSAY: Others, COUNT: 2 ASSAY: Others, COUNT: 10 ASSAY: Xpert HCV VL, COUNT: 57 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1739 | 2023-11-14 | GX4 (10C) | 110001647 | Due for maintenance and xpertcheck | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance and xpertcheck | B1 failed xpertcheck B2 and B3 newly replaced module,B4 passed xpertcheck |  | Xpertcheck | 10826 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759838275517881183/5759838275517881183_signature_28.png | Angelyn Langreo | https://www.jotform.com/uploads/macarewecare/222927978372067/5759838275517881183/5759838275517881183_signature_31.png
| 2023-11-16 | 2023-11-16 13:58 | Amang Rodriguez Memorial Medical Center | Marikina-Sumulong Highway | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26481 | PBSP | maricel_c27@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3999 |  | GX-IV-D-6C | 807982 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759245519027494122/5759245519027494122_signature_28.png | Maricel Calderon | https://www.jotform.com/uploads/macarewecare/222927978372067/5759245519027494122/5759245519027494122_signature_31.png
| 2023-11-16 | 2023-11-16 13:48 | Eastern Visayas Regional Medical Center | Bagacay, Tacloban City | 8 | Leyte | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | merl_agner62@yahoo.com |  | 26480 | PBSP | shyzamoras@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 150 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4400 ASSAY: Xpert HIV-1 Viral Load, COUNT: 250 | 2023-11-16 | GXIV | 808631 | Module B4/681450 error 2032 | #NAME? | #NAME? | Unit is Operational Module B4 disabled (Defective) | Module B4 for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759236013126535140/5759236013126535140_signature_28.png | Shirlyn C . ZAMORAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5759236013126535140/5759236013126535140_signature_31.png
| 2023-11-16 | 2023-11-16 13:24 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26479 | COMMERCIAL | jonnifercodoyceniza@gmail.com | 28 |  |  | Gx 4 | 829419 | Cannot open windows of Laptop | Restart Laptop, clean fan filter | Restart Laptop, clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759220029226766686/5759220029226766686_signature_28.png | Jonnifer C. Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5759220029226766686/5759220029226766686_signature_31.png
| 2023-11-16 | 2023-11-16 12:06 | Talakag Rural Health Unit | Poblacion, Talakag | 10 | Bukidnon | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26478 | DOH | mhotalakaglaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 219 | 2023-08-03 | GX-4 | 838566 | None | None | XpertCheck Calibration Preventive Maintenance | All modules are functional |  | XpertCheck CE-IVD 5 Cartridges Lot No. 1000410214 | 10454 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759176042524083558/5759176042524083558_signature_28.png | Alanna Eureka D. Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/5759176042524083558/5759176042524083558_signature_31.png
| 2023-11-16 | 2023-11-16 11:45 | Marikina City Health Office | Shoe Ave. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26477 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1041 ASSAY: Xpert HIV-1 Viral Load, COUNT: 161 |  | GX-IV-D-6C | 805211 | For network switch connection | Connected to network switch | Performed maintenance Connected 2 units of GeneXpert in one desktop (805211/831588) Functionality test Generated reports | Unit is operational. Module B3/657408 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759165769021852400/5759165769021852400_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5759165769021852400/5759165769021852400_signature_31.png
| 2023-11-16 | 2023-11-16 11:50 | Marikina City Health Office | Shoe Avenue | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26476 | MACARE-HBDC | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 45 |  | GX-IV-D-10C | 110016107 | For relocation | Transferred the GeneXpert System | Performed maintenance Transferred the GeneXpert system to other location (same room) Functionality test Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759164639024590649/5759164639024590649_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5759164639024590649/5759164639024590649_signature_31.png
| 2023-11-16 | 2023-11-16 11:22 | Amma Jadsac District Hospital | Pudtol, Apayao | CAR | Apayao | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26475 | PBSP | ajdh.hospital@gmail.com | 25 | ASSAY: , COUNT: 2833 | 2023-11-16 | Genexpert IV | 807609 | Module 3 Hardware Failure | Performed Module Replacement on Module 3 - (915564 - 690475) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Module - 690475 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759145646939498892/5759145646939498892_signature_28.png | Maybel Abuan | https://www.jotform.com/uploads/macarewecare/222927978372067/5759145646939498892/5759145646939498892_signature_31.png
| 2023-11-16 | 2023-11-16 08:45 | Notre Dame De Chartres Hospital | Baguio City | CAR | Benguet | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26474 | COMMERCIAL | srnormavspc@gmail.com | 25 |  |  | MCS+ | 10M103 | Unit for calibration | #NAME? | Performed calibration | Unit is operational. All calibration parameters passed. | w/ Engr. Salih |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5759077108317955602/5759077108317955602_signature_28.png | Norma Vicentillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5759077108317955602/5759077108317955602_signature_31.png
| 2023-11-15 | 2023-11-15 23:41 | Jose Abad Santos Rural Health Unit | Caburan, Jose Abad Santos | 11 | Davao Occidental | TB-DOTS | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 26473 | DOH | bevespinosa168@gmail.com | 28.6 |  | 2023-08-03 | GX-IV 4d | 838460 | Due for Calibration and Preventive Maintenance CPU System Unit not working | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Troubleshoot the system unit by reseating and cleaning the ram, cmos battery and solid state drive. Run XPC Calibration | Perform Preventive Maintenance, and XpertCheck Calibration Troubleshooting the System Unit | After thoroughly isolation while troubleshooting the System Unit, hence it was defective for not detecting the SSD. No OS was detected. All Modules Passed the Calibration | System Unit was damaged due to leakage of the ceiling of the site. And XpertCheck Calibration was done on a backup/spare unit of the Laboratory | Xpertcheck CE 5 | 10439 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758724836412950693/5758724836412950693_signature_28.png | Beverly Espinosa, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5758724836412950693/5758724836412950693_signature_31.png
| 2023-11-15 | 2023-11-15 18:08 | Klinika Novaliches Sundown Clinic | 2nd flr, Bautista Bldg, Novaliches Proper, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26472 | MACARE-HBDC | klinikanovaliches2022@gmail.com | 29 |  |  | GX IV 10C | 120000140 | For installation | Setup GX machine and accessories | Installed GeneXpert | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758661519211807841/5758661519211807841_signature_28.png | Michelle Louise D. Perono, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5758661519211807841/5758661519211807841_signature_31.png
| 2023-11-15 | 2023-11-15 16:54 | V. Luna Medical Center | V. Luna Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Carlos Leniel Origenes |  | carlos.leniel.origenes@gmail.com |  |  | 26471 | COMMERCIAL | laboratorynco@gmail.com | 25 |  |  | MCS+ | 17H636SPG | One calibration parameter for adjustment | Adjusted Bowl Optics offset | Adjusted Bowl Optics offset Performed functionality test and simulation run | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758488443691594693/5758488443691594693_signature_28.png | Jem Gagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5758488443691594693/5758488443691594693_signature_31.png
| 2023-11-15 | 2023-11-15 15:06 | Molave Rural Health unit | Molave, Zamboanga del sur | 9 | Zamboanga del Sur | Labia | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia54@gmail.com |  | 26470 | DOH | aldrinacain@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 780 | 2023-11-15 | GX 4 | 838233 | Module 3 h/w failed | Preventive Maintenence while module 2 undergone Xpertcheck calibration and passed | Preventive Maintenence while module 2 undergone Xpertcheck calibration and passed | Module 1, 2 and 4 are available | Module 2 for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758415290147786536/5758415290147786536_signature_28.png | Aldrin L. Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/5758415290147786536/5758415290147786536_signature_31.png
| 2023-11-15 | 2023-11-15 14:17 | Cardinal Santos Medical Center | Wilson St. Greenhills | NCR | NCR 2nd - San Juan | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 26469 | COMMERCIAL | gabbyg.carag@gmail.com | 25 |  |  | MCS+ | 08M461 | For calibration | Calibrated the unit | Performed maintenance Checked calibration parameters Checked rotor pumps Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758393440321481050/5758393440321481050_signature_28.png | Ma Gabriela Carag | https://www.jotform.com/uploads/macarewecare/222927978372067/5758393440321481050/5758393440321481050_signature_31.png
| 2023-11-15 | 2023-11-15 13:57 | Siquijor Provincial Health Office | Caipilan, Siquijor | 7 | Siquijor | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  | ramasmarkefren@gmail.com | glinniericamor73@gmail.com |  | 26468 | PBSP | glinniericamor73@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1471 |  | GX4 | 808638 | B1 SN # 667651 and B3 SN # 210068080 failed during calibration last October | Replaced new module , generate data IQR, gxx giles , system log , assay statistis, clean filter and fan , check voltage supply | Module replacement in B1 New module 706452 and B3 New Module 807134 | All module are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758376752225081090/5758376752225081090_signature_28.png | Glinnie Ricamor, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5758376752225081090/5758376752225081090_signature_31.png
| 2023-11-15 | 2023-11-15 13:47 | CHONG HUA HOSPITAL - MANDAUE | MANDAUE CITY | 7 | Cebu | BLOOD BANK | NO | REPAIR MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Joseph Acebedo |  | jacebedo.macare@gmail.com |  |  | 26467 | COMMERCIAL | bloodbankmandaue@gmail.com | 22 |  |  | MCS+ | 11J217 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758376418433492151/5758376418433492151_signature_28.png | CHH MANDAUE BLOODBANK | https://www.jotform.com/uploads/macarewecare/222927978372067/5758376418433492151/5758376418433492151_signature_31.png
| 2023-11-15 | 2023-11-15 13:44 | Cardinal Santos Medical Center | Wilson St. Greenhills | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com | biomed@csmc.ph |  | 26466 | COMMERCIAL | microbiology.csmc@gmail.com | 24 | ASSAY: Xpert HBV VL, COUNT: 82 ASSAY: Xpert HPV, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1005 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. Module B1 not operational. Reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758371018355761462/5758371018355761462_signature_28.png | Paula Beatrice N. Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5758371018355761462/5758371018355761462_signature_31.png
| 2023-11-15 | 2023-11-15 13:43 | Mahayag rural health unit | Mahayag, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 26465 | PBSP | cavanricky@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1095 | 2023-11-15 | GX 4 | 834436 | Module 1, 3 and 4 Failed in xpertcheck calibration | Replaced modules 778048, 778490 and 778141 with modules 719869, 788463 and 723243 respectively | Replaced module and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758366040147244442/5758366040147244442_signature_28.png | Ricky M. Cavan | https://www.jotform.com/uploads/macarewecare/222927978372067/5758366040147244442/5758366040147244442_signature_31.png
| 2023-11-15 | 2023-11-15 13:37 | Maddela District Hospital | Maddela, Quirino | 2 | Quirino | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | roxasmatthew659@yahoo.com.ph |  | 26464 | DOH | roxasmatthew659@yahoo.com.ph | 27 | ASSAY: , COUNT: 262 | 2023-10-27 | Genexpert IV | 838469 | UPS not working Monitor half display | Replaced Monitor Replaced UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | UPS and Monitor | 10787 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758365120321050176/5758365120321050176_signature_28.png | Vanessa Roxanne Balbin | https://www.jotform.com/uploads/macarewecare/222927978372067/5758365120321050176/5758365120321050176_signature_31.png
| 2023-11-15 | 2023-11-15 11:50 | Bukidnon Provincial Hospital - Kibawe | C.M Recto St. Palma, Kibawe | 10 | Bukidnon | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26463 | DOH | juliealte10@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 615 | 2023-08-03 | GX-4 | 838547 | None | All modules passed in XPC Calibration | Preventive Maintenance XpertCheck Calibration | All modules are functional |  | XpertCheck CE-IVD 5 Cartridges Lot No. 1000410214 | 10451 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758304867318842944/5758304867318842944_signature_28.png | John Rhel T. Denque | https://www.jotform.com/uploads/macarewecare/222927978372067/5758304867318842944/5758304867318842944_signature_31.png
| 2023-11-15 | 2023-11-15 10:44 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | IPMM | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26462 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 23 |  |  | MCS+ | 04H020 | For calibration | Calibrated the unit. | Performed maintenance Checked calibration parameters Checked rotor pumps Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758267366395467300/5758267366395467300_signature_28.png | Maria Zoraida Caina | https://www.jotform.com/uploads/macarewecare/222927978372067/5758267366395467300/5758267366395467300_signature_31.png
| 2023-11-15 | 2023-11-15 10:42 | Lavezares Rural Health Unit | Lavezares Northern Samar | 8 | Northern Samar | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  | johnpaullabutap@gmail.com | zexuzahmpaul22@gmail.com |  | 26461 | PBSP | jotchan3@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 340 | 2023-11-15 | GXIV-2D | 834821 | Module B2/787189 for replacement | #NAME? | - Install new module in B2 w/ SN: 791454 - Performed Functionality test - Performed Preventive maintenance - Generate reports - Pullout defective modules SN:787189 | Unit is Operational |  | 1 GeneXpert Module 6 color SN: 791454 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758258405515880124/5758258405515880124_signature_28.png | Jocelyn T. Chan | https://www.jotform.com/uploads/macarewecare/222927978372067/5758258405515880124/5758258405515880124_signature_31.png
| 2023-11-15 | 2023-11-15 10:25 | The Medical City | Ortigas ave. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26460 | COMMERCIAL | moleculardx@themedicalcity.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 23 ASSAY: Xpert CT/NG, COUNT: 40 ASSAY: Xpert HBV VL, COUNT: 338 ASSAY: Xpert HPV, COUNT: 261 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2697 ASSAY: Xpert HCV VL, COUNT: 51 ASSAY: Xpert HIV-1 Viral Load, COUNT: 476 ASSAY: Xpert SARS-CoV-2, COUNT: 65 |  | GX-XVI-D-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758247456392842787/5758247456392842787_signature_28.png | Paolo Crispin | https://www.jotform.com/uploads/macarewecare/222927978372067/5758247456392842787/5758247456392842787_signature_31.png
| 2023-11-15 | 2023-11-15 09:26 | The Medical City | Ortigas Ave. | NCR | NCR 2nd - Pasig | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  | jmdoctor@macare-medicals.com |  |  | 26459 | COMMERCIAL | bloodbank@themedicalcity.com | 24 |  |  | MCS+ | 02F070 | For Calibration | Calibrated the unit | Performed maintenance Checked calibration parameters Checked rotor pumps Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5758232006393523504/5758232006393523504_signature_28.png | Daryl Tio | https://www.jotform.com/uploads/macarewecare/222927978372067/5758232006393523504/5758232006393523504_signature_31.png
| 2023-11-14 | 2023-11-14 19:54 | Bicol Medical Center | Naga | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com | danemilbodino23@gmail.com |  | 26458 | PBSP | cristinerod211@gmail.com | 27 | ASSAY: , COUNT: 4 ASSAY: Xpert HPV, COUNT: 1473 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7161 ASSAY: Xpert SARS-CoV-2, COUNT: 2341 ASSAY: Xpert HIV-1 Viral Load, COUNT: 346 |  | GX16 | 842225 | B3/640895, D1/719871 for module replacement | Replaced Module: B3-640895>>210013422, D1-719871>>909656 Cleaned cartridge bays, fan, plunger rod Generate reports | Replace Module: B3/640895, D1/719871 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757790875118594229/5757790875118594229_signature_28.png | Cristine Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5757790875118594229/5757790875118594229_signature_31.png
| 2023-11-14 | 2023-11-08 13:04 | Quezon Medical Center | Quezon Ave., Lucena, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26457 | PBSP | joannpalermo0613@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 337 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2121 ASSAY: Xpert HIV-1 Viral Load, COUNT: 95 |  | GX IV 6C | 804692 | CPU not working | Setup 2 units Combined using network switch Perfomed maintenance Generated report | Combined 2 GX machine temporarily using netw9rk switch | All modules working Temporary combined while waiting for the replacement cpu |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757761479218829639/5757761479218829639_signature_28.png | Jo-Ann Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5757761479218829639/5757761479218829639_signature_31.png
| 2023-11-14 | 2023-11-14 14:37 | AFPMC - AFRIMS | V. Luna, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26456 | PBSP | susieleonardia@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 691 ASSAY: Xpert SARS-CoV-2, COUNT: 2469 ASSAY: Xpert HCV VL, COUNT: 58 ASSAY: Others, COUNT: 48 ASSAY: Xpert HIV-1 Viral Load, COUNT: 74 |  | GX IV 10C | 845462 | Database error upon opening GX software | Checked and cleaned cartridge bay, plunger rod, filter and icore slot Troubleshoot GX software Restart laptop multiple times Performed database management task Archived and purged Jan 2023 - June 2023 reports Generated report | Troubleshooting GX software | Unit is operational. All modules working | w/ sir noel |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757751369211445279/5757751369211445279_signature_28.png | Susie Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5757751369211445279/5757751369211445279_signature_31.png
| 2023-11-14 | 2023-11-14 16:18 | V. Luna Medical Center | V. Luna, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26455 | MACARE-HBDC | laboratory.vlmc@yahoo.com | 27 |  |  | MCS+ | 17H636SP6 | Unit due for Calibration | Cleaned Unit Checked program parameters Checked and adjust line sensor value Checked rollers, dpm, spm, centrifuge Performed functionality test Performed isolation running | Performed Calibration | Bowl optics potentiometer lost connection 6 rollers for replacement | w/ Sir Noel Floresca |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757750219216871367/5757750219216871367_signature_28.png | VICTORIANO LUNA MEDICAL CENTER | https://www.jotform.com/uploads/macarewecare/222927978372067/5757750219216871367/5757750219216871367_signature_31.png
| 2023-11-14 | 2023-11-14 15:30 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26454 | PBSP | bmclab_micro@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1500 ASSAY: Xpert HCV VL, COUNT: 25 ASSAY: Xpert HIV-1 Viral Load, COUNT: 318 |  | GXIV | 805112 | Unit for calibration | #NAME? | Performed calibration | Unit is operational. B1: passed B2: passed B3: passed B4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757576510396341066/5757576510396341066_signature_28.png | Alyssa Ico | https://www.jotform.com/uploads/macarewecare/222927978372067/5757576510396341066/5757576510396341066_signature_31.png
| 2023-11-14 | 2023-11-14 15:11 | The Medical City - Clark | 100 Gatwick Gateway Sabah Al-Ahmad Global Gateway Logistics City Industrial Estate 5 Clark Freeport Zone | 3 | Pampanga | Microbiology | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 26453 | PBSP | remortajessarae1998@gmail.com | 56 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 31 ASSAY: Xpert HPV, COUNT: 27 ASSAY: Xpert MTB/RIF Ultra, COUNT: 187 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-11-14 | Genexpert IV | 839235 | Barcode Scanner not working | Replaced Barcode Scanner | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Barcode Scanner - 1618000503884 | 10824 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757564028226864876/5757564028226864876_signature_28.png | Jessa Rae Remorta | https://www.jotform.com/uploads/macarewecare/222927978372067/5757564028226864876/5757564028226864876_signature_31.png
| 2023-11-14 | 2023-11-14 14:26 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 26452 | PBSP | bmcgenexpertlab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA | 2023-11-14 | GXIV | 884623 | Back up laptop for installation | #NAME? | Installed laptop | Unit is operational. |  | Lenovo Laptop SN: CPH1019F03266 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757550990395270117/5757550990395270117_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5757550990395270117/5757550990395270117_signature_31.png
| 2023-11-14 | 2023-11-14 14:30 | Quezon Rural Health Unit | Poblacion, Quezon | 10 | Bukidnon | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26451 | DOH | quenxlopez@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 780 |  | GX-4 | 838440 | Module B2 found service required in XPC | None | XpertCheck Calibration Preventive Maintenance | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757531808987939798/5757531808987939798_signature_28.png | Quenie Rose Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5757531808987939798/5757531808987939798_signature_31.png
| 2023-11-14 | 2023-11-14 14:11 | Kumalarang rural health unit | Kumalarang, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  | niagnozia54@gmail.com | niagnozia@gmail.com |  | 26450 | PBSP | lisaleebindab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 940 | 2023-11-14 | GX 4 | 815278 | None | Replaced module 695634 with Module 695631. Preventive Maintenance | Replaced module 695644 with Module 695631. Preventive Maintenance | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757519640028306105/5757519640028306105_signature_28.png | Lisa Lee B. Indab | https://www.jotform.com/uploads/macarewecare/222927978372067/5757519640028306105/5757519640028306105_signature_31.png
| 2023-11-14 | 2023-11-14 11:42 | Pamplona Rural Health Unit | Pamplona | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  | danemilbodino23@gmail.com |  |  | 26449 | PBSP | francislozano226@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 304 |  | GX4 | 812774 | A2/647604 for module replacement | Replaced Module: A2 647604>>874672 Cleaned cartridge bays, plunger rod, fan Generate reports | Replaced module A2 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757432122717531630/5757432122717531630_signature_28.png | Francis S. Lozano | https://www.jotform.com/uploads/macarewecare/222927978372067/5757432122717531630/5757432122717531630_signature_31.png
| 2023-11-14 | 2023-11-14 11:33 | Passi City Health Office | Iloilo city | 6 | Iloilo | Laboratory | YES | REPAIR | CEPHEID | Jovy Noel |  | jdnoel.eisi@gmail.com | jdnoel.eisi@gmail.com |  | 26448 | PBSP | passichoab@gmail.com | 23 |  | 2023-11-14 | Gx4 | 804701 | Calibration fail | Replace module B1 and B4 with new serial number B1-210009712 B4-646647 | Replace module B1 and B4 | Module B1 ok to use module B4 not ok with o use h/w error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757428366666567053/5757428366666567053_signature_28.png | John Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5757428366666567053/5757428366666567053_signature_31.png
| 2023-11-14 | 2023-11-14 10:47 | Pangantucan Rural Health Unit | Poblacion, Pangantucan | 10 | Bukidnon | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  | avenido.faifai@gmail.com |  |  | 26447 | PBSP | melita.alas@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 630 |  | GX-2 | 835191 | Stuck Cartridge at Module B2 | None | Removing Stuck Cartridge Preventive Maintenance XpertCheck Calibration | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5757399526025407216/5757399526025407216_signature_28.png | Melita P. Alas | https://www.jotform.com/uploads/macarewecare/222927978372067/5757399526025407216/5757399526025407216_signature_31.png
| 2023-11-13 | 2023-11-07 16:02 | Masambong Health Center | 8 Malac St. Masambong, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26446 | PBSP | michellebonrostro@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1715 | 2023-11-07 | GXIV 6C | 814160 | Module B4 for replacement | Performed PM Replaced Module b4 Generated report | Replaced Module B4 210063218 -> 732791 | Unit is operational. All modukes working |  | GX Module 6C | G 6404 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756743183614123735/5756743183614123735_signature_28.png | Michelle Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/5756743183614123735/5756743183614123735_signature_31.png
| 2023-11-13 | 2023-11-10 10:42 | Alabat Island District Hospital | National road, Brgy. 4, Poblacion, Alabat, Quezon Province | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26445 | DOH | letpiquero@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 759 | 2023-11-10 | GX IV 6C | 831614 | Unit due for xpertcheck | Checked and cleaned cartridge bay plunger rod, icore slot, and filter Performed xpertcheck All modules passed xpertcheck Generated report | Performed XpertCheck | Unit is operational All modules working | All modules passed xpertcheck | XpertCheck kit | DR 10778 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756741983612097842/5756741983612097842_signature_28.png | Carmelita Piquero | https://www.jotform.com/uploads/macarewecare/222927978372067/5756741983612097842/5756741983612097842_signature_31.png
| 2023-11-13 | 2023-11-13 14:48 | Bacolod City Health Office | Burgos St., Bacolod City | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 26444 | PBSP | cyrilcordova0515@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6490 |  | GX4 | 812020 | N/A | Pulled out the defective module for D3 with SN 864267 and save IQR file before and after replacement. | Installed the new module for D3 with SN 667778. | All 4 modules are now working in good condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756682717519443459/5756682717519443459_signature_28.png | Mr. Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5756682717519443459/5756682717519443459_signature_31.png
| 2023-11-13 | 2023-11-10 10:01 | Lingap Kalusugan CHU I - Alima | Evangelista, Alima, Bacoor | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26441 | PBSP | eng112803@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1727 |  | GXIV-D | 838185 | Module 4 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 4 Generate Reports | 3 Modules are Operational | Module 1 for Reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756518193612037980/5756518193612037980_signature_28.png | Aide E. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5756518193612037980/5756518193612037980_signature_31.png
| 2023-11-13 | 2023-11-09 09:41 | Biñan City Health Office I | Sto. Domingo, Biñan, Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26443 | PBSP | annefriam.almazora@yahoo.com | 26 | ASSAY: Xpert HPV, COUNT: 149 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1901 |  | GXIV-D | 812780 | Module 3 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 2 Checked Functionality Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756518193612267364/5756518193612267364_signature_28.png | Annefriam Almazora | https://www.jotform.com/uploads/macarewecare/222927978372067/5756518193612267364/5756518193612267364_signature_31.png
| 2023-11-13 | 2023-11-07 10:47 | Manila Doctors Hospital | U.N. Ave, Manila | NCR | NCR 1st - Manila | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  |  |  |  | 26442 | MACARE-HBDC | albeaniza@maniladoctors.com.ph | 24 |  |  | MCS+ | 02H078 | Unit is due for Maintenance Unit is due for Calibration | Preventive Maintenance Calibration | Cleaned Surface Area. Checked and Cleaned Rollers. Cleaned Valves and Sensors. Checked Centrifuge Function. Checked Pressures. Checked A/D Values. Performed Simulation Tests. Performed Calibration | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756518193613230361/5756518193613230361_signature_28.png | Adonis L. Beaniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5756518193613230361/5756518193613230361_signature_31.png
| 2023-11-13 | 2023-11-09 15:45 | Madridejos Rural Health Unit | Madridejos, Bantayan Island | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE ON-SITE ASSISTANCE Install new CPU | CEPHEID | Raymond Laurito |  | daymon899@gmail.com |  |  | 26440 | PBSP | cabasevilma@gmail.com | 27 |  |  | Gx 4 | 835057 | Defective CPU | Install New CPU, xpert check calibration | Install New CPU | Systems operational, all 2 modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756506553915495871/5756506553915495871_signature_28.png | Vilma E. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5756506553915495871/5756506553915495871_signature_31.png
| 2023-11-13 | 2023-11-13 08:39 | Marawi City Health Office | Marawi City | BARMM | Lanao del Sur | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26439 | DOH | faridahmamao@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 850 | 2023-08-08 | GX-IV | 838260 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | Xpertcheck Kit | 10477 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756455393911361413/5756455393911361413_signature_28.png | Faridah M. Busara, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5756455393911361413/5756455393911361413_signature_31.png
| 2023-11-13 | 2023-11-10 14:48 | Philippine Heart Center | East avenue,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION | CEPHEID | Noel Floresca |  |  |  |  | 26438 | MACARE-HBDC | phcmolecularlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 860 ASSAY: Others, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 1645 |  | GX16 | 838655 | Due for xpertcheck,module b4,a3,c3,a4,c4,d2 reported already for replacement | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Xpertcheck and maintenance | 8 module operational,8 module unoperational | B3,C1 failed in xpertcheck for replacement,D4 unable to performed xpertcheck due to lack of kits, passed xpertcheck a1,a2,b1,b2,c2,d1,d3 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756447993618968840/5756447993618968840_signature_28.png | ISON SAN JOSE | https://www.jotform.com/uploads/macarewecare/222927978372067/5756447993618968840/5756447993618968840_signature_31.png
| 2023-11-13 | 2023-11-10 10:25 | University of the Philippines | U.P diliman QC | NCR | NCR 2nd - Quezon City | Institute of biology Regen lab | NO | MAINTENANCE | OTHERS | Noel Floresca |  |  |  | Plasmair | 26437 | MACARE-HBDC | ladiaz@up.edu.ph | 26 |  |  | 20006-S | 2006-TS-833 | For maintenance/ dirty filter | Cleaned filter Checked filters Performed functionality test Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756445243614054139/5756445243614054139_signature_28.png | Leomir Diaz / Cielo Marquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5756445243614054139/5756445243614054139_signature_31.png
| 2023-11-13 | 2023-11-10 09:40 | University of the Philippines (Institute of Biology) | U.P diliman,QC | NCR | NCR 2nd - Quezon City | Biology | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  |  |  | Plasmair | 26436 | MACARE-HBDC | sciarm.ib.upd@up.edu.ph | 25 |  |  | Sentinel | 2006-TS-0015 | Error b encounter during startup of the unit module stage 1 weak current alarm,corrupted programs | Performed reset to default value Checked connection Checked filters Reset filter Restart system. Replaced bios battery Update software | Troubleshooting | Unit unoperational | For further troubleshooting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5756444773611838412/5756444773611838412_signature_28.png | Anna Isabel Navarro | https://www.jotform.com/uploads/macarewecare/222927978372067/5756444773611838412/5756444773611838412_signature_31.png
| 2023-11-10 | 2023-11-10 11:07 | Western Pangasinan District Hospital | Alaminos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26435 | PBSP | wpdhlab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1753 | 2023-11-10 | GX4 | 807598 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 884568 >>> 807790, 889789 >>> 685784 - Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 807790, 685784 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5754360897064371199/5754360897064371199_signature_28.png | Czarina Catabay | https://www.jotform.com/uploads/macarewecare/222927978372067/5754360897064371199/5754360897064371199_signature_31.png
| 2023-11-10 | 2023-11-10 19:28 | AMAI PAKPAK MEDICAL CENTER | MARAWI CITY | BARMM | Lanao del Sur | MOLECULAR LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26434 | PBSP | apmcmollab@gmail.com | 30 | ASSAY: Xpert SARS-CoV-2, COUNT: 7410 |  | GX-IV | 834444 | FOR MODULE B1 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE AND CLEANING OF GX WORKSTATION -PERFORM MODULE B1 REPLACEMENT: DEF SN: 210068644 NEW SN: 768575 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5754263427928239690/5754263427928239690_signature_28.png | Rayan S. Cosari, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5754263427928239690/5754263427928239690_signature_31.png
| 2023-11-10 | 2023-11-10 15:34 | AFP Hospital | V. Luna QC | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Harold Toco |  |  |  |  | 26433 | COMMERCIAL | laboratorynco@gmail.com | 30 |  | 2023-11-10 | MCS+ | 17H636SPG | None. For maintenance | Checked pump rotors. Verify calibration parameters. | Performed maintenance. | Unit is operational. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5754112494517029746/5754112494517029746_signature_28.png | Jem Gagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5754112494517029746/5754112494517029746_signature_31.png
| 2023-11-10 | 2023-11-10 14:38 | Buhangin Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26432 | DOH | buhanginlaboratory@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 869 |  | GX-IV 4D | 838232 | Due for Preventive Maintenance, and UPS Problem | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Troubleshoot the UPS | Perform Preventive Maintenance and Troubleshoot the UPS | All Modules are operational UPS is Defective at Battery Mode, |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5754081975972319762/5754081975972319762_signature_28.png | Allysa Thea C. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5754081975972319762/5754081975972319762_signature_31.png
| 2023-11-10 | 2023-11-10 13:27 | Morong Rural Health Unit | Gov. Tupas St., Morong | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26431 | DOH | morongrhu@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 538 |  | GX-IV-D-6C | 821213 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. Module A1/717338 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5754064414411191277/5754064414411191277_signature_28.png | Necitas V Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5754064414411191277/5754064414411191277_signature_31.png
| 2023-11-10 | 2023-11-09 13:51 | Makati Medical Center | Amorsolo st ,Makati City | NCR | NCR 4th - Makati | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 26430 | COMMERCIAL | majarocon022@gmail.com | 26 |  | 2023-11-07 | MCS+ | 97H330 | Due for maintenance and calibration,open wire barcode scanner | Cleaned unit Checked program parameters Checked and adjust line sensor value Checked rollers,dpm,spm,centrifuge Performed functionality test Performed isolation running Performed calibration | Maintenance and calibration | Operational | For replacement of barcode scanner cable | Urethane roller assy | 10801 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5754056207409308928/5754056207409308928_signature_28.png | Mark Dexter Majarocon | https://www.jotform.com/uploads/macarewecare/222927978372067/5754056207409308928/5754056207409308928_signature_31.png
| 2023-11-10 | 2023-11-08 16:14 | National Childrens Hospital | E.rodriguez QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26429 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 23 |  | 2023-11-07 | GX16 | 110014171 | For module replacement b3 898183 - 210067748 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance and module replacement | Operational |  | 900-0650rs | 10799 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5754054418815648676/5754054418815648676_signature_28.png | National Children's Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/5754054418815648676/5754054418815648676_signature_31.png
| 2023-11-10 | 2023-11-10 13:43 | AFRIMS | V.Luna QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Harold Toco |  |  |  |  | 26428 | MACARE-HBDC | mtvalderama@gmail.com | 28 |  | 2023-11-10 | Gx4 | 845462 | Stuck cartridge removal/maintenance. | Checked module door and hinges.. | Performed cleaning and maintenance.Removed cartridge. | All 4 modules operational. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5754050986022181341/5754050986022181341_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/5754050986022181341/5754050986022181341_signature_31.png
| 2023-11-10 | 2023-11-10 13:40 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | INSTALLATION REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26427 | PBSP | cdh.laboratorydept@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1096 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GXIV-D | 807562 | C4 for module replacement Due for calibration | Module replacement Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Installation of new HIV VL Assay Performed module replacement Gathered reports Performed functionality test | Unit is operational | All module passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5754041210639884537/5754041210639884537_signature_28.png | Maria Rowena T. Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5754041210639884537/5754041210639884537_signature_31.png
| 2023-11-10 | 2023-11-10 11:09 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 26426 | DOH | ccmclaboratory01@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1916 |  | GXIV-4D | 838561 | None | None | #NAME? | Module A4 disabled, waiting for replacement. Other modules are functional. | None |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753954113489872451/5753954113489872451_signature_28.png | Sheila Mae B. Lariosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5753954113489872451/5753954113489872451_signature_31.png
| 2023-11-10 | 2023-11-10 10:39 | Amang Rodriguez Memorial Medical Center | Marikina-Infanta Highway | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26425 | PBSP | pathologymole.armmc@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 3669 ASSAY: Others, COUNT: 19 |  | GX-IV-D-6C | 809257 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753933684411747983/5753933684411747983_signature_28.png | Racquel Rhys C. Verzosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5753933684411747983/5753933684411747983_signature_31.png
| 2023-11-10 | 2023-11-10 09:49 | Marikina City Health Office | Shoe Avenue | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26424 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 930 ASSAY: Xpert HIV-1 Viral Load, COUNT: 128 |  | GX-IV-L-6C | 805211 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753922514412518686/5753922514412518686_signature_28.png | Amelia C. Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5753922514412518686/5753922514412518686_signature_31.png
| 2023-11-10 | 2023-11-10 09:40 | Marikina City Health Office | Shoe Avenue | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26423 | MACARE-HBDC | amymitra22@gmail.com | 24 |  |  | GX-IV-D-10C | 120000613 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753922104413181059/5753922104413181059_signature_28.png | Amelia C. Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5753922104413181059/5753922104413181059_signature_31.png
| 2023-11-10 | 2023-11-10 10:10 | Marikina City Health Office | Shoe Ave. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26422 | DOH | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1045 ASSAY: Xpert HIV-1 Viral Load, COUNT: 163 |  | GX-IV-D-6C | 831588 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. Module A3 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753921594415488216/5753921594415488216_signature_28.png | Amelia C. Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5753921594415488216/5753921594415488216_signature_31.png
| 2023-11-10 | 2023-11-10 10:12 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26421 | MACARE-HBDC | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 10 |  | GX-IV-D-10C | 110016107 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753920574411032262/5753920574411032262_signature_28.png | Amelia C. Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5753920574411032262/5753920574411032262_signature_31.png
| 2023-11-09 | 2023-11-09 09:06 | Tagkawayan Rural Health Unit | Poblacion, Tagkawayan, Quezon Province | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26419 | DOH | rvteopy@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 621 | 2023-10-25 | GX IV 6C | 831916 | Unit due for XpertCheck | Checked and cleaned cartridge bay, plunger rod, icore slot and filter Performed Xpertcheck All passed xpertcheck Generated report | Performed XpertCheck | Unit is operational All modules working | 2 new modules. 2 passed xpertcheck | XpertCheck kit | DR 10781 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753377390113129385/5753377390113129385_signature_28.png | Rodelo V Teopy | https://www.jotform.com/uploads/macarewecare/222927978372067/5753377390113129385/5753377390113129385_signature_31.png
| 2023-11-09 | 2023-11-09 14:15 | Lopez Rural Health Unit | Magsaysay Center, Lopez, Quezon | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26420 | DOH | janette_villasanta@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 780 | 2023-10-25 | GX IV 6C | 831889 | Unit due for XpertCheck | Checked and cleaned cartridge bay, plunger rod, icore slot and filter Performed XpertCheck C1 failed XpertCheck Generated report | Performed XpertCheck | Unit is operational 3 modules passed xpertcheck, 1 module failed | C1/783221 Service Required | XpertCheck Kit | DR 10782 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753377390115679054/5753377390115679054_signature_28.png | Janette Villasanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5753377390115679054/5753377390115679054_signature_31.png
| 2023-11-09 | 2023-11-09 14:15 | Lopez Rural Health Unit | Magsaysay Center, Lopez, Quezon | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26418 | DOH | janette_villasanta@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 780 | 2023-10-25 | GX IV 6C | 831889 | Unit due for XpertCheck | Checked and cleaned cartridge bay, plunger rod, icore slot and filter Performed XpertCheck C1 failed XpertCheck Generated report | Performed XpertCheck | Unit is operational 3 modules passed xpertcheck, 1 module failed | C1/783221 Service Required | XpertCheck Kit | DR 10782 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753377320112502454/5753377320112502454_signature_28.png | Janette Villasanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5753377320112502454/5753377320112502454_signature_31.png
| 2023-11-09 | 2023-11-09 09:06 | Tagkawayan Rural Health Unit | Poblacion, Tagkawayan, Quezon Province | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26418 | DOH | rvteopy@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 621 | 2023-10-25 | GX IV 6C | 831916 | Unit due for XpertCheck | Checked and cleaned cartridge bay, plunger rod, icore slot and filter Performed Xpertcheck All passed xpertcheck Generated report | Performed XpertCheck | Unit is operational All modules working | 2 new modules. 2 passed xpertcheck | XpertCheck kit | DR 10781 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753377320113183093/5753377320113183093_signature_28.png | Rodelo V Teopy | https://www.jotform.com/uploads/macarewecare/222927978372067/5753377320113183093/5753377320113183093_signature_31.png
| 2023-11-09 | 2023-11-09 09:57 | BURAUEN DISTRICT HOSPITAL | Burauen Leyte | 8 | Leyte | Tb-dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26417 | DOH | burauenlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 427 | 2023-11-09 | GXIV | 838551 | N/a | N/A | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753321497819350336/5753321497819350336_signature_28.png | Mary Anne Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5753321497819350336/5753321497819350336_signature_31.png
| 2023-11-09 | 2023-11-09 15:18 | MR Flores Health Center | P. Herrera St. | NCR | NCR 4th - Pateros | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26416 | PBSP | paterosclinicallaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 | 2023-11-07 | GX-IV-D-6C | 804254 | Module D3 for replacement | Replaced module D3 | Performed maintenance Replaced module D3 Generated reports | All modules operational. |  | Old: D3/709295 New: D3/889948 | G 6405 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753238475692175154/5753238475692175154_signature_28.png | Jose Mari Zoilo B. De Borja, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5753238475692175154/5753238475692175154_signature_31.png
| 2023-11-09 | 2023-11-09 11:00 | Tinagacan Rural Health Unit | Brgy. Tinagacan, General Santos City, South Cotabato | 12 | South Cotabato | Tb-lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26415 | PBSP | kolleenarcega24@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 968 |  | GX-IV | 838196 | Stuck cartridge | Perform Genexpert machine Preventive maintenance | Fixed/repair stuck cartridge problem | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753095046414776856/5753095046414776856_signature_28.png | Mary Faith Kolleen Arcega | https://www.jotform.com/uploads/macarewecare/222927978372067/5753095046414776856/5753095046414776856_signature_31.png
| 2023-11-09 | 2023-11-09 10:48 | Manuel J. Santos Hospital | 554 Montilla Blvd. Butuan City | 13 | Agusan del Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 26414 | COMMERCIAL | molecularlab@mjsantos.com.ph | 24.95 | ASSAY: Xpert SARS-CoV-2, COUNT: 1421 |  | GX-4 | 831292 | None | None | Preventive Maintenance Installation of MTB Ultra Assay | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753076315323277724/5753076315323277724_signature_28.png | Raulyn Queen Tabañag | https://www.jotform.com/uploads/macarewecare/222927978372067/5753076315323277724/5753076315323277724_signature_31.png
| 2023-11-09 | 2023-11-09 10:06 | Antipolo City Health Office | M. Santos St., Antipolo | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26413 | PBSP | acho.laboratory@gmail.com | 24 |  |  | GX-XVI-D-6C | 842226 | Defective CPU | Replaced defective CPU | Replaced defective CPU with backup laptop Import assays Performed maintenance Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753056191192363812/5753056191192363812_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5753056191192363812/5753056191192363812_signature_31.png
| 2023-11-09 | 2023-11-08 15:06 | Sta. Rosa City Health Office 1 | Sta. Rosa City Hall Compound | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26412 | PBSP | santarosacho1clinlab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 900 | 2023-11-07 | GX-IV-D-6C | 804839 | Module A1 and A3 for replacement | Replaced module A1 and A3 | Performed maintenance Replaced module A1 and A3 Generated reports | Unit is operational. |  | Old: A1/910573, A3/733542 New: A1/644705, A3/788671 | G 6412 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5753045741195055508/5753045741195055508_signature_28.png | Marivic Batallones | https://www.jotform.com/uploads/macarewecare/222927978372067/5753045741195055508/5753045741195055508_signature_31.png
| 2023-11-08 | 2023-11-08 13:53 | Maria Aurora Community Hospital | Barangay 1, Maria Aurora, Aurora | 3 | Aurora | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26411 | PBSP | mariabelenth@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2983 ASSAY: Xpert HIV-1 Viral Load, COUNT: 26 | 2023-11-08 | Genexpert IV | 807565 | Module 3 error | Performed Module Replacement = B3(628831 - 210057609) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5752327053329418963/5752327053329418963_signature_28.png | Maria Belen T. Teh | https://www.jotform.com/uploads/macarewecare/222927978372067/5752327053329418963/5752327053329418963_signature_31.png
| 2023-11-08 | 2023-11-08 13:28 | Sto. Tomas City Health Office | Sto. Tomas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26410 | PBSP | nvsj12@gmail.com | 25 |  | 2023-11-07 | GX-IV-D-6C | 812786 | For module replacement and CPU replacement | Replaced module B1 and CPU | Performed maintenance Replaced module B1 Replaced CPU Generated reports | Unit is operational. |  | Old: B1/701245, Backup laptop: CPH1019F03472, CPU: 2UA6511RLL New: B1/754477 CPU: CPH0723F43754 | G 6408 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5752322382448357722/5752322382448357722_signature_28.png | VICTORIA B. MENOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5752322382448357722/5752322382448357722_signature_31.png
| 2023-11-08 | 2023-11-07 16:46 | Lung Center of the Philippines | Quezon avenue,QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26409 | MACARE-HBDC | lcpmicrolab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 155 ASSAY: Xpert SARS-CoV-2, COUNT: 47 |  | GX4 | 804205 | Due for maintenance, disabled 2,3,4 due to error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | Module 2,3,4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5752267733615141838/5752267733615141838_signature_28.png | Laurence M.Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5752267733615141838/5752267733615141838_signature_31.png
| 2023-11-08 | 2023-11-07 16:43 | Lung Center of the Philippines | Quezon avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26408 | MACARE-HBDC | lcpmicrolab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 387 ASSAY: Xpert SARS-CoV-2, COUNT: 1723 ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 ASSAY: Others, COUNT: 11 |  | GX4 | 838837 | Due for maintenance | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | B4 disabled h/w failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5752267263611994307/5752267263611994307_signature_28.png | Laurence M. Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5752267263611994307/5752267263611994307_signature_31.png
| 2023-11-08 | 2023-11-07 16:32 | Lung Center of the Philippines | Quezon avenue QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26407 | PBSP | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4583 ASSAY: Xpert SARS-CoV-2, COUNT: 3385 | 2023-11-07 | GX16 | 801080 | For module replacement C3 624937 - 662268 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance and module replacement | Operational |  | 900-0491rs | G6403 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5752266763617387371/5752266763617387371_signature_28.png | Laurence M.Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5752266763617387371/5752266763617387371_signature_31.png
| 2023-11-08 | 2023-11-08 11:20 | Isabela Municial Health Unut | Municipality of Isabela | 6 | Negros Occidental | TB DOTS | YES | INSTALLATION | CEPHEID | Emerson Bejemino |  |  |  |  | 26406 | PBSP | rhofelgareza@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1146 |  | GX4 | 838855 | Defective CPU due to damage in power supply. | Pulled out the defective CPU of Isabela MHU with SN CPH0420F07047. | Installed the borrowed CPU from FPOP Community in Iloilo with SN CPH0420F07737. | All modules are now detected by the new CPU. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5752233551129400473/5752233551129400473_signature_28.png | Rhofel Marie L. Gareza | https://www.jotform.com/uploads/macarewecare/222927978372067/5752233551129400473/5752233551129400473_signature_31.png
| 2023-11-08 | 2023-11-08 08:50 | Romana Pangan District Hosoital | Floridablanca Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26405 | PBSP | rpangandh@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 947 |  | GXIV-D | 814378 | B1 and B3 for module replacement Due for calibration | Module replacement Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Performed module replacement Gathered reports Performed functionality test | Unit is operational | 2 modules passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5752157051119431742/5752157051119431742_signature_28.png | Hayfa Mae Almario | https://www.jotform.com/uploads/macarewecare/222927978372067/5752157051119431742/5752157051119431742_signature_31.png
| 2023-11-07 | 2023-11-07 16:20 | Hagonoy Rural Health Unit | Hagonoy, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 26404 | PBSP | lanielyncastro@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1822 | 2023-11-07 | GXIV | 814158 | Modules B2, B3, B4 for replacement | #NAME? | Replaced modules | Unit is operational. |  | GeneXpert Module SN: 642246, 615348, 630858 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5751552582913800471/5751552582913800471_signature_28.png | Maricar Godoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5751552582913800471/5751552582913800471_signature_31.png
| 2023-11-07 | 2023-11-07 15:47 | Catmon Super Health Center | Sanciangco St., Catmon | NCR | NCR 3rd - Malabon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26403 | PBSP | mariates0218@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3039 ASSAY: Xpert HIV-1 Viral Load, COUNT: 128 | 2023-11-07 | GX-IV-D-6C | 838880 | Module B2 for replacement | Replaced module B2 | Performed maintenance Replaced module B2 Generated reports | Unit is operational. |  | Old: B2/863567 New: B2/716021 | G 6401 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5751526879235074455/5751526879235074455_signature_28.png | Ma. Theresa O. Nuñez | https://www.jotform.com/uploads/macarewecare/222927978372067/5751526879235074455/5751526879235074455_signature_31.png
| 2023-11-07 | 2023-11-07 13:59 | Masinloc Rural Health Unit | Madinloc Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26402 | DOH | mariapaz.ferrer08@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2160 |  | GXIV-D | 831608 | A4 for re-calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | Used 1 xpertcheck kit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5751470988824345781/5751470988824345781_signature_28.png | Maria Paz E. Ferrer | https://www.jotform.com/uploads/macarewecare/222927978372067/5751470988824345781/5751470988824345781_signature_31.png
| 2023-11-07 | 2023-11-07 13:18 | Cebu TB Reference Lab | SRP-CTRL | 7 | Cebu | Laboratory | NO | INSTALLATION | CEPHEID | Raymond Laurito |  |  |  |  | 26401 | PBSP | asmantiquilla.pbsp@gmail.com | 28 |  |  | Gx 4 | 808620 | Defective CPU | New CPU installed | New CPU installed | Systems operational, white lines in down side of monitor |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5751447148814505464/5751447148814505464_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5751447148814505464/5751447148814505464_signature_31.png
| 2023-11-07 | 2023-11-07 13:28 | Norzagaray Rural Health Unit | Norzagaray, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 26400 | PBSP | chelseajempolicarpio@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1237 | 2023-11-07 | GXIV | 838167 | Module B2 for replacement | #NAME? | Replaced B2 | Unit is operational. |  | GeneXpert Module SN: 737171 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5751444470121331576/5751444470121331576_signature_28.png | Chelsea Policarpio | https://www.jotform.com/uploads/macarewecare/222927978372067/5751444470121331576/5751444470121331576_signature_31.png
| 2023-11-07 | 2023-11-07 11:12 | Sarangani Islands Rural Health Unit | Mabila, Sarangani Islands | 11 | Davao Occidental | TB-DOTS | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26399 | DOH | lovitosquennie26@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 173 | 2023-08-03 | GX-IV 4D | 837614 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration UPS intended to the GeneXpert System is Defective All filters are damaged and subjected for replacement |  | Xpertcheck CE 5 | DR 10446 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5751438822292042595/5751438822292042595_signature_28.png | Marilyn A. Astang | https://www.jotform.com/uploads/macarewecare/222927978372067/5751438822292042595/5751438822292042595_signature_31.png
| 2023-11-07 | 2023-11-07 12:44 | Dalaguete Rural Health Unit | Poblacion Dalaguete Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 26398 | PBSP | ajreef5517@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1622 |  | Gx4 | 815464 | Module 4 SN 210073850 encountered Error of 2032 during 30 mins of running | Replaced new module | Replaced new module SN 210027279 Generate IQR, Sys Log, Gxx files, Assay Stat | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5751420973522134477/5751420973522134477_signature_28.png | Allyza Rivera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5751420973522134477/5751420973522134477_signature_31.png
| 2023-11-07 | 2023-11-07 11:38 | Valenzuela Social Hygiene Clinic | G. Lazaro Rd. Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | YES | INSTALLATION | CEPHEID | John Monsour Doctor |  |  |  |  | 26397 | DOH | reyesmarlette030319@gmail.c9m | 26 |  |  | GX-IV-D-10C | 110020699 | For Installation | Installed the GeneXpert System | Installed the GeneXpert System: Installed GeneXpert machine, printer, UPS desktop and monitor Installed HIV-VL assay Functionality Test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5751395676816359384/5751395676816359384_signature_28.png | Marlette Delino | https://www.jotform.com/uploads/macarewecare/222927978372067/5751395676816359384/5751395676816359384_signature_31.png
| 2023-11-07 | 2023-11-07 11:03 | Sub8c Rural Health Unit | Subic, Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26396 | PBSP | sanchezmacristina313@gmail.comb | 29 | ASSAY: Xpert MTB/RIF, COUNT: 51 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2967 ASSAY: Xpert HIV-1 Viral Load, COUNT: 67 |  | GXIV-D | 812788 | B2 for module replacement | Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5751357517913018824/5751357517913018824_signature_28.png | Ma. Cristina D. Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/5751357517913018824/5751357517913018824_signature_31.png
| 2023-11-06 | 2023-11-06 14:19 | Cristino Paragas Memorial Hospital | Quiniput, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26395 | PBSP | sarahtabudjul@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 507 | 2023-11-06 | GX 4 | 815269 | Module 3 door cannot close | Re align module valve and Preventive Maintenence | Re align module valve and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750677679127076237/5750677679127076237_signature_28.png | Sarah Gail A. Tabudjul | https://www.jotform.com/uploads/macarewecare/222927978372067/5750677679127076237/5750677679127076237_signature_31.png
| 2023-11-06 | 2023-11-06 14:37 | Moalboal Rural Health Unit | Poblacion Moalboal Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 26394 | PBSP | roavaul@gmail.com | 638 |  | 2023-11-06 | Gx4 | 838892 | Display Monitor defective | Replaced new monitor | Replace new monitor Thinkvision Flat SN# :V5FTH187 | All modules are operational with New ThinkVision SN V5FTH187 |  | Thinkvision Flat Panel Monitor SN V5FTH187 | SN # V5FTH187 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750622082666657124/5750622082666657124_signature_28.png | Ava Anciano Gako | https://www.jotform.com/uploads/macarewecare/222927978372067/5750622082666657124/5750622082666657124_signature_31.png
| 2023-11-06 | 2023-11-06 14:02 | San Antonio District Hospital | San Mariano, San Antonio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26393 | PBSP | garciaadonisanthony@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 588 | 2023-11-06 | Genexpert IV | 812832 | Module 4 error | Performed Module Replacement B4(685801 - 681375) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750602618222516235/5750602618222516235_signature_28.png | Ana Charize G. Manalastas | https://www.jotform.com/uploads/macarewecare/222927978372067/5750602618222516235/5750602618222516235_signature_31.png
| 2023-11-06 | 2023-10-31 13:10 | Pangil Rural Health Unit | Pakil-Pangil-Mabitac | 4A | Rizal | Laboratory | NO | MAINTENANCE Checkup | CEPHEID | John Monsour Doctor |  |  |  |  | 26392 | PBSP | tacile10@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 399 |  | GX-IV-D-6C | 807661 | Error 5007 | Checked modules | Performed maintenance Checked modules Generated reports | All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750474473613800815/5750474473613800815_signature_28.png | Cecile Sagritalo | https://www.jotform.com/uploads/macarewecare/222927978372067/5750474473613800815/5750474473613800815_signature_31.png
| 2023-11-06 | 2023-11-06 09:37 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 26391 | DOH | laureanjo@yahoo.com | 28 |  |  | GX4 | 838527 | SITE visit | Clean fan filter, plunger rod maintenance, clean module bay | Clean fan filter, plunger rod maintenance, clean module bay | System operational.. waiting for module D3 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750445776376434311/5750445776376434311_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5750445776376434311/5750445776376434311_signature_31.png
| 2023-11-06 | 2023-10-25 09:57 | Tanza Rural Health Unit | Municipal Compound, Tanza, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26390 | PBSP | gutierrezrhian@gmail.com | 24 |  | 2023-10-16 | GXIV-D | 812835 | Module 1 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 1 Generate Reports | All Modules are Operational |  | GX Module | G 6355 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750429233616017053/5750429233616017053_signature_28.png | MARY ANN C. GUTIERREZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5750429233616017053/5750429233616017053_signature_31.png
| 2023-11-06 | 2023-10-13 13:44 | Carmona Rural Health Unit | Burgos St., Carmona, Cavite | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26388 | DOH | genexpert.carmona@gmail.com | 25 |  | 2023-10-09 | GXIV-D | 821216 | UPS is Defective | Replacement | Replaced UPS Checked Functionality of UPS Generate Reports | All Modules are Operational |  | UPS 1.5kVA | DR 10722 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750428313611276525/5750428313611276525_signature_28.png | Noreen Fortu | https://www.jotform.com/uploads/macarewecare/222927978372067/5750428313611276525/5750428313611276525_signature_31.png
| 2023-11-06 | 2023-10-17 13:16 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Rd, Las Piñas City | NCR | NCR 4th - Las Piñas | Blood Bank | NO | MAINTENANCE Troubleshoot | HAEMONETICS | Brenguiler Bagbaguen |  |  |  |  | 26389 | MACARE-HBDC | uphdmc_lab@yahoo.com | 24 |  |  | MCS+ | 11J220 | Centrifuge Locking Alarm | Troubleshoot | Performed Maintenance Checked Transducers Cleaned Sensors Checked Functionality | UPS is Defective | Centrifuge Locking for Observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750428313614463100/5750428313614463100_signature_28.png | MARISSA MACHADO | https://www.jotform.com/uploads/macarewecare/222927978372067/5750428313614463100/5750428313614463100_signature_31.png
| 2023-11-06 | 2023-10-18 09:37 | Trece Martires City Health Office | Governor's Drive, Trece Martires, Cavite | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26387 | PBSP | loricavivian@gmail.com | 22 |  | 2023-10-16 | GXIV-D | 838160 | Module 1 with Temperature Error Unit is due for XpertCheck | Module Replacement XpertCheck | Performed Maintenance Replaced Module 1 Performed Self-test Performed XpertCheck Generate Reports | All Modules are Operational | 806865 -> 210100586 | 900-0508 | G 6354 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750428313615365830/5750428313615365830_signature_28.png | Vivian D. Lorica | https://www.jotform.com/uploads/macarewecare/222927978372067/5750428313615365830/5750428313615365830_signature_31.png
| 2023-11-06 | 2023-10-25 09:57 | Tanza Rural Health Unit | Municipal Compound, Tanza, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26387 | PBSP | gutierrezrhian@gmail.com | 24 |  | 2023-10-16 | GXIV-D | 812835 | Module 1 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 1 Generate Reports | All Modules are Operational |  | GX Module | G 6355 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5750428313617199631/5750428313617199631_signature_28.png | MARY ANN C. GUTIERREZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5750428313617199631/5750428313617199631_signature_31.png
| 2023-11-05 | 2023-11-03 13:21 | Commonwealth Health Center | 2 katuparan st, commonwealth, quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26386 | DOH | commonwealthclinical@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 243 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2808 |  | GX I V 6C | 831615 | Module A1 frequent error | Performed PM Checked system log No current error on module A1 system log Checked module a1 report Error 5007 frequent problem found Cleaned optics thoroughly Advised to do Reagent run Module 1 passed reagent run Made A1 available again Generated report | Troubleshoot A1 | Unit is operational 3 modules working. Module 2 already reported waiting for replacment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5749977008324268956/5749977008324268956_signature_28.png | Iron Capistrano | https://www.jotform.com/uploads/macarewecare/222927978372067/5749977008324268956/5749977008324268956_signature_31.png
| 2023-11-05 | 2023-11-03 10:36 | National Government Center Super Health Center | MRB Compound, Pilot Drive, Quezon City | NCR | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26385 | DOH | tmcolisao@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1634 |  | GX IV 6C | 831586 | Module 2 frequent error | Performed PM Checked system log Run XpertCheck Module passed xpert check Generated report | Troubleahoot module 2 | Unit is operational Module 2 under monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5749977008325416085/5749977008325416085_signature_28.png | Trecia Marie Colisao | https://www.jotform.com/uploads/macarewecare/222927978372067/5749977008325416085/5749977008325416085_signature_31.png
| 2023-11-05 | 2023-11-05 20:07 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank Service Unit | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 26384 | COMMERCIAL | vsmmcbloodservices@gmail.com | 25 |  |  | MCS+ | 02A090 | Procedure UPP, encounter noisy in the bowl during 1st draw UPP bowl problem | Diagnostics procedure in Centrifuge motor assembly Check kit of UPP install in the machine | Check UPS supply, check centrifuge and centrifuge assembly | MCS+ unit are operational during change of another kit |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5749955543913294539/5749955543913294539_signature_28.png | Elyn Resuento, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5749955543913294539/5749955543913294539_signature_31.png
| 2023-11-04 | 2023-11-03 10:57 | Ilocos Training and Regional Medical Center | San Fernando, La Union | 1 | La Union | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26383 | WHO | culomariejennifer08@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 335 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1708 ASSAY: Others, COUNT: 687 ASSAY: Xpert SARS-CoV-2, COUNT: 441 | 2023-11-03 | GX4 10C | 848249 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 210067419>>> 210056220 - Generated reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210056220 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5748872764406802705/5748872764406802705_signature_28.png | Marie Jennifer P. Culo | https://www.jotform.com/uploads/macarewecare/222927978372067/5748872764406802705/5748872764406802705_signature_31.png
| 2023-11-04 | 2023-11-03 11:01 | Ilocos Training and Regional Medical Center | San Fernando, La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26382 | PBSP | culomariejennifer08@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 103 ASSAY: Xpert HPV, COUNT: 357 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2255 ASSAY: Xpert SARS-CoV-2, COUNT: 427 ASSAY: Xpert HIV-1 Viral Load, COUNT: 174 |  | GX4 6C | 707524 | Due for Maintenance | #NAME? | Performed Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5748870684407301766/5748870684407301766_signature_28.png | Marie Jennifer P. Culo | https://www.jotform.com/uploads/macarewecare/222927978372067/5748870684407301766/5748870684407301766_signature_31.png
| 2023-11-04 | 2023-11-04 11:29 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank Service Unit | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 26381 | COMMERCIAL | vsmmcbloodservices@gmail.com | 25 |  |  | MCS+ | 02A090 | Machine will turn off during start up and self test | Transfer UPS to other UPS MCS+ are operational UPS need to replaced as soon as possible | Check up fuse and voltage supply AC 220 Check UPS output 0 Volts AC Transfer UPS to other UPS unit | MCS+ are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5748796573915974669/5748796573915974669_signature_28.png | Elyn Resuento , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5748796573915974669/5748796573915974669_signature_31.png
| 2023-11-04 | 2023-11-04 11:03 | Calamba Medical Center | Calamba Crossing,Calamba laguna | 4A | Laguna | Molecular | YES | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 26380 | COMMERCIAL | cmc.rtpcr@gmail.com | 20 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HBV VL, COUNT: 20 ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF Ultra, COUNT: 600 | 2023-11-03 | GX2 | 812916 | For module replacement b2 686424 - 210096466, b1 encounter continous h/w failed. | Performed module replacement Cleaned plunger,cartridge bay,filter Gathered reports Performed selt test multiple times but b1 error still occur. Test run unit | Module replacement | B2 operational ready to use, b1 h/w failed for reporting |  | 900-0508 | 10789 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5748771671025773197/5748771671025773197_signature_28.png | Andrea Gail Orgesvik | https://www.jotform.com/uploads/macarewecare/222927978372067/5748771671025773197/5748771671025773197_signature_31.png
| 2023-11-03 | 2023-11-03 14:38 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 26379 | MACARE-HBDC | baluyotjeric@gmail.com | 26 |  |  | MCS + | 20168 | Defective and pixelated display | Swap lcd temporary to mcs s/n 01M201. Lcd. Swap lcd display Test all button function. Checked program protocol. Performed functionality test. | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5748040397619147781/5748040397619147781_signature_28.png | Jeric C. Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5748040397619147781/5748040397619147781_signature_31.png
| 2023-11-03 | 2023-11-03 11:59 | Tacloban City Health Office | Tacloban City | 8 | Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26378 | DOH | marteetingzon20@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 43 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1289 | 2023-11-03 | GXIV | 838421 | N/a | N/a | #NAME? | Unit is Operational (B1 and B4 disabled - Waiting for module replacement ) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5747940889319318655/5747940889319318655_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5747940889319318655/5747940889319318655_signature_31.png
| 2023-11-03 | 2023-11-03 09:27 | Binalbagan Municipal Health Unit | Binalbagan, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | CALIBRATION ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 26377 | DOH | anne_galilea@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 951 |  | GX-4 | 838280 | E5006 on module B1 | - | #NAME? | - B1 passed XPC, for observation for recurrence of E5006. - 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5747850580415997143/5747850580415997143_signature_28.png | XERXANNE A. GALILEA | https://www.jotform.com/uploads/macarewecare/222927978372067/5747850580415997143/5747850580415997143_signature_31.png
| 2023-11-02 | 2023-11-02 15:31 | Calamba Medical Center | Calamba, Laguna | 4A | Laguna | Molecular Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Harold Toco |  |  |  |  | 26376 | COMMERCIAL | cmc.rtpcr@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 604 ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert HBV VL, COUNT: 20 ASSAY: Xpert CT/NG, COUNT: 2 | 2023-11-02 | Gx-II | 812916 | Error 4017 Temperature related issue. | Performed maintenance..gather data.run functionality test.Reliability test completed. | Performed unit maintenance.Clean fan filter..performed plunger maintenance. Clean cartridge bay.. Run reliability test | Module B1 operational but for monitoring. |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5747210096711722018/5747210096711722018_signature_28.png | Andrea Gail Orgesvik | https://www.jotform.com/uploads/macarewecare/222927978372067/5747210096711722018/5747210096711722018_signature_31.png
| 2023-11-02 | 2023-10-26 10:14 | Philippine Heart Center | East Avenue,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26375 | PBSP | molecularlab@phc.gov.ph | 26 |  | 2023-10-23 | GX4 | 834443 | Due for xpertcheck and maintenance | Performed xpertcheck Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Mainyenance and xpertcheck | All passed in xpertcheck |  | Xpertcheck | G6382 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5747175316012434679/5747175316012434679_signature_28.png | Kenneth Melvin N. Casidsid | https://www.jotform.com/uploads/macarewecare/222927978372067/5747175316012434679/5747175316012434679_signature_31.png
| 2023-11-02 | 2023-10-26 10:17 | Philippine Heart Center | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26374 | MACARE-HBDC | molecularlab@phc.gov.ph | 28 |  |  | GX4 | 804260 | Due for maintenance, | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Maintenance | 2 module operational | C2 and C4 disabled encounter intermittent temp error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5747174866017680612/5747174866017680612_signature_28.png | Kenneth Melvin N. Casidsid | https://www.jotform.com/uploads/macarewecare/222927978372067/5747174866017680612/5747174866017680612_signature_31.png
| 2023-11-02 | 2023-10-25 13:16 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Laboratory (Koica) | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26373 | PBSP | nilabis.pbsp@gmail.com | 27 |  | 2023-10-23 | GX4 R1 | 704584 | For module replacement B4 761791 - 210068678 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | Operational |  | 900 0491RS | G6379 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5747174206013476481/5747174206013476481_signature_28.png | Norman Labis | https://www.jotform.com/uploads/macarewecare/222927978372067/5747174206013476481/5747174206013476481_signature_31.png
| 2023-11-02 | 2023-11-02 08:19 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Molecular Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 26372 | COMMERCIAL | cduhlabresults2@yahoo.com | 19 | ASSAY: Xpert HCV VL, COUNT: 1299 |  | GX4 | 110008524 | Error 2097 in HCV Assay cartridge problem | Performed maintenance and plunger , Icore clean up Required end user to report in jotform regarding the persistence Error 2097 | Replaced new assay in HCV Viral Load and install with new assay Performed Icore clean up and plunger maintenance | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5746941233915722660/5746941233915722660_signature_28.png | Adrian Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5746941233915722660/5746941233915722660_signature_31.png
| 2023-10-31 | 2023-10-27 11:21 | Tinagacan Rural Health Unit | Brgy. Tinagacan, General Santos City | 12 | South Cotabato | Tb-Lab | YES | REPAIR TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 26371 | PBSP | kolleenarcega24@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 963 |  | Gx-IV | 838196 | Module 4 stuck cartridge Screw Stand problem | Perform Genexpert machine Preventive maintenance | Reschedule stuck cartridge repair due to screw problem | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5745700904314570841/5745700904314570841_signature_28.png | Mary Faith Kolleen Arcega | https://www.jotform.com/uploads/macarewecare/222927978372067/5745700904314570841/5745700904314570841_signature_31.png
| 2023-10-31 | 2023-10-31 14:59 | St. Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 26370 | COMMERCIAL | biomed.qc@stlukes.com.ph | 29 | ASSAY: Xpert CT/NG, COUNT: 266 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert HPV, COUNT: 220 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3200 ASSAY: Xpert HCV VL, COUNT: 5 |  | GX4 | 807155 | As per end user they encounter 2008 error | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5745452643097647985/5745452643097647985_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5745452643097647985/5745452643097647985_signature_31.png
| 2023-10-31 | 2023-10-31 10:27 | Lumban Rural Health Unit | Gen Luna St., Lumban | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26369 | PBSP | rizalynesacdalan@gmail.com | 29 |  |  | GX-IV-D-6C | 838148 | Stuck cartridge on module B4 | Manual removal of stuck cartridge | Performed maintenance Manual removal of stuck cartridge Checked module B3 Generated reports | Unit is operational. Module B3/630625 not operational, reported already. Room temperature is high due to faulty aircon. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5745300719104816442/5745300719104816442_signature_28.png | Rizalyne Sacdalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5745300719104816442/5745300719104816442_signature_31.png
| 2023-10-31 | 2023-10-25 16:06 | Pilar Rural Health Unit | Pilar | 5 | Sorsogon | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 26368 | DOH | manalangmark88@gmail.com | 24 |  |  | GX4 | 838266 | CPU hard to press-open Keyboard mot turning off when shutdown | Check power option setup Check ports placement Generate reports | Performed troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5745239125211766226/5745239125211766226_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5745239125211766226/5745239125211766226_signature_31.png
| 2023-10-31 | 2023-10-27 17:14 | Tabaco City Health Unit | Tabaco | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26367 | PBSP | elenocbina@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3278 |  | GX4 | 804257 | D2/210089677 for module replacement | Replaced module: D2 - 210089677>>779135 Cleaned cartridge bays, plunger rod, icore, fan Generate reports | Replaced D2 module Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5745236295217744362/5745236295217744362_signature_28.png | Eleanor Ocbina | https://www.jotform.com/uploads/macarewecare/222927978372067/5745236295217744362/5745236295217744362_signature_31.png
| 2023-10-31 | 2023-10-31 08:39 | Patnanungan Rurahl Health Unit | Patnanungan, Quezon | 4A | Quezon | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Cienne Mae Valerio |  |  |  |  | 26366 | DOH | rhu_patnanungan@yahoo.com.ph | 26 |  | 2023-10-20 | GXIV-4L-6C | 831611 | Defective CPU | Replaced Defective CPU with Back up Laptop Old CPU 2UA8471W5V Back Up Laptop 5CG5040H6L 3 Spare Xpertcheck for Module Troubleshooting | Perfomed Maintenance Cleaned Cartridge Bay Plunger Rod Maintenance Cleaned Filter Gather Reports | Unit is Operational |  | Back up Laptop | 10765 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5745229414761434344/5745229414761434344_signature_28.png | John Michael Bantucan | https://www.jotform.com/uploads/macarewecare/222927978372067/5745229414761434344/5745229414761434344_signature_31.png
| 2023-10-28 | 2023-10-27 15:51 | Sariaya Rural Health Unit | Sariaya, Quezon Province | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26365 | DOH | marielinnetabril@yahoo.com | 24 |  |  | GX IV 6C | 831890 | Unit due for xpertcheck | Performed PM Performed XpertCheck Generated report | Performed xpertcheck | Unit is operationaL Module 2 and 3 service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5742670576218260396/5742670576218260396_signature_28.png | Marie linnet abril | https://www.jotform.com/uploads/macarewecare/222927978372067/5742670576218260396/5742670576218260396_signature_31.png
| 2023-10-27 | 2023-10-27 13:44 | Eversly Childs Sanitarium | Jagobiao, mandaue city | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 26364 | PBSP | alimurungjane@gmail.com | 28 |  |  | Gx 4 | 805270 | Defective Module A4 753612 | Install New Module A4 694399/ clean fan filter | Install New Module A4 694399 | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741983121026260863/5741983121026260863_signature_28.png | Juanita V. Alimurung | https://www.jotform.com/uploads/macarewecare/222927978372067/5741983121026260863/5741983121026260863_signature_31.png
| 2023-10-27 | 2023-10-27 14:01 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 26363 | DOH | cauayanmho@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1266 |  | GX-4 | 838477 | E1002 and E4014 encountered on modules B1 and B2. | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741971137228054434/5741971137228054434_signature_28.png | Cauayan Municipal Health Office | https://www.jotform.com/uploads/macarewecare/222927978372067/5741971137228054434/5741971137228054434_signature_31.png
| 2023-10-27 | 2023-10-27 14:10 | Antipolo City Health Office | M. Santos St. | 4A | Rizal | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26362 | PBSP | acho.laboratory@gmail.com | 25 |  |  | GX-16-D-6C | 842226 | CPU always shutting down after booting | Checked CPU | Checked the connections Reseat internal connections Checked RAM Checked HDD Swap RAM to other slot | Unit is not operational. CPU SN: 2UA3520VTR |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741965317711872412/5741965317711872412_signature_28.png | Gwendolyn C. Batingan | https://www.jotform.com/uploads/macarewecare/222927978372067/5741965317711872412/5741965317711872412_signature_31.png
| 2023-10-27 | 2023-10-27 11:30 | Quezon Medical Center | Quezon Ave, Lucena | 4A | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26361 | PBSP | joannpalermo0613@gmail.com | 29 |  |  | GX IV 6C | 804692 | GX machine for packing and CPU troubleshooting | Checked CPU Checked power cable Checked and cleaned ram, and ssd Transferred Ram Problem persists CPU for reporting Packed unit for transfer to new laboratory | Troubleshoot cpu, packed unit for transfer to new laboratory | Gx is operational. Cpu not working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741893790157657478/5741893790157657478_signature_28.png | Jo-Ann Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5741893790157657478/5741893790157657478_signature_31.png
| 2023-10-27 | 2023-10-27 12:09 | Tanay Rural Health Unit | Manila E Rd, Tanay | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26360 | PBSP | maryjunnette@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1470 ASSAY: Xpert HIV-1 Viral Load, COUNT: 21 | 2023-10-23 | GX-IV-D-6C | 807593 | Module A1 for replacement | Replaced module A1 | Performed maintenance Replaced module A1 Generated reports | All modules operational. |  | Old: A1/628393 New: A1/725380 | G 6378 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741893527713752329/5741893527713752329_signature_28.png | Mary Junnette Fulgueras | https://www.jotform.com/uploads/macarewecare/222927978372067/5741893527713752329/5741893527713752329_signature_31.png
| 2023-10-27 | 2023-10-27 11:38 | PILAR RURAL HEALTH UNIT | PILAR SIARGAO | 10 | Surigao del Norte | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26359 | PBSP | DexterJohnTDotillos@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 390 |  | GX-II | 834833 | FOR MODULE B1 REPLACEMENT | #NAME? | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE B1 REPLACEMENT: DEFECTIVE SN: 787231 NEW MODULE SN: 791649 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741879726164692506/5741879726164692506_signature_28.png | Dexter John T. Dotillos | https://www.jotform.com/uploads/macarewecare/222927978372067/5741879726164692506/5741879726164692506_signature_31.png
| 2023-10-27 | 2023-10-27 10:58 | Pililla Rural Health Unit | MA Roxas St., Pililla | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26358 | DOH | tbdotspililla@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 703 |  | GX-IV-D-6C | 821293 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. Module B2/687232 not operational, reported already. Modules B1 and B3 passed. Module B4 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741865105389394272/5741865105389394272_signature_28.png | Karren Abegail Olaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5741865105389394272/5741865105389394272_signature_31.png
| 2023-10-26 | 2023-10-26 13:40 | Macalelon Rural Health Unit | Peñalosa St., Macalelon, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26357 | DOH | cheyster7388@gmail.com | 28 |  |  | GX IV 6C | 831636 | To perform additional troubleshooting on module 4 | Performed PM Checked and cleaned icore slot, filter, plunger rod and cartridge bay Checked modules valve Reseat Gateway board Generated report | Performed additional reoubleshooting | Unit is operational | Module 4 error 2035 again after troubleshooting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741361610151776280/5741361610151776280_signature_28.png | Cheyster O. Agustin | https://www.jotform.com/uploads/macarewecare/222927978372067/5741361610151776280/5741361610151776280_signature_31.png
| 2023-10-26 | 2023-10-24 12:06 | Masambong Health Center | Capoas St., Masambong, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26356 | PBSP | michellebonrostro@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1668 |  | GX IV 6C | 814160 | Module 4 frequent error | Performed PM Troubleshoot module 4 Checked system log Performed reliability test; HW Failed module 4 Generated report | Troubleshoot module 4 | Unut is operational. 3 modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741360190159688831/5741360190159688831_signature_28.png | Cristy cortezano | https://www.jotform.com/uploads/macarewecare/222927978372067/5741360190159688831/5741360190159688831_signature_31.png
| 2023-10-26 | 2023-10-26 16:52 | Basud Rural Health Unit | Maharlika Highway, Poblacion 2, Basud | 5 | Camarines Norte | RTDL Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26355 | PBSP | gbechano@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1077 |  | GX4 | 807600 | B4/210070836 for Module Replacement | Replaced Module: B4 - 210070836>>210101552 | Replaced B4 Module Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741198922324909979/5741198922324909979_signature_28.png | Grace B. Echano | https://www.jotform.com/uploads/macarewecare/222927978372067/5741198922324909979/5741198922324909979_signature_31.png
| 2023-10-26 | 2023-10-26 14:57 | San Fernando Rural Health Unit I | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26354 | PBSP | celpineda1973@yahoo.com | 29 |  | 2023-10-26 | Genexpert IV | 804837 | Module 2,4 error | Performed Module Replacement (690343 - 628383) (631936 - 210019566) (666782 - 638450) | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Modules | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741130714428918332/5741130714428918332_signature_28.png | Celia T.Pineda | https://www.jotform.com/uploads/macarewecare/222927978372067/5741130714428918332/5741130714428918332_signature_31.png
| 2023-10-26 | 2023-10-26 14:34 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 26353 | DOH | laguitannataniel@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 912 |  | GXIV | 831846 | Module B1, B2 for reXpertCheck | #NAME? | Performed XpertCheck | Unit is operational. B1: B2: |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741122512521947964/5741122512521947964_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5741122512521947964/5741122512521947964_signature_31.png
| 2023-10-26 | 2023-10-26 14:11 | Caraga Regional Hospital | Brgy. Washington, Rizal Street, Surigao City | 13 | Surigao del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 26352 | PBSP | abpabellan.pbsp@gmail.com | 23.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5501 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 |  | GX-4 | 815489 | None | None | Replacing Module D2 with SN: 798340 to SN: 210044001 Preventive Maintenance | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741112311913669266/5741112311913669266_signature_28.png | Marie June Lastimosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5741112311913669266/5741112311913669266_signature_31.png
| 2023-10-26 | 2023-10-26 13:42 | Tomas Claudio Health Center | Corner E. Quirino Street, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26351 | DOH | arleen_p_alconera@yahoo.com | 25.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 801 |  | GX-IV 4D | 838492 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Three Modules are operational | Module B2 disabled due to error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741084173718502815/5741084173718502815_signature_28.png | Arleen P. Alconera,. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5741084173718502815/5741084173718502815_signature_31.png
| 2023-10-26 | 2023-10-26 13:36 | Binan City Health Office 1 | Binan City, Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 26350 | PBSP | annefriam.almazora@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 149 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1848 | 2023-10-23 | GXIV | 812780 | Unit for PM. For module replacement of module B4. | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational |  | GX module B4 (210060063) in replacement to defective B4 (698927). | G6375 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741082402482342205/5741082402482342205_signature_28.png | Criselda F.Gallo | https://www.jotform.com/uploads/macarewecare/222927978372067/5741082402482342205/5741082402482342205_signature_31.png
| 2023-10-26 | 2023-10-26 13:17 | San Carlos City Health Office | S. Carmona-Ylagan St., San Carlos City, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 26349 | PBSP | minozac450@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2335 |  | GX-4 | 805033 | Defective modules 687124/B3. | #NAME? | - Replaced modules 687124/B3 and with new modules with SN 668567/B3. - Generated Pre and Post IQRs. | All 4 modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741075576037695118/5741075576037695118_signature_28.png | Christine Miñoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5741075576037695118/5741075576037695118_signature_31.png
| 2023-10-26 | 2023-10-26 12:13 | Cabuyao City Health Office 2 | Cabuyao, Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 26348 | DOH | irahchristine.scd@gmail.com | 25 |  |  | GXIV | 821003 | Unit for PM. For Xpertcheck of modules A1 and A3. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Modules A1 and A3, passed Xpertcheck. | Used 2 spare Xpertcheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741034482483563946/5741034482483563946_signature_28.png | Irah Christine Dimatulac | https://www.jotform.com/uploads/macarewecare/222927978372067/5741034482483563946/5741034482483563946_signature_31.png
| 2023-10-26 | 2023-10-26 11:18 | Mabalacat Rural Health Unit | Mabalacat Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26347 | PBSP | perezriyadh@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1436 ASSAY: Xpert HIV-1 Viral Load, COUNT: 61 |  | GXIV-D | 807564 | Hard drive not detected Slow Boot up( 15-20 mins) to GX application | Troubleshooting Preventive maintenance | Performed preventive maintenance Performed RAM reset and RAM cleaning Performed Hard drive reset Performed hard drive defragment Gathered reports Performed functionality test | Unit is operational | For CPU replacement ( Slow CPU) since 2015 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741015990129924675/5741015990129924675_signature_28.png | Riyadh Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5741015990129924675/5741015990129924675_signature_31.png
| 2023-10-26 | 2023-10-26 11:26 | Carcar City Health Office | Valladolid Carcar City | 7 | Cebu | Tb Laboratory | NO | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 26346 | DOH | chachaska.lp14@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1649 |  | Gx4 | 838283 | 2008 for B1, B2 B4 | Orient end user for proper sampling preparation and reagents mixing. Generate gxx files, patient report | Performed Maintenance clean Icore, Valve, Plunger No more filter damage all | All units are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5741004313711792397/5741004313711792397_signature_28.png | Lucille Mae A. Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5741004313711792397/5741004313711792397_signature_31.png
| 2023-10-26 | 2023-10-26 09:52 | Camp Capinpin Army Station Hospital | Camp Capinpin Compound | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26345 | DOH | ccashlaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 148 |  | GX-IV-D-6C | 831896 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740958780217531267/5740958780217531267_signature_28.png | Maquilan, Khimberly C | https://www.jotform.com/uploads/macarewecare/222927978372067/5740958780217531267/5740958780217531267_signature_31.png
| 2023-10-26 | 2023-10-25 08:33 | Mangaldan Rural Health Unit | Mangaldan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26344 | DOH | arleneabad32@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 576 |  | GX4 | 838288 | Due for ReXpertCheck | #NAME? | Performed ReXpertCheck | Unit is Operational | Module B2 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740892554409065286/5740892554409065286_signature_28.png | Arlene U. Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5740892554409065286/5740892554409065286_signature_31.png
| 2023-10-26 | 2023-10-25 09:48 | San Jacinto Rural Health Unit | San Jacinto, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26343 | DOH | rhusanjacinto16@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 328 |  | GX4 | 838518 | Due for Troubleshooting | #NAME? | Performed Troubleshooting | Unit is operational | For observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740891544408703320/5740891544408703320_signature_28.png | Shara Mae Dela Masa | https://www.jotform.com/uploads/macarewecare/222927978372067/5740891544408703320/5740891544408703320_signature_31.png
| 2023-10-25 | 2023-10-25 19:25 | Tuburan Rural Health Unit | Poblacion Tuburan Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 26342 | PBSP | marianmelodyrabaca37@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 337 |  | GX4 | 815485 | B2 Defective SN # 791790 error 4012 repeated during start up and running | Replaced new module SN # 717505 | Replacement of New module SN# 717505 | All module are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740427583915975342/5740427583915975342_signature_28.png | Mariel Reguero, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5740427583915975342/5740427583915975342_signature_31.png
| 2023-10-25 | 2023-10-25 14:58 | Catmon Health Center | Catmon, Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 26341 | PBSP | mariathes0218@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-10-25 | GXIV | 838880 | Module B1, B3, B4, for replacement | #NAME? | Replaced modules | Unit is operational. |  | GeneXpert Module SN: 653728, 710644, 655300 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740265891311725055/5740265891311725055_signature_28.png | Maria Theresa Nunez | https://www.jotform.com/uploads/macarewecare/222927978372067/5740265891311725055/5740265891311725055_signature_31.png
| 2023-10-25 | 2023-10-25 14:45 | Antipolo City Health Office | M. Santos St. | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26340 | PBSP | acho.laboratory@gmail.coml | 25 | ASSAY: Xpert CT/NG, COUNT: 80 ASSAY: Xpert MTB/RIF Ultra, COUNT: 9475 ASSAY: Xpert HIV-1 Viral Load, COUNT: 625 | 2023-10-23 | GX-XVI-D-6C | 842226 | Module A2 for replacement | Replaced module A2 | Performed maintenance Replaced module A2 Generated reports | All modules operational. |  | Old: A2/885372 New: A2/636902 | G 6377 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740264858355828570/5740264858355828570_signature_28.png | Gwendolyn C. Batingan | https://www.jotform.com/uploads/macarewecare/222927978372067/5740264858355828570/5740264858355828570_signature_31.png
| 2023-10-25 | 2023-10-25 13:58 | Davao Chest Center | Villa Abrille St. Poblacion, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26339 | PBSP | dcclabdept@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF, COUNT: 2222 |  | GX-IV 4D | 804835 | Module A1 and A2 for replacement | Replaced the Defective Module A1 SN: 645083 with new Module SN: 915216 Replaced the Defective Module A2 SN: 610272 with new Module SN 702694 Generate IQR, Assay Stat and System Log | Perform Module Replacement | Three Modules are operational | Modules A3 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740233144713423916/5740233144713423916_signature_28.png | CRIZZ VALEREY BAUTISTA | https://www.jotform.com/uploads/macarewecare/222927978372067/5740233144713423916/5740233144713423916_signature_31.png
| 2023-10-25 | 2023-10-25 13:35 | Pulupandan Rural Health Unit | Poblacion, Pulupandan, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Emerson Bejemino |  |  |  |  | 26338 | WHO | pulupandantbdots@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 482 |  | GX-4 | 847255 | Defective modules 910236/B2 and 910237/B3 | #NAME? | - Replaced modules 910236/B2 and 910237/B3 with new modules with SN 886683/B2 and 720511/B3. - Generated Pre and Post IQRs. | All 4 modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740214725217756052/5740214725217756052_signature_28.png | Arriane Gaye Tambolero | https://www.jotform.com/uploads/macarewecare/222927978372067/5740214725217756052/5740214725217756052_signature_31.png
| 2023-10-25 | 2023-10-25 12:02 | Rizal Medical Center | Pasig Blvd, Pasig City | NCR | NCR 2nd - Pasig | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Gian Paolo Magnaye |  |  |  |  | 26337 | COMMERCIAL | rmcbloodbank@gmail.com | 25 |  |  | MCS+ | 15H045G | Unit for preventive maintenance | Performed preventive maintenance Run diagnostic test Generated report | Performed preventive maintenance | unit is operational | With Sir Monsour |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740160130427316369/5740160130427316369_signature_28.png | Ferdinand B. Barroga | https://www.jotform.com/uploads/macarewecare/222927978372067/5740160130427316369/5740160130427316369_signature_31.png
| 2023-10-25 | 2023-10-25 11:30 | Tipo-Tipo Rural Health Unit | Tipo-Tipo proper, Tipo-Tipo Basilan | BARMM | Basilan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26336 | PBSP | fatmaaunal84@gmail.com | 29 |  | 2023-10-25 | GX 4 | 812773 | Module 3 h/w failed | All modules undergone Xpertcheck calibration and Preventive Maintenence | All modules undergone Xpertcheck calibration and Preventive Maintenence | Module 1, 2 and 4 passed the calibration while module 3 failed | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740157266812483615/5740157266812483615_signature_28.png | Fatma Aunal | https://www.jotform.com/uploads/macarewecare/222927978372067/5740157266812483615/5740157266812483615_signature_31.png
| 2023-10-25 | 2023-10-25 09:13 | Sto. Domingo District Hospital | Sto Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26335 | DOH | lhaydeesel@yahoo.com | 27 |  | 2023-10-25 | Genexpert IV | 831582 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740148612912079937/5740148612912079937_signature_28.png | Haydee Alejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5740148612912079937/5740148612912079937_signature_31.png
| 2023-10-25 | 2023-10-25 11:12 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26334 | PBSP | molecularlab@rmc.doh.gov.ph | 25 |  | 2023-10-23 | GX-IV-D-6C | 814374 | Module B2 for replacement | Replaced module B2 | Performed maintenance Replaced module B2 Generated reports | All modules operational. |  | Old: B2/210068691 New: B2/210070889 | G 6380 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5740143788359252849/5740143788359252849_signature_28.png | Erika Mae M. Memdoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5740143788359252849/5740143788359252849_signature_31.png
| 2023-10-24 | 2023-10-24 16:08 | MR Flores Health Center | P. Herrera St. | NCR | NCR 4th - Pateros | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26333 | PBSP | paterosclinicallaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 17 | 2023-10-23 | GX-IV-D-6C | 804254 | Module D2 for replacement. | Replaced module D2 | Performed maintenance Replaced module D2 Run reliability test on module D3 Generated reports | Unit is operational. Module D3 operational but under observation. | Unit from Marikina CHO | Old: D2/690385 New: D2/654888 | G 6374 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5739450098621354214/5739450098621354214_signature_28.png | John Karl Jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/5739450098621354214/5739450098621354214_signature_31.png
| 2023-10-24 | 2023-10-24 15:26 | Muñoz City Health Office | Science City of Munoz, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26332 | DOH | scomrtdl@gmail.com | 19 |  | 2023-10-24 | Genexpert IV | 831673 | UPS no backup | Replaced battery of UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Battery 2pcs | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5739420164914896968/5739420164914896968_signature_28.png | Catherine Celestino | https://www.jotform.com/uploads/macarewecare/222927978372067/5739420164914896968/5739420164914896968_signature_31.png
| 2023-10-24 | 2023-10-24 14:45 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 26331 | PBSP | bmclab_micro@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV | 805112 | Unit for maintenance | #NAME? | Performed maintenance | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5739397783817407544/5739397783817407544_signature_28.png | Mary Joy Marquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5739397783817407544/5739397783817407544_signature_31.png
| 2023-10-24 | 2023-10-24 14:34 | Adela Serra Ty Memorial Medical Center | Brgy. Telaje, Capitol Hills, Tandag City | 13 | Surigao del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 26330 | PBSP | tbdotsppmdunit@gmail.com | 27.15 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1175 |  | GX-4 | 804255 | None | None | Replacing Module A2 with SN: 623381 to SN: 638076 A3 with SN: 632682 to SN: 210009389 A4 with SN: 681588 to SN: 700114 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5739389601215418435/5739389601215418435_signature_28.png | Lorene Dagasdas | https://www.jotform.com/uploads/macarewecare/222927978372067/5739389601215418435/5739389601215418435_signature_31.png
| 2023-10-24 | 2023-10-24 14:21 | Tuburan rural health | Tables used, Tuburan | 9 | Tawi-Tawi | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26329 | DOH | rjakilan20@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 276 | 2023-08-03 | GX 4 | 838832 | None | All 4 modules undergone Xpertcheck calibration | Preventive Maintenence and xpertcheck calibration | All modules passed the calibration. | None | XPERTCHECK-CE-5 | 10459 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5739387554567171461/5739387554567171461_signature_28.png | Raiyana Hassan Jakilan | https://www.jotform.com/uploads/macarewecare/222927978372067/5739387554567171461/5739387554567171461_signature_31.png
| 2023-10-24 | 2023-10-24 14:07 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 26328 | PBSP | bmcgenexpertlab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 884623 | - Stuck cartridge on module 1 - CPU not powering up | #NAME? | #NAME? | Unit is not operational. CPU for reporting Model: Lenovo ThinkCentre SN: CPH0520F08487 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5739377993814104467/5739377993814104467_signature_28.png | Michaela Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5739377993814104467/5739377993814104467_signature_31.png
| 2023-10-24 | 2023-10-24 09:05 | Banlat Health Center | Daisy St., cor Daffodil St., Banlat Road, Tandang Sora, Quezon City, 1116 | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26327 | PBSP | mfgadayrmt17@gmail.com | 29 |  |  | GX IV 6C | 804688 | GX unit for PM | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed self test; all passed Performed preventive maintenance; all passed Generated report | Performed preventive maintenance | Unit is operational. All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5739290150423854883/5739290150423854883_signature_28.png | Maria Adaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5739290150423854883/5739290150423854883_signature_31.png
| 2023-10-24 | 2023-10-24 11:00 | Corazon Locsin Montelibano Memorial Regional Hospital | Bacolod City | 6 | Negros Occidental | Laboratory | NO | ON-SITE ASSISTANCE Unit transfer | CEPHEID | Emerson Bejemino |  |  |  |  | 26326 | MACARE-HBDC | burtyboy22@yahoo.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 48 |  | GX-4 | 110020743 | Unit foe tranfer to another location in ithe laboratory. | n/a | #NAME? | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5739260184796810503/5739260184796810503_signature_28.png | Mr. Burt Padronia | https://www.jotform.com/uploads/macarewecare/222927978372067/5739260184796810503/5739260184796810503_signature_31.png
| 2023-10-24 | 2023-10-24 10:56 | Central Luzon Doctor's Hospital | Hospital Drive San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26325 | COMMERCIAL | pjlsantos@cldh.com.ph | 24 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 166 ASSAY: Xpert SARS-CoV-2, COUNT: 1352 |  | GXIV-D | 839232 and 830ñ0841 | Machine relocation | Machine relocation Preventive maintenance | Machine relocation from Molecular Laboratory to Main Laboratory Installation and setup of Genexpert machine Setup daisy chain connection Gathered reports Performed functionality test | Unit is operational | For quotation of xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5739258634213549030/5739258634213549030_signature_28.png | Precious Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5739258634213549030/5739258634213549030_signature_31.png
| 2023-10-23 | 2023-10-23 18:58 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | REPAIR | HAEMONETICS | Joseph Acebedo |  |  |  |  | 26324 | COMMERCIAL | rodaflores1038@yahoo.com.ph | 27 |  |  | MCS+ | 11H307 | UNIT HAS NO POWER | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738680719425979787/5738680719425979787_signature_28.png | Roda Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5738680719425979787/5738680719425979787_signature_31.png
| 2023-10-23 | 2023-10-18 10:03 | Zamboanga del norte medical center | Dipolog, Zamboanga del norte | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26323 | PBSP | ferranbryder@gmail.com | 28 |  | 2023-10-18 | GX 4 | 815285 | Module 3 stuck cartridge | Removed stuck cartridge and all 4 modules undergone Xpertcheck calibration all modules passed. | Removed stuck cartridge and all 4 modules undergone Xpertcheck calibration | All 4 modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738603746418356586/5738603746418356586_signature_28.png | Ferrari Bryder Tomoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5738603746418356586/5738603746418356586_signature_31.png
| 2023-10-23 | 2023-10-23 14:17 | Mabinay Medicre Community Hospital | Mabinay, Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 26322 | DOH | mmchlaboratory2017@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 786 |  | GX-4 | 838481 | E5011 in module 769404/B2 | #NAME? | #NAME? | - 4 modules operational, - module B2 for observation and parallel run on other modules in case of E5011 recurrence. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738516772012960387/5738516772012960387_signature_28.png | DIANNE MI-OT | https://www.jotform.com/uploads/macarewecare/222927978372067/5738516772012960387/5738516772012960387_signature_31.png
| 2023-10-23 | 2023-10-18 14:24 | Bicol Medical Center | Naga | 5 | Camarines Sur | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Dan Emil Bodino |  |  |  |  | 26321 | MACARE-HBDC | josefreyes0715@gmail.com | 25 |  |  | MCS+ | 15H027G | For Preventive Maintenance | Cleaned rollers, surface Checked fuse, system | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738384192813749273/5738384192813749273_signature_28.png | Josef D. Reyes, RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5738384192813749273/5738384192813749273_signature_31.png
| 2023-10-23 | 2023-10-18 14:22 | Bicol Medical Center | Naga | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  |  |  |  | 26320 | MACARE-HBDC | josefreyes0715@gmail.com | 25 |  |  | MCS+ | 15H031G | MCS+ not powering on | Replaced fuse Cleaned rollers, surface | Checked fuse Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738379372819750906/5738379372819750906_signature_28.png | Josef D. Reyes, RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5738379372819750906/5738379372819750906_signature_31.png
| 2023-10-23 | 2023-10-18 14:20 | Bicol Medical Center | Naga | 5 | Camarines Sur | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Dan Emil Bodino |  |  |  |  | 26319 | MACARE-HBDC | josefreyes0715@gmail.com | 25 |  |  | MCS+ | 11J221 | For preventive maintenance | Cleaned rollers, surface. checked fuse | Performed Preventive Maintenance | Unit is Operational | For adjustment of SPM tolerance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738376212814518426/5738376212814518426_signature_28.png | Josef D. Reyes, RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5738376212814518426/5738376212814518426_signature_31.png
| 2023-10-23 | 2023-10-18 10:32 | Bicol Medical Center | Naga | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26318 | MACARE-HBDC | princesscalingacion@y7mail.com | 30 | ASSAY: Xpert HPV, COUNT: 512 ASSAY: Xpert MTB/RIF Ultra, COUNT: 614 ASSAY: Xpert SARS-CoV-2, COUNT: 488 |  | GX4 | 110008458 | For preventive maintenance | Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738312252816284815/5738312252816284815_signature_28.png | Princess Jastine Calingacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5738312252816284815/5738312252816284815_signature_31.png
| 2023-10-23 | 2023-10-18 10:28 | Bicol Medical Center | Naga | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 26317 | PBSP | princesscalingacion@y7mail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1593 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6956 ASSAY: Xpert SARS-CoV-2, COUNT: 2216 ASSAY: Xpert HIV-1 Viral Load, COUNT: 359 |  | GX16 | 842225 | B3, D1 hardware failed | Cleaned cartridge bays, plunger rod Generate report | Performed troubleshooting Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738302082811308181/5738302082811308181_signature_28.png | Princess Calingacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5738302082811308181/5738302082811308181_signature_31.png
| 2023-10-23 | 2023-10-20 11:46 | Dagupan City Health Office | Dagupan City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26316 | DOH | kath101278@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1387 |  | GX4 | 838286 | Due for ReXpertCheck | #NAME? | Performed ReXpertCheck | Unit is operational | Module B3 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738296404403303109/5738296404403303109_signature_28.png | Katherine Guarin | https://www.jotform.com/uploads/macarewecare/222927978372067/5738296404403303109/5738296404403303109_signature_31.png
| 2023-10-23 | 2023-10-20 09:35 | Lingayen Rural Health Unit | Lingayen, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26315 | PBSP | michelledoriarmt@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1125 | 2023-10-20 | GX4 | 834629 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 786169 >>>210101549 - Generated reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210101549 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738294054401492702/5738294054401492702_signature_28.png | Susan c.delacruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5738294054401492702/5738294054401492702_signature_31.png
| 2023-10-23 | 2023-10-19 08:37 | Ilocos Training and Medical Center | San Fernando, La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26314 | WHO | itrmc.mbl@gmail.com | 21 | ASSAY: Xpert HPV, COUNT: 365 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1723 ASSAY: Others, COUNT: 719 ASSAY: Xpert SARS-CoV-2, COUNT: 401 |  | GX4 | 848249 | Due for Maintenance | #NAME? | Performed Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738291774402984245/5738291774402984245_signature_28.png | Romnick Conui | https://www.jotform.com/uploads/macarewecare/222927978372067/5738291774402984245/5738291774402984245_signature_31.png
| 2023-10-23 | 2023-10-19 08:30 | Ilocos Training and Regional Medical Center | San Fernando, La Union | 1 | La Union | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26313 | PBSP | itrmc.mbl@gmail.com | 20 | ASSAY: Xpert HBV VL, COUNT: 97 ASSAY: Xpert HPV, COUNT: 357 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2189 ASSAY: Xpert SARS-CoV-2, COUNT: 401 ASSAY: Xpert HIV-1 Viral Load, COUNT: 174 | 2023-10-19 | GX4 | 707524 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 616730 >>> 210102068 - Generated reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210102068 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738288924405771382/5738288924405771382_signature_28.png | Romnick Conui | https://www.jotform.com/uploads/macarewecare/222927978372067/5738288924405771382/5738288924405771382_signature_31.png
| 2023-10-23 | 2023-10-17 09:56 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26312 | DOH | gbmmbg@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 571 | 2023-10-12 | GX4 | 838437 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 210097620 >>> 662094 - Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 662094 | DR 10730 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5738286074406946559/5738286074406946559_signature_28.png | Jeanette C.Trinidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5738286074406946559/5738286074406946559_signature_31.png
| 2023-10-21 | 2023-10-20 16:24 | St. Luke's Extension Clinic | J. Bocobo St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26311 | MACARE-HBDC | laboratory@slec.ph | 25 |  | 2023-10-17 | GX Infinity 6C | 809643 | Unit due for xpertcheck | Performed PM Checked and cleaned plunger rod, icore slot, filter, cartridge bays | Performed XpertCheck | Passed - A4, B4, C1, C4, D2, D3, D4, E3, F1, G1, G4 Failed - A1, A3, B1, B2, B3, C2, C3, D1, E1, E2, E4, F2, F3, F4, G2, G3 Not Available - A2 | With Sir Pat, Maam Cienne, Maam Rhea & Mike | Xpertcheck kit | DR 10748 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5737119397516518937/5737119397516518937_signature_28.png | MA. SARAH R. SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5737119397516518937/5737119397516518937_signature_31.png
| 2023-10-20 | 2023-10-20 15:57 | Philippine Heart Center | East Avenue,QC | NCR | NCR 2nd - Quezon City | Molecular | YES | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26310 | MACARE-HBDC | molecularlab@phc.gov.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 947 ASSAY: Others, COUNT: 7 ASSAY: Xpert SARS-CoV-2, COUNT: 1785 |  | Gx16 | 838655 | A3 644007 /A4 634060/ C3 645559/C4 628454/ D2 653968 encounter continous temp error. | Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Checked up unit | B3,B4,C1 disabled for reporting | A3 644007 /A4 634060/ C3 645559/C4 628454/ D2 653968 encounter continous temp error. For replacement of module reported already |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5736171926018550719/5736171926018550719_signature_28.png | Kenneth Melvin N. Casidsid | https://www.jotform.com/uploads/macarewecare/222927978372067/5736171926018550719/5736171926018550719_signature_31.png
| 2023-10-20 | 2023-10-20 14:49 | National Kidney and Transplant Institute | East,Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26309 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 888 ASSAY: Others, COUNT: 29 ASSAY: Xpert HBV VL, COUNT: 611 ASSAY: Xpert HPV, COUNT: 117 ASSAY: Xpert MTB/RIF, COUNT: 1115 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1307 ASSAY: Xpert SARS-CoV-2, COUNT: 14074 ASSAY: Xpert HCV VL, COUNT: 59 ASSAY: Xpert SARS-CoV-2, COUNT: 9 | 2023-10-19 | GX16 | 110005587 | For module replacement A1 620546 - 896997 | Performed module replacement Cleaned unit Cleaned plunger rod,filter,icore,cartridge bay Gathered reports Test run unit | Module replacement and maintenance | 15 modules operational | A4 h/w failed reported already | 900-0491rs | 10752 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5736171496014452546/5736171496014452546_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5736171496014452546/5736171496014452546_signature_31.png
| 2023-10-20 | 2023-10-20 11:22 | AFP Health Service Command | AFP Medical Center, V. Luna Road,Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26308 | MACARE-HBDC | mtvalderama@gmail.com | 26 |  |  | GX IV 6C | 805294 | Module 1 and 3 frquent problem | Performed PM Checked system logs Performed reliability test Module 1 and 3 HW failed Gathered data Generated report | Checked module 1 and 3 | U ut is operational Module 1 and 3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735992032616459868/5735992032616459868_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/5735992032616459868/5735992032616459868_signature_31.png
| 2023-10-20 | 2023-10-20 13:55 | Angono Rural Health Unit | P. Tolentino St. Angono | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26307 | DOH | siennablanco@yahoo.com | 25 |  |  | GX-IV-D-6C | 819864 | For XpertCheck, for monitor replacement | Run XpertCheck, replaced monitor | Performed maintenance Replaced monitor Run XpertCheck Generated reports | Unit is operational. Module A2/784784 failed in XpertCheck. Module A4/719514 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735965126815591410/5735965126815591410_signature_28.png | Sienna Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5735965126815591410/5735965126815591410_signature_31.png
| 2023-10-20 | 2023-10-19 13:48 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26306 | PBSP | lcpmicrolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4429 ASSAY: Xpert SARS-CoV-2, COUNT: 3368 | 2023-10-19 | GX16 | 801080 | Due for xpertcheck | Cleaned plunger rod,i core, cartridge bay,filter Gathered reports Performed xpertcheck | Xpertcheck and maintenance | 15 module passed xpertcheck | C3 failed in xpertcheck | Xpertcheck | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735780813615367655/5735780813615367655_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5735780813615367655/5735780813615367655_signature_31.png
| 2023-10-20 | 2023-10-20 09:54 | Cardona Rural Health Unit | Dr. Estancio St., Cardona | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26305 | DOH | renalyn_villaran@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 554 |  | GX-IV-D-6C | 821215 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735776032421301804/5735776032421301804_signature_28.png | Renalyn Villaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5735776032421301804/5735776032421301804_signature_31.png
| 2023-10-20 | 2023-10-17 16:55 | Nabua Rural Health Unit | Nabua | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26304 | PBSP | mongealyzza@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 58 |  | GX4 | 834620 | B2/653893 for Module Replacement | Module replaced: B2 - 653893>>210101613 Cleaned Gx fan Generate reports | Replaced Module B2/653893 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735775839226133660/5735775839226133660_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5735775839226133660/5735775839226133660_signature_31.png
| 2023-10-20 | 2023-10-20 09:42 | NGC Super Health Center | MRB Compound Pilot Drive, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26303 | DOH | tmcolisao@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1570 |  | GX IV 6C | 831586 | Unit for preventive maintenance | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed reliability test; all passed Generated report | Performed preventive maintenance | Unit is operational All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735758692612780138/5735758692612780138_signature_28.png | TRECIA MARIE D. COLISAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5735758692612780138/5735758692612780138_signature_31.png
| 2023-10-20 | 2023-10-19 15:08 | Dr. Elvira Lagrosa (Duay) Health Center | Libertad, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  | gpdmagnaye30@gmail.com |  |  | 26302 | PBSP | rhemejetbotaquia@gmail.com | 29 |  |  | GX IV 6C | 812833 | Module 2 invalid results | Performed PM Checked module 2 Performed reliability test Noticed that the valve is not in line with the case Gathered data for reporting | Troubleshoot module 2 | Unit is operational. Module 2 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735751212611001141/5735751212611001141_signature_28.png | Rheme Tañeza | https://www.jotform.com/uploads/macarewecare/222927978372067/5735751212611001141/5735751212611001141_signature_31.png
| 2023-10-20 | 2023-10-20 08:07 | DATU HALUN SAKILAN MEMORIAL HOSPITAL | BONGAO, TAWI-TAWI | BARMM | Tawi-Tawi | MOLECULAR LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26301 | IOM | joel_kamlian2003@yahoo.com | 30 | ASSAY: Xpert SARS-CoV-2, COUNT: 18 |  | GX-IV | 110008652 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735699001958826557/5735699001958826557_signature_28.png | Joel A. Kamlian | https://www.jotform.com/uploads/macarewecare/222927978372067/5735699001958826557/5735699001958826557_signature_31.png
| 2023-10-19 | 2023-10-19 17:57 | Roxas City Health Office | Roxas city | 6 | Capiz | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 26300 | PBSP | ruby_camacho@ymail.com | 24 |  |  | Gx 4 | 805333 | Defective module b4 sn 210100878 | New module installed b4 sn 210101550, clean fan and fan filter | New module installed b4 sn 210101550 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735189181595398332/5735189181595398332_signature_28.png | RUBY C. LUT | https://www.jotform.com/uploads/macarewecare/222927978372067/5735189181595398332/5735189181595398332_signature_31.png
| 2023-10-19 | 2023-10-19 15:01 | Polanco Rural health unit | Pblacion South polanco | 9 | Zamboanga del Norte | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26299 | PBSP | rhupolancocl@gmail.com | 27 |  | 2023-10-19 | GX 4 | 838539 | Module 3 error 2008 | Run some sample test on module 3, undergone Xpertcheck on module 3 and conduct preventive Maintenence on all modules. | Preventive Maintenence and xpertcheck calibration for module 3 | Module 3 failed xpertcheck calibration and needs to be replaced | Module 1 and 3 are pending for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735138467116596925/5735138467116596925_signature_28.png | Gergean luna | https://www.jotform.com/uploads/macarewecare/222927978372067/5735138467116596925/5735138467116596925_signature_31.png
| 2023-10-19 | 2023-10-19 16:24 | Governor Generoso Rural Health Unit | Poblacion, Governor Generoso | 11 | Davao Oriental | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26298 | DOH | josephmagallanes72@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 454 | 2023-08-03 | GX-IV 4d | 838490 | Due for Calibration and Preventive Maintenance | Run XpertCheck Calibration on Four Modules Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and XpertCheck Calibration | Modules B3 and B4 passed the Calibration Modules B1 and B2 skipped due to retries | Module B1 and B2 for under observation | Xpertcheck CE 5 | 10438 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735133716816428612/5735133716816428612_signature_28.png | Charlene Magallanes RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5735133716816428612/5735133716816428612_signature_31.png
| 2023-10-19 | 2023-10-19 15:42 | Teresa Rural Health Unit | ML Quezon St., Teresa | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26297 | DOH | teresa_rhu@yahoo.com | 25 |  |  | GX-IV-D-6C | 820710 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational. All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735123216312137869/5735123216312137869_signature_28.png | Rosalyn F. San Pedro | https://www.jotform.com/uploads/macarewecare/222927978372067/5735123216312137869/5735123216312137869_signature_31.png
| 2023-10-19 | 2023-10-19 15:06 | Tacloban City Health Office | Tacloban City | 8 | Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26296 | DOH | marteetingzon@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 69 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1271 | 2023-10-19 | GXIV | 838421 | Unit for Preventive Maintenance | #NAME? | #NAME? | -Unit is Operational -Module B1 and B4 disabled (waiting for the |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735092981216849472/5735092981216849472_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5735092981216849472/5735092981216849472_signature_31.png
| 2023-10-19 | 2023-10-19 13:39 | Catmon Health Center | Catmon, Malabon | NCR | NCR 3rd - Malabon | Laboratory l | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 26295 | PBSP | mariathes0218@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3041 ASSAY: Xpert HIV-1 Viral Load, COUNT: 144 |  | GXIV | 838880 | Unit for additional troubleshooting | #NAME? | Gathered data required by Cepheid | Unit is not operational. Data gathered for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735053121426645748/5735053121426645748_signature_28.png | Maria Theresa Nunez | https://www.jotform.com/uploads/macarewecare/222927978372067/5735053121426645748/5735053121426645748_signature_31.png
| 2023-10-19 | 2023-10-19 13:50 | Datu Halun Sakilan Memorial Hospital | Bongao, Tawi-Tawi | BARMM | Tawi-Tawi | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 26294 | DOH | joel_kamlian2003@yahoo.com | 28 |  |  | GXIV-4D | 838834 | None | None | #NAME? | Module B1 H/W failed. Other modules are functional. | Waiting for module B1 replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735041017425757082/5735041017425757082_signature_28.png | Joel A. Kamlian | https://www.jotform.com/uploads/macarewecare/222927978372067/5735041017425757082/5735041017425757082_signature_31.png
| 2023-10-19 | 2023-10-19 13:18 | Jala-Jala Rural Health Unit | Pililla-Jala-Jala-Pakil Rd. | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26293 | DOH | jjrhu@yahoo.com | 25 |  |  | GX-IV-D-6C | 820184 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735037266311740704/5735037266311740704_signature_28.png | Angeline minebe | https://www.jotform.com/uploads/macarewecare/222927978372067/5735037266311740704/5735037266311740704_signature_31.png
| 2023-10-19 | 2023-10-18 15:48 | University of the Philippines | U.P campus | NCR | NCR 2nd - Quezon City | Biology | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  |  |  | Plasmair | 26292 | MACARE-HBDC | palim1@up.educ.ph | 26 |  |  | Airinspace | 2006-TS-0015 | Encounter no power,no display,on and off running of fan motor | Checked electrical Checked controller Checked control panel Checked fan motor Checked connector and filters Test run unit | Troubleshooting | Unit unoperational,during test run encounter intermittent no display and on and off fan motor running | For observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735007764416609404/5735007764416609404_signature_28.png | Patricia Anne Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5735007764416609404/5735007764416609404_signature_31.png
| 2023-10-19 | 2023-10-17 10:37 | Las piñas general hospital and trauma center | Zapote st las piñas | NCR | NCR 4th - Las Piñas | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26291 | PBSP | lpghstc.molecular@gmail.com | 31 | ASSAY: Xpert SARS-CoV-2, COUNT: 32330 | 2023-10-16 | GX4 | 834628 | For module replacement B1 785012 -210101669 B3 786171 - 210101532 B4 785093 - 210101897 | Cleaned unit Cleaned filter,plunger rod, cartridge bay,icore Gathered reports Performed module replacement Test run unit | Module replacement and maintenance | Operational | With bren bagbaguen | 900-0508 | G6352 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5735007304419208657/5735007304419208657_signature_28.png | Vivien Angelu J. Panganiban | https://www.jotform.com/uploads/macarewecare/222927978372067/5735007304419208657/5735007304419208657_signature_31.png
| 2023-10-19 | 2023-10-19 10:46 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | Tbdots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26290 | DOH | lornsabaigar@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 681 | 2023-10-19 | GXIV | 838409 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734930930829749186/5734930930829749186_signature_28.png | Lorna Abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/5734930930829749186/5734930930829749186_signature_31.png
| 2023-10-19 | 2023-10-19 09:16 | Sta. Ignacia Rural Health Unit | Sta. Ignacia Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26289 | DOH | jdi0973@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 499 |  | GXIV-D | 831909 | A2 error 2035 UPS not charging | Xpertcheck Troubleshooting | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Performed xpertcheck Troubleshoot UPS Generate reports Archived test results Performed functionality test | Unit is operational,A2 passed calibration under observation | UPS for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734917844747567279/5734917844747567279_signature_28.png | Jess D. Ibarra Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5734917844747567279/5734917844747567279_signature_31.png
| 2023-10-19 | 2023-10-19 09:53 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26288 | PBSP | crmctblab@gmail.com | 23.5 | ASSAY: Xpert HBV VL, COUNT: 16 ASSAY: Xpert HPV, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 65 ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 | 2023-10-19 | Gx-IV | 807463 | Module 4 H/W failed | Perform Genexpert machine Preventive maintenance | Replace/repair defective module 4 SN: 655757 with new module SN:210102218 | 4 operational modules |  | Genexpert module 6 color | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734905473121462985/5734905473121462985_signature_28.png | Shaneen Meditar | https://www.jotform.com/uploads/macarewecare/222927978372067/5734905473121462985/5734905473121462985_signature_31.png
| 2023-10-19 | 2023-10-19 09:14 | Binan City Health Office 2 | Binan City | 4A | Laguna | Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  |  |  |  | 26287 | DOH | kensarinas28@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 126 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1887 ASSAY: Xpert HIV-1 Viral Load, COUNT: 400 | 2023-10-16 | GXIV | 821278 | For installation of new monitor. | Performed installation of monitor. Generated reports | Installation. | Unit is operational |  | Nvision Monitor (H22V5SG81A230601341) in replacement to defective HP ProDisplay P19A monitor (CNK629119V). | DR#10733 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734876509817049258/5734876509817049258_signature_28.png | Kenneth Adrian Sarinas | https://www.jotform.com/uploads/macarewecare/222927978372067/5734876509817049258/5734876509817049258_signature_31.png
| 2023-10-18 | 2023-10-18 17:01 | DATU HALUN SAKILAN MEMORIAL HOSPITAL | BONGAO, TAWI TAWI | BARMM | Tawi-Tawi | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26286 | PBSP | wmhadjula.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 885 |  | GX-IV | 838454 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734519882224960458/5734519882224960458_signature_28.png | Wardasa M. Hadjula | https://www.jotform.com/uploads/macarewecare/222927978372067/5734519882224960458/5734519882224960458_signature_31.png
| 2023-10-18 | 2023-10-18 16:58 | Datu Halun Sakilan Memorial Hospital | Bongao, Tawi-Tawi | BARMM | Tawi-Tawi | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 26285 | PBSP | wmhadjula.pbsp@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 783 | 2023-10-18 | GXIV-4D | 808643 | Module B1 defective | -Replace defective module B1/605709 with new module SN: 210102185 | #NAME? | Module B3 encountering H/W fail. All other modules are functional. |  | GeneXpert Module 6-Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734294422327613001/5734294422327613001_signature_28.png | Wardasa M. Hadjula | https://www.jotform.com/uploads/macarewecare/222927978372067/5734294422327613001/5734294422327613001_signature_31.png
| 2023-10-18 | 2023-10-18 16:50 | Dr Rafael S Tumbukon Memorial Hospital - HACT | Kalibo Aklan | 6 | Aklan | Laboratory | YES | INSTALLATION | CEPHEID | Raymond Laurito |  |  |  |  | 26284 | MACARE-HBDC | drstmhrhivda@gmail.com | 28 |  |  | Gx 4 | 110020745 | New installation | New installation | New installation | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734289397326809173/5734289397326809173_signature_28.png | Chrizeus R. Mayiong | https://www.jotform.com/uploads/macarewecare/222927978372067/5734289397326809173/5734289397326809173_signature_31.png
| 2023-10-18 | 2023-10-18 16:23 | Leyte Provincial Hospital | Palo, Leyte | 8 | Leyte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26283 | PBSP | LPHLabdeptcovid19@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 101 ASSAY: Xpert MTB/RIF Ultra, COUNT: 749 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2023-10-18 | GXIV | 815267 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734268684428133956/5734268684428133956_signature_28.png | Janelle Sevillana | https://www.jotform.com/uploads/macarewecare/222927978372067/5734268684428133956/5734268684428133956_signature_31.png
| 2023-10-18 | 2023-10-18 15:50 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26282 | PBSP | crmcmolecularlab@gmail.com | 24.6 | ASSAY: Xpert SARS-CoV-2, COUNT: | 2023-10-18 | Gx-IV | 834441 | Module 4 H/W failed | Perform Genexpert machine Preventive maintenance | Replace defective module 4 SN:805746 with new module SN:210102154 | 4 operational modules |  | Genexpert module 6 color SN: 210102154 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734257907228943629/5734257907228943629_signature_28.png | Darianne Mae Tubilan | https://www.jotform.com/uploads/macarewecare/222927978372067/5734257907228943629/5734257907228943629_signature_31.png
| 2023-10-18 | 2023-10-18 15:56 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26281 | PBSP | crmcmolecularlab@gmail.com | 25.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 7835 | 2023-10-18 | Gx-IV | 834442 | Module 1 H/W failed | Perform Genexpert machine Preventive maintenance | Replace defective module 1 SN:733420 with new module SN:210102042 | 4 operational modules |  | Genexpert module 6 color SN:210102042 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734256837226823833/5734256837226823833_signature_28.png | Darianne Mae Tubilan | https://www.jotform.com/uploads/macarewecare/222927978372067/5734256837226823833/5734256837226823833_signature_31.png
| 2023-10-18 | 2023-10-17 12:50 | Tupi Rural Health Unit | Tupi, South Cotabato | 12 | South Cotabato | Tb-Lab | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26280 | PBSP | pedzella@gmail.com | 28.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 804 | 2023-10-17 | Gx-IV (2 modules) | 835184 | Printer cartridge not detected | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine annual Xpertcheck/Calibration Troubleshooting and repair printer problem | 2 operational modules |  | Xpertcheck CE-5 cartridge/kit | Lot no.: 1000907498 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734245977223785475/5734245977223785475_signature_28.png | Ella Esmailene S. Pedregosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5734245977223785475/5734245977223785475_signature_31.png
| 2023-10-18 | 2023-10-18 15:26 | Dr. Rafael S. Tumbokon Memorial Hospital | Kalibo, Aklan | 6 | Aklan | Laboratory | NO | New CPU installation | CEPHEID | Raymond Laurito |  |  |  |  | 26279 | PBSP | gonzagaevarmt@gmail.com | 25 |  |  | Gx 4 | 804680 | Defective CPU | New CPU installation, retrive data from old CPU | New CPU installation | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734241576313231651/5734241576313231651_signature_28.png | Eva Raquel D. Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5734241576313231651/5734241576313231651_signature_31.png
| 2023-10-18 | 2023-10-18 15:20 | Quezon City Jail | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 26278 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA | 2023-10-17 | GXIV | 838881 | Modules B3, B3 for replacement | #NAME? | Replaced B2, B3 | Unit is operational. |  | GeneXpert Module SN: 210101856, 210101622 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734230123619112043/5734230123619112043_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5734230123619112043/5734230123619112043_signature_31.png
| 2023-10-18 | 2023-10-18 14:31 | Macalelon Rural Health Unit | Macalelon, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26277 | DOH | cheyster7388@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 407 |  | GX IV 6C | 831636 | CPU for temporary replacement | Performed PM Checked CPU Archived results Replaced CPU Retrieved results Installed drivers and softwares Generated report | Replaced CPU | Unit is operational | Actual date of service Oct. 6,2023. Jotform error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734207933619310697/5734207933619310697_signature_28.png | Cheyster O. Agustin | https://www.jotform.com/uploads/macarewecare/222927978372067/5734207933619310697/5734207933619310697_signature_31.png
| 2023-10-18 | 2023-10-18 14:11 | Guinayangan Munucipal Health Office | Guinayangan, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26276 | WHO | lala0117r.kampitan@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 554 | 2023-09-22 | GX IV 6C | 847334 | Module B3 for replacement | Checked and cleaned cartridge bays, plunger rod, i-core slot and filter Replaced module B3 Generated report | Replaced module B3 908422 -> 210101928 | Unit is operational All modules working | Actual date of service October 5, 2023. Jotform error | GX module 6C | DR 43550 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734193533618848215/5734193533618848215_signature_28.png | Leonila K. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5734193533618848215/5734193533618848215_signature_31.png
| 2023-10-18 | 2023-10-18 10:54 | Philippine Tuberculosis Society Inc. | E Rodriguez Sr. Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26275 | PBSP | sannymisorena@gmail.com | 22 |  |  | GX IV 6C | 707527 | Module 1 results for technical checking | Performed PM Archived results Performed reliability tests: Passed Gathered report | Archived results for checking | Unit ia operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734184983619451103/5734184983619451103_signature_28.png | Sanny Isorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5734184983619451103/5734184983619451103_signature_31.png
| 2023-10-18 | 2023-10-18 13:38 | Malita Rural Health Unit | Poblacion, Malita | 11 | Davao Occidental | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26274 | PBSP | mariastellarmt@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 860 |  | GX-IV 4D | 805258 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | All Modules Are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734168444027497139/5734168444027497139_signature_28.png | Maria Stella Solitana RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5734168444027497139/5734168444027497139_signature_31.png
| 2023-10-18 | 2023-10-18 12:43 | Dr. Jose Rizal Memorial Hospital | Junction Polo-Dapitan Park National Rd, Lungsod ng Dapitan, 7101 Zamboanga del Norte | 9 | Zamboanga del Norte | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26273 | PBSP | laboratorydjrmh@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1066 | 2023-10-17 | GX 4 | 804698 | Module 3 h/w failed | Module 3 undergone sample testing | Preventive Maintenence, checked module and room temperature. Site inspection. | Module 1,2 and 4 are available while module 3 is disabled at the moment. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734138719714637179/5734138719714637179_signature_28.png | Antonieta Lovino | https://www.jotform.com/uploads/macarewecare/222927978372067/5734138719714637179/5734138719714637179_signature_31.png
| 2023-10-18 | 2023-10-18 11:23 | San Antonio District Hospital | San Mariano, San Antonio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26272 | PBSP | garciaadonisanthony@gmail.com | 28 |  | 2023-10-18 | Genexpert IV | 812832 | Module 4 Hardware Failure, UPS not working | Troubleshoot UPS | Performed Maintenance Gathered Reports Installed Replacement Monitor Performed Functionality Test | Unit is Operational | UPS for Replacement, Module 4 for Reporting | Monitor - V5GAH837 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734092978911475550/5734092978911475550_signature_28.png | Adonis Anthony Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5734092978911475550/5734092978911475550_signature_31.png
| 2023-10-18 | 2023-10-18 10:28 | Pintuyan District Hospital | Pintuyan, Southern Leyte | 8 | Southern Leyte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26271 | PBSP | oclaritcleofedel25@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 527 | 2023-10-18 | GXIV-2D | 834825 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed (B1 and B2) | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734061826464529405/5734061826464529405_signature_28.png | CLEOFEDEL MARIE OCLARIT | https://www.jotform.com/uploads/macarewecare/222927978372067/5734061826464529405/5734061826464529405_signature_31.png
| 2023-10-18 | 2023-10-18 10:27 | Lambunao RHU | Lambunao | 6 | Iloilo | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 26270 | PBSP | gabatanelson@gmail.com | 27 |  |  | Gx IV | 815279 | Defective module B2 657139 | New module Installed B2 210101903, clean fan filter | New module Installed B2 210101903 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734056710213192955/5734056710213192955_signature_28.png | Nelson Gabata | https://www.jotform.com/uploads/macarewecare/222927978372067/5734056710213192955/5734056710213192955_signature_31.png
| 2023-10-18 | 2023-10-13 10:09 | Maluso Rural Health Unit | Maluso, Basilan | BARMM | Basilan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26269 | DOH | sharminahussin@gmail.com | 35 |  | 2023-08-03 | GX 4 | 838835 | None | Preventive Maintenence and Xpertcheck Calibration all modules passed | Modules undergone Xpertcheck calibration | All modules are available | None | XPERTCHECK-CE-5 | 1045 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734037727766021656/5734037727766021656_signature_28.png | Sharmina A. Hussin | https://www.jotform.com/uploads/macarewecare/222927978372067/5734037727766021656/5734037727766021656_signature_31.png
| 2023-10-18 | 2023-10-18 09:45 | Macare In-house | 1005 Aurora Blvd. Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | OTHERS | John Monsour Doctor |  |  |  | GeneXpert laptop | 26268 | MACARE-HBDC | htoco@macare-medicals.com | 25 |  |  | HP Probook | 5CG5040H6L | No CMOS battery | Replaced CMOS battery | Replaced CMOS battery Functionality test Connected to the GeneXpert unit Run reliability test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5734034183619880945/5734034183619880945_signature_28.png | Harold B. Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5734034183619880945/5734034183619880945_signature_31.png
| 2023-10-17 | 2023-10-17 16:43 | St Lukes Extension Clinic | J. Bocobo St. | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  |  |  |  | 26267 | MACARE-HBDC | laboratory@slec.ph | 23 |  |  | Infinity 48s | 809643 | Unable to connect to Infinity server error on startup | #NAME? | Performed troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733422995782860487/5733422995782860487_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5733422995782860487/5733422995782860487_signature_31.png
| 2023-10-17 | 2023-10-17 15:38 | San Jose Del Monte City Health Center VI | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 26266 | PBSP | krizetteamorio@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1895 ASSAY: Xpert HIV-1 Viral Load, COUNT: 107 | 2023-10-16 | GXIV | 814161 | Module B1 for replacement | #NAME? | Replaced B1 | Unit is operational. |  | GeneXpert Module SN: 210101583 | G 6353 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733377641895985222/5733377641895985222_signature_28.png | Krizette Marie Amorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5733377641895985222/5733377641895985222_signature_31.png
| 2023-10-17 | 2023-10-17 14:43 | West Crame Health Center | Rd 10., West Crame | NCR | NCR 2nd - San Juan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26265 | PBSP | azel1101@yahoo.com | 24 |  | 2023-10-17 | GX-IV-D-6C | 812007 | Module B3 for replacement, barcode scanner won't turn on. | Replaced module B3 | Performed maintenance Replaced module B3 Replaced filter Checked barcode scanner Generated reports | Unit is operational. Barcode scanner not operational, recommend to replace. | Barcode scanner SN: 1624400501383 | Old: B3/682815 New: B3/210101893 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733351687591646493/5733351687591646493_signature_28.png | Azazel Alpad | https://www.jotform.com/uploads/macarewecare/222927978372067/5733351687591646493/5733351687591646493_signature_31.png
| 2023-10-17 | 2023-10-17 14:31 | Gabaldon Rural Health Unit | South Poblacion, Gabaldon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 26264 | PBSP | chinablue_12@yahoo.com | 26 |  | 2023-10-17 | Genexpert IV | 834804 | UPS not working | Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733340982314975973/5733340982314975973_signature_28.png | Helen Grace Parungao | https://www.jotform.com/uploads/macarewecare/222927978372067/5733340982314975973/5733340982314975973_signature_31.png
| 2023-10-17 | 2023-10-17 13:45 | Dingalan Community Hospital | Caragsacan | 3 | Aurora | Laboratory | NO | CALIBRATION MAINTENANCE Troubleshooting | CEPHEID | Matthew Roxas |  |  |  |  | 26263 | DOH | bnjlvrz9@gmail.com | 28 |  | 2023-10-17 | Genexpert IV | 831621 | Module 2 and Module 3 Hardware Failure | Performed Xpertcheck on Module 2 and 3 | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = (A1, A4 = Passed) & (A2, A3 = Service Required) | Xpertcheck 4pc | N/w |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733310792312085860/5733310792312085860_signature_28.png | Benjo Alvarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5733310792312085860/5733310792312085860_signature_31.png
| 2023-10-17 | 2023-10-17 11:43 | Hagonoy Rural Health Unit | Hagonoy, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 26262 | PBSP | lanielyncastro@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1909 |  | GXIV | 814158 | #NAME? | - checked and cleaned cartridge bay, plunger rod, filter - relocated GeneXpert since ambient temp. of room exceeds 30 degrees C. - checked UPS' back up mode (<5min. back up & defective power switch) - performed XpertCheck - generated reports | Performed troubleshooting | Unit is operational. B1: passed B2: service required B3: service required B4: service required UPS is recommended for replacement/repair. GeneXpert is under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733275961891622857/5733275961891622857_signature_28.png | Lanielyn Castro Abadam | https://www.jotform.com/uploads/macarewecare/222927978372067/5733275961891622857/5733275961891622857_signature_31.png
| 2023-10-17 | 2023-10-17 11:32 | Sagay City Health Office | Sagay City | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 26261 | PBSP | maravelesdebra@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1510 |  | GX4 | 838144 | N/A | Pulled out the old module for B4 with SN 806359 and save IQR files before and after replacement. | Installed the new module for B4 with SN 210101994. | All modules are now working. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733230692172912561/5733230692172912561_signature_28.png | Debra Abellar | https://www.jotform.com/uploads/macarewecare/222927978372067/5733230692172912561/5733230692172912561_signature_31.png
| 2023-10-17 | 2023-10-17 11:10 | Dr. Rafael S. Tumbokon Memorial Hospital | Poblacion | 6 | Aklan | Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 26260 | PBSP | gonzagaevarmt@gmail.com | 21 |  | 2023-10-17 | GX4 | 804680 | #NAME? | #NAME? | - Perform troubleshooting on temperature related errors. - Reset modules gateway board. - Check new CPU recived from cephied for installation. - New CPU SN. CPH0623F42287. - No available VGA port for existing LCD monitor. - Canvas area for HDMI-to-VGA or DP-to-VGA connector. No available adapters around facility and within town area. New CPU to be installed on 10/18/2023. | 4 modules operational, for further observation for temperature related errors. | - No available HDMI-VGA adaptor in the vicinity, will continue installation of new CPU on 10/18/2023. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733227054229095894/5733227054229095894_signature_28.png | Eva Raquel Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5733227054229095894/5733227054229095894_signature_31.png
| 2023-10-17 | 2023-10-17 11:10 | Dr. Rafael S. Tumbokon Memorial Hospital | Poblacion | 6 | Aklan | Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 26259 | PBSP | gonzagaevarmt@gmail.com | 21 |  | 2023-10-17 | GX4 | 804680 | Defective CPU Due to Preventive Maintenance | #NAME? | #NAME? | All module are functional. | - No available HDMI-VGA adaptor in the vicinity, will continue installation of new CPU on 10/18/2023. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733225524228188060/5733225524228188060_signature_28.png | Eva Raquel Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5733225524228188060/5733225524228188060_signature_31.png
| 2023-10-17 | 2023-10-17 10:07 | Cadiz City Health Office | Cadiz City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 26258 | DOH | zhevebacomo@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1804 |  | GX4 | 838499 | Module B4 with error 4018 with valve label version is out of range configuration setting of 0. | Module B1, B2 and B3 passed the Xpert Check while module B4 is not detectable due to Harware Failed. | Performed Xpert check calibration for B1, B2 and B3. | Module B1, B2 and B3 are in good working condition while module B4 is need service. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733187294178686301/5733187294178686301_signature_28.png | Zherlyn Eve Bacomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5733187294178686301/5733187294178686301_signature_31.png
| 2023-10-17 | 2023-10-17 09:10 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | YES | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 26257 | COMMERCIAL | cduhlabresults2@yahoo.com | 25 |  |  | MCS+ | 17H639SPG | Purple valve problem during self test error | Test valve in diagnostics mode and for observation | Diagnose valve , check compressor and run valve test | MCS+ are working properly |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733145513919653561/5733145513919653561_signature_28.png | Kyyara Lunzaga, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5733145513919653561/5733145513919653561_signature_31.png
| 2023-10-17 | 2023-10-17 08:57 | CEBU DOCTORS UNIVERSITY HOSPITAL | OSMENA BLVD. CAPITOL SITE CEBU CITY | 7 | Cebu | Bloodbank | YES | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 26256 | COMMERCIAL | cduhlabresults2@yahoo.com | 23 |  |  | M | 14F014 | CENTRIFUGE TOO LOW DURING RETURNS CYCLE | Performed diagnostics in centrifuge, clean motor pumps red, white and blue pump, | Performed diagnostics in centrifuge speed test for 45 minutes no problem exist | The Mcs+ is operational, |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5733142243914212778/5733142243914212778_signature_28.png | Kyyara Lunzaga, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5733142243914212778/5733142243914212778_signature_31.png
| 2023-10-16 | 2023-10-16 19:21 | Guihulngan City Helth Office | Poblacion | 7 | Negros Occidental | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 26255 | PBSP | ayceepadillo@gmail.com | 23 |  | 2023-10-12 | Gx4 | 815463 | Module 2 | Replace module 2 | Replace module 2 with new serial number. | All moduleok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732648430698684574/5732648430698684574_signature_28.png | Hyacinth Padellio | https://www.jotform.com/uploads/macarewecare/222927978372067/5732648430698684574/5732648430698684574_signature_31.png
| 2023-10-16 | 2023-10-16 16:50 | Talisay City City Health Unit | Cansojong, Talisay City Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 26254 | PBSP | marydelgella@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1938 |  | GX4 | 815384 | Module replacement of B2 SN# 705605 due to error of 1001 / 1002 | Replaced new module of B2 and informed end user to maintain aircon temperature at 24 to 26 Celcius | Replaced new module of B2- 210101904 | All 4 modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732558163913571374/5732558163913571374_signature_28.png | Marydel Gella, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5732558163913571374/5732558163913571374_signature_31.png
| 2023-10-16 | 2023-10-13 17:52 | Iriga City Health Office | Iriga City | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26253 | DOH | mongealyzza@yahoo.com.ph | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1915 |  | GX4 | 838391 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732528293423848153/5732528293423848153_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5732528293423848153/5732528293423848153_signature_31.png
| 2023-10-16 | 2023-10-13 17:51 | Nabua Rural Health Unit | Nabua | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26252 | PBSP | mongealyzza@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 25 |  | GX4 | 834620 | For CPU Replacement GX for Calibration | CPU Replaced: CPH0623F42286 Performed XpertCheck Cleaned cartridge bays, plunger rod, icore Generate reports | Replaced CPU Performed XpertCheck Performed Preventive Maintenance | Unit is operational | XpertCheck: ,B1,B2,B3 passed, B4 for retry |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732526343429335329/5732526343429335329_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5732526343429335329/5732526343429335329_signature_31.png
| 2023-10-16 | 2023-10-13 11:09 | Ragay Rural Health Unit | Ragay | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26251 | DOH | phylgomez0115@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 637 |  | GX4 | 838560 | Gx for Callibration | Performed XpertCheck Cleaned cartridge bays, plunger rod, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit Is Operational | B2/801695, B4/862868 - Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732522103421344633/5732522103421344633_signature_28.png | PHYL FRANCES C. GOMEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5732522103421344633/5732522103421344633_signature_31.png
| 2023-10-16 | 2023-10-12 16:17 | Pamplona Rural Health Unit | Pamplona | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26250 | PBSP | francislozano226@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 218 |  | GX4 | 812774 | A1/686139 for module replacement | Module replaced: A1/686139>>210100892 Cleaned cartridge bays, plunger rod, icore, fan Generate reports | Replaced module A1/686139 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732516523422152509/5732516523422152509_signature_28.png | Francis S Lozano | https://www.jotform.com/uploads/macarewecare/222927978372067/5732516523422152509/5732516523422152509_signature_31.png
| 2023-10-16 | 2023-10-16 15:05 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26249 | PBSP | elma.bragat@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4731 ASSAY: Xpert HIV-1 Viral Load, COUNT: 159 |  | GX-IV | 815275 | FOR MODULE D2 REPLACEMENT | COLLECT IQR, SYTEM LOG, AND ARCHIVE | -PERFORM PREVENTIVE MAINTENANCE -PERFORM MODULE D2 REPLACEMENT: DEF SN: 629321 NEW MODULE SN:210101403 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732496535107055784/5732496535107055784_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5732496535107055784/5732496535107055784_signature_31.png
| 2023-10-16 | 2023-10-16 13:00 | Bansalan Rural Health Unit | Poblacion Dos, Bansalan | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26248 | PBSP | ellen68nebrada@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1010 |  | GX-IV 4D | 838181 | Module B1 for replacement Due for Preventive Maintenance | Replace the defective Module B1 SN 768092 with new Module SN 210102149 Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732419273415453889/5732419273415453889_signature_28.png | Ellen N. Mayor RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5732419273415453889/5732419273415453889_signature_31.png
| 2023-10-16 | 2023-10-09 13:21 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26247 | PBSP | dmgonzales.pbsp@gmail.com | 25 |  | 2023-10-09 | GX4 10C | 120000171 | Due for Module Replacement Modules B1 and B3 for Troubleshooting | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B4 220011690 >>>210067568 - Performed Reliability Test on Module B3 - Generated Reports | Performed Module Replacement Performed Troubleshooting | Unit is operational | Modules B1 and B3 for reporting | GeneXpert Module 210067568 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732280843618866987/5732280843618866987_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5732280843618866987/5732280843618866987_signature_31.png
| 2023-10-16 | 2023-10-09 13:27 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26246 | PBSP | dmgonzales.pbsp@gmail.com | 28 |  |  | GX4 6C | 805212 | For ReXpertCheck on Modules B1, B2, and B3 | #NAME? | Performed ReXpertCheck | Unit is operational | Module B1 for Reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732279303613805884/5732279303613805884_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5732279303613805884/5732279303613805884_signature_31.png
| 2023-10-16 | 2023-10-09 13:30 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26245 | PBSP | dmgonzales.pbsp@gmail.com | 29 |  |  | GX4 6C | 834685 | For Preventive Maintenance | #NAME? | Performed Preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732278043617368197/5732278043617368197_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5732278043617368197/5732278043617368197_signature_31.png
| 2023-10-16 | 2023-10-12 08:30 | Laoag City Health Office | Laoag, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26244 | DOH | denver.mahor@gmail.com | 28 |  |  | GX4 | 838416 | Due for ReXpertCheck | #NAME? | Performed ReXpertCheck | Unit is operational | Module B3 failed ReXpertCheck; For Reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732276603615115487/5732276603615115487_signature_28.png | Rodenver D. Mahor | https://www.jotform.com/uploads/macarewecare/222927978372067/5732276603615115487/5732276603615115487_signature_31.png
| 2023-10-16 | 2023-10-12 08:30 | Laoag City Health Office | Laoag, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26243 | DOH | denver.mahor@gmail.com | 28 |  |  | GX4 | 838416 | Due for ReXpertCheck | #NAME? | Performed ReXpertCheck | Unit is operational | Module B3 failed ReXpertCheck; For Reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732276173616161997/5732276173616161997_signature_28.png | Rodenver D. Mahor | https://www.jotform.com/uploads/macarewecare/222927978372067/5732276173616161997/5732276173616161997_signature_31.png
| 2023-10-16 | 2023-10-12 10:05 | Dingras District Hospital | Dingras, Ilocos Norte | 1 | Ilocos Norte | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26242 | PBSP | kenacudapas1022@gmail.com | 33 |  |  | GX4 | 807597 | Due for Troubleshooting | #NAME? | Performed Troubleshooting | Unit is operational | Modules B1, B2 and B3 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732275563619951924/5732275563619951924_signature_28.png | Kenadine Cudapas | https://www.jotform.com/uploads/macarewecare/222927978372067/5732275563619951924/5732275563619951924_signature_31.png
| 2023-10-16 | 2023-10-12 15:42 | Mariano Marcos Memorial Hospital and Medical Center | Batac, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26241 | DOH | dennisoncatindig@gmail.com | 26 |  |  | GX4 | 831908 | Due for ReXpertCheck | #NAME? | Performed ReXpertCheck | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732274323612657333/5732274323612657333_signature_28.png | Dennison Catindig | https://www.jotform.com/uploads/macarewecare/222927978372067/5732274323612657333/5732274323612657333_signature_31.png
| 2023-10-16 | 2023-10-12 15:47 | Mariano Marcos Memorial Hospital and Medical Center | Batac, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26240 | PBSP | dennisoncatindig@gmail.com | 25 |  |  | GX4 | 838168 | For Preventive Maintenance | #NAME? | Performed Preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732271433612238947/5732271433612238947_signature_28.png | Dennison Catindig | https://www.jotform.com/uploads/macarewecare/222927978372067/5732271433612238947/5732271433612238947_signature_31.png
| 2023-10-16 | 2023-10-13 09:51 | Bantay Municipal Health Office | Bantay, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26239 | PBSP | arvibration@gmail.com | 27 |  | 2023-10-13 | GX4 | 807585 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module A4 663731 >>> 210100212 - Generated Reports | Performed Module Replacement | Unit is operational | Module A2 for reporting | GeneXpert Module 210100212 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5732270463611159558/5732270463611159558_signature_28.png | Mel Arvin Pagalanan | https://www.jotform.com/uploads/macarewecare/222927978372067/5732270463611159558/5732270463611159558_signature_31.png
| 2023-10-15 | 2023-10-12 10:11 | Binan City Health Office 2 | Binan Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 26238 | DOH | montanezaira016@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 126 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1858 ASSAY: Xpert HIV-1 Viral Load, COUNT: 396 |  | GXIV | 821278 | Unit for PM. For Xpertcheck of module A3. For checking of monitor. | Performed PM. Performed Xpertcheck. Genersted reports. | PM and Xpertcheck. | Unit is operational. Module A3, service required. Monitor for replacement due to lower part display has pixel damage. | Used 2 spare Xpertcheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728863514718698687/5728863514718698687_signature_28.png | Jasmine Aira Montañez | https://www.jotform.com/uploads/macarewecare/222927978372067/5728863514718698687/5728863514718698687_signature_31.png
| 2023-10-13 | 2023-10-13 11:21 | Nasugbu Rural Health Unit | Nasugbu-Ternate Highway | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  |  |  |  | 26237 | PBSP | lucenavicente15@gmail.com | 26 |  |  | GX-IV-D-6C | 807581 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5730121114406799396/5730121114406799396_signature_28.png | Lucena I. Vicente | https://www.jotform.com/uploads/macarewecare/222927978372067/5730121114406799396/5730121114406799396_signature_31.png
| 2023-10-13 | 2023-10-13 16:00 | Ospital ng Palawan | San Juan Rd. cor. Malvar St. Puerto Princesa City, Palawan | 4B | Palawan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26236 | PBSP | amabduljalil.pbsp@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3032 ASSAY: Xpert SARS-CoV-2, COUNT: 1500 ASSAY: Xpert HIV-1 Viral Load, COUNT: 174 | 2023-10-04 | GX IV 6C | 842893 | Module C2 and D4 for replacement | Performed PM Checked all modules Checked system log Gathered data Replaced 2 modules Generated report | Replaced Modules C2: 885614 -> 210102179 D4: 877188 -> 210101886 | unit is operational |  | GX modules 6C | DR 43709 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729959301421823270/5729959301421823270_signature_28.png | Arabia M. Abduljalil | https://www.jotform.com/uploads/macarewecare/222927978372067/5729959301421823270/5729959301421823270_signature_31.png
| 2023-10-13 | 2023-10-13 13:22 | Puerto Princesa City Health Office | Valencia St., Puerto Princesa, Palawan | 4B | Palawan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26235 | PBSP | jocal0324@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1405 |  | GXIV 6C | 832026 | Module B2 and B3 for troubleshooting | Performed PM Checked all modules system log Module B2 frequent temp error Module B3 hardware failed Gathered data Generated report | Checked system log, performed reliability test | unit is operational 2 modules for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729842352522434570/5729842352522434570_signature_28.png | Jocelyn L.Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5729842352522434570/5729842352522434570_signature_31.png
| 2023-10-13 | 2023-10-13 13:09 | Puerto Princesa City Health Office | Valencia St., Puerto Princesa, Palawan | 4B | Palawan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26234 | PBSP | jocal0324@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1451 |  | GX IV 6C | 804272 | Modules D2, D3, D4 for troubleshooting | Performed PM Checked all modules Performed reliability test Modules D1, D2, D3 hardware failed, for reporting Module D4-ok Generated report | Checked Modules and performed reliability test | Unit is operational 1 module working, 3 modules hw failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729834352525873279/5729834352525873279_signature_28.png | Jocelyn L. Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5729834352525873279/5729834352525873279_signature_31.png
| 2023-10-13 | 2023-10-13 12:11 | Sagay City Health Office | Sagay City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 26233 | PBSP | maravelesdebra@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1497 |  | GX4 | 838144 | Module B4 already reported and has allocated new module for replacement. | Module B2 and B3 passed the XPC. Save IQR and GXC File. | Run Xpert Check calibration for module B2 and B3 only. | Module B1, B2 and B3 are in good condition. While module B4 is in disable mode. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729805561708521288/5729805561708521288_signature_28.png | Debra Abellar | https://www.jotform.com/uploads/macarewecare/222927978372067/5729805561708521288/5729805561708521288_signature_31.png
| 2023-10-13 | 2023-10-13 11:50 | St. Elizabeth Hospital Inc. | General Santos City, South Cotabato | 12 | South Cotabato | Molecular lab | NO | INSTALLATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26232 | COMMERCIAL | laboratory@sehi.ph | 21.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 120 |  | Gx-IV | 831313 | N/a | Genexpert machine Preventive maintenance | Transfer Genexpert machine from molecular lab to micrology room Install Genexpert Computer set up and UPS to micrology room Install Genexpert machine to micrology room | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729788757305528396/5729788757305528396_signature_28.png | SEHI LABORATORY | https://www.jotform.com/uploads/macarewecare/222927978372067/5729788757305528396/5729788757305528396_signature_31.png
| 2023-10-13 | 2023-10-13 11:31 | DAVAO DEL NORTE HOSPITAL-IGACOS ZONE | SAMAL, DAVAO DEL NORTE | 11 | Davao del Norte | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26231 | PBSP | allensoriano92@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 |  | GX-IV | 815383 | MODULE B1 DEFECTIVE | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | -PERFORM MAINTENANCE -PERFORM MODULE B1 REPLACEMENT: DEF SN:913008 NEW MODULE SN:210102030 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729775092311067603/5729775092311067603_signature_28.png | ALLEN GLENN SORIANO | https://www.jotform.com/uploads/macarewecare/222927978372067/5729775092311067603/5729775092311067603_signature_31.png
| 2023-10-13 | 2023-10-13 11:24 | Makati Medical Center (Medilinx) | Amorsolo at., Makati | NCR | NCR 4th - Makati | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 26230 | COMMERCIAL | clairemariejuguiad.medilinx@outlook.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 120 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 555 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 583 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2178 ASSAY: Others, COUNT: 2 ASSAY: Others, COUNT: 10 ASSAY: Xpert HCV VL, COUNT: 51 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1567 | 2023-10-13 | GX4(10C) | 110001647 | For Module replacement B2 210066496 - 210010556 | Performed Module replacement Cleaned plunger Rod, filter cartridge bay Gathered reports Test run unit | Module replacement, maintenance | Operational |  | 900-0650rs | G6346 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729774950149090451/5729774950149090451_signature_28.png | Claire Marie C. Juguiad | https://www.jotform.com/uploads/macarewecare/222927978372067/5729774950149090451/5729774950149090451_signature_31.png
| 2023-10-13 | 2023-10-13 10:26 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 26229 | DOH | rsjalova1974@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2429 |  | GXIV | 831596 | Stuck cartridge on C1 | #NAME? | Removed stuck cartridge | Unit is operational. C1 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729733365224011504/5729733365224011504_signature_28.png | Teresita Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5729733365224011504/5729733365224011504_signature_31.png
| 2023-10-13 | 2023-10-11 12:14 | Bulusan Rural Health Unit | Bulusan | 5 | Sorsogon | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26228 | PBSP | erminokarlian@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 711 |  | GX4 | 835197 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore, fan filter Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729716450201537810/5729716450201537810_signature_28.png | Karl Ian Ermino | https://www.jotform.com/uploads/macarewecare/222927978372067/5729716450201537810/5729716450201537810_signature_31.png
| 2023-10-13 | 2023-10-13 09:23 | MAYOR HILARION A RAMIRO SR MEDICAL CENTER | Maningcol, Ozamiz-Oroquieta National Highway 7200 Ozamis City Misamis Occidental · | 10 | Misamis Occidental | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 26227 | PBSP | mharsmicrobio@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2619 |  | GX4 | 804765 | Module 2 Hardware failure SN 662968 due to temperature problem | Replaced new module in B2 SN # 210100203 | Replaced new module in B2 SN 210100203 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729698283913331508/5729698283913331508_signature_28.png | Felipe Jr. Saren, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5729698283913331508/5729698283913331508_signature_31.png
| 2023-10-13 | 2023-10-13 09:12 | MAYOR HILARION A RAMIRO SR MEDICAL CENTER | Maningcol, Ozamiz-Oroquieta National Highway 7200 Ozamis City Misamis Occidental · | 10 | Misamis Occidental | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 26226 | PBSP | mharsmicrobio@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 474 ASSAY: Xpert SARS-CoV-2, COUNT: 5251 |  | GX4 | 834437 | Module 3 SN # 784125 encountered hardware failure due to error 1002/1001 | Replaced new module and performed preventive maintenance | Replaced Module 3 with new module SN# 210060931 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729692893917036450/5729692893917036450_signature_28.png | Felipe Jr. Saren, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5729692893917036450/5729692893917036450_signature_31.png
| 2023-10-13 | 2023-10-12 10:07 | East Avenue Medical Center | East avenue, QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26225 | MACARE-HBDC | eamcmolecularlab@gmail.com | 25 |  |  | GX4 | 110014085 | Due for maintenance and for network of their new gx | Cleaned unit Cleaned plunger Rod, filter, cartridge bay, icore Gathered reports Test run unit Performed network | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729644123614337802/5729644123614337802_signature_28.png | Jev Ryan Bonifacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5729644123614337802/5729644123614337802_signature_31.png
| 2023-10-13 | 2023-10-12 10:17 | East Avenue Medical Center | East Avenue, QC | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26224 | MACARE-HBDC | eamcmolecularlab@gmail.com | 26 |  |  | GX4 | 110016781 | For maintenance and for network | Cleaned unit Cleaned plunger Rod, filter, cartridge bay, icore Gathered reports Test run unit Performed network | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729643723613440596/5729643723613440596_signature_28.png | Jev Ryan Bonifacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5729643723613440596/5729643723613440596_signature_31.png
| 2023-10-13 | 2023-10-12 15:22 | Philippine Blood Center | Quezon Ave, QC | NCR | NCR 2nd - Quezon City | Blood collection | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 26223 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15H028G | CRC Bios, drain bios battery | Checked battery voltage 0 volts Swap CPU board frm MCS 15H029G Temporary installed CPU board frm 15H029G. Performed calibration Test run unit Checked program parameters | Troubleshooting | Operational | For quotation of CPU Board |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729643153615997378/5729643153615997378_signature_28.png | Ivan Ritz Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/5729643153615997378/5729643153615997378_signature_31.png
| 2023-10-13 | 2023-10-13 07:24 | Bago City Health Office | Bago City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 26222 | DOH | swabteambago@gmail.coml | 4 | ASSAY: Xpert MTB/RIF, COUNT: 259 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2598 |  | GX4 | 838486 | Module B3 and B4 already reported for module replacement due to failed in XPC. | Module B2 failed in XPC and save IQR & calibration info. | Run XPC for module B2 with SN 711021 due to error 4012 & 1002. | Module B1 and B3 are the functional module. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729628406122084446/5729628406122084446_signature_28.png | Glenly Casona | https://www.jotform.com/uploads/macarewecare/222927978372067/5729628406122084446/5729628406122084446_signature_31.png
| 2023-10-13 | 2023-10-12 11:04 | Datu Odin Sinsuat District Hospital | Datu Odin Sinsuat, Maguindanao | BARMM | Maguindanao | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26221 | DOH | rowenasalazar1234@gmail.com | 28.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 558 | 2023-10-12 | Gx-IV | 838205 | Module 2 under observation due to temperature errors | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine annual Xpertcheck/Calibration Module 1,2 and 3 passed Module 4 skipped - new module Perform actual sample process to module 2 for troubleshooting | 4 operational modules |  | Xpertcheck CE-5 cartridge/kit | DR no.: 10417 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729611425508669953/5729611425508669953_signature_28.png | Rowena B. Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5729611425508669953/5729611425508669953_signature_31.png
| 2023-10-12 | 2023-10-12 16:02 | Liloan Rural Health Unit | Liloan, Cebu City | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 26220 | PBSP | vickcabahugmt@gmail.com | 28 |  |  | Gx 4 | 835189 | Due date for xpertcheck calibration | Xpertcheck calibration for module 1 & 2 only | Xpertcheck calibration for module 1 & 2 only | Systems operational | All modules passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5729087212226290413/5729087212226290413_signature_28.png | Vick Cabahug | https://www.jotform.com/uploads/macarewecare/222927978372067/5729087212226290413/5729087212226290413_signature_31.png
| 2023-10-12 | 2023-10-12 13:40 | Bangkal Health Center | Makati | NCR | NCR 4th - Makati | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  |  |  |  | 26219 | DOH | kapid5sepe@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1716 |  | GXIV | 831875 | Unit for installation | #NAME? | Connected GX to switch | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728985644559180820/5728985644559180820_signature_28.png | Ed Sepe | https://www.jotform.com/uploads/macarewecare/222927978372067/5728985644559180820/5728985644559180820_signature_31.png
| 2023-10-12 | 2023-10-12 11:53 | Tumauini District Hospital | Centro, Tumauini, Isabela | 2 | Nueva Ecija | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26218 | MACARE-HBDC | kayecatembung2911@gmail.com | 25 |  | 2023-10-12 | Genexpert GXIV-10c | 110013709 | N/a | N/a | Installed Genexpert Installed Printer Installed UPS Installed PC Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728957175414867587/5728957175414867587_signature_28.png | Kathrine Mae Catembung | https://www.jotform.com/uploads/macarewecare/222927978372067/5728957175414867587/5728957175414867587_signature_31.png
| 2023-10-12 | 2023-10-12 12:28 | Las Nieves Rural Health Unit | Poblacion, Las Nieves | 13 | Agusan del Norte | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 26217 | WHO | Daisermt@gmail.com | 27.03 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 170 |  | GX-4 | 847338 | Stuck Cartridges at Module B1 and B2 | None | Removing Stuck Cartridges Preventive Maintenance | All Modules are Functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728953229016216961/5728953229016216961_signature_28.png | Daisy Siega | https://www.jotform.com/uploads/macarewecare/222927978372067/5728953229016216961/5728953229016216961_signature_31.png
| 2023-10-12 | 2023-10-11 11:27 | Tatalon health center | Victory Ave. QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26216 | DOH | annieduhino15@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 892 |  | GX4 | 831610 | Due for maintenance | Cleaned unit Cleaned filter, cartridge bay, icore, plunger Rod Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728931420142527585/5728931420142527585_signature_28.png | Annie D. Repayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5728931420142527585/5728931420142527585_signature_31.png
| 2023-10-12 | 2023-10-10 13:43 | Lung Center of the Philippines | Quezon Avenue, QC | NCR | NCR 2nd - Quezon City | TB lab | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26215 | PBSP | nilabis.pbsp@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 348 |  | GX4 R1 | 704584 | Due for maintenance, B4 encounter temp error 4009 | Cleaned plunger, i-core, cartridge bay, Gathered reports Test run unit | Maintenance | B4 disabled, 3 Module operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728930980143325149/5728930980143325149_signature_28.png | Norman Labis | https://www.jotform.com/uploads/macarewecare/222927978372067/5728930980143325149/5728930980143325149_signature_31.png
| 2023-10-12 | 2023-10-10 13:38 | Lung Center of the Philippines | Quezon Avenue, QC | NCR | NCR 2nd - Quezon City | TB lab | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26214 | PBSP | nilabis.pbsp@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 910 | 2023-10-10 | GX4 R1 | 707450 | For Module replacement C4 686446 - 210102141 | Performed Module replacement Cleaned plunger, i-core, cartridge bay, Gathered reports Test run unit | Module replacement and maintenance | Operational |  | 900-0508 | O |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728930410146644952/5728930410146644952_signature_28.png | Norman Labis | https://www.jotform.com/uploads/macarewecare/222927978372067/5728930410146644952/5728930410146644952_signature_31.png
| 2023-10-12 | 2023-10-10 10:17 | Philippine heart center | East avenue, QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 26213 | MACARE-HBDC | baluyotjeric@gmail.com | 28 |  |  | MCS + | 11H204 | Due for maintenance | Cleaned unit Inspect roller Checked program parameters Adjust line sensor value Checked dpm, spm, valves, centrifuge Performed simulation running Performed functionality test Performed maintenance | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728929870145160260/5728929870145160260_signature_28.png | Jeric Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5728929870145160260/5728929870145160260_signature_31.png
| 2023-10-12 | 2023-10-10 09:47 | Philippine heart center | East Avenue, QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 26212 | MACARE-HBDC | baluyotjeric@gmail.com | 28 |  |  | MCS+ | 20168 | Due for maintenance | Cleaned unit Inspect roller Checked program parameters Adjust line sensor value Checked dpm, spm, valves, centrifuge Performed simulation running Performed functionality test Performed maintenance | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728928850143316649/5728928850143316649_signature_28.png | Jeric Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5728928850143316649/5728928850143316649_signature_31.png
| 2023-10-12 | 2023-10-12 10:21 | Lanao Del Norte Provincial Hospital | Lower Sagadan, Baroy Lanao Del Norte | 10 | Lanao del Norte | Tb Laborartory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 26211 | PBSP | annamariebroce31@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1720 |  | Gx4 | 807673 | Module 1 - 662770 Module 2- 662772 and Module 3 - 662773 encounter hardware failure Error 1001 and 1002 | Replaced module | Replaced new module 1 - 210102178, module 2 - 210102162 and module 3 - 210101969 all new | All units modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728868867217868544/5728868867217868544_signature_28.png | Anne Marie B. Tubio | https://www.jotform.com/uploads/macarewecare/222927978372067/5728868867217868544/5728868867217868544_signature_31.png
| 2023-10-12 | 2023-10-12 10:15 | Matalom Community Hospital | Matalom Leyte | 8 | Leyte | TB-DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26210 | WHO | fegorne@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 112 ASSAY: Xpert MTB/RIF Ultra, COUNT: 340 | 2023-10-12 | GXIV | 847227 | #NAME? | #NAME? | #NAME? | Unit is Operational |  | LCD monitor BRAND: Lenovo MODEL: ThinkVision S22e-20 SERIAL NUMBER: V5FT169 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728866569514822627/5728866569514822627_signature_28.png | FE GORNE | https://www.jotform.com/uploads/macarewecare/222927978372067/5728866569514822627/5728866569514822627_signature_31.png
| 2023-10-12 | 2023-10-12 10:11 | Binan City Health Office 2 | Binan Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 26209 | DOH | montanezaira016@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 126 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1858 ASSAY: Xpert HIV-1 Viral Load, COUNT: 396 |  | GXIV | 821278 | Unit for PM. For Xpertcheck of module A3. | Performed PM. Performed Xpertcheck. Genersted reports. | PM and Xpertcheck. | Unit is operational. Module A3, service required. | Used 2 spare Xpertcheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728863514718698687/5728863514718698687_signature_28.png | Jasmine Aira Montañez | https://www.jotform.com/uploads/macarewecare/222927978372067/5728863514718698687/5728863514718698687_signature_31.png
| 2023-10-12 | 2023-10-10 15:42 | Caibiran Rural Health Unit | Caibiran, Biliran | 8 | Biliran | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26208 | PBSP | ednaobaletran07@gmail.com | 29 |  | 2023-10-11 | GXIV | 834815 | Module B2/642814 for replacement | #NAME? | - Install new module on B2 w/ SN: 210102188 - Performed Preventive maintenance - Performed Functionality test - Generate reports - Pullout defective module SN 642814 | Unit is Operational |  | Genexpert Module 6 color SN: 210102188 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728859719518314132/5728859719518314132_signature_28.png | Edna O Letran | https://www.jotform.com/uploads/macarewecare/222927978372067/5728859719518314132/5728859719518314132_signature_31.png
| 2023-10-12 | 2023-10-12 09:35 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26207 | PBSP | cdh.laboratorydept@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1144 ASSAY: Xpert HIV-1 Viral Load, COUNT: 27 |  | GXIV-D | 807562 | System lag while running Unable to stop tests | Troubleshooting Preventive maintenance | Performed preventive maintenance Troubkeshoot error Gathered reports Performed functiinality test | Unit is operational except module 4 | Module 4 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728853680525002745/5728853680525002745_signature_28.png | Rowena Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5728853680525002745/5728853680525002745_signature_31.png
| 2023-10-12 | 2023-10-10 14:24 | Bacolod City Health Office | Burgos St., Bacolod City | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 26206 | PBSP | cyrilcordova0515@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6285 |  | GX4 | 812020 | N/A | Pulled out the old module for D1 with SN 640561. Save IQR files before and after replacement. | Installed the new module for D1 with SN 210102186. | All units are now functionall. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728772027915027708/5728772027915027708_signature_28.png | Mr. Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5728772027915027708/5728772027915027708_signature_31.png
| 2023-10-12 | 2023-10-10 15:42 | Caibiran Rural Health Unit | Caibiran, Biliran | 8 | Biliran | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26205 | PBSP | ednaobaletran07@gmail.com | 29 |  | 2023-10-11 | GXIV | 834815 | Module B2/642814 for replacement | #NAME? | - Install new module on B2 w/ SN: 210102188 - Performed Preventive maintenance - Performed Functionality test - Generate reports - Pullout defective module SN 642814 | Unit is Operational |  | Genexpert Module 6 color SN: 210102188 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728662029872337385/5728662029872337385_signature_28.png | Edna O Letran | https://www.jotform.com/uploads/macarewecare/222927978372067/5728662029872337385/5728662029872337385_signature_31.png
| 2023-10-11 | 2023-10-11 15:41 | Philippine Red Cross Ormoc Chapter | Ormoc, Leyte | 8 | Leyte | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  |  |  |  | 26204 | COMMERCIAL | ormoc@redcross.org.ph | 28 |  |  | MCS + | 15H43 | Defective Line Sensor | perform Diagnostics and Install Line Sensor | Install Temporary Line Sensor | Systems Operational | Subject for Qoutation (out of warranty) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728257194225032436/5728257194225032436_signature_28.png | Venus D. Coching | https://www.jotform.com/uploads/macarewecare/222927978372067/5728257194225032436/5728257194225032436_signature_31.png
| 2023-10-11 | 2023-10-11 15:49 | Lubang Rural Health Unit | Binacas Rd. | 4B | Occidental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26203 | DOH | lalaine.batalla0821@gmail.com | 25 |  | 2023-07-03 | GX-IV-D-6C | 838228 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Run XpertCheck Generated reports | All modules operational. All modules passed the XpertCheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728219063504367447/5728219063504367447_signature_28.png | Lalaine.V.Batalla | https://www.jotform.com/uploads/macarewecare/222927978372067/5728219063504367447/5728219063504367447_signature_31.png
| 2023-10-11 | 2023-10-11 14:40 | Caraga Regional Hospital | Rizal Street, Brgy. Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 26202 | PBSP | abpabellan.pbsp@gmail.com | 24.08 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5369 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 |  | GX-4 | 815489 | D1 skipped in XPC Calibration (not detected) D2 retry then service required | Reported to Cepheid | XPC Calibration Preventive Maintenance | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728170163433200572/5728170163433200572_signature_28.png | Andrew B. Pabellan | https://www.jotform.com/uploads/macarewecare/222927978372067/5728170163433200572/5728170163433200572_signature_31.png
| 2023-10-11 | 2023-10-11 14:38 | Mlang Rural Health Unit | Mlang, North Cotabato | 12 | Cotabato | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26201 | PBSP | kathlynlacabe7@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1325 | 2023-10-11 | Gx-IV | 815434 | Module 3 - H/W failed | Perform Genexpert machine Preventive maintenance | Replace defective module 3 SN: 640319 with new module SN:210102175 Genexpert repositioning due to high temperature | 4 operational modules |  | Genexpert module 6 color SN:210102175 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728168405181414978/5728168405181414978_signature_28.png | Kathlyn mae y, lacabe | https://www.jotform.com/uploads/macarewecare/222927978372067/5728168405181414978/5728168405181414978_signature_31.png
| 2023-10-11 | 2023-10-11 13:11 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | TB Dots Laboratory | YES | MAINTENANCE TROUBLESHOOTING Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 26200 | DOH | zionclydematote@gmail.com | 23 |  | 2023-10-11 | Genexpert IV | 831718 | Module A1 temperature error | Performed xpertcheck on module A1 = Passed | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module A1 for observation | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728113760474955687/5728113760474955687_signature_28.png | Sunshine Matote | https://www.jotform.com/uploads/macarewecare/222927978372067/5728113760474955687/5728113760474955687_signature_31.png
| 2023-10-11 | 2023-10-11 12:57 | Alabat Island District Hospital | Perez Rd., Alabat, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26199 | DOH | letpiquero@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 700 |  | GX IV 6C | 831614 | Module C4 stuck cartridge | Performed PM Removed stuck cartridge Generated report | Removed stuck cartridge | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728102155899087248/5728102155899087248_signature_28.png | Carmelita Piquero | https://www.jotform.com/uploads/macarewecare/222927978372067/5728102155899087248/5728102155899087248_signature_31.png
| 2023-10-11 | 2023-10-11 12:18 | Quezon City Jail | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  |  |  |  | 26198 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 838881 | Unit for installation | #NAME? | Installed unit | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728072252517688654/5728072252517688654_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5728072252517688654/5728072252517688654_signature_31.png
| 2023-10-11 | 2023-10-11 11:35 | Calaanan Health Center | Nadurata St. Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 26197 | DOH | rosaleslorina0917@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1325 |  | GXIV | 831892 | Unit for PM. For Xpertcheck of module A4. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck | Unit is operational. | Used 2 spare Xpertchecks. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728063694612244988/5728063694612244988_signature_28.png | Lorina Rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/5728063694612244988/5728063694612244988_signature_31.png
| 2023-10-11 | 2023-10-11 11:26 | West Visayas State University Medical Center | Jaro Iloilo City | 6 | Iloilo | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 26196 | PBSP | tvanrlujan@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2219 |  | GX4 | 707531 | N/A | Pulled out the old module for C4 with SN 696407. Save IQR file before and after replacement. | Installed the new module for C4 with SN 210102140 | All modules are now functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728044572427437858/5728044572427437858_signature_28.png | Van Ryan Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5728044572427437858/5728044572427437858_signature_31.png
| 2023-10-11 | 2023-10-11 10:42 | Dr. Paulino J. Garcia Memorial Research MC | Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26195 | PBSP | jaballakarla@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 822 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4508 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 251 |  | GXXVI-D | 844111 | B3 for module replacement | Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728024823988304216/5728024823988304216_signature_28.png | Karla Mae Jaballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5728024823988304216/5728024823988304216_signature_31.png
| 2023-10-11 | 2023-10-11 10:51 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26194 | MACARE-HBDC | ruellamorena@gmail.com | 22 |  | 2023-09-28 | Genexpert IV | 110008457 | Module 3 error 1002 | Performed Module Replacement | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module Replacement = (210034758 - 614353) | Genexpert Module | 10694 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5728022998223859527/5728022998223859527_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5728022998223859527/5728022998223859527_signature_31.png
| 2023-10-11 | 2023-10-11 09:50 | Cagayan Valley Center for Health Development (DOH R02) | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26193 | PBSP | mctorio.pbsp@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1365 ASSAY: Xpert SARS-CoV-2, COUNT: 3170 | 2023-10-11 | Genexpert IV | 834434 | Module B3 error | Performed Module Replacement | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module Replacement = (616313 - 210102183) | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727991161519783078/5727991161519783078_signature_28.png | Mikeala Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5727991161519783078/5727991161519783078_signature_31.png
| 2023-10-11 | 2023-10-11 09:23 | Alfonso Rural Health Unit | Burgos St., Alfonso, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26191 | DOH | aim_1225@yahoo.com | 25 |  | 2023-09-19 | GXIV-D | 821279 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | All Modules are Operational | All Modules Passed | XpertChecl Kit | DR 10677 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727970755795705780/5727970755795705780_signature_28.png | Aim Arroyo | https://www.jotform.com/uploads/macarewecare/222927978372067/5727970755795705780/5727970755795705780_signature_31.png
| 2023-10-11 | 2023-10-10 11:07 | Jose Reyes Memorial Medical Center | Rizal Ave., Manila | NCR | NCR 1st - Manila | Laboratory | NO | Software Upgrade | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26192 | DOH | eeps101487sambolledo@yahoo.com | 28 |  |  | GXIV-D | 831530 | Unit is due for Software Upgrade | Software Upgrade | Upgrade Software from 4.8 to 6.5 Performed Self-test Checked Functionality Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727970765794498716/5727970765794498716_signature_28.png | Ervin Emmett Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/5727970765794498716/5727970765794498716_signature_31.png
| 2023-10-10 | 2023-10-10 16:28 | Bangkal Health Center | Rodriguez Ave. | NCR | NCR 4th - Makati | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26190 | PBSP | kapid5sepe@yahoo.com | 24 |  |  | GX-IV-D-6C | 838875 | For Installation. | Installed the GeneXpert unit. | Installed the GeneXpert unit. (Unit only) Performed maintenance Functionality test Generated reports | Unit is operational. | Unit from West Rembo. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727366357128288446/5727366357128288446_signature_28.png | Eduardo O. Sepe | https://www.jotform.com/uploads/macarewecare/222927978372067/5727366357128288446/5727366357128288446_signature_31.png
| 2023-10-10 | 2023-10-10 15:59 | Bangkal Health Center | Rodriguez Ave | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26189 | DOH | kapid5sepe@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1695 |  | GX-IV-D-6C | 831875 | Module A1 and A2 has errors 2014 | Run XpertCheck on modules A1 and A2 | Performed maintenance Run XpertCheck on modules A1 and A2 Generated reports | Unit is operational. Module A1/210020089 and A2/862667 failed in XpertCheck, not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727365257122897337/5727365257122897337_signature_28.png | Eduardo O. Sepe | https://www.jotform.com/uploads/macarewecare/222927978372067/5727365257122897337/5727365257122897337_signature_31.png
| 2023-10-10 | 2023-10-10 15:52 | Angeles Ryral Health Unit 3 | Angeles City Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26188 | PBSP | femiepangilinan@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2566 ASSAY: Xpert HIV-1 Viral Load, COUNT: 683 | 2023-09-29 | GXIV-D | 807569 | B2 for module replacement | Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 43633 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727337717222908899/5727337717222908899_signature_28.png | Femie Pangilinan | https://www.jotform.com/uploads/macarewecare/222927978372067/5727337717222908899/5727337717222908899_signature_31.png
| 2023-10-10 | 2023-10-10 15:22 | Matnog Medicare Community Hospital | Purok 3, Gadgaron | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26187 | PBSP | mtmh@sorsogon.gov.ph | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 797 |  | GX4 | 835200 | B1/210091134 & B2/789045 for module replacement | Module replaced: B1 - 210091134>>210101968, B2 - 789045>>210100204 Cleaned fan Generate reports | Replaced B1 & B2 module Performed Preventive Maintenance | Unit is operational 1 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727320995214053679/5727320995214053679_signature_28.png | Mernadith Garcera | https://www.jotform.com/uploads/macarewecare/222927978372067/5727320995214053679/5727320995214053679_signature_31.png
| 2023-10-10 | 2023-10-10 15:09 | Amma Jadsac District Hospital | Poblacion, Pudtol, Apayao | CAR | Apayao | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 26186 | PBSP | rheamarierodriguez.con@gmail.com | 24 |  | 2023-10-10 | Genexpert IV | 807609 | Module 3 not detected, module 3 error 4008 | Reset Gateway board Swap module 2 and 3 Test the Fuse 1 of module 3 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 3 for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727313412169297080/5727313412169297080_signature_28.png | Rhea Marie Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5727313412169297080/5727313412169297080_signature_31.png
| 2023-10-10 | 2023-10-10 13:51 | Compostela Rural Health Unit | Poblacion, Compostela Davao De Oro | 11 | Davao de Oro | TB-DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26185 | PBSP | nielmaquiling20@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1212 |  | GX-IV 4D | 807360 | Error at system unit, user is stucked at startup boot only Error at UPS, Due for Preventive Maintenance | Replaced the CMOS battery, Clean the ram slot Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses | System unit troubleshooting Perform Preventive Maintenance | Error at System Unit still persist after how many retry to replaced new cmos battery. UPS Battery is defective , GeneXpert System Immediately shutoff when at battery mode |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727277801025952401/5727277801025952401_signature_28.png | Niel Maquiling | https://www.jotform.com/uploads/macarewecare/222927978372067/5727277801025952401/5727277801025952401_signature_31.png
| 2023-10-10 | 2023-10-10 13:49 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb-Lab | YES | CALIBRATION TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 26184 | PBSP | Celinejean16@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 853 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 |  | Gx-IV | 815428 | Genexpert CPU Problem | Perform Genexpert machine CPU troubleshooting | Check CPU power cable - functional Replace CMOS battery Unplug Ram and move to other pin CPU problem reported | Genexpert CPU subject for replacement due to power supply problem | Set-up and Network two Genexpert Unit to other Computer |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727276067715898141/5727276067715898141_signature_28.png | Celine Jean villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5727276067715898141/5727276067715898141_signature_31.png
| 2023-10-10 | 2023-10-10 13:36 | Bondoc Peninsula District Hospital | Brgy. 8, Catanauan, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26183 | PBSP | trining136@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 956 | 2023-10-04 | GX IV 6C | 805339 | Module B3 for replacement | Performed PM Replaced module B3 Generated report | Replaced B3 210068501 -> 210102144 | Unit is operational All modules working |  | GX module 6C | DR 43712 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727259790614136535/5727259790614136535_signature_28.png | Trinity Carisse Puyal | https://www.jotform.com/uploads/macarewecare/222927978372067/5727259790614136535/5727259790614136535_signature_31.png
| 2023-10-10 | 2023-10-10 13:35 | Orani Rural Health Unit | Orani Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26182 | DOH | oranirhu@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 560 |  | GXIV-D | 831668 | A1 for Re-xpertcheck | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except A1 | A1 failed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727254069816703053/5727254069816703053_signature_28.png | Belinda C. Guinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5727254069816703053/5727254069816703053_signature_31.png
| 2023-10-10 | 2023-10-10 13:31 | Rosario Super Health Center | Dr. Sixto Avenue | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26181 | WHO | eden.cawayan@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2028 | 2023-10-10 | GX-IV-D-6C | 843924 | Module B4 for replacement | Replaced module B4 | Performed maintenance Replaced module B4 Generated reports | All modules operational. |  | Old: B4/210068082 New: B4/210102067 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727253927129892238/5727253927129892238_signature_28.png | Eden R. Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5727253927129892238/5727253927129892238_signature_31.png
| 2023-10-10 | 2023-10-10 11:25 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26180 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 25 |  |  | GX-XVI-D-6C | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod. Generated reports Checked module A2 | Unit is operational. Modules A1, B2, B3, B4 and C2 not operational, reported already. Module A2 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727249917127427045/5727249917127427045_signature_28.png | Erika Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5727249917127427045/5727249917127427045_signature_31.png
| 2023-10-10 | 2023-10-10 11:20 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26179 | PBSP | molecularlab@rmc.doh.gov.ph | 25 |  | 2023-10-10 | GX-IV-D-6C | 81437 | Module B1 for replacement | Replaced module B1/210100887 | Performed maintenance Replaced module B1 Generated reports | Unit is operational. Module B2 not operational, reported already. |  | Old: B1: 210100887 New: B1/743374 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727249897128599779/5727249897128599779_signature_28.png | Erika Mendoza, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5727249897128599779/5727249897128599779_signature_31.png
| 2023-10-10 | 2023-10-10 11:00 | San Fernando Rural Health Unit 2 | Sindalan San Fernando Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26178 | PBSP | punoricahrd@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1186 | 2023-10-04 | GXIV-D | 838879 | B3 for module replcement | Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 43711 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727246449813202524/5727246449813202524_signature_28.png | Richard M. Puno | https://www.jotform.com/uploads/macarewecare/222927978372067/5727246449813202524/5727246449813202524_signature_31.png
| 2023-10-10 | 2023-10-10 12:27 | Sorsogon Provincial Hospital | Sorsogon City | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26177 | PBSP | wizard682000@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 995 |  | GX4 | 835213 | C2/633491 for module replacement | Replaced module: C2 - 633491>>210102145 Cleaned cartridge bays, plunger rod, fan, icore Generate reports | Replaced C2 module Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727215269809046988/5727215269809046988_signature_28.png | Reynaldo B. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5727215269809046988/5727215269809046988_signature_31.png
| 2023-10-10 | 2023-10-09 16:56 | Malilipot Rural Health Unit | Malilipot | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26176 | PBSP | fiath.buban2124@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1900 |  | GX4 | 808280 | B2/625608 for Module Replacement | Module Replaced: B2 - 625608>>210102150 Cleaned Cartridge bays, plunger rod, fan, icore Generate reports | Replaced B2 Module | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5727211949807604804/5727211949807604804_signature_28.png | Rechel B. Canomoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5727211949807604804/5727211949807604804_signature_31.png
| 2023-10-09 | 2023-10-09 16:56 | Tanay Rural Health Unit | East Rd. | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26175 | PBSP | maryjunnette@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1359 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-IV-D-6C | 807593 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726511493614934645/5726511493614934645_signature_28.png | Mary Junette Fulgueras | https://www.jotform.com/uploads/macarewecare/222927978372067/5726511493614934645/5726511493614934645_signature_31.png
| 2023-10-09 | 2023-10-09 16:48 | Department of Health - Treatment and Rehabilitation Center | Camp Bagong Diwa | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26174 | PBSP | dohtrcbicutan.lab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2167 |  | GX-IV-D-6C | 807452 | Module A2 not detected | Change module A2's lan cable | Performed maintenance Checked module A2 Checked connections Swap module A2 to A1 Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726508703619711081/5726508703619711081_signature_28.png | April Chakee Javier | https://www.jotform.com/uploads/macarewecare/222927978372067/5726508703619711081/5726508703619711081_signature_31.png
| 2023-10-09 | 2023-10-09 16:42 | International Organization for Migration | HV Dela Costa | NCR | NCR 4th - Makati | Laboratory | NO | Relocation | CEPHEID | John Monsour Doctor |  |  |  |  | 26173 | IOM | jsabalboro@iom.int | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 986 ASSAY: Others, COUNT: 142 |  | GX-IV-L-6C | 110002868 | For relocation | Relocated the unit | Relocated the GeneXpert unit Performed maintenance Functionality test Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726504163619872453/5726504163619872453_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5726504163619872453/5726504163619872453_signature_31.png
| 2023-10-09 | 2023-10-09 16:34 | West Rembo Health Center | JP Rizal St. | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26172 | DOH | danise_er_19@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1174 |  | GX-IV-D-6C | 831712 | Module A2 for XpertCheck | Run XpertCheck on module A2 | Performed maintenance Run XpertCheck on module A2 Generated reports | Unit is operational. Module A2 passed in XpertCheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726498953618377511/5726498953618377511_signature_28.png | Danis Maebel Calimag | https://www.jotform.com/uploads/macarewecare/222927978372067/5726498953618377511/5726498953618377511_signature_31.png
| 2023-10-09 | 2023-10-09 16:27 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26171 | DOH | gemmponce11@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 622 |  | GX-IV-D-6C | 821292 | CPU does not power on sometimes | Checked CPU | Checked CPU Checked connections Checked UPS Performed maintenance Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726495493616147610/5726495493616147610_signature_28.png | Gemma Ponce | https://www.jotform.com/uploads/macarewecare/222927978372067/5726495493616147610/5726495493616147610_signature_31.png
| 2023-10-09 | 2023-10-09 16:14 | International Organization for Migration | HV Dela Costa | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26170 | IOM | jsabalboro@iom.int | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 986 ASSAY: Others, COUNT: 142 ASSAY: Xpert SARS-CoV-2, COUNT: 16 |  | GX-IV-L-6C | 805288 | For relocation | Relocated the unit | Relocated the GeneXpert unit Performed maintenance Functionality test Generate reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726491043613778249/5726491043613778249_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5726491043613778249/5726491043613778249_signature_31.png
| 2023-10-09 | 2023-10-09 16:14 | International Organization for Migration | HV Dela Costa | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26169 | IOM | jsabalboro@iom.int | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 986 ASSAY: Others, COUNT: 142 ASSAY: Xpert SARS-CoV-2, COUNT: 16 |  | GX-IV-L-6C | 805288 | For relocation | Relocated the unit | Relocated the GeneXpert unit Performed maintenance Functionality test Generate reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726490283613812838/5726490283613812838_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5726490283613812838/5726490283613812838_signature_31.png
| 2023-10-09 | 2023-10-09 14:38 | Department of Health - Treatment and Rehabilitation Center | Camp Bagong Diwa | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26168 | PBSP | dohtrcbicutan.lab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2122 | 2023-09-05 | GX-IV-D-6C | 807452 | For module C3 for replacement | Replaced module C3 | Performed maintenance Replaced defective module Generated reports | Unit is operational. |  | Old: C3/801912 New: C3/210102242 | G 6265 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726435253614733966/5726435253614733966_signature_28.png | April Chakee Javier | https://www.jotform.com/uploads/macarewecare/222927978372067/5726435253614733966/5726435253614733966_signature_31.png
| 2023-10-09 | 2023-10-02 10:40 | South Cotabato Provincial Hospital | Koronadal City, South Cotabato | 12 | South Cotabato | Tb-lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26166 | PBSP | anamariermt22@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2102 ASSAY: Xpert HIV-1 Viral Load, COUNT: 501 | 2023-10-02 | Gx-IV | 815380 | Module 4 H/W failed | Perform Genexpert machine Preventive maintenance | Replace module 4 SN:684397 with new module SN: 210060050 | 4 operational modules |  | Genexpert module 6 color SN: 210060050 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726345684311488500/5726345684311488500_signature_28.png | ANA MARIE GRACE EGASAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5726345684311488500/5726345684311488500_signature_31.png
| 2023-10-09 | 2023-10-05 13:43 | Lake Sebu Rural Health Unit | Lake Sebu, South Cotabato | 12 | South Cotabato | Tb-Lab | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 26167 | DOH | jennypingkian23@yahoo.com | 28.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 504 |  | Gx-IV | 838569 | Module 4 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance | Perform Xpertcheck at module 4 for troubleshooting Report module 4 problem | 3 operational modules | Module 4 for Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726345684317576232/5726345684317576232_signature_28.png | Jenny A. Pingki an | https://www.jotform.com/uploads/macarewecare/222927978372067/5726345684317576232/5726345684317576232_signature_31.png
| 2023-10-09 | 2023-10-06 10:48 | Silang Rural Health Unit | Del Pilar St., Silang, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26164 | DOH | bencq78@gmail.com | 25 |  | 2023-09-19 | GXIV-D | 820708 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed Self-test Performed XpertCheck Generate Reports | 3 Modules are Operational, 1 for Replacement |  | XpertCheck Kit | DR 10675 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726225603619013378/5726225603619013378_signature_28.png | Benessie Madlangsakay | https://www.jotform.com/uploads/macarewecare/222927978372067/5726225603619013378/5726225603619013378_signature_31.png
| 2023-10-09 | 2023-10-06 12:42 | Carmona Rural Health Unit | Loyola St, Carmona,Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26165 | DOH | genexpert.carmona@gmail.com | 25 |  | 2023-09-19 | GXIV-D | 821216 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed Self-test Performed XpertCheck Checked UPS Function Generate Reports | UPS is Defective. All Modules are Operational. |  | XpertChecl Kit | DR 10674 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5726225603619613240/5726225603619613240_signature_28.png | Noreen Fortu | https://www.jotform.com/uploads/macarewecare/222927978372067/5726225603619613240/5726225603619613240_signature_31.png
| 2023-10-08 | 2023-10-06 08:27 | Aguilar Rural Health Unit | Aguilar, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26163 | PBSP | ndpaguilar2014@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 428 | 2023-10-06 | GX4 2 placer | 834369 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B2 784155 >>> 210102173 - Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210102173 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5725682258718083388/5725682258718083388_signature_28.png | Vilma D. Pastor | https://www.jotform.com/uploads/macarewecare/222927978372067/5725682258718083388/5725682258718083388_signature_31.png
| 2023-10-06 | 2023-10-06 15:35 | National kidney and transplant institute | East avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 26162 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 852 ASSAY: Others, COUNT: 23 ASSAY: Xpert HBV VL, COUNT: 577 ASSAY: Xpert HPV, COUNT: 113 ASSAY: Xpert MTB/RIF, COUNT: 1161 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1209 ASSAY: Xpert SARS-CoV-2, COUNT: 14570 ASSAY: Xpert HCV VL, COUNT: 53 ASSAY: Others, COUNT: 9 |  | GX16 | 110005587 | A4 unable close the door and encounter 4017 temp error | Checked door latch, Performed self test multiple time but error 4017 still occur Gathered reports | Troubleshooting | 14 module operational A1 reported,A3 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723875379833706499/5723875379833706499_signature_28.png | Danielle Marie M. Cortes | https://www.jotform.com/uploads/macarewecare/222927978372067/5723875379833706499/5723875379833706499_signature_31.png
| 2023-10-06 | 2023-10-06 14:04 | Catmon Health Center | Catmon, Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 26161 | DOH | mariathes0218@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2982 ASSAY: Xpert HIV-1 Viral Load, COUNT: 149 | 2023-10-06 | GXIV | 831903 | Module C2 for re XpertCheck | #NAME? | Performed XpertCheck on C2 | Unit is operational. C2: service required |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723852368705697584/5723852368705697584_signature_28.png | Maria Theresa Nunez | https://www.jotform.com/uploads/macarewecare/222927978372067/5723852368705697584/5723852368705697584_signature_31.png
| 2023-10-06 | 2023-10-06 14:09 | Lower Bicutan Health Center | Lower Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26160 | PBSP | ellainemagtoto35@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2967 |  | GX-IV-D-6C | 838858 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723840247114224208/5723840247114224208_signature_28.png | Ellaine Magtoto | https://www.jotform.com/uploads/macarewecare/222927978372067/5723840247114224208/5723840247114224208_signature_31.png
| 2023-10-06 | 2023-10-06 11:11 | Tropical disease foundation | Amorsolo st | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION | CEPHEID | Noel Floresca |  |  |  |  | 26159 | PBSP | aageronimo@tdf.org.ph | 22 | ASSAY: Xpert MTB/RIF, COUNT: 200 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1466 |  | GX4 | 838134 | Due for xpert check | Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Test run unit | Xpertcheck | 3 module operation b4 failed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723823499831500745/5723823499831500745_signature_28.png | Anthony Geronimo | https://www.jotform.com/uploads/macarewecare/222927978372067/5723823499831500745/5723823499831500745_signature_31.png
| 2023-10-06 | 2023-10-06 10:40 | Tropical disease foundation | Amorsolo st | NCR | NCR 4th - Makati | Laboratory | NO | INSTALLATION | CEPHEID | Noel Floresca |  |  |  |  | 26158 | COMMERCIAL | aageronimo@tdf.org.ph | 23 |  |  | GX4 10C | 120000022 | For installation | Installed unit,complete accessories | Installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723819329834773146/5723819329834773146_signature_28.png | Anthony Geronimo | https://www.jotform.com/uploads/macarewecare/222927978372067/5723819329834773146/5723819329834773146_signature_31.png
| 2023-10-06 | 2023-10-05 08:38 | Rodriguez RHU | J.P Rizal Rodriguez,Montalban Rizal | 4A | Rizal | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 26157 | PBSP | cchlaboratory2018@yahoo.com | 26 |  | 2023-10-03 | GX4 | 838894 | For module replacement B1 863306 - 210102189 | Performed module replacement Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Test run unit | Module replacement | Operational |  | 900-0491rs | G6327 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723818669839741552/5723818669839741552_signature_28.png | Mykaela Josepha De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5723818669839741552/5723818669839741552_signature_31.png
| 2023-10-06 | 2023-10-06 11:58 | Bongabon District Hospital | Curva, Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 26156 | DOH | tsantuajr0128@gmail.com | 26 |  | 2023-10-06 | Genexpert IV | 831562 | Module 1 error 2037 | Performed Xpertcheck on module 1 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck Result = Passed | Xpertcheck 2pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723781859319432631/5723781859319432631_signature_28.png | Teodorico Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5723781859319432631/5723781859319432631_signature_31.png
| 2023-10-06 | 2023-10-06 09:40 | Corazon Locsin Montelibano Memorial Regional Hospital | Lacson St., Bacolod City | 6 | Negros Occidental | Immunology/Serology Department | NO | INSTALLATION | CEPHEID | Emerson Bejemino |  |  |  |  | 26155 | MACARE-HBDC | hact.clmmrh@gmail.com | 24 |  |  | GX-4 10C | 110020743 | New Genexpert 10c unit for installation. | n/a | - Install new 10c unit and accessories (Desktop, LCD Monitor, AVR, UPS, Printer) - CPU SN CPH0723F43822 - System configuration - Import Assay Definition Files (ADF) - Generate Installation Qualification Report - Orientation of end users on machine operation and maintenance. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723746554217963178/5723746554217963178_signature_28.png | CHERRY ANNE IBAÑEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5723746554217963178/5723746554217963178_signature_31.png
| 2023-10-06 | 2023-10-06 10:29 | Taguig Pateros District Hospital | East Service Rd | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26154 | DOH | kristinejane.kjd25@gmail.com | 25 |  |  | GX-IV-D-6C | 831638 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723690882616225863/5723690882616225863_signature_28.png | Kristine Jane D. Base | https://www.jotform.com/uploads/macarewecare/222927978372067/5723690882616225863/5723690882616225863_signature_31.png
| 2023-10-06 | 2023-10-06 09:51 | Carcar City Health Office | Valladollid, Carcar City | 7 | Cebu | Tb Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 26153 | DOH | chaska.lp14@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1644 |  | Gx4 | 838283 | All modules encountered Error 2008 as od Sept 2023 and 6 errors starting October 2023 | Performed maintenance, and reorient end user for specimen and reagent preparation Check the cartridge before using | Performed maintenance, valve cleaning plunger maintenance, Icore clean up | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723666725961827739/5723666725961827739_signature_28.png | Lucille Mae A. Potoy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5723666725961827739/5723666725961827739_signature_31.png
| 2023-10-05 | 2023-10-05 20:52 | SEBASTE RURAL HEALTH UNIT | SEBASTE, ANTIQUE | 6 | Antique | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26152 | PBSP | francisconaciongayo@gmail.com | 39 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1785 |  | GX-IV | 808287 | FOR MODULE B3 REPLACEMENT | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | -PERFORM MODULE B3 REPLACEMENT: DEF SN:688184 NEW MODULE SN:210100218 | EQUIPMENT IS OPERATIONAL | FOR REQUEST OF GX FAN FILTER |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5723199906416271465/5723199906416271465_signature_28.png | Francisco S. Naciongayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5723199906416271465/5723199906416271465_signature_31.png
| 2023-10-05 | 2023-10-05 14:17 | Sumilang Super Health Center | Sumilang | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26151 | PBSP | maricor.obnamia05@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 150 ASSAY: Xpert MTB/RIF Ultra, COUNT: 714 | 2023-10-03 | GX-IV-D_6C | 814613 | CPU replacement | Replaced CPU | Performed maintenance Pull out backup laptop and defective CPU Install new CPU Generated reports | All modules operational. |  | Defective CPU: 2UA7021MQ1 New CPU: CPH0623F42285 | G 6329 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722970505621250713/5722970505621250713_signature_28.png | Maria Corazon Obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/5722970505621250713/5722970505621250713_signature_31.png
| 2023-10-05 | 2023-09-27 11:39 | Detoxicare | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  |  |  |  | 26150 | CONSORTIUM | imeehabaluyas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 |  | GXIV | 110018312 | Unit for installation | #NAME? | Installed unit, GeneXpert GX-IV Dim: 28.2 cm. wide, 30.5 cm. high, 29.7 cm. deep Weight: 11.4 kg. Power Supply: Auto-ranging Rated AC Voltage Range: 100-240V~, 50-60Hz. Mains Supply Fluctuations: up to +-10% of the nominal voltage Transient Over-Voltages: up to 2500V peak Rated Current: 1.4A @100V Fuse Rating: 250V~ T3A | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715958641625990333/5715958641625990333_signature_28.png | Imelda Habaluyas | https://www.jotform.com/uploads/macarewecare/222927978372067/5715958641625990333/5715958641625990333_signature_31.png
| 2023-10-05 | 2023-10-05 13:22 | Imus City Health Office II | Bayan Luma, Imus, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26149 | DOH | amielynnmangulabnan1217@gmail.com | 26 |  | 2023-09-19 | GXIV-D | 821284 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed Selft-test Performed XpertCheck Generate Reports | All Modules are Operational | 2 Modules are Newly Replaced, | XpertCheck Kit | DR 10672 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722938326224166845/5722938326224166845_signature_28.png | Amielynn Mangulabnan | https://www.jotform.com/uploads/macarewecare/222927978372067/5722938326224166845/5722938326224166845_signature_31.png
| 2023-10-05 | 2023-10-04 09:44 | Philippine General Hospital | Pedro Gil, Taft Ave, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26148 | PBSP | amreyes8@up.edu.ph | 24 |  |  | GXIV-D | 834624 | Unit is for Relocation | Relocation | Relocate GeneXpert System Check Functionality Performed Maintenance Generate Reports | 2 are Operational, 2 Modules for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722938326225929631/5722938326225929631_signature_28.png | Angela Ricamela Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5722938326225929631/5722938326225929631_signature_31.png
| 2023-10-05 | 2023-09-18 13:44 | Tuburan Rural Health Unit | Poblacion | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 26147 | PBSP | marielgracia20@gmail.com | 23 |  | 2023-09-14 | Gx4 | 815485 | Module B2 Hardware fail. | Clean all modules | Perform maintenance for all modules | Module B2 with serail number 791790 hardware fail. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722918303917380460/5722918303917380460_signature_28.png | Mariel Gracia | https://www.jotform.com/uploads/macarewecare/222927978372067/5722918303917380460/5722918303917380460_signature_31.png
| 2023-10-05 | 2023-10-05 12:35 | Northern Mindanao Medical Center | Cagayan De Oro City | 10 | Misamis Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 26146 | PBSP | nmmcgenexpress@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 744 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7499 ASSAY: Xpert SARS-CoV-2, COUNT: 10304 ASSAY: Xpert HIV-1 Viral Load, COUNT: 477 |  | GX-16 | 844118 | None | None | Replacing module D3 with SN: 674972 to SN: 210102225 Preventive Maintenance | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722901945806484507/5722901945806484507_signature_28.png | Vidal G. Villas III | https://www.jotform.com/uploads/macarewecare/222927978372067/5722901945806484507/5722901945806484507_signature_31.png
| 2023-10-05 | 2023-10-05 11:47 | Mandaluyong City Health Office | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 26145 | DOH | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2666 ASSAY: Xpert HIV-1 Viral Load, COUNT: 159 |  | GXIV | 831596 | Stuck cartridge on C1 | #NAME? | Manually removed stuck cartridge | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722870752523998037/5722870752523998037_signature_28.png | Rosevi Jalova | https://www.jotform.com/uploads/macarewecare/222927978372067/5722870752523998037/5722870752523998037_signature_31.png
| 2023-10-05 | 2023-10-05 09:30 | Sagay City Health Office | Sagay City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 26144 | PBSP | maravelesdebra@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1499 |  | GX-IV | 838144 | Defective module 210091070/B1 | #NAME? | - Replaced module 210091070/B1 with new module SN 210100321 - Generate Pre- and Post-IQR | - Module 4 encountering temperature-related errors - for report to Cepheid. - Modules 1, 2, and 3 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722858681025717041/5722858681025717041_signature_28.png | DEBRA ABELLAR, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5722858681025717041/5722858681025717041_signature_31.png
| 2023-10-05 | 2023-10-05 10:27 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26143 | DOH | jazer.palasigue30@gmail.com | 27 |  | 2023-09-29 | Genexpert IV | 831637 | UPS not working | Replaced UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | UPS | 10702 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722836655521734871/5722836655521734871_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5722836655521734871/5722836655521734871_signature_31.png
| 2023-10-05 | 2023-10-05 09:04 | Ramos Rural Health Unit | Ramos Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | jadeeh8143@yahoo.com |  | 26142 | DOH | jadeeh8143@yahoo.com | 28 | ASSAY: , COUNT: 19 ASSAY: Xpert MTB/RIF Ultra, COUNT: 484 |  | GXIV-D | 831721 | A3 temperature error | Xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational A3 Passed calibration, m9dupe under observation | Used 1 xpertcheck kit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722779311527142014/5722779311527142014_signature_28.png | Emily Jade Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/5722779311527142014/5722779311527142014_signature_31.png
| 2023-10-05 | 2023-10-05 08:44 | Samboan Rural Health Unit | Samboan Poblacion, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 26141 | DOH | sheneannesenagan@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 387 |  | Gx4 | 838513 | Error 2008 for all modules | Performed maintenance, Archive test, generate Gxx files, IQR, Sys log, assay statistics Orient enduser for specimen handling, proper Reagent time preparation | Performed preventive maintenance, clean Plunger, Icore, and ultrasonic horn and valve Clean filter, check voltage ups and AVR | All modules are available |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722762562631944681/5722762562631944681_signature_28.png | Shene Anne A. Senagan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5722762562631944681/5722762562631944681_signature_31.png
| 2023-10-04 | 2023-10-04 14:26 | Amparo Health Center | Amparo, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 26140 | PBSP | jjpsjasperjoyce@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1374 | 2023-10-03 | GXIV | 838873 | Module B3 for replacement Unit for calibration | #NAME? | Replaced module B3 Performed XpertCheck | Unit is operational. B1: passed B2: insufficient cartridge B4: insufficient cartridge |  | GeneXpert Module SN: 210101927 | G 6332 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722119404226673512/5722119404226673512_signature_28.png | Jasper Joyce Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/5722119404226673512/5722119404226673512_signature_31.png
| 2023-10-04 | 2023-10-04 11:49 | Midsalip Rural health unit | Midsalip, Zamboanga del Sur | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26139 | DOH | genaliner@rocketmail.com | 26 |  | 2023-08-23 | GX 4 | 837664 | Module 4 h/w failed | Replaced module 4 and module 1, 2 and 3 undergone Xpertcheck calibration all 3 modules passed | Replaced module 769597 with Module 700515 | All modules are available | None | 900-0491RS / XPERTCHECK-CE-5 | 10533 / 10397 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722108821814522756/5722108821814522756_signature_28.png | Genaline A Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5722108821814522756/5722108821814522756_signature_31.png
| 2023-10-04 | 2023-10-04 14:23 | St Lukes Medical Center | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Cellular Immunology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26138 | COMMERCIAL | biomed.qc@stlukes.com.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HPV, COUNT: 261 ASSAY: Xpert HCV VL, COUNT: 3 |  | GX4 | 839234 | Due for maintenance | Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Test run unit | Maintenance | D1,D2 operational,D3,D4 disable failed xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722103148845359376/5722103148845359376_signature_28.png | Nicole Ty | https://www.jotform.com/uploads/macarewecare/222927978372067/5722103148845359376/5722103148845359376_signature_31.png
| 2023-10-04 | 2023-10-04 14:07 | Candaba Rural Health Unit | Laboratory | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26137 | DOH | rochelannlontoc@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 467 | 2023-09-29 | Genexpert IV | 831873 | UPS Problem | Replaced UPS | Performed Maintenance Gathered Reports Performef Functionality Test | Unit is Operational | N/A | UPS | 10701 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722090045522661119/5722090045522661119_signature_28.png | Rochell Ann Lontoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5722090045522661119/5722090045522661119_signature_31.png
| 2023-10-04 | 2023-10-04 13:53 | St.Lukes Medical Center | E.rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26136 | COMMERCIAL | biomed.qc@stlukes.com.ph | 29 | ASSAY: Xpert CT/NG, COUNT: 263 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert HPV, COUNT: 219 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3091 ASSAY: Xpert HCV VL, COUNT: 5 |  | GX4 | 807155 | Due for maintenance | Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722085168843598294/5722085168843598294_signature_28.png | Ma.Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5722085168843598294/5722085168843598294_signature_31.png
| 2023-10-04 | 2023-10-04 13:40 | Toril Health Center | Toril, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26135 | PBSP | leabuelis@yahoo.com | 25.8 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1489 |  | GX-IV 4D | 838865 | Reported Stuck Cartridge at Module B3 | Troubleshoot the Module 3 to remove the cartridge | Perform Module Troubleshooting | Stuck cartridge successfully removed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722075193227513593/5722075193227513593_signature_28.png | Lea M. Buelis, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5722075193227513593/5722075193227513593_signature_31.png
| 2023-10-04 | 2023-10-04 12:55 | Medellin Rural Health Unit | Poblacion, Medellin Cebu | 7 | Cebu | Tb laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 26134 | DOH | medellintdc@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1387 |  | Gx4 | 838506 | Error 1001 in Module 1 and Module 3 | Performed self test in Module 1 and 3 Conduct preventive maintenance | Conduct preventive maintenance in Module 2 and Module 4, Clean Filter, Icore, plunger and cartridge bay Clean filter, fan and casing | All modules are available but only Module 2 and 4 are operating consistently. Module 1 and 3 terminated during loading of cartridges, the medtech suspend the usage to avoid wastage of cartridge Awaiting for module 1 and 3 for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722049855329645163/5722049855329645163_signature_28.png | Raul M. Galano, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5722049855329645163/5722049855329645163_signature_31.png
| 2023-10-04 | 2023-10-04 12:49 | LAMBUNAO RURAL HEALTH UNIT | LAMBUNAO, ILOILO | 6 | Iloilo | LABORATORY | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  |  |  |  | 26133 | PBSP | gabatanelson@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1816 |  | GX-IV | 815279 | MODULE B2 ERROR 2032 | #NAME? | -PERFORM GX EQUIPMENT MAINTENANCE -PERFORM ERROR 2032 ULTRASONIC HORN TROUBLESHOOTING PROCEDURE | EQUIPMENT IS OPERATIONAL | MODULE B2 FOR REPORTING |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722048520414037477/5722048520414037477_signature_28.png | Nelson Gabata, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5722048520414037477/5722048520414037477_signature_31.png
| 2023-10-04 | 2023-10-04 12:39 | National Children's Hospital | E.rodriguez QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | HAEMONETICS | Noel Floresca |  |  |  |  | 26132 | MACARE-HBDC | zarabellehilario@yahoo.com | 28 |  | 2023-09-27 | MCS + | 17H632SPG | For roller replacement | Replaced roller Checked programs parameters. Adjust line sensor value 3071 Performed functional test Performed isolation running | Repair | Unit operational |  | Urethane roller | 10693 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722044208841227232/5722044208841227232_signature_28.png | national children's hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/5722044208841227232/5722044208841227232_signature_31.png
| 2023-10-04 | 2023-10-04 12:02 | Jose B. Lingad Memorial Regional Hospital | City of San Fernando, Pampanga | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26131 | PBSP | jblmghmpl@gmail.com | 24 |  | 2023-10-04 | Genexpert IV | 834445 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5722028325528156882/5722028325528156882_signature_28.png | Kimberly Gulapa | https://www.jotform.com/uploads/macarewecare/222927978372067/5722028325528156882/5722028325528156882_signature_31.png
| 2023-10-04 | 2023-10-04 10:56 | Iligan Society of Internists | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB-Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 26130 | PBSP | dayolaermajessica@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2591 ASSAY: Xpert HIV-1 Viral Load, COUNT: 79 | 2023-10-04 | GXIV-4D | 805090 | Module C2 hardware failed. | -Replace defective module C2/644791 with new module SN: 210102201 | #NAME? | All modules are functional. |  | GeneXpert Module 6-Colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721976188023822188/5721976188023822188_signature_28.png | Erma Jessica Dayola | https://www.jotform.com/uploads/macarewecare/222927978372067/5721976188023822188/5721976188023822188_signature_31.png
| 2023-10-04 | 2023-10-04 10:49 | Iligan Society of Internists | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB-Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 26129 | PBSP | dayolaermajessica@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2591 ASSAY: Xpert HIV-1 Viral Load, COUNT: 79 | 2023-10-04 | GXIV-4D R1 | 707842 | Module D2 and D4 defective | -Replace defective module D2/684440 with new module SN: 210102206 -Replace defective module D4/750600 with new module SN: 210102257 | #NAME? | All modules are functional. |  | 2 GeneXpert Module 6-Colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721973678025699585/5721973678025699585_signature_28.png | Erma Jessica Dayola | https://www.jotform.com/uploads/macarewecare/222927978372067/5721973678025699585/5721973678025699585_signature_31.png
| 2023-10-04 | 2023-10-04 09:55 | Kidapawan City Health Office | Kidapawan City, North Cotabato | 12 | Cotabato | Tb-lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26128 | DOH | sg93rmt@gmail.com | 28.4 |  |  | Gx-IV | 838545 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721948294428803969/5721948294428803969_signature_28.png | Tessa Mae E. Saguindan | https://www.jotform.com/uploads/macarewecare/222927978372067/5721948294428803969/5721948294428803969_signature_31.png
| 2023-10-04 | 2023-10-04 09:28 | Philippine General Hospital | Taft. Ave., Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 26127 | MACARE-HBDC | amreyes8@up.edu.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX16 | 844975 | Unit for PM. For relocation of Genexpert system because of renovation. | Performed PM. Performed installation to new location (Molecular Laboratory Office). Generated reports. | PM and relocation. | Unit is operational. For reporting of modules A3, B1 and D2 with temperature errors. Modules A3, B1 and D2 are confirmed for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721923419463151788/5721923419463151788_signature_28.png | Angela Ricamela Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5721923419463151788/5721923419463151788_signature_31.png
| 2023-10-04 | 2023-10-04 08:49 | Apalit Rural Health Unit | Apalit Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26126 | DOH | duenasblesilda@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 903 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 | 2023-09-28 | GXIV-D | 831900 | A1 for module replacement A3 for stuck cartridge removal | Module replacement Stuck cartridge removal | Performed preventive maintenance Removed stuck cartridge on A3 Performed moduoe replacement on A1 Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 10695 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721900647612414475/5721900647612414475_signature_28.png | Blesilda Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5721900647612414475/5721900647612414475_signature_31.png
| 2023-10-04 | 2023-10-04 07:42 | Vicente Sotto Memorial Medical Center | B Rodriguez St. Capitol, Cebu | 7 | Cebu | Bloodbank Service Unit | YES | REPAIR TROUBLESHOOTING | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 26125 | COMMERCIAL | ivyzeabaid1994@gmail.com | 24 |  |  | MCS+ | 02A090 | UPS encountered long beep sound drain battery | Check ups performed line voltage checking Transfer UPS to other user owned UPS unit | Check Ups and test charge voltage | MCS+ operational with their owned unit. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721861535328554924/5721861535328554924_signature_28.png | IVY ZEA C. BAID | https://www.jotform.com/uploads/macarewecare/222927978372067/5721861535328554924/5721861535328554924_signature_31.png
| 2023-10-03 | 2023-10-03 16:27 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | ILOILO CITY | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26124 | PBSP | tvanrlujan@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2118 |  | GX-IV | 707531 | FOR MODULE C3 REPLACEMENT | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | PERFORM MODULE C3 REPLACEMENT: DEF SN: 677627 NEW MODULE SN: 210102111 | EQUIPMENT IS OPERATIONAL | MODULE C4 PENDING |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721310610411137351/5721310610411137351_signature_28.png | VAN RYAN LUJAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5721310610411137351/5721310610411137351_signature_31.png
| 2023-10-03 | 2023-10-03 16:13 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | ILOILO CITY | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26123 | PBSP | tvanrlujan@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2118 |  | GX-IV | 804699 | FOR MODULE B1 REPLACEMENT | #NAME? | PERFORM MODULE B1 REPLACEMENT : DEF SN: 807206 NEW MODULE SN: 210101844 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721308240419741244/5721308240419741244_signature_28.png | VAN RYAN LUJAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5721308240419741244/5721308240419741244_signature_31.png
| 2023-10-03 | 2023-10-03 15:26 | Altavas Rural Health Unit | Altavas, Aklan | 6 | Aklan | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Emerson Bejemino |  |  |  |  | 26122 | DOH | nycolebenedicto@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 236 |  | GX-4 | 838458 | Due for Calibration/ Xpertcheck. | Preventive Maintenance | - Performed Xpertcheck on all 4 modules | All 4 modules Passed Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721297536268943252/5721297536268943252_signature_28.png | MA. NYCOLE S. BENEDICTO | https://www.jotform.com/uploads/macarewecare/222927978372067/5721297536268943252/5721297536268943252_signature_31.png
| 2023-10-03 | 2023-09-28 11:44 | Department of Health Treatment and Rehabilitation Center | Camp Bagong Diwa | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26121 | PBSP | dohtrcbicutan.lab@gmail.com | 25 |  |  | GX-IV-D-6C | 807452 | Module C2/210100544 not detected | Replaced lan cable of module C2 | Performed maintenance Checked internal connection Swap module C2 to module C1 Replaced lan cable of module C2 | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721289355217696397/5721289355217696397_signature_28.png | hellen joy rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5721289355217696397/5721289355217696397_signature_31.png
| 2023-10-03 | 2023-10-03 15:47 | International Organization for Migration | HV Dela Costa | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26120 | IOM | jsabalboro@iom.int | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1040 ASSAY: Others, COUNT: 152 ASSAY: Xpert SARS-CoV-2, COUNT: 13 |  | GX-IV-L-10C | 110004816 | GeneXpert unit does not power on | Replaced power cord | Checked UPS Checked cable Performed maintenance Functionality test Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721286125211900556/5721286125211900556_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5721286125211900556/5721286125211900556_signature_31.png
| 2023-10-03 | 2023-10-03 13:56 | Adela Serra Ty Memorial Medical Center | Brgy. Telaje, Capitol Hills, Tandag City | 13 | Surigao del Sur | TB-DOTS | NO | REPAIR TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 26119 | PBSP | tbdotsppmdunit@gmail.com | 25 |  |  | GX-4 | 804255 | No display on monitor upon turning on CPU Communication and booting errors of modules | None | Reseating RAM on CPU Replacing CMOS battery Directly attach the lan cable from machine to cpu without a network. | 1 gx machine out of service. Only backup unit working. CPU not functional. Backup unit used (Laptop: ACER E5-471-32HE) | Due to electricty fluctuation in the site, Highly recommend to have a stable electricity. UPS must be replaced due to lack of functionality. (Recommended UPS Brand: Powercom 1KVA UPS Raptor Series) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721284608025405175/5721284608025405175_signature_28.png | Editha M. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5721284608025405175/5721284608025405175_signature_31.png
| 2023-10-03 | 2023-10-03 15:16 | Cotabato Provincial Hospital | Brgy. Amas, Kidapawan City, North Cotabato | 12 | Cotabato | Tb-lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26118 | PBSP | ellainebiongan1203@gmail.com | 25.7 |  | 2023-10-03 | Gx-IV | 807615 | Module 4 for service Required | Perform Genexpert machine Preventive maintenance | Replace defective module 4 SN: 690460 with new module SN: 210102256 | 4 operational modules |  | Genexpert module 6 color SN:210102256 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721272223125195188/5721272223125195188_signature_28.png | Ellaine Grace Biongan | https://www.jotform.com/uploads/macarewecare/222927978372067/5721272223125195188/5721272223125195188_signature_31.png
| 2023-10-03 | 2023-10-03 13:29 | President Ramon Magsasay Memorial Hospital | Iba Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26117 | PBSP | prmmhmoleclab@gmail.com | 30 | ASSAY: Xpert SARS-CoV-2, COUNT: 3566 | 2023-09-22 | GXIV-D | 834622 | B3 for module replqcement | Module replacemeny | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert moduke | 43564 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721202237615764552/5721202237615764552_signature_28.png | Angelina Cuaton | https://www.jotform.com/uploads/macarewecare/222927978372067/5721202237615764552/5721202237615764552_signature_31.png
| 2023-10-03 | 2023-09-07 07:48 | Bataan Genaral Hospital and Medical Center | Balanga Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26116 | PBSP | honeyeichel@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 84 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2481 ASSAY: Xpert HCV VL, COUNT: 80 ASSAY: Xpert HIV-1 Viral Load, COUNT: 224 |  | GXIV-D | 805124 | B3 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721199767619180247/5721199767619180247_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5721199767619180247/5721199767619180247_signature_31.png
| 2023-10-03 | 2023-10-03 12:46 | Mahayag rural health unit | Poblacion mahayag, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26115 | PBSP | cavanricky@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1090 | 2023-10-03 | GX 4 | 834436 | Module 3 and 4 error 4012 and 1002 | Module 2 passed while module 1, 3 and 4 failed | All modules undergone Xpertcheck calibration and Preventive Maintenence | Module 2 available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721179385312438733/5721179385312438733_signature_28.png | Ricky Cavan | https://www.jotform.com/uploads/macarewecare/222927978372067/5721179385312438733/5721179385312438733_signature_31.png
| 2023-10-03 | 2023-10-03 12:41 | Alang-alang Rural Health Unit | Alang alang leyte | 8 | Leyte | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  |  |  |  | 26114 | WHO | vinagatela@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 71 ASSAY: Xpert MTB/RIF Ultra, COUNT: 308 | 2023-10-03 | GXIV | 847249 | #NAME? | #NAME? | - Performed Preventive maintenance - Troubleshoot UPS (battery defective) - Checked voltage output (220) - Generate reports | Unit is Operational | UPS for replacement (Battery defective) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721177097218844313/5721177097218844313_signature_28.png | Vina Milagros G. Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5721177097218844313/5721177097218844313_signature_31.png
| 2023-10-03 | 2023-10-03 11:51 | Bagong Pag-Asa Health Cente4 | 35 Rd., Bagong Pag-Asa, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26113 | DOH | deandomingo12@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 162 |  | GXIV 6C | 831597 | Module A1 temp problem | Performed PM Re-XpertCheck Module A1 Module A1 failed re-xpertcheck Module A1 for reporting Generated report | Re-XpertCheck A1 | Unit is operational. 3 Modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721147606716149625/5721147606716149625_signature_28.png | Dean Antonio B. DOMINGO | https://www.jotform.com/uploads/macarewecare/222927978372067/5721147606716149625/5721147606716149625_signature_31.png
| 2023-10-03 | 2023-10-03 09:49 | Kawit Rural Health Unit | Tabon 2, Kawit, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26111 | DOH | eleanor.perona@yahoo.com | 26 |  |  | GXIV-D | 820705 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed Self-test Performed XpertCheck Generate Reports | 3 modules are Operational, 1 for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141852312806893/5721141852312806893_signature_28.png | Eleanor C.Perona | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141852312806893/5721141852312806893_signature_31.png
| 2023-10-03 | 2023-10-03 11:40 | Molino III City Health Unit | Molino Road, Molino, Bacoor, Cavite | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26112 | DOH | joanrnd@gmail.com | 24 |  | 2023-09-19 | GXIV-D | 821221 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed Self-Test Performed XpertCheck Generate Reports | 3 Modules are Operational |  | XpertCheck Kit | DR 10671 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141852317485657/5721141852317485657_signature_28.png | Jonalyn Aranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141852317485657/5721141852317485657_signature_31.png
| 2023-10-03 | 2023-09-28 12:50 | Trece Martires City Health Office | Governor's Drive, Trece Martires | 4A | Cavite | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26107 | PBSP | lorocavivian@gmail.com | 21 |  |  | GXIV-D | 838160 | Modules 2 and 3 with Temperature Errors | Module Replacement | Replaced Modules 2 and 3 Generate Reports Performed Self-Test | 3 Modules are Operational. 1 for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141802311366291/5721141802311366291_signature_28.png | Vivian D. Lorica | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141802311366291/5721141802311366291_signature_31.png
| 2023-10-03 | 2023-09-28 12:03 | Las Piñas General Hospital and Satellite Trauma Center | Alabang-Zapote Rd, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26109 | DOH | lpghstc.microbiology@gmail.com | 26 |  | 2023-09-19 | GXIV-D | 831847 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports Performed Self-Test | All Modules are Operational | All Modules Passed | XpertCheck Kit | DR 10676 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141802313785199/5721141802313785199_signature_28.png | Julie Jane Andal | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141802313785199/5721141802313785199_signature_31.png
| 2023-10-03 | 2023-09-26 12:01 | UST Hospital | Dapitan St., Manila | NCR | NCR 1st - Manila | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  |  |  |  | 26108 | MACARE-HBDC | usth_bloodbank@yahoo.com | 24 |  |  | MCS+ | 02F069 | Unit is due for Maintenance | Preventive Maintenance | Performed Simulation Test Cleaned Pumps and Valves Checked and Cleaned Sensors Checked Parameters | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141802316564761/5721141802316564761_signature_28.png | Mylene Balboa | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141802316564761/5721141802316564761_signature_31.png
| 2023-10-03 | 2023-09-28 13:05 | Las Piñas General Hospital and Satellite Trauma Center | Alabang-Zapote Rd, Las Piñas | NCR | NCR 4th - Las Piñas | Molecular | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26110 | PBSP | lpghstc.molecular@gmail.com | 24 |  |  | GXIV-D | 834628 | Module 2 with Temperature Error Unit is due for XpertCheck | Module Replacement XpertCheck | Replaced Module 2 Performed Maintenance Performed XpertCheck Performed Self-test Generate Reports | 3 Modules for Reporting(Failed in XpertCheck) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141802316970851/5721141802316970851_signature_28.png | Jennalyn Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141802316970851/5721141802316970851_signature_31.png
| 2023-10-03 | 2023-09-27 10:56 | Tanza Rural Health Unit | Municipal Compound, Tanza | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26106 | PBSP | gutierrezrhian@gmail.com | 25 |  |  | GXIV-D | 812835 | Module 4 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 4 Generate Reports | 3 Modules are Operational. Module 1 for Reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141792319365360/5721141792319365360_signature_28.png | Mary Ann C. Gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5721141792319365360/5721141792319365360_signature_31.png
| 2023-10-03 | 2023-10-03 11:21 | Dalaguete Rural Health Unit | Poblacion, Dalaguete Cebu | 7 | Cebu | Tb laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 26105 | PBSP | ajreef5517@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1490 |  | Gx4 | 815464 | Module 4 SN 210073850 encountered an Error of 2032 started July 12 up to August 22, 2023 until disabled by Medtech up to present | Reset module and conduct preventive maintenance. | Clean Up ultrasonic horn reset module, Clean up plunger, and Icore | All modules are available except Module 4 during running Error 20232 will occured |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721130344795887453/5721130344795887453_signature_28.png | Allaiza Jane Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5721130344795887453/5721130344795887453_signature_31.png
| 2023-10-03 | 2023-10-03 10:39 | Eversly Childs Sanitarium | Jagobiao mandaue City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 26104 | PBSP | alimurungjane@gmail.com | 28 |  |  | Gx 4 | 805270 | Defective module 2 sn 620606 | Clean fan filter, New module 2 installed sn 210100224 | New module 2 installed sn 210100224 | Systems operational, module 4 defective |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721130799314536072/5721130799314536072_signature_28.png | Juanita V. Alimurong | https://www.jotform.com/uploads/macarewecare/222927978372067/5721130799314536072/5721130799314536072_signature_31.png
| 2023-10-03 | 2023-10-03 10:16 | GUIMBAL RURAL HEALTH UNIT | GUIMBAL, ILOILO | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 26102 | PBSP | guimbalrhulaboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1853 ASSAY: Xpert HIV-1 Viral Load, COUNT: 371 |  | GX-IV | 807587 | MODULE B3 REPLACEMENT | #NAME? | PERFORM MODULE B3 REPLACEMENT: DEF SN: 210068670 NEW MODULE SN: 210100427 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721090059606572171/5721090059606572171_signature_28.png | Jackie lou encanto | https://www.jotform.com/uploads/macarewecare/222927978372067/5721090059606572171/5721090059606572171_signature_31.png
| 2023-10-03 | 2023-10-03 10:16 | Banlat Health Center | Daisy St., Banlat, Tandang Sora , Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26103 | PBSP | reyzelmariesantos@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1629 ASSAY: Xpert HPV, COUNT: 95 | 2023-10-03 | GXIV 6C | 804688 | Module A2 for replacement | Performed PM Replaced Module A2 Generated report | Replaced module A2 654654->210102176 | Unit is operarional All modules working |  | GX module 6C | G6326 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721090065183299973/5721090065183299973_signature_28.png | Reyzel marie s. Dulay | https://www.jotform.com/uploads/macarewecare/222927978372067/5721090065183299973/5721090065183299973_signature_31.png
| 2023-10-03 | 2023-10-03 10:05 | Abucay Rural Health Unit | Abucay Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26101 | DOH | parreraledelina22176@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 474 | 2023-08-29 | GXIV-D | 831878 | Bqrcode scanner for replacement | Bqrcode scanner replacement | Performed preventive maintenance Performed barcode scanner Gathered reports Performed functionality test | Unit is operational |  | Barc9de scanner | 10700 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721080727619796247/5721080727619796247_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5721080727619796247/5721080727619796247_signature_31.png
| 2023-10-03 | 2023-10-03 09:16 | Bataan General Hospital and Medical Center | Balanga City Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26100 | MACARE-HBDC | honeyeichel@gmail.com | 22 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 11 ASSAY: Xpert SARS-CoV-2, COUNT: 45 ASSAY: Xpert HBV VL, COUNT: 27 ASSAY: Xpert HCV VL, COUNT: 7 | 2023-09-28 | GXIV-D 10C | 110017894 | B1 for module replacement | Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 Genexpert module | 10698 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5721060707619547969/5721060707619547969_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5721060707619547969/5721060707619547969_signature_31.png
| 2023-10-02 | 2023-10-02 14:11 | Talisay City Health Office | Talisay City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 26099 | WHO | hadmeoffguard@gmail.con | 23 |  |  | GX-4 | 847222 | Printer USB cable short for current printer location. | #NAME? | #NAME? | 4 modules operational, printer operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5720558288111768604/5720558288111768604_signature_28.png | Hadney Besa | https://www.jotform.com/uploads/macarewecare/222927978372067/5720558288111768604/5720558288111768604_signature_31.png
| 2023-10-02 | 2023-10-02 15:48 | Quirino Memorial Medical Center | J.P. Rizal St., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Gian Paolo Magnaye |  |  |  |  | 26098 | MACARE-HBDC | macnellmalapit@gmail.com | 25 |  |  | MCS+ | 19G059SPG | Unit due for calibration and maintenance | Performed Maintenance and Calibration Checked programs parameters Checked rollers ok Checked and test centrifuge Checked and adjust dpm and spm pressure to 300mmhg Checked and adjust line sensor Checked and test pinch valve Checked and test display and buttons Performed functionality test | Performed calibration and maintenance | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5720424615788144331/5720424615788144331_signature_28.png | macnell malapit | https://www.jotform.com/uploads/macarewecare/222927978372067/5720424615788144331/5720424615788144331_signature_31.png
| 2023-10-02 | 2023-10-02 11:50 | Love Yourself | Gorrodo Ave, Cebu City | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  |  |  |  | 26097 | COMMERCIAL | paulo14reforzado@gmail.com | 23 |  | 2023-09-28 | Gx42 | 814150 | Due for Xpertcheck | Perform xpertcheck | Perform xpertcheck | Module 1,2,3 passed 4 fail. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5720282563915517015/5720282563915517015_signature_28.png | Paulo Reforzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5720282563915517015/5720282563915517015_signature_31.png
| 2023-10-02 | 2023-10-02 09:36 | Talisay CIty Health Unit | Talisay | 7 | Cebu | Tb Lan | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 26096 | PBSP | marydelgella@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1840 |  | Gx4 | 815384 | Error 2014 and 4009 Module 2 | Performed maintenance, valve, Icore and plunger | Performed maintenance and reset module Generate data Sys log, IQR, Gxx Files, Assay Statistics | All modues are available but Module terminate during running |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5720200811465053571/5720200811465053571_signature_28.png | Marydel C. Gella | https://www.jotform.com/uploads/macarewecare/222927978372067/5720200811465053571/5720200811465053571_signature_31.png
| 2023-10-01 | 2023-09-29 13:28 | National Kidney and Transplant Institute | East Avenue,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 26095 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 826 ASSAY: Others, COUNT: 23 ASSAY: Xpert HBV VL, COUNT: 564 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1157 ASSAY: Xpert SARS-CoV-2, COUNT: 14675 ASSAY: Xpert HCV VL, COUNT: 48 ASSAY: Others, COUNT: 9 |  | Gx16 | 110005587 | A1 module unable to detect | Checked unit Checked i.o address ok Checked and replaced cable Swap location A1 to A2 but module still underected Reset gateway board | Troubleshooting | 15 module operational A1 undetected |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5719572274212285812/5719572274212285812_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5719572274212285812/5719572274212285812_signature_31.png
| 2023-10-01 | 2023-09-27 16:11 | National Kidney and Transplant Institute | East avenue ,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 26094 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 821 ASSAY: Others, COUNT: 23 ASSAY: Xpert HBV VL, COUNT: 556 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1134 ASSAY: Xpert SARS-CoV-2, COUNT: 14672 ASSAY: Xpert HCV VL, COUNT: 47 ASSAY: Others, COUNT: 9 | 2023-09-26 | GX16 | 110005587 | For module replacement b1 669719 - 700023 | Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Test run unit Performed module replacement | Module replacement | Operational |  | 900-0491RS | 10689 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5719571274214281297/5719571274214281297_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5719571274214281297/5719571274214281297_signature_31.png
| 2023-09-29 | 2023-09-29 16:04 | Macare In House | Aurora Grand, 1005 Aurora Blvd, Brgy. Duyan-Duyan, Project 3, QC | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26093 | DOH | gpdmagnaye30@gmail.com | 25 |  |  | CPU | 2UA70329DC | Macalelon RHU CPU defective | Checked CPU Upgrade software to 6.5 Run Reliability Test | Repair and tested CPU for replacement | CPU working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717843973616083691/5717843973616083691_signature_28.png | Gian Paolo Magnaye | https://www.jotform.com/uploads/macarewecare/222927978372067/5717843973616083691/5717843973616083691_signature_31.png
| 2023-09-29 | 2023-09-29 14:44 | St. Elizabeth Hospital Inc. | General Santos City, South Cotabato | 12 | South Cotabato | Molecular Lab | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26092 | COMMERCIAL | kccrisostomo@sehi.ph | 20.83 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 157 | 2023-09-29 | Gx-IV 10C | 110003713 | Module 2, 3, and 4- modules H/W failed, Subject for module replacement due to multiple errors | Perform Genexpert machine Preventive maintenance | Replace module 1 SN 210066469 with new module SN:210082723 | 1 operational module |  | Genexpert module 10 color SN:210082723 | DR number: 10502 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717809269651666145/5717809269651666145_signature_28.png | SEHI LABORATORY | https://www.jotform.com/uploads/macarewecare/222927978372067/5717809269651666145/5717809269651666145_signature_31.png
| 2023-09-29 | 2023-09-29 14:54 | Poro Rural Health Unit | Poblacion Poro Camotes | 7 | Cebu | Tb Laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 26091 | DOH | avilacharmaine6@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 374 |  | Gx4 | 838279 | Due for Calibration | Performed calibration and maintenance | Performed calibration, Maintenance, Generate data IQR, Sys log, Assay Stst, Gxx files | Module 2 failed and Module 1 Passed. Unitsnare operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717808537321445885/5717808537321445885_signature_28.png | Charmaine A. Apat, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5717808537321445885/5717808537321445885_signature_31.png
| 2023-09-29 | 2023-09-29 14:54 | St. Elizabeth Hospital Inc . | General Santos City, South Cotabato | 12 | South Cotabato | Molecular lab | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26090 | COMMERCIAL | kccrisostomo@sehi.ph | 26.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 83 |  | Gx-IV | 831313 | Module 3 H/W failed due to multiple errors | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 3 operational modules | Module 3 subject for module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717806609656343772/5717806609656343772_signature_28.png | Norisa Mae A. Perea | https://www.jotform.com/uploads/macarewecare/222927978372067/5717806609656343772/5717806609656343772_signature_31.png
| 2023-09-29 | 2023-09-29 14:24 | West Crame Health Center | Rd. 10 West Crame | NCR | NCR 2nd - San Juan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26089 | PBSP | azel1101@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 221 |  | GX-IV-D-6C | 812007 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. Module B3/682815 failed in XpertCheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717785595118939231/5717785595118939231_signature_28.png | Azazel Alpad | https://www.jotform.com/uploads/macarewecare/222927978372067/5717785595118939231/5717785595118939231_signature_31.png
| 2023-09-29 | 2023-09-29 13:34 | VICENTE SOTTO MEMORIAL MEDICAL CENTER | CEBU CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  |  |  |  | 26088 | DOH | myraestose@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 80 ASSAY: Xpert HPV, COUNT: 233 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2894 ASSAY: Xpert HIV-1 Viral Load, COUNT: 81 |  | GXIV | 820206 | FOR MODULE A1 CALIBRATION | COLLECT IQR, SYSYEM LOG, AND ARCHIVE | PERFORM MODULE A1 CALIBRATION, PASSED | EQUIPMENT IS OPERATIONAL | MODULE B2 PENDING |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717762743821901033/5717762743821901033_signature_28.png | Myra Riccil Estose | https://www.jotform.com/uploads/macarewecare/222927978372067/5717762743821901033/5717762743821901033_signature_31.png
| 2023-09-29 | 2023-09-29 13:28 | VICENTE SOTTO MEMORIAL MEDICAL CENTER | CEBU CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  |  |  |  | 26087 | DOH | myraestose@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 96 ASSAY: Xpert HPV, COUNT: 232 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2767 ASSAY: Xpert HIV-1 Viral Load, COUNT: 100 |  | GX-IV | 838825 | FOR MODULE A4 CALIBRATION | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | PERFORM MODULE A4 CALIBRATION-PASSED | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717761333828593752/5717761333828593752_signature_28.png | Myra Riccil Estose | https://www.jotform.com/uploads/macarewecare/222927978372067/5717761333828593752/5717761333828593752_signature_31.png
| 2023-09-29 | 2023-09-29 13:09 | San Luis Rural Health Unit | San Luis, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 26086 | DOH | mohan_santos@yahoo.com | 24 |  | 2023-09-29 | Genexpert IV | 831714 | Error 2008 | Gathered Reports | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module for Monitoring | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717752692214032840/5717752692214032840_signature_28.png | Mohan S. Pangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5717752692214032840/5717752692214032840_signature_31.png
| 2023-09-29 | 2023-09-29 12:56 | San Jose Del Monte City Health Center VII | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 26085 | PBSP | krizetteamorio@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3127 ASSAY: Xpert HIV-1 Viral Load, COUNT: 144 |  | GXIV | 807620 | CPU problem | #NAME? | Performed troubleshooting | Unit is operational. CPU under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717731357611273406/5717731357611273406_signature_28.png | Krizette Marie Amorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5717731357611273406/5717731357611273406_signature_31.png
| 2023-09-29 | 2023-09-29 11:02 | Bagong Silang Ph. 1 Health Center | Bagong Silang, Caloocan | NCR | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 26084 | PBSP | alcaydeirene07@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2041 | 2023-09-29 | GXIV | 838856 | Module B2 for replacement | #NAME? | Replaced module B2 | Unit is operational. |  | GeneXpert Module SN: 210101970 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717691847618397602/5717691847618397602_signature_28.png | Irene Alcayde | https://www.jotform.com/uploads/macarewecare/222927978372067/5717691847618397602/5717691847618397602_signature_31.png
| 2023-09-29 | 2023-09-29 11:06 | Batis Health Center | F. Manalo St. | NCR | NCR 2nd - San Juan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26083 | DOH | almacenaio16@gmail.com | 24 |  | 2023-09-05 | GX-IV-D-6C | 831598 | For monitor replacement. Module A4 has error 1001/1002/2014 | Replaced monitor Run XpertCheck on module A4 | Performed maintenance Replaced defective monitor Run XpertCheck on module A4 Generated reports | Unit is operational. Module A4/879974 passed in XpertCheck, under observation. |  | Old monitor: CNK832243Z New monitor: H22V5SG81A230601754 | DR# 10642 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717665555116889073/5717665555116889073_signature_28.png | JENICA ROSE A. ORTIZO | https://www.jotform.com/uploads/macarewecare/222927978372067/5717665555116889073/5717665555116889073_signature_31.png
| 2023-09-29 | 2023-09-28 07:35 | Bayambang District Hospital | Bayambang, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26082 | PBSP | bevzneri@gmail.com | 24 |  | 2023-09-28 | GX4 | 812830 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 861757 >>>210100433 | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210100433 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717649620329788643/5717649620329788643_signature_28.png | Beverly A. Neri | https://www.jotform.com/uploads/macarewecare/222927978372067/5717649620329788643/5717649620329788643_signature_31.png
| 2023-09-29 | 2023-09-29 09:45 | Medilinx | Amorsolo St.,Makati City | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 26081 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 113 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 526 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 555 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2057 ASSAY: Others, COUNT: 2 ASSAY: Others, COUNT: 9 ASSAY: Xpert HCV VL, COUNT: 49 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1455 |  | GX4 | 110001647 | H/Wfailed Module 2 | Checked unit Checked filter ok Checked room temp ok Gathered reports Disabled module 2 | Troubleshooting | Module 2 disabled | For reporting module 2 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5717620494553605254/5717620494553605254_signature_28.png | Aiza Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5717620494553605254/5717620494553605254_signature_31.png
| 2023-09-28 | 2023-09-28 15:18 | Dra. Elvira Lagrosa Health Center (Duay) | Dolores Hidalgo St., Pasay | NCR | NCR 4th - Pasay | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26080 | PBSP | rhemejetbotaquia@gmail.com | 32 |  |  | GXIV 6C | 812833 | Module 2 frequent invalid result GX software slow | Performed PM Upgraded software version: 6.1 to 6.4 Troubleshoot module 2 Run reliability test twice: ok Performed reagent run: ok Tried running 2 samples: ok Generated report | Upgraded software version Troubleshoot module 2 | Unit is operational Module 2 for monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716981265336259104/5716981265336259104_signature_28.png | Rheme Tañeza | https://www.jotform.com/uploads/macarewecare/222927978372067/5716981265336259104/5716981265336259104_signature_31.png
| 2023-09-28 | 2023-09-28 15:18 | Dra. Elvira Lagrosa Health Center (Duay) | Dolores Hidalgo St., Pasay | NCR | NCR 4th - Pasay | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26079 | PBSP | rhemejetbotaquia@gmail.com | 32 |  |  | GXIV 6C | 812833 | Module 2 frequent invalid result GX software slow | Performed PM Upgraded software version: 6.1 to 6.4 Troubleshoot module 2 Run reliability test twice: ok Performed reagent run: ok Tried running 2 samples: ok Generated report | Upgraded software version Troubleshoot module 2 | Unit is operational Module 2 for monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716980195339086343/5716980195339086343_signature_28.png | Rheme Tañeza | https://www.jotform.com/uploads/macarewecare/222927978372067/5716980195339086343/5716980195339086343_signature_31.png
| 2023-09-28 | 2023-09-28 15:03 | Nueva Valencia Rural Health Unit | Nueva Valencia, Guimaras | 6 | Guimaras | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 26078 | DOH | nuevarhu@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 74 | 2023-07-27 | GXIV-4D | 838291 | Due for XPC | #NAME? | #NAME? | All modules are functional. Module B3 failed XPC. | Module B3 for replacement. | XpertCheck-CE-5 | 10312 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716960426618393010/5716960426618393010_signature_28.png | Rhea April Jaleco | https://www.jotform.com/uploads/macarewecare/222927978372067/5716960426618393010/5716960426618393010_signature_31.png
| 2023-09-28 | 2023-09-28 14:50 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  |  |  |  | 26077 | PBSP | elma.bragat@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4537 ASSAY: Xpert HIV-1 Viral Load, COUNT: 189 |  | GX-IV | 815275 | -MODULE D2 STUCK CARTRIDGE -MODULE D2 ERROR 2016 AND 4004 | #NAME? | #NAME? | MODULE D2 DISABLED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716932063385476670/5716932063385476670_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5716932063385476670/5716932063385476670_signature_31.png
| 2023-09-28 | 2023-09-28 14:27 | Claver Rural Health Unit | Poblacion, Claver | 13 | Surigao del Norte | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 26076 | DOH | rhodaruthdalugdog@yahoo.com.ph | 30 | ASSAY: Xpert MTB/RIF, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1336 | 2023-08-01 | GX-4 | 838514 | Module B3 found service required thru XPC | Reported to Cepheid | XpertCheck Calibration Preventive Maintenance | 3 modules working, 1 module for replacement |  | Xpertcheck, CE-IVD 5 Cartridges/Kit | 10386 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716928436681072831/5716928436681072831_signature_28.png | Joyce E. Andoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5716928436681072831/5716928436681072831_signature_31.png
| 2023-09-28 | 2023-09-28 14:32 | North Daang Hari Super Health Center | North Daang Hari | NCR | NCR 4th - Taguig | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26075 | PBSP | onlychumz@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2950 |  | GX-IV-D-6C | 804702 | Module A1 has error 1002 module A2 has error 2120/2126 | Checked modules A1 and A2 | Performed maintenance Checked defective modules Checked connections of module A2 Generated reports | Module A1/661743 and A2/713660 not operational. Modules A4 and A5 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716921681814594523/5716921681814594523_signature_28.png | Charmaine D. Trinidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5716921681814594523/5716921681814594523_signature_31.png
| 2023-09-28 | 2023-09-28 14:23 | Candaba Rural Health Unit | Candaba, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 26074 | DOH | rochelannlontoc@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 461 | 2023-09-28 | Genexpert IV | 831873 | UPS not working | Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716916018503406326/5716916018503406326_signature_28.png | Rochelle Ann Lontoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5716916018503406326/5716916018503406326_signature_31.png
| 2023-09-28 | 2023-09-28 13:10 | Tungawan rural health unit | Tungawan, Zamboanga Sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26073 | PBSP | punkypink29@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 486 | 2023-09-28 | GX 2 | 834830 | None | Module 1 and 2 undergone Xpertcheck calibration and Preventive Maintenence. All modules passed | Module 1 and 2 undergone Xpertcheck calibration and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716904799988166911/5716904799988166911_signature_28.png | Love Faith Montilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5716904799988166911/5716904799988166911_signature_31.png
| 2023-09-28 | 2023-09-28 13:21 | Buenavista Rural Health Unit | Buenavista, Guimaras | 6 | Guimaras | TB Dots | YES | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 26072 | PBSP | clys0910@gmail.com | 21 |  | 2023-09-28 | GXIV-4D | 838861 | Module B3 defective | -Replace defective module B3/914644 with new module SN: 210102196 | #NAME? | All modules are functional. |  | GeneXpert Module 6-colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716888836775626087/5716888836775626087_signature_28.png | Raymund Salapio | https://www.jotform.com/uploads/macarewecare/222927978372067/5716888836775626087/5716888836775626087_signature_31.png
| 2023-09-28 | 2023-09-28 12:24 | Kamuning Health Center | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 26071 | DOH | vicvic.chua25@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 674 ASSAY: Xpert HPV, COUNT: 482 |  | GXIV | 831612 | Stuck cartridge on A2 | #NAME? | Removed stuck cartridge | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716844244212946587/5716844244212946587_signature_28.png | Jesusa Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5716844244212946587/5716844244212946587_signature_31.png
| 2023-09-28 | 2023-09-28 09:27 | Tayasan Rural Health Unit | National Highway, Tayasan, Negros Oriental | 7 | Negros Oriental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 26070 | PBSP | porshcabz151@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 545 |  | GX-4 6c | 835097 | - Defective module 787973/B1. | #NAME? | - Replaced module 787973/B1 with new module with SN 210101910 - Generate Pre- and Post- IQR - Pullout defective module 787973/B1 for return to Cepheid | All 2 modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716740005963860208/5716740005963860208_signature_28.png | Portia Cabanting | https://www.jotform.com/uploads/macarewecare/222927978372067/5716740005963860208/5716740005963860208_signature_31.png
| 2023-09-28 | 2023-09-28 08:42 | Schistosomiasis Research and Training Center | Palo Leyte | 8 | Leyte | Tb-dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26069 | PBSP | mcesplago.pbsp@gmai.com | 20 |  | 2023-09-28 | GXIV | 707521 | -Module B2/210060753 defective | #NAME? | - Module Replacement - Install new module in B2 w/ SN 210102177 - Performed Functionality test - Performed Preventive maintenance - Generate reports - Pullout defective module SN/210060753 | Unit is Operational |  | Genexpert Module 6 Color SN/210102177 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716714503814585884/5716714503814585884_signature_28.png | Arabelle Joy Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5716714503814585884/5716714503814585884_signature_31.png
| 2023-09-28 | 2023-09-28 08:42 | Schistosomiasis Research and Training Center | Palo Leyte | 8 | Leyte | Tb-dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26068 | PBSP | mbesplago.pbsp@gmail.com | 20 |  | 2023-09-28 | GXIV | 707521 | -Module B2/210060753 defective | #NAME? | - Module Replacement - Install new module in B2 w/ SN 210102177 - Performed Functionality test - Performed Preventive maintenance - Generate reports - Pullout defective module SN/210060753 | Unit is Operational |  | Genexpert Module 6 Color SN/210102177 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716714123816132510/5716714123816132510_signature_28.png | Arabelle Joy Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5716714123816132510/5716714123816132510_signature_31.png
| 2023-09-27 | 2023-09-27 21:48 | Philippine Tuberculosis Society, Inc. | E. Rodriguez Sr. Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26067 | PBSP | sannymisorena@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1212 | 2023-09-26 | GX IV 6C | 707527 | Module A1 for replacement | Perfomed preventive maintenance Replaced module A1 Generated report | Replaced Module A1 882275 -> 210102193 | Unit is operational All modules working |  | GX Module6C | G6652 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716324567512632670/5716324567512632670_signature_28.png | Peter Samonte | https://www.jotform.com/uploads/macarewecare/222927978372067/5716324567512632670/5716324567512632670_signature_31.png
| 2023-09-27 | 2023-09-27 16:24 | Taguig City Health Office | T&D Village, cor Col P Cruz, St, Taguig, 1637 Kalakhang Maynila | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26066 | PBSP | dumawalgina20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3013 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-IV-D-6C | 814372 | For maintenance | Performed maintenance | Performed maintenance Relocated the GeneXpert to other room Functionality test Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716125899502035979/5716125899502035979_signature_28.png | Gina B. Dumawal | https://www.jotform.com/uploads/macarewecare/222927978372067/5716125899502035979/5716125899502035979_signature_31.png
| 2023-09-27 | 2023-09-26 08:33 | Region 1 Medical Center | Dagupan City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26065 | PBSP | r1mcxpertlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1940 ASSAY: Xpert SARS-CoV-2, COUNT: 63 ASSAY: Xpert HIV-1 Viral Load, COUNT: 249 |  | GX4 | 804726 | For Maintenance | #NAME? | Performed maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716085030327398617/5716085030327398617_signature_28.png | Kristin Deinielle S. Banaag | https://www.jotform.com/uploads/macarewecare/222927978372067/5716085030327398617/5716085030327398617_signature_31.png
| 2023-09-27 | 2023-09-26 08:40 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 26064 | PBSP | r1mcxpertlab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1505 ASSAY: Xpert SARS-CoV-2, COUNT: 1641 |  | GX 16 | 842230 | For Maintenance | #NAME? | For Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716082820328512335/5716082820328512335_signature_28.png | Kirstin Deinielle S. Banaag | https://www.jotform.com/uploads/macarewecare/222927978372067/5716082820328512335/5716082820328512335_signature_31.png
| 2023-09-27 | 2023-09-27 13:24 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Yasser Salih |  |  |  |  | 26063 | DOH | gbmmbg@yahoo.com | 29 |  |  | GX4 | 838437 | Module B1 recurring Error 2008 | #NAME? | Performed Troubleshooting and Test Instructions | Unit is operational | Module B1 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716077280328175333/5716077280328175333_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5716077280328175333/5716077280328175333_signature_31.png
| 2023-09-27 | 2023-09-27 13:47 | Bayugan City Health Office | Poblacion, Bayugan City | 13 | Agusan del Sur | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 26062 | PBSP | lourdescayao@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2383 |  | GX-4 | 815391 | Module B3 cant be run through XPC | Reported to Cepheid | XpertCheck Calibration Preventive Maintenance | 3 modules working, 1 module for observation (B1), 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716057496225825305/5716057496225825305_signature_28.png | Lourdes Diola | https://www.jotform.com/uploads/macarewecare/222927978372067/5716057496225825305/5716057496225825305_signature_31.png
| 2023-09-27 | 2023-09-27 13:40 | Monkayo Rural Health Unit | Poblacion, Monkayo | 11 | Davao de Oro | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26061 | DOH | kendralois730@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1454 | 2023-08-03 | GX-IV 4D | 838487 | Due for Calibration and Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration | Perform Calibration and Preventive Maintenance | Module B1 for Service Required Modules B2, B3 and B4 Passed ghe Calibration | High Room Temperature, Enduser is advice to have better cooling to achieve the optimal room temperature for GeneXpert. | Xpertcheck CE - 5 | DR 10442 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716055219028042877/5716055219028042877_signature_28.png | Karen S. Anuta | https://www.jotform.com/uploads/macarewecare/222927978372067/5716055219028042877/5716055219028042877_signature_31.png
| 2023-09-27 | 2023-09-27 12:27 | Mandaluyong City Health Office | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 26060 | PBSP | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2509 | 2023-09-27 | GXIV | 805062 | Modules A3, A4 for replacement | #NAME? | Replaced A3, A4 | Unit is operational. |  | GeneXpert Module SN: 210101792, 210102029 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716035001627123407/5716035001627123407_signature_28.png | Teresita Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5716035001627123407/5716035001627123407_signature_31.png
| 2023-09-27 | 2023-09-27 13:01 | Estevez Memorial Hospital, Inc. | Don Juan St., Legazpi City | 5 | Albay | Laboratory | YES | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26059 | COMMERCIAL | armenaemmanuel@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 10 ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 118 ASSAY: Xpert MTB/RIF Ultra, COUNT: 118 ASSAY: Others, COUNT: 14 |  | GX4-10C | 110017942 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore, fan filter Generate Reports | Performed Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716026337813794060/5716026337813794060_signature_28.png | Emmanuel M Armena | https://www.jotform.com/uploads/macarewecare/222927978372067/5716026337813794060/5716026337813794060_signature_31.png
| 2023-09-27 | 2023-09-27 13:14 | MR Flores Health Center | P. Herrera St. | NCR | NCR 4th - Pateros | Laboratory | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26058 | PBSP | paterosclinicallaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 790 |  | GX-IV-D-6C | 838881 | Module B2 and Be has error 1001/1002 | Checked module B2 and B3 | Performed maintenance Checked module B2 and B3 Generated reports | Unit is operational. Module B2/861632 and B3/861628 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716015409503921356/5716015409503921356_signature_28.png | Jose Mari Zoilo B. De Borja, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5716015409503921356/5716015409503921356_signature_31.png
| 2023-09-27 | 2023-09-27 13:17 | MR Flores Health Center | P. Herrera St. | NCR | NCR 4th - Pateros | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26057 | PBSP | paterosclinicallaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 711 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GX-IV-D-6C (2 MODULES) | 834803 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716014829508989977/5716014829508989977_signature_28.png | Jose Mari Zoilo B. De Borja, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5716014829508989977/5716014829508989977_signature_31.png
| 2023-09-27 | 2023-09-27 12:43 | Masinloc Rural Health Unit | Masinloc Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26056 | DOH | rodenne@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2259 |  | GXIV-D | 831608 | A2 for re-xpertcheck temperature error | Xpertcheck | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Performed xpertcheck Generate reports Archived test results Performed functionality test | Unit us operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5716000441814837665/5716000441814837665_signature_28.png | Rodenn Eduvala | https://www.jotform.com/uploads/macarewecare/222927978372067/5716000441814837665/5716000441814837665_signature_31.png
| 2023-09-27 | 2023-09-27 12:22 | Sta Rosa Community District Hospital | Sta Rosa Olango Island Lapulapu City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 26055 | DOH | ariestumz@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 186 |  | Gx4 | 838534 | Due for Maintenance, activate code, Clean Icore, Valve, plunger | Performed maintenance, activate code | Performed maintenance, generate data, IQR and sys log | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715980341218278235/5715980341218278235_signature_28.png | Ariston Tumulak | https://www.jotform.com/uploads/macarewecare/222927978372067/5715980341218278235/5715980341218278235_signature_31.png
| 2023-09-27 | 2023-09-27 12:16 | President Diosdado Macapagal District Hospital | Tobias Fornier, Antique | 6 | Antique | Microbiology | YES | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 26054 | PBSP | nerielserapion@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 281 |  | GXIV-2D | 835096 | None | None | -Preventive Maintenance -Archive tests since May 2023 -Install XpertCheck version 1.5a | All two modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715977450617292920/5715977450617292920_signature_28.png | Neriel Serapion | https://www.jotform.com/uploads/macarewecare/222927978372067/5715977450617292920/5715977450617292920_signature_31.png
| 2023-09-27 | 2023-09-27 11:47 | Calubian Rural Health Unit | Calubian, Leyte | 8 | Leyte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26053 | DOH | mjcabardo06@gmail.com | 27 |  | 2023-09-27 | GXIV | 838399 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | Module B1 and B4 for service required | Xpertcheck Calibration | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715966466716988094/5715966466716988094_signature_28.png | Mighty Joy Cabardo | https://www.jotform.com/uploads/macarewecare/222927978372067/5715966466716988094/5715966466716988094_signature_31.png
| 2023-09-27 | 2023-09-27 11:39 | Detoxicare | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  |  |  |  | 26052 | CONSORTIUM | imeehabaluyas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 |  | GXIV | 110018312 | Unit for installation | #NAME? | Installed unit | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715958641625990333/5715958641625990333_signature_28.png | Imelda Habaluyas | https://www.jotform.com/uploads/macarewecare/222927978372067/5715958641625990333/5715958641625990333_signature_31.png
| 2023-09-27 | 2023-09-27 11:38 | San Remegio RHU | San Remegio, Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 26051 | WHO | rodacanetecamay@gmail.com | 25 |  |  | Gx 4 | 847333 | Defective module b3 sn 910902 | Install new module sn 210102232, clean fan filter | Install new module sn 210102232 | System operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715954165469646391/5715954165469646391_signature_28.png | Roda Camay | https://www.jotform.com/uploads/macarewecare/222927978372067/5715954165469646391/5715954165469646391_signature_31.png
| 2023-09-27 | 2023-09-27 10:33 | Bangkal Health Center | Rodriguez St. | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26050 | DOH | mhdbangkalsatlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1559 | 2023-09-04 | GX-IV-D-6C | 831875 | For monitor replacement | Replaced defective monitor | Performed maintenance Replaced defective monitor Generated reports | All modules operational. |  | Old monitor: CNK83224MZ New monitor: H22V5SG81A230601873 | DR# 10635 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715928407995158543/5715928407995158543_signature_28.png | Ghemyer Hope M. Jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/5715928407995158543/5715928407995158543_signature_31.png
| 2023-09-27 | 2023-09-27 09:49 | Quezon City General Hospital | Seminary Rd., Bahay Torro, EDSA, Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE Gathered Data | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26049 | MACARE-HBDC | treatmenthub.qcgh@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GXIV 6C | 110012766 | OS version and Hard drive details needed by city hall | Performed PM Gathered data Generated report | Gathered needed data | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715892865121496359/5715892865121496359_signature_28.png | Elena G. Ventura, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5715892865121496359/5715892865121496359_signature_31.png
| 2023-09-27 | 2023-09-27 09:33 | Navotas Medical and Wellness Center | M. Naval St. Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 26048 | PBSP | alejandrinoknowell@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 62 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1728 | 2023-09-26 | GXIV | 807592 | Unit for PM. For module replacement of module B2. | Performed PM. Performed module replacement. Generated reports | PM and module replacement. | Unit is operational. |  | GX module B2 (210102146) in replacement to defective B2 (634856). | G6305 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715887686814827134/5715887686814827134_signature_28.png | Noel Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5715887686814827134/5715887686814827134_signature_31.png
| 2023-09-27 | 2023-09-27 08:28 | Olongapo City Health Office | Olongapo City, Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com | olongapochotmlandrtdl@gmail.com |  | 26047 | PBSP | olongapochotmlandrtdl@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 346 ASSAY: Xpert MTB/RIF Ultra, COUNT: 964 ASSAY: Xpert HIV-1 Viral Load, COUNT: 167 |  | GXIV-2 | 834807 | C2 for module replacement | Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715842811818316332/5715842811818316332_signature_28.png | Roselyn M. Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5715842811818316332/5715842811818316332_signature_31.png
| 2023-09-27 | 2023-09-07 07:48 | Bataan Genaral Hospital and Medical Center | Balanga Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26046 | PBSP | honeyeichel@gmail.com | 29 |  |  | GXIV-D | 805124 | B3 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715834721816917071/5715834721816917071_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5715834721816917071/5715834721816917071_signature_31.png
| 2023-09-27 | 2023-09-26 14:02 | Eastern Visayas Regional Medical | Bagacay, Tacloban City | 8 | Leyte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26045 | PBSP | shyzamoras@gmail.com | 23 |  | 2023-09-26 | GXIV | 808631 | Module B4 keep booting/reset | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715809571316719760/5715809571316719760_signature_28.png | SHIRLYN C.ZAMORAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5715809571316719760/5715809571316719760_signature_31.png
| 2023-09-26 | 2023-09-26 14:58 | East Rembo Health Center | 23rd Ave., Taguig, Metro Manila | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26044 | PBSP | rtmendros@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 420 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1666 |  | GXIV 6C | 807591 | GX machine for packing | Performed PM Archived and purge all reports Saved all files to USB Packed GX and accessories Generated report | Packed GX machine and all its accessories | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715436167518897523/5715436167518897523_signature_28.png | Roxan T. Mendros | https://www.jotform.com/uploads/macarewecare/222927978372067/5715436167518897523/5715436167518897523_signature_31.png
| 2023-09-26 | 2023-09-26 15:30 | Jose B. Lingad Memorial Regional Hospital | San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26043 | PBSP | jan.sky1623@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2525 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 21 |  | GXIV-D | 804763 | Due for calibration | Xpertcheck Preventive maintenace | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Performed calibration Generate reports Archived test results Performed functionality test | Unit is opertional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715251044344093242/5715251044344093242_signature_28.png | Janice Condes | https://www.jotform.com/uploads/macarewecare/222927978372067/5715251044344093242/5715251044344093242_signature_31.png
| 2023-09-26 | 2023-09-26 12:28 | West Rembo Health Center | J.P. Rizal Ext., Makati, Metro Manila | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26042 | PBSP | danise_er_19@yahoo.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1179 |  | GX IV 6C | 831712 | Machine for pack up | Performed PM Archived and purged results Saved all files to USB Generated report | Packed up GX and all accessories | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715235457453668008/5715235457453668008_signature_28.png | Danis Maebel Calimag | https://www.jotform.com/uploads/macarewecare/222927978372067/5715235457453668008/5715235457453668008_signature_31.png
| 2023-09-26 | 2023-09-26 14:21 | Hinigaran Rural Health Unit | Municipality of Hinigaran | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 26041 | PBSP | mfeliprada@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 931 |  | GX4 | 815283 | None | Pulled out the old module for B3 with SN 609119. Save IQR file before and after replacement. | Installed the new module for A3 with SN 210102147 | All modules are in good conditions. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715199209915496934/5715199209915496934_signature_28.png | Mary Flor Feliprada | https://www.jotform.com/uploads/macarewecare/222927978372067/5715199209915496934/5715199209915496934_signature_31.png
| 2023-09-26 | 2023-09-26 14:01 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | TB-Dots | YES | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 26040 | PBSP | sanjoseantiquegxsite@gmail.com | 28 |  |  | GXIV-4D | 804714 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715182970391220254/5715182970391220254_signature_28.png | Oscar Delano Salazar IV | https://www.jotform.com/uploads/macarewecare/222927978372067/5715182970391220254/5715182970391220254_signature_31.png
| 2023-09-26 | 2023-09-26 14:05 | San Jose Municipal Health Office | San Jose de Buenavista, Antique | 6 | Antique | TB-Dots | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 26039 | DOH | sanjoseantiquegxsite@gmail.com | 28 |  | 2023-08-23 | GXIV-4D | 838412 | Module B2 and B4 H/W fail | -Replace defective module B2/861117 with new module SN: 892457 -Replace defective module B4/861168 with new module SN: 790369 | #NAME? | All modules are functional. |  | GeneXper Module 6-Colors | 10542, 10543 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715182970391752165/5715182970391752165_signature_28.png | Oscar Delano Salazar, IV | https://www.jotform.com/uploads/macarewecare/222927978372067/5715182970391752165/5715182970391752165_signature_31.png
| 2023-09-26 | 2023-09-26 13:53 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 26038 | DOH | abigailgsula@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV-D | 831912 | A2 for re-xpertcheck | Xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is partially operational | A2 failed xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715174836425923821/5715174836425923821_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5715174836425923821/5715174836425923821_signature_31.png
| 2023-09-26 | 2023-09-26 10:19 | Carcar City Health Office | Valladolid, Carcar Poblacion | 7 | Cebu | Tb laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 26037 | DOH | abellana.jilliannadine@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1629 |  | Gx4 | 838283 | Error 2008 module 1 started on 8-04-2023 | Conduct preventive maintenance, generate data, Orient end user for proper sample processing | Clean up plunger, Icore, Valve, conduct self test Generate IQR, Sys log, Assay Statistics, Gxx files | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715053480511245717/5715053480511245717_signature_28.png | JILLIAN NADINE ABELLANA | https://www.jotform.com/uploads/macarewecare/222927978372067/5715053480511245717/5715053480511245717_signature_31.png
| 2023-09-26 | 2023-09-26 10:10 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 26036 | MACARE-HBDC | fabella.mollabgene@gmail.com | 22 |  | 2023-09-20 | GXIV-D-10-C | 110015482 | Modules 2 and 4 Failed in XpertCheck | Module Replacement | Replaced Modules 2 and 4 Generate Reports Performed Self-Test | All Modules are Operational |  | GX Module | DR 10679 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5715038903533716356/5715038903533716356_signature_28.png | Christianne Bernadette Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5715038903533716356/5715038903533716356_signature_31.png
| 2023-09-25 | 2023-09-25 14:38 | Bogo City Health Office | Poblacion | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 26035 | PBSP | bogocholab@yahoo.com | 23 |  | 2023-09-19 | GX4 | 808630 | UPS installation | Install UPS with new serial number 120423 | Install UPS | All are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5714335613911166628/5714335613911166628_signature_28.png | Rege | https://www.jotform.com/uploads/macarewecare/222927978372067/5714335613911166628/5714335613911166628_signature_31.png
| 2023-09-24 | 2023-09-20 14:53 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26034 | PBSP | lcpmicrolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4441 ASSAY: Xpert SARS-CoV-2, COUNT: 3961 | 2023-09-20 | GX16 | 801080 | For module replacement B3 210070829 - 210102258 | Performed module replacement Cleaned filter,cartridge bay,plunger rod,icore Gathered reports Test run unit | Module replacement and maintenance | Operational |  | 9000491RS | G6300 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5713548662609265535/5713548662609265535_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5713548662609265535/5713548662609265535_signature_31.png
| 2023-09-24 | 2023-09-21 15:46 | National Kidney and Transplant Institute | East Avenue,QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 26033 | MACARE-HBDC | molecularlabnkti@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 804 ASSAY: Others, COUNT: 23 ASSAY: Xpert HBV VL, COUNT: 532 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1078 ASSAY: Xpert SARS-CoV-2, COUNT: 14655 ASSAY: Xpert HCV VL, COUNT: 47 | 2023-09-19 | GX16 | 110005587 | For module replacement A1 210029613 - 620546 B2 700412 - 685498 B4 636241 - 686979 C2 210028938 - 801983 | Performed module replacement, Cleaned filter,plunger,cartridge bay | Module replacement | 15 module operational B1 waiting for module arrival |  | 900 - 0491rs | 10678 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5713519482604397498/5713519482604397498_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5713519482604397498/5713519482604397498_signature_31.png
| 2023-09-22 | 2023-09-22 10:28 | Eastern Pangasinan District Hospital | Tayug, Pangasinan | 1 | Pangasinan | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26032 | PBSP | epdhlaboratory@gmail.com | 26 |  |  | GX4 | 815394 | CPU for further troubleshooting | Performed Troubleshooting | Performed Troubleshooting | Unit is not operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5711971670117082523/5711971670117082523_signature_28.png | Justine Joy Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5711971670117082523/5711971670117082523_signature_31.png
| 2023-09-22 | 2023-09-22 10:28 | Eastern Pangasinan District Hospital | Tayug, Pangasinan | 1 | Pangasinan | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 26031 | PBSP | epdhlaboratory@gmail.com | 26 |  |  | GX4 | 815394 | CPU for further troubleshooting | Performed Troubleshooting | Performed Troubleshooting | Unit is not operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5711971260112957756/5711971260112957756_signature_28.png | Justine Joy Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5711971260112957756/5711971260112957756_signature_31.png
| 2023-09-22 | 2023-09-22 10:58 | Lorenzo D. Zayco District Hospital | Kabankalan City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | CALIBRATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Emerson Bejemino |  |  |  |  | 26030 | PBSP | beuagravante@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1812 |  | GX-4 | 807455 | Due for Xpertcheck. | #NAME? | #NAME? | 4 modules passed Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5711611784612785299/5711611784612785299_signature_28.png | Beulah V.Agravante | https://www.jotform.com/uploads/macarewecare/222927978372067/5711611784612785299/5711611784612785299_signature_31.png
| 2023-09-22 | 2023-09-20 08:22 | Kawayan Rural Health Unit | Kawayan, Biliran | 8 | Biliran | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26029 | PBSP | clytiebernadas@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 87 | 2023-09-20 | GXIV-2D | 835196 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed (B1,B2) | Xpertcheck Cartridge | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5711579137435742017/5711579137435742017_signature_28.png | Clytie Bernadas | https://www.jotform.com/uploads/macarewecare/222927978372067/5711579137435742017/5711579137435742017_signature_31.png
| 2023-09-22 | 2023-09-21 11:26 | Caibiran Rural Health Unit | Caibiran, Biliran | 8 | Biliran | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26028 | PBSP | ednaobaletran07@gmail.com | 31 |  | 2023-09-21 | GXIV-2D | 834815 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | Module B2 for service required | Xpertcheck kit | PBSB |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5711578157435446533/5711578157435446533_signature_28.png | Edna Letran | https://www.jotform.com/uploads/macarewecare/222927978372067/5711578157435446533/5711578157435446533_signature_31.png
| 2023-09-22 | 2023-09-21 11:26 | Caibiran Rural Health Unit | Caibiran, Biliran | 8 | Biliran | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26027 | PBSP | ednaobaletran07@gmail.com | 31 |  | 2023-09-21 | GXIV-2D | 834815 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | Module B2 for service required | Xpertcheck kit | PBSB |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5711577617436445015/5711577617436445015_signature_28.png | Edna Letran | https://www.jotform.com/uploads/macarewecare/222927978372067/5711577617436445015/5711577617436445015_signature_31.png
| 2023-09-22 | 2023-09-22 09:52 | Tacloban City Health Office | Tacloban City | 8 | Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 26026 | DOH | marteetingzon@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 97 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1221 | 2023-09-22 | GXIV | 838421 | Module B1, B4 defective | N/a | Preventive maintenance | Unit is Operational (B1, B4 waiting for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5711574500899026151/5711574500899026151_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5711574500899026151/5711574500899026151_signature_31.png
| 2023-09-21 | 2023-09-21 15:33 | Dr. M. Icasiano Health Center | 1806 Pedro Gil St., Paco, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26025 | DOH | fabroslorrainelucila@gmail.com | 26 | ASSAY: Others, COUNT: 4 | 2023-09-15 | GX IV 6C | 831684 | CPU for replacment | Performed PM Archived reports Installed new CPU Setup configuration Installed printer driver Run sample Generated report | Replaced CPU 2UA8471W40 -> CPH1119F03925 | Unit is operational 4 modules working |  | GeneXpert Desktop | DR 10665 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710912792422885111/5710912792422885111_signature_28.png | Lorraine Lucila L Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5710912792422885111/5710912792422885111_signature_31.png
| 2023-09-21 | 2023-09-21 15:21 | General Tinio Rural Health Unit | General Tinio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 26024 | DOH | jadelyn_pajarillaga@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 827 | 2023-09-21 | Genexpert IV | 831695 | Module 3 error | Gathered Gxx file | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 3 for observation | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710902818128801896/5710902818128801896_signature_28.png | Jadelyn Pajarillaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5710902818128801896/5710902818128801896_signature_31.png
| 2023-09-21 | 2023-09-21 13:54 | Bagong Silang Ph.2 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 26023 | DOH | laguitannataniel@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 987 | 2023-09-21 | GXIV | 831845 | Modules A1, A2, A3 for replacement | #NAME? | Performed re XpertCheck on A1, A2, A3 | Unit is operational. A1: passed A2: passed A3: service required |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710864959151228876/5710864959151228876_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5710864959151228876/5710864959151228876_signature_31.png
| 2023-09-21 | 2023-09-21 13:49 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 26022 | DOH | gemmponce11@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 623 |  | GX-IV-D-6C | 821393 | CPU shutting down | Checked connections | Performed maintenance Checked connections Checked UPS and backup Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710849948121146551/5710849948121146551_signature_28.png | Genma Ponce | https://www.jotform.com/uploads/macarewecare/222927978372067/5710849948121146551/5710849948121146551_signature_31.png
| 2023-09-21 | 2023-09-21 13:35 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 26021 | DOH | ccmclaboratory01@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1616 |  | GXIV-4D | 838561 | Module A4 - H/W Fail | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710840756378842551/5710840756378842551_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5710840756378842551/5710840756378842551_signature_31.png
| 2023-09-21 | 2023-09-21 13:22 | Toril Health Center | Toril, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26020 | PBSP | leabuelis@yahoo.com | 24.8 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1417 |  | GX-IV 4D | 838865 | Stuck Cartridge, and For Preventive Maintenance | Removal the stuck cartridge, Clean Surface Area, Cartridge Bay bay, Fan Filters and Icore Lenses. Generate IQR, Assay Stat, and System Log | Module Troubleshooting and Perform Preventive Maintenance | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710836536714351655/5710836536714351655_signature_28.png | Lea Buelis, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5710836536714351655/5710836536714351655_signature_31.png
| 2023-09-21 | 2023-09-21 12:35 | Columbio Rural Health Unit | Poblacion, Columbio, Sultan Kudarat | 12 | Sultan Kudarat | Tb-lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 26019 | DOH | covivido211@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 504 | 2023-09-21 | Gx-IV | 838556 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine annual Xpertcheck/Calibration | 4 operational modules | All module passed the annual Calibration/Xpertcheck | Xpertcheck CE-5 cartridge/kits | DR no.: 10405 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710813774917721012/5710813774917721012_signature_28.png | Rolan Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/5710813774917721012/5710813774917721012_signature_31.png
| 2023-09-21 | 2023-09-21 11:59 | Talisay City Health Office | Talisay City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Gerald Fritz Selga |  |  |  |  | 26018 | PBSP | hadmeoffguard@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-4 | 847222 | Defective UPS | Preventive Maintenance (optics cleaning, module bay cleaning, filter cleaning) | - Install new UPS (model RPT-2000AP) and AVR (model ATR-5000 Pro) - Functionality testing | Unit operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710782537126229096/5710782537126229096_signature_28.png | Hadney P. Besa | https://www.jotform.com/uploads/macarewecare/222927978372067/5710782537126229096/5710782537126229096_signature_31.png
| 2023-09-21 | 2023-09-21 11:38 | Sta. Cruz Rural Health Unit | Pio Valenzuela St. | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26017 | DOH | rhustacruz2@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1322 | 2023-09-05 | GX-IV-D-6C | 820229 | For monitor replacement | Replaced defective monitor | Performed maintenance Replaced defective monitor Generated reports | Unit is operational. Module D2/722095 not operational, for reporting. |  | Old monitor: CNK831071H New monitor: H22V5SG81A230602173 | DR# 10641 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710773798125898180/5710773798125898180_signature_28.png | Evelyn T. RAMOS | https://www.jotform.com/uploads/macarewecare/222927978372067/5710773798125898180/5710773798125898180_signature_31.png
| 2023-09-21 | 2023-09-21 10:00 | Caloocan City Quality Assurance RTDL | Sangandaan, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26016 | PBSP | crissyname@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3112 ASSAY: Xpert HIV-1 Viral Load, COUNT: 529 |  | GX IV 6C | 814154 | CPU problem for additional troubleshooting from cepheid | Performed additional troubleshooting Problem still persists CPU for reporting Generated report | Performed additional troubleshooting | GX operational CPU not working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710726262511009503/5710726262511009503_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5710726262511009503/5710726262511009503_signature_31.png
| 2023-09-21 | 2023-09-21 07:47 | Carrascal Rural Health Unit | Poblacion, Carrascal | 13 | Surigao del Sur | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 26014 | PBSP | rhucarrascal2016@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 945 |  | GX-4 | 835098 | None | For reporting to Cepheid | XpertCheck Calibration Preventive Maintenance | All modules are operational | All modules passed in XPC Calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710657906019332148/5710657906019332148_signature_28.png | Roela O. Garay | https://www.jotform.com/uploads/macarewecare/222927978372067/5710657906019332148/5710657906019332148_signature_31.png
| 2023-09-20 | 2023-09-20 17:20 | Datu halun sakilan Memorial Hospital | Lamion, datu halun sakilan Memorial hospital | BARMM | Tawi-Tawi | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26013 | DOH | dhsmhmolecularlaboratory@gmail.com | 27 |  | 2023-08-02 | GX 4 | 838834 | Module 1 hardware failed and disabled for xpertcheck calibration | Module 2, 3 and 4 undergone Xpertcheck. While module 1 undergone documentation for its status. | Troubleshoot module 1 for its disabled state during xpertcheck calibration | Module1 hardware failed and unable to conduct xpertcheck in the said module. | None | XPERTCHECK-CE-5 | 10402 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710115295918914515/5710115295918914515_signature_28.png | Almu-eed M. Abdilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5710115295918914515/5710115295918914515_signature_31.png
| 2023-09-20 | 2023-09-20 16:20 | Madrid District Hospital | Poblacion, Madrid | 13 | Surigao del Sur | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 26012 | DOH | madrid_dh@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 911 |  | GX-4 | 838439 | Module B3 found Service Required thru XPC Calibration | For reporting to Cepheid | Replacing Module B2 with SN: 800944 to SN: 731240 XpertCheck Calibration Preventive Maintenance | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710083040612135321/5710083040612135321_signature_28.png | Nhilrose S. Bontor | https://www.jotform.com/uploads/macarewecare/222927978372067/5710083040612135321/5710083040612135321_signature_31.png
| 2023-09-20 | 2023-09-20 15:27 | Banlat Health Center | Daisy St., Cor Daffodil St., Banlat Road, Tandang Sora, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26011 | PBSP | mfgadayarmt17@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 95 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1599 |  | GX IV 6C | 804688 | Module 2 Hardware Failed | Performed PM Checked Module 2 Error 1002 on Module 2 Module 2 for reporting Disable module 2 Generated report | Troubleshoot Module 2 | 3 modules operational, 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710060849919840520/5710060849919840520_signature_28.png | Maria Adaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5710060849919840520/5710060849919840520_signature_31.png
| 2023-09-20 | 2023-09-20 15:29 | Datu halun Sakilan Memorial Hospital | Lamion, datu halun sakilan Memorial Hospital | BARMM | Tawi-Tawi | Laboratory | NO | CALIBRATION | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26010 | PBSP | wmhadjula.pbsp@gmail.com | 28 |  | 2023-09-20 | GX 4 | 808643 | Module 1 hardware failed | Module 1 undergone Xpertcheck calibration and it requires service. | Module 1 undergone Xpertcheck calibration | Module 1 , 2 and 3 are available | Module 1 for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710058707283248050/5710058707283248050_signature_28.png | Wardasa M. Hadjula | https://www.jotform.com/uploads/macarewecare/222927978372067/5710058707283248050/5710058707283248050_signature_31.png
| 2023-09-20 | 2023-09-20 15:22 | Datu halun Sakilan Memorial Hospital | Lamion, Datu Halun Memorial hospital | 9 | Tawi-Tawi | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 26009 | PBSP | wmhadjula.pbsp@gmail.com | 28 |  | 2023-09-08 | GX 4 | 838454 | None | Replaced module 801136 and 802379 with Module 210100362 and 210100363 respectively. | Preventive Maintenence and Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710058697287433137/5710058697287433137_signature_28.png | Warsash M. Hadjula | https://www.jotform.com/uploads/macarewecare/222927978372067/5710058697287433137/5710058697287433137_signature_31.png
| 2023-09-20 | 2023-09-20 15:24 | Rosario Super Health Center | Sixto Ave. | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 26008 | WHO | eden.cawayan@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1982 | 2023-09-20 | GX-IV-D-6C | 843924 | Module B1 for replacement | Replaced module B1 | Performed maintenance Replaced module B1 Generated reports | Unit is operational. Module B4/210068082 not operational, for reporting. |  | Old: B1/684120 New: B1/210101926 | G 6301 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710041374414336979/5710041374414336979_signature_28.png | Eden R. Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5710041374414336979/5710041374414336979_signature_31.png
| 2023-09-20 | 2023-09-20 14:37 | Belmonte Health Center | Don Quijote St. Sampaloc | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 26007 | PBSP | spsurell@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1190 | 2023-09-20 | GXIV | 808619 | Unit for PM. For module replacement of modules B1, B2 and B4. | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | GX Module B1 (210101692), B2 (210059617) and B4 (210060024) in replacement to defectivr B1 (667470), B2 (629980) and B4 (699604). | G6298 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5710034327212056907/5710034327212056907_signature_28.png | Kristine Escarlan | https://www.jotform.com/uploads/macarewecare/222927978372067/5710034327212056907/5710034327212056907_signature_31.png
| 2023-09-20 | 2023-09-20 08:24 | Aguilar Rural Health Unit | Aguilar, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  | yass.salih97@gmail.com |  |  | 26006 | PBSP | rheynava440@gmail.com | 35 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 413 |  | GX4 | 834369 | Module B2 for additional troubleshooting | #NAME? | Performed Troubleshooting | Unit is operational | Module B2 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709987552221115111/5709987552221115111_signature_28.png | Reynaldo a. Nava | https://www.jotform.com/uploads/macarewecare/222927978372067/5709987552221115111/5709987552221115111_signature_31.png
| 2023-09-20 | 2023-09-20 13:26 | Domingo Santiago Health | 844 Domingo Santiago, 581, Manila, Metro Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26005 | DOH | fei_paas@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 999 |  | GX IV 6C | 831527 | Module 4 hardware failed | Perfomed PM Performed rexpert check on module 4 Module 4 failed xpert check Generated report | Performed Rexpert check on module 4 | 3 modules working, 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709977066011894140/5709977066011894140_signature_28.png | Flory Fe Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5709977066011894140/5709977066011894140_signature_31.png
| 2023-09-20 | 2023-09-20 12:55 | Sorsogon Provincial Hospital | Macabog, West District, Sorsogon City | 5 | Sorsogon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 26004 | PBSP | wizard682000@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 965 |  | GX-4 | 835213 | GX for Calibration | Performed XpertCheck Cleaned cartridge bays, plunger rod, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational | XpertCheck Result: C2/633491 - Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709966438917368219/5709966438917368219_signature_28.png | Reynaldo B. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5709966438917368219/5709966438917368219_signature_31.png
| 2023-09-20 | 2023-09-20 11:30 | Matanao Rural Health Unit | Poblacion, Matanao Davao Del Sur | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26003 | DOH | pamsustento@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 478 ASSAY: Xpert MTB/RIF, COUNT: 11 | 2023-08-03 | GX-IV 4D | 838456 | For XpertCheck Calibration | Run XpertCheck Calibration on Four Modules | Perform XpertCheck Calibration | Module B1 and B2 for Service Required Module B3 and B4 passed the Calibration | Modules B1 and B2 for under observation | Xpertcheck CE 5 | DR 10441 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709960042413404330/5709960042413404330_signature_28.png | Pamela Joyce P. Sustento | https://www.jotform.com/uploads/macarewecare/222927978372067/5709960042413404330/5709960042413404330_signature_31.png
| 2023-09-20 | 2023-09-20 10:28 | Commonwealth Health Center | Katuparan, Quezon City, Metro Manila | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 26002 | PBSP | commonwealthclinical@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 426 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2606 |  | GXIV 6C | 831615 | Module 2 frequent error 2014 | Performed PM Performed rexpert check on module 2 Module 2 failed xpert check Generated report | Performed Rexpert check on module 2 | 3 modules operational, 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709865756015871202/5709865756015871202_signature_28.png | Rizza Clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5709865756015871202/5709865756015871202_signature_31.png
| 2023-09-20 | 2023-09-20 09:08 | Western Visayas Medical Center | Q. Abeto St, Mandurriao, Iloilo City | 6 | Iloilo | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 26001 | PBSP | mitorilla.pbsp@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5231 |  | GX-16 | 839316 | Defective 723641/D2 due to temperature-related errors. | - Performed Preventive Maintenace (optics cleaning, module bay cleaning, plunger rod maintenance, filters cleaning). - Archive tests. | - Replaced module 723641/D2 with new module with SN 210102239. - Generated Pre- and Post -IQR. | 16 nodules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709815014669100831/5709815014669100831_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5709815014669100831/5709815014669100831_signature_31.png
| 2023-09-19 | 2023-09-19 16:13 | Antipas Rural Health Unit | Poblacion Antipas, North Cotabato | 11 | Cotabato | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 26000 | DOH | rhuantipas@gmail.com | 23.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 671 | 2023-08-03 | GX-IV 4D | 838259 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration |  | Xpertcheck CE 5 | DR 10450 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709212340411295469/5709212340411295469_signature_28.png | Alvin A. Codilla, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5709212340411295469/5709212340411295469_signature_31.png
| 2023-09-19 | 2023-09-19 16:05 | Rapu-Rapu Rural Health Unit | Purok 4, Poblacion, Rapu-Rapu | 5 | Albay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25999 | DOH | suzetteoguan1@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 388 |  | GX4 | 838565 | For Callibration | XpertCheck Result: B2/737128 - service required | Performed XpertCheck Performed Preventive Maintenance | Unit is operational | B2/73728 - Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709203671419928226/5709203671419928226_signature_28.png | Suzette C. Oguan | https://www.jotform.com/uploads/macarewecare/222927978372067/5709203671419928226/5709203671419928226_signature_31.png
| 2023-09-19 | 2023-09-19 14:55 | Santa Rosa Rural Health Unit | Rizal, Santa Rosa, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25998 | DOH | addc1995@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 205 | 2023-09-19 | Genexpert IV | 831599 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709159114249157968/5709159114249157968_signature_28.png | Angelo Espino | https://www.jotform.com/uploads/macarewecare/222927978372067/5709159114249157968/5709159114249157968_signature_31.png
| 2023-09-19 | 2023-09-19 14:20 | International Organization for Migration | HB Dela Costa, Makati | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE Relocation | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25997 | IOM | jsabalboro@iom.int | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 986 ASSAY: Others, COUNT: 142 ASSAY: Xpert SARS-CoV-2, COUNT: 16 |  | GXIV-L-10C | 110004816 | Unit is for Relocation | Relocate | Relocate GeneXpert System Performed Maintenance Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709142421515814714/5709142421515814714_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5709142421515814714/5709142421515814714_signature_31.png
| 2023-09-19 | 2023-09-19 13:11 | Dra. Elvira Lagrosa Health Center | Dolores Hidalgo St., Pasay, Metro Manila | NCR | NCR 4th - Pasay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25996 | PBSP | rhemejetbotaquia@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1488 |  | GX IV 6C | 812833 | Module 2 results always invalid | Performed PM Checked system log Only 1 error recorded on module 2 Checked reports on samples run in module 2 Most of the results are invalid End user waits 1hr then the result comes out as Invalid Archived results for further inspection of technical specialist UPS is still operational Generated report | Troubleshoot module 2 Checked UPS | Unit is operational Module 2 for checking by specualist |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709136332341940046/5709136332341940046_signature_28.png | Rheme Tañeza | https://www.jotform.com/uploads/macarewecare/222927978372067/5709136332341940046/5709136332341940046_signature_31.png
| 2023-09-19 | 2023-09-19 14:04 | Isabela Municipal Health Office | Gomez St., Isabela, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25995 | PBSP | laurelmarjorie1@gmail.com | 25 |  |  | GX-IV | 838855 | Defective CPU - No Power | #NAME? | - Checked line voltage at facility - 230 V - Check continuity of CPU power cord - good - Reset power supply connections to the motherboard - Swapped power cords - Checked power supply output - Cleaned RAM pcb connectors | CPU for replacement, for report to Cepheid |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709130346914063388/5709130346914063388_signature_28.png | Marjorie Laurel | https://www.jotform.com/uploads/macarewecare/222927978372067/5709130346914063388/5709130346914063388_signature_31.png
| 2023-09-19 | 2023-09-19 13:37 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25994 | PBSP | bryanmichaelbalunes@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1166 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4124 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 238 | 2023-09-19 | Genexpert XVI | 844111 | Module B3 broken Module Hinge | Disassemble Genexpert Take Picture of broken module door hinge Assemble Genexpert | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module B3 broken Hinge | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709115724202382334/5709115724202382334_signature_28.png | Bryan Michael Balunes | https://www.jotform.com/uploads/macarewecare/222927978372067/5709115724202382334/5709115724202382334_signature_31.png
| 2023-09-19 | 2023-09-19 12:44 | Magsaysay Rural Health Unit | Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25993 | DOH | jkakilgan12@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 264 |  | GX-4 | 838557 | Module B1 and B4 found service required thru XPC Calibration | For Reporting to Cepheid | XpertCheck Calibration Preventive Maintenance | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709107313953946279/5709107313953946279_signature_28.png | Jesusa N. Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/5709107313953946279/5709107313953946279_signature_31.png
| 2023-09-19 | 2023-09-18 13:59 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25992 | DOH | gbmmbg@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 547 |  | GX4 | 838437 | Error 2008 recurring after new module installation | #NAME? | #NAME? | Unit is operational | Module B1 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709085852227859062/5709085852227859062_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5709085852227859062/5709085852227859062_signature_31.png
| 2023-09-19 | 2023-09-19 12:24 | Guiguinto Rural Health Unit II | Guiguinto, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 25991 | PBSP | joie_calica@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA | 2023-09-14 | GXIV | 814165 | Monitor for replacement | #NAME? | Replaced defective monitor | Unit is not operational. Desktop is not powering on. (Reported) |  | GeneXpert Monitor SN: V5GAH812 | G 6281 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709070900191602292/5709070900191602292_signature_28.png | Ligaya Calica | https://www.jotform.com/uploads/macarewecare/222927978372067/5709070900191602292/5709070900191602292_signature_31.png
| 2023-09-19 | 2023-09-19 10:38 | Bagumbayan Rural Health Unit | Poblacion Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25990 | DOH | marizalenoid@yahoo.com | 28.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1174 | 2023-09-19 | Gx-IV | 838261 | Module 2 for service Required | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine annual Xpertcheck/ Calibration | 3 operational modules | Module 1,3 and 4 passed the annual Xpertcheck/Calibration | Xpertcheck CE-5 cartridge/kit | DR no.:10403 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5709023870125805292/5709023870125805292_signature_28.png | MA. RIZA R. DIONELA | https://www.jotform.com/uploads/macarewecare/222927978372067/5709023870125805292/5709023870125805292_signature_31.png
| 2023-09-19 | 2023-09-19 09:46 | Cebu TB Reference Lab-SRP | SRP-Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION MAINTENANCE ON-SITE ASSISTANCE Installation under shell foundation 10c unit | CEPHEID | Raymond Laurito |  |  |  |  | 25989 | PBSP | asmantiquilla.pbsp@gmail.com | 20 |  |  | Gx 4 | 120000133 | Genexpert installation | Installing genexpert machine | Installing genexpert machine | Systems operational | No printer installed |  |  |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5708987825612997959/5708987825612997959_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5708987825612997959/5708987825612997959_signature_31.png
| 2023-09-19 | 2023-09-19 09:46 | Cebu TB Reference Lab-SRP | SRP-Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25988 | PBSP | asmantiquilla.pbsp@gmail.com | 20 |  |  | Gx 4 | 120000133 | The unit was borrowed for training in quest hotel | Installing genexpert machine | Installing genexpert machine | Systems operational | No Charger for Laptop, No UTP cable |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708987825612997959/5708987825612997959_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5708987825612997959/5708987825612997959_signature_31.png
| 2023-09-19 | 2023-09-19 08:28 | Hinigaran Rural Health Unit | Aguilar St., Hinigaran, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25987 | PBSP | mfeliprada@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 905 |  | GX-IV | 815283 | Defective AVR and UPS | #NAME? | - Checked line voltage at facility - 234 V - Install new UPS (manufacturer Powercom, model RPT-2000AP, SN 40162162112) - Install new AVR (manufacturer RegVolt , model ATVR-5000 Pro, SN 1707111105-00023 ) - Functionality testing of system with new AVR and UPS installed - Replaced defective Keyboard and Mouse with new USB-connected Keyboard and mouse. | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708930855819956265/5708930855819956265_signature_28.png | MayFlor M. Feliprada | https://www.jotform.com/uploads/macarewecare/222927978372067/5708930855819956265/5708930855819956265_signature_31.png
| 2023-09-19 | 2023-09-14 20:42 | Josefina Belmonte Duran Albay Provincial Hospital | Ligao | 5 | Albay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25986 | PBSP | bermundocristine@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1828 |  | GX4 | 814371 | CPU - hardrive not found | Perform hardrive diagnostic Reset CPU | Perform hardrive diagnostic Reset CPU | Unit is not operational | CPU for Replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708921832911565067/5708921832911565067_signature_28.png | Cristine Bermundo,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5708921832911565067/5708921832911565067_signature_31.png
| 2023-09-19 | 2023-09-15 12:54 | Virac Rural Health Unit | San isidro village, Virac | 5 | Catanduanes | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25985 | DOH | limarmariano50@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1743 |  | GX4 | 838563 | B4/863009 for Module Replacement | Replaced Module: B4-863009>>785028 Re-XpertCheck Result: Cleaned cartridge bays, plunger rod, icore, fan Generate reports | Replaced B4/863009 Module Performed Re-XpertCheck for B2/862611 Performed Preventive Maintenance | Unit is operational | Re-XpertCheck B2/862611 - |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708918492914790979/5708918492914790979_signature_28.png | Ramil A. Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/5708918492914790979/5708918492914790979_signature_31.png
| 2023-09-18 | 2023-09-14 09:21 | De La Salle Health and Sciences Institute | D. Mangubat Ave., Dasmariñas City | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25984 | CONSORTIUM | acpabruada@dlsmhsi.edu.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1350 | 2023-09-11 | GXXVI-D | 842166 | Module C2 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module C2 Generate Reports | All Modules are Operational | C2: 880228 -> 639226 | GX Module | DR 10653 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349453619905015/5708349453619905015_signature_28.png | Angie Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349453619905015/5708349453619905015_signature_31.png
| 2023-09-18 | 2023-09-12 11:55 | Asian Hospital and Medical Center | Civic Dr., Filinvest | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  |  |  |  | 25979 | MACARE-HBDC | bloodbank@asianhospital.com | 24 |  |  | MCS+ | 10B129 | Unit is due for Preventive Maintenance | Preventive Maintenance | Checked Parameters and Sensors Value Checked A/D Values Cleaned Rotors | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383611705544/5708349383611705544_signature_28.png | Maria Vanessa Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383611705544/5708349383611705544_signature_31.png
| 2023-09-18 | 2023-09-12 11:13 | Asian Hospital and Medical Center | Civic Dr., Filinvest | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Brenguiler Bagbaguen |  |  |  |  | 25980 | MACARE-HBDC | bloodbank@asianhospital.com | 24 |  |  | MCS+ | 01F116 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Simulation Checked Parameters and Sensors Value Cleaned Rotors | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383612231692/5708349383612231692_signature_28.png | Maria Vanessa Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383612231692/5708349383612231692_signature_31.png
| 2023-09-18 | 2023-09-13 08:10 | Seamen's Hospital | General Luna St., Intramuros | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25981 | MACARE-HBDC | shmolecularlab@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 697 |  | GXXVI-D | 838777 | Unit is due for Maintenance | Preventive Maintenance | Performed Maintenance Performed Self-test Checked Module D4 Performed Functionality Test Generate Reports | 14 Modules are Operational | Module D4 with Temperature Related Errors |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383615291645/5708349383615291645_signature_28.png | Queenie Lea Serquina | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383615291645/5708349383615291645_signature_31.png
| 2023-09-18 | 2023-09-13 11:03 | Alabang Health Center | Liwasan ng mga Bayani, Alabang | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25983 | DOH | cathycarasig@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1833 |  | GXIV-D | 831913 | Modules 3 and 4 with Hardware Problems | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 2 Modules are Operational, 2 are Defective | Modules 3 and 4 Failed in XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383615736027/5708349383615736027_signature_28.png | Catherine carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383615736027/5708349383615736027_signature_31.png
| 2023-09-18 | 2023-09-13 08:44 | Seamen's Hospital | General Luna St, Intramuros | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25982 | MACARE-HBDC | shmolecularlab@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 21 |  | GXIV-D | 838836 | Unit is due for Maintenance | Preventive Maintenance | Performed Maintenance Performed Self-test Performed Functionality Test Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383616259345/5708349383616259345_signature_28.png | Queenie Lea Serquina | https://www.jotform.com/uploads/macarewecare/222927978372067/5708349383616259345/5708349383616259345_signature_31.png
| 2023-09-18 | 2023-09-18 13:56 | Davao del Norte Provincial Hospital-Samal | Brgy. Peñaplata, Island Garden City of Samal | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25978 | PBSP | allensoriano92@yahoo.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 875 |  | GX-IV 4D | 815383 | Module B1 not detected by Module | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Troubleshoot the Module B1 by reseating the components, still error exist having infinite loop of the base. Tried to switch places with other module, still error exist; also cannot be detected regardless where to swap the module. | Module Troubleshooting | Module B1 cannot be detected by the Gx Software. Only three Modules are operational | Errors and documentation sent to Cepheid for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708265028219461983/5708265028219461983_signature_28.png | ALLEN GLENN B. SORIANO | https://www.jotform.com/uploads/macarewecare/222927978372067/5708265028219461983/5708265028219461983_signature_31.png
| 2023-09-18 | 2023-09-18 13:25 | Nasipit Rural Health Unit | Brgy. Talisay, Nasipit | 13 | Agusan del Norte | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25977 | DOH | peterdolotina@gmail.com | 26.15 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1443 |  | GX-4 | 837611 | B4 was found service required thru XPC Calibration | None | Replacing Module B2 with SN: 698547 to SN: 673655 Preventive Maintenance XpertCheck Calibration | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708246260175538971/5708246260175538971_signature_28.png | Peter Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/5708246260175538971/5708246260175538971_signature_31.png
| 2023-09-18 | 2023-09-18 11:23 | Bacolod City Health Office | Galo St., Bacolod City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25976 | PBSP | cyrilcordova0515@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6657 |  | GX-IV | 812020 | - Module 700546/D2 HW fail with temperature-related errors. | #NAME? | - Replaced 700546/D2 with new module with SN 210102239. - Generate Pre- and Post- IQR | - Module 640561/D1 waiting for allocated module. - Modules D2, D3 and D4 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708169445417755317/5708169445417755317_signature_28.png | CYRIL CORDOVA | https://www.jotform.com/uploads/macarewecare/222927978372067/5708169445417755317/5708169445417755317_signature_31.png
| 2023-09-18 | 2023-09-18 10:10 | Negros Oriental Provincial Hospital RTDL | Dunamguety City | 7 | Negros Oriental | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 25975 | PBSP | rezelmarielagos@gmail.com | 23 |  | 2023-09-05 | Gx4 | 805275 | Calibration fail | Replace Module C3 with new serial number 210102245 | Replace Module C3 with serial number 648367 | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708126223916835010/5708126223916835010_signature_28.png | Rezel Marie Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5708126223916835010/5708126223916835010_signature_31.png
| 2023-09-18 | 2023-09-18 09:33 | GOVERNOR CELESTINO GALLARES MEMORIAL MEDICAL CENTER | TAGBILARAN, BOHOL | 7 | Bohol | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 25974 | PBSP | majieocmen@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 54 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2212 ASSAY: Xpert HIV-1 Viral Load, COUNT: 110 |  | GX-IV | 838878 | FOR MODULE B4 REPLACEMENT | COLLECT IQR, ARCHIVE, AND SYSTEM LOG | PERFORM MODULE B4 DEF SN:618997 REPLACEMENT WITH NEW SN:210102255 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708108943913951435/5708108943913951435_signature_28.png | Ma. Jernai Sosoban, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5708108943913951435/5708108943913951435_signature_31.png
| 2023-09-18 | 2023-09-18 08:34 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25973 | COMMERCIAL | jonnifercodoyceniza@gmail.com | 28 |  |  | Gx 4 | 829419 | Borrowed machine to quest hotel representation | Re installaton of genexpert machine | Re installaton of genexpert machine | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5708070418225901700/5708070418225901700_signature_28.png | Jonnifer C. Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5708070418225901700/5708070418225901700_signature_31.png
| 2023-09-16 | 2023-09-16 09:06 | Philippine Heart Center | East,Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 25972 | MACARE-HBDC | phcmolecularlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1085 ASSAY: Others, COUNT: 12 ASSAY: Xpert SARS-CoV-2, COUNT: 4082 |  | GXXVI | 838655 | Workstation unable to boot 3 short beep and 1 long beep | Checked workstation Cleaned terminal RAM Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Test run unit | Troubleshooting | Workstation operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5706447232414410545/5706447232414410545_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5706447232414410545/5706447232414410545_signature_31.png
| 2023-09-16 | 2023-09-15 11:41 | Antipolo CHO | M.Santos St.,Barangay San Roque ,Antipolo | 4A | Rizal | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25971 | PBSP | zeraus.tine@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 80 ASSAY: Xpert MTB/RIF Ultra, COUNT: 8907 ASSAY: Xpert HIV-1 Viral Load, COUNT: 744 | 2023-09-14 | GX XVI | 842226 | For module replacement and maintenance A3 885783 - 210102202 | Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Performed module replacement Test run unit | Module replacement and maintenance | A2 encounter 1002 and 4012 |  | 900-0508 | G6278 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5706445532415969780/5706445532415969780_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5706445532415969780/5706445532415969780_signature_31.png
| 2023-09-16 | 2023-09-14 14:26 | International Organization for Migration | H.V Dela Costa,Makati | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25970 | IOM | jsabalboro@iom.int | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 992 ASSAY: Others, COUNT: 142 ASSAY: Xpert SARS-CoV-2, COUNT: 16 | 2023-09-14 | GX4 | 805288 | For Module Replacement D3 680216 - 631961 | Performed module replacement, Cleaned filter,plunger rod,cartridge bay Gathered reports Test run unit | Module Replacement | Operational |  | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5706444412418078603/5706444412418078603_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5706444412418078603/5706444412418078603_signature_31.png
| 2023-09-15 | 2023-09-15 15:31 | Angono Rural Health Unit | Angono | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25969 | DOH | siennablanco@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 945 | 2023-09-05 | GX-IV-D-6C | 819864 | UPS for replacement | Replaced defective UPS | Replaced defective UPS Performed maintenance Generated reports | Unit is operational. Module A4/719514 not operational, reported already. | Monitor not powering on. Monitor SN: CNK831082Y | UPS SN: AT1500211227ST0025 | DR# 10644 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705730520526431714/5705730520526431714_signature_28.png | Rose Sienna Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5705730520526431714/5705730520526431714_signature_31.png
| 2023-09-15 | 2023-09-15 09:24 | Eastern Visayas Regional Medical Center | Bagacay, Tacloban city | 8 | Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25968 | PBSP | shyzamoras@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 500 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4412 ASSAY: Xpert HIV-1 Viral Load, COUNT: 283 | 2023-09-15 | GXIV | 808631 | N/a | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705727438216420130/5705727438216420130_signature_28.png | Shirlyn C. Zamoras | https://www.jotform.com/uploads/macarewecare/222927978372067/5705727438216420130/5705727438216420130_signature_31.png
| 2023-09-15 | 2023-09-15 13:55 | Tagum City Health Office | Mabini St., Poblacion Tagum City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25967 | DOH | tagumcho.mabini@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2693 |  | GX-IV 4D | 838553 | Error occurred to B1 and B3 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses. Reposition the GeneXpert Setup to optimal area. Generate IQR, Assay Stat, and System Log | Performed Module Troubleshooting | Module B1 for under observation Module B3 report to Cepheid for further Assessment | B3 currently disabled due to frequent H/W Failure also Enduser is advice to have supplementary cooling system in meet the optimal room temperature prescribed by GeneXpert |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705689222012859590/5705689222012859590_signature_28.png | Lorena penas | https://www.jotform.com/uploads/macarewecare/222927978372067/5705689222012859590/5705689222012859590_signature_31.png
| 2023-09-15 | 2023-09-15 14:20 | Puerto Princesa CHO | Puerto Princesa, Palawan | 4B | Palawan | TB DOTS | NO | MAINTENANCE Gather data | CEPHEID | Harold Toco |  |  |  |  | 25966 | PBSP | jocal0324@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1676 | 2023-09-15 | Gx4 | 804272 | For unit checkbup and maintenance | Performed unit maintenance | Performed unit maintenance. | All 4 modules operational (Availabe ststus) | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705685080901984598/5705685080901984598_signature_28.png | JOCELYN L. CALALIN | https://www.jotform.com/uploads/macarewecare/222927978372067/5705685080901984598/5705685080901984598_signature_31.png
| 2023-09-15 | 2023-09-15 14:10 | Puerto Princesa CHO | Puerto Princesa Palawan | 4B | Palawan | TB DOTS | NO | MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 25965 | MACARE-HBDC | jocal0324@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1157 | 2023-09-15 | Gx4 | 832026 | For check up and maintenance | Performed maintenance | Performed maintenance | 3 modules operational 1 disable status | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705677330908359354/5705677330908359354_signature_28.png | Jocelyn L.Calalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5705677330908359354/5705677330908359354_signature_31.png
| 2023-09-15 | 2023-09-15 13:33 | Candelaria Rural Health Unit | Rizal St., Brgy. Poblacion, Candelaria, Quezon City | 4A | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25964 | PBSP | shirleyzrubz@gmail.com | 32 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1659 |  | GX IV 6C | 814155 | Defective Monitor | Checked and cleaner cartridge bays, icore slot, plunger rod and filter Troubleshoot monitor Monitor black screen with white lines Power chord ok, vga cable ok; problem persists Monitor for reporting Generated report | Checked Monitor | GX is operational Uses back up monitor for tge meantime |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705664091615493132/5705664091615493132_signature_28.png | Shirley R. Dollente | https://www.jotform.com/uploads/macarewecare/222927978372067/5705664091615493132/5705664091615493132_signature_31.png
| 2023-09-15 | 2023-09-15 13:19 | East Rembo Health Center | 23rd Ave | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25963 | PBSP | eastrembosatellitelaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 468 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1606 | 2023-09-14 | GX-IV-D-6C | 807591 | Module B3 for replacement | Replaced module B3 | Performed maintenance Replaced module B3 Generated reports | All modules operational |  | Old: B3/602562 New: B3/210102221 | G 6280 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705647054826091767/5705647054826091767_signature_28.png | Nesty Bañaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5705647054826091767/5705647054826091767_signature_31.png
| 2023-09-15 | 2023-09-15 12:47 | Mandaluyong City Health Office | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25962 | PBSP | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2539 ASSAY: Xpert HIV-1 Viral Load, COUNT: 141 | 2023-08-14 | GXIV | 805062 | Module A2 for replacement | #NAME? | Replaced A2 | Unit is operational. |  | GeneXpert Module SN: 210102224 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705628890125058418/5705628890125058418_signature_28.png | Rosevi Jalova | https://www.jotform.com/uploads/macarewecare/222927978372067/5705628890125058418/5705628890125058418_signature_31.png
| 2023-09-15 | 2023-09-15 11:54 | Quirino Province Medical Center | Mangandingay, Cabarroguis, Quirino | 2 | Quirino | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25961 | PBSP | ylbumidang0817@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1335 | 2023-09-15 | Genexpert IV | 814166 | Module 3 error 2005 | Enabled Module 3 for observation | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 690618 - 210100889 | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705601053323204249/5705601053323204249_signature_28.png | Yvone Bumidang | https://www.jotform.com/uploads/macarewecare/222927978372067/5705601053323204249/5705601053323204249_signature_31.png
| 2023-09-15 | 2023-09-15 11:26 | Ospital Ng Palawan | Puerto Princesa, Palawan | 4B | Palawan | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 25960 | PBSP | amabduljalil.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2886 ASSAY: Xpert SARS-CoV-2, COUNT: 1602 ASSAY: Xpert HIV-1 Viral Load, COUNT: 174 | 2023-09-15 | Gx16 | 842893 | For Module replacement.Modules C3 and C4 | Replaced module C3: 888456 and C4: 888295 | Performed maintenance.. Performed module replacement | 14 modules operational..2 modules disabled status due to Hardware issue.. | New modules C3: 210102265 and C4: 210102269 | Stated | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705583460903892335/5705583460903892335_signature_28.png | Arabia Abduljalil | https://www.jotform.com/uploads/macarewecare/222927978372067/5705583460903892335/5705583460903892335_signature_31.png
| 2023-09-15 | 2023-09-15 09:11 | Isabela city health office | Isabela city, Basilan | 9 | Basilan | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25959 | PBSP | reynantemontano@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 870 ASSAY: Xpert CT/NG, COUNT: 1 | 2023-09-15 | GX 4 | 807595 | None | Replaced module 662307 with Module 210100722 | Module Replacement and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705500676296381175/5705500676296381175_signature_28.png | Reynante L Montaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5705500676296381175/5705500676296381175_signature_31.png
| 2023-09-14 | 2023-09-14 20:42 | Basilan General hospital | Kilometre 2 Binaungan Isabela city, Basilan | BARMM | Basilan | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25958 | PBSP | ell2080@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 359 | 2023-09-14 | GX 2 | 707515 | None | Replaced module 694833 with Module 210100452 | Module Replacement and Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5705053212402047447/5705053212402047447_signature_28.png | Allene Ventura | https://www.jotform.com/uploads/macarewecare/222927978372067/5705053212402047447/5705053212402047447_signature_31.png
| 2023-09-14 | 2023-09-14 15:34 | Rogaciano M. Mercado Memorial General Hospital | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25957 | PBSP | alona0276@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2160 |  | GXIV | 805252 | Stuck cartidge on A1 | #NAME? | Removed stuck cartridge | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704893837227364908/5704893837227364908_signature_28.png | Alona Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5704893837227364908/5704893837227364908_signature_31.png
| 2023-09-14 | 2023-09-14 14:54 | Malungon Rural Health Unit | Poblacion, Malungon, Sarangani Province | 12 | Sarangani | Tb-lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25956 | DOH | ulbenarioe@yahoo.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 493 | 2023-09-14 | Gx-IV | 838546 | Module 2 and 3 for service required | Perform Genexpert machine Preventive maintenance | Replace defective module 2 SN: 862856 with new module SN: 210074760 Replace defective module 3 SN: 862854 with new module SN: 664419 | 4 operational modules |  | Genexpert module 6 color SN:210074760 and SN:664419 | DR no.: 10541 and Dr no.:10535 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704868701217120034/5704868701217120034_signature_28.png | Ermie D Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5704868701217120034/5704868701217120034_signature_31.png
| 2023-09-14 | 2023-09-14 15:24 | Esperanza Rural Health Unit | Poblacion, Esperanza | 13 | Agusan del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 25955 | WHO | lilibethcacal123@gmail.com | 27.45 | ASSAY: Xpert MTB/RIF, COUNT: 99 ASSAY: Xpert MTB/RIF Ultra, COUNT: 816 |  | GX-4 | 847258 | None | None | Replacing Module B4 with SN: 910241 to SN: 210101929 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704860472021617430/5704860472021617430_signature_28.png | Lilibeth Cacal | https://www.jotform.com/uploads/macarewecare/222927978372067/5704860472021617430/5704860472021617430_signature_31.png
| 2023-09-14 | 2023-09-14 14:43 | Baras Rural Health Unit | JP Rizal St. Baras | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25954 | PBSP | mesu.baras@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 664 | 2023-09-14 | GX-IV-D-6C | 815389 | Module B3 for replacement | Replaced module B3 | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Replaced module B3 Generated reports | All modules operational. |  | Old: B3/696221 New: B3/210102200 | G 6277 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704848555219258236/5704848555219258236_signature_28.png | MICHELLE M. AVELA | https://www.jotform.com/uploads/macarewecare/222927978372067/5704848555219258236/5704848555219258236_signature_31.png
| 2023-09-14 | 2023-09-14 09:21 | Sogod District Hospital | Sogod Southern Leyte | 8 | Southern Leyte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25953 | PBSP | sogoddistricthospital@gmail.com | 25 |  | 2023-09-14 | GXIV | 807465 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704834860515416393/5704834860515416393_signature_28.png | Ruñitha Simeona G. Olais | https://www.jotform.com/uploads/macarewecare/222927978372067/5704834860515416393/5704834860515416393_signature_31.png
| 2023-09-14 | 2023-09-11 13:26 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25952 | DOH | gbmmbg@yahoo.com | 28 |  |  | GX4 | 838437 | Recurring Error 2008 | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Relayed Cephid's instructions for sample prep - Generated Reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704776342224394223/5704776342224394223_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5704776342224394223/5704776342224394223_signature_31.png
| 2023-09-14 | 2023-09-14 12:42 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 25951 | PBSP | bmcgenexpertlab@yahoo.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 5571 |  | GXIV | 834623 | Various errors on GeneXpert | - scanned system log and test reports (found 5007, 1012, 2008, 2125) - performed self test on modules (good) - performed reliabilty test on C1 to test if error 1012 persists (passed) - checked and cleaned cartridge bay, plunger rod, filter - generated reports | Performed troubleshooting | Unit is operational. 5007 - sample related error 2008 - cartridge error 2125 - insufficient sample volume 1012 - cartridge / module error No hardware problems found. GeneXpert is in good working environment. All 4 modules are in good condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704776287226433846/5704776287226433846_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5704776287226433846/5704776287226433846_signature_31.png
| 2023-09-14 | 2023-09-11 13:26 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25950 | DOH | gbmmbg@yahoo.com | 28 |  |  | GX4 | 838437 | Recurring Error 2008 | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Relayed Cephid's instructions for sample prep - Generated Reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704768592223995405/5704768592223995405_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5704768592223995405/5704768592223995405_signature_31.png
| 2023-09-14 | 2023-09-13 11:29 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25949 | PBSP | jeffreyvrumbaoa@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 815 |  | GX4 | 838473 | New Module B1 all results Invalid. Module B3 for ReXpertCheck | #NAME? | Performed Troubleshooting and ReXpertCheck | Unit is operational | Module B3 passed, for Monitoring |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704766482225959195/5704766482225959195_signature_28.png | Jeffrey V. Rumbaoa | https://www.jotform.com/uploads/macarewecare/222927978372067/5704766482225959195/5704766482225959195_signature_31.png
| 2023-09-14 | 2023-09-13 11:29 | Umingan Community Hospital | Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25948 | PBSP | jeffreyvrumbaoa@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 815 |  | GX4 | 838473 | New Module B1 all results Invalid. Module B3 for ReXpertCheck | #NAME? | Performed Troubleshooting and ReXpertCheck | Unit is operational | Module B3 passed, for Monitoring |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704765932229587969/5704765932229587969_signature_28.png | Jeffrey V. Rumbaoa | https://www.jotform.com/uploads/macarewecare/222927978372067/5704765932229587969/5704765932229587969_signature_31.png
| 2023-09-14 | 2023-09-13 14:16 | Eastern Pangasinan District Hospital | Tayug, Pangasinan | 1 | Pangasinan | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25947 | PBSP | epdhlaboratory@gmail.com | 27 |  |  | GX4 | 815394 | CPU Red screen upon start up | #NAME? | Performed Troubleshooting | Unit is not operational, | CPU for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704762292225884212/5704762292225884212_signature_28.png | Justine Joy C. Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5704762292225884212/5704762292225884212_signature_31.png
| 2023-09-14 | 2023-09-08 13:47 | Cauayan Municipal Health Office | Municipality of Cauayan | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 25946 | DOH | djeatrmt@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1268 |  | GX4 | 838477 | Module B1 with SN 862002 with hardware failed. | Generate files(IQR & System Log). Condu t PM | Install the new module for B1 with SN 862945. | All modules are now in good running condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704732178425396152/5704732178425396152_signature_28.png | Doriza Joy A. Tizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5704732178425396152/5704732178425396152_signature_31.png
| 2023-09-14 | 2023-09-08 13:47 | Cauayan Municipal Health Office | Municipality of Cauayan | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 25945 | DOH | djeatrmt@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1268 |  | GX4 | 838477 | Module B1 with SN 862002 with hardware failed. | Generate files(IQR & System Log). Condu t PM | Install the new module for B1 with SN 862945. | All modules are now in good running condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704731868423250710/5704731868423250710_signature_28.png | Doriza Joy A. Tizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5704731868423250710/5704731868423250710_signature_31.png
| 2023-09-14 | 2023-09-14 11:49 | Macalelon Rural Health Unit | Brgy. Pinagbayanan, Macalelon, Quezon | NCR | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25944 | DOH | yesha_7388@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 404 | 2023-09-12 | GX IV 6C | 831636 | Unit for XpertCheck | Checked and cleaned cartridge bays, plunger rod, icore slot and filter Performed XpertCheck Modules for monitoring incase error occurs again Generated report | Performed XpertCheck All modules passed | Unit is operational All modules passed xpertcheck |  | XpertCheck Kit | DR 10656 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704730616218944260/5704730616218944260_signature_28.png | Cheyster O. Agustin | https://www.jotform.com/uploads/macarewecare/222927978372067/5704730616218944260/5704730616218944260_signature_31.png
| 2023-09-14 | 2023-09-14 11:24 | Pulupandan RHU | Pulupandan, Negros Occ. | 6 | Negros Occidental | TB-DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25943 | WHO | jackieorcajada@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 381 |  | GX4 | 847255 | Module B2 with SN 910236 already reported for module replacement. | Conduct PM and generate files(IQR, System Log ang GXC File) | Run Xpert check calibration for module B1, B3 and B4. Installed their newly purchased AVR with 3KVA capacity. | Module B1 and B4 passed the calibration while module B3 with SN 910237 failed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704725918427124405/5704725918427124405_signature_28.png | Jackie Lou Orcajada | https://www.jotform.com/uploads/macarewecare/222927978372067/5704725918427124405/5704725918427124405_signature_31.png
| 2023-09-14 | 2023-09-14 11:07 | Mallig Super Health Center | Centro 2, Mallig, Isabela | 2 | Isabela | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25942 | PBSP | rhumallig@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 891 | 2023-09-14 | Genexpert IV | 834686 | N/a | N/a | Installed Genexpert Installed PC Installed Printer Installed UPS Installed Monitor Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704705220913127620/5704705220913127620_signature_28.png | Hanna Jane Respicio | https://www.jotform.com/uploads/macarewecare/222927978372067/5704705220913127620/5704705220913127620_signature_31.png
| 2023-09-14 | 2023-09-14 08:33 | Cawayan Rural Health Unit | Poblacion, Cawayan | 5 | Masbate | Laboratory | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25941 | DOH | gellatisimo@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 216 ASSAY: Xpert MTB/RIF Ultra, COUNT: 848 |  | GX4 | 838267 | C1/807602 for Re-XpertCheck | Re-XpertCheck Result: C1/807602 Passed Generate Reports | Performed Re-XpertCheck for C1 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704617395455228531/5704617395455228531_signature_28.png | Gellie Jane Gader | https://www.jotform.com/uploads/macarewecare/222927978372067/5704617395455228531/5704617395455228531_signature_31.png
| 2023-09-14 | 2023-09-13 17:08 | Cawayan Rural Health Unit | Poblacion, Cawayan | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25940 | DOH | gellatisimo@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 216 ASSAY: Xpert MTB/RIF Ultra, COUNT: 848 |  | GX4 | 838267 | C4/807594 for Module Replacement | Replaced Module: C4-807594>>621794 Cleaned cartridge bays, plunger rod, icore, fan, fan filter Generate reports | Replaced Module C4 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704607845451451798/5704607845451451798_signature_28.png | Gellie Jane Gader | https://www.jotform.com/uploads/macarewecare/222927978372067/5704607845451451798/5704607845451451798_signature_31.png
| 2023-09-14 | 2023-09-13 12:47 | Aroroy Rural Health Unit | Baga-uma | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25939 | DOH | danemilbodino23@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 197 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1109 |  | GX4 | 838268 | B2-681268 for Module Replacement | Replaced Module: B2-681268>>879219 Cleaned cartridge bays, plunger rod, icore, fan, fan filter Generate reports | Replaced B2 Module Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704599675451610217/5704599675451610217_signature_28.png | Rhena M. Etcobanez | https://www.jotform.com/uploads/macarewecare/222927978372067/5704599675451610217/5704599675451610217_signature_31.png
| 2023-09-13 | 2023-09-13 13:34 | St. Bernard Rural Health Unit | St. Bernard Southern Leyte | 8 | Southern Leyte | TB-Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25938 | PBSP | gmccap2@yahoo.com | 29 |  | 2023-09-13 | GXIV | 835062 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed (C1, C2) | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704094892011780087/5704094892011780087_signature_28.png | Gary M. Capili | https://www.jotform.com/uploads/macarewecare/222927978372067/5704094892011780087/5704094892011780087_signature_31.png
| 2023-09-13 | 2023-09-13 15:40 | Rodriguez Rhu | J.P Rizal,Rodriguez,Montalban | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25937 | PBSP | karlbanal25@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2528 ASSAY: Xpert HIV-1 Viral Load, COUNT: 36 |  | GX4 | 807580 | Due for Maintenance | Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704006697815211282/5704006697815211282_signature_28.png | Karl Michael Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5704006697815211282/5704006697815211282_signature_31.png
| 2023-09-13 | 2023-09-12 11:32 | Biopath Makati | Medical Plaza Makati Amorsolo st,Makati City | NCR | NCR 4th - Makati | Laboratry | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25936 | COMMERCIAL | laboratory@biopath.ph | 33 |  |  | GX4 | 812572 | Due for maintenance, CPU unable to boot | Performed Maintenance Clleaned Filter,plunger rod,cartridge bay,icore, Gathered reports | Maintenance,and troubleshooting | GX operational,CPU booting problem, Mafunction of ups switch |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704005497819372402/5704005497819372402_signature_28.png | Biopath Clinical Diagnostics, Inc. | https://www.jotform.com/uploads/macarewecare/222927978372067/5704005497819372402/5704005497819372402_signature_31.png
| 2023-09-13 | 2023-09-13 15:08 | Rodriguez RHU | J.P Rizal, Rodriguez,Montalban | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25935 | PBSP | karlbanal25@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2497 ASSAY: Xpert HIV-1 Viral Load, COUNT: 115 |  | GX4 | 838894 | Due for Maintenance | Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Test run unit | Maintenance | B1 encounter temp error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5704003077818016238/5704003077818016238_signature_28.png | Karl Michael M. Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5704003077818016238/5704003077818016238_signature_31.png
| 2023-09-13 | 2023-09-13 11:19 | Democrito O. Plaza Memorial Hospital | Brgy. Patin-ay, Prosperidad | 13 | Agusan del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 25934 | PBSP | lermabpabillan@yahoo.com | 20.57 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 336 |  | GX-4 | 805335 | None | None | Replacing Module B3 with SN: 210067890 to SN: 210101815 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703910015202898651/5703910015202898651_signature_28.png | Lerma B. Pabillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5703910015202898651/5703910015202898651_signature_31.png
| 2023-09-13 | 2023-09-13 11:19 | M. Icasiano Health Center | 1806 Pedro Gil St., Paco, Manila | NCR | NCR 1st - Manila | Laboratoy | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25933 | DOH | lorrainefabros@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 652 | 2023-09-06 | GXIV 6C | 831684 | Monitor for replacement | Performed PM Replaced Monitor Intermittent problem on CPU CPU for reporting Generated report | Replaced Monitor: CNK8322465 -> H22V5SG81A230602126 Troubleshoot CPU | Monitor and GX working CPU for replacement |  | NVisi9n Monitor | DR10646 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703849989512146341/5703849989512146341_signature_28.png | Maruja valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/5703849989512146341/5703849989512146341_signature_31.png
| 2023-09-13 | 2023-09-13 09:24 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25932 | PBSP | crmctblab@gmail.com | 25.6 | ASSAY: Xpert HBV VL, COUNT: 203 ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2945 ASSAY: Xpert HIV-1 Viral Load, COUNT: 269 | 2023-09-13 | Gx-XVI | 839307 | Module 2 service required | Perform Genexpert machine Preventive maintenance | Replace defective module SN: 675715 with new module SN: 210101739 | 16 operational modules |  | Genexpert module 6 color SN: 210101739 | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703809232809996516/5703809232809996516_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5703809232809996516/5703809232809996516_signature_31.png
| 2023-09-13 | 2023-09-13 10:02 | Kalinga Provincial Hospital | Laboratory | CAR | Kalinga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25931 | DOH | leombelgica139@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1166 | 2023-08-24 | Genexpert IV | 838466 | Module 1 and 2 error | Performed Xpertcheck on Module 2 | Performed Maintenance Gathered Reports Performed Module Replacement (210000689 & 861497 - 766317 & 888009) Performed Functionality Test | Unit is Operational | Xpertcheck Result B2 = Service Required | Genexpert Module & Xpertcheck 2pc | 10573 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703800744911307148/5703800744911307148_signature_28.png | Leo Mark Belgica | https://www.jotform.com/uploads/macarewecare/222927978372067/5703800744911307148/5703800744911307148_signature_31.png
| 2023-09-13 | 2023-09-13 09:42 | Kalinga Provincial Hospital | Bulanao, Tabuk City, Kalinga | CAR | Kalinga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25930 | PBSP | leombelgica139@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1864 | 2023-09-13 | Genexpert IV | 807622 | Module 2 error 2008 | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 2 for Reporting | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703789974919231459/5703789974919231459_signature_28.png | Leo Mark Belgica | https://www.jotform.com/uploads/macarewecare/222927978372067/5703789974919231459/5703789974919231459_signature_31.png
| 2023-09-12 | 2023-09-12 15:40 | Masbate City Health Office | City Hall Compound, Brgy. Calipay, Masbate | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25929 | PBSP | desrmt2017@gmail.com | 26 |  |  | GX4 | 808633 | Module B2/730126, B3/662583, B4/635443 for replacement | Replaced module: B2-730126>>210101870, B3-662583>>210101899, B4-635443>>210101900 Cleaned fan, fan filter B1 cartridge bays, plunger rod, Icore Generate reports | Replaced B2,B3, B4 modules Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703402900812739009/5703402900812739009_signature_28.png | Desiree Mae D Dela Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5703402900812739009/5703402900812739009_signature_31.png
| 2023-09-12 | 2023-09-11 16:49 | Pilar Rural Health Office | Pilar | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25928 | DOH | manalangmark88@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1411 |  | GX4 | 838266 | B2/877102, B3/807693 for Module Replacement | Replaced Module: B2/877102>>650253, B3/807693>>880428 Cleaned cartridge bays, plunger rod, icore, fan, fan filter Generate reports | Replaced Module B2, B3 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703402300816812817/5703402300816812817_signature_28.png | Mark Ramir C Manalang Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5703402300816812817/5703402300816812817_signature_31.png
| 2023-09-12 | 2023-09-12 14:57 | Quezon City General Hospital | Seminary Rd., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25927 | MACARE-HBDC | treatmenthub.qcgh@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GXIV 10C | 110012766 | Printer not setup yet | Performed PM Download and Installed driver for printer Print test page Generated report | Setup printer | Unit is operational Printer working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703113566917115324/5703113566917115324_signature_28.png | Leagin palmejar | https://www.jotform.com/uploads/macarewecare/222927978372067/5703113566917115324/5703113566917115324_signature_31.png
| 2023-09-12 | 2023-09-12 14:43 | Mindanao central sanitarium | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25926 | DOH | ybioamie@gmail.com | 18 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1974 | 2023-08-02 | GX 4 | 821214 | None | Module 3 and 4 undergone Xpertcheck calibration | Module 3 and 4 undergone Xpertcheck calibration | All modules are available module 3 and 4 passes the calibration | None | XPERTCHECK-CE-5 | 10398 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703112076375195372/5703112076375195372_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5703112076375195372/5703112076375195372_signature_31.png
| 2023-09-12 | 2023-09-12 14:37 | Amang Rodriguez Memorial Medical Center | Sumulong Highway | NCR | NCR 2nd - Marikina | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25925 | PBSP | pathologymole.armmc@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 3632 |  | GX-IV-D-6C | 809257 | For Installation | Installed the GeneXpert unit | Installed the GeneXpert unit Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. | Unit from NBP |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703105377716239394/5703105377716239394_signature_28.png | Vita Salen | https://www.jotform.com/uploads/macarewecare/222927978372067/5703105377716239394/5703105377716239394_signature_31.png
| 2023-09-12 | 2023-09-12 13:52 | Angeles Rural Health Unit 3 | Angeles Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25924 | PBSP | femiepangilinan@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2674 ASSAY: Xpert HIV-1 Viral Load, COUNT: 741 |  | GXIV-D | 807569 | B2 for Re-xpertcheck | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | B2 failed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703076200894736599/5703076200894736599_signature_28.png | Femie Pangilinan | https://www.jotform.com/uploads/macarewecare/222927978372067/5703076200894736599/5703076200894736599_signature_31.png
| 2023-09-12 | 2023-09-12 11:48 | Caloocan City Health Office | Sangandaan, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25923 | PBSP | caloocancentrallaboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3112 ASSAY: Xpert HIV-1 Viral Load, COUNT: 529 |  | GXIV 6C | 814154 | CPU does not boot up | Performed PM Checked CPU Documented CPU problem for reporting Generated report | Documented the problem for reporting | CPU for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703008493903234087/5703008493903234087_signature_28.png | Albert Igama | https://www.jotform.com/uploads/macarewecare/222927978372067/5703008493903234087/5703008493903234087_signature_31.png
| 2023-09-12 | 2023-09-12 11:35 | Amma Jadsac District Hospital | Pudtol, Apayao | CAR | Apayao | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25922 | PBSP | rheamarierodriguez.con@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2622 | 2023-09-12 | Genexpert IV | 807609 | Printer not working as intended | Troubleshoot printer | Performed Maintenance Gathered Reports Performed Module Replacement (210046343 - 210102203) Performed Functionality Test | Unit is Operational | Printer for Replacement | Unit is Operational | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5703000107123244605/5703000107123244605_signature_28.png | RHEA MARIE S. RODRIGUEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5703000107123244605/5703000107123244605_signature_31.png
| 2023-09-12 | 2023-09-06 08:19 | Candijay Community Hospital | Candijay, bohol | 7 | Bohol | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 25921 | DOH | angelinaredulla@yahoo.com.ph | 23.5 |  |  | Gx 4 | 838396 | Defective module B4 860995 | New module installed B4 681484, | New module installed B4 681484, clean fan and filter, get copy of IQR, system log and assay stats | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702981973913349669/5702981973913349669_signature_28.png | ANGELINA O. REDULLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5702981973913349669/5702981973913349669_signature_31.png
| 2023-09-12 | 2023-09-06 08:19 | Candijay Community Hospital | Candijay, bohol | 7 | Bohol | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 25920 | DOH | angelinaredulla@yahoo.com.ph | 23.5 |  |  | Gx 4 | 838396 | Defective module B4 860995 | New module installed B4 681484, | New module installed B4 681484, clean fan and filter, get copy of IQR, system log and assay stats | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702981753911256471/5702981753911256471_signature_28.png | ANGELINA O. REDULLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5702981753911256471/5702981753911256471_signature_31.png
| 2023-09-12 | 2023-09-12 10:27 | Bagong Pag-Asa Health Center | Rd 10, Bagong Pag-Asa, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25919 | DOH | erlindalorilla026@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 174 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1522 | 2023-09-06 | GXIV 6C | 831597 | CPU Intermittent problem not booting up Monitor for back up only | Performed PM Archived tests and saved all reports to new CPU Replaced monitor and CPU Generated report | Replaced monitor and CPU MONITOR: CNK83224Q2 -> H22V5SG81A230603426 CPU: 2UA8471VRW -> CPH1119F03692 | Unit it operational All modules working |  | GX Lenovo CPU and Nvision Monitor | DR 10647, 10648 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702960115024638901/5702960115024638901_signature_28.png | Erlinda Lorilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5702960115024638901/5702960115024638901_signature_31.png
| 2023-09-12 | 2023-09-12 09:22 | Lasam District Hospital | IBJ, Lasam, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25918 | DOH | lasamdh.laboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 420 | 2023-08-24 | Genexpert IV | 838530 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement (672411 - 210020302) | Genexpert Module | 10571 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702912691131245012/5702912691131245012_signature_28.png | Verbell Manuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5702912691131245012/5702912691131245012_signature_31.png
| 2023-09-11 | 2023-09-08 09:52 | De Los Santos Medical Center | E. Rodriguez Sr. Boulevard | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25916 | CONSORTIUM | msmababa@dlsmc.ph | 24 |  | 2023-09-06 | GXIV-D-10C | 110020700 | Unit is due for Installation | Installation | Installed GeneXpert System Checked Outlet Voltage, UPS Voltage Output. Generate Reports | All Modules are Operational |  | GeneXpert System | SI 96262 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702212663615069684/5702212663615069684_signature_28.png | Marie Joy Mababa | https://www.jotform.com/uploads/macarewecare/222927978372067/5702212663615069684/5702212663615069684_signature_31.png
| 2023-09-11 | 2023-09-08 15:49 | De La Salle Health and Science Institute | D. Mangubat Ave., Dasmariñas City | 4A | Cavite | Laboratory | NO | MAINTENANCE Software Upgrade | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25917 | CONSORTIUM | acpabruada@dlsmhsi.edu.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1332 |  | GXXVI-D | 842166 | Unit is for GX Software Upgrade | Software Upgrade | Performed Maintenance Install Latest GX Software Generate Reports | 15 Modules are Operational, 1 Module Reported |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702212663617419631/5702212663617419631_signature_28.png | Angelita Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5702212663617419631/5702212663617419631_signature_31.png
| 2023-09-11 | 2023-09-08 15:01 | De La Salle Health and Science Institute | D. Mangubat Ave., Dasmariñas City | 4A | Cavite | Molecular | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25915 | MACARE-HBDC | jvfernandez@my.dlshsi.edu.ph | 24 |  |  | GXXVI-8-10C | 110012951 | Unit is due for Installation | Installation | Install GeneXpert System Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702212663619585376/5702212663619585376_signature_28.png | Juvy V.Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5702212663619585376/5702212663619585376_signature_31.png
| 2023-09-11 | 2023-09-11 13:31 | Alfonso Ponce Enrile Memorial District Hospital | Gonzaga, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25914 | DOH | apemdh_laboratory@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1532 | 2023-08-24 | Genexpert IV | 838474 | Module 1 and 3 error | Peformed Xpertcheck on Module 1 and 3 | Performed Maintenance Gathered Reports Performed Module Replacement (862047 - 719422) Performed Functionality Test | Unit is Operational | Xpertcheck Result = (B1 & B3 = Passed) | Genexpert Module | 10566 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702201780242442118/5702201780242442118_signature_28.png | Regina Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/5702201780242442118/5702201780242442118_signature_31.png
| 2023-09-11 | 2023-09-11 10:55 | Alcala Municipal Health Office | Centro Norte, Alcala, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25913 | DOH | rhualcala@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 323 | 2023-08-24 | Genexpert IV | 838531 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 210006638 - 692359 | Genexpert Module | 10572 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702103577025545664/5702103577025545664_signature_28.png | Clarisa Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/5702103577025545664/5702103577025545664_signature_31.png
| 2023-09-11 | 2023-09-11 09:41 | Cagayan Valley Center for Health Development (DOH R02) | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25912 | PBSP | mctorio.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2618 ASSAY: Xpert SARS-CoV-2, COUNT: 6730 | 2023-09-11 | Genexpert XVI | 842894 | Booting issue | Installed GX Software 6.4 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702058955628122362/5702058955628122362_signature_28.png | Mikeala Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5702058955628122362/5702058955628122362_signature_31.png
| 2023-09-11 | 2023-09-11 09:17 | Cagayan Valley Center for Health Development (DOH R02) | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25911 | PBSP | mctorio.pbsp@gmail.com | 27 |  | 2023-09-11 | Genexpert IV | 834434 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 696059 - 210102238 | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702053175623711617/5702053175623711617_signature_28.png | Mikeala Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5702053175623711617/5702053175623711617_signature_31.png
| 2023-09-11 | 2023-09-08 13:48 | Department of Health Treatment and Rehabilitation Center - Bicutan | Camp Bagong Diwa | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25910 | PBSP | dohtrcbicutan.lab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2122 | 2023-09-05 | GX-IV-D-6C | 807452 | Module C3 for replacement | Replaced module C3 | Performed maintenance Replaced module C3 Generated reports | All modules operational. |  | Old: C3/801912 New: C3/210102242 | G 6265 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702020863614647586/5702020863614647586_signature_28.png | Hellen joy Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5702020863614647586/5702020863614647586_signature_31.png
| 2023-09-11 | 2023-09-08 13:48 | Department of Health Treatment and Rehabilitation Center - Bicutan | Camp Bagong Diwa | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25909 | PBSP | dohtrcbicutan.lab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2122 | 2023-09-05 | GX-IV-D-6C | 807452 | Module C3 for replacement | Replaced module C3 | Performed maintenance Replaced module C3 Generated reports | All modules operational. |  | Old: C3/801912 New: C3/210102242 | G 6265 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5702020773612298998/5702020773612298998_signature_28.png | Hellen joy Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5702020773612298998/5702020773612298998_signature_31.png
| 2023-09-11 | 2023-09-05 13:09 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 25908 | PBSP | lcpmicrolab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4527 ASSAY: Xpert SARS-CoV-2, COUNT: 4316 |  | GXXVI | 801080 | For additional troubleshooting, gathered version of gateway board | Performed troubleshooting Get additional info of gateway board Gathered reports Test run unit | Troubleshooting | B3 reported, 15 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5701999218021216679/5701999218021216679_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5701999218021216679/5701999218021216679_signature_31.png
| 2023-09-11 | 2023-09-08 10:24 | Carigara District Hospital | Carigara Leyte | 8 | Leyte | TB dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25907 | PBSP | profetanaerlarhyn@yahoo.com.ph | 28 |  | 2023-09-08 | GXIV | 807619 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5701974536114705363/5701974536114705363_signature_28.png | Erla Rhyn Profetana | https://www.jotform.com/uploads/macarewecare/222927978372067/5701974536114705363/5701974536114705363_signature_31.png
| 2023-09-10 | 2023-09-10 22:26 | San Jose Municipal Health Office | San Jose | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25906 | DOH | jhuneric2008@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1709 |  | GX4 | 838389 | B1/210015304, B3/861164 for Module Replacement | Replaced Module: B1-210015304>>676102, B3-861164>>739734 Cleaned cartridge bays, plunger rod, icore, fan, fan filter Generate reports | Replaced B1,B3 Modules Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5701655659223005718/5701655659223005718_signature_28.png | Rosmar Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5701655659223005718/5701655659223005718_signature_31.png
| 2023-09-09 | 2023-09-09 10:52 | Lanao del Norte Provincial Hospital | Sagadan, Baroy Lanao Del Norted | 10 | Lanao del Norte | Tb Laboratory | YES | REPAIR CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25905 | PBSP | annamariebroce31@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1779 |  | Gx4 | 807673 | B1, B2, B3 Failed during calibration only B4 passed the calibration | Performed calibration of genexpert machine | Performed calibration, conduct Preventive maintenance, clean Icore, clean valve and plunger | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5700376394111540043/5700376394111540043_signature_28.png | Anna Marie Tubio | https://www.jotform.com/uploads/macarewecare/222927978372067/5700376394111540043/5700376394111540043_signature_31.png
| 2023-09-08 | 2023-09-08 20:58 | San Narciso Rural Health Unit | Maayas-as St., San Naeciso, Quezon Province | 4A | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25904 | PBSP | morena_veluz@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 768 |  | GX IV 6C | 838163 | Module B1 hinge not going down motion | Checked and cleaned cartridge bays, icore slot, plunger rod and filter Troubleshoot Module 1 Checked hinge up and down motions Problem solved Generated report | Troubleshoot module B1 | unit is operational all modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699879800615754512/5699879800615754512_signature_28.png | Ma. Morena V. Dimaculangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5699879800615754512/5699879800615754512_signature_31.png
| 2023-09-08 | 2023-09-08 16:55 | Cordlife Medical Philippines | UP Ayala Technohub, Commonwealth Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION CALIBRATION | CYTIVA | Patrick Los Baños |  |  |  |  | 25903 | COMMERCIAL | patricia.perfecto@cordlife.com | 25 |  | 2023-09-04 | Sepax 2 | 6114.0142 | None, back-up unit for installation | None | Installed back-up unit Performed unit calibration | Unit is operational |  | Sepax 2 S-100 | 10634 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699735749419859331/5699735749419859331_signature_28.png | Patricia Hannah Perfecto | https://www.jotform.com/uploads/macarewecare/222927978372067/5699735749419859331/5699735749419859331_signature_31.png
| 2023-09-08 | 2023-09-08 16:22 | MAYOR HILARION A RAMIRO SR MEDICAL CENTER | Maningcol, Ozamiz-Oroquieta National Highway 7200 Ozamis City Misamis Occidental · | 10 | Lanao del Norte | Tb laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25902 | PBSP | mharsmicrobio@gmail.com | 34 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2421 |  | GX4 | 804765 | B2 SN 662968 encountered multiple error of 1002 and 4009 | Reset module 2 SN-662968 , clean up Icore, cartridge bay, valve and plunger | Reset module 2 SN-662968 , clean up Icore, cartridge bay, valve and plunger | All modules are operational except module 2. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699709054112046209/5699709054112046209_signature_28.png | Felipe Jr. Saren, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5699709054112046209/5699709054112046209_signature_31.png
| 2023-09-08 | 2023-09-08 16:13 | MAYOR HILARION A RAMIRO SR MEDICAL CENTER | Ozamiz-Oroquieta National Highway 7200 Ozamis City Misamis Occidental · | 10 | Misamis Occidental | Covid Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25901 | PBSP | mharsmicrobio@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 217 ASSAY: Xpert SARS-CoV-2, COUNT: 5396 |  | GX4 | 834437 | Module 3 encountered an error of 4019 thermistor and heater problem | Reset module 3 SN- 784125 ,clean up Icore, cartridge bay, valve and plunger | Reset module 3 SN- 784125 ,clean up Icore, cartridge bay, valve and plunger | All modules are operational except module 3 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699705094117425577/5699705094117425577_signature_28.png | Felipe Jr. Saren, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5699705094117425577/5699705094117425577_signature_31.png
| 2023-09-08 | 2023-09-08 16:06 | Kapatagan Rural Health Unit | Kapatagan , Lanao del Norte | 10 | Lanao del Norte | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25900 | DOH | jhayrajaculbe@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 703 |  | GX4 | 838224 | Due for Calibration and preventive maintenance | Performed maintenance and calibration of GeneXpert machine | Performed calibration in all modules , clean up Icore, cartridge bay, valve and plunger | B1-Passed , B2-Retry and Failed Service Required, B3- Passed, B4-Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699699744111930840/5699699744111930840_signature_28.png | Jhayra Jaculbe, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5699699744111930840/5699699744111930840_signature_31.png
| 2023-09-08 | 2023-09-08 16:01 | Lala Rural Health Unit | Lala, Lanipao Lanao del Norte | 10 | Lanao del Norte | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25899 | DOH | villagonzalojoyb@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 900 |  | GX4 | 838229 | Due for Calibration | Performed maintenance and calibration of GeneXpert machine | Performed calibration in all modules , clean up Icore, cartridge bay, valve and plunger | B1 - Passed , B2 - Passed , B3 - Passed, B4 - Retry and Service Required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699695484119223604/5699695484119223604_signature_28.png | Maria Glenda Villagonzalo | https://www.jotform.com/uploads/macarewecare/222927978372067/5699695484119223604/5699695484119223604_signature_31.png
| 2023-09-08 | 2023-09-08 15:52 | Kauswagan Rural Health Unit -Tb Dots | Kauswagan Lanao del Norted | 10 | Lanao del Norte | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25898 | DOH | dextertinoy@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 341 |  | GX4 | 838402 | Due for annual calibration | Performed maintenance and calibration of GeneXpert machine | Performed calibration in all modules , clean up Icore, cartridge bay, valve and plunger | All modules passed during the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699692464119439593/5699692464119439593_signature_28.png | Dexter Tinoy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5699692464119439593/5699692464119439593_signature_31.png
| 2023-09-08 | 2023-09-08 15:29 | Quezon City General Hospital | Semenary Rd., Brgy Bahay Turo, QC | NCR | NCR 2nd - Quezon City | Pathology | NO | ON-SITE ASSISTANCE Unit relocation. | CEPHEID | Harold Toco |  |  |  |  | 25897 | MACARE-HBDC | treatmenthub.qcgh@gmail.com | 25 |  | 2023-09-08 | GXIV 10C | 110012766 | None unit relocation | Performed unit relocation | Perforned unit relocation | All 4 modules operational. | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699677206892136630/5699677206892136630_signature_28.png | Milma Laxina | https://www.jotform.com/uploads/macarewecare/222927978372067/5699677206892136630/5699677206892136630_signature_31.png
| 2023-09-08 | 2023-09-08 14:55 | Makati Medical Center | Amorsolo st., Makati | NCR | NCR 4th - Makati | Stem Cell | YES | CALIBRATION MAINTENANCE | OTHERS | Noel Floresca |  |  |  | Miltenyi | 25896 | COMMERCIAL | arvingregorio71@yahoo.com | 28 |  |  | Gentlemacs octo Dissociator | 1261 | Due for maintenance | Performed Maintenance Cleaned unit Checked and test lcd display Checked program parameters Cheked tube holder wear and tear Cheked heater film Perfomed functionality test Checked temperature and calibration | Preventive maintenance,Calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699656940425603760/5699656940425603760_signature_28.png | Arvin Gregorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5699656940425603760/5699656940425603760_signature_31.png
| 2023-09-08 | 2023-09-08 14:50 | University of the Philippines - National Institutes of Health | UP Manila, Pedri Gil, Ermita | NCR | NCR 1st - Manila | Central Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25895 | COMMERCIAL | vvluzentales@up.edu.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 12 ASSAY: Xpert HBV VL, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Others, COUNT: MTB-XDR - 1 ASSAY: Xpert HCV VL, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 151 | 2023-09-04 | GC16-10C | 110014186 | Unit for PM. For module replacement of module D2. | Performed PM. Performed module replacement on D2. Generated reports. | PM and module replacement. | Unit is operational. All modules operational. |  | GX module D2 (210010323) in replacement to defective D2 (210070086). | DR#10629 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699655292841690157/5699655292841690157_signature_28.png | Vivienne V Luzentales | https://www.jotform.com/uploads/macarewecare/222927978372067/5699655292841690157/5699655292841690157_signature_31.png
| 2023-09-08 | 2023-09-08 14:34 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25894 | PBSP | jvdiaz.pbsp@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 893 ASSAY: Xpert SARS-CoV-2, COUNT: 2426 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 | 2023-09-05 | GXIV | 834431 | Module D4 for replacement | #NAME? | Replaced D4 | Unit is operational. |  | GeneXpert Module SN: 210102195 | G 6263 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699653663252057009/5699653663252057009_signature_28.png | Jim Vianne Diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5699653663252057009/5699653663252057009_signature_31.png
| 2023-09-08 | 2023-09-08 13:47 | Cauayan Municipal Health Office | Municipality of Cauayan | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 25893 | DOH | djeatrmt@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1268 |  | GX4 | 838477 | Module B1 with SN 862002 with hardware failed. | Generate files(IQR & System Log). Condu t PM | Install the new module for B1 with SN 862945. | All modules are now in good running condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699619540654293764/5699619540654293764_signature_28.png | Doriza Joy A. Tizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5699619540654293764/5699619540654293764_signature_31.png
| 2023-09-08 | 2023-09-08 12:56 | Labo District Hospital | Talobatib, Labo | 5 | Camarines Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25892 | DOH | ldh.lab17@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1179 |  | GX4 | 838207 | B2/807309 for Module Replacement | Replaced Module: B2 - 807309>>878474 Cleaned cartridge bays, plunger rod, fan filter, icore, fan Generate reports | Replaced B2 module Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699614821333598505/5699614821333598505_signature_28.png | Janela Joyce Esteves | https://www.jotform.com/uploads/macarewecare/222927978372067/5699614821333598505/5699614821333598505_signature_31.png
| 2023-09-08 | 2023-09-08 12:53 | Padada Rural Health Unit | Rizal Street, Padada Davao del Sur | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com | zhakie.marvz@gmail.com |  | 25891 | DOH | rhupadada8007@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 311 | 2023-08-03 | GX-IV 4d | 838511 | Due for Calibration | Run Calibration on four Modules | Perform Calibration | Modules B2 B3 B4 passed the Calibration Module B1 for service required | Module B1 for under observation | Xpertcheck CE 5 | DR 10443 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699601318024525625/5699601318024525625_signature_28.png | Jimmy Cabungcal | https://www.jotform.com/uploads/macarewecare/222927978372067/5699601318024525625/5699601318024525625_signature_31.png
| 2023-09-08 | 2023-09-08 11:06 | Jose Panganiban Rural Health Unit I | Jose Panganiban | 5 | Camarines Sur | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25890 | PBSP | materesitapenas@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1674 |  | GX4 | 808641 | Date and Time resetting | Replaced Bios/CMOS Battery Generate reports | Replace Bios/CMOS Battery | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699579451333278176/5699579451333278176_signature_28.png | Ma.Teresita G, Peñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5699579451333278176/5699579451333278176_signature_31.png
| 2023-09-08 | 2023-09-08 11:03 | Navotas Medical and Wellness Center | M. Naval St. Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25889 | DOH | alejandrinoknowell@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 35 ASSAY: , COUNT: 1642 |  | GXIV | 831626 | Unit for PM. For troubleshooting of module A4. | Performed PM. Performed Xpertcheck on module A4. Generated reports. | PM and Xpertcheck. | Unit is operational. Module A4 | Used 2 spare Xpertchecks. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699529341813158200/5699529341813158200_signature_28.png | Noel C. Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5699529341813158200/5699529341813158200_signature_31.png
| 2023-09-08 | 2023-09-08 11:13 | Navotas Medical and Wellness Center | M. Naval St. Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25888 | PBSP | alejandrinoknowell@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1642 |  | GXIV | 807592 | GeneXpert was not detected on software. Unit for PM. | Performed troubleshooting. Installed new lan cable. Performed PM. Generated reports. | PM and troubleshooting | Unit is now operational. | Used 1 spare lan cable |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699528431816833982/5699528431816833982_signature_28.png | Noel C. Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5699528431816833982/5699528431816833982_signature_31.png
| 2023-09-08 | 2023-09-08 10:38 | Sta Cruz Rural Health Unit | Poblacion, Sta. Cruz Davao Del Sur | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25887 | DOH | chendalagan72@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF, COUNT: 32 ASSAY: Xpert MTB/RIF Ultra, COUNT: 565 | 2023-08-03 | GX-IV 4D | 838505 | Due for Calibration | Run XPC Calibration of Four Modules | Perform XPC Calibration | Modules B1 and B3 for service Required Modules B2 and B4 passed the Calibration | Modules b1 and b3 are for ubder observation | Xpertcheck CE-5 | DR 10447 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699522497325918147/5699522497325918147_signature_28.png | MAY ANN E CALAÑO | https://www.jotform.com/uploads/macarewecare/222927978372067/5699522497325918147/5699522497325918147_signature_31.png
| 2023-09-08 | 2023-09-08 10:29 | Apalit Rural Health Unit | Apalit Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25886 | DOH | duenasblesilda@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 863 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 |  | GXIV-D | 831900 | A1 for additional troubleshooting | Troubleshooting Preventive maintenance | Performed preventive maintenance Performed Cepheid troubleshooting procedure Archived test results Gathered reports Performed functionality test | Unit us operational except A1 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699495809426939205/5699495809426939205_signature_28.png | Bkesilda G. Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5699495809426939205/5699495809426939205_signature_31.png
| 2023-09-08 | 2023-09-08 09:54 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Tb-lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25885 | PBSP | renlizros11@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 700 ASSAY: Xpert HIV-1 Viral Load, COUNT: 33 | 2023-09-08 | Gx-IV | 815428 | Module 4 for Service Required | Perform Genexpert machine Preventive maintenance | Replace defective module 4 SN:637815 with new module SN:210102259 | 4 operational modules |  | Genexpert module 6 color SN:210102259 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699480301217623894/5699480301217623894_signature_28.png | Renee Liz M. Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5699480301217623894/5699480301217623894_signature_31.png
| 2023-09-08 | 2023-09-07 14:28 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Tb -lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25884 | DOH | jan19.cam@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1362 | 2023-09-07 | Gx-IV | 838504 | Module 1, 2, 3 and 4 for service Required | Perform Genexpert machine Preventive maintenance | Replace defective module 1 SN: 862407 with new module SN: 693761 Replace defective module 4 SN: 861946 with new module SN: 913187 | 3 operational modules |  | Genexpert module 6 color | DR no.: 10532 and 10544 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699464971218286242/5699464971218286242_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5699464971218286242/5699464971218286242_signature_31.png
| 2023-09-08 | 2023-09-08 09:18 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | Relocation | CEPHEID | John Monsour Doctor |  |  |  |  | 25883 | COMMERCIAL | moleculardx@themedicalcity.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 16 ASSAY: Xpert CT/NG, COUNT: 20 ASSAY: Xpert HBV VL, COUNT: 239 ASSAY: Xpert HPV, COUNT: 182 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1770 ASSAY: Others, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 35 ASSAY: Xpert HIV-1 Viral Load, COUNT: 330 ASSAY: Xpert SARS-CoV-2, COUNT: 5 |  | GX-XVI-D-10C | 110015767 | For relocation | Relocate the GeneXpert unit | Transferred the GeneXpert unit Performed maintenance Functionality test | Unit is operational. Module D1/210082300 has error 2014 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699462229017455010/5699462229017455010_signature_28.png | Xchosa Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/5699462229017455010/5699462229017455010_signature_31.png
| 2023-09-08 | 2023-09-07 10:17 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25880 | MACARE-HBDC | fabella.mollabgene@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 37 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 4831 | 2023-09-04 | GXIV-D-10C | 110015482 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed Self-Test Performed XpertCheck Generate Reports | Modules 2 and 4 for Replacement | 2 and 4 Failed in XpertCheck | XpertCheck Kit | 10632 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173611690414/5699399173611690414_signature_28.png | CHRISTIANNE BERNADETTE BAUTISTA | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173611690414/5699399173611690414_signature_31.png
| 2023-09-08 | 2023-09-07 08:34 | Bagong Brgy Health Center | W. Zamora, Pandacan | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25882 | DOH | thea.peneyra@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1095 | 2023-08-25 | GXIV-D | 831694 | Module 1 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 1 Generate Reports | All Modules are Operational |  | GX Module | 10597 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173612949713/5699399173612949713_signature_28.png | Maria Anthea S. Peneyra | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173612949713/5699399173612949713_signature_31.png
| 2023-09-08 | 2023-09-06 08:43 | Dasmariñas City Health Office I | D. Mangubat Ave., Dasmariñas City | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25880 | PBSP | cho1lab@yahoo.com | 24 | ASSAY: Xpert HPV, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2565 ASSAY: Xpert HIV-1 Viral Load, COUNT: 122 |  | GXIV-D | 838893 | Module 1 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 1 Generate Reports | All Modules are Operational | 863786->210102263 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173618969395/5699399173618969395_signature_28.png | Elma R.Belostrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173618969395/5699399173618969395_signature_31.png
| 2023-09-08 | 2023-09-06 11:17 | De La Salle Health and Science Institute | D. Mangubat Ave., Dasmariñas City | 4A | Cavite | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25879 | CONSORTIUM | acpabruada@dlsmhsi.edu.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1323 |  | GXXVI-D | 842166 | GX Software Lagging | Troubleshoot | Clear RAM Usage Purged GX Tests Generate Reports | Unit for Software Upgrade |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173619107695/5699399173619107695_signature_28.png | Angelita C. Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173619107695/5699399173619107695_signature_31.png
| 2023-09-08 | 2023-09-06 09:58 | Dasmariñas Ciry Health Office II | DBB-E, Dasmariñas City | 4A | Cavite | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25881 | DOH | hazelmilan0777@gmail.com | 26 |  |  | GXIV-D | 838535 | CPU Not Booting-Up | Troubleshoot | Reset Connections Cleaned RAM and SSD Checked Voltage Outputs of AVR, Power Cables, UPS, Outlet | CPU is Defective. | CPH0520F08551 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173619869849/5699399173619869849_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5699399173619869849/5699399173619869849_signature_31.png
| 2023-09-07 | 2023-09-07 14:39 | Tinambac RHU | Tinambac | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25878 | PBSP | allancaballero623@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1036 |  | GX4 | 808285 | B1/665879 for module replacement | Replaced module: B1 - 665879>>210100891 Cleaned cartridge bays, plunger rod, fan, icore Generate reports | Replaced module B1 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699061124224790952/5699061124224790952_signature_28.png | Allan F. Caballero | https://www.jotform.com/uploads/macarewecare/222927978372067/5699061124224790952/5699061124224790952_signature_31.png
| 2023-09-07 | 2023-09-06 23:03 | Lagonoy RHU | Lagonoy | 5 | Camarines Sur | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25877 | WHO | barbacenam128@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1527 |  | GX4 | 846355 | For installation of new UPS | Check all port connections going to UPS Generate reports | Installed UPS | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699057784227806053/5699057784227806053_signature_28.png | Marietta Barbacena | https://www.jotform.com/uploads/macarewecare/222927978372067/5699057784227806053/5699057784227806053_signature_31.png
| 2023-09-07 | 2023-09-06 23:03 | Tigaon MHO | Tigaon | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25876 | DOH | jingimperial248@gmail.com | 30 |  |  | GX4 | 838558 | B2/862621 for Module Replacement | Replaced Module: B2 - 862621>>653396 Cleaned cartridge bays, plunger rod, icore, fan and fan filter Generate reports | Replaced B2 Module Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5699052234222166276/5699052234222166276_signature_28.png | Maria Jesusa Imperial | https://www.jotform.com/uploads/macarewecare/222927978372067/5699052234222166276/5699052234222166276_signature_31.png
| 2023-09-07 | 2023-09-07 16:47 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25875 | PBSP | bryanmichaelbalunes@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1349 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3968 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 242 | 2023-09-07 | Genexpert IV | 844111 | Broken Module B3 | Troubleshoot Module B3 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Broken Hinge | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698859928704501342/5698859928704501342_signature_28.png | Bryan Michael Balunes | https://www.jotform.com/uploads/macarewecare/222927978372067/5698859928704501342/5698859928704501342_signature_31.png
| 2023-09-07 | 2023-09-07 15:06 | Northern Mindanao Medical Center | Cagayan de Oro City | 10 | Misamis Oriental | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25874 | PBSP | nmmcgenexpress@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 744 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7455 ASSAY: Xpert SARS-CoV-2, COUNT: 10693 ASSAY: Xpert HIV-1 Viral Load, COUNT: 621 |  | GX-16 | 844118 | D3 was found error | None | Replacing Module B1 with SN: 685159 to SN: 210102207 Replacing Module B4 with SN: 677191 to SN: 210102261 | Only 1 module having H/W Failed | Upcoming module for D3 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698803508213961788/5698803508213961788_signature_28.png | Alexandra O. Lomongo | https://www.jotform.com/uploads/macarewecare/222927978372067/5698803508213961788/5698803508213961788_signature_31.png
| 2023-09-07 | 2023-09-07 14:56 | Muñoz City Health Office | Science City of Muñoz, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25873 | DOH | scomrtdl@gmail.com | 21 |  | 2023-09-07 | Genexpert IV | 831673 | UPS not working | Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698793012837645119/5698793012837645119_signature_28.png | Catherine I Celestino | https://www.jotform.com/uploads/macarewecare/222927978372067/5698793012837645119/5698793012837645119_signature_31.png
| 2023-09-07 | 2023-09-07 14:50 | CHONG HUA HOSPITAL | Fuente Osmeña, Cebu City | 7 | Cebu | TB Laboratory | NO | Pullout | CEPHEID | Jovy Noel |  |  |  |  | 25872 | COMMERCIAL | chhc_laboratory_@chonghua.com.ph | 21 | ASSAY: Xpert MTB/RIF, COUNT: 11 |  | GX-2 | 802314 | None | Pullout of loan Unit | Pullout the following loan unit: 1. GenXpert Machine, SN 802314; 2. Barcode Scanner SN 1214400502325; 3. Acer Laptop with Charger, Model Aspire E14, SN NXGCUSP001624150A07600 | 2 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698788928329372565/5698788928329372565_signature_28.png | Reynaldo Manzanares, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5698788928329372565/5698788928329372565_signature_31.png
| 2023-09-07 | 2023-09-07 14:24 | Ramon magsaysay rural health unit | Ramon magsaysay, Zamboanga del sur | 8 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25871 | DOH | rhuramonmagsaysay7@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 631 | 2023-08-02 | GX 4 | 838443 | Module 2 undetected | Swapped module 2 to different slots of module to isolate the problem | Preventive Maintenence, Xpertcheck calibration and troubleshooting | Module 1 passed, module 3 service required, module 4 passed, module 2 still undetected | None | XPERTCHECK-CE-5 | 10400 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698778166124520274/5698778166124520274_signature_28.png | Gina B. Borinaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5698778166124520274/5698778166124520274_signature_31.png
| 2023-09-07 | 2023-09-07 14:29 | CHONG HUA HOSPITAL | Fuente Osmeña, Cebu City | 7 | Cebu | TB Laboratory | NO | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 25870 | COMMERCIAL | chhc_laboratory_@chonghua.com.ph | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-4 | 110020777 | None | New Unit Installation | #NAME? | All 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698777478324130534/5698777478324130534_signature_28.png | Reynaldo Manzanares | https://www.jotform.com/uploads/macarewecare/222927978372067/5698777478324130534/5698777478324130534_signature_31.png
| 2023-09-07 | 2023-09-07 13:44 | Tagum City Health Office | Poblacion Tagum City | 11 | Davao del Norte | TB-DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25869 | DOH | tagumcho.mabini@gmail.com | 28.5 |  |  | GX-IV 4D | 838553 | Module B1 and B3 encounter temperature error after Calibration | Gather data, Generate IQR, Assay Stat and System Log | Module Troubleshooting | Data will be sent to Cepheid for further Assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698749263325494995/5698749263325494995_signature_28.png | Cristina Tiara F. Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5698749263325494995/5698749263325494995_signature_31.png
| 2023-09-07 | 2023-09-07 13:28 | Bago City Health Office | Bago City | 6 | Negros Occidental | TB DOTS | YES | REPAIR CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25868 | DOH | swabteambago@gmail.com | 4 |  |  | GX4 | 838486 | Module B3 already reported for replacement due to failed in Xpert check last Aug.8 and module B4 withe error 2014 & hardware failed. | Conduct PM & Save generate files(IQR, System log and GXC File. | Run Xpert check calibration for module B4 with SN 862207. Ang installed the new module for B1 with SN 749095. | Module B4 require service due to failed in Xpert check. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698742588911155448/5698742588911155448_signature_28.png | Glenly C. Casona | https://www.jotform.com/uploads/macarewecare/222927978372067/5698742588911155448/5698742588911155448_signature_31.png
| 2023-09-07 | 2023-09-06 16:27 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | YES | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25867 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 752 ASSAY: Others, COUNT: 18 ASSAY: Xpert HBV VL, COUNT: 496 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1206 ASSAY: Xpert MTB/RIF Ultra, COUNT: 965 ASSAY: Xpert SARS-CoV-2, COUNT: 15336 ASSAY: Xpert HCV VL, COUNT: 45 ASSAY: Others, COUNT: 9 | 2023-08-31 | GXXVI | 110005586 | Due for maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed xpertcheck,performed archiving. | Maintenance and xpertcheck | A1, B1,B2,B4,C2, failed xpertcheck 9 module passed | C3 and D3 skip xpertcheck new replaced module | Xpertcheck-ce-5 | 96225 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698732200428573635/5698732200428573635_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5698732200428573635/5698732200428573635_signature_31.png
| 2023-09-07 | 2023-09-06 16:27 | National Kidney and Transplant Institute | East Avenue QC | NCR | NCR 2nd - Quezon City | Molecular | YES | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25866 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 752 ASSAY: Others, COUNT: 18 ASSAY: Xpert HBV VL, COUNT: 496 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1206 ASSAY: Xpert MTB/RIF Ultra, COUNT: 965 ASSAY: Xpert SARS-CoV-2, COUNT: 15336 ASSAY: Xpert HCV VL, COUNT: 45 ASSAY: Others, COUNT: 9 | 2023-08-31 | GXXVI | 110005586 | Due for maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed xpertcheck,performed archiving. | Maintenance and xpertcheck | A1, B1,B2,B4,C2, failed xpertcheck 9 module passed | C3 and D3 skip xpertcheck new replaced module | Xpertcheck-ce-5 | 96225 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698731860421802735/5698731860421802735_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5698731860421802735/5698731860421802735_signature_31.png
| 2023-09-07 | 2023-09-06 09:31 | Northern Benguet District Hospital | Buguias, Benguet | 1 | Benguet | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25865 | PBSP | kym20nbdh@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 8 | 2023-09-06 | GX4 | 814157 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module A3 686904 >>> 210102244 | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210102244 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698725317325727503/5698725317325727503_signature_28.png | Kim Antonio | https://www.jotform.com/uploads/macarewecare/222927978372067/5698725317325727503/5698725317325727503_signature_31.png
| 2023-09-07 | 2023-09-06 09:31 | Northern Benguet District Hospital | Buguias, Benguet | 1 | Benguet | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25864 | PBSP | kym20nbdh@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 8 | 2023-09-06 | GX4 | 814157 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module A3 686904 >>> 210102244 | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210102244 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698724827327337861/5698724827327337861_signature_28.png | Kim Antonio | https://www.jotform.com/uploads/macarewecare/222927978372067/5698724827327337861/5698724827327337861_signature_31.png
| 2023-09-07 | 2023-09-07 12:39 | Davao Regional Medical Center | Apokon Road, Tagum City, Davao del Norte | 11 | Davao del Norte | RHIVDA Section | NO | INSTALLATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25863 | COMMERCIAL | nimezsheilamarie@gmail.com | 25.6 |  |  | GX-IV 10COLORS 4L | 120000127 | For Installation | Setup the GeneXpert System, import assay required and input user/admin details. Refresher tips on enduser on basic troubleshooting Maintenance | Perform GeneXpert Installation and Importing Assay of VL | Successfully Installed and setup | GeneXpert is from Shell Foundation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698720520324803298/5698720520324803298_signature_28.png | Sheila Marie Nimez | https://www.jotform.com/uploads/macarewecare/222927978372067/5698720520324803298/5698720520324803298_signature_31.png
| 2023-09-07 | 2023-09-07 11:40 | Calauag Municipal Hospital | New Calauag Government Complex, Calauag, Quezon Province | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25862 | DOH | jamila17carmel@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 802 |  | GXIV 6C | 831910 | Unit for XpertCheck | Cleaned cartridge bays, icore slot, plunger rod and filter Performed XpertCheck Generated report Perfor | Performed XpertCheck A3: Sevice Required | Unit is operational. A3 failed XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698681237308725050/5698681237308725050_signature_28.png | jamila pedrezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5698681237308725050/5698681237308725050_signature_31.png
| 2023-09-07 | 2023-09-07 09:30 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25861 | PBSP | benguetphortdl@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 949 | 2023-09-07 | GX4 | 807621 | Due fo Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 686447 >>>210101785 - Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 10101785 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667967327430166/5698667967327430166_signature_28.png | Tillber Baniaga Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667967327430166/5698667967327430166_signature_31.png
| 2023-09-07 | 2023-09-07 08:51 | Baguio Health Services Office | Baguio City | CAR | Benguet | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25860 | DOH | sherryannemarrero@gmail.com | 21 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 834 | 2023-08-24 | GX4 | 838281 | For Module Replacement For Troubleshooting of module B2 | #NAME? | #NAME? | Unit is operational | Module B2 Error after ReXpertCheck on 8-9-23, For reporting | GeneXpert Module 678653 | DR 10570 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667837323233830/5698667837323233830_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667837323233830/5698667837323233830_signature_31.png
| 2023-09-07 | 2023-09-06 14:21 | Baguio General Hospital and Medical Center | Baguio City | 1 | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25859 | PBSP | keithtagudar@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 354 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1699 |  | GX4 | 805212 | For Preventive Maintenance | #NAME? | Performed Preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667837324211284/5698667837324211284_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667837324211284/5698667837324211284_signature_31.png
| 2023-09-07 | 2023-09-06 14:17 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25858 | PBSP | keithtagudar@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 685 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1993 |  | GX4 | 834685 | For Preventive Maintenance | #NAME? | Performed Preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667717327498854/5698667717327498854_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667717327498854/5698667717327498854_signature_31.png
| 2023-09-07 | 2023-09-06 13:57 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25857 | PBSP | keithtagudar@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 63 ASSAY: Xpert MTB/RIF Ultra, COUNT: 458 ASSAY: Others, COUNT: 13 |  | GX4 10C | 120000171 | For Troubleshooting of module B4 | #NAME? | Performed Troubleshooting | Unit is operational | Module B4 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667717325006629/5698667717325006629_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5698667717325006629/5698667717325006629_signature_31.png
| 2023-09-07 | 2023-09-07 10:17 | Cainta City Health Office | Karangalan Dr. | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 25856 | DOH | ballesterrichard18@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 925 ASSAY: Xpert HIV-1 Viral Load, COUNT: 84 |  | GX-IV-D-6C | 831623 | Module A2 has error 4009 | Run XpertCheck on module A2 | Performed maintenance Rin XpertCheck on module A2 Generated reports | Unit is operational. Module A2/754519 passed in XpertCheck but under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698648944328466405/5698648944328466405_signature_28.png | Richard A. Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5698648944328466405/5698648944328466405_signature_31.png
| 2023-09-07 | 2023-09-07 10:55 | Cong. Simeon Toribio memorial hospital | Carmen, Bohol | 7 | Bohol | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 25855 | DOH | cristalroy@yahoo.com | 28 |  |  | Gx 4 | 838394 | Defective module 861074 | New module installed sn 786289 | New module installed sn 786289 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698648834165378611/5698648834165378611_signature_28.png | Roy Autentico Cristal | https://www.jotform.com/uploads/macarewecare/222927978372067/5698648834165378611/5698648834165378611_signature_31.png
| 2023-09-07 | 2023-09-07 10:16 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25854 | DOH | abigailgsula@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1387 |  | GXIV-D | 831912 | A4 HW Failed | Xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is ooerational except A3 and A4 | A4 service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698625826124219907/5698625826124219907_signature_28.png | Michaella David | https://www.jotform.com/uploads/macarewecare/222927978372067/5698625826124219907/5698625826124219907_signature_31.png
| 2023-09-07 | 2023-09-07 09:55 | NAZARETH HEALTH CENTER | NAZARETH CAGAYAN DE ORO, MISAMIS ORIENTAL | 10 | Misamis Oriental | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25853 | DOH | nanettebtolentino@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1219 | 2023-07-31 | GX-IV | 838269 | FOR MAINTENANCE AND CALIBRATION | COLLECT IQR, SYSTEM LOG, AND CALIBRATION DATA | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK KIT | 10368 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698615270223219361/5698615270223219361_signature_28.png | Nanette Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/5698615270223219361/5698615270223219361_signature_31.png
| 2023-09-07 | 2023-09-07 09:57 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | Tb-dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25852 | DOH | burauenlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 393 | 2023-09-07 | GXIV | 838551 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All Module passed | Xpertcheck kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698615036425238060/5698615036425238060_signature_28.png | Mary Ann Sodario | https://www.jotform.com/uploads/macarewecare/222927978372067/5698615036425238060/5698615036425238060_signature_31.png
| 2023-09-07 | 2023-09-06 10:25 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25851 | DOH | claritomirasol@gmail.com | 28 |  | 2023-09-06 | GXIV | 838548 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed for Calibration | Xpertcheck kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698592976423176140/5698592976423176140_signature_28.png | Mirasol Clarito | https://www.jotform.com/uploads/macarewecare/222927978372067/5698592976423176140/5698592976423176140_signature_31.png
| 2023-09-07 | 2023-09-07 08:29 | Abucay Rural Health Unit | Abucay Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25850 | DOH | parreraledelina22176@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 |  | GXIV-D | 831878 | Barcod scanner not working | Troubleshooting | Performe preventive maintenance Troubleshoot barcode scanner Gathered reports Performed functionality test | Unit is operational except barcode scanner | Barcode scanner for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698563966125992871/5698563966125992871_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5698563966125992871/5698563966125992871_signature_31.png
| 2023-09-07 | 2023-09-07 07:48 | Bataan Genaral Hospital and Medical Center | Balanga Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25849 | PBSP | honeyeichel@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 84 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2481 ASSAY: Xpert HCV VL, COUNT: 80 ASSAY: Xpert HIV-1 Viral Load, COUNT: 224 |  | GXIV-D | 805124 | B3 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698536006122507352/5698536006122507352_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5698536006122507352/5698536006122507352_signature_31.png
| 2023-09-06 | 2023-09-06 17:25 | RURAL HEALTH UNIT MANOLO FORTICH | MANOLO FORTICH, BUKIDNON | 10 | Bukidnon | LABORATORY | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  |  |  |  | 25848 | PBSP | mfrhutbdots@gmail.com | 815286 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 |  | GX-IV | 815286 | FOR MODULE B4 CALIBRATION | COLLEC IQR, SYSTEM LOG, AND CALIBRATION DATA | PERFORM MODULE B4 CALIBRATION : PASSED | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698019286221019992/5698019286221019992_signature_28.png | Elaine Chung | https://www.jotform.com/uploads/macarewecare/222927978372067/5698019286221019992/5698019286221019992_signature_31.png
| 2023-09-06 | 2023-09-06 17:12 | VALENCIA CITY HEALTH OFFICE | VALENCIA CITY, BUKIDNON | 10 | Bukidnon | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25847 | DOH | caluyajopay@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1836 | 2023-07-31 | GX-IV | 838271 | FOR MAINTENANCE AND CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK KIT | 10371 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5698011226222092442/5698011226222092442_signature_28.png | JOEFELYN R. CALUYA | https://www.jotform.com/uploads/macarewecare/222927978372067/5698011226222092442/5698011226222092442_signature_31.png
| 2023-09-06 | 2023-09-06 16:45 | Tomas Cabili Health Center | Tomas Cabili, Iligan City, Lanao del Norte | 10 | Lanao del Norte | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25846 | DOH | mariacielovenezuela@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 969 |  | GX4 | 838294 | Due for calibration | Performed maintenance and calibration of GeneXpert machine | Performed calibration in all modules , clean up Icore, cartridge bay, valve and plunger | All modules passed during the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697993983711748414/5697993983711748414_signature_28.png | Maria Cielo C. Venezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5697993983711748414/5697993983711748414_signature_31.png
| 2023-09-06 | 2023-09-06 16:37 | Hinaplanon RTDL | Upper Hinaplanon, Iligan City, Lanao Del Norte | 10 | Lanao del Norte | Tb laboratory | NO | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25845 | DOH | Villacortaamy50@gmail.com | 28 |  |  | GX4 | 838227 | Module 2 SN# 807236 and Module 4 SN# 807440 | Performed maintenance and calibration of | Performed calibration in all modules , clean up Icore, cartridge bay, valve and plunger | Only Module 1 and Module 3 passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697990663713720434/5697990663713720434_signature_28.png | Amy S. Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/5697990663713720434/5697990663713720434_signature_31.png
| 2023-09-06 | 2023-09-06 15:16 | Calauan Rural Health Unit | Calauan | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25844 | DOH | jphilipsegui@yahoo.com | 25 |  | 2023-08-24 | GXIV | 821164 | Unit for PM. For module replacement | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module 671935 in replacement to defective 783229. | DR#10564 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697953843819632950/5697953843819632950_signature_28.png | John philip velecina | https://www.jotform.com/uploads/macarewecare/222927978372067/5697953843819632950/5697953843819632950_signature_31.png
| 2023-09-06 | 2023-09-06 15:30 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 25843 | PBSP | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2600 ASSAY: Xpert HIV-1 Viral Load, COUNT: 144 |  | GXIV | 805062 | For additional troubleshooting | #NAME? | Performed recommended troubleshooting steps from Cepheid | Unit is operational. A3 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697949402525865496/5697949402525865496_signature_28.png | Revena Sibayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5697949402525865496/5697949402525865496_signature_31.png
| 2023-09-06 | 2023-09-06 15:27 | Tagkawayan Rural Health Unit | Tagkawayan, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25842 | DOH | rvteopy@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 578 | 2023-08-24 | GXIV 6C | 831916 | Modules B2 and B3 for replacement | Performed PM Replaced modules B2 and B3 Generated report | Replaced 2 modules B2: 888663 -> 863440 B3: 758385 -> 678522 | Unit is operational All modules working |  | 2 GX Modules 6C | DR 10553, 10556 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697947980129428785/5697947980129428785_signature_28.png | Rodelo Teopy | https://www.jotform.com/uploads/macarewecare/222927978372067/5697947980129428785/5697947980129428785_signature_31.png
| 2023-09-06 | 2023-09-06 13:52 | Panabo City Health Office | Poblacion Panabo City | 11 | Davao del Norte | TB-DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25841 | DOH | nadineeust67@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2600 | 2023-08-23 | GX-IV 4D | 838295 | Due for Preventive Maintenance, Calibration Error at Module B3 | Replaced the Defective Module SN 808195 with new Module SN: 719047 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration | Perform Module Replacement, Preventive Maintenance, andXpertCheck Calibration | Module B4 for service required Module B4 for under observation |  | XPERTCHECK CE 5, GeneXpert 900-0491rs | DR 10444 and DR 10526 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697924060327751845/5697924060327751845_signature_28.png | Louella Jane G. Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5697924060327751845/5697924060327751845_signature_31.png
| 2023-09-06 | 2023-09-06 14:18 | Iba Rural Health Unit | Iba Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25840 | PBSP | garadoshanna@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 150 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1880 |  | GXIV-D | 805070 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697919286122642094/5697919286122642094_signature_28.png | Haana C. Garados | https://www.jotform.com/uploads/macarewecare/222927978372067/5697919286122642094/5697919286122642094_signature_31.png
| 2023-09-06 | 2023-09-06 14:19 | Nagpayong Super Health Center | Centennial 2 | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25839 | DOH | michelindg23@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 630 | 2023-08-24 | GX-IV-D-6C | 831905 | Module B2 for replacement | Replaced module B2 | Performed maintenance Replaced defective module Generated reports | All modules operational |  | Old: B2/758210 New: B2/724966 | DR# 10555 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697907137112805805/5697907137112805805_signature_28.png | Michelle De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5697907137112805805/5697907137112805805_signature_31.png
| 2023-09-06 | 2023-09-06 13:50 | Western Visayas Sanitarium | Santa Barbara, Iloilo | 6 | Iloilo | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25838 | DOH | migzko@yahoo.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 538 |  | GX4 | 838393 | Module B4 with error 4009 and hardware failed. | Conduct PM and generate files(IQR, System Log and GXC File) | Run Xpert check for module B4 with SN 692166. | Module B4 failed in Xpert check and module B1, B2 and B3 are in good condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697891201816897265/5697891201816897265_signature_28.png | Ruth Dequilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5697891201816897265/5697891201816897265_signature_31.png
| 2023-09-06 | 2023-09-06 13:53 | San Luis Rural Health Unit | Santa Cruz Poblacion, San Luis, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25837 | DOH | mohan_santos@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 285 | 2023-08-24 | Genexpert IV | 831714 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 784272 - 909584 | Genexpert Module | 10552 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697890275025159533/5697890275025159533_signature_28.png | Mohan S. Pangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5697890275025159533/5697890275025159533_signature_31.png
| 2023-09-06 | 2023-09-06 13:40 | A. H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25836 | PBSP | dinajaen1990@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3196 | 2023-09-05 | GX XVI | 839292 | Modules B1, B2, B3, C1, C4 for replacement | #NAME? | Replaced modules | Unit is operational. All modules functional. |  | GeneXpert Module SN: 210102249, 210102246, 210102254, 210102087, 210102214 | G 6262 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697885162525982086/5697885162525982086_signature_28.png | Dina Jaen | https://www.jotform.com/uploads/macarewecare/222927978372067/5697885162525982086/5697885162525982086_signature_31.png
| 2023-09-06 | 2023-09-06 13:15 | Cebu City Medical Center | Cebu City, Cebu | 7 | Cebu | Micriobiology | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25835 | DOH | ccmclaboratory01@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1579 | 2023-08-23 | GXIV | 838561 | Module A1 defective. Module A4 frequent h/w fail. | -Replace defective module A1/862882 with new module SN: 879218 -Run XpertCheck on module A4 | #NAME? | All modules are operational. | Module A4 passed XPC, for observation. | GeneXpert Module 6C | 10527 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697882663229848620/5697882663229848620_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5697882663229848620/5697882663229848620_signature_31.png
| 2023-09-06 | 2023-08-24 14:49 | Taytay Rural Health Unit 5B | Ynares Ave, Sta. Ana, Taytay | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25834 | PBSP | zenresurreccion@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1420 | 2023-08-24 | GX-IV-D | 838136 | Module B2/806712 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | Unit is operational. |  | Old: B2/806712 New: B2/210100668 | G 6111 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697873674018976503/5697873674018976503_signature_28.png | Zenilynn r. Bobiles | https://www.jotform.com/uploads/macarewecare/222927978372067/5697873674018976503/5697873674018976503_signature_31.png
| 2023-09-06 | 2023-08-24 14:49 | Taytay Rural Health Unit 5B | Ynares Ave, Sta. Ana, Taytay | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25833 | PBSP | zenresurreccion@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1420 | 2023-08-24 | GX-IV-D | 838136 | Module B2/806712 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | Unit is operational. |  | Old: B2/806712 New: B2/210100668 | G 6111 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697872224019616943/5697872224019616943_signature_28.png | Zenilynn r. Bobiles | https://www.jotform.com/uploads/macarewecare/222927978372067/5697872224019616943/5697872224019616943_signature_31.png
| 2023-09-06 | 2023-09-06 12:29 | President Ramon Magsaysay Memorial Hospital | Iba, Zambales | 3 | Zambales | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25832 | PBSP | prmmhmoleclab@gmail.com | 30 | ASSAY: Xpert SARS-CoV-2, COUNT: 3507 |  | GXIV-D | 834622 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is fully operational, 3 modules passed B3 failed calibration for reporting | B3 service required failed calibration, used 6 cartridges |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697853006123163044/5697853006123163044_signature_28.png | Angelina Cuaton | https://www.jotform.com/uploads/macarewecare/222927978372067/5697853006123163044/5697853006123163044_signature_31.png
| 2023-09-06 | 2023-09-06 11:41 | Kalilangan Rural Health Unit | Poblacion, Kalilangan | 10 | Bukidnon | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25831 | DOH | myrnasuyao17@gmail.com | 23.27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 464 | 2023-08-23 | GX-4 | 838571 | Module B3 found Service Required through XPC Calibration | None | Replacing Module B1 with SN: 633020 to SN: 647747 | 3 modules working, 1 module defective (for replacement) |  | GeneXpert Module 6-color | 10536 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697816244112087507/5697816244112087507_signature_28.png | Aida M. Macarimbang | https://www.jotform.com/uploads/macarewecare/222927978372067/5697816244112087507/5697816244112087507_signature_31.png
| 2023-09-06 | 2023-09-06 11:21 | Rosario Rural Health Unit | Poblacion, Rosario | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25830 | DOH | jcinciong_james@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 715 | 2023-08-24 | GXIV | 821421 | Unit for PM. For module replacement of module A4. | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module A4 (674210) in replacement to defective A4 (761196). | DR#10561 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697810543599775041/5697810543599775041_signature_28.png | James C. Inciong | https://www.jotform.com/uploads/macarewecare/222927978372067/5697810543599775041/5697810543599775041_signature_31.png
| 2023-09-06 | 2023-09-06 10:52 | Norala District Hospital | Poblacion Norala, South Cotabato | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25829 | DOH | rezielrds0709@gmail.com | 21.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 975 | 2023-09-06 | Gx-IV | 838543 | Module 2 H/W failed Module 2 service required Module 3 service required | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine annual Calibration/Xpertcheck | 3 operational modules | Module 1 and 4 passed the annual Xpertcheck/Calibration | Xpertcheck CE-5 cartridge/kit | DR no.: 10411 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697804733005352906/5697804733005352906_signature_28.png | Reziel D. Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/5697804733005352906/5697804733005352906_signature_31.png
| 2023-09-06 | 2023-09-06 10:24 | Santolan Super Health Center | Evangelista Ave. | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25828 | PBSP | samandlet17@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4057 | 2023-09-05 | GX-IV-D-6C | 838898 | Module B4 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | All modules operational |  | Old: B4/619237 New: B4/210102253 | G 6266 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697777084016899014/5697777084016899014_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5697777084016899014/5697777084016899014_signature_31.png
| 2023-09-06 | 2023-09-06 09:37 | Guimbal Rural Health Unit | Guimbal, Iloilo | 6 | Iloilo | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 25827 | PBSP | percival_girado@yahoo.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1803 ASSAY: Xpert HIV-1 Viral Load, COUNT: 371 |  | GX4 | 807587 | Module B4 with SN 656706 that failed in Xpert check last July 2023 | Conduct PM and generate IQR files before and after replacement. | Replacement of new module for B4 with SN 210100877. | Module B3 has an error of harware failed error 4012. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697771769316128098/5697771769316128098_signature_28.png | Percival Girado | https://www.jotform.com/uploads/macarewecare/222927978372067/5697771769316128098/5697771769316128098_signature_31.png
| 2023-09-06 | 2023-09-06 08:27 | Lubao Rural Health Unit | Lubao Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25826 | PBSP | nathaliecoleencovar@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 898 |  | GXIV-D | 812834 | Stuck cartridge on module 1 | Repair Preventive maintenance | Performed preventive maintenance Manually remove cartridge Gathered reports Performed functionality test | Unit is operational | Broken cartridge valve |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697696193113435960/5697696193113435960_signature_28.png | Nathalie Covar | https://www.jotform.com/uploads/macarewecare/222927978372067/5697696193113435960/5697696193113435960_signature_31.png
| 2023-09-05 | 2023-09-05 18:01 | Calamba Medical Center | Crossing, Real Calamba | 4A | Laguna | Molecular Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25825 | COMMERCIAL | cmc.rtpcr@gmail.com | 25 |  |  | GXII | 812916 | H/W Failure on module B2. For replacement of AVR and UPS. | Generating of iqr and system log for reporting. | Troubleshooting. | Unit is operational. Module B2 for replacement. For replacement of UPS (1kVA) and AVR (1kVA). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697176726282289278/5697176726282289278_signature_28.png | Andea Gail | https://www.jotform.com/uploads/macarewecare/222927978372067/5697176726282289278/5697176726282289278_signature_31.png
| 2023-09-05 | 2023-09-05 16:46 | Mindanao central sanitarium | Pasabolong Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25824 | DOH | ybioamie@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1933 | 2023-08-23 | GX IV | 821214 | Module 2 error 2014 and 1001 | Replaced module 620758 | Module Replacement and Preventive Maintenence | All modules are available | None | 900-0491R | 10534 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697130048525512714/5697130048525512714_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5697130048525512714/5697130048525512714_signature_31.png
| 2023-09-05 | 2023-09-05 16:20 | Batasan Super Health Center | Batasan Rd, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25823 | PBSP | ashwinkevin@gmail.com | 34 | ASSAY: Xpert HPV, COUNT: 868 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2277 | 2023-09-05 | GXIV 6C | 805254 | Module D4 for replacement | Performed PM Replaced Module D4 Generated report | Replaced module D4 685505 -> 210102209 | Unit is operational All modules working |  | GX module 6C | G 6257 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697117729814524559/5697117729814524559_signature_28.png | Kevin Orlando D. Cabuso | https://www.jotform.com/uploads/macarewecare/222927978372067/5697117729814524559/5697117729814524559_signature_31.png
| 2023-09-05 | 2023-09-05 15:42 | Bale Angeleño | C. Surla St. Balibago Angeles City, Pampanga | 3 | Pampanga | Laboratory | NO | INSTALLATION | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25822 | MACARE-HBDC | cistron0511@gmail.com | 27 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 |  | GXIV-D | 120000128 | Installation | Installation and setup | Performed installation and setup of Genexpert machine Gathered reports Performed functionality test | Unit us operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697102437573937559/5697102437573937559_signature_28.png | Paul Brian Simbulan | https://www.jotform.com/uploads/macarewecare/222927978372067/5697102437573937559/5697102437573937559_signature_31.png
| 2023-09-05 | 2023-09-05 13:52 | Meycauayan City Health Center VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 25821 | DOH | ronamarieiris_25@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1379 ASSAY: Xpert HIV-1 Viral Load, COUNT: 85 |  | GXIV | 831904 | UPS back up mode for check up | - checked UPS' parameters (Line Voltage: 232V, Back-up Mode Voltage: 237V) - checked and cleaned cartridge bay, plunger rod, filter - performed 5 min. battery back up mode (success) - generated reports | Checked UPS' back up mode | Unit is operational. Under observation. Battery for request. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697047923319773969/5697047923319773969_signature_28.png | Rona Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5697047923319773969/5697047923319773969_signature_31.png
| 2023-09-05 | 2023-09-05 14:18 | Marikina City Health Office | Shoe Avenue | NCR | NCR 2nd - Marikina | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  |  |  |  | 25820 | MACARE-HBDC | amymitra22@gmail.com | 25 |  |  | GX-IV-D-10C | 120000613 | For installation | Installed the GeneXpert System | Installed the GeneXpert System: Printer UPS Desktop Barcode scanner GeneXpert unit Installed MTB assay Functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697041338015851218/5697041338015851218_signature_28.png | Amelia C. Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5697041338015851218/5697041338015851218_signature_31.png
| 2023-09-05 | 2023-09-05 14:09 | Palayan City Health Office | Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25819 | DOH | riazarizzamae@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1022 | 2023-09-05 | Genexpert IV | 831637 | UPS no backup | Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697037485027523899/5697037485027523899_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5697037485027523899/5697037485027523899_signature_31.png
| 2023-09-05 | 2023-09-01 10:54 | Marilao Rural Health Unit | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25818 | DOH | jmjenilo27@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1547 ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 | 2023-08-25 | GXIV | 831897 | Module A3 for replacement | #NAME? | Replaced A3 | Unit is operational. |  | GeneXpert Module SN: 632641 | 10578 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697027743316831922/5697027743316831922_signature_28.png | Merlie De Los Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5697027743316831922/5697027743316831922_signature_31.png
| 2023-09-05 | 2023-09-01 10:54 | Marilao Rural Health Unit | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25817 | DOH | jmjenilo27@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1547 ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 | 2023-08-25 | GXIV | 831897 | Module A3 for replacement | #NAME? | Replaced A3 | Unit is operational. |  | GeneXpert Module SN: 632641 | 10578 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697027443317974157/5697027443317974157_signature_28.png | Merlie De Los Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5697027443317974157/5697027443317974157_signature_31.png
| 2023-09-05 | 2023-09-05 13:34 | Angeles Rural Health Unit 3 | Fajardo St. Lourdes sur east, Angeles City | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25816 | PBSP | femiepangilan@yahoo.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2627 ASSAY: Xpert HIV-1 Viral Load, COUNT: 715 |  | GXIV-D | 807569 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is fully operational except B2 for reporting | B2 failed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697024407222791281/5697024407222791281_signature_28.png | Femie Pangilinan | https://www.jotform.com/uploads/macarewecare/222927978372067/5697024407222791281/5697024407222791281_signature_31.png
| 2023-09-05 | 2023-09-05 12:26 | Punturin 3S Health Center | Punturin, Valenzuela | 3 | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 25815 | DOH | lukegonzaga09@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1014 |  | GXIV | 831914 | Module B3 for re XpertCheck | - checked and cleaned cartridge bay, plunger rod, filter, icore slot - performed XpertCheck on B3 - generated reports - | Performed XpertCheck on B3 | Unit is operational. B3 passed XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697023473318455774/5697023473318455774_signature_28.png | Judelyn Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5697023473318455774/5697023473318455774_signature_31.png
| 2023-09-05 | 2023-09-05 13:21 | Domingo Santiago Health Center | 844 Domingo Santiago St., Brgy. 571, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25814 | DOH | fei_paas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1009 |  | GXIV 6C | 831527 | Module 2 door broken | Performed PM Replaced door of Module 2 Generated report | Replaced module 2 door | Unit is operational. 3 modules working, 1 module hardware failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5697009469819955019/5697009469819955019_signature_28.png | Flory Fe Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5697009469819955019/5697009469819955019_signature_31.png
| 2023-09-05 | 2023-09-05 12:26 | Philippine Red Cross -General Santos City & Sarangani Province chapter | General Santos City, South Cotabato | 12 | South Cotabato | Blood Bank | YES | MCS+ | HAEMONETICS | Jesson Tupas |  |  |  |  | 25813 | COMMERCIAL | prc_gensanbloodservices@yahoo.com.ph | 23 |  |  | MCS+ | 03J030 | For pull out | Pull out MCS+ and UPS shipped back to Macare Q.C | Pull out MCS+ and UPS shipped back to Macare Q.C | Successfuly pull out |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696983394526211729/5696983394526211729_signature_28.png | JESTRIL JINGCO | https://www.jotform.com/uploads/macarewecare/222927978372067/5696983394526211729/5696983394526211729_signature_31.png
| 2023-09-05 | 2023-09-05 11:25 | Garcia Memorial Provincial Hospital | Talibon, Bohol | 7 | Bohol | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 25812 | DOH | ybanezrosalina19@gmail.com | 23 |  |  | Gx 4 | 838203 | Defective module B1 sn 730958 | New module installed B1 884653 | New module installed B1 884653 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696941704379897645/5696941704379897645_signature_28.png | Rosalina A Ybañez | https://www.jotform.com/uploads/macarewecare/222927978372067/5696941704379897645/5696941704379897645_signature_31.png
| 2023-09-05 | 2023-09-05 10:56 | General Tinio Rural Health Unit | Poblacion Central, General Tinio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25811 | DOH | jadelyn_pajarillaga@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 801 | 2023-08-24 | Genexpert IV | 831695 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 757236 - 694981 | Genexpert Module | 10558 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696920406955806919/5696920406955806919_signature_28.png | Jadelyn B. Pajarillaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5696920406955806919/5696920406955806919_signature_31.png
| 2023-09-05 | 2023-09-05 10:03 | Amang Rodriguez Memorial Medical Center | Sumulong Highway | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25810 | PBSP | pathologymole.armmc@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 3605 ASSAY: Others, COUNT: 10 | 2023-09-05 | GX-XVI-D | 844113 | Modules A1, B2, B4 and C3 for replacement | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. |  | Old: A1/891760, B2/628603, B4/804084, C2/891680 New: A1/210101827, B2/210101813, B4/210102220, C2/210102270 | G 6256 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696909815826080458/5696909815826080458_signature_28.png | Racquel Rhys C. Verzosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5696909815826080458/5696909815826080458_signature_31.png
| 2023-09-05 | 2023-09-05 09:45 | Tagoloan Rural Health Unit | Balagtas St. Poblacion, Tagoloan | 10 | Misamis Oriental | RTDL | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25809 | DOH | macalisangglezeal@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1233 |  | GX-4 | 838574 | None | None | Replacing Module B4 with SN: 863010 to SN: 802736 Preventive Maintenance XpertCheck Calibration | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696888942929137633/5696888942929137633_signature_28.png | Glezeal G. Macalisang | https://www.jotform.com/uploads/macarewecare/222927978372067/5696888942929137633/5696888942929137633_signature_31.png
| 2023-09-05 | 2023-09-05 09:24 | CAGAYAN DE ORO CITY HEALTH OFFICE | CAGAYAN DE ORO, MISAMIS ORIENTAL | 10 | Misamis Oriental | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25808 | DOH | chocdotbdots@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3152 | 2023-07-31 | GX-IV | 838435 | FOR MAINTENANCE AND CALIBRATION | #NAME? | PERFORM XPERTCHECK CALIBRATION: -MODULE C1:PASSED -MODULE C2:PASSED -MODULE C3:PASSED | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK KIT | 10359 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696864406225982684/5696864406225982684_signature_28.png | CHARISSA B. GALACIO,RMT,MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/5696864406225982684/5696864406225982684_signature_31.png
| 2023-09-05 | 2023-09-05 09:15 | CAGAYAN DE ORO CITY HEALTH OFFICE | CAGAYAN DE ORO, MISAMIS ORIENTAL | 10 | Misamis Oriental | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25807 | DOH | chocdotbdots@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3152 | 2023-07-31 | GX-IV | 838441 | FOR MAINTENANCE AND CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | MODULE B2 AND MODULE B3 DISABLED | XPERTCHECK KIT | 10360 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696861166226892603/5696861166226892603_signature_28.png | CHARISSA B. GALACIO,RMT,MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/5696861166226892603/5696861166226892603_signature_31.png
| 2023-09-04 | 2023-09-04 15:28 | Iriga City Health Office | Iriga | 5 | Camarines Sur | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25806 | DOH | mongealyzza@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1850 | 2023-08-25 | GX4 | 838391 | B2/628901, B3/654462, B4/695599 for Module Replacement | Replaced Module: B2 - 698901>>210044677 B3 - 654462>>669045 B4 - 695599>>210040454 Performed Re-XpertCheck for B1/780882 - service required Generate Reports | Replaced B2,B3,B4 reported modules Performed Re-XpertCheck for B1/780882 | Unit is Operational | B1/780882 - Service Required | Module: 210040454 669045 210044677 | DR 10580, 10581, 10598 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696428685214178609/5696428685214178609_signature_28.png | Alyzza Rosabelle M. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5696428685214178609/5696428685214178609_signature_31.png
| 2023-09-04 | 2023-09-04 16:31 | Philippine Red Cross | Port,Area Manila | NCR | NCR 1st - Manila | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 25805 | MACARE-HBDC | dolly.banares@redcross.org.ph | 28 |  | 2023-09-04 | MCS+ | 02A087 | Encounter safety system error during start of procedure | Performed troubleshooting Replaced safety board but error still persist. Swap CPU board, drive and motor board erros still persist. Performed Calibration after performed the calibration error still persist. Checked all connection. Checked program parameters. Cannot able to do functionality test because of error appear on the display safety system error. | Troubleshooting | Unoperational for additional troubleshooting and repair | With Ronnel Belizar | 86158-01 | 10628 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696264252631134755/5696264252631134755_signature_28.png | Dolly banares | https://www.jotform.com/uploads/macarewecare/222927978372067/5696264252631134755/5696264252631134755_signature_31.png
| 2023-09-04 | 2023-09-04 16:09 | Cebu City Jail Male Dorm | Kalunasan, Cebu City | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 25804 | PBSP | janalbrienevendiola@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1917 |  | GX4 | 805214 | B1-792113, B2-862144, B-668142 defective failed during calibration | Replaced and conduct corrective maintenance in GeneXpert machine | Replaced new module B1 210101936,B2 210102197,B4 210102208 | All module are operational. | Replaced accompanied by : Engr. Jovy D. Noel |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696246633918115511/5696246633918115511_signature_28.png | Jan Albriene Vendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5696246633918115511/5696246633918115511_signature_31.png
| 2023-09-04 | 2023-09-01 12:35 | Dasmariñas City Health Office II | DBB-E, Dasmariñas City | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25800 | DOH | hazelmilan0777@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2291 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 | 2023-08-25 | GXIV-D | 838535 | Modules 1 and 4 with Temperature Errors | Module Replacement | Performed Maintenance Replaced Modules 1 and 4 | All Modules are Operational |  | GX Module 6-Colors | DR # 10594 & 10588 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953611071678/5696241953611071678_signature_28.png | Hazel Tabucol | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953611071678/5696241953611071678_signature_31.png
| 2023-09-04 | 2023-09-01 09:29 | General Trias City Health Office | Hospital Rd., General Trias | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25803 | PBSP | chogentri.lab@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2285 ASSAY: Xpert HIV-1 Viral Load, COUNT: 214 | 2023-08-29 | GXIV-D | 804691 | Module 1 Failed in XpertCheck | Module Replacement | Performed Maintenance Replaced Module 1 Generate Reports | All Modules are Operational | 746455 -> 210101925 | GeneXpert Module 6-C | G 6243 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953614285540/5696241953614285540_signature_28.png | Mack Alfred L. Chico, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953614285540/5696241953614285540_signature_31.png
| 2023-09-04 | 2023-09-01 10:55 | Imus City Health Office I | Velarde Compound, Imus | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25801 | PBSP | armandlasq@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1316 ASSAY: Xpert HIV-1 Viral Load, COUNT: 288 | 2023-08-29 | GXIV-D | 814168 | 3 Modules Failed in XpertCheck | Module Replacement | Performed Maintenance Replaced Modules 1,2 and 4 Generate Reports | All Modules are Operational |  | GX Module 6-C | G 6242 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953616088890/5696241953616088890_signature_28.png | Armand G. Lasquete | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953616088890/5696241953616088890_signature_31.png
| 2023-09-04 | 2023-08-31 14:59 | Molino III City Health Unit | Malinis Compound, Brgy. Molino 3, Bacoor, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25802 | DOH | joanrnd@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-08-25 | GXIV-D | 821221 | Module 1 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 1 Generate Reports | 3 Modules are Operational |  | GX Module 6-C | DR 10582 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953617851922/5696241953617851922_signature_28.png | Jonalyn Aranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953617851922/5696241953617851922_signature_31.png
| 2023-09-04 | 2023-08-31 10:15 | Imus City Health Office II | Bayan Luma, Imus, Cavite | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25799 | DOH | amielynnmangulabnan1217@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1135 | 2023-08-25 | GXIV-D | 821284 | Module 3 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 3 Generate Reports | All Modules are Operational | A3:730981-> 710267 | Gx Module 6-C | DR 10590 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953617908397/5696241953617908397_signature_28.png | Amielynn Mangulabnan | https://www.jotform.com/uploads/macarewecare/222927978372067/5696241953617908397/5696241953617908397_signature_31.png
| 2023-09-04 | 2023-09-04 13:30 | Koronadal City Health Office | Koronadal City, South Cotabato | 12 | South Cotabato | TB-lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25798 | PBSP | calvokenneth12@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2169 | 2023-09-04 | Gx-IV | 707529 | Module 3 - H/W failed | Perform Genexpert machine Preventive maintenance | Replace defective module 3 SN:210019125 with new module SN: 210101819 | 4 operational modules |  | Genexpert module 6 color SN:210101819 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696158879919730771/5696158879919730771_signature_28.png | Kenneth Paolo T. Calvo | https://www.jotform.com/uploads/macarewecare/222927978372067/5696158879919730771/5696158879919730771_signature_31.png
| 2023-09-04 | 2023-09-04 13:32 | CARMEN HEALTH CENTER | CARMEN, CAGAYAN DE ORO CITY | 10 | Misamis Oriental | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25797 | DOH | carmenhealthcenterlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 716 | 2023-07-31 | GX-IV | 838442 | FOR MAINTENANCE AND CALIBRATION | COLLECT IQR,SYSTEM LOG, AND CALIBRATION DATA | PERFORM XPERTCHECK CALIBRATION: CARMEN HC MODULE 1-RETRY THEN SERVICE MODULE 2-PASSED MODULE 3-RETRY THEN SERVICE MODULE 4-PASSED | EQUIPMENT IS OPERATIONAL | MODULE B4 DISABLED | XPERTCHECK KIT | 10361 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696149726224550992/5696149726224550992_signature_28.png | Juvy Madarang | https://www.jotform.com/uploads/macarewecare/222927978372067/5696149726224550992/5696149726224550992_signature_31.png
| 2023-09-04 | 2023-09-04 13:27 | LAGUINDINGAN RURAL HEALTH UNIT | LAGUINDINGAN, MISAMIS ORIENTAL | 10 | Misamis Oriental | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25796 | DOH | chellesambaan@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 919 | 2023-07-31 | GX-IV | 838575 | FOR MAINTENANCE AND CALIBRATION | #NAME? | PERFORM XPERTCHECK CALIBRATION: MODULE B1:FOR SERVICE MODULE B2:PASSED MODULE B3:PASSED MODULE B4:PASSED | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK CARTRIDGES KIT | 10365 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696147256225996973/5696147256225996973_signature_28.png | Rochelle Sambaan | https://www.jotform.com/uploads/macarewecare/222927978372067/5696147256225996973/5696147256225996973_signature_31.png
| 2023-09-04 | 2023-09-04 13:26 | Mexico Rural Health Unit I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25795 | DOH | jasperdizon20@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 507 | 2023-09-04 | Genexpert IV | 831877 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 757999 - 629235 | Genexpert Module | 10568 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696146932395537971/5696146932395537971_signature_28.png | Jasper Dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5696146932395537971/5696146932395537971_signature_31.png
| 2023-09-04 | 2023-09-04 13:20 | LUGAIT RURAL HEALTH UNIT | LUGAIT MISAMIS ORIENTAL | 10 | Misamis Oriental | TB DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25794 | DOH | lugaitmaternity@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 | 2023-08-23 | GX-IV | 838562 | #NAME? | #NAME? | -PERFORM MAINTENANCE AND CLEANING OF GX EQUIPMENT -PERFORM MODULE B4 DEFECTIVE SN:862890 WITH NEW SN:647906 -PERFROM MODULE B1 AND B3 CALIBRATION:PASSED | EQUIPMENT IS OPERATIONAL |  | GENEXPERT MODULE 6-COLOR | 10538 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696144126221053936/5696144126221053936_signature_28.png | Rodelyn B. Limocon | https://www.jotform.com/uploads/macarewecare/222927978372067/5696144126221053936/5696144126221053936_signature_31.png
| 2023-09-04 | 2023-09-04 10:20 | Lake Sebu Rural Health Unit | Porok rosas, Lake Sebu, South Cotabato | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25793 | DOH | jennypingkian23@yahoo.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 496 | 2023-09-04 | Gx-IV | 838569 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine annual Xpertcheck/Calibration | 4 operational modules | All modules passed the annual Xpertcheck/Calibration | Xpertcheck CE-5 cartridge/kit | DR no.: 10408 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696054065971636723/5696054065971636723_signature_28.png | Jenny A. Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/5696054065971636723/5696054065971636723_signature_31.png
| 2023-09-04 | 2022-12-01 15:30 | National Kidney and Transplant Institute | East Avenue Q.C | NCR | NCR 2nd - Quezon City | Molecular | NO | Gathered reports | CEPHEID | Noel Floresca |  |  |  |  | 25792 | MACARE-HBDC | molecularlabnkti@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 743 ASSAY: Others, COUNT: 18 ASSAY: Xpert HBV VL, COUNT: 484 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1213 ASSAY: Xpert MTB/RIF Ultra, COUNT: 940 ASSAY: Xpert SARS-CoV-2, COUNT: 15747 ASSAY: Xpert HCV VL, COUNT: 45 ASSAY: Others, COUNT: 9 |  | GXVI | 110005587 | B1 module encounter continous temperature error 1002 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Troubleshooting | B1 disabled, 15 module operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5696034213619500851/5696034213619500851_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5696034213619500851/5696034213619500851_signature_31.png
| 2023-09-03 | 2023-09-03 12:15 | St. Luke's Medical Center - Global City | 5th Avenue | NCR | NCR 4th - Taguig | Cellular Immunology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 25791 | COMMERCIAL | ci.bgc@stlukes.com.ph | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 154 ASSAY: Xpert CT/NG, COUNT: 1325 ASSAY: Xpert HBV VL, COUNT: 73 ASSAY: Xpert HPV, COUNT: 1045 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2216 ASSAY: Xpert HCV VL, COUNT: 14 ASSAY: Xpert HIV-1 Viral Load, COUNT: 39 |  | GX-IV-D | 839231 | UPS has no backup power | Troubleshoot UPS | Checked UPS for backup power Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. Module C1/866691 has error 4009, not operational and recommend to replace. UPS has no backup power, recommend to replace. | UPS SN: AT1000190402ST0056 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5695243473095498778/5695243473095498778_signature_28.png | Jane Katherina Navora | https://www.jotform.com/uploads/macarewecare/222927978372067/5695243473095498778/5695243473095498778_signature_31.png
| 2023-09-01 | 2023-08-29 17:04 | Aguilar Rural Health Unit | Aguilar, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25790 | PBSP | riverabrendalyn@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 419 |  | GX4 | 834369 | Module B2 Repeated booting | #NAME? | Performed Troubleshooting | Unit is operational | Work with Engineer Mark Joseph Venzon; Module B2 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693728524018160812/5693728524018160812_signature_28.png | Brendalyn Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5693728524018160812/5693728524018160812_signature_31.png
| 2023-09-01 | 2023-09-01 10:28 | Calasiao Rural Health Unit | Calasiao, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25789 | PBSP | calasiaoclinicallab@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 558 |  | GX4 | 834664 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693723874014680858/5693723874014680858_signature_28.png | John John Bauzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5693723874014680858/5693723874014680858_signature_31.png
| 2023-09-01 | 2023-09-01 09:11 | Pangasinan Provincial Hospital | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25788 | PBSP | judzfm@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 32 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2870 | 2023-09-01 | GX4 | 808283 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 666118>>>210102223 - Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210102223 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693720954019419952/5693720954019419952_signature_28.png | Judy Ann F. Mabunga | https://www.jotform.com/uploads/macarewecare/222927978372067/5693720954019419952/5693720954019419952_signature_31.png
| 2023-09-01 | 2023-08-29 13:46 | Umingan Community Hospital | Umingam, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25787 | DOH | uminganch@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 | 2023-08-24 | GX4 | 838473 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Modules: 745274>>>614912 862038>>>724823 - Generated Reports | Replaced Defective Modules | Unit is operational | Work with Engineer Mark Joseph Venzon | GeneXpert Modules 614912, 724823 | DR 10559 10560 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693720474012558220/5693720474012558220_signature_28.png | Agnes S. Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/5693720474012558220/5693720474012558220_signature_31.png
| 2023-09-01 | 2023-08-29 13:46 | Umingan Community Hospital | Umingam, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25786 | DOH | uminganch@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 | 2023-08-24 | GX4 | 838473 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Modules: 745274>>>614912 862038>>>724823 - Generated Reports | Replaced Defective Modules | Unit is operational | Work with Engineer Mark Joseph Venzon | GeneXpert Modules 614912, 724823 | DR 10559 10560 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693719624014910173/5693719624014910173_signature_28.png | Agnes S. Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/5693719624014910173/5693719624014910173_signature_31.png
| 2023-09-01 | 2022-12-29 11:39 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25785 | DOH | allan.vinluan2013@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 154 ASSAY: Xpert MTB/RIF Ultra, COUNT: 639 | 2023-08-24 | GX4 | 838472 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 862035 >>> 807925 | Performed Module Replacement | Unit is operational | Work with Engineer Mark Joseph Venzon | GeneXpert Module 807925 | DR 10569 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693718564018310931/5693718564018310931_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5693718564018310931/5693718564018310931_signature_31.png
| 2023-09-01 | 2022-12-29 11:39 | Rosales Rural Health Unit | Rosales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25784 | DOH | allan.vinluan2013@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 154 ASSAY: Xpert MTB/RIF Ultra, COUNT: 639 | 2023-08-24 | GX4 | 838472 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 862035 >>> 807925 | Performed Module Replacement | Unit is operational | Work with Engineer Mark Joseph Venzon | GeneXpert Module 807925 | DR 10569 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693718564018677825/5693718564018677825_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5693718564018677825/5693718564018677825_signature_31.png
| 2023-09-01 | 2023-09-01 13:15 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 25783 | DOH | laureanjo@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1582 |  | Gx 4 | 838527 | Defective module D1 sn 644949 and D2 710902 | New module installed D1 719233 and D2 686101 clean fan filter upgrade software version to 5.1 | New module installed D1 719233 and D2 686101.. upgrade software version to 5.1 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693589483017360835/5693589483017360835_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5693589483017360835/5693589483017360835_signature_31.png
| 2023-09-01 | 2023-09-01 14:11 | El Nido MHO | El Nido, Palawan | 4B | Palawan | Laboratorh | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 25782 | PBSP | genesorianorodriguez@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 534 | 2023-09-01 | GxIV | 807453 | For module replacement D1 and D3 | Performed module replacement | Performed module replacement. | All 4 modules operational | Defective mods D1: 690232 D3: 661836 | New modules 210100690 210100704 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693584715119126512/5693584715119126512_signature_28.png | Gene S. Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5693584715119126512/5693584715119126512_signature_31.png
| 2023-09-01 | 2023-09-01 14:03 | Teresa Rural Health Unit | M. Garovilla St., Teresa | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25781 | DOH | teresa_rhu@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 741 |  | GX-IV-D | 820710 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693578080329279953/5693578080329279953_signature_28.png | Kamille Peñaranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5693578080329279953/5693578080329279953_signature_31.png
| 2023-09-01 | 2023-09-01 13:16 | Tagum City Health Office | Poblacion, Tagum | 11 | Davao del Norte | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25780 | DOH | tagumcholab.mabini@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2815 | 2023-08-03 | GX-IV 4D | 838553 | Due for Calibration and Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XPC Calibration on four Modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration | Modules B1 and B3 are for observation | XPERTCHECK CE 5 | DR 10448 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693571596221977344/5693571596221977344_signature_28.png | Jesusa Bayato, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5693571596221977344/5693571596221977344_signature_31.png
| 2023-09-01 | 2023-09-01 13:49 | Moncada Rural Health Unit | Moncada Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25779 | PBSP | moncadarhulaboratory.2023@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 339 | 2023-08-29 | GXIV-D | 808292 | A2 for module replacement | Module replacement | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is operational |  |  | 42617 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693569190693336671/5693569190693336671_signature_28.png | Jessa Pabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5693569190693336671/5693569190693336671_signature_31.png
| 2023-09-01 | 2023-09-01 12:51 | Macare Medicals Inc. | 1005 Aurora Grand | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25778 | COMMERCIAL | florescanoel77@gmail.com | 25 |  | 2023-09-01 | GX4 | 803870 | For module replacement B1 628733 - 210072949,B2 651690 - 210072950, B3 210073345, B4 210072884 | Module replacement | Module replacement | GX4 all module installed |  | 900- 0508 | . |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693533053614335188/5693533053614335188_signature_28.png | Macare Medicals Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/5693533053614335188/5693533053614335188_signature_31.png
| 2023-09-01 | 2023-09-01 12:45 | Macare Medicals Inc. | 1005 Aurora Grand | NCR | NCR 2nd - Quezon City | In house | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25777 | COMMERCIAL | florescanoel77@gmail.com | 25 |  | 2023-09-01 | GX4 | 801889 | For module replacement A1 724867 - 210072812, A2 611367 - 210072803, A3 725299 - 210072801, A4 724873 - 210072872 | Module replacement | Module replacement | GX4 installed all module |  | 900-0508 | , |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693529753616725172/5693529753616725172_signature_28.png | Macare Medicals Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/5693529753616725172/5693529753616725172_signature_31.png
| 2023-09-01 | 2023-09-01 11:50 | Zamboanga del sur medical center | Barangay daog, pagadian city | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25776 | DOH | limboemillie@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1262 | 2023-08-02 | GX IV | 838204 | Module 2 and 3 temperature error | Preventive Maintenence and all 4 modules undergone Preventive Maintenence | Preventive Maintenence and all 4 modules undergone Preventive Maintenence | Module 1 ,2 and 4 passed while module 3 service required | None | XPERTCHECK-CE-5 | 10401 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693514509019986971/5693514509019986971_signature_28.png | Emillie limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/5693514509019986971/5693514509019986971_signature_31.png
| 2023-09-01 | 2023-09-01 10:23 | Bacolod City Health Office | Burgos Street, Bacolod City | 6 | Negros Occidental | TB DOTS | YES | TROUBLESHOOTING | CEPHEID | Emerson Bejemino |  |  |  |  | 25775 | PBSP | cyrilcordova0515@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6545 |  | GX IV | 812020 | Module D1 with SN 640561 with error code 4017 due to Hardware failed. | Perform Self Test for module D1 and PM. | Perform Self test and generate files(IQR & System Log) | Module D2 already reported for module replacement. Module 3 and 4 are in good working condition. | Request for additional UPS and AVR for 4 units of GeneXpert IV machine due to fluctuation of electricity in their area. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693472959869800886/5693472959869800886_signature_28.png | Mr. Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5693472959869800886/5693472959869800886_signature_31.png
| 2023-09-01 | 2023-09-01 11:05 | Morong Rural Health Unit | Morong, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25774 | DOH | morongrhu@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 569 |  | GX-IV-D | 821213 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693472080324171047/5693472080324171047_signature_28.png | Necitas V Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5693472080324171047/5693472080324171047_signature_31.png
| 2023-09-01 | 2023-08-23 11:54 | Cauayan Municipal Health Office | Municipality of Cauayan | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25773 | DOH | djeatrmt@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1289 |  | GX4 | 838477 | Module B1 is Harware Failed but has already allocated new module. | Conduct PM and Generate files(IQR and GXC File) | Run Xpert check calibration for module B2, B3 and B4 | Modules B2, B3 and B4 passed the Xpert Check. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693470789862295476/5693470789862295476_signature_28.png | Doriza Joy Tizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5693470789862295476/5693470789862295476_signature_31.png
| 2023-09-01 | 2023-08-31 18:16 | Claveria Unit | Claveria, Burias Island | 5 | Masbate | RTDL Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25772 | DOH | mgrmartinez757@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 514 |  | GX4 | 838572 | For Callibration | XpertCheck Result - All Modules Passed Cleaned cartridge bays, plunger rod, icore | Performed XpertCheck Performed Preventive Maintenance | Unit is operational All modules passed the Callibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693420196805897208/5693420196805897208_signature_28.png | Mary Grace R. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5693420196805897208/5693420196805897208_signature_31.png
| 2023-09-01 | 2023-08-30 10:29 | Malilipot RHU | Malilipot | 5 | Albay | RTDL Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25771 | PBSP | rechelbiscocho80@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1968 |  | GX4 | 808280 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693415976801095721/5693415976801095721_signature_28.png | Rechel canomoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5693415976801095721/5693415976801095721_signature_31.png
| 2023-09-01 | 2023-08-29 15:21 | Tabaco City Health Unit | Taihi, Tabaco City | 5 | Albay | RTDL Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25770 | PBSP | elyjoy17rmt@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3142 |  | GX4 | 804257 | Monitor for replacement | Monitor Replaced - V5GAH857 Cleaned cartridge bays, plunger rod, icore Generate reports | Replacement of New Monitor Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693413456805218745/5693413456805218745_signature_28.png | Ely Joy Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5693413456805218745/5693413456805218745_signature_31.png
| 2023-09-01 | 2023-09-01 09:04 | Southern Isabela Medical Center | Santiago, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25769 | PBSP | simcmicrobiology@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3806 ASSAY: Xpert HIV-1 Viral Load, COUNT: 174 | 2023-09-01 | Genexpert IV | 805104 | Module 4 error | Performed Module Replacement on module 4 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module Replacement = (862169 - 210100854) | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693397847264634744/5693397847264634744_signature_28.png | Cass Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5693397847264634744/5693397847264634744_signature_31.png
| 2023-09-01 | 2023-09-01 08:39 | Medellin Rural Health Unit | Poblacion Medellin, Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 25768 | DOH | medellindtc@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1309 |  | Gx4 | 838506 | Module 2 SN 862245 and Module 4 SN 633038 failed during calibration and error 1002 | Replaced new module | Replaced with New module 2 Sn 883447 and Module 4 210009597 | All. Modules are available but due recurrent occurrence of Module 1 and module 3 error of 1002 awaiting the replacement hereafter. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693383816325230624/5693383816325230624_signature_28.png | Raul Galano. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5693383816325230624/5693383816325230624_signature_31.png
| 2023-09-01 | 2023-08-31 10:48 | Laoang Rural Health Unit | Laoang, Northern Samar | 8 | Northern Samar | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25766 | DOH | roseadarayan26@gmail.com | 27 |  | 2023-08-31 | GXIV | 838202 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | Xpertcheck kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693377919419122857/5693377919419122857_signature_28.png | Rosario B. Adarayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5693377919419122857/5693377919419122857_signature_31.png
| 2023-09-01 | 2023-09-01 08:21 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | Tb-dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25767 | DOH | sarah_yangzon@yahoo.com | 26 |  | 2023-09-01 | GXIV | 838406 | Due for Calibration | #NAME? | #NAME? | Unit is Operational |  | Xpertcheck kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5693377929412032896/5693377929412032896_signature_28.png | Ma.teresa E aliman | https://www.jotform.com/uploads/macarewecare/222927978372067/5693377929412032896/5693377929412032896_signature_31.png
| 2023-08-31 | 2023-08-31 16:51 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Bloodbank | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Matthew Roxas |  |  |  |  | 25765 | MACARE-HBDC | lendmelad@yahoo.com | 27 |  | 2023-08-24 | MCS+ | 15H033G | Line sensor problem | Replaced Line Sensor | Performed Maintenance Performed Calibration Performed Functionality Test | Unit is Operational | N/a | Line Sensor | 10574 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692815532112203502/5692815532112203502_signature_28.png | Lolenia Melad | https://www.jotform.com/uploads/macarewecare/222927978372067/5692815532112203502/5692815532112203502_signature_31.png
| 2023-08-31 | 2023-08-31 16:35 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | TB Dots | NO | REPAIR MAINTENANCE Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 25764 | DOH | olivesalvador10@gmail.com | 27 |  | 2023-08-31 | Genexpert IV | 831718 | Module 3 error | Performed Xpertcheck on module 3 = passed | Performed Maintenance Gathered Reports Performed Module Replacement (757493 - 620570 Performed Functionality Test | Unit is Operational | Performed on August 8, 2023 | Genexpert Module = 620570 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692808572117703096/5692808572117703096_signature_28.png | Olive A.Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5692808572117703096/5692808572117703096_signature_31.png
| 2023-08-31 | 2021-12-31 15:14 | National Crop Protection Center | Los Baños, Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 25763 | COMMERCIAL | nmbarbecho@up.edu.ph | 25 |  |  | Primo Star HD | 3148001025 | None | Adjusted Kohler | Light cleaning Check up of unit | Unit is operational. Both eyepieces have some mold |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692777625129595745/5692777625129595745_signature_28.png | Normandy Barbecho | https://www.jotform.com/uploads/macarewecare/222927978372067/5692777625129595745/5692777625129595745_signature_31.png
| 2023-08-31 | 2021-12-31 15:09 | National Crop Protection Center | Los Baños, Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 25761 | COMMERCIAL | nmbarbecho@up.edu.ph | 25 |  |  | Primo Star HD | 3148011024 | Does not power on | Swapped power supply of working unit to isolate problem | Light cleaning Check-up of unit Adjusted Kohler Checked Power Supply output | Unit is operational but need to replace power supply | Eyepieces have slight mold |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692777615124909370/5692777615124909370_signature_28.png | Normandy Barbecho | https://www.jotform.com/uploads/macarewecare/222927978372067/5692777615124909370/5692777615124909370_signature_31.png
| 2023-08-31 | 2022-12-31 15:16 | National Crop Protection Center | Los Baños, Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 25760 | COMMERCIAL | nmbarbecho@up.edu.ph | 25 |  |  | Primo Star HD | 3148001108 | None | Adjusted Kohler | Light cleaning Check up of unit | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692777615126753972/5692777615126753972_signature_28.png | Normandy Barbecho | https://www.jotform.com/uploads/macarewecare/222927978372067/5692777615126753972/5692777615126753972_signature_31.png
| 2023-08-31 | 2023-08-31 15:05 | National Crop Protection Center | Los Baños, Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 25762 | COMMERCIAL | nmbarbecho@up.edu.ph | 25 |  |  | Stemi 305 | 3947001216 | None | None | Light Cleaning Checking of unit | Unit is Operational. Left eyepiece has mold. | For Preventive Maintenance |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692777625125979043/5692777625125979043_signature_28.png | Normandy Barbecho | https://www.jotform.com/uploads/macarewecare/222927978372067/5692777625125979043/5692777625125979043_signature_31.png
| 2023-08-31 | 2020-12-31 13:37 | Baganga Rural Health Unit | Poblacion Central Baganga | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25759 | DOH | rhubaganga@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 577 | 2023-08-02 | GX-IV 4D | 838395 | Due for Calibration and Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Perform XPERTCHECK Calibration | Perform Preventive Maintenance and XpertCheck Calibration | Module B2 for Service Required Modules B1 B3 and B4 Passed the Calibration |  | XPERTCHECK CE 5 | DR 10428 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692763137652641050/5692763137652641050_signature_28.png | Liza N. Narciso RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5692763137652641050/5692763137652641050_signature_31.png
| 2023-08-31 | 2023-08-31 14:45 | Metro Antipolo Hospital | Marikina-Infanta Hwy, Antipolo | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25758 | COMMERCIAL | mahmci_labcmt@yahoo.com | 20 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 102 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-IV-D | 110005265 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692754412332034534/5692754412332034534_signature_28.png | Graziela Marie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/5692754412332034534/5692754412332034534_signature_31.png
| 2023-08-31 | 2023-08-31 14:35 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25757 | DOH | laguitannataniel@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 | 2023-08-25 | GXIV | 831677 | Module A1 for replacement | #NAME? | Replaced A1 | Unit is operational. |  | GeneXpert Module SN: 766558 | 10586 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692745834126651891/5692745834126651891_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5692745834126651891/5692745834126651891_signature_31.png
| 2023-08-31 | 2023-08-31 13:37 | Sta. Ignacia Rural Health Unit | Sta. Ignacia Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25756 | DOH | lysasome98@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 496 |  | GXIV-D | 831909 | A2 for re-xpertcheck | Xpertcheck Preventive maintenance | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Performed calibration Generate reports Archived test results Performed functionality test | Unit is partially operational, A2 under observation | Used 1 xpertcheck kit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692706262434363754/5692706262434363754_signature_28.png | Alyssa Jayne Antonio | https://www.jotform.com/uploads/macarewecare/222927978372067/5692706262434363754/5692706262434363754_signature_31.png
| 2023-08-31 | 2023-08-31 12:47 | Dinas Rural Health Unit | Dinas, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25755 | DOH | bernantereynacres@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 349 | 2023-08-02 | GX IV | 838533 | None | All modules undergone Xpertcheck calibration and Preventive Maintenence | All modules undergone Xpertcheck calibration and Preventive Maintenence | Module 2, 3 and 4 passed Module 1 service required | None | XPERTCHECK-CE-5 | 10394 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692689491125207173/5692689491125207173_signature_28.png | Reyna Cres | https://www.jotform.com/uploads/macarewecare/222927978372067/5692689491125207173/5692689491125207173_signature_31.png
| 2023-08-31 | 2023-08-31 12:58 | San Remigio Rhu | Poblacion, San Remigio | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  |  |  |  | 25754 | WHO | rodacanetecamay@gmail.com | 22 |  | 2023-08-30 | GX4 | 847333 | Due for Calibration | Perform Calibration | Perform calibration | Module B1, B2, B4 passed for calibration module B3 fail for calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692673620224776846/5692673620224776846_signature_28.png | Roday C. Camay | https://www.jotform.com/uploads/macarewecare/222927978372067/5692673620224776846/5692673620224776846_signature_31.png
| 2023-08-31 | 2023-08-31 09:26 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25753 | DOH | onehope834@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1566 | 2023-08-31 | Gx-IV | 838528 | Module 2 and 4 waiting for replacement modules | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine annual Xpertcheck/Calibration | 2 operational modules | Module 1 and 3 passed the annual Calibration/Xpertcheck | Xpertcheck CE-5 cartridge/kit | DR no.: 10416 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692663211022204330/5692663211022204330_signature_28.png | Emelyn Agustin Cadiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5692663211022204330/5692663211022204330_signature_31.png
| 2023-08-31 | 2023-08-31 09:26 | Cotabato City Health Office | Cotabato City, Maguindanao | BARMM | Maguindanao | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25752 | DOH | onehope834@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1566 | 2023-08-31 | Gx-IV | 838528 | Module 2 and 4 waiting for replacement modules | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine annual Xpertcheck/Calibration | 2 operational modules | Module 1 and 3 passed the annual Calibration/Xpertcheck | Xpertcheck CE-5 cartridge/kit | DR no.: 10416 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692663181027877981/5692663181027877981_signature_28.png | Emelyn Agustin Cadiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5692663181027877981/5692663181027877981_signature_31.png
| 2023-08-31 | 2023-08-31 11:31 | Bureau of Plant Industry Central Post Entry Quarantine Section | BPI Compound Economic Garden, Brgy. Timugan, Los Baños, Laguna | 4A | Laguna | Entom Laboratory | NO | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 25749 | COMMERCIAL | cpeqsprocurement.losbanos@gmail.com | 25 |  | 2023-08-30 | Stemi 305 | 3947020836 | None | None | Installed and setup new Stemi 305 unit | Unit is operational |  | Stemi 305, Stand, Spotlight | DR 10615, SI 96205 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692652805121083161/5692652805121083161_signature_28.png | VICKY S. MANAYON | https://www.jotform.com/uploads/macarewecare/222927978372067/5692652805121083161/5692652805121083161_signature_31.png
| 2023-08-31 | 2023-08-31 11:36 | Bureau of Plant Industry Central Post Entry Quarantine Section | BPI Compound Economic Garden, Brgy. Timugan, Los Baños, Laguna | 4A | Laguna | Entom Laboratory | NO | TROUBLESHOOTING | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 25751 | COMMERCIAL | cpeqsprocurement.losbanos@gmail.com | 25 |  |  | Stemi 305 | 3947019527 | Spotlight not working | Reset spotlight connections | Checked connections | Unit is operational but spotlight works intermittently | Spotlight needs to be replaced |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692652805122557271/5692652805122557271_signature_28.png | VICKY S. MANAYON | https://www.jotform.com/uploads/macarewecare/222927978372067/5692652805122557271/5692652805122557271_signature_31.png
| 2023-08-31 | 2023-08-31 11:38 | Bureau of Plant Industry Central Post Entry Quarantine Section | BPI Compound Economic Garden, Brgy. Timugan, Los Baños, Laguna | 4A | Laguna | Plant Pathology Laboratory | NO | INSTALLATION | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 25750 | COMMERCIAL | cpeqsprocurement.losbanos@gmail.com | 25 |  | 2023-08-30 | Stemi 305 | 3947020834 | None | None | Installed and setup new Stemi 305 unit | Unit is operational |  | Stemi 305, Stand, Spotlight | DR 10615, SI 96205 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692652805128459970/5692652805128459970_signature_28.png | VICKY S. MANAYON | https://www.jotform.com/uploads/macarewecare/222927978372067/5692652805128459970/5692652805128459970_signature_31.png
| 2023-08-31 | 2023-08-31 12:17 | Tanauan City Health Office 1 | Tanauan, Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25748 | DOH | ramilomher27@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1793 | 2023-08-24 | GXIV | 821299 | Unit for PM. For module replacement of module A3. | Performed PM. Performed module replacement. Generated reports. Troubleshoot module A3 due to H/W Failure after installation. | PM and module replacement. | Unit is operational. Module A3 (731300) out of box H/W Failure. For reporting of module A3. |  | Gx module A3 (731300) in replacement to defective A3 (731305). | DR#10557 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692652449519174300/5692652449519174300_signature_28.png | EDITA D. RAMILO | https://www.jotform.com/uploads/macarewecare/222927978372067/5692652449519174300/5692652449519174300_signature_31.png
| 2023-08-31 | 2023-08-31 12:17 | Kamuning Health Center | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25747 | DOH | vicvic.chua25@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 520 ASSAY: Xpert MTB/RIF Ultra, COUNT: 578 | 2023-08-24 | GXIV | 831612 | Modules A2, A4 for replacement | #NAME? | Replaced A2, A4 | Unit is operational. |  | GeneXpert Module SN: 699275, 621383 | 10562, 10563 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692649344129211601/5692649344129211601_signature_28.png | Jesusa Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5692649344129211601/5692649344129211601_signature_31.png
| 2023-08-31 | 2023-08-31 12:04 | Sara Rural Health Unit | Sara, Iloilo | 6 | Iloilo | TB-Dots | YES | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 25746 | PBSP | candyparrenas@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 31 ASSAY: Xpert MTB/RIF Ultra, COUNT: 229 | 2023-08-31 | GXIV | 838200 | Module B4 - XpertCheck Service Required | -Replace defective module B4/653895 with new module SN: 210100890 | #NAME? | All modules are functional. |  | GeneXpert Module 6-Colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692642366029106311/5692642366029106311_signature_28.png | Ma. Candelaria V. Parreñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5692642366029106311/5692642366029106311_signature_31.png
| 2023-08-31 | 2023-08-31 11:40 | Cardona Rural Health Unit | Cardona | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 25745 | DOH | rhucardona@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 487 |  | GX-IV-D | 821215 | For maintenance Error 2008 on module B1 | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. Module B1/730608 operational but under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692629502339756214/5692629502339756214_signature_28.png | Khawla punzalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5692629502339756214/5692629502339756214_signature_31.png
| 2023-08-31 | 2021-12-31 10:58 | Davao Oriental Provincial Hospital - Cateel Zone | Poblacion, Cateel Davao Oriental | 11 | Davao Oriental | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25744 | PBSP | minozaeljohnmark@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 660 |  | GX-IV 4D | 807432 | Module B1 for replacement due XPC result is for service reuqired | Replaced the Defective Module SN 661873 with new Module SN: 210100409 | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692611001263020440/5692611001263020440_signature_28.png | ELJOHN MARK A. MIŇOZA | https://www.jotform.com/uploads/macarewecare/222927978372067/5692611001263020440/5692611001263020440_signature_31.png
| 2023-08-31 | 2023-08-31 10:48 | Laoang Rural Health Unit | Laoang, Northern Samar | 8 | Northern Samar | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25743 | DOH | roseadarayan26@gmail.com | 27 |  | 2023-08-31 | GXIV | 838202 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | All module passed | Xpertcheck kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692602630811767036/5692602630811767036_signature_28.png | Rosario B. Adarayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5692602630811767036/5692602630811767036_signature_31.png
| 2023-08-31 | 2023-08-31 10:37 | Talisay Rural Health Unit | Talisay, Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25742 | DOH | medcertrhutalisay@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 506 | 2023-08-25 | GXIV | 820201 | Unit for PM. For module replacement of module A4. | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module A4 (706165) in replacement to defective A4 (792112). | DR#10595 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692598304235775161/5692598304235775161_signature_28.png | Arien Joy Festijo | https://www.jotform.com/uploads/macarewecare/222927978372067/5692598304235775161/5692598304235775161_signature_31.png
| 2023-08-30 | 2022-12-30 16:07 | Payatas B Health Center | Bulacan st, Payatas | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25741 | DOH | racquelmrocero@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1632 | 2023-08-25 | GX4 | 831899 | For module replacement and maintenance | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | Operational |  | 900-0491RS | 10601 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692046552649122631/5692046552649122631_signature_28.png | Racquel M. Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5692046552649122631/5692046552649122631_signature_31.png
| 2023-08-30 | 2022-12-30 16:07 | Payatas B Health Center | Bulacan st, Payatas | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25740 | DOH | racquelmrocero@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1632 | 2023-08-25 | GX4 | 831899 | For module replacement and maintenance | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | Operational |  | 900-0491RS | 10601 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692046012645478504/5692046012645478504_signature_28.png | Racquel M. Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5692046012645478504/5692046012645478504_signature_31.png
| 2023-08-30 | 2023-08-30 16:07 | Payatas B HC | Bulacan st, Payatas | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25739 | DOH | racquelmrocero@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1632 | 2023-08-25 | GX4 | 831899 | For module replacement and maintenance | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | Operational |  | 900-0491RS | 10601 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5692044892644192602/5692044892644192602_signature_28.png | Racquel M. Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5692044892644192602/5692044892644192602_signature_31.png
| 2023-08-30 | 2021-12-30 11:29 | Maco Rural Health Unit | Poblacion Maco | 11 | Davao de Oro | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25738 | DOH | edlienkai26@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF, COUNT: 174 ASSAY: Xpert MTB/RIF Ultra, COUNT: 711 | 2023-08-03 | GX-IV 4D | 838500 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XPC Calibration | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration and Operational | High Room Temperature | XPERTCHECK CE 5 | DR 10440 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691981666107148435/5691981666107148435_signature_28.png | Edlien Kai Prado | https://www.jotform.com/uploads/macarewecare/222927978372067/5691981666107148435/5691981666107148435_signature_31.png
| 2023-08-30 | 2023-08-30 15:39 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25737 | PBSP | crmctblab@gmail.com | 22.3 | ASSAY: Xpert HBV VL, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 42 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2023-08-30 | Gx-IV | 807607 | Module 3 for service Required | Perform Genexpert machine Preventive maintenance | Replace module SN:796163 with new module SN:210100875 | 4 operational modules |  | Genexpert module 6 color SN:210100875 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691924809028943238/5691924809028943238_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5691924809028943238/5691924809028943238_signature_31.png
| 2023-08-30 | 2023-08-30 15:33 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25736 | PBSP | crmctblab@gmail.com | 22.1 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 42 ASSAY: Xpert HBV VL, COUNT: 7 | 2023-08-30 | Gx-IV | 807463 | Module 3 for service Required Module 4 H/W failed | Perform Genexpert machine Preventive maintenance | Replace module 3 SN:653874 with new module SN:210100874 Report module 4 problem/error | 3 operational modules |  | Genexpert module 6 color SN: 210100874 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691924149023039836/5691924149023039836_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5691924149023039836/5691924149023039836_signature_31.png
| 2023-08-30 | 2023-08-30 14:27 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25735 | PBSP | jvdiaz.pbsp@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 27 ASSAY: Xpert HBV VL, COUNT: 2 ASSAY: Others, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2410 ASSAY: Xpert SARS-CoV-2, COUNT: 8848 ASSAY: Xpert HIV-1 Viral Load, COUNT: 158 ASSAY: Others, COUNT: 9 | 2023-08-30 | GXIV | 834432 | Modules A1, A2, A3 for replacement | #NAME? | Replaced modules A1, A2 A3 | Unit is operational. |  | GeneXpert Module SN: 210101937, 210101939, 210101938 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691892842429609983/5691892842429609983_signature_28.png | Jim Vianne Diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5691892842429609983/5691892842429609983_signature_31.png
| 2023-08-30 | 2023-08-30 14:53 | Rosario Super Health Center | Dr. Sixto Avenue | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 25734 | WHO | eden.cawayan@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1948 |  | GX-IV-D | 843924 | For maintenance Module B1 has error 4012 | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports. Checked module B1 | Unit is operational. Module B1/684120 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691881348715814627/5691881348715814627_signature_28.png | Eden R. Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5691881348715814627/5691881348715814627_signature_31.png
| 2023-08-30 | 2023-08-30 14:47 | Victorias City Health Office | Victorias City | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 25733 | PBSP | vincent_llamoleonor@yahoo.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1468 |  | GX4 | 838141 | Module B3 is Hardware failed with SN 210070812 | Save IQR file before and after replacement | Replaced new module for B3 with SN 210100935 | All modules are now functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691876504427959753/5691876504427959753_signature_28.png | Vincent Leonor | https://www.jotform.com/uploads/macarewecare/222927978372067/5691876504427959753/5691876504427959753_signature_31.png
| 2023-08-30 | 2023-08-29 16:37 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25732 | PBSP | crmcmolecularlab@gmail.com | 25.6 | ASSAY: Xpert SARS-CoV-2, COUNT: 5633 | 2023-08-29 | Gx-IV | 834441 | Module 1,2 and 3 for service Required | Perform Genexpert machine Preventive maintenance | Replace module 1 SN:784179 with new module SN: 210100871 Replace module 2 SN:784177 with new module SN: 210100872 Replace module 3 SN:784088 with new module SN: 210100873 | 4 operational modules |  | Genxpert Module 6 color SN:210100871, 210100872, and 210100873 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691857829026083691/5691857829026083691_signature_28.png | Chareasse love yu utto | https://www.jotform.com/uploads/macarewecare/222927978372067/5691857829026083691/5691857829026083691_signature_31.png
| 2023-08-30 | 2023-08-29 16:23 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25731 | DOH | crmcmolecularlab@gmail.com | 24.9 | ASSAY: Xpert SARS-CoV-2, COUNT: 5633 | 2023-08-29 | Gx-IV | 837662 | Module C4 under observation due to temperature error | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules | Module 2 and 4 passed the annual Xpertcheck/Calibration | Xpertcheck CE-5 cartridge/kit | DR no.:10406 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691856449027522437/5691856449027522437_signature_28.png | Chareasse love yu utto | https://www.jotform.com/uploads/macarewecare/222927978372067/5691856449027522437/5691856449027522437_signature_31.png
| 2023-08-30 | 2023-08-30 13:52 | San Roque Rural Health Unit | San Roque, Northern Samar | 8 | Northern Samar | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25730 | DOH | lynnjarito@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: | 2023-08-30 | GXIV | 838498 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | Module C4 - service required | Xpertcheck kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691842164119225749/5691842164119225749_signature_28.png | Lynn Jarito | https://www.jotform.com/uploads/macarewecare/222927978372067/5691842164119225749/5691842164119225749_signature_31.png
| 2023-08-30 | 2023-08-30 13:20 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | IPMM | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  |  |  |  | 25729 | COMMERCIAL | regenerativemedicine@themedicalcity.com | 24 |  |  | MCS+ | 04H020 | For maintenance | Performed maintenance | Performed maintenance Checked rotor pumps Verified calibration parameters Simulation run Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691833048717839192/5691833048717839192_signature_28.png | Maria Zoraida Caiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5691833048717839192/5691833048717839192_signature_31.png
| 2023-08-30 | 2023-08-30 12:32 | Tagbina Rural Health Unit | Brgy. Poblacion, Tagbina | 13 | Surigao del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 25728 | PBSP | ellalaguros@yahoo.com | 25.75 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 803 |  | GX-4 | 838877 | No problems found | None | Replacing Module B2 with SN: 863430 to SN: 210100857 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691795818612707500/5691795818612707500_signature_28.png | Ellaquim L. Albarracin | https://www.jotform.com/uploads/macarewecare/222927978372067/5691795818612707500/5691795818612707500_signature_31.png
| 2023-08-30 | 2023-08-30 12:01 | Velasquez Health Center | Varona St. Tondo | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25727 | DOH | wcpatron11x4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1033 | 2023-08-24 | GXIV | 831534 | Unit for PM. For module replacement of module B3. | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B3 (676004) in replacement to defective B3 (876732). | DR#10583 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691777387912500132/5691777387912500132_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5691777387912500132/5691777387912500132_signature_31.png
| 2023-08-30 | 2023-08-30 11:32 | Cotabato Sanitarium | Cotabato City, Maguindanao | BARMM | Maguindanao | Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 25726 | PBSP | cotsanlaboratory@gmail.com | 28.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 853 |  | Gx-IV | 815277 | Module 2 stuck cartridge Module 1 temperature error Request for Genexpert fan filter | Perform Genexpert machine Preventive maintenance | Fixed/repair module 2 stuck cartridge Genexpert machine Cleaning Troubleshoot module 1 problem Import HIV viral load Assay | 4 operational modules | Module 1 under observation due to temperature error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691769033011109694/5691769033011109694_signature_28.png | Albarah Meditar | https://www.jotform.com/uploads/macarewecare/222927978372067/5691769033011109694/5691769033011109694_signature_31.png
| 2023-08-30 | 2023-08-30 11:26 | Pontevedra Rural Health Unit | Pontevedra, Capiz | 6 | Capiz | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25725 | DOH | rtabngjr@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 556 | 2023-07-27 | GXIV | 838296 | Due for XPC | #NAME? | #NAME? | All modules are functional. | Module B2 and B4 passed XPC, used three XPCs. | XpertCheck-CE-5 | 10313 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691753357225765351/5691753357225765351_signature_28.png | Reynaldo Tabang, Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5691753357225765351/5691753357225765351_signature_31.png
| 2023-08-30 | 2023-08-30 11:03 | Northern Samar Provincial Hospital | Catarman Northern Samar | 8 | Northern Samar | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25724 | PBSP | dinahada70@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2233 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 | 2023-08-30 | GXIV | 807566 | Module B1/661952 defective | #NAME? | - Install new module on B1 w/ SN: 210100895 - Performed Preventive maintenance - Performed Functionality test - Generate reports | Unit is Operational |  | Genexpert Module 6 color SN 210100895 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691742068857230662/5691742068857230662_signature_28.png | Dinah F. Ada | https://www.jotform.com/uploads/macarewecare/222927978372067/5691742068857230662/5691742068857230662_signature_31.png
| 2023-08-30 | 2023-08-30 10:46 | Parola Health Center | Parola, Tondo | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25723 | DOH | friaslynette@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1081 | 2023-08-25 | GXIV | 831871 | Unit for PM. For module replacement of module B4. | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module B4 (636390) in replacement to defective B4 (757238). | DR#10587 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691741518425579160/5691741518425579160_signature_28.png | Lynette Olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5691741518425579160/5691741518425579160_signature_31.png
| 2023-08-30 | 2023-08-30 10:58 | The medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25722 | COMMERCIAL | moleculardx@themedicalcity.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 20 ASSAY: Others, COUNT: 27 ASSAY: Xpert CT/NG, COUNT: 18 ASSAY: Xpert HBV VL, COUNT: 264 ASSAY: Xpert HPV, COUNT: 182 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1432 ASSAY: Others, COUNT: 28 ASSAY: Xpert SARS-CoV-2, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 34 ASSAY: Xpert HIV-1 Viral Load, COUNT: 316 |  | GX-IV-L | 802630 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. Module B4/210071314 not operational. | Module B4 has error 4012 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691739168713164246/5691739168713164246_signature_28.png | Stephanie Obida | https://www.jotform.com/uploads/macarewecare/222927978372067/5691739168713164246/5691739168713164246_signature_31.png
| 2023-08-30 | 2023-08-30 09:52 | Sultan Mastura Rural Health Unit | Tapayan, Sultan Mastura | BARMM | Maguindanao | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25721 | DOH | salehshahinaz@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 391 | 2023-08-30 | Gx-IV | 837613 | N/A | Genexpert machine Preventive maintenance | Genexpert machine Cleaning Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules | All modules passed for annual Calibration | Xpertcheck CE-5 cartridge/kit | DR no.: 10419 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691711197325323785/5691711197325323785_signature_28.png | Shahinaz Saleh | https://www.jotform.com/uploads/macarewecare/222927978372067/5691711197325323785/5691711197325323785_signature_31.png
| 2023-08-30 | 2023-08-30 09:37 | Roxas City Health Office | Roxas City, Capiz | 6 | Capiz | TB-Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25720 | PBSP | ruby_camacho@ymail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1840 ASSAY: Xpert HIV-1 Viral Load, COUNT: 148 | 2023-08-30 | GXIV | 805333 | #NAME? | -Remove stuck cartridge on module B2/648722 -Replace defective module B4/805093 with new module SN: 210100878 | #NAME? | All modules are functional. |  | GeneXpert Module 6-Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691690728963363198/5691690728963363198_signature_28.png | Ruby Camacho Lut | https://www.jotform.com/uploads/macarewecare/222927978372067/5691690728963363198/5691690728963363198_signature_31.png
| 2023-08-30 | 2023-08-30 09:18 | Tondo Medical Center | H. Lopez Blvd. Tondo | NCR | NCR 1st - Manila | Covid laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25719 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 10 ASSAY: Xpert HPV, COUNT: 210 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1758 ASSAY: Xpert SARS-CoV-2, COUNT: 9966 ASSAY: Xpert HIV-1 Viral Load, COUNT: 7 | 2023-08-29 | GX16 | 844121 | Unit for PM. For module replacement of modules A2, A4, B1 and C1 | Performed PM. Performed module replacement. Generated reports. | PM and module replacement. | Unit is operational. |  | GX modules A2 (210102219) A4 (210102199) B1 (210102262) and C1 (210102234) in replacement to defective A2 (210057331) A4 (220003328) B1 (802228) and C1 (210070820) | G6240 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691682609317005408/5691682609317005408_signature_28.png | Vernelee panlaqui | https://www.jotform.com/uploads/macarewecare/222927978372067/5691682609317005408/5691682609317005408_signature_31.png
| 2023-08-30 | 2022-12-29 15:16 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | RHIVDA Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25718 | MACARE-HBDC | spmcrhivda@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 57 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1397 |  | GX-IV 4D | 839152 | Error Encounter on Module B2 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XPC for error Verification | Perform Preventive Maintenance and Module Troubleshooting | Module 2 for service required Data will be sent to Cepheid for further Assessment | Module B2 SN 866069 temporarily disabled |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691471370013725091/5691471370013725091_signature_28.png | Ma. Flordeliza Go, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5691471370013725091/5691471370013725091_signature_31.png
| 2023-08-29 | 2023-08-29 13:49 | Roxas Memorial Provincial Hospital | Brgy. Lanot, Roxas City, Capiz | 6 | Capiz | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25717 | DOH | fanugon.grace@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 546 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 | 2023-07-27 | GXIV | 837615 | Due for XPC | #NAME? | #NAME? | All modules are functional. | All modules passed Xpert Check. | XpertCheck-CE-5 | 10315 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691082765316139978/5691082765316139978_signature_28.png | Joyce B. Bercadez | https://www.jotform.com/uploads/macarewecare/222927978372067/5691082765316139978/5691082765316139978_signature_31.png
| 2023-08-29 | 2023-08-29 15:36 | Labuan General hospital | Labuan Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25716 | DOH | laboratory.lph@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2395 | 2023-08-29 | GX IV | 838568 | Module 4 stucked cartridge | Removed stucked cartridge in module 4 | Removed stucked cartridge in module 4 | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691040211026937483/5691040211026937483_signature_28.png | Irekatte L. Acierto | https://www.jotform.com/uploads/macarewecare/222927978372067/5691040211026937483/5691040211026937483_signature_31.png
| 2023-08-29 | 2023-08-29 14:43 | Cabanatuan City Health Office | Kapt. Pepe Subdivision, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25715 | DOH | cholab2021@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2209 ASSAY: Xpert MTB/RIF, COUNT: 200 | 2023-08-24 | Genexpert IV | 831592 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Perfomred Functionality Test | Unit is Operational | Module Replacement = 786121 - 862394 | Genexpert Module = 862394 | 10551 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5691013140354528199/5691013140354528199_signature_28.png | Maricris So | https://www.jotform.com/uploads/macarewecare/222927978372067/5691013140354528199/5691013140354528199_signature_31.png
| 2023-08-29 | 2023-08-29 12:54 | San Remihio Rural Health | Poblacion , San Remigion | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 25714 | WHO | rodacanetecamay@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 46 | 2023-08-24 | GX4 | 847333 | Preventive maintenance | Perform preventive maintenance for all module's | Perform preventive maintenance | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5690945803919582014/5690945803919582014_signature_28.png | Roda Cañete Camay | https://www.jotform.com/uploads/macarewecare/222927978372067/5690945803919582014/5690945803919582014_signature_31.png
| 2023-08-29 | 2023-08-29 12:30 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25713 | DOH | vanessadelmonte401@gmail.com | 28 |  | 2023-08-29 | GXIV | 838409 | Due for Calibration | #NAME? | #NAME? | Unit is Operational | #NAME? | Xpertcheck kit | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5690934181418950209/5690934181418950209_signature_28.png | Vanessa Mabag | https://www.jotform.com/uploads/macarewecare/222927978372067/5690934181418950209/5690934181418950209_signature_31.png
| 2023-08-29 | 2023-08-29 12:03 | Talacogon District Hospital | Purok Bayabas, San Isidro, Talacogon | 13 | Agusan del Sur | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25712 | DOH | shyratortor@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2300 | 2023-08-01 | GX-4 | 838275 | Module C1, C2 and C4 were found service required thru XPC | XPC Result for reporting to Cepheid | Xpertcheck Calibration Preventive Maintenance | All modules are functional | 1 modules passed in XPC and 3 modules prior for replacement as for now modules can be used | XpertCheck CE-IVD 5 Cartridges/Kit | 10392 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5690918322718825199/5690918322718825199_signature_28.png | Stella Rica Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5690918322718825199/5690918322718825199_signature_31.png
| 2023-08-29 | 2023-08-29 10:21 | Paniqui Rural Health Unit | Oaniqui Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25711 | PBSP | forylab76@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV-D | 812792 | CPU ERROR (3F0) Monitor dead pixel | Troubleahooting | Performed troubleshooting Performed preventive maintenance Performed functionality test | Unit is not operational CPU for repalacement | CPU error 3FO for replacement, monitor dead pixel |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5690860030222010850/5690860030222010850_signature_28.png | Fortune Abutin | https://www.jotform.com/uploads/macarewecare/222927978372067/5690860030222010850/5690860030222010850_signature_31.png
| 2023-08-29 | 2023-08-24 09:59 | Cebu TB Reference Lab | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION ON-SITE ASSISTANCE Installation from cebu cho to ctrl srp | CEPHEID | Raymond Laurito |  |  |  |  | 25710 | PBSP | asmantiquilla.pbsp@gmail.com | 26 |  |  | Gx4 | 838882 | Transfer from Cebu CHO to CTRL SRP | Installation of Genexpert to CTRL SRP | Installation of Genexpert to CTRL SRP | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5690839866026126057/5690839866026126057_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5690839866026126057/5690839866026126057_signature_31.png
| 2023-08-29 | 2023-08-23 09:51 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE Pull out genexpert machine | CEPHEID | Raymond Laurito |  |  |  |  | 25709 | PBSP | laureanjo@yahoo.com | 25 |  |  | Gx 4 | 808620 | For Pull out genexpert machine | Packing genexpert machine for pull out | Packing genexpert machine for pull out | Defective CPU |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5690834366028016049/5690834366028016049_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5690834366028016049/5690834366028016049_signature_31.png
| 2023-08-29 | 2023-08-25 15:00 | Angono Rural Health Unit | P. Tolentino St. | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 25708 | DOH | siennablanco@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 928 |  | GX-IV-D | 819864 | Module A4/719564 has error 2014/4012 | Run XpertCheck on module A4. | Performed maintenance Run XpertCheck on module A4 Generated reports | Unit is operational. Module A4 failed in XpertCheck. UPS not operational for replacement. | UPS SN: AT1000180910ST429 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5690831443611418836/5690831443611418836_signature_28.png | Rose Sienna Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5690831443611418836/5690831443611418836_signature_31.png
| 2023-08-29 | 2023-08-29 08:55 | Victoria Rural Health Unit | Brgy. Bulo Victoria Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25707 | DOH | diazrmt@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 895 | 2023-08-22 | GXIV-D | 831669 | A1 and A2 for module replacement | Module replacement Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Performed Module replacement Gathered reports Performed functionality test | Unit is fully operational |  | 2 genexpert modukes | 10516 and 10517 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5690801060227180068/5690801060227180068_signature_28.png | Alma D. Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5690801060227180068/5690801060227180068_signature_31.png
| 2023-08-28 | 2023-08-28 21:22 | Cuyapo Rural Health Unit | Cuyapo Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25706 | DOH | jolenerodico09171998@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1223 |  | GXIV-D | 831672 | A2:757048 temperature error A4:757050 temperature error | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is partially operational, A2 and A4 failed xpertcheck | Used 2 xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5690383870323034413/5690383870323034413_signature_28.png | Jolene Anthoinette Rodico | https://www.jotform.com/uploads/macarewecare/222927978372067/5690383870323034413/5690383870323034413_signature_31.png
| 2023-08-27 | 2022-12-23 09:30 | Lung Center of the Philippines | Quezon ave., Quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 25705 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | MCS + | 08H188 | Due for maintenance | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,checked program parameters,performed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5689486294483500686/5689486294483500686_signature_28.png | Janella madulid | https://www.jotform.com/uploads/macarewecare/222927978372067/5689486294483500686/5689486294483500686_signature_31.png
| 2023-08-27 | 2022-12-23 10:01 | Lung Center of the Philippines | Quezon ave.,QC | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 25704 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | MCS + | 02D166 | Due for maintenance | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,checked program parameters,performed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5689485384489626656/5689485384489626656_signature_28.png | Janella Madulid | https://www.jotform.com/uploads/macarewecare/222927978372067/5689485384489626656/5689485384489626656_signature_31.png
| 2023-08-27 | 2022-12-23 10:19 | Lung Center of the Philippines | Quezon avenue Q.C | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25703 | PBSP | lcpmicrolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4543 ASSAY: Xpert SARS-CoV-2, COUNT: 4520 |  | GXXVI | 801080 | B3 module encounter temp error | Checked unit, gathered reports,test run unit | Check up | 15 modules operational and b3 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5689484394486564351/5689484394486564351_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5689484394486564351/5689484394486564351_signature_31.png
| 2023-08-27 | 2022-12-24 13:59 | Makati Medical Center (medilinx) | Amorsolo st Makati City | NCR | NCR 4th - Makati | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25702 | COMMERCIAL | aiza.pelayo@medilinx.com.ph | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 100 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 452 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 484 ASSAY: Xpert BCR-ABL Ultra, COUNT: 1764 ASSAY: Others, COUNT: 2 ASSAY: Others, COUNT: 6 ASSAY: Xpert HCV VL, COUNT: 41 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1280 | 2023-08-23 | GX4 (10C) | 110001647 | Broken uniframe for module replacement B3 210066486 - 210061783 | Performed Module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | Operational |  | 900-0650 RS | 10550 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5689483654482144525/5689483654482144525_signature_28.png | Aiza G. Pelayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5689483654482144525/5689483654482144525_signature_31.png
| 2023-08-27 | 2022-12-25 15:37 | Casimiro Ynares Sr Memorial Hospital | Becco St., Burgos Rodriguez | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25701 | DOH | mariannegervacio1@gmail.com | 32 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 443 |  | GX4 | 831642 | Due for Maintenance | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5689482754487357847/5689482754487357847_signature_28.png | Marianne Gervacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5689482754487357847/5689482754487357847_signature_31.png
| 2023-08-26 | 2023-08-25 12:59 | Malapatan Rural Health Unit | Poblacion Malapatan, Sarangani Province | 12 | Sarangani | Laboratory | YES | MAINTENANCE Transfer Genexpert machine to new building/room | CEPHEID | Jesson Tupas |  |  |  |  | 25700 | DOH | kamilahnichol@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 400 |  | Gx-IV | 838419 | N/A | Genexpert machine Re-installation | Re-install and transfer of Genexpert machine to new building/room | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5688691856617342228/5688691856617342228_signature_28.png | Kamilah Nichol Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/5688691856617342228/5688691856617342228_signature_31.png
| 2023-08-25 | 2022-12-24 11:34 | St.lukes Medical Center Global City | Rizal drive,corner 5th and 32th Global City | NCR | NCR 4th - Taguig | Cellular Immunology | YES | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25699 | COMMERCIAL | biomedreports@stlukes.com.ph | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 152 ASSAY: Xpert CT/NG, COUNT: 1349 ASSAY: Xpert HBV VL, COUNT: 73 ASSAY: Xpert HPV, COUNT: 1051 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2235 ASSAY: Xpert HCV VL, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 39 | 2023-08-18 | GX4 | 839231 | For module replacement and maintenance C4 866803(def) - 210096427 (new) | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | C1 hw failed,c2,c3,c4 operational |  | 900-0508 | 96151 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5687668735914042953/5687668735914042953_signature_28.png | JOSHUA AYA-AY | https://www.jotform.com/uploads/macarewecare/222927978372067/5687668735914042953/5687668735914042953_signature_31.png
| 2023-08-25 | 2023-08-25 14:11 | Concepcion District Hospital | L.Cortes St. Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25698 | PBSP | cdh.laboratorydept@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1164 ASSAY: Xpert HIV-1 Viral Load, COUNT: 26 |  | GXIV-D | 807562 | A1 error 5011 | Troubleshooting Preventive maintenance | Performed preventive maintennce Installation of assay(CTNG) Archived test result Gathered reports Performed functionality test Jotform reporting training | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5687540983022860402/5687540983022860402_signature_28.png | Rowena Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5687540983022860402/5687540983022860402_signature_31.png
| 2023-08-25 | 2021-12-25 12:47 | Davao del Sur Provincial Hospital | Poblacion Digos City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25697 | PBSP | shammyj0308@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 334 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2306 |  | GX-IV 4D | 807464 | Module B2 for replacement | Replace the defective Module SN 672190 with new Module SN 210100861 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Module Replacement and Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5687498030012211803/5687498030012211803_signature_28.png | Sharmella Joy D. Pinote | https://www.jotform.com/uploads/macarewecare/222927978372067/5687498030012211803/5687498030012211803_signature_31.png
| 2023-08-25 | 2023-08-25 13:00 | Wenceslao Dela Paz Health Center | SB Park, Tandang Sora, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25696 | DOH | ameliafclemente2001@yahoo.com | 29 | ASSAY: Xpert HPV, COUNT: 492 ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2517 |  | GX IV 6C | 831876 | Module 2 error 2005 problem | Performed PM Checked module B2 B2 passed re-xpertcheck Advised end user to still use module Generated report. | Troubleshoot module B2 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5687492806971751135/5687492806971751135_signature_28.png | Camille Joy Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5687492806971751135/5687492806971751135_signature_31.png
| 2023-08-25 | 2022-12-23 16:40 | Bicol Regional General Hospital and Geriatric Medical Center | Camarines Sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25695 | PBSP | jlvdvera_rmt@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1820 |  | GX4 | 807601 | B4 613032 for module replacement | Module replaced: B4 613032>>210100893 | Replaced module B4 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5687485765427687312/5687485765427687312_signature_28.png | Jessie lowell de vera | https://www.jotform.com/uploads/macarewecare/222927978372067/5687485765427687312/5687485765427687312_signature_31.png
| 2023-08-25 | 2022-12-24 12:38 | Lagonoy Rural Health Unit | Lagonoy | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25694 | WHO | dyanrespeto@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1577 |  | GX4 | 846355 | CPU, UPS not working | Check power cables, ports, UPS Cleaned cartridge bays, plunger rod, icore Generate reports | Performed Troubleshooting Performed Preventive Maintenance | Unit is operational UPS bypass connection. To be reported by end-user to WHO |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5687483245426094566/5687483245426094566_signature_28.png | DYAN P. RESPETO | https://www.jotform.com/uploads/macarewecare/222927978372067/5687483245426094566/5687483245426094566_signature_31.png
| 2023-08-25 | 2023-08-25 11:06 | Dinagat District Hospital | Brgy. White Sand, Dinagat, Dinagat Islands | 13 | Surigao del Norte | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 25693 | PBSP | idhona@yahoo.com | 27.63 | ASSAY: Xpert MTB/RIF, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1145 |  | GX-4 | 807611 | None | None | Replacing Module B2 with SN: 609280 to SN: 210100425 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5687425983854554390/5687425983854554390_signature_28.png | Idhona N. Tillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5687425983854554390/5687425983854554390_signature_31.png
| 2023-08-25 | 2023-08-25 10:02 | Sta. Lucia Health Center | Sta. Lucia Ave, Sta. Lucia, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25692 | DOH | josiedingle2004@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1429 |  | GXIV 6C | 831901 | Module A4 Hardware failed problem | Performed PM Checked module A4 Module A4 failed re-xpertcheck Genereated report | Troubleshoot Module A4 | Unit is operational. Module A4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5687390026975354607/5687390026975354607_signature_28.png | Josephine L. Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5687390026975354607/5687390026975354607_signature_31.png
| 2023-08-25 | 2023-08-25 09:01 | Gerona Rural Health Unit | Brgy. Pinasling Geron Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25691 | DOH | huwakie15@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 22 ASSAY: Xpert MTB/RIF Ultra, COUNT: 577 | 2023-08-22 | GXIV-D | 831693 | A1 for module replacement A3 for re-xpertcheck | Module replacement Re-xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Performed module replacement Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | A3 passed calibration | 1 module 1 xpertcheck kit | 10519 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5687350343028861610/5687350343028861610_signature_28.png | Artemio I. Nool III | https://www.jotform.com/uploads/macarewecare/222927978372067/5687350343028861610/5687350343028861610_signature_31.png
| 2023-08-24 | 2023-08-23 12:49 | Magsingal District Hospital | Magsingal, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25690 | PBSP | jeanbuena08@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 289 |  | GX4 | 835190 | Due for Calibration | #NAME? | Performed XpertCheck | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686946265241797860/5686946265241797860_signature_28.png | Jean Michelle Buena | https://www.jotform.com/uploads/macarewecare/222927978372067/5686946265241797860/5686946265241797860_signature_31.png
| 2023-08-24 | 2023-08-23 15:31 | Bantay Municipal Health Office | Bantay, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25689 | PBSP | lornarabara_23@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 330 |  | GX4 | 807585 | Stuck Cartridge | #NAME? | Removed Stuck Cartridge | Unit is operational | Module A4 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686944265243585961/5686944265243585961_signature_28.png | Lorna R Rabara | https://www.jotform.com/uploads/macarewecare/222927978372067/5686944265243585961/5686944265243585961_signature_31.png
| 2023-08-24 | 2023-08-24 10:03 | Candon City Health Office | Candon, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25688 | DOH | candoncholaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 499 | 2023-07-24 | GX4 | 838273 | For Monitor Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Monitor CNK9430C2G >>>MMA1PSS00221507D963W01 - Generated Reports | Replaced Monitor | Unit is operational |  | Monitor MMA1PSS00221507D963W01 | DR 10246 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686943375246300458/5686943375246300458_signature_28.png | Michelle Bucalen | https://www.jotform.com/uploads/macarewecare/222927978372067/5686943375246300458/5686943375246300458_signature_31.png
| 2023-08-24 | 2023-08-23 08:38 | Bacarra Rural Health Unit | Bacarra, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25687 | PBSP | davidangelogamidobarrozo@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 60 | 2023-08-23 | GX4 | 807430 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module B4 606461 >>> 21010043 | Performed Module Replacement | Unit is operational |  | Module 210100443 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686941955247274954/5686941955247274954_signature_28.png | David Angelo Barrozo | https://www.jotform.com/uploads/macarewecare/222927978372067/5686941955247274954/5686941955247274954_signature_31.png
| 2023-08-24 | 2023-08-24 16:12 | International Organization for Migration | H.V Dela Costa,Makati | NCR | NCR 4th - Makati | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  |  |  |  | 25686 | COMMERCIAL | jsabalboro@iom.int | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 969 ASSAY: Others, COUNT: 139 ASSAY: Xpert SARS-CoV-2, COUNT: 18 |  | GX4 (10c) | 110004816 | For transfer in another location | Transfer and set up unit Test run unit | Transfer GX4 110002868, 805288 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686775737286603560/5686775737286603560_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5686775737286603560/5686775737286603560_signature_31.png
| 2023-08-24 | 2021-12-24 16:01 | Tomas Claudio (District A) Health Center | Corner E. Quirino - Camus Street, Davao City | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  | rkfagutao.macare@gmail.com |  |  | 25685 | DOH | arleen_p_alconera@yahoo.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 135 ASSAY: Xpert MTB/RIF Ultra, COUNT: 648 | 2023-08-02 | GX-IV 4D | 838492 | Due for Calibration | Run XPC Calibration | Perform XpertCheck Calibration | Module B2 for Service Required 3 Modules Passed the Calibration |  | Xpertcheck CE 5 | DR 10430 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686751706324224321/5686751706324224321_signature_28.png | Arleen P. Alconera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5686751706324224321/5686751706324224321_signature_31.png
| 2023-08-24 | 2023-08-24 16:11 | Balamban Rural Health Unit I | Balamban, Cebu | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25684 | DOH | climaco.mildred@yahoo.com | 24 |  |  | GXIV | 838231 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686742831229732674/5686742831229732674_signature_28.png | Mildred E. Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5686742831229732674/5686742831229732674_signature_31.png
| 2023-08-24 | 2023-08-24 16:14 | Balamban Rural Health Unit I | Balamban, Cebu | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25683 | DOH | climaco.mildred@yahoo.com | 24 |  |  | GXIV | 838564 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686742821223960483/5686742821223960483_signature_28.png | Mildred E. Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5686742821223960483/5686742821223960483_signature_31.png
| 2023-08-24 | 2023-08-24 15:07 | Bataan General Hospital and Medical Center | Balanga, Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25682 | PBSP | melodyjoy16@yahoo.com | 31 | ASSAY: Xpert HBV VL, COUNT: 66 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2489 ASSAY: Xpert HCV VL, COUNT: 77 ASSAY: Xpert HIV-1 Viral Load, COUNT: 214 |  | GXIV-D | 805124 | Due for Calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except B3 | B3:686088 Failed xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686713593592888244/5686713593592888244_signature_28.png | Melody Joy C. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5686713593592888244/5686713593592888244_signature_31.png
| 2023-08-24 | 2023-08-24 14:58 | Commonwealth Supern Health Center | Katuparan St., Commonwealth, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25681 | DOH | commonwealthclinical@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 426 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2425 | 2023-08-23 | GXIV 6C | 831615 | Module A3 and A4 for replacement | Performed PM Replaced Module A3 and A4 Generated report | Replaced module A3 and A4 A3: 625759 -> 210060068 A4: 754439 -> 798928 | Unit is operational. Module A2 for troubleshooting |  | GX modules 6c - 2pcs | 10548, 10549 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686703032326275015/5686703032326275015_signature_28.png | Rizza Clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5686703032326275015/5686703032326275015_signature_31.png
| 2023-08-24 | 2019-12-24 13:21 | Talomo North Health Center | S.I.R Matina, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25680 | PBSP | alfredojrsantos@gmail.com | 22.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1175 |  | GX-IV 4D | 815385 | Module B1 for replacement | Replaced the Defective Module SN 630927 with new Module SN 210100861 | Perform Module Replacement | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686695186321622557/5686695186321622557_signature_28.png | Alfredo Santos Jr., RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5686695186321622557/5686695186321622557_signature_31.png
| 2023-08-24 | 2023-08-24 13:46 | Tampakan Rural Health Unit | Poblacion Tampakan, South Cotabato | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25679 | DOH | jmedtech_lab@yahoo.com | 28.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 853 | 2023-08-24 | Gx-IV | 838544 | Module 1 for Service Required | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 3 operational modules |  | Xpertcheck CE-5 cartridge/kit | DR no.: 10414 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686671025022344964/5686671025022344964_signature_28.png | Jovi G. Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/5686671025022344964/5686671025022344964_signature_31.png
| 2023-08-24 | 2023-08-23 09:48 | Cotabato Provincial Hospital | Brgy. Amas, Kidapawan City, North Cotabato | 12 | Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25678 | PBSP | januarybing1206@gmail.com | 22.75 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 664 ASSAY: Xpert HIV-1 Viral Load, COUNT: 15 | 2023-08-23 | Gx-IV | 807615 | Module 4 H/W failed due to multiple temperature error Module 4 for Service Required | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration Cleaning and reset all cables and gatewayboard Report module problem | 3 operational modules | Module 1,2 and 3 passed the Xpertcheck/Calibration | Xpertcheck CE-5 cartridge/kit | Lot 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686651185029718574/5686651185029718574_signature_28.png | Hermosilla Enero | https://www.jotform.com/uploads/macarewecare/222927978372067/5686651185029718574/5686651185029718574_signature_31.png
| 2023-08-24 | 2023-08-24 13:12 | Maysan 3S Health Center | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 25677 | PBSP | neldimatera@gmail.com | 35 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1388 ASSAY: Xpert HIV-1 Viral Load, COUNT: 139 | 2023-08-24 | GXIV | 814151 | PSU and Gateway Board for replacement | #NAME? | Replaced parts | Unit is operational. |  | Gateway Board Power Supply Unit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686640847225259098/5686640847225259098_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/5686640847225259098/5686640847225259098_signature_31.png
| 2023-08-24 | 2023-08-24 11:34 | Cebu TB Reference Lab - SRP | Cebu City, SRP | 7 | Cebu | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE Installation From Cebu CHO to CTRL - SRP | CEPHEID | Raymond Laurito |  |  |  |  | 25676 | PBSP | tivalerio.pbsp@gmail.com | 27 |  |  | Gx 4 | 808620 | Defective CPU | Transfer from Cebu CHO to CTRL-SRP.. Installation using with same CPU and Monitor of Gx 838882.. connected with router | Transfer from Cebu CHO to CTRL-SRP.. Installation using with same CPU of Gx 838882.. connected with router | Systems operational (router connected gx 838882 & 808620) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686612005882477864/5686612005882477864_signature_28.png | Titus Ian Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5686612005882477864/5686612005882477864_signature_31.png
| 2023-08-24 | 2023-08-24 12:25 | Quirino Memorial Medical Center | J.P. Rizal, Project 4, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25675 | PBSP | carlopalomera58@gmail.com | 33 | ASSAY: Xpert CT/NG, COUNT: 11 ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2979 ASSAY: Xpert HIV-1 Viral Load, COUNT: 55 ASSAY: Others, COUNT: 21 | 2023-03-13 | GX IV 6C | 815484 | Unit for XpertCheck | Performed PM Performed XpertCheck Generated report | Performed XpertCheck | Unut is operational. All modules working | All modules passed XpertCheck | XpertCheck Kit | PL 9904 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686610382329577244/5686610382329577244_signature_28.png | Carlo Palomera | https://www.jotform.com/uploads/macarewecare/222927978372067/5686610382329577244/5686610382329577244_signature_31.png
| 2023-08-24 | 2023-08-24 11:57 | Subic Rural Health Unit | Manggahan, Subic, Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25674 | PBSP | salonga.jrs97@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 529 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2348 ASSAY: Xpert HIV-1 Viral Load, COUNT: 68 | 2023-08-24 | Genexpert IV | 812788 | Module 1 stuck cartridge | Manually open module door | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686593598011001427/5686593598011001427_signature_28.png | Jessa Rowe Salonga | https://www.jotform.com/uploads/macarewecare/222927978372067/5686593598011001427/5686593598011001427_signature_31.png
| 2023-08-24 | 2022-12-24 10:59 | Atang dela Rama Health Center | Pampanga St. Gagalangin | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25673 | DOH | monskie_22@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1420 |  | GXIV | 831644 | Unit for PM. For Xpertcheck of modules B1 and B3. | Performed PM. Performed Xpertcheck of B1 and B3. Generated reports. | PM and Xpertcheck. | Unit is operational. Modules B1 (756758) and B3 (614658), service required. | Used 2 spare Xpertcheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686560365213927567/5686560365213927567_signature_28.png | Monaliza Balcos | https://www.jotform.com/uploads/macarewecare/222927978372067/5686560365213927567/5686560365213927567_signature_31.png
| 2023-08-24 | 2023-08-24 09:33 | Tarlac City Health Center 2 | Brgy. Matatalaib, Tarlac City | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25672 | DOH | clarizannebeltran1996@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2023-08-22 | GXIV-D | 831917 | Module 4 and CPU for replacement UPS not backing up | Module and CPU replacement Preventive maintenance | Performed preventive maintenance Performed CPU installation and setup Performed module replacement Gathered reports Performed functionality test | Unit is operational | UPS for replacement | Genexpert module CPU | 10521 and 10518 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686506217313685056/5686506217313685056_signature_28.png | Clariz Anne Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5686506217313685056/5686506217313685056_signature_31.png
| 2023-08-24 | 2023-08-24 09:19 | Matalom Community Hospital | Matalom leyte | 8 | Leyte | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25671 | WHO | fegorne@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 139 ASSAY: Xpert MTB/RIF Ultra, COUNT: 275 | 2023-08-24 | GXIV | 847227 | - Module B2/910532 defective (for replacement) | #NAME? | - Replace module B2 w/ new SN: 210100730 - Performed Preventive maintenance - Performed Functionality test - Generate reports | Unit is Operational (waiting for LCD monitor replacement) |  | Genexpert Module 6 color SN: 210100730 | WHO |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5686495939424804898/5686495939424804898_signature_28.png | FE D. GORNE | https://www.jotform.com/uploads/macarewecare/222927978372067/5686495939424804898/5686495939424804898_signature_31.png
| 2023-08-23 | 2020-12-23 16:01 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25670 | DOH | dyrosemarie71@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1171 | 2023-08-03 | GX-IV 4D | 838401 | Due for Calibration and Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XPC Calibration in four modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration and operational |  | XPERTCHECK CE-5 | DR 10449 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685913080311004428/5685913080311004428_signature_28.png | Rosemarie B. Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/5685913080311004428/5685913080311004428_signature_31.png
| 2023-08-23 | 2023-08-23 16:54 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE Transfer Genexpert Unit | CEPHEID | Raymond Laurito |  |  |  |  | 25669 | PBSP | laureanjo@yahoo.com | 28 |  |  | GX 4 | 838882 | For Transfer to CTRL SRP | Packaging genexpert unit ready for deliver to SRP CTRL | Packaging genexpert unit ready for deliver to SRP CTRL | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685904870311985323/5685904870311985323_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5685904870311985323/5685904870311985323_signature_31.png
| 2023-08-23 | 2023-08-23 16:15 | Mandaluyong City Health Office | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25668 | PBSP | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2649 ASSAY: Xpert HIV-1 Viral Load, COUNT: 149 | 2023-08-23 | GXIV | 805062 | Module A1 for replacement | #NAME? | Replaced A1 | Unit is operational. |  | GeneXpert Module SN: 210101730 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685879729406712706/5685879729406712706_signature_28.png | Rosevi Jalova | https://www.jotform.com/uploads/macarewecare/222927978372067/5685879729406712706/5685879729406712706_signature_31.png
| 2023-08-23 | 2023-08-23 15:43 | Jose B. Lingad Memorial Hospital | San Fernando Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25667 | PBSP | valerioregine@yahoo.com | 30 | ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2448 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 21 |  | GXIV-D | 804763 | Monitor for replacement | Monitor replacement | Performed preventive maintenance Performed monitor replacement Archived test results Gathered reports Performed functionality test | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685874649322088694/5685874649322088694_signature_28.png | Regine V. Manalastas | https://www.jotform.com/uploads/macarewecare/222927978372067/5685874649322088694/5685874649322088694_signature_31.png
| 2023-08-23 | 2023-08-23 15:45 | Muñoz City Health Office | Science City of Muñoz, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 25666 | DOH | scomrtdl@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 694 | 2023-08-23 | Genexpert IV | 831673 | Module 1 error | Performed Xpertcheck on module 1 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck Result = Passed | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685868522839839658/5685868522839839658_signature_28.png | Ronalyn Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5685868522839839658/5685868522839839658_signature_31.png
| 2023-08-23 | 2023-08-23 15:03 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25665 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 351 ASSAY: Xpert MTB/RIF Ultra, COUNT: 328 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2133 | 2023-08-23 | GXIV | 807112 | Unit for XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. C1: passed C2: passed C3: passed C4: retry |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685840759408311225/5685840759408311225_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/5685840759408311225/5685840759408311225_signature_31.png
| 2023-08-23 | 2023-08-23 14:46 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25664 | PBSP | laboratory@loveyourself.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 351 ASSAY: Xpert MTB/RIF Ultra, COUNT: 328 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2133 | 2023-08-23 | GXIV | 808649 | Modules B2, B3, B4 for replacement. | #NAME? | Replaced modules | Unit is operational. |  | GeneXpert Module SN: 210101291, 210101170, 21011317 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685839779404209128/5685839779404209128_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/5685839779404209128/5685839779404209128_signature_31.png
| 2023-08-23 | 2023-08-23 15:07 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25663 | MACARE-HBDC | laboratory@loveyourself.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 351 ASSAY: Xpert MTB/RIF Ultra, COUNT: 328 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2133 | 2023-08-23 | GXIV | 807956 | Unit for Xpertcheck | #NAME? | Performed XpertCheck | Unit is operational. A1: failed A2: failed A3: failed A4: retry |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685839169404416604/5685839169404416604_signature_28.png | Raecel Lipang | https://www.jotform.com/uploads/macarewecare/222927978372067/5685839169404416604/5685839169404416604_signature_31.png
| 2023-08-23 | 2023-08-23 14:58 | SAMBOAN RURAL HEALTH UNIT | SAMBOAN CEBU | 7 | Cebu | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25662 | DOH | sheneannesenagan@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 86 ASSAY: Xpert MTB/RIF Ultra, COUNT: 200 | 2023-07-31 | GX-IV | 838513 | FOR MAINTENANCE AND CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK-CE-5 | 10352 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685833782132648950/5685833782132648950_signature_28.png | Shene Anne Señagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5685833782132648950/5685833782132648950_signature_31.png
| 2023-08-23 | 2023-08-23 14:05 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25661 | DOH | suzzettekho34@gmail.com | 31.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1477 | 2023-08-01 | GX-4 | 838554 | B3 found Service Required thru XPC | Reported to Cepheid | XpertCheck Calibration Preventive Maintenance | 3 modules working 1 module for replacement (Service Required) |  | XpertCheck, CE-IVD 5 Cartridges/Kit | 10387 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685803984225679212/5685803984225679212_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/5685803984225679212/5685803984225679212_signature_31.png
| 2023-08-23 | 2023-08-23 14:02 | Masambong Health Center | 8 Malac St., Masambong, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25660 | PBSP | michellebonrostro@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1296 | 2023-08-23 | GX IV 6C | 814160 | Old monitor have 2 lines and screen is flickering | Performed PM Replaced Monitor Generated report | Replaced monitor | Unit is operational. New monitor working. All modules working |  | GX Monitor Display | G 6154 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685802702326490057/5685802702326490057_signature_28.png | Michelle D. Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/5685802702326490057/5685802702326490057_signature_31.png
| 2023-08-23 | 2023-08-23 11:56 | Calinan Health Center | Calinan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25659 | PBSP | mailfegelacio@gmail.com | 24.5 |  |  | GX-IV 4D | 815390 | Infinite Windows Booting | Reset ram, hard disk drive and replace cmos battery | System Unit Troubleshooting | Hard disk failure error (3F0) | Unit is reported to Cepheid for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685733101317572810/5685733101317572810_signature_28.png | MAIL FE GELACIO | https://www.jotform.com/uploads/macarewecare/222927978372067/5685733101317572810/5685733101317572810_signature_31.png
| 2023-08-23 | 2023-08-23 11:57 | Philippine Tuberculosis Society Inc. | Quezon Institute Compound, E. Rodriguez Sr. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25658 | PBSP | sannymisorena@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 150 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1084 | 2023-08-22 | GXIV 6C | 707527 | Module A4 for replacement | Performed PM Replaced Module A4 Generated report | Replaced Module A4 621280 -> 210100954 | Unit is operational. All modules working |  | GX module 6Colors | G 6150 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685732262325018139/5685732262325018139_signature_28.png | sanny isorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5685732262325018139/5685732262325018139_signature_31.png
| 2023-08-22 | 2023-08-22 13:56 | Kidapawan City Health Office | Kidapawan City, North Cotabato | 12 | Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25657 | DOH | sg93rmt@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2170 | 2023-08-22 | Gx-IV | 838545 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules | Module 1 and 2 passed the Calibration | Xpertcheck CE-5 cartridge/kit | DR no.: 10407 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5685134635518841642/5685134635518841642_signature_28.png | Tessa Mae E. Saguindan | https://www.jotform.com/uploads/macarewecare/222927978372067/5685134635518841642/5685134635518841642_signature_31.png
| 2023-08-22 | 2023-08-22 14:59 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25656 | PBSP | laureanjo@yahoo.com | 29 |  |  | Gx 4 | 808620 | Defective CPU of 80820 gx machine | Switch hub connections 2 genexpert machine sn 838527 & 808620 | Switch hub connections 2 genexpert machine sn 838527 & 808620 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5684971656715655276/5684971656715655276_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5684971656715655276/5684971656715655276_signature_31.png
| 2023-08-22 | 2023-08-22 12:29 | Giporlos Rural Health Unit | Giporlos E. Samar | 8 | Eastern Samar | Tb-dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25655 | PBSP | ningrediang@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 532 | 2023-08-22 | GXIV-2D | 835060 | Module B2/787878 defective (for replacement) | #NAME? | - Module replacement - Install new module on B2 w/ SN 210100869 - Performed Preventive maintenance - Performed Functionality test - Generate reports | Unit is Operational |  | Genexpert Module 6 color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5684880922519242225/5684880922519242225_signature_28.png | Luningning R. Abuda | https://www.jotform.com/uploads/macarewecare/222927978372067/5684880922519242225/5684880922519242225_signature_31.png
| 2023-08-22 | 2023-08-22 11:56 | Overseas Filipino Workers Hospital | San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25654 | MACARE-HBDC | ofwhlaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 7 |  | GXIV-D | 110018556 | Printer ink cartridge | Troubleshooting | Performed preventive maintenance Clean ink cartridge Gathered reports Performed functionality test | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5684860283617280999/5684860283617280999_signature_28.png | Francisca Tiqui | https://www.jotform.com/uploads/macarewecare/222927978372067/5684860283617280999/5684860283617280999_signature_31.png
| 2023-08-22 | 2023-08-22 11:14 | Pioduran Rural Health Unit | Caratagan, Pioduran | 5 | Albay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25653 | DOH | mitzijc@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 591 |  | GX4 | 838264 | GX For Callibration | XpertCheck Result: all modules passed Cleaned cartridge bays, plunger rod, fan filter, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5684835758423578643/5684835758423578643_signature_28.png | Mitzi Cedeño | https://www.jotform.com/uploads/macarewecare/222927978372067/5684835758423578643/5684835758423578643_signature_31.png
| 2023-08-22 | 2023-08-22 11:06 | Bicol Medical Center | Naga City | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25652 | PBSP | cristinerod211@gmail.com | 30 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HPV, COUNT: 1499 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6495 ASSAY: Xpert SARS-CoV-2, COUNT: 2367 ASSAY: Xpert HIV-1 Viral Load, COUNT: 359 |  | GX16 | 842225 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, fan filter, icore Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5684832058422325795/5684832058422325795_signature_28.png | Cristine Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5684832058422325795/5684832058422325795_signature_31.png
| 2023-08-22 | 2023-08-22 09:02 | Jose Panganiban Rural Health Unit | Jose Panganiban | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25651 | PBSP | materesitapenas1@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1677 |  | GX4 | 808641 | B3/784530 for Module Replacement | Module replaced: B3 784530>>210100859 Cleaned cartridge bay, plunger rod, fan, fan filter, icore Generate reports | Replaced B3 module Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5684759098425378064/5684759098425378064_signature_28.png | Janela Joyce Esteves | https://www.jotform.com/uploads/macarewecare/222927978372067/5684759098425378064/5684759098425378064_signature_31.png
| 2023-08-22 | 2023-08-17 11:33 | Philippine heart center | East avenue q.c | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 25650 | MACARE-HBDC | baluyotjeric@gmail.com | 28 |  |  | MCS + | 20168 | As per end user they encounter perform self test error | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,checked program parameters,performed functionality test | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5684714642717864303/5684714642717864303_signature_28.png | Jeric Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5684714642717864303/5684714642717864303_signature_31.png
| 2023-08-22 | 2023-08-17 11:01 | Philippine heart center | East avenue q.c | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 25649 | MACARE-HBDC | baluyotjeric@gmail.com | 28 |  |  | MCS + | 11H204 | As per end user they encounter DPM out of tolerance | checked valve,rollers,line sensor,display panel,dpm,spm,cuff,checked program parameters,performed functionality test, adjust dpm and spm value | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5684714642719843149/5684714642719843149_signature_28.png | Jeric Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5684714642719843149/5684714642719843149_signature_31.png
| 2023-08-21 | 2023-08-18 12:59 | Baguio General Hospital | Governors pack rd., Baguio City | CAR | Benguet | Tb laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25648 | PBSP | dmgonzales.pbsp@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 338 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1670 |  | GX4 | 805212 | Due for maintenance and xpertcheck | Performed xpertcheck Cleaned unit Cleaned filter,icore,plunger rod,cartridge bay Gathered reports Test run unit | Maintenance and xpertcheck | B1 and B2,B4 failed in xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5684256997222404686/5684256997222404686_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5684256997222404686/5684256997222404686_signature_31.png
| 2023-08-18 | 2023-08-18 18:58 | Cebu TB Reference Laboratory | South Road Property, Cebu City | 7 | Cebu | Tb Laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25647 | DOH | asmantiquilla.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2390 |  | Gx4 | 838263 | Due for Calibration of the genexpert | Conduct calibration for B1 | Performed calibration of the genxpert | B1 Service Required and Module B2 encountered 4012 self test error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681659553916024168/5681659553916024168_signature_28.png | Allan Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5681659553916024168/5681659553916024168_signature_31.png
| 2023-08-18 | 2023-08-18 18:46 | MADRIDEJOS RURAL HEALTH UNIT | MADRIDEJOS, BANTAYAN ISLAND CEBU | 7 | Cebu | LABORATORY | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25646 | PBSP | cabasevilma@gmail.com | 0 |  |  | GX-IV | 835057 | SYTEM UNIT NO POWER | #NAME? | PERFORM SYSTEM UNIT SN: CPH1119F04296 TROUBLESHOOTING: - CHECK POWER SOURCE AND CABLE FUNCTIONALITY - INSPECT COMPONENTS | SYSTEM UNIT CONFIRMED DEFECTIVE |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681651013119243099/5681651013119243099_signature_28.png | Vilma Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5681651013119243099/5681651013119243099_signature_31.png
| 2023-08-18 | 2023-08-18 16:06 | St. Luke's Extension Clinic | Bocobo Street, Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Carlos Leniel Origenes |  |  |  |  | 25645 | MACARE-HBDC | laboratory@slec.ph | 23.2 | ASSAY: Xpert CT/NG, COUNT: 4921 ASSAY: Xpert MTB/RIF, COUNT: 138 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5134 | 2023-08-14 | GeneXpert Infinity | 809643 | Module A2 for Replacement | Replaced module 616088 with 220001426 | Performed PM(cleaning of icores, plungers, filters, and cartrodge bays) Generated reports Archived Results Performed module replacement Performed functionality test | Unit is operational | Work with Sir Patrick and Rhea | 900-0491RS GeneXpert Module 6C 220001426 | DR 10501 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681560736604997280/5681560736604997280_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5681560736604997280/5681560736604997280_signature_31.png
| 2023-08-18 | 2023-08-18 15:48 | BANTAYAN RURAL HEALTH UNIT | BINAOBAO BANTAYAN | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25644 | DOH | mcdelle916bausin@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 51 ASSAY: Xpert MTB/RIF Ultra, COUNT: 974 | 2025-07-27 | GX-IV | 838509 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  | XPERTCHECK-CE-5 | 10322 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681543709573392665/5681543709573392665_signature_28.png | Roxanne Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/5681543709573392665/5681543709573392665_signature_31.png
| 2023-08-18 | 2023-08-18 14:53 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | rHIVda | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25643 | MACARE-HBDC | jblmgh.is@gmail.com | 28 |  | 2023-08-18 | Genexpert IV | 110013748 | N/a | N/a | Installed Genexpert Installed PC Installed Printer Installed Drivers Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681515843414754809/5681515843414754809_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5681515843414754809/5681515843414754809_signature_31.png
| 2023-08-18 | 2023-08-18 13:55 | Bicol South Luzon Sub National Reference Laboratory | BRTTH Compound, Rizal St., Legazpi City | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25642 | PBSP | rbalzaga.pbsp@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 6 ASSAY: Others, COUNT: 507 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1463 ASSAY: Xpert HIV-1 Viral Load, COUNT: 197 |  | GX4 | 808636 | B3/862525 for Module Replacement | Replaced module: B3 862525>>210100862 Cleaned cartridge bays, plunger rod, fan, icore Generate reports | Replaced module B3 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681476970505818891/5681476970505818891_signature_28.png | Richard B. Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5681476970505818891/5681476970505818891_signature_31.png
| 2023-08-18 | 2023-08-18 13:22 | Polomolok Rural Health Unit | Poblacion, Polomolok, South Cotabato | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25641 | DOH | define.awkward90@gmail.com | 27.1 |  | 2023-08-18 | Gx-IV | 838265 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules | All modules passed the Xpertcheck/Calibration | Xpertcheck CE-5 cartridge/Kit | Dr no.: 10412 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681469870423940491/5681469870423940491_signature_28.png | Jethro Camaganacan | https://www.jotform.com/uploads/macarewecare/222927978372067/5681469870423940491/5681469870423940491_signature_31.png
| 2023-08-18 | 2023-08-18 13:34 | Zamboanga City Health Office | Pettit barracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25640 | PBSP | floresteresita25@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 604 | 2023-08-18 | GX 4 | 818902 | None | Module 1- Service required, module 2 Module 3 - passed, module 4 - passed. | Preventive Maintenence and xpertcheck calibration | Module 1, 2, 3 and 4 are available | Module 3 keeps on giving error 4009 and 1001 but passed in xpertcheck calibration | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681465050424389829/5681465050424389829_signature_28.png | Teresita Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5681465050424389829/5681465050424389829_signature_31.png
| 2023-08-18 | 2023-08-18 10:09 | Calumpang Rural Health Unit | Brgy. Calumpang General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25639 | DOH | jan19.cam@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1438 | 2023-08-18 | Gx-IV | 838504 | Module 1 and 4 -waiting for replacement module Module 2 service Required Module 3 multiple temperature error | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 1 operational modules | Xpertcheck/Calibration results module 2 service Required and module 3 passed | Xpertcheck CE-5 cartridge/kit | DR no.: 10404 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681453620424882805/5681453620424882805_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5681453620424882805/5681453620424882805_signature_31.png
| 2023-08-18 | 2023-08-18 10:09 | Calumpang Rural Health Unit | Brgy. Calumpang General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25638 | DOH | jan19.cam@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1438 | 2023-08-18 | Gx-IV | 838504 | Module 1 and 4 -waiting for replacement module Module 2 service Required Module 3 multiple temperature error | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 1 operational modules | Xpertcheck/Calibration results module 2 service Required and module 3 passed | Xpertcheck CE-5 cartridge/kit | DR no.: 10404 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681453390423981987/5681453390423981987_signature_28.png | Janice R Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5681453390423981987/5681453390423981987_signature_31.png
| 2023-08-18 | 2023-08-18 11:55 | Baguio General Hospital | Baguio City | CAR | Benguet | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25637 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 4 | 2023-08-16 | GX4 (10 c) | 110016112 | For module replacement b2 220007130 - 210065580, printer unable to print black color clogged black cartridge | Performed module replacement, gathered reports,checked and troubleshoot printer | Module replacement | Unit operational |  | 900-0650rs | G6135 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681406693701996436/5681406693701996436_signature_28.png | Syxerlyne Dang-ay | https://www.jotform.com/uploads/macarewecare/222927978372067/5681406693701996436/5681406693701996436_signature_31.png
| 2023-08-18 | 2023-08-18 11:27 | Jose B. Lingad Memorial Regional Hospital | City of San Fernando, Pampanga | 3 | Pampanga | Molecular Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25636 | MACARE-HBDC | cudiaaurita@gmail.com | 25 |  | 2023-08-17 | Genexpert IV | 110012825 | N/a | N/a | Installed Genexpert Installed PC Installed Printer Installed Drivers Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Package | 10511 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681393761016545163/5681393761016545163_signature_28.png | Aurita Cudio | https://www.jotform.com/uploads/macarewecare/222927978372067/5681393761016545163/5681393761016545163_signature_31.png
| 2023-08-18 | 2023-08-18 11:11 | Pablo N. Marquez Health & Diagnostic Center | Isok I, Marinduque | 4B | Marinduque | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25635 | PBSP | navarrolizabeth93@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1683 | 2023-08-11 | GX IV 6C | 807563 | Modules A1, A3 and A4 for replacement | Performed PM Replaced Modules Generated Report | Replaced A1, A3, A4 A1: 661909 -> 210100863 A3: 661907 -> 210100864 A4: 661906 -> 210100865 | Unit is operational. All modules working |  | 3 GX modules 6C | DR 0042504 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681383011819843788/5681383011819843788_signature_28.png | Lizabeth L.Navarro | https://www.jotform.com/uploads/macarewecare/222927978372067/5681383011819843788/5681383011819843788_signature_31.png
| 2023-08-18 | 2023-08-18 10:41 | TACLOBAN CITY HEALTH OFFICE | Tacloban City | 8 | Leyte | Tb-dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25634 | DOH | ammegzedd@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 97 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1235 | 2023-08-18 | GXIV | 838421 | Unit for Preventive Maintenance | N/A | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681359499197924418/5681359499197924418_signature_28.png | Gemma Malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5681359499197924418/5681359499197924418_signature_31.png
| 2023-08-18 | 2023-08-18 10:22 | Sto. Tomas City Health Office | Purok 7, San Miguel, Sto. Tomas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25633 | PBSP | nvsj12@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: New Laptop | 2023-08-16 | GXIV | 812786 | Unit for PM. For module replacement of module B4. For installation of backup laptop. | Performed PM. Performed installation of backup laptop. Performed module replacement on B4. Generated reports. | PM and installation. | Unit is operational. |  | GX module B4 (210100855) in replacment to defective B4 (805766). Backup Laptop Lenovo Thinkpad (CPH1019F03472) | G6132, G6129 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681351131514844839/5681351131514844839_signature_28.png | Victoria B. Menor | https://www.jotform.com/uploads/macarewecare/222927978372067/5681351131514844839/5681351131514844839_signature_31.png
| 2023-08-18 | 2023-08-18 09:29 | Marinduque Provincial Hospital | Santol, Boac, Marinduque | 4B | Marinduque | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25632 | PBSP | anarvas73@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 306 ASSAY: Xpert SARS-CoV-2, COUNT: 714 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 | 2023-07-27 | GXIV 6C | 835084 | Module B2 for replacement | Performed PM Replaced Module B2 Generated Report | Replaced module B2 210091205 -> 210100432 | Unit is operational. Both modules working |  | Gx Module 6 colors | DR 0042200 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681319185327760411/5681319185327760411_signature_28.png | Anthony Narvas | https://www.jotform.com/uploads/macarewecare/222927978372067/5681319185327760411/5681319185327760411_signature_31.png
| 2023-08-18 | 2023-08-18 09:12 | Naguilian District Hospital | Naguilian, La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25631 | DOH | marjelynrayray01@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 722 |  | GX4 | 838426 | For Maintenance | #NAME? | Performed Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681305315249398847/5681305315249398847_signature_28.png | Marjelyn Rayray | https://www.jotform.com/uploads/macarewecare/222927978372067/5681305315249398847/5681305315249398847_signature_31.png
| 2023-08-18 | 2023-08-18 09:08 | La Union Medical Center | Agoo, La Union | 1 | La Union | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25630 | PBSP | karizzamilomyg@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1863 |  | GX4 | 707400 | No modules detected | Reset IP address and Reconnect all LAN cables | Performed Troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681303435248957783/5681303435248957783_signature_28.png | Karizza Cyriene Milo | https://www.jotform.com/uploads/macarewecare/222927978372067/5681303435248957783/5681303435248957783_signature_31.png
| 2023-08-18 | 2023-08-11 12:07 | Northern Benguet District Hospital | Buguias, Benguet | CAR | Benguet | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25629 | PBSP | kym20nbdh@gmail.com | 26 |  | 2023-08-11 | GX4 | 814157 | Due for Computer Replacement Due for Calibration | - Replaced Defective Computer 2Ua7021 MPS >>> MJ0H94BW - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Performed XpertCheck - Generated Reports | Replaced Defective Computer Performed XpertCheck | Unit is operational | Module B3 failed, for reporting | GeneXpert Computer MJ0H94BW | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681298805244985181/5681298805244985181_signature_28.png | Kim Antonio | https://www.jotform.com/uploads/macarewecare/222927978372067/5681298805244985181/5681298805244985181_signature_31.png
| 2023-08-18 | 2023-08-11 16:38 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25628 | PBSP | benguetphortdl@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 |  | GX4 | 807621 | Due for troubleshooting | #NAME? | Performed Troubleshooting | Unit is operational | Module B3 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681296835242042294/5681296835242042294_signature_28.png | Frannie G. Calixto | https://www.jotform.com/uploads/macarewecare/222927978372067/5681296835242042294/5681296835242042294_signature_31.png
| 2023-08-18 | 2023-08-14 13:14 | Pangasinan Provincial Hospital | Bulingit, San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25627 | PBSP | judzfm@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 32 ASSAY: , COUNT: 2997 |  | GX4 | 808283 | Due for Calibration | #NAME? | Performed XpertCheck | Unit is operational | Module A2 failed; For reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681293585246982442/5681293585246982442_signature_28.png | Judy Ann Mabunga | https://www.jotform.com/uploads/macarewecare/222927978372067/5681293585246982442/5681293585246982442_signature_31.png
| 2023-08-18 | 2023-08-11 12:07 | Northern Benguet District Hospital | Buguias, Benguet | CAR | Benguet | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25626 | PBSP | kym20nbdh@gmail.com | 26 |  | 2023-08-11 | GX4 | 814157 | Due for Computer Replacement Due for Calibration | - Replaced Defective Computer 2Ua7021 MPS >>> MJ0H94BW - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Performed XpertCheck - Generated Reports | Replaced Defective Computer Performed XpertCheck | Unit is operational | Module B3 failed, for reporting | GeneXpert Computer MJ0H94BW | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681292625245577355/5681292625245577355_signature_28.png | Kim Antonio | https://www.jotform.com/uploads/macarewecare/222927978372067/5681292625245577355/5681292625245577355_signature_31.png
| 2023-08-18 | 2023-08-17 09:47 | Calasiao Rural Health Unit | Calasiao, Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25625 | PBSP | calasiaoclinicallab@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 536 | 2023-08-17 | GX4 | 834664 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods, I-CORE - Replaced Defective Module B4 786385 >>> 210100866 | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210100866 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681291515249193318/5681291515249193318_signature_28.png | John John Bauzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5681291515249193318/5681291515249193318_signature_31.png
| 2023-08-18 | 2023-08-14 14:35 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25624 | DOH | gbmmbg@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 565 |  | GX4 | 838437 | For Troubleshooting | #NAME? | Performed Troubleshooting | Unit is operational | Module 1 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5681290255242730738/5681290255242730738_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5681290255242730738/5681290255242730738_signature_31.png
| 2023-08-17 | 2023-08-17 13:37 | Malungon Rural Health Unit | Poblacion, Malungon, Sarangani Province | 12 | Sarangani | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25623 | DOH | madameerrmie@gmail.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 396 | 2023-08-17 | Gx-IV | 838546 | Module 2 and 3 h/w failed Module 4 service Required | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 2 operational modules | Xpertcheck result- Module 1 passed and module 4 service Required | Xpertcheck CE-5 cartridge/kit | DR no.: 10410 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680875158128953736/5680875158128953736_signature_28.png | Lois Caitlin Villacarlos | https://www.jotform.com/uploads/macarewecare/222927978372067/5680875158128953736/5680875158128953736_signature_31.png
| 2023-08-17 | 2023-08-17 19:50 | Buhangin Health Center | Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25622 | DOH | buhanginlaboratory@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 914 | 2023-08-02 | GX-IV 4D | 838232 | Due for Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Run XPC Calibration in four modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration |  | Xpertcheck CE 5 | DR 10429 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680824138905043637/5680824138905043637_signature_28.png | Alyssa Thea Santos RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5680824138905043637/5680824138905043637_signature_31.png
| 2023-08-17 | 2023-08-17 16:09 | San Isidro Rural Health Unit | Poblacion, San Isidro, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25621 | PBSP | urian_marlyn@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 802 | 2023-08-17 | Genexpert IV | 805107 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 714934 - 210100709 | Genexpert Module = 210100709 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680695854024218382/5680695854024218382_signature_28.png | Marlyn M. Urian | https://www.jotform.com/uploads/macarewecare/222927978372067/5680695854024218382/5680695854024218382_signature_31.png
| 2023-08-17 | 2023-08-17 14:29 | Medellin Rural Health Unit | Medellin Poblacion Cebu | 7 | Cebu | Tb Laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25620 | DOH | medellindtc@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1353 |  | Gx4 | 838506 | Due for Calibration Module B1 and B3 | Performed calibration for B1 an B3 | Performed calibration in B1 and B3 | B1 and B3 Failed during calibration, B2 and Waiting for Module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680634076411327761/5680634076411327761_signature_28.png | Raul Galano, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5680634076411327761/5680634076411327761_signature_31.png
| 2023-08-17 | 2023-08-17 13:24 | Candaba Rural Health Unit | Poblacion, Candaba, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25619 | DOH | rochelannlontoc@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 446 | 2023-07-20 | Genexpert IV | 831873 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 663001 - 210034408 | Genexpert Module = 210034408 | 10210 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680595673401691759/5680595673401691759_signature_28.png | Rochel Ann Lontoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5680595673401691759/5680595673401691759_signature_31.png
| 2023-08-17 | 2023-08-17 13:09 | A.H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 25618 | PBSP | dinajaen1990@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1880 |  | GXXVI | 839292 | Module B1, B2, B3, C1, C4 for troubleshooting | #NAME? | Performed XpertCheck on modules | Unitis operational. ReXC results: all said modules failed. For reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680583945196463676/5680583945196463676_signature_28.png | Dina Jaen | https://www.jotform.com/uploads/macarewecare/222927978372067/5680583945196463676/5680583945196463676_signature_31.png
| 2023-08-17 | 2023-08-17 12:22 | Region 1 Medical Center | Dagupan City, Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25617 | PBSP | r1mcxpertlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1499 ASSAY: Xpert SARS-CoV-2, COUNT: 1764 |  | GXIV-D | 738529 | B1 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is fully operational | Work with Engr. Yasser Salih |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680559230948649366/5680559230948649366_signature_28.png | Kirstin Banaag | https://www.jotform.com/uploads/macarewecare/222927978372067/5680559230948649366/5680559230948649366_signature_31.png
| 2023-08-17 | 2023-08-17 10:12 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25616 | DOH | renlizros11@gmail.com | 23.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 765 | 2024-04-28 | Gx-IV | 838570 | Module 2 and 3 for Service Required | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 2 operational modules |  | Xpertcheck CE-5 cartridge/kit | DR no.: 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680496918122028664/5680496918122028664_signature_28.png | Renee Liz M. Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5680496918122028664/5680496918122028664_signature_31.png
| 2023-08-17 | 2023-08-17 10:41 | Cong. Lamberto L. Macias Memorial Hospital | Siaton, Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25615 | DOH | honiemacaday619@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 511 |  | GX-IV | 838482 | - Due for calibration/ Xpertcheck. - Module 1 with temperature -related errors. | #NAME? | Performed Xpertcheck in all 4 modules. | module 1 failed XPC, modules 2, 3 4 passed Xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680496372716694460/5680496372716694460_signature_28.png | Honie Dy G. Macaday, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5680496372716694460/5680496372716694460_signature_31.png
| 2023-08-17 | 2023-08-17 10:35 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25614 | PBSP | renlizros11@gmail.com | 23.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 599 ASSAY: Xpert HIV-1 Viral Load, COUNT: 33 | 2023-08-17 | Gx-IV | 815428 | Module 4 for Service Required | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 3 operational modules | Module 1, 2 and 3 passed the Xpertcheck/Calibration | Xpertcheck CE-5 cartridge/kit | Lot 1000725070 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680495058128620149/5680495058128620149_signature_28.png | Renee Liz M. Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5680495058128620149/5680495058128620149_signature_31.png
| 2023-08-17 | 2023-08-17 10:23 | Villareal Rural Health Unit | Villareal samar | 8 | Samar | Tb-dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25613 | PBSP | lourdenmaynavallo@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 496 | 2023-08-17 | GXIV-2D | 834823 | Due for Calibration | #NAME? | #NAME? | Unit is Operational |  | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680485221145947777/5680485221145947777_signature_28.png | Arman l navallo | https://www.jotform.com/uploads/macarewecare/222927978372067/5680485221145947777/5680485221145947777_signature_31.png
| 2023-08-17 | 2023-08-16 14:22 | Malapatan Rural Health Unit | Poblacion, Malapatan, Sarangani | 12 | Sarangani | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25612 | DOH | rhu.malapatan12@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 379 | 2023-08-16 | Gx-IV | 838419 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules | All modules passed the Xpertcheck/Calibration | Xpertcheck CE-5 cartridge/ kit | DR no.: 10409 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680475028122540621/5680475028122540621_signature_28.png | Ceres Mae Jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/5680475028122540621/5680475028122540621_signature_31.png
| 2023-08-17 | 2023-08-17 09:59 | Best Diagnostic | Masikap st,qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25611 | COMMERCIAL | valbuenashy@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 72 ASSAY: Xpert HBV VL, COUNT: 90 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GX4 | 110008529 | Due for maintenance | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680472857403540157/5680472857403540157_signature_28.png | EDGINICA SHANNEN VALBUENA | https://www.jotform.com/uploads/macarewecare/222927978372067/5680472857403540157/5680472857403540157_signature_31.png
| 2023-08-17 | 2023-08-17 09:25 | Moncada Rural Health Unit | Moncada Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25610 | PBSP | moncadarhulaboratory.2023@gmail.com | 29 |  |  | GXIV-D | 808292 | UPS Moduke B2 for troubleshooting | Troubleshooting Preventive maintenance | Performed preventive maintenance Performed cepheid troubkeshooting procedure Archived test results Gathered reports Performed functionality test | Unit is oartially operational | Defective UPS |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680456600943441616/5680456600943441616_signature_28.png | Jessa Pabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5680456600943441616/5680456600943441616_signature_31.png
| 2023-08-17 | 2023-08-17 09:24 | Liloy rural health unit | Liloy, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25609 | PBSP | gisellefipulab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 949 | 2023-08-02 | GX4 | 838436 | Module 3 temperature error | Module 1, 2, 3 and 4 undergone Xpertcheck calibration and Preventive Maintenence. Module 1 passed, module 2 passed, module 3 service required, module 4 passed | Xpertcheck calibration and Preventive Maintenence | Module 1, 2 and 4 are available, while module 3 is not in use. | None | XPERTCHECK-CE-5, CE-IVD 5 cartridges/kit 1 ( 1000410214) | 10396 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680450278125883721/5680450278125883721_signature_28.png | Giselle Gipulan | https://www.jotform.com/uploads/macarewecare/222927978372067/5680450278125883721/5680450278125883721_signature_31.png
| 2023-08-17 | 2023-08-17 08:32 | Ramos Rural Health Unit | Poblacion South, Ramos, Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25608 | DOH | jadeeh8143@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 388 | 2023-07-26 | GXIV-D | 831721 | Defective UPS | UPS Replacement Preventive maintenance | Performed preventive maintenance Performed UPS replacement Archived test results Gathered reports Performed functionality test | Unit is fully operational |  | UPS | 10294 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680426535936784415/5680426535936784415_signature_28.png | Emily Jade Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/5680426535936784415/5680426535936784415_signature_31.png
| 2023-08-17 | 2023-08-17 07:35 | Camalig Rural Health Unit | Brgy. 2, Camalig | 5 | Albay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25607 | PBSP | leilanigrageda@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 900 |  | GX4 | 835187 | GX for Callibration | XpertCheck Result: All modules passed Cleaned cartridge bays, plunger rod, fan filter, icore Generate reports | Performed XpertCheck | Unit is operational B1 - Passed B2 -Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680389853509898030/5680389853509898030_signature_28.png | Leilani M Grageda | https://www.jotform.com/uploads/macarewecare/222927978372067/5680389853509898030/5680389853509898030_signature_31.png
| 2023-08-16 | 2023-08-16 22:21 | Negros Oriental Provincial Hospital | Dumaguete City Negros Oriental | 7 | Negros Oriental | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25606 | DOH | rezelmarielagos@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2540 ASSAY: Xpert HIV-1 Viral Load, COUNT: 157 |  | GX-IV | 838525 | Due for Xpertcheck/Calibration | #NAME? | Performed Xpertcheck on modules 2 and 4. | 4 modules operational, modules 2 and 4 passed Xpertcheck. Already done XPC with modules 1 and 3. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680051930514163693/5680051930514163693_signature_28.png | Rezel Marie Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5680051930514163693/5680051930514163693_signature_31.png
| 2023-08-16 | 2023-08-16 22:16 | Negros Oriental Provincial Hospital | Dumaguete City, Negros Oriental | 7 | Negros Oriental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25605 | PBSP | rezelmarielagos@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2540 ASSAY: Xpert HIV-1 Viral Load, COUNT: 157 |  | GX-IV | 815433 | Defective module D2/757101 | #NAME? | Replaced module D2 with new module with SN 210100867. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680048130513104919/5680048130513104919_signature_28.png | Rezel Marie Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5680048130513104919/5680048130513104919_signature_31.png
| 2023-08-16 | 2023-08-16 22:08 | Negros Oriental Provincial Hospital | Dumaguete City, Negros Oriental | 7 | Negros Oriental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25604 | PBSP | rezelmarielagos@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2540 ASSAY: Xpert HIV-1 Viral Load, COUNT: 157 |  | GX-IV | 805275 | Defective CPU | #NAME? | Replaced defective CPU with new CPU with SN CPH0622F37114. | Unit Operational, 4 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5680044990514199756/5680044990514199756_signature_28.png | Rezel Marie Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5680044990514199756/5680044990514199756_signature_31.png
| 2023-08-16 | 2023-08-16 20:30 | National children's hospital | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 25603 | MACARE-HBDC | zarabellehilario@yahoo.com | 28 |  |  | MCS + | 17H632SPG | Due for maintenance | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,checked program parameters,performed functionality test | Maintenance | For quotation of rollers need to be replace due to wear and tear |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679989005616804287/5679989005616804287_signature_28.png | Zarabelle Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5679989005616804287/5679989005616804287_signature_31.png
| 2023-08-16 | 2023-08-16 20:08 | Cebu City Jail Male Dorm | Kalunasan, Cebu City | 7 | Cebu | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25602 | PBSP | janalbrienevendiola@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1437 |  | GX4 | 805214 | Due for Calibration and B1 encountered Hardware failure due to error of 4017 during self test | Performed calibration in genexpert machine | Performed maintenance, clean Icore, cartridge bay, reset module and calibrate all modules | B1 - Hardware failure and failed during calibration B2- Failed during calibration B3- Passed B4-Failed during calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679973983919174002/5679973983919174002_signature_28.png | Jan Albriene Vendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5679973983919174002/5679973983919174002_signature_31.png
| 2023-08-16 | 2023-08-16 17:03 | District C (Mini Forest) Health Center | Quezon Boulevard, Barangay 23-C, Davao City | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25601 | DOH | miniforestlaboratory@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF, COUNT: 174 ASSAY: Xpert MTB/RIF Ultra, COUNT: 327 | 2023-08-03 | GX-IV 4d | 838507 | Due for Calibration | Run XpertCheck on Three Modules | Perform XpertCheck Calibration | All Modules Passed the Calibration |  | Xpertcheck CE 5 | DR 10437 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679860164819538903/5679860164819538903_signature_28.png | Joy Nio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5679860164819538903/5679860164819538903_signature_31.png
| 2023-08-16 | 2023-08-16 13:56 | Department of Health Bicutan | Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25600 | PBSP | dohtrcbicutan.lab@yahoo.com | 25 |  |  | GXIV | 807452 | Unit for relocation | #NAME? | Relocated unit to nearby work desk | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679840693619856904/5679840693619856904_signature_28.png | Joanna Carlota Rufino | https://www.jotform.com/uploads/macarewecare/222927978372067/5679840693619856904/5679840693619856904_signature_31.png
| 2023-08-16 | 2023-08-16 16:13 | Quezon City General Hospital | Seminary Road, Edsa, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25599 | MACARE-HBDC | treatmenthub.qcgh@gmail.com | 27 |  | 2023-08-09 | GXIV 10C | 110012766 | GX unit for installation | Unboxed GX machine and accessories Setup GX machine and accessories Generated report | Installed GX and accessories | Unit is operational. |  | GX machine, UPS, CPU, Barcode Scanner, Printer, Monitor | DR 10487 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679833207429190699/5679833207429190699_signature_28.png | Milma B. Laxina | https://www.jotform.com/uploads/macarewecare/222927978372067/5679833207429190699/5679833207429190699_signature_31.png
| 2023-08-16 | 2023-08-15 13:20 | Apalit Rural Health Unit | Apalit Pampanga | 3 | Pampanga | Laboraory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25598 | DOH | duenasblesilda@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 835 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 |  | GXIV-D | 831900 | A1 error 1006,1001,4011,4017 | Troubleshooting Preventive maintenance | Performed preventive maintenance Archived test results Gathered reports Performed functionality test | Unit is partially operational | A1 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679769139413058324/5679769139413058324_signature_28.png | Blesilda G. Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5679769139413058324/5679769139413058324_signature_31.png
| 2023-08-16 | 2023-08-16 14:03 | Mexico Rural Health Unit IV | Pandacaqui, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 25597 | DOH | annbathan79@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 512 | 2023-08-16 | Genexpert IV | 831902 | Module 3 error | Performed Xpertcheck on Module 3 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck Result = Passed | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679758499326822101/5679758499326822101_signature_28.png | Mary Ann D Carreon | https://www.jotform.com/uploads/macarewecare/222927978372067/5679758499326822101/5679758499326822101_signature_31.png
| 2023-08-16 | 2023-08-16 14:10 | Reproductive Health and Wellness Center | Jacinto Street, Davao City | 11 | Davao del Sur | TB-HIV | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25596 | DOH | rhwclab@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 144 ASSAY: Xpert HIV-1 Viral Load, COUNT: 919 | 2023-08-03 | GX IV 4D | 838262 | Due for Calibration Module B2 encounter temperature related Error | Perform XPC Calibration and Functionality Test for Module B2 | Troubleshoot Module B2 | All Modules Passed the Calibration | Module B2 for under observation | XPERTCHECK CE 5 | DR 10445 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679756841413713350/5679756841413713350_signature_28.png | Ellaine Malano RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5679756841413713350/5679756841413713350_signature_31.png
| 2023-08-16 | 2023-08-16 13:17 | Candijay Community Hospital | Poblacion, Candijay | 7 | Bohol | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25595 | DOH | senenmilo24@gmail.com | 22.125 | ASSAY: Xpert MTB/RIF, COUNT: 265 ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 | 2023-07-27 | GX-4 | 838396 | B1 found Service Required B4 reported through Jotform | B4 reported through jotform | XpertCheck Calibration Preventive Maintenance | 3 modules working 1 module reported | B1 for replacement due to XPC Result | XpertCheck, CE-IVD 5 Cartridges/Kit | 10324 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679729745029305750/5679729745029305750_signature_28.png | Lian Milo | https://www.jotform.com/uploads/macarewecare/222927978372067/5679729745029305750/5679729745029305750_signature_31.png
| 2023-08-16 | 2023-08-16 12:38 | Sarangani Health Care Facility | Alabel, Sarangani Province | 12 | Sarangani | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25594 | DOH | ronayadao@yahoo.com | 20.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 704 | 2023-08-16 | Gx-IV | 838542 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/Calibration | 4 operational modules |  | Xpertcheck CE-5 cartridge/kit 1000410214 | DR no. 10413 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679727758125813910/5679727758125813910_signature_28.png | Ronora Yadao | https://www.jotform.com/uploads/macarewecare/222927978372067/5679727758125813910/5679727758125813910_signature_31.png
| 2023-08-16 | 2023-08-16 12:53 | Leyte Provincial Hospital | Palo, campetic | 8 | Leyte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25593 | PBSP | lphlabdeptcovid19@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 599 ASSAY: Xpert HPV, COUNT: 8 | 2023-08-16 | GXIV | 815267 | - CPU error (operating system not booting) SN / 2UA7021M0Z | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679714940326620287/5679714940326620287_signature_28.png | Ameryl Mostales | https://www.jotform.com/uploads/macarewecare/222927978372067/5679714940326620287/5679714940326620287_signature_31.png
| 2023-08-16 | 2023-08-16 11:46 | Gov. Celestino Gallares Memorial Hospital | Tagbilaran, Bohol | 7 | Bohol | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25592 | DOH | majieocmen@gmail.com | 27 |  |  | Gx 4 | 820228 | Duedate for xpertcheck module A1 & A3 | Xpertcheck calibration module A1 & A3 | Xpertcheck calibration module A1 & A3 | Module a1 & a3 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679670621716799164/5679670621716799164_signature_28.png | Ma. Jernai Sosoban | https://www.jotform.com/uploads/macarewecare/222927978372067/5679670621716799164/5679670621716799164_signature_31.png
| 2023-08-16 | 2023-08-16 11:36 | Gov. Celestino gallares memorial hospital | Tagbilaran, bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 25591 | PBSP | majieocmen@gmail.com | 27 |  |  | Gx 4 | 838878 | Defective module b3 sn 634295 | New module b3 installed sn 210100707 | New module b3 installed sn 210100707 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679667245224743380/5679667245224743380_signature_28.png | Ma. Jernai Sosoban | https://www.jotform.com/uploads/macarewecare/222927978372067/5679667245224743380/5679667245224743380_signature_31.png
| 2023-08-16 | 2023-08-16 11:26 | Escalante City Health Office | Escalante City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25590 | DOH | sheila.negrido@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1070 |  | GX4 | 838523 | None | Conduct PM and generate files. | Run Xpert check calibration for module B2 and B4 only. | Module B2 and B4 passed the calibration. All modules are in good conditions. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679666254308930189/5679666254308930189_signature_28.png | Sheila Mae Negrido | https://www.jotform.com/uploads/macarewecare/222927978372067/5679666254308930189/5679666254308930189_signature_31.png
| 2023-08-16 | 2023-08-16 10:19 | Teodoro B. Galagar District Hospital | Poblacion, Jagna | 7 | Bohol | RTDL | NO | CALIBRATION | CEPHEID | Faizar Avenido |  |  |  |  | 25589 | DOH | applecpalma09@yahoo.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 230 ASSAY: Xpert MTB/RIF Ultra, COUNT: 557 | 2023-07-31 | GX-4 | 838292 | B2 Found Service Required | None | XpertCheck Calibration Preventive Maintenance | All modules are operational | B2 - Service Required (For Replacement) | Xpertcheck, CE-IVD 5 Cartridges/Kit | 10354 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679649269412219806/5679649269412219806_signature_28.png | Therese Arjelin C. Palima | https://www.jotform.com/uploads/macarewecare/222927978372067/5679649269412219806/5679649269412219806_signature_31.png
| 2023-08-16 | 2023-08-16 11:06 | Caloocan City Quality Assurance Center RTDL | Sangandaan, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25588 | PBSP | caloocancentrallaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 180 ASSAY: Xpert HIV-1 Viral Load, COUNT: 41 |  | GXIV 6C | 805271, 814154 | CPU of GX with SN 814154 is for replacement | Connected two GX in network switch Setup printer Generated report | Temporarily connected 2 GX to 1 CPU using network switch | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679649155758884088/5679649155758884088_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5679649155758884088/5679649155758884088_signature_31.png
| 2023-08-16 | 2023-08-16 10:46 | Jose dalman rural health unit | Jose dalman, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | CALIBRATION MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25587 | PBSP | crizzycristina@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 346 | 2023-08-16 | GX2 | 834827 | Module 1 hardware failed | Replaced module 787266 with Module 210100394. Module 2 undergone Xpertcheck calibration and passed | Preventive maintenence, Module 1 Replacement and xpertcheck calibration | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679638582326975536/5679638582326975536_signature_28.png | Ma Cristina M. Alquiza | https://www.jotform.com/uploads/macarewecare/222927978372067/5679638582326975536/5679638582326975536_signature_31.png
| 2023-08-16 | 2023-08-16 10:46 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25586 | DOH | luzvimindanavidad2@gmail.com | 26 |  |  | GXIV-4D | 838297 | None | None | #NAME? | All modules are functional. Module B4 for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679634588466534370/5679634588466534370_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5679634588466534370/5679634588466534370_signature_31.png
| 2023-08-16 | 2023-08-16 10:43 | University of the Philipines - National Institutes of Health | G/F National Institutes of Health Building, Pedro Gil Street, Ermita | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25585 | COMMERCIAL | nihcentrallab@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 12 ASSAY: Xpert HBV VL, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 ASSAY: Xpert HCV VL, COUNT: 13 ASSAY: Xpert HIV-1 Viral Load, COUNT: 124 |  | GX16-10C | 110014186 | Unit for PM. Module D2, H/W Failure. | Performed PM. Performed troubleshooting. Performed selftest on D2. Generated reports. | Preventive maintenance and troubleshooting. | Unit is operational. Module D2 still has H/W Failure and has an error of 4017. Module D2 (210070086) for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679634065416798144/5679634065416798144_signature_28.png | Mary Grace Hernaez | https://www.jotform.com/uploads/macarewecare/222927978372067/5679634065416798144/5679634065416798144_signature_31.png
| 2023-08-16 | 2023-08-16 09:48 | Macare In-House | Quezon City | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR TROUBLESHOOTING | OTHERS | Ronnel Belizar |  |  |  | Lenovo Desktop | 25584 | MACARE-HBDC | rcb.macare@gmail.com | 25 |  |  | ThinkCentre M720s | CPH0520F08201 | For repair | #NAME? | Performed torubleshooting | Unit is operational. Unit is ready for dispatch. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679599223616958099/5679599223616958099_signature_28.png | Patrick Los Banos | https://www.jotform.com/uploads/macarewecare/222927978372067/5679599223616958099/5679599223616958099_signature_31.png
| 2023-08-16 | 2023-08-16 09:45 | Bindoy District Hospital | Bindoy, Negros Oriental | 7 | Negros Oriental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25583 | PBSP | bindoydistricthosp.laboratory@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 162 ASSAY: Xpert MTB/RIF Ultra, COUNT: 747 |  | GX4 | 838475 | None | #NAME? | - Perform Xpertcheck on 4 modules | 4 modules passed Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679598889128958210/5679598889128958210_signature_28.png | Lemuel Quirit, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5679598889128958210/5679598889128958210_signature_31.png
| 2023-08-16 | 2023-08-16 09:43 | Minglanilla Rural Health Unit | Ward 1, Pider Minglanilla Cebu | 7 | Cebu | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25582 | DOH | brendagonzagaparedes@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1437 |  | GX4 | 838512 | Due for calibration of genexpert unit and B3 encountered a problem of 4017 during start up | Conduct calibration for B1, B3 and B4 | Performed calibration of genexpert machine | B1 and B4 passed calibration and B3 failed during calibration for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679597463915955208/5679597463915955208_signature_28.png | Brenda Paredes , RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5679597463915955208/5679597463915955208_signature_31.png
| 2023-08-16 | 2023-08-16 09:45 | Macare In-House | Quezon City | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR TROUBLESHOOTING | OTHERS | Ronnel Belizar |  |  |  | Lenovo Desktop | 25581 | MACARE-HBDC | rcb.macare@gmail.com | 25 |  |  | ThinkCentre M720s | CPH0420F08168 | For repair | #NAME? | Performed troubleshooting | Unit is operational. Ready for dispatch. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679596693612030257/5679596693612030257_signature_28.png | Patrick Los Banos | https://www.jotform.com/uploads/macarewecare/222927978372067/5679596693612030257/5679596693612030257_signature_31.png
| 2023-08-16 | 2023-08-16 09:29 | San Lazaro Hospital | Quiricada St. Sta Cruz | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25580 | PBSP | acong.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 850 ASSAY: Xpert HIV-1 Viral Load, COUNT: 600 | 2023-08-16 | GXIV | 707843 | Unit for PM. For module replacement of modules C1 and C4. | Performed PM. Performed module replacement on C1 and C4. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module. C1 (210100881) and C4 (210100884) in replacement to defective C1 (662707) and C4 (699744). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679587861122735011/5679587861122735011_signature_28.png | Ashley Cynna Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5679587861122735011/5679587861122735011_signature_31.png
| 2023-08-16 | 2023-08-16 09:20 | AGUSAN DEL NORTE PROVINCIAL HEALTH OFFICE | BUTUAN | 13 | Agusan del Norte | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25579 | PBSP | ronniebalansag74@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 31 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1710 |  | GX-IV | 807461 | FOR MODULE A1 REPLACEMENT | #NAME? | -PERFORM GX EQUIPMENT PM AND CLEANING -PERFORM MODULE A1 DEF SN : 210068636 REPLACEMENT WITH NEW SN: 210100555 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679582895676833775/5679582895676833775_signature_28.png | Ronnie Balansag | https://www.jotform.com/uploads/macarewecare/222927978372067/5679582895676833775/5679582895676833775_signature_31.png
| 2023-08-15 | 2023-08-15 19:16 | BISLIG DISTRICT HOSPITAL | BISLIG SURIGAO DEL SUR | 13 | Surigao del Sur | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25578 | PBSP | bdhlaboratory2021@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2368 |  | GX IV | 805280 | FOR MODULE B4 REPLACEMENT | #NAME? | -PERFORM GX MAINTENANCE AND CLEANING -PERFORM MODULE B4 REPLACEMENT DEF SN: 674605 WITH NEW SN: 210100329 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679076585815829380/5679076585815829380_signature_28.png | Georgina V. Novero | https://www.jotform.com/uploads/macarewecare/222927978372067/5679076585815829380/5679076585815829380_signature_31.png
| 2023-08-15 | 2023-08-15 16:40 | Bicol Medical Center | Panganiban Drive, Naga City | 5 | Camarines Sur | Microbiology Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25577 | MACARE-HBDC | cristinerod211@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 566 ASSAY: Xpert MTB/RIF Ultra, COUNT: 460 ASSAY: Xpert SARS-CoV-2, COUNT: 393 |  | GX4 | 110008458 | B3/210034756 for Module Replacement | Replaced module: B3 210034756>>635083 Cleaned cartridge bays, plunger rod, fan, fan filter, icore Generate reports | B3 module replaced Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679022242614776599/5679022242614776599_signature_28.png | Cristine Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5679022242614776599/5679022242614776599_signature_31.png
| 2023-08-15 | 2023-08-15 16:25 | San Leonardo Rural Health Unit | Diversion, San Leonardo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25574 | PBSP | ligayadomingo20@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 ASSAY: Xpert MTB/RIF, COUNT: 50 | 2023-08-15 | Genexpert IV | 846042 | Computer not turning on UPS not charging | Troubleshoot the computer Troubleshoot UPS UPS not functioning | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679013023611002195/5679013023611002195_signature_28.png | Ligaya Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5679013023611002195/5679013023611002195_signature_31.png
| 2023-08-15 | 2023-08-15 16:25 | San Leonardo Rural Health Unit | Diversion, San Leonardo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25575 | PBSP | ligayadomingo20@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 ASSAY: Xpert MTB/RIF, COUNT: 50 | 2023-08-15 | Genexpert IV | 846042 | Computer not turning on UPS not charging | Troubleshoot the computer Troubleshoot UPS UPS not functioning | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679013023613212785/5679013023613212785_signature_28.png | Ligaya Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5679013023613212785/5679013023613212785_signature_31.png
| 2023-08-15 | 2023-08-15 16:25 | San Leonardo Rural Health Unit | Diversion, San Leonardo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25576 | PBSP | ligayadomingo20@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 ASSAY: Xpert MTB/RIF, COUNT: 50 | 2023-08-15 | Genexpert IV | 846042 | Computer not turning on UPS not charging | Troubleshoot the computer Troubleshoot UPS UPS not functioning | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679013023613295574/5679013023613295574_signature_28.png | Ligaya Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5679013023613295574/5679013023613295574_signature_31.png
| 2023-08-15 | 2023-08-15 16:25 | San Leonardo Rural Health Unit | Diversion, San Leonardo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25573 | PBSP | ligayadomingo20@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 ASSAY: Xpert MTB/RIF, COUNT: 50 | 2023-08-15 | Genexpert IV | 846042 | Computer not turning on UPS not charging | Troubleshoot the computer Troubleshoot UPS UPS not functioning | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5679013023616003440/5679013023616003440_signature_28.png | Ligaya Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5679013023616003440/5679013023616003440_signature_31.png
| 2023-08-15 | 2023-08-15 14:05 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25572 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1108 |  | GXIV | 831691 | B3 for re XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. B3: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678890470113782703/5678890470113782703_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5678890470113782703/5678890470113782703_signature_31.png
| 2023-08-15 | 2023-08-15 13:10 | Cong. Simeon G. Toribio Memorial Hospital | Carmen, bohol | 7 | Bohol | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25571 | DOH | himaya.vale@yahoo.com | 27 |  |  | Gx 4 | 838394 | Duedate for xpertcheck module b2 and b4 | Xpertcheck b2 and b4 | Xpertcheck b2 and b4 | Module b2 service require and b4 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678863698027484979/5678863698027484979_signature_28.png | Himaya Vale | https://www.jotform.com/uploads/macarewecare/222927978372067/5678863698027484979/5678863698027484979_signature_31.png
| 2023-08-15 | 2023-08-15 12:57 | Sto Domingo District Hospital | Pulong Buli, Sto. Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25570 | DOH | lhaydeesel@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1037 | 2023-08-15 | Genexpert IV | 831582 | Module 2 error 5007 | Performed Xpertcheck on Module 2 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck Result = Service Required | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678852157329491122/5678852157329491122_signature_28.png | Haydee Alejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5678852157329491122/5678852157329491122_signature_31.png
| 2023-08-15 | 2023-08-15 12:54 | Tondo Foreshore Super Health Center | Pacheco St. Tondo | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25569 | MACARE-HBDC | msff-tondo-labtech@paris.msf.org | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1887 |  | GX16-10C | 110016343 | Unit for preventive maintenance. | Performed preventive maintenance. Generated reports. | Preventive maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678848098916562183/5678848098916562183_signature_28.png | Dennis Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/5678848098916562183/5678848098916562183_signature_31.png
| 2023-08-15 | 2023-08-15 12:48 | Tondo Foreshore Super Health Center | Pacheco St. Tondo | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25568 | MACARE-HBDC | carina_moncada@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2107 ASSAY: Others, COUNT: XDR 436 |  | GXIV-10C | 110001650 | Unit for preventive maintenance. | Performed preventive maintenance. Generated reports. | Preventive maintenance. | Unit is operational. For request of Xpertcheck (Due Date: September 2023). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678844228918455780/5678844228918455780_signature_28.png | Carina Moncada | https://www.jotform.com/uploads/macarewecare/222927978372067/5678844228918455780/5678844228918455780_signature_31.png
| 2023-08-15 | 2023-08-15 12:33 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25567 | DOH | sajecopooh20@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 385 ASSAY: Xpert MTB/RIF Ultra, COUNT: 968 | 2023-07-31 | GXIV-4D | 838555 | Due for XPC. | #NAME? | #NAME? | Module A1 and A3 - for replacement. Module A2 and A4 - passed XPC. Module A2, A3, and A4 are functional. |  | XPERTCHECK-CE-5 | 10357 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678838187429959256/5678838187429959256_signature_28.png | Analiza Bianca Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/5678838187429959256/5678838187429959256_signature_31.png
| 2023-08-15 | 2023-08-15 11:57 | Caloocan City Quality Assurance Center RTDL | Sangandaan, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25566 | PBSP | caloocancentrallaboratory@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 156 ASSAY: Xpert HIV-1 Viral Load, COUNT: 41 |  | GXIV 6C | 805271 | CPU Real Time Clock Power Loss | Checked CPU Replaced BIOS battery; problem solved Performed PM Generated report | Troubleshoot CPU | Unit is operational. CPU working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678813505651300640/5678813505651300640_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5678813505651300640/5678813505651300640_signature_31.png
| 2023-08-15 | 2023-08-15 11:35 | Caloocan City Health Office | Sangandaan, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25565 | PBSP | caloocancentrallaboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3112 ASSAY: Xpert HIV-1 Viral Load, COUNT: 529 |  | GXIV 6C | 814154 | CPU does not boot up Boot device not found | Checked CPU Changed BIOS Battery; problem persists Updated date & time; problem persists CPU for replacement Performed PM Generated report | Troubleshoot CPU | GX is operational CPU for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678804095651712662/5678804095651712662_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5678804095651712662/5678804095651712662_signature_31.png
| 2023-08-15 | 2023-08-15 09:53 | Garcia memorial provincial hospital | Talibon, bohol | 7 | Bohol | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25564 | DOH | ybanezrosalina19@gmail.com | 24 |  |  | Gx4 | 838203 | Duedate for xpertcheck module B4 | Xpertcheck calibration module B4 | Xpertcheck calibration module B4 | Module 4 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678753773843462173/5678753773843462173_signature_28.png | Rosalina A. Ybañez | https://www.jotform.com/uploads/macarewecare/222927978372067/5678753773843462173/5678753773843462173_signature_31.png
| 2023-08-15 | 2023-08-15 09:19 | Mabinay Community Medicare Hospital | Mabinay, Negros Oriental | 7 | Negros Oriental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25563 | DOH | diannecm8@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 108 ASSAY: Xpert MTB/RIF Ultra, COUNT: 562 |  | GX4 | 838481 | Due for calibration/ Xpertcheck | #NAME? | - Perform Xpertcheck on all 4 modules | 4 Modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678735290324876863/5678735290324876863_signature_28.png | Dianne Mi-ot | https://www.jotform.com/uploads/macarewecare/222927978372067/5678735290324876863/5678735290324876863_signature_31.png
| 2023-08-15 | 2023-08-15 08:20 | Tondo Medical Center | Balut, Tondo | NCR | NCR 1st - Manila | Covid laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25562 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 65 ASSAY: Xpert MTB/RIF Ultra, COUNT: 260 ASSAY: Xpert SARS-CoV-2, COUNT: 1377 ASSAY: Xpert HIV-1 Viral Load, COUNT: 5 |  | GXIV | 834430 | Unit for PM. | Performed preventive maintenance. Generated reports. | Preventive maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678690608918425374/5678690608918425374_signature_28.png | Analyn N. Facto RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5678690608918425374/5678690608918425374_signature_31.png
| 2023-08-15 | 2023-08-15 08:26 | Tondo Medical Center | Balut, Tondo | NCR | NCR 1st - Manila | Covid laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25561 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 9 ASSAY: Xpert HPV, COUNT: 210 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1730 ASSAY: Xpert SARS-CoV-2, COUNT: 10378 ASSAY: Xpert HIV-1 Viral Load, COUNT: 7 |  | GX16 | 844121 | Unit for PM. For troubleshooting of disabled modules. | Performed PM. Performed troubleshooting of modules. Generated reports. | Preventive maintenance and troubleshooting. | Unit is operational. For reporting of A1, A2, A4, B1, C1 and D2. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678689868917986259/5678689868917986259_signature_28.png | Analyn N. Facto RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5678689868917986259/5678689868917986259_signature_31.png
| 2023-08-14 | 2023-08-14 13:04 | DINAGAT DISTRICT HOSPITAL | DINAGAT ISLAND | 13 | Dinagat Islands | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25560 | PBSP | ddhlaboratory30@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1164 |  | GX-IV | 807611 | PERFORM MODULE B3 AND B4 REPLACEMENT | COLLECT IQR, SYSTEM LOG,AND ARCHIVE | -PERFORM MODULE B3 DEF SN:662495 WITH NEW SN: 210100428 AND MODULE B4 DEF SN: 662496 WITH NEW SN: 210100429 -PERFORM PM AND GX CLEANING | EQUIPMENT IS OPERATIONAL | MODULE B2 PENDING |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678176870478101953/5678176870478101953_signature_28.png | Idhona Nuevas-Tillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5678176870478101953/5678176870478101953_signature_31.png
| 2023-08-14 | 2023-08-14 13:18 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25559 | DOH | luzvimindanavidad2@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1000 | 2023-07-31 | GXIV-4D | 838297 | Due for XPC | #NAME? | #NAME? | Module B4 -retry then Xpertcheck service required; for replacement Module B1, B2, and B3 - passed |  | XPERTCHECK-CE-5 | 10358 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678016072127819071/5678016072127819071_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5678016072127819071/5678016072127819071_signature_31.png
| 2023-08-14 | 2023-08-14 13:27 | Catigbian District Hospital | Catigbian, Bohol | 7 | Bohol | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25558 | DOH | rheranes@gmail.com | 28 |  |  | Gx 4 | 838494 | Due for xpertcheck calibration | Xpertcheck calibration clean fan filter | Xpertcheck calibration | System operational all modules passed calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5678007490169937766/5678007490169937766_signature_28.png | Rheazel G. Ranes | https://www.jotform.com/uploads/macarewecare/222927978372067/5678007490169937766/5678007490169937766_signature_31.png
| 2023-08-14 | 2023-08-10 14:45 | Milagros RHU | Baklay, Hills, Brgy. Bacolod, Milagros | 5 | Masbate | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25557 | PBSP | gracemarymesa1029@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 243 ASSAY: Xpert MTB/RIF Ultra, COUNT: 849 |  | GX4 | 838244 | For XpertCheck | XpertCheck: Cleaned cartridge bays, plunger rod mfan filter, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677934293019417544/5677934293019417544_signature_28.png | Mary grace Mesa | https://www.jotform.com/uploads/macarewecare/222927978372067/5677934293019417544/5677934293019417544_signature_31.png
| 2023-08-14 | 2023-08-14 10:30 | Toledo City Health Office | Ilihan Toledo City, Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 25556 | PBSP | loriemaediapen@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 |  | Gx4 | 808640 | Error 1001 and Error 4017 B1 several times daily Sn 667772 | Replaced new module Generate data IQR, Sys log, and Gxx files | Replaced B1 with new module B1 SN 210100860 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677897626788645070/5677897626788645070_signature_28.png | Lorie Mae C. Diapen | https://www.jotform.com/uploads/macarewecare/222927978372067/5677897626788645070/5677897626788645070_signature_31.png
| 2023-08-14 | 2023-08-14 10:19 | Tubigon Rural Health Unit | Tubigon, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25555 | PBSP | zaecin@gmail.com | 25 |  |  | Gx 4 | 815431 | Dedective module B1 sn 695935 | Install new module B1 sn 210100856 | Install new module B1 sn 210100856 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677890411661558981/5677890411661558981_signature_28.png | Niceas Joseph Boligao | https://www.jotform.com/uploads/macarewecare/222927978372067/5677890411661558981/5677890411661558981_signature_31.png
| 2023-08-14 | 2023-08-14 09:38 | Bogo City Health Office | Poblacion, Bogo City | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 25554 | PBSP | bogocholab@gmail.com | 23 |  | 2023-08-11 | GX4 | 808630 | Calibration Fail Module D2, D4 | Replace module D2, D4 with new serial number 210100787,210100803. | Replace module D2, D4 | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677867993911982954/5677867993911982954_signature_28.png | Regie Dahili | https://www.jotform.com/uploads/macarewecare/222927978372067/5677867993911982954/5677867993911982954_signature_31.png
| 2023-08-14 | 2023-08-14 09:30 | LA CARLOTA CITY HEALTH OFFICE | La Carlota City, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25553 | PBSP | tinbellotinbello@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1797 |  | GX4 | 807594 | Modules A2/627187 and A4/692185 failed latest XPC. | Preventive Maintenance. Optics cleaning, plunger rod maintenance, module bay cleaning, filter cleaning. | Replaced modules A2 and A4 with new modules with SN 210100815 and 210100849, respectively. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677860502175490695/5677860502175490695_signature_28.png | VALERIE CHRISTINE G. BELLO | https://www.jotform.com/uploads/macarewecare/222927978372067/5677860502175490695/5677860502175490695_signature_31.png
| 2023-08-14 | 2023-08-08 15:14 | Medilinx | Amorsolo st makati city | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 25552 | COMMERCIAL | reinaluz.bactol@medilinx.com.ph | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 91 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 421 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 457 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1645 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 34 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1228 |  | GX4 10C | 110001647 | Broken door himge module 3 | Check module b3 , cleaned plunger,cartridge bay, gathered reports | Troubleshooting and checked up | B3 broken door hinge |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677844883618740510/5677844883618740510_signature_28.png | Maria Reinaluz Bactol | https://www.jotform.com/uploads/macarewecare/222927978372067/5677844883618740510/5677844883618740510_signature_31.png
| 2023-08-14 | 2023-08-09 11:26 | General Nakar MHO | General Nakar,Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25551 | PBSP | buendichol@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 217 | 2023-08-09 | GX4 | 838180 | For module replacement module 2 807109 - 210101376 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | Operational |  | 900-0508 | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677843873614265628/5677843873614265628_signature_28.png | LYDELYN B. COSTUMBRADO | https://www.jotform.com/uploads/macarewecare/222927978372067/5677843873614265628/5677843873614265628_signature_31.png
| 2023-08-14 | 2023-08-09 13:13 | Cllaro M.Recto Memorial District Hospital | Infanta, Quezon | 4A | Quezon | Tb dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25550 | PBSP | katelynaveryumerez@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1366 | 2023-08-09 | GX4 | 807608 | Due for maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed Xpertcheck | Xpertcheck and maintenance | All module passed |  | XPERTCHECK-CE-5 | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677842983615227122/5677842983615227122_signature_28.png | Lorelyn Umerez Miras | https://www.jotform.com/uploads/macarewecare/222927978372067/5677842983615227122/5677842983615227122_signature_31.png
| 2023-08-14 | 2023-08-10 11:35 | Tropical Disease foundation inc | Amorsolo st. Makati city | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25549 | COMMERCIAL | krcueno@tdf.org.ph | 27 | ASSAY: Xpert MTB/RIF, COUNT: 32 ASSAY: Xpert MTB/RIF Ultra, COUNT: 35 ASSAY: Xpert HIV-1 Viral Load, COUNT: 959 | 2023-08-08 | GX4 | 705175 | Due for maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed xpertcheck | Maintenance and xpertcheck | Operational |  | Xpertcheck-ce-5 | 96085 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677842383615327386/5677842383615327386_signature_28.png | Karl Francis R. Cueno | https://www.jotform.com/uploads/macarewecare/222927978372067/5677842383615327386/5677842383615327386_signature_31.png
| 2023-08-14 | 2023-08-10 14:45 | Lung center of the Philippines | Quezon avenue | NCR | NCR 2nd - Quezon City | Stem Cell | NO | MAINTENANCE | OTHERS | Noel Floresca |  |  |  | Plasmair | 25548 | MACARE-HBDC | lcp.mdctl2022@gmail.com | 28 |  |  | Airinspace | 2006- T - 1188 | Due for maintenance | Cleaned unit Cleaned filter, and visual check Checked electrical and connection Checked program parameters Performed functionality test | Maintenance | Operationa. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677839253619292460/5677839253619292460_signature_28.png | Jennifer p. Panuelos | https://www.jotform.com/uploads/macarewecare/222927978372067/5677839253619292460/5677839253619292460_signature_31.png
| 2023-08-14 | 2023-08-11 13:53 | AFRIMS | V.luna,QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | OTHERS | Noel Floresca |  |  |  | Thermo Scientific | 25547 | COMMERCIAL | dionesps@afrims.org | 26 |  |  | Forma 700 series | 823102 - 1736 | Unable desire temp. -80 | Checked filter Checked Condensing fan motor,door gasket Checked program parameters Test and observed unit | Troubleshooting | For observation unit reach -70 deg. Set temp -80 deg. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677838333618450282/5677838333618450282_signature_28.png | Paula corazon diones | https://www.jotform.com/uploads/macarewecare/222927978372067/5677838333618450282/5677838333618450282_signature_31.png
| 2023-08-14 | 2023-08-11 15:45 | National kidney and transplant institute | East avenue q.c | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25546 | MACARE-HBDC | molecularlabnkti@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 703 ASSAY: Others, COUNT: 13 ASSAY: Xpert HBV VL, COUNT: 444 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1219 ASSAY: Xpert MTB/RIF Ultra, COUNT: 821 ASSAY: Xpert SARS-CoV-2, COUNT: 16376 ASSAY: Xpert HCV VL, COUNT: 40 ASSAY: Others, COUNT: 9 | 2023-08-10 | GXXVI | 110005587 | For module replacement and maintenance C3 210028988 -796840 D3 800477 - 722098 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | Operational |  | 900-0491RS | 10490 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677837683618609574/5677837683618609574_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5677837683618609574/5677837683618609574_signature_31.png
| 2023-08-14 | 2023-08-11 17:17 | East Avenue Medical Center | East avenue q.c | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25545 | MACARE-HBDC | eamcmolecularlab@gmail.com | 22 | ASSAY: Xpert CT/NG, COUNT: 277 ASSAY: Xpert HPV, COUNT: 496 ASSAY: Xpert SARS-CoV-2, COUNT: 2979 ASSAY: Others, COUNT: 2 | 2023-08-10 | GX4 | 110014085 | Due for maintenance and for module replacement B2 210069925 - 210076858 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Maintenance and Module replacement | Operational |  | 900-0650 | 10491 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677836043612385781/5677836043612385781_signature_28.png | Jev Ryan Bonifacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5677836043612385781/5677836043612385781_signature_31.png
| 2023-08-13 | 2023-08-13 19:02 | Teresita L. Jalandoni Provincial Hospital | Silay City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25544 | PBSP | mvillan.pbsp@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3196 |  | GX4 | 805093 | None | Conduct PM and generate files | Run Xpert check calibration for 4 modules. | All modules passed the calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677339368611449527/5677339368611449527_signature_28.png | Arna Villan | https://www.jotform.com/uploads/macarewecare/222927978372067/5677339368611449527/5677339368611449527_signature_31.png
| 2023-08-13 | 2023-08-11 16:16 | Teresita L. Jalandoni Provincial Hospital | Silay City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25543 | DOH | mvillan.pbsp@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3196 |  | GX4 | 838464 | Module B1 with SN 802411 failed in Xpert check. | Conduct PM and generate files | Run Xpert check calibration for 4 modules. | Modules B2, B3 & B4 passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677333508615849249/5677333508615849249_signature_28.png | Arna Villan | https://www.jotform.com/uploads/macarewecare/222927978372067/5677333508615849249/5677333508615849249_signature_31.png
| 2023-08-13 | 2023-08-13 12:02 | Vicente Sotto Memorial Medical Center | B Rodriguez St, Fuente Cebu | 7 | Cebu | Bloodbank | YES | REPAIR | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 25542 | COMMERCIAL | vsmmcbloodservices@gmail.com | 25 |  |  | Mcs+ | 02A090 | Valve test Failure | Replaced motor compressor | Replaced air motor compressor | Unit mcs+ operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5677089128013830101/5677089128013830101_signature_28.png | Maria Catherine Milani Malusay | https://www.jotform.com/uploads/macarewecare/222927978372067/5677089128013830101/5677089128013830101_signature_31.png
| 2023-08-12 | 2023-08-12 09:29 | Cebu South Medical Center | San Isidro Talisay City, Cebu | 7 | Cebu | Tb Laboratory | YES | CALIBRATION TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25541 | DOH | csmclaboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 653 ASSAY: Xpert HIV-1 Viral Load, COUNT: 32 |  | Gx4 | 838559 | Due for Calibratio | Performed calibration | Performed calibration in genexpert mschjnr | All modules are operational and passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5676135652323653898/5676135652323653898_signature_28.png | Gerald Solana, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5676135652323653898/5676135652323653898_signature_31.png
| 2023-08-11 | 2023-08-11 15:53 | Zamboanga City Medical Centre | Dr. Evangelista Street, Snta Catalina Rd. Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25540 | PBSP | jbcua.pbsp@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5070 ASSAY: Xpert SARS-CoV-2, COUNT: 13596 ASSAY: Xpert HIV-1 Viral Load, COUNT: 304 | 2023-08-11 | GX 16 | 839309 | Hardware failed module A1 and D2 | Replaced Modules 210068683, 210068687, 210068682, 653849 and 210068699 with new Modules 210100699, 210100728, 210100740, 210101371 and 210101373 respectively. | Preventive Maintenence and Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675526408418882283/5675526408418882283_signature_28.png | James peniel b. Cua | https://www.jotform.com/uploads/macarewecare/222927978372067/5675526408418882283/5675526408418882283_signature_31.png
| 2023-08-11 | 2023-08-11 15:53 | Zamboanga City Medical Centre | Dr. Evangelista Street, Snta Catalina Rd. Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25539 | PBSP | jbcua.pbsp@gmail.com | 28 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5070 ASSAY: Xpert SARS-CoV-2, COUNT: 13596 ASSAY: Xpert HIV-1 Viral Load, COUNT: 304 | 2023-08-11 | GX 16 | 839309 | Hardware failed module A1 and D2 | Replaced Modules 210068683, 210068687, 210068682, 653849 and 210068699 with modules 210100699, 210100728, 210100740, 210101371 and 210101373 respectively. | Preventive Maintenence and Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675526408418882283/5675526408418882283_signature_28.png | James peniel b. Cua | https://www.jotform.com/uploads/macarewecare/222927978372067/5675526408418882283/5675526408418882283_signature_31.png
| 2023-08-11 | 2023-08-11 15:01 | San Manuel Rural Health Unit | District 3, San Manuel, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25538 | PBSP | tonyantoniojr100@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1013 | 2023-08-11 | Genexpert IV | 834686 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 786817 - 210092911 | Genexpert Module = 210092911 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675483510118263616/5675483510118263616_signature_28.png | Tony Antonio Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5675483510118263616/5675483510118263616_signature_31.png
| 2023-08-11 | 2023-08-11 15:17 | Philippine Blood Center | Q. Ave, Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  |  |  |  | 25537 | MACARE-HBDC | rcb.macare@gmail.com | 25 |  |  | MCS+ | 15H029G | Display panel buttons not working | #NAME? | Performed troubleshooting | Unit is not operational. Display panel ribbon for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675482372518331749/5675482372518331749_signature_28.png | Ivan Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/5675482372518331749/5675482372518331749_signature_31.png
| 2023-08-11 | 2023-08-11 15:20 | Philippine Blood Center | Q. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  |  |  |  | 25536 | MACARE-HBDC | strobecatcher@gmail.com | 25 |  |  | MCS+ | 15H028G | Unit for maintenance | #NAME? | Performed maintenance | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675481352513395690/5675481352513395690_signature_28.png | Ivan Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/5675481352513395690/5675481352513395690_signature_31.png
| 2023-08-11 | 2023-08-11 15:22 | Philippine Blood Center | Q. Ave., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  |  |  |  | 25535 | MACARE-HBDC | strobecatcher@gmail.com | 25 |  |  | MCS+ | 15H030G | Unit for maintenance | #NAME? | Performed maintenance | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675479652514466342/5675479652514466342_signature_28.png | Ivan Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/5675479652514466342/5675479652514466342_signature_31.png
| 2023-08-11 | 2023-08-11 14:12 | Ospital ng Palawan | 220 Malvar Road, San Miguel, Puerto Princesa, Palawan | 4B | Palawan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25534 | PBSP | svsilva.pbsp@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 3016 ASSAY: Xpert SARS-CoV-2, COUNT: 1901 | 2022-12-12 | GX XVI 6C | 842893 | Unit for XpertCheck | Performed PM Performed XpertCheck Module C3 and C4; Service Required Generated report | Performed XpertCheck | Unit is operational. AC needs to be fixed as soon as possible to avoid further problems on other modules |  | XpertCheck Kit | DR 37730 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675450119921925686/5675450119921925686_signature_28.png | Shayne Marie Silva | https://www.jotform.com/uploads/macarewecare/222927978372067/5675450119921925686/5675450119921925686_signature_31.png
| 2023-08-11 | 2023-08-11 13:22 | Subic Rural Health Unit | National Rora, Mangan-vaca Subic Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25533 | PBSP | sanchezmacristina313@gmail.com | 29 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 529 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2349 ASSAY: Xpert HIV-1 Viral Load, COUNT: 68 | 2023-08-07 | GXIV-D | 812788 | B3 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is fully operational |  | 1 genexpert module | 42436 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675416052815868797/5675416052815868797_signature_28.png | Ma. Cristina D. Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/5675416052815868797/5675416052815868797_signature_31.png
| 2023-08-11 | 2023-08-11 13:22 | Subic Rural Health Unit | National Rora, Mangan-vaca Subic Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25532 | PBSP | sanchezmacristina313@gmail.com | 29 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 529 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2349 ASSAY: Xpert HIV-1 Viral Load, COUNT: 68 | 2023-08-07 | GXIV-D | 812788 | B3 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is fully operational |  | 1 genexpert module | 42436 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675416022813328086/5675416022813328086_signature_28.png | Ma. Cristina D. Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/5675416022813328086/5675416022813328086_signature_31.png
| 2023-08-11 | 2023-08-11 13:18 | Esperanza Rural Health Unit | Poblacion, Esperanza | 13 | Agusan del Sur | RTDL | NO | CALIBRATION | CEPHEID | Faizar Avenido |  |  |  |  | 25531 | WHO | arleneo1970@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 99 ASSAY: Xpert MTB/RIF Ultra, COUNT: 766 |  | GX-4 | 847258 | Module B4 found Retry in XPC | None | XpertCheck Calibration Preventive Maintenance | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675408822214027169/5675408822214027169_signature_28.png | Arlene Ocampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5675408822214027169/5675408822214027169_signature_31.png
| 2023-08-11 | 2023-08-11 10:51 | St.lukes Medical Center QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | YES | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25530 | COMMERCIAL | megapo@stlukes.com.ph | 25 | ASSAY: Xpert HPV, COUNT: 263 | 2023-08-08 | GX4 | 839234 | For module replacement defective D1 module 866693 (def) - 210072911 new module, D2 module 862368 (def) - 210096421 new module and for maintenance | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | D1 and D2 module operational ready to use,D3 and D4 disabled |  | 900-0508 | 96091 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675354705327093582/5675354705327093582_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5675354705327093582/5675354705327093582_signature_31.png
| 2023-08-11 | 2023-08-11 10:51 | St.lukes Medical Center QC | E.Rodriguez QC | NCR | NCR 2nd - Quezon City | Microbiology | YES | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25529 | COMMERCIAL | biomed.qc@stlukes.com.ph | 25 | ASSAY: Xpert HPV, COUNT: 263 | 2023-08-08 | GX4 | 839234 | For module replacement defective D1 module 866693 (def) - 210072911 new module, D2 module 862368 (def) - 210096421 new module and for maintenance | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | D1 and D2 module operational ready to use,D3 and D4 disabled |  | 900-0508 | 96091 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5675354365324120522/5675354365324120522_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5675354365324120522/5675354365324120522_signature_31.png
| 2023-08-10 | 2023-08-10 17:08 | Bontoc General Hospital | Bontoc, Mountain Province | CAR | Mountain Province | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25528 | PBSP | mffanglayan@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 741 | 2023-08-10 | GX4 | 807582 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Modules 662554>>>210093389 662555>>>210092946 662553>>>210092871 - Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Modules 210092871, 210092946, 210093389 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674684009718593715/5674684009718593715_signature_28.png | Valentina Bagano | https://www.jotform.com/uploads/macarewecare/222927978372067/5674684009718593715/5674684009718593715_signature_31.png
| 2023-08-10 | 2023-08-09 13:59 | Baguio Health Services Office | Baguio City, Benguet | CAR | Benguet | TB DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25527 | DOH | sherryannemarrero@gmail.com | 23 |  | 2023-07-20 | GX4 | 838281 | Due for Calibration Due for Troubleshooting | - Cleaned Cartridge Bays, Plunger Rods, I-CORE - Replaced Defective module B4 807865>>>210005712 - Performed ReXpertCheck - Generated Reports | Performed Module Replacement Performed ReXpertCheck | Unit is operational |  | GeneXpert Module 210005712 | DR 10206 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674674219716668709/5674674219716668709_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5674674219716668709/5674674219716668709_signature_31.png
| 2023-08-10 | 2023-08-10 14:56 | Southern Palawan Provincial Hospital | Brooke's Point, Southern Palawan | 4B | Palawan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25526 | PBSP | roderick.conte@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1202 | 2023-07-24 | GXIV 6C | 804271 | Module B2 for replacement Unit for XpertCheck Module B3 not detected | Checked and tightened power cord and cables of B3; problem persists Straightened bent power cable; problem fixed,module b3 detected Replaced module B2 ( 761283 -> 210100382 ) Performed XpertCheck on B1,B3 and B4 Generated report | Troubleshoot Module B3 Replaced Module B2 Performed XpertCheck on B1, B3 and B4 | Unit is operational |  | GX module 6C & XpertCheck kit | DR#0042148, DR#0041936 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674670008781919715/5674670008781919715_signature_28.png | RODERICK D CONTE | https://www.jotform.com/uploads/macarewecare/222927978372067/5674670008781919715/5674670008781919715_signature_31.png
| 2023-08-10 | 2023-08-10 16:13 | Naujan Municipal Health Office | Naujan | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25525 | DOH | hmaryann524@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 550 | 2023-07-25 | GX-IV-D | 838226 | Modules B2 and B3 for replacement. | Replaced defective modules | Performed maintenance Replaced defective modules Generates reports | Unit is operational |  | Old: B2/807209, B3/807207 New: B2/677540, B3/ 676201 | DR# 10280 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674653948884210945/5674653948884210945_signature_28.png | Mary Ann Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5674653948884210945/5674653948884210945_signature_31.png
| 2023-08-10 | 2023-08-10 15:41 | Badian Rural Health Unit | Badian Poblacio | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25524 | PBSP | geecube23@gmail.com | 38 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 792 |  | Gx4 | 807588 | Module 4 encounter an error of 2008 | Orient end user to check sample during pipette | Clean Up machine, generate data, gxx fis, assay statistics, iQR, sys log | All units encounter hot temp in the area. Provide additional aircon in the room |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674629395173124982/5674629395173124982_signature_28.png | Avenizer Absin, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5674629395173124982/5674629395173124982_signature_31.png
| 2023-08-10 | 2023-08-10 14:53 | Zamboanga City Medical Centre | Evangelista Street, Snta Catalina Road, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25523 | WHO | jbcua.pbsp@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 285 ASSAY: Xpert SARS-CoV-2, COUNT: 1423 ASSAY: Xpert HIV-1 Viral Load, COUNT: 23 | 2023-08-10 | GX 4 | 848250 | Module 1 hardware failed | Replaced module 210002709 and 210002724 with modules 210031453 and 210032985 respectively | Preventive Maintenence and Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674604740511930393/5674604740511930393_signature_28.png | James peniel b. Cua | https://www.jotform.com/uploads/macarewecare/222927978372067/5674604740511930393/5674604740511930393_signature_31.png
| 2023-08-10 | 2023-08-10 14:06 | Pinamalayan Rural Health Unit | Pinamalayan | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25522 | DOH | mholaboratory.pinamalayan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1329 | 2023-07-25 | GX-IV-D | 838230 | Module D1 for replacement module D3 for XpertCheck | Replaced module D1 Run XpertCheck on module D3 | Performed maintenance Replaced defective module Run XpertCheck on module D3 Generated reports | Unit is operational. Module D3 passed in XpertCheck | Used 2 spare XpertCheck | Old: D1/655304 New: D1/685066 | DR# 10282 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674598020317405958/5674598020317405958_signature_28.png | Dave Erwin M. Festin, RN, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5674598020317405958/5674598020317405958_signature_31.png
| 2023-08-10 | 2023-08-10 13:57 | Moalboal Rural Health Unit | Moalboal, Poblacion Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25521 | MACARE-HBDC | roavaul@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 543 |  | Gx4 | 838892 | Monitor problem | Troubleshoot and recommend the monitor for replacement | Swap monitor of other unit and found out the problem exist in the other cpu unit.. | All modules are operational and only the monitor lcd problem still exist during 2 weeks of observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674568657819815773/5674568657819815773_signature_28.png | Ava A. Gako, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5674568657819815773/5674568657819815773_signature_31.png
| 2023-08-10 | 2023-08-10 13:12 | Las Nieves Rural Health Unit | Purok 5, Brgy. Poblacion, Las Nieves | 13 | Agusan del Norte | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 25520 | WHO | Daisermt@yahoo.com | 27.87 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 199 |  | GX-4 | 847338 | 4 Modules Stuck Cartridges | None | Removing Stuck Cartridges | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674541511813678623/5674541511813678623_signature_28.png | Daisy Siega | https://www.jotform.com/uploads/macarewecare/222927978372067/5674541511813678623/5674541511813678623_signature_31.png
| 2023-08-10 | 2023-08-10 12:55 | Kawayan Rural Health Unit | Kawayan, Biliran | 8 | Biliran | TB Dots | YES | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25519 | PBSP | lpvidal@yahoo.com.ph | 25 |  | 2023-08-10 | GXII | 835196 | Defective CPU | #NAME? | #NAME? | All modules are functional. |  | Lenovo ThinkCentre SN: CPH0622F37221 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674529316489858060/5674529316489858060_signature_28.png | Leo P. Vidal | https://www.jotform.com/uploads/macarewecare/222927978372067/5674529316489858060/5674529316489858060_signature_31.png
| 2023-08-10 | 2023-08-10 12:07 | Balud Rural Health Unit | Purok Everlasting, Poblacion, Balud | 5 | Masbate | RTDL Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25518 | PBSP | shao_shaol@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 348 |  | GX4 | 835199 | B1/788613 for module replacement | Replaced module: 788613>>210101750 Performed XpertCheck Cleaned cartridge bays, plunger rod, fan, fan filter, icore Generate reports | Replaced B1 module Performed XpertCheck Performed Preventive Maintenance | Unit is operational B1/210101750 - new module B2/789015 - |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674505408224731112/5674505408224731112_signature_28.png | Sharon Oliva | https://www.jotform.com/uploads/macarewecare/222927978372067/5674505408224731112/5674505408224731112_signature_31.png
| 2023-08-10 | 2023-08-10 10:52 | San Mariano Rural Health Unit | Sta. Filomena, San Mariano, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25517 | PBSP | jonalyndumlao78@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1549 | 2023-08-10 | Genexpert IV | 812789 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 686199 - 210061010 | Genexpert Module = 210061010 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674457582845909739/5674457582845909739_signature_28.png | Jonalyn F. Dumlao | https://www.jotform.com/uploads/macarewecare/222927978372067/5674457582845909739/5674457582845909739_signature_31.png
| 2023-08-10 | 2023-08-10 10:52 | San Mariano Rural Health Unit | Sta. Filomena, San Mariano, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25516 | PBSP | jonalyndumlao78@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1549 | 2023-08-10 | Genexpert IV | 812789 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 686199 - 210061010 | Genexpert Module = 210061010 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674457552841073527/5674457552841073527_signature_28.png | Jonalyn F. Dumlao | https://www.jotform.com/uploads/macarewecare/222927978372067/5674457552841073527/5674457552841073527_signature_31.png
| 2023-08-10 | 2023-08-10 10:26 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25515 | DOH | lornsabaigar@gmaip.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 615 |  | GXIV | 838409 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674438377425823643/5674438377425823643_signature_28.png | Lorna C. Abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/5674438377425823643/5674438377425823643_signature_31.png
| 2023-08-10 | 2023-08-10 10:09 | Tarlac City Health Center 2 | Brgy. Matatalaib, Tarlac City | 3 | Tarlac | Laboratory | NO | CPU pulled-out | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25514 | DOH | clarizannebeltran1996@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV-D | 831917 | CPU for troubleshooting UPS not backing-up for more than 1 minute | Troubleshooting | Cpu troubleshooting and pulled-out UPS troubleshooting | Unit is not operational, defective CPU and UPS |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674430462835246058/5674430462835246058_signature_28.png | Clariz Anne Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5674430462835246058/5674430462835246058_signature_31.png
| 2023-08-10 | 2023-08-10 10:09 | Tarlac City Health Center 2 | Brgy. Matatalaib, Tarlac City | 3 | Tarlac | Laboratory | NO | CPU pulled-out | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25513 | DOH | clarizannebeltran1996@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV-D | 831917 | CPU for troubleshooting UPS not backing-up for more than 1 minute | Troubleshooting | Cpu troubleshooting and pulled-out UPS troubleshooting | Unit is not operational, defective CPU and UPS |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674430432833484824/5674430432833484824_signature_28.png | Clariz Anne Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5674430432833484824/5674430432833484824_signature_31.png
| 2023-08-10 | 2023-08-10 09:19 | Danao City Health Office | Danao City, Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25512 | PBSP | chockenriquez@gmail.com | 28 |  |  | Gx 4 | 808286 | Defective Module B1 644046 | Clean fan filter, New module installed B1 210100551 | New module installed B1 210100551 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674425848724266510/5674425848724266510_signature_28.png | Nicanor Enriquez Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5674425848724266510/5674425848724266510_signature_31.png
| 2023-08-10 | 2023-08-10 08:41 | San Sebastian District Hospital | Sablayan | 4B | Occidental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25510 | PBSP | ssdhlaboratory@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2303 | 2023-08-10 | GX-IV-D | 812791 | Module B2 for replacement | Replaced module B2 | Performed maintenance Replaced defective module Generated reports | All modules operational. |  | Old: B2/693248 New: B2/210100332 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674381556211558910/5674381556211558910_signature_28.png | Anthony Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5674381556211558910/5674381556211558910_signature_31.png
| 2023-08-10 | 2023-08-10 08:41 | San Sebastian District Hospital | Sablayan | 4B | Occidental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25511 | PBSP | ssdhlaboratory@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2303 | 2023-08-10 | GX-IV-D | 812791 | Module B2 for replacement | Replaced module B2 | Performed maintenance Replaced defective module Generated reports | All modules operational. |  | Old: B2/693248 New: B2/210100332 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674381566214649563/5674381566214649563_signature_28.png | Anthony Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5674381566214649563/5674381566214649563_signature_31.png
| 2023-08-09 | 2023-08-09 10:50 | Lambayong Rural Health Unit | Poblacion, Lambayong, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25509 | PBSP | jeanalonzo73@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 492 |  | Gx-IV | 838159 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5674044019545635728/5674044019545635728_signature_28.png | Jean K. Alonzo | https://www.jotform.com/uploads/macarewecare/222927978372067/5674044019545635728/5674044019545635728_signature_31.png
| 2023-08-09 | 2023-08-09 21:34 | Masbate City Health Office | Brgy. Kalipay, PHO Rd., Masbate City | 5 | Masbate | RTDL Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25508 | PBSP | desrmt2017@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 32 ASSAY: Xpert MTB/RIF Ultra, COUNT: 532 |  | GX4 | 808633 | For XpertCheck and Preventive Maintenance | XpertCheck Result: B2, B3, B4 Service Required Cleaned cartridge bays, plunger rod, fan filter, fan, icore Generate reports | Performed XpertCheck Performed Preventive Maintenance | Unit is Operational B1/687101 - Passed B2/730126 - Service Required B3/662583 - Service Required B4/635443 - Service Required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673977803423352401/5673977803423352401_signature_28.png | Desiree Mae D Dela Rosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5673977803423352401/5673977803423352401_signature_31.png
| 2023-08-09 | 2023-07-26 12:12 | Cataingan Rural Health Unit | Quezon st., Poblacion, Cataingan | 5 | Masbate | RTDL Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25507 | PBSP | catainganrhu@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 198 ASSAY: Xpert MTB/RIF Ultra, COUNT: 469 |  | GX4 | 838164 | B2/806995 for module replacement | Replaced module B2 806995>>210093390 Performed XpertCheck: all module passed Cleaned cartridge bays, plunger rod, fan, fan filter, icore Generate reports | Replaced B2 module Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673972043423237139/5673972043423237139_signature_28.png | Rowena c alicante | https://www.jotform.com/uploads/macarewecare/222927978372067/5673972043423237139/5673972043423237139_signature_31.png
| 2023-08-09 | 2023-08-09 17:05 | Sto Rosa Community Hospital | Sto Rosa, Olango Island, Lapulapu City | 7 | Cebu | Tb Laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 25506 | DOH | ariestumz@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 178 |  | Gx4 | 838534 | Due for Calibration | Clean Up Icore, cartridge bay and performed calibrate all modules | Performed Calibration | All modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673813633911338692/5673813633911338692_signature_28.png | Arieston Tumulak | https://www.jotform.com/uploads/macarewecare/222927978372067/5673813633911338692/5673813633911338692_signature_31.png
| 2023-08-09 | 2023-08-09 15:28 | Ormoc City Health Office | Ormoc City, Leyte | 8 | Leyte | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25505 | PBSP | theresesaladaga@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 912 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2390 ASSAY: Xpert HIV-1 Viral Load, COUNT: 59 | 2023-08-09 | GXIV-4D | 804253 | Module B3 and B4 - Xpertcheck service required | -Replace defective module B3/654651 with new module SN: 210100530 -Replace defective module B4/698275 with new module SN: 210100543 | #NAME? | All modules are functional. |  | 900-0508 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673771650014469278/5673771650014469278_signature_28.png | Raina Zerene P. Ablen | https://www.jotform.com/uploads/macarewecare/222927978372067/5673771650014469278/5673771650014469278_signature_31.png
| 2023-08-09 | 2023-08-09 14:40 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 25504 | PBSP | laboratory@loveyourself.ph | 25 |  |  | GXIV | 808649 | Laptop CMOS battery for replacement | #NAME? | Replaced CMOS battery | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673769761317572470/5673769761317572470_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5673769761317572470/5673769761317572470_signature_31.png
| 2023-08-09 | 2023-08-09 15:39 | Labuan General hospital | Labuan General hospital | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25503 | DOH | jenne060213@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-07-21 | GX 4 | 838568 | Module 1 hardware failed | Replaced module 862914 with Module 906356, module 2 passed, 3 service required and 4 passed, undergone Xpertcheck calibration | Module Replacement, Preventive maintenence and xpertcheck calibration | All modules are available. | None | 900-0491R genexpert module 6 - colors serial number 906356 | 10225 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673764819415937144/5673764819415937144_signature_28.png | Jennerose Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5673764819415937144/5673764819415937144_signature_31.png
| 2023-08-09 | 2023-08-09 14:43 | Balamban Rural Health Unit | Poblacion, Balamban | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  |  |  |  | 25502 | DOH | climaco.mildred@yahoo.com | 22 |  | 2023-08-09 | Gx4 | 838231 | Due for calibration | Perform calibration | Perform Calibration for module B2, B3, B4 | All modules passed for calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673736630426666340/5673736630426666340_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5673736630426666340/5673736630426666340_signature_31.png
| 2023-08-09 | 2023-08-09 14:36 | Balamban Rural Health Unit | Poblacion, Balamban | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  |  |  |  | 25501 | DOH | climaco.mildred@yahoo.com | 22 |  | 2023-08-09 | Gx4 | 838564 | Due for calibration. | Perform calibration | Perform calibration for all module's | All module passed for calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673725170426275542/5673725170426275542_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5673725170426275542/5673725170426275542_signature_31.png
| 2023-08-09 | 2023-08-09 14:12 | Placer Rural Health Unit | Poblacion, placer | 5 | Masbate | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25500 | PBSP | rhuplacermasbatefiles@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 380 |  | GX4 | 835198 | For preventive maintenance | Cleaned cartridge bays,plunger rod, fan filter Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673715837506691955/5673715837506691955_signature_28.png | Ma. Serecia B. Santor | https://www.jotform.com/uploads/macarewecare/222927978372067/5673715837506691955/5673715837506691955_signature_31.png
| 2023-08-09 | 2023-08-09 13:49 | Jose R. Reyes Memorial Medical Center | Rizal Ave., Sta. Cruz, Manila | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25499 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 137 ASSAY: Xpert MTB/RIF Ultra, COUNT: 164 ASSAY: Xpert SARS-CoV-2, COUNT: 306 |  | GXIV 10C | 110014099 | Module B4 door broken | Checked Module B4 Replaced Module B4 door Performed PM Generated report | Replaced Module B4 door | Unit is operational. All Modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673703129712488290/5673703129712488290_signature_28.png | Ervin Emmett Sambolledo | https://www.jotform.com/uploads/macarewecare/222927978372067/5673703129712488290/5673703129712488290_signature_31.png
| 2023-08-09 | 2023-08-09 12:55 | Caridad Health Center | Cavite City | 4A | Cavite | Laboratory | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25498 | DOH | seyerlanie85@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: N/A |  | GXIV | 821227 | CPU not powering on. | Performed troubleshooting and repair. Resets internal components. Replaced CMOS battery. Tested other display ports. Cleaned RAM and tested other ram slots. Generated reports. | On-site troubleshooting and repair. | Unit is operational. Workstation is still not powering On and is defective and subject for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673664776547818828/5673664776547818828_signature_28.png | Melanie Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5673664776547818828/5673664776547818828_signature_31.png
| 2023-08-09 | 2023-08-09 12:08 | Malay Municipal Health Office | Malay, Aklan | 6 | Aklan | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25497 | DOH | rhedal30@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 692 |  | GX4 | 838457 | None | Run Xpert check calibration for 4 modules.. | Run Xpert check calibration | All modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673651893323221368/5673651893323221368_signature_28.png | Redal D. Taytay | https://www.jotform.com/uploads/macarewecare/222927978372067/5673651893323221368/5673651893323221368_signature_31.png
| 2023-08-09 | 2023-08-09 12:08 | Malay Municipal Health Office | Malay, Aklan | 6 | Aklan | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25496 | DOH | rhedal30@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 692 |  | GX4 | 838457 | None | Run Xpert check calibration for 4 modules.. | Run Xpert check calibration | All modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673651873325476309/5673651873325476309_signature_28.png | Redal D. Taytay | https://www.jotform.com/uploads/macarewecare/222927978372067/5673651873325476309/5673651873325476309_signature_31.png
| 2023-08-09 | 2023-08-09 12:08 | Cataingan Rural Health Unit | Quezon st., Poblacion, Cataingan | 5 | Masbate | RTDL Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25495 | PBSP | catainganrhu@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 177 |  | GX4 | 812787 | For Preventive Maintenance | Cleaned cartridge bays, plunger rod, fan filter, icore Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673634864615003875/5673634864615003875_signature_28.png | Rowena C. Alicante | https://www.jotform.com/uploads/macarewecare/222927978372067/5673634864615003875/5673634864615003875_signature_31.png
| 2023-08-09 | 2023-08-09 11:53 | Cabadbaran City Health Office | Brgy. Mabini, Cabadbaran City | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25494 | DOH | samstercem@gmail.com | 27.27 | ASSAY: Xpert MTB/RIF, COUNT: 221 ASSAY: Xpert MTB/RIF Ultra, COUNT: 893 | 2023-07-21 | GX-4 | 838550 | None | None | Replacing Module B2 with SN: 862860 to SN: 698222 Preventive Maintenance | All modules are operational |  | GeneXpert Module 6-Color | 10229 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673632525304649106/5673632525304649106_signature_28.png | Sammy Cempron | https://www.jotform.com/uploads/macarewecare/222927978372067/5673632525304649106/5673632525304649106_signature_31.png
| 2023-08-09 | 2023-08-09 11:17 | Braulio E. Dujali Rural Health Unit | Poblacion, Braulio E. Dujali, Davao Del Norte | 11 | Davao del Norte | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25493 | PBSP | sblacorda@yahoo.com | 25.9 |  |  | GX-IV 2D | 835216 | For Calibration | Run XPC Calibration on Two Modules | Perform XPC Calibration | All Modules Passed the Calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673627793811708817/5673627793811708817_signature_28.png | Sharon Rose B. Lacorda RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5673627793811708817/5673627793811708817_signature_31.png
| 2023-08-09 | 2023-08-09 11:29 | Delfin Albano Memorial Hospital | Ragan Sur, Delfin Albano, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph | maryrosebacanirmt@gmail.com |  | 25492 | DOH | damh_da@yahoo.com | 28 |  | 2023-02-14 | Genexpert IV | 838208 | Module 3, HW Failure | Performed Xpertcheck on Module 3 | Performed Maintenance Gathered Reports Performed Calibration {Xpertcheck Result = (B2, B4 - Passed) & (B1, B3 - Service Required)} Performed Functionality Test | Unit is Operational | Aircon Needs Replacement, 1.5hp - 2hp aircon | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673616283014264109/5673616283014264109_signature_28.png | STEPHEN VALOROZO | https://www.jotform.com/uploads/macarewecare/222927978372067/5673616283014264109/5673616283014264109_signature_31.png
| 2023-08-09 | 2023-08-09 11:29 | Delfin Albano Memorial Hospital | Ragan Sur, Delfin Albano, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25491 | DOH | damh_da@yahoo.com | 28 |  | 2023-02-14 | Genexpert IV | 838208 | Module 3, HW Failure | Performed Xpertcheck on Module 3 | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = (B2, B4 - Passed) & (B1, B3 - Service Required) | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673616263019700127/5673616263019700127_signature_28.png | STEPHEN VALOROZO | https://www.jotform.com/uploads/macarewecare/222927978372067/5673616263019700127/5673616263019700127_signature_31.png
| 2023-08-09 | 2023-08-09 11:30 | Pulang Lupa Dos Health Center | Pulang Lupa Dos, Las Pinas | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25490 | PBSP | elvimir08@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1883 | 2023-08-07 | GXIV | 838183 | Unit for PM. For module replacement of module B3. | Performed PM. Performed module replacement in B3. Generated reports | PM and module replacement. | Unit is operational. |  | GX Module B3 (210100852) in replacement to defective B3 (807120). | G6109 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673613346547346913/5673613346547346913_signature_28.png | Elvira alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5673613346547346913/5673613346547346913_signature_31.png
| 2023-08-09 | 2023-08-09 11:19 | Gandara Rural Health Unit | Gandara samar | 8 | Samar | Tb-dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25489 | DOH | sarah_yangzon@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 797 | 2023-08-09 | GXIV | 838406 | Module B1 SN: 766195 Defective | #NAME? | - Replaced module B1 w/ new SN: 760084 - Performed Preventive maintenance - Performed Functionality test - Generate reports | Unit is Operational |  | 1 Genexpert Module 6 Color SN: 760084 | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673607144021730744/5673607144021730744_signature_28.png | Sarah V. Yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5673607144021730744/5673607144021730744_signature_31.png
| 2023-08-09 | 2023-08-09 10:36 | Nabas Rural Health Unit | Nabas, Aklan | 6 | Aklan | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 25488 | PBSP | lambito.godelyn61@bmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 268 |  | GX4 | 815387 | Module B1 failed in last calibration | Conduct PM and generate files. | Replaced the old module for B1 with SN 696206 to new module with SN 210100410 | All modules are now functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5673580888015708211/5673580888015708211_signature_28.png | Godelyn A Lambito | https://www.jotform.com/uploads/macarewecare/222927978372067/5673580888015708211/5673580888015708211_signature_31.png
| 2023-08-08 | 2023-08-08 15:59 | Overseas Filipino Worker Hospital | San Fernando Pampanga | 3 | Pampanga | Laboratory | NO | ON-SITE ASSISTANCE Relocation and setup | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25487 | MACARE-HBDC | ofwhlaboratory@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 7 |  | GXIV-D | 110018556 | For relocation | Preventive maintenance | Relocate machine from Laboratory to 6TH Floor(training purposes) Performed preventive maintenance Gathered reports Relocate from 6th floor to laboratory Setup and installation of genexpert machine Archived test results Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672917936215942797/5672917936215942797_signature_28.png | Jayanara Capulong | https://www.jotform.com/uploads/macarewecare/222927978372067/5672917936215942797/5672917936215942797_signature_31.png
| 2023-08-08 | 2023-08-08 15:59 | Overseas Filipino Worker Hospital | San Fernando Pampanga | 3 | Pampanga | Laboratory | NO | ON-SITE ASSISTANCE Relocation and setup | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25486 | MACARE-HBDC | ofwhlaboratory@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 7 |  | GXIV-D | 110018556 | For relocation | Preventive maintenance | Relocate machine from Laboratory to 6TH Floor(training purposes) Performed preventive maintenance Gathered reports Relocate from 6th floor to laboratory Setup and installation of genexpert machine Archived test results Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672917916218796655/5672917916218796655_signature_28.png | Jayanara Capulong | https://www.jotform.com/uploads/macarewecare/222927978372067/5672917916218796655/5672917916218796655_signature_31.png
| 2023-08-08 | 2023-08-08 15:32 | Cagayan Valley Center for Health Development (DOH R02) | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25485 | PBSP | mctorio.pbsp@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2360 ASSAY: Xpert SARS-CoV-2, COUNT: 7155 | 2023-08-08 | Genexpert XVI | 842894 | Module A3 error | For Observation | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = (638053 & 888719 - 210100393 & 210100208) | Genexpert Module = (210100208 & 210100393) | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672898568221146551/5672898568221146551_signature_28.png | Mikeala Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5672898568221146551/5672898568221146551_signature_31.png
| 2023-08-08 | 2023-08-08 15:37 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25484 | DOH | elma.bragat@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3910 ASSAY: Xpert HIV-1 Viral Load, COUNT: 195 | 2023-07-31 | GX-IV | 838510 | FOR CALIBRATION | #NAME? | #NAME? | #NAME? |  | XPERTCHECK CE-IVD 5 CARTRIDGES/KIT | 10347 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672897002192086832/5672897002192086832_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5672897002192086832/5672897002192086832_signature_31.png
| 2023-08-08 | 2023-08-08 15:29 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25483 | DOH | elma.bragat@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3910 ASSAY: Xpert HIV-1 Viral Load, COUNT: 195 | 2023-07-31 | GX-IV | 838508 | FOR CALIBRATION | #NAME? | #NAME? | #NAME? |  | XPERTCHECK CE-IVD 5 CARTRIDGES/KIT | 10346 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672894182191762730/5672894182191762730_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5672894182191762730/5672894182191762730_signature_31.png
| 2023-08-08 | 2023-08-08 11:52 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | INSTALLATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25482 | MACARE-HBDC | lisandrogacias@gmail.com | 24 |  | 2023-08-08 | Genexpert IV | 110008457 | N/a | N/a | Performed Maintenance Installed Genexpert Performed Functionality Test | Unit is Operational, outlet no power because of renovation | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672892968226191153/5672892968226191153_signature_28.png | Lisandro Gacias | https://www.jotform.com/uploads/macarewecare/222927978372067/5672892968226191153/5672892968226191153_signature_31.png
| 2023-08-08 | 2023-08-08 11:52 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | INSTALLATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25481 | MACARE-HBDC | lisandrogacias@gmail.com | 24 |  | 2023-08-08 | Genexpert IV | 110008457 | N/a | N/a | Performed Maintenance Installed Genexpert Performed Functionality Test | Unit is Operational, outlet no power because of renovation | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672892438223731388/5672892438223731388_signature_28.png | Lisandro Gacias | https://www.jotform.com/uploads/macarewecare/222927978372067/5672892438223731388/5672892438223731388_signature_31.png
| 2023-08-08 | 2023-08-08 15:06 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25480 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 230 ASSAY: Xpert MTB/RIF, COUNT: 367 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2299 ASSAY: Xpert HIV-1 Viral Load, COUNT: 59 | 2023-07-31 | GXIV-4D | 838285 | Due for calibration. | #NAME? | #NAME? | All modules are functional. |  | XpertCheck-CE-5 | 10355 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672892265654882802/5672892265654882802_signature_28.png | Marc Rodney Manaban | https://www.jotform.com/uploads/macarewecare/222927978372067/5672892265654882802/5672892265654882802_signature_31.png
| 2023-08-08 | 2023-08-08 15:23 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25479 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 229 ASSAY: Xpert MTB/RIF, COUNT: 339 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2411 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 | 2023-07-31 | GXIV-4D | 820206 | Due for calibration | -Run XpertCheck calibration on module A1, A2, & A4 | #NAME? | Module A2 xpertcheck service required, for replacement. Module A1 and A4 passed and are functional. |  | XpertCheck-CE-5 | 10356 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672889865656729840/5672889865656729840_signature_28.png | Marc Rodney Manaban | https://www.jotform.com/uploads/macarewecare/222927978372067/5672889865656729840/5672889865656729840_signature_31.png
| 2023-08-08 | 2023-08-08 14:54 | Batasan Super Health Center | Batasan Rd., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25478 | PBSP | ashwinkevin@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 835 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2178 |  | GXIV 6C | 805254 | Module 4 temp problem | Performed preventive maintenance Checked all modules Module 4 for reporting Generated report | Performed PM Gathered report | Unit is operational. Only module 4 have temp problem |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672886833184097281/5672886833184097281_signature_28.png | Kevin Orlando D. Cabuso | https://www.jotform.com/uploads/macarewecare/222927978372067/5672886833184097281/5672886833184097281_signature_31.png
| 2023-08-08 | 2023-08-08 14:29 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25477 | PBSP | jvdiaz.pbsp@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 27 ASSAY: Xpert HPV, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2370 ASSAY: Others, COUNT: 8058 ASSAY: Xpert HIV-1 Viral Load, COUNT: 159 ASSAY: Others, COUNT: 9 | 2023-08-08 | GXIV | 834432 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. A1: failed A2: failed A3: retry A4: pass |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672875951317794600/5672875951317794600_signature_28.png | Vianne Diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5672875951317794600/5672875951317794600_signature_31.png
| 2023-08-08 | 2023-08-08 14:35 | West Visayas State University Medical Center | Jaro, Iloilo City | 6 | Iloilo | Molecular Laboratory | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25476 | DOH | wvsumc.genex@gmail.com | 4 | ASSAY: Xpert HPV, COUNT: 403 |  | GX4 | 838290 | None | Conduct PM and generate files. | Run Xpert check calibration for 4 modules | All modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672875676194990081/5672875676194990081_signature_28.png | Karla Danielle T. Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/5672875676194990081/5672875676194990081_signature_31.png
| 2023-08-08 | 2023-08-08 14:33 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25475 | PBSP | jvdiaz.pbsp@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Xpert HBV VL, COUNT: 1 ASSAY: Xpert HPV, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 967 ASSAY: Xpert SARS-CoV-2, COUNT: 2803 ASSAY: Xpert HIV-1 Viral Load, COUNT: 59 | 2023-08-08 | GXIV | 834431 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational D1: retry D2: passed D3: passed D4: failed |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672874811319234761/5672874811319234761_signature_28.png | Vianne Diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5672874811319234761/5672874811319234761_signature_31.png
| 2023-08-08 | 2023-08-08 14:07 | Matanao Rural Health Unit | Poblacion, Matanao, Davao del Sur | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25474 | DOH | pamsustento@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 493 |  | GX-IV 4D | 838456 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672844423117480345/5672844423117480345_signature_28.png | Pamela Joyce P. Sustento RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5672844423117480345/5672844423117480345_signature_31.png
| 2023-08-08 | 2023-08-08 13:36 | Balanga Rural Health Unit | Sampaguita St. San Jose, Balanga City, Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25473 | PBSP | cho.melaniealonzo@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 752 | 2023-07-07 | GXIV-D | 835102 | B1 for module replacement Date and time not synching CPU hard booting up | Module replacement Troubleshooting | Performed preventive maintenance Troubleshoot CPU Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is operational | CPU BIOS for replacement | 1 Genexpert module | 42435 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672834153634358290/5672834153634358290_signature_28.png | Melanie H. Alonzoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5672834153634358290/5672834153634358290_signature_31.png
| 2023-08-08 | 2023-08-08 13:44 | Democrito O. Plaza Memorial Hospital RTDL | Brgy. Patin-ay, Properidad | 13 | Agusan del Sur | RTDL | NO | INSTALLATION REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 25472 | PBSP | lermabpabillan@yahoo.com | 23.86 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1419 |  | GX-4 | 805335 | Module B3 found E2014 | None | Replacing B1 with SN: 803321 with SN: 210100436 Replacing Module B2 with SN: 695878 to SN: 210100437 Installation of New Lenovo CPU with SN: CPH0522F37028 and peripherals. Installation of drivers and assays. | 3 modules working, 1 module reported to jotform |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672832376811538951/5672832376811538951_signature_28.png | Lerma B. Pabillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5672832376811538951/5672832376811538951_signature_31.png
| 2023-08-08 | 2023-08-08 13:20 | Carcar City Health Office | Valladollid Carcar City | 7 | Cebu | Tb Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25471 | DOH | chaska.lp14@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1570 |  | Gx4 | 838283 | Due for Calibration | Conduct calibration | Performed calibration for B1, B3 and B4 | All modules are operational and B1, B3, and B4 passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672818013752136575/5672818013752136575_signature_28.png | Lucille Mae A. Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5672818013752136575/5672818013752136575_signature_31.png
| 2023-08-08 | 2023-08-08 12:18 | Camarin Health Center | Camarin, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25470 | PBSP | jasminfalala@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1730 | 2023-08-08 | GXIV | 831895 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. B1: passed B2: passed B3: passed B4: passed |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672805011314067346/5672805011314067346_signature_28.png | Arianne Jasmin Roma | https://www.jotform.com/uploads/macarewecare/222927978372067/5672805011314067346/5672805011314067346_signature_31.png
| 2023-08-08 | 2023-08-08 10:11 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25469 | DOH | laureanjo@yahoo.com | 28 |  |  | Gx 4 | 838527 | Due date for xpertcheck calibration | Running xpertcheck calibration for module D3 & D4 | Running xpertcheck calibration for module D3 & D4 | Module D3 service required, module D4 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672779163856076233/5672779163856076233_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5672779163856076233/5672779163856076233_signature_31.png
| 2023-08-08 | 2023-08-08 12:10 | Iloilo Provincial Hospital | Pototan, Iloilo | 6 | Iloilo | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25468 | DOH | djpragale92@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 48 |  | GX4 | 838289 | None | Conduct PM and generate files | Run Xpert check calibration for 4 modules | All modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672778936198672082/5672778936198672082_signature_28.png | Rachel Pragale | https://www.jotform.com/uploads/macarewecare/222927978372067/5672778936198672082/5672778936198672082_signature_31.png
| 2023-08-08 | 2023-08-08 10:46 | Don Jose S. Monfort Medical Center Ext. Hospital | Barotac Nuevo | 6 | Iloilo | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25467 | DOH | shirleymaepecto@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 635 |  | GX4 | 838501 | None | Conduct PM and Generate files | Run Xpert check calibration for 4 modules | All modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672736646197552184/5672736646197552184_signature_28.png | SHIRLEY MAE G. PECTO | https://www.jotform.com/uploads/macarewecare/222927978372067/5672736646197552184/5672736646197552184_signature_31.png
| 2023-08-08 | 2023-08-08 10:17 | Northern Samar Provincial Hospital | Catarman | 8 | Northern Samar | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25466 | PBSP | dinahada70@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 110 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2194 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 | 2023-08-08 | GXIV | 807566 | Unit for Calibration | #NAME? | #NAME? | Unit is Operational (Module B1 - service required) |  | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672707848854469728/5672707848854469728_signature_28.png | Dinah F. Ada | https://www.jotform.com/uploads/macarewecare/222927978372067/5672707848854469728/5672707848854469728_signature_31.png
| 2023-08-08 | 2023-08-08 09:47 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25465 | DOH | ccmclaboratory01@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 146 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1563 | 2023-07-27 | GXIV-4D | 838561 | Due for calibration | #NAME? | #NAME? | -Module A3 passed XPC calibration -Module A1 disabled; waiting for replacement -Other modules are functional. |  | XpertCheck-CE-5 | 10328 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672688125651914264/5672688125651914264_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5672688125651914264/5672688125651914264_signature_31.png
| 2023-08-07 | 2023-08-07 15:55 | Bago City Health Office | Bago City | 6 | Negros Occidental | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25464 | DOH | swabteambago@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 259 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2311 |  | GX4 | 838486 | Module B1 already reported for replacement and waiting for allocation. | Conduct PM and generate files. | Run Xpert check calibration for module B2, B3 and B4. | Module B3 failed in calibration and module B2 & B4 passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672059160528174953/5672059160528174953_signature_28.png | Glenly Cansona | https://www.jotform.com/uploads/macarewecare/222927978372067/5672059160528174953/5672059160528174953_signature_31.png
| 2023-08-07 | 2023-08-04 16:17 | National kidney and transplant institute | East ave QC | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 25463 | MACARE-HBDC | molecularlabnkti@gmail.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 690 ASSAY: Others, COUNT: 11 ASSAY: Xpert HBV VL, COUNT: 422 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1230 ASSAY: Xpert MTB/RIF Ultra, COUNT: 753 ASSAY: Xpert SARS-CoV-2, COUNT: 16565 ASSAY: Xpert HCV VL, COUNT: 37 ASSAY: Others, COUNT: 9 |  | GXXVI | 110005587 | B1 encounter loss communication error | Replaced cable Swap module location Performed self test multiple times Gathered reports Test run unit | Troubleshooting | B1 no error loss communication encounter during module swapping location |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672041863615479849/5672041863615479849_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5672041863615479849/5672041863615479849_signature_31.png
| 2023-08-07 | 2023-08-03 14:12 | Philippine Heart Center | East Ave. QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25462 | PBSP | phcmolecularlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 444 ASSAY: Others, COUNT: 5 ASSAY: Xpert SARS-CoV-2, COUNT: 2487 | 2023-07-31 | GX4 | 834443 | For module replacement B2 210068732 - 210100435 | Cleaned unit Cleaned filter,plunger rod ,cartridge bay,icore Gathered reports Test run unit | Module replacement and Maintenance | Operational |  | 900-0508 | G6011 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672034763612357339/5672034763612357339_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5672034763612357339/5672034763612357339_signature_31.png
| 2023-08-07 | 2023-08-03 10:54 | East Avenue Medical Center | East Ave.QC | NCR | NCR 2nd - Quezon City | Tb dots | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 25461 | PBSP | roseannparnes@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3614 |  | GX4 | 803087 | C2 and C3 encounter ng continous 5007 error | Cleaned plunger rod Cleaned icore,cartridge bay Gathered reports | Troubleshooting | C2 and C3 for observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672031353618685143/5672031353618685143_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/5672031353618685143/5672031353618685143_signature_31.png
| 2023-08-07 | 2023-08-02 15:23 | The Medical City | Ortigas ave.Pasig City | NCR | NCR 2nd - Pasig | Molecular | NO | Relocation | CEPHEID | Noel Floresca |  |  |  |  | 25460 | COMMERCIAL | moleculardx@themedicalcity.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 10 ASSAY: Xpert CT/NG, COUNT: 13 ASSAY: Xpert HBV VL, COUNT: 175 ASSAY: Xpert HPV, COUNT: 133 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1284 ASSAY: Others, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 23 ASSAY: Xpert HIV-1 Viral Load, COUNT: 238 |  | GXXVI | 110015767 | For transfer | Transfer unit from another table, reinstall and set up unit | Transfer unit | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5672029833612218458/5672029833612218458_signature_28.png | Kristine Chelle Parilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5672029833612218458/5672029833612218458_signature_31.png
| 2023-08-07 | 2023-08-07 14:05 | Quirino Memorial Medical Center | P. Tuazon, Quezon City | NCR | NCR 2nd - Quezon City | Blood Bano | NO | TROUBLESHOOTING | HAEMONETICS | Ronnel Belizar |  |  |  |  | 25459 | COMMERCIAL | moncastillo11@gmail.com | 25 |  |  | MCS+ | 19G059 | Line Sensor Error | - checked calibration parameters (passed) - cleaned line sensor (value: 2900) - performed functionality test | Cleaned Line Sensor | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671978816318645459/5671978816318645459_signature_28.png | Raymond Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5671978816318645459/5671978816318645459_signature_31.png
| 2023-08-07 | 2023-08-07 13:40 | Binalbagan Rural Health Unit | Binalbagan, Negros Occidental | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25458 | DOH | ntpbinalbagan@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 969 |  | GX4 | 838280 | None | Conduct PM and generate files | Run Xpert check calibration for 4 modules | All modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671976778027948696/5671976778027948696_signature_28.png | Mary Flor Egca | https://www.jotform.com/uploads/macarewecare/222927978372067/5671976778027948696/5671976778027948696_signature_31.png
| 2023-08-07 | 2023-07-26 13:27 | Nagpayong Super Health Center | Nagpayong, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25457 | DOH | michelindg23@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 567 | 2023-07-24 | GX-IV-D | 831905 | Module B2 and B4 for replacement | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | Unit is operational. Module B2 not operational reported already. |  | Old: B3/786244, B4/758212 New: B3/210017770, B4/210019231 | DR# 10254, 10253 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671966063612289278/5671966063612289278_signature_28.png | Michelle D. De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5671966063612289278/5671966063612289278_signature_31.png
| 2023-08-07 | 2023-07-26 13:27 | Nagpayong Super Health Center | Nagpayong, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25456 | DOH | michelindg23@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 567 | 2023-07-24 | GX-IV-D | 831905 | Module B2 and B4 for replacement | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | Unit is operational. Module B2 not operational reported already. |  | Old: B3/786244, B4/758212 New: B3/210017770, B4/210019231 | DR# 10254, 10253 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671966043618222324/5671966043618222324_signature_28.png | Michelle D. De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5671966043618222324/5671966043618222324_signature_31.png
| 2023-08-07 | 2023-08-04 11:37 | GMA Rural Health Unit II | Brgy. Memije, Congressional Road, General Mariano Alvarez | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25455 | PBSP | ldlconti@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1735 |  | GXIV-D | 814152 | Module 4 with Temperarure Errors | Module Replacement | Performed Maintenance Replaced Module 4 Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959443617839097/5671959443617839097_signature_28.png | Lorena L. Conti | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959443617839097/5671959443617839097_signature_31.png
| 2023-08-07 | 2023-08-04 09:37 | Dasmariñas City Health Office II | DBB-E, Dasmariñas | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25452 | DOH | hazelmilan0777@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2331 |  | GXIV-D | 838535 | Module 3 with Tempetarure Error | Module Replacement | Performed Maintenance Replaced Module 3 Generate Reports | 3 Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373611413300/5671959373611413300_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373611413300/5671959373611413300_signature_31.png
| 2023-08-07 | 2023-08-03 13:38 | General Trias City Health Office | Hospital Rd, General Trias | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25453 | DOH | loremontoya1169@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2122 ASSAY: Xpert HIV-1 Viral Load, COUNT: 211 | 2023-07-25 | GXIV-D | 821340 | Module 4 with Temperature Error Module 1 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 4 Generate Reports | Mod 1 Not Available in XpertCheck, 3 Modules are Operational |  | 900-0491R | DR 10283 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373612646756/5671959373612646756_signature_28.png | Lorena R Montoya | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373612646756/5671959373612646756_signature_31.png
| 2023-08-07 | 2023-08-03 13:18 | General Trias City Health Office | Hospital Road, General Trias | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25454 | PBSP | loremontoya1169@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2116 ASSAY: Xpert HIV-1 Viral Load, COUNT: 211 |  | GXIV-D | 804691 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 1 Failed, 3 Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373614897951/5671959373614897951_signature_28.png | Lorena R. Montoya | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373614897951/5671959373614897951_signature_31.png
| 2023-08-07 | 2023-08-03 11:33 | Rosario Rural Health Unit | Scollt Torillo, Rosario | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25454 | DOH | rhu.rosario@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 997 | 2023-07-25 | GXIV-D | 838537 | Modules 1 and 4 with Temperature Errors | Module Replacement | Performed Maintenance Replaced Modules 1 and 4 Generate Reportd | 3 Modules are Operational |  | 900-0491R | DR's 10278,10250 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373615983084/5671959373615983084_signature_28.png | Jenilyn dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373615983084/5671959373615983084_signature_31.png
| 2023-08-07 | 2023-08-02 11:28 | Kawit Rural Health Unit | Tabon 2, Kawit | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25451 | DOH | mariellepanganiban1004@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1025 | 2023-07-25 | GXIV-D | 820705 | Modules 3 and 4 with Temperature Error | Module Replacement | Performed Maintenance Replaced Modules 3 and 4 | All Modules are Operational |  | 900-0491R | DR's 10285, 10288 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373618469057/5671959373618469057_signature_28.png | Marielle Panganiban | https://www.jotform.com/uploads/macarewecare/222927978372067/5671959373618469057/5671959373618469057_signature_31.png
| 2023-08-07 | 2023-08-07 13:24 | Molave Rural Health unit | Molave, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25450 | DOH | aldrinacain@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 774 | 2023-07-21 | GX 4 | 838233 | Module 1, 3 and 4 temperature error | Replaced module 807256 and Module 807261 with Module 883311 and 890829 respectively. Module 3 undergone Xpertcheck calibration and the result is service required, must be replaced. | Module Replacement and Xpertcheck calibration | Module 1, 2 and 4 are available, module 3 must be repaired. | None | 900-0491R Genexpert Module 6 color S.N (883311 / 890829) | 10223/ 10224 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671955791315582823/5671955791315582823_signature_28.png | Aldrin Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/5671955791315582823/5671955791315582823_signature_31.png
| 2023-08-07 | 2023-08-07 13:15 | Mindanao central sanitarium | Pasabolong Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25449 | DOH | ybioamie@gmail.com | 23 |  | 2023-07-21 | GX 4 | 821214 | Module 1 hardware failed | Replaced module 695973 with Module 909139 | Module Replacement and Preventive Maintenence | Module 1, 3 and 4 are available. Module 2 must be replaced | Module 2 waiting for module Replacement | 900-0491R Genexpert Module 6 S.N. 909139 | 10226 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671950211311862502/5671950211311862502_signature_28.png | Amie b. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5671950211311862502/5671950211311862502_signature_31.png
| 2023-08-07 | 2023-08-07 11:47 | Pontevedra Rural Health Unit | Pontevedra, Negros Occidental | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25448 | DOH | pontevedra.laboratory@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 747 |  | GX4 | 838491 | None | Conduct PM and generate files(IQR & GXC File) | Run Xpert check calibration for module B3 and B4 only. | Module B3 and B4 passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671909398025888886/5671909398025888886_signature_28.png | Mae Ann Bonilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5671909398025888886/5671909398025888886_signature_31.png
| 2023-08-07 | 2023-08-02 10:41 | Corazon Locsin Montelibano Memorial Regional Hospital | Lacson Street, Bacolod City | 6 | Negros Occidental | Micro Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25447 | DOH | jennybellegallego@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 803 ASSAY: Xpert HIV-1 Viral Load, COUNT: 833 |  | GX4 | 821219 | Module A3 with error 4009 | Conduct PM and save files | Run Xpert check for 4 modules | All modules passed the Xpert check calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671898258027877687/5671898258027877687_signature_28.png | Jennybelle Gallego | https://www.jotform.com/uploads/macarewecare/222927978372067/5671898258027877687/5671898258027877687_signature_31.png
| 2023-08-07 | 2023-08-02 10:41 | Corazon Locsin Montelibano Memorial Regional Hospital | Lacson Street, Bacolod City | 6 | Negros Occidental | Micro Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25446 | DOH | jennybellegallego@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 803 ASSAY: Xpert HIV-1 Viral Load, COUNT: 833 |  | GX4 | 821219 | Module A3 with error 4009 | Conduct PM and save files | Run Xpert check for 4 modules | All modules passed the Xpert check calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671897938026619561/5671897938026619561_signature_28.png | Jennybelle Gallego | https://www.jotform.com/uploads/macarewecare/222927978372067/5671897938026619561/5671897938026619561_signature_31.png
| 2023-08-06 | 2023-08-04 15:31 | La Union Medical Center | Agoo, La Union | 1 | La Union | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25445 | PBSP | karizzamilomyg@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1911 | 2023-08-04 | GX4 | 707400 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective Module 628666 >>> 210100327 - Generated Reports | Performed Module Replacement and Maintenance | Unit is operational |  | GeneXpert Module 210100327 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671096038424317508/5671096038424317508_signature_28.png | Karizza Cyriene Milo | https://www.jotform.com/uploads/macarewecare/222927978372067/5671096038424317508/5671096038424317508_signature_31.png
| 2023-08-06 | 2023-08-04 12:07 | Naguilian District Hospital | Naguilian, La Union | 1 | La Union | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25444 | DOH | marjelynrayray01@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 725 | 2023-06-13 | GX4 | 838426 | Due for Calibration | #NAME? | Performed XpertCheck | Unit is Operational | All modules passed | XpertCheck Kits | DR 5340 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5671095098421137018/5671095098421137018_signature_28.png | Marjelyn Rayray | https://www.jotform.com/uploads/macarewecare/222927978372067/5671095098421137018/5671095098421137018_signature_31.png
| 2023-08-06 | 2023-08-06 08:20 | Eastern Samar Provincial Hospital | Borongan city | 8 | Eastern Samar | Tb - dots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  |  |  |  | 25443 | PBSP | elvalynlom@gmail.com | 33 | ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1099 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 | 2023-08-03 | GXIV | 804256 | Module B2 received multiple temperature error | #NAME? | Troubleshoot - check the ambient temperature (above 30°C) - check the room air conditioning ( not working properly due to low voltage and current fluctuating) - Performed Preventive maintenance - Generate reports | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5670909072815400524/5670909072815400524_signature_28.png | Elvalyn Gracia G. Lomuntad | https://www.jotform.com/uploads/macarewecare/222927978372067/5670909072815400524/5670909072815400524_signature_31.png
| 2023-08-04 | 2023-08-04 16:03 | Makati Medical Center (medilinx) | Amorsolo st., Makati city | NCR | NCR 4th - Makati | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25442 | COMMERCIAL | john.manzano@sodexo.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 89 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 416 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 448 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1616 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 34 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1215 |  | GX4 (10c) | 110001647 | Due for Maintenance, Broken door hinges of B3 module | Cleaned unit, Cleaned filter,plunger rod,icore,cartridge bay Gathered reports Checked b3 module door hinge | Maintenance | B3 broken door hinge for repair |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669458380019340411/5669458380019340411_signature_28.png | Medi Linx Laboratory Inc | https://www.jotform.com/uploads/macarewecare/222927978372067/5669458380019340411/5669458380019340411_signature_31.png
| 2023-08-04 | 2023-08-04 15:28 | Ospital ng Lungsod ng San Jose Del Monte | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25441 | PBSP | arvalerio21@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1809 | 2023-08-04 | GXIV | 807624 | Unit for XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. B1: failed B2: passed B3: failed B4: failed |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669449734259921554/5669449734259921554_signature_28.png | Arnold Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5669449734259921554/5669449734259921554_signature_31.png
| 2023-08-04 | 2023-08-04 15:31 | Gubat Rural Health Unit | Highway 59, Brgy. Ariman, Gubat | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25440 | PBSP | catafescoto@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1218 |  | GX4 | 814370 | B2 for module replacement | Replaced module: 691584>>210100697 Cleaned cartridge bays, plunger rod, icore,fan, replace fan filter Generate reports | B2 replaced module Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669445339124951002/5669445339124951002_signature_28.png | Catalina F.Escoto | https://www.jotform.com/uploads/macarewecare/222927978372067/5669445339124951002/5669445339124951002_signature_31.png
| 2023-08-04 | 2023-08-04 15:02 | Asian Hospital And Medical Center | Alabang, Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular Laboratory | NO | Pull Out GX Machine | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25439 | COMMERCIAL | molecularpathology@asianhospital.com | 22 | ASSAY: Xpert HBV VL, COUNT: 12 ASSAY: Xpert HPV, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1021 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 ASSAY: Others, COUNT: 32 |  | GXIV 6C | 803868 | GX machine for pull out | Archive Test Generated report Disassemble GX Packed GX machine and laptop | Pulled out GX machine | Unit to be returned to office |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669423961818663420/5669423961818663420_signature_28.png | Trisha Faye Pastor | https://www.jotform.com/uploads/macarewecare/222927978372067/5669423961818663420/5669423961818663420_signature_31.png
| 2023-08-04 | 2023-08-04 14:11 | Casiguran Rural Health Unit | Casiguran Aurora | 3 | Aurora | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25438 | DOH | cezalen@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 233 ASSAY: Xpert MTB/RIF Ultra, COUNT: 536 |  | GXIV-D | 831844 | A1,A2 AND A3 For re-calibration and troubleshooting | Re-xpertcheck Module inspection Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is fully operational, module A1, A2 and A3 passed calibration | Used 5 xpertcheck kits, A1, A2 and A3 for observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669404472899871075/5669404472899871075_signature_28.png | Princess Alen | https://www.jotform.com/uploads/macarewecare/222927978372067/5669404472899871075/5669404472899871075_signature_31.png
| 2023-08-04 | 2023-08-04 14:27 | DOH TRC | Camp bagong diwa bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25437 | PBSP | dohtrcbicutan.lab@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2115 | 2023-07-31 | GX4 | 807452 | For module replacement C2 609195 - 210100544 C4 666170 - 210100547 | Performed module replacement Cleaned filter,plunger rod,cartridge bay,icore Gathered reports Performed troubleshooting C3 module | Module replacement and maintenance | C3 disabled |  | 900-0508 | G6016 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669399540013506293/5669399540013506293_signature_28.png | April Chakee Javier | https://www.jotform.com/uploads/macarewecare/222927978372067/5669399540013506293/5669399540013506293_signature_31.png
| 2023-08-04 | 2023-08-04 13:46 | Bagumbong Health Center | Bagumbong, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25436 | DOH | jenilyn_duenas@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2435 | 2023-07-26 | GXIV | 831907 | GeneXpert monitor for replacement | #NAME? | Replaced monitor | Unit is operational. |  | GeneXpert Monitor SN: MMA1PSS00221507B763W01 | 10292 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669375884611126152/5669375884611126152_signature_28.png | Jenilyn Duenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5669375884611126152/5669375884611126152_signature_31.png
| 2023-08-04 | 2023-08-04 12:04 | Saguiaran Rural Health Unit | Saguiaran | BARMM | Lanao del Sur | TB DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25435 | DOH | saharaalimaradial.muti@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 369 |  | GeneXpert | 838298 | None | None | Preventive maintenance Module replacement Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669353531021744066/5669353531021744066_signature_28.png | Sahara Muti | https://www.jotform.com/uploads/macarewecare/222927978372067/5669353531021744066/5669353531021744066_signature_31.png
| 2023-08-04 | 2023-08-04 13:08 | Tomas Cabili Health Center | Tomas Cabili, Iligan City | 10 | Lanao del Norte | TB DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25434 | DOH | mariacielovenezuela@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 932 |  | GeneXpert | 838294 | None | None | Preventive Maintenance Module replacement Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669351391021699985/5669351391021699985_signature_28.png | Maria Cialo C Venezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5669351391021699985/5669351391021699985_signature_31.png
| 2023-08-04 | 2023-08-04 11:29 | Sorsogon Medical Mission Group Hospital | Pampang, West District, Sorsogon City | 5 | Sorsogon | RTDL Culture Lab | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25433 | PBSP | lotino.ana@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1324 |  | GX4 R1 | 707452 | B4 for module replacement | Replaced module: 210068733>>210101745 Performed XpertCheck Cleaned cartridge bays, plunger rod, icores Generate reports | Replaced module B4 Performed XpertCheck Performed Preventive Maintenance | Unit is operational B1 - Passed B2 - Passed B3 - Passed B4 - New Module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669313062424565991/5669313062424565991_signature_28.png | Ana Liza Lotino | https://www.jotform.com/uploads/macarewecare/222927978372067/5669313062424565991/5669313062424565991_signature_31.png
| 2023-08-04 | 2023-08-03 13:05 | Bicol Regional General Hospital and Geriatric Medical Center | San Pedro, Cabusao | 5 | Camarines Sur | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25432 | PBSP | jlvdvera_rmt@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1807 |  | GX4 | 807601 | B2/662117 for module replacement | Replaced module: 662117>>210100434 Performed XpertCheck Cleaned cartridge bays, plunger rod, icore,fan,fan filter Generate reports | Replaced module B2 Performed XpertCheck Performed preventive maintenance | Unit is operational B1 - Passed B2 - New Module B3 - Passed B4 - Service Required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669311732428770431/5669311732428770431_signature_28.png | Jessie Lowell de Vera | https://www.jotform.com/uploads/macarewecare/222927978372067/5669311732428770431/5669311732428770431_signature_31.png
| 2023-08-04 | 2023-08-04 11:24 | Narra MHO | Narra Palawan | 4B | Palawan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 25431 | PBSP | esg110478@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 839 | 2023-08-04 | Gx4 | 807454 | Module B3 for replacement. | Replaced module B3. SN: 661895 | Performed maintenance. Replaced module B3 | All 4 modules operational |  | New mod SN: 210092904 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669290260412019717/5669290260412019717_signature_28.png | ELIZABETH S. GAPULAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5669290260412019717/5669290260412019717_signature_31.png
| 2023-08-04 | 2023-08-03 13:40 | Bugasong Rural Health Unit | Bugasong, Antique | 6 | Antique | Tb Dots | YES | INSTALLATION TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 25430 | PBSP | jennifer.pasaquian2485@gmail.com | 25 |  | 2023-08-03 | GXIV-4D | 815427 | CPU for replacement. Module B1 and B4 encountering ultrasonic horn error. | #NAME? | #NAME? | All modules are functional. |  | 91337 Cepheid GX Global M70S Lenovo ThinkCentre | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5669162135572407770/5669162135572407770_signature_28.png | Jennifer Contante | https://www.jotform.com/uploads/macarewecare/222927978372067/5669162135572407770/5669162135572407770_signature_31.png
| 2023-08-03 | 2023-08-03 12:56 | Southern Philippines Medical Center | J.P Laurel, Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25429 | PBSP | spmcrhivda@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1610 ASSAY: Xpert HIV-1 Viral Load, COUNT: 584 |  | GX-IV 4D | 808298 | Due for Maintenance and Calibration | Clean Surface Area, Cartridge Bays, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log Run Xpertcheck on Modules | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668775855715178413/5668775855715178413_signature_28.png | Flordeliza Go. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5668775855715178413/5668775855715178413_signature_31.png
| 2023-08-03 | 2023-08-03 15:00 | Kidapawan City Health Office | Kidapawan City, North Cotabato | 12 | Cotabato | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25428 | DOH | manz.franz@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2187 | 2023-08-03 | Gx-IV | 838545 | Module 3 H/W failed | Replace module 3 SN: 807311 with new module SN: 775505 | Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 color SN:775505 | Dr number: 10234 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668724824126170861/5668724824126170861_signature_28.png | Ma. Luisa Francisco | https://www.jotform.com/uploads/macarewecare/222927978372067/5668724824126170861/5668724824126170861_signature_31.png
| 2023-08-03 | 2023-08-03 15:59 | Pasig City General Hospital | F. Legazpi Extension, Maybunga | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25427 | PBSP | vilmarumadhay@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3207 | 2023-08-03 | GXIV 6C | 805010 | Gateway Board for upgrade | Performed PM Replaced Gateway board Changed Gateway board serial number Generated report | Replaced gateway board with a new one | Unit is operational |  | Gateway Board Kit | G 6099 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668593970891046793/5668593970891046793_signature_28.png | Vilma R. Umadhay | https://www.jotform.com/uploads/macarewecare/222927978372067/5668593970891046793/5668593970891046793_signature_31.png
| 2023-08-03 | 2023-08-03 14:23 | Bustos Rural Health Unit | Bustos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25426 | DOH | ronald_galvez2002@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 645 | 2023-07-24 | GXIV | 831874 | Module B2 for replacement | #NAME? | Replaced B2 | Unit is operational. |  | GeneXpert Module SN: 619121 | 10264 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668532399326182198/5668532399326182198_signature_28.png | Ronel Julian | https://www.jotform.com/uploads/macarewecare/222927978372067/5668532399326182198/5668532399326182198_signature_31.png
| 2023-08-03 | 2023-08-03 11:49 | Norzagaray Rural Health Unit | Norzagaray, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25425 | PBSP | chelseajempolicarpio@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1101 | 2023-07-24 | GXIV | 838167 | Module B4 for replacement | #NAME? | Replace module B4 | Unit is operational. |  | GeneXpert Module SN: 210100379 | G 5896 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668468199327880313/5668468199327880313_signature_28.png | Chelsea Policarpio | https://www.jotform.com/uploads/macarewecare/222927978372067/5668468199327880313/5668468199327880313_signature_31.png
| 2023-08-03 | 2023-08-03 11:58 | Tanauan City Health Office | Tanauan Batangas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25424 | DOH | ramilomher27@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1663 | 2023-07-24 | GXIV | 821299 | Unit for PM. For module replacement of modules A1 and A4. | Performed PM. Performed module replacement on A1 and A4. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module A1 (210012231) and A4 (210012390) in replacement to defective A1 (731303) and A4 (760160). | DR#s 10249 & 10251 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668446171024125245/5668446171024125245_signature_28.png | Maria Teresita Labsan | https://www.jotform.com/uploads/macarewecare/222927978372067/5668446171024125245/5668446171024125245_signature_31.png
| 2023-08-03 | 2023-08-03 11:12 | Lapulapu City Health Office | Lapulapu City, Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 25423 | PBSP | tanya_pax@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2124 |  | Gx4 | 808637 | Error 2126 module is keep reseting unavailable, botting | Reset module B1 | Reset module B1 and reset power supply and gateway board | B1 is restored. All modules are available except B4 waiting for next module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668443396421743214/5668443396421743214_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5668443396421743214/5668443396421743214_signature_31.png
| 2023-08-03 | 2023-08-03 11:06 | Angeles Rural Health Unit VI | EPZA Pulung Cacutod,Angeles City Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25422 | PBSP | rosieendozo@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 819 | 2023-07-06 | GXIV-D | 838870 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration | 4 xpertcheck kots | 41814 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668425204615242019/5668425204615242019_signature_28.png | Rosita C. Endozo | https://www.jotform.com/uploads/macarewecare/222927978372067/5668425204615242019/5668425204615242019_signature_31.png
| 2023-08-03 | 2023-08-03 10:21 | Philippine Tuberculosis Society, Inc. | PTSI Compound, E. Rodriguez Sr. Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25421 | PBSP | PTSICENLAB@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 219 | 2023-07-07 | GXIV 6C | 707527 | Unit for XpertCheck | Performed PM Perfomed XpertCheck Generated Report | Performed XpertCheck | Unit is operational Module A4(621280) : Failed XpertCheck |  | XpertCheck Kit | DR 41854 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668398680213267745/5668398680213267745_signature_28.png | Peter Samonte | https://www.jotform.com/uploads/macarewecare/222927978372067/5668398680213267745/5668398680213267745_signature_31.png
| 2023-08-03 | 2023-08-03 09:37 | Eversly Childs Sanitarium | Jagobiao mandaur city | 7 | Cebu | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25420 | PBSP | alimurungjane@gmail.com | 28.7 |  |  | Gx 4 | 805270 | Module A3, Error 2126 module was reset | Reseting ,cleaning and swapping UTP cable for module 3 | Reseting ,cleaning and swapping UTP cable for module 3 | Systems operational, under obeservation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668365585115988720/5668365585115988720_signature_28.png | Juanita v. Alimurung | https://www.jotform.com/uploads/macarewecare/222927978372067/5668365585115988720/5668365585115988720_signature_31.png
| 2023-08-03 | 2023-08-03 09:28 | Bauan Rural Health Unit II | La Paz Ext. Road, Bauan | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25419 | DOH | rhubauan2@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1273 | 2023-07-24 | GXIV | 820204 | Unit for PM. For module replacement of modules A1 and A4. | Performed PM. Performed module replacement on A1 and A4. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module A1 (210037059) and A4 (210039499) in replacement to defective A1 (721909) and A4 (721920). | DR#s 10256 & 10257 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5668357896167546604/5668357896167546604_signature_28.png | Jeanette C De Chavez | https://www.jotform.com/uploads/macarewecare/222927978372067/5668357896167546604/5668357896167546604_signature_31.png
| 2023-08-02 | 2023-08-02 12:25 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | PCR Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25418 | PBSP | honeyeichel@gmail.com | 26 | ASSAY: Xpert SARS-CoV-2, COUNT: 16777 |  | GXIV-D | 834433 | C3 for module replacement | Module replacement | Performed preventive maintenance Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667854727615605486/5667854727615605486_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5667854727615605486/5667854727615605486_signature_31.png
| 2023-08-02 | 2023-08-02 16:38 | Pamplona RHU | Zone-2, Tambo, Pamplona | 5 | Camarines Sur | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25417 | PBSP | francislozano226@gmail.com | 26 |  |  | GX4 | 812774 | For CPU Replacement | CPU Replaced: CPH0622F38599 Performed XpertCheck Cleaned cartridge bays, plunger rod, icore Generate Reports | Replaced CPU PERFORMED XPERTCHECK PERFORMED PREVENTIVE MAINTENANCE | Unit is Operational A1 - Service Required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667760798426440559/5667760798426440559_signature_28.png | Francis Lozano | https://www.jotform.com/uploads/macarewecare/222927978372067/5667760798426440559/5667760798426440559_signature_31.png
| 2023-08-02 | 2023-08-02 15:35 | Morong Rural Health Unit | Burgps St. Poblacion, Morong, Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25416 | DOH | lielyn.lim@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 43 ASSAY: Xpert MTB/RIF Ultra, COUNT: 383 | 2023-07-17 | GXIV-D | 831607 | A1 for module replacement | Module replacement | Performed preventive maintenance Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is fully operational |  | 1 genexpert module | 10139 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667710419747493949/5667710419747493949_signature_28.png | Lielyn M. Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5667710419747493949/5667710419747493949_signature_31.png
| 2023-08-02 | 2023-08-02 14:55 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25415 | DOH | laboratory@loveyourself.ph | 26 | ASSAY: Xpert CT/NG, COUNT: 313 ASSAY: Xpert MTB/RIF Ultra, COUNT: 296 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1884 | 2023-08-02 | GXIV | 808649 | Unit for XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. B1: passed B2: insufficient cartridge B3: failed B4: insufficient cartridge |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667704133518294174/5667704133518294174_signature_28.png | Nicole Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5667704133518294174/5667704133518294174_signature_31.png
| 2023-08-02 | 2023-08-02 14:07 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 25414 | PBSP | laureanjo@yahoo.com | 28 |  |  | Gx 4 | 804840 | Defective Module A1 210092143 | new module install A1 210100395 & clean fan filter | Install new module A1 210100395 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667693040543371635/5667693040543371635_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5667693040543371635/5667693040543371635_signature_31.png
| 2023-08-02 | 2023-08-02 14:29 | San Jose Rural Health Unit | San Jose de Buenavista, Antique | 6 | Antique | Tb Dots | YES | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 25413 | PBSP | sanjoseantiquegxsite@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2519 ASSAY: Xpert HIV-1 Viral Load, COUNT: 60 | 2023-08-02 | GXIV-4D | 804714 | Module C2 defective, gateway board for replacement due to ultrasonic horn error occuring in module C3. | -Replace defective module C2/789684 with new module SN: 210100165 -Replace gateway board -Configure gateway board serial number using DTerm | #NAME? | All modules are functional. |  | 1 unit 700-6140S Gateway Panel PCB Kit, 4 site; 1 unit 900-0508 GeneXpert Module 6-Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667687636714819131/5667687636714819131_signature_28.png | Hubert Spencer Francis Ang | https://www.jotform.com/uploads/macarewecare/222927978372067/5667687636714819131/5667687636714819131_signature_31.png
| 2023-08-02 | 2023-08-02 14:29 | San Jose Rural Health Unit | San Jose de Buenavista, Antique | 6 | Antique | Tb Dots | YES | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 25412 | PBSP | sanjoseantiquegxsite@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2519 ASSAY: Xpert HIV-1 Viral Load, COUNT: 60 | 2023-08-02 | GXIV-4D | 804714 | Module C2 defective, gateway board for replacement due to ultrasonic horn error occuring in module C3. | -Replace defective module C2/789684 with new module SN: 210100165 -Replace gateway board -Configure gateway board serial number using DTerm | #NAME? | All modules are functional. |  | 1 unit 700-6140S Gateway Panel PCB Kit, 4 site; 1 unit 900-0508 GeneXpert Module 6-Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667687466716373920/5667687466716373920_signature_28.png | Hubert Spencer Francis Ang | https://www.jotform.com/uploads/macarewecare/222927978372067/5667687466716373920/5667687466716373920_signature_31.png
| 2023-08-02 | 2023-08-02 13:21 | Dra. Elvira Lagrosa Health Center | R. Layug, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25410 | PBSP | LBCANICULA@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2315 | 2023-07-31 | GXIV 6C | 805106 | Module B2 and B4 for replacement UPS for troubleshooting | Performed preventive maintenance Replaced modules B2 & B4 Checked UPS cable: ok Checked if UPS do back up: Does not back up Generated report | Replaced Modules B2: 644514 -> 210100506 B4: 604721 -> 210100517 Troubleshoot UPS | Unit is operational. All modules working. UPS does not back up, to be reported for replacement. |  | GX modules 6C | G 6012 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667637649696985141/5667637649696985141_signature_28.png | Lorna Canicula | https://www.jotform.com/uploads/macarewecare/222927978372067/5667637649696985141/5667637649696985141_signature_31.png
| 2023-08-02 | 2023-08-02 12:19 | Tagum City Health Office | Poblacion, Tagum City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25409 | DOH | tagumcholab.mabini@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2765 |  | GX-IV 4D | 838553 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Perform Functionality Test | Perform Preventive Maintenance | All Modules are operational | Module B3 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667635401127959532/5667635401127959532_signature_28.png | Jesusa Bayato | https://www.jotform.com/uploads/macarewecare/222927978372067/5667635401127959532/5667635401127959532_signature_31.png
| 2023-08-02 | 2023-08-02 12:58 | Juan Posadas Health Center | Rodriguez St. Balut | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25408 | PBSP | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1247 |  | GXIV | 805060 | For onsite troubleshooting of Cepheid directions. | Performed troubleshooting. Verified installed applications. Simulated slow response of the workstation. Resets internal components including RAM. Archieved test reports. Generated reports. | Troubleshooting | Unit is operational. Workstation is slow even after troubleshooting and clearing of datas. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667618882215083746/5667618882215083746_signature_28.png | Marlinson Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5667618882215083746/5667618882215083746_signature_31.png
| 2023-08-02 | 2023-08-02 12:36 | Datu Odin Sinsuat District Hospital | Datu Odin Sinsuat, Maguindanao | BARMM | Maguindanao | Tb - Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25407 | DOH | rowenasalazar1234@gmail.com | 32 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 508 | 2023-08-02 | Gx-IV | 838205 | Genexpert room Fluctuations of electrical/power Supply Module 2 multiple temperature error | Replace module 4 SN: 807329 with new module SN: 685820 | Perform Genexpert machine Preventive maintenance | 3 operational modules |  | Genexpert module 6 color SN:685820 | Dr no: 10233 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667613001421684255/5667613001421684255_signature_28.png | Rowen B. Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5667613001421684255/5667613001421684255_signature_31.png
| 2023-08-02 | 2023-08-02 12:15 | Graman Medical And Maternity Hospital Inc. | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25406 | COMMERCIAL | gramanmolecularlab@gmail.com | 19 | ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 238 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 7 |  | GXIV | 110008530 | Unit for maintenance | - checked and cleaned cartridge bay, plunger rod, filter - checked system logs for hardware errors (none) - archived tests from Dec. 2022 to Jul. 2023 - generated IQR, System Log, Assay Statistic reports - performed functionality test | Performed maintenance | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667592543511941986/5667592543511941986_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5667592543511941986/5667592543511941986_signature_31.png
| 2023-08-02 | 2023-08-02 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  | mark10venzon@gmail.com |  |  | 25405 | PBSP | honeyeichel@gmail.com | 38 | ASSAY: Xpert HBV VL, COUNT: 46 ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2443 ASSAY: Xpert HCV VL, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 191 |  | GXIV-D | 805124 | Monitor for replacement Defective monitor:307NDJXME852 | Monitor replacement Preventive maintenance | Performed preventive maintenance Performed monitor replacement Archived test results Gathered reports Performed functionality test | Unit is operational, UPS not backing up New Monitor SN:V5GAK167 Lenovo ThinkVision S22e-20 | UPS for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667581093713269620/5667581093713269620_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5667581093713269620/5667581093713269620_signature_31.png
| 2023-08-02 | 2023-08-02 11:14 | Sulat Rural Health Unit | Sulat Eastern Samar | 8 | Eastern Samar | TB - DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25404 | PBSP | katherinelupos@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 704 | 2023-08-02 | GXIV-2D | 835100 | #NAME? | #NAME? | - Replaced Module B1/787996 into 210100353 - Xpertcheck Calibration - Performed Preventive maintenance -Performed Functionality test - Generate reports | Unit is Operational |  | 1 - Genexpert Module 6 color SN/210100352 1 - Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667558866513380418/5667558866513380418_signature_28.png | Katherine Ann L. Alday | https://www.jotform.com/uploads/macarewecare/222927978372067/5667558866513380418/5667558866513380418_signature_31.png
| 2023-08-02 | 2023-08-02 10:36 | Abucay Rural Health Unit | Devena Ave., Sailan, Abucay, Bataan | 3 | Bataan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25403 | DOH | parreraledelina22176@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 39 ASSAY: Xpert MTB/RIF Ultra, COUNT: 441 | 2023-07-17 | GXIV-D | 831878 | A1 and A2 for module replacement | Module replacement | Performed preventive maintenance Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is operational |  | 2 genexpert modules | 10138 and 10140 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667530942421310797/5667530942421310797_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5667530942421310797/5667530942421310797_signature_31.png
| 2023-08-02 | 2023-08-02 09:56 | Chinese General Hospital | Blumentritt, Sta. Cruz | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25402 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 21 ASSAY: Xpert CT/NG, COUNT: 38 ASSAY: Xpert HBV VL, COUNT: 86 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1504 ASSAY: Xpert HCV VL, COUNT: 20 ASSAY: Xpert HIV-1 Viral Load, COUNT: 21 |  | GXIV | 835604 | Errors on module B1. Unit for Preventive maintenance. | Performed preventive maintenance. Generated reports. Generated gxx files. | Preventive maintenance and troubleshooting. | Unit is operational. For monitoring of module B1. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667513062213594422/5667513062213594422_signature_28.png | Maria Odessa Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/5667513062213594422/5667513062213594422_signature_31.png
| 2023-08-02 | 2023-08-02 08:48 | Bagong Pag-Asa Health Center | 35 Road 9 cor. Road 10, Brgy. Bagong Pag-Qsa | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25401 | DOH | erlindalorilla026@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 174 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1554 | 2023-07-27 | GXIV 6C | 831597 | Monitor Display defective | Replaced Monitor Performed PM Generated report | Replaced Monitor CNK8322441 -> CNK83224Q2 | Unit is operational. Monitor working. |  | GX Monitor Display HP | G 5971 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5667484847119056794/5667484847119056794_signature_28.png | Erlinda Lorilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5667484847119056794/5667484847119056794_signature_31.png
| 2023-08-01 | 2023-08-01 16:09 | Pulilan Rural Health Unit | Pulilan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25400 | DOH | celsptcruz19@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 627 | 2023-07-24 | GXIV | 831915 | Module A2 for replacement | #NAME? | Replaced A2 | Unit is operational. |  | GeneXpert Module SN: 697523 | 10265 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666867103511187385/5666867103511187385_signature_28.png | Cecile Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5666867103511187385/5666867103511187385_signature_31.png
| 2023-08-01 | 2023-08-01 15:47 | Mangatarem Rural Health Unit | Mangatarem, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25399 | DOH | rhumangatarem@yahoo.com | 288 | ASSAY: Xpert MTB/RIF, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 680 | 2023-07-19 | GX4 | 838284 | Due for Module Replacement | - Cleaned Cartridge Bays, Plunger Rods and I-CORE - Replaced Defective module B3 685919 >>> 210006666 - Generated Reports | Performed Module Replacement | Unit is operational |  | Module 210006666 | DR 10205 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666855660217507354/5666855660217507354_signature_28.png | Ken Lee Cesar D. Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/5666855660217507354/5666855660217507354_signature_31.png
| 2023-08-01 | 2023-08-01 15:05 | Dimasalang Health Center | Isagani St. Sta Cruz | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25398 | PBSP | joneevamarie@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1411 | 2023-07-31 | GXIV | 805111 | Unit for PM. For module replacement of module B3. For Xpertcheck of module B1 and B4. | Performed PM. Performed module replacement on B3. Performed Xpertcheck of B1 and B4. Generated reports. | PM, module replacement and Xpertcheck. | Unit is operational. Module B4, service required. | Used 2 spare Xpertcheck kit | Gx module B3 (210100519) in replacement to defective B3 (641256). | G6014 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666844042217913603/5666844042217913603_signature_28.png | Eva Marie C. Jone | https://www.jotform.com/uploads/macarewecare/222927978372067/5666844042217913603/5666844042217913603_signature_31.png
| 2023-08-01 | 2023-08-01 14:11 | San Antonio District Hospital | San Mariano, San Antonio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25397 | PBSP | garciaadonisanthony@gmail.com | 29 |  | 2023-07-25 | Genexpert IV | 812832 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Installed Laptop Performed Functionality Test | Unit is Operational | Xpertcheck = All Modules Passed | Genexpert Laptop - CPH1019F03275 | G 5917 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666803142815294611/5666803142815294611_signature_28.png | Adonis Anthony Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5666803142815294611/5666803142815294611_signature_31.png
| 2023-08-01 | 2023-08-01 14:05 | Tunasan Health Center | E. Rodriguez St., Tunasan, Muntinlupa | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25396 | PBSP | gigietan0911@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1660 ASSAY: Xpert HIV-1 Viral Load, COUNT: 195 ASSAY: Xpert HPV, COUNT: 100 | 2023-07-24 | GXIV 6C | 807613 | Module B3 for replacement | Performed PM Replaced Module B3 | Replaced module B3 678800 -> 210100380 | Unit is operational. All modules working |  | GX module 6C | G 5895 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666796824316642729/5666796824316642729_signature_28.png | geraldine trazo-tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5666796824316642729/5666796824316642729_signature_31.png
| 2023-08-01 | 2023-08-01 13:43 | Meycauayan City Health Center VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25395 | DOH | ronamarieiris_25@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1419 ASSAY: Xpert HIV-1 Viral Load, COUNT: 85 | 2023-07-24 | GXIV | 831904 | #NAME? | #NAME? | Replaced A1, A4, Monitor | Unit is operational. |  | GeneXpert Module SN: 601011, 909140 GeneXpert Monitor SN: MMA1PSS00221507F3E3W01 | 10261, 10262, 10291 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666781743512566439/5666781743512566439_signature_28.png | Rona Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5666781743512566439/5666781743512566439_signature_31.png
| 2023-08-01 | 2023-08-01 13:16 | Reproductive Health and Wellness Center | Jacinto St. Poblacion Davao City | 11 | Davao del Sur | RHIVDA Section | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25394 | MACARE-HBDC | rhwclab@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 67 ASSAY: Xpert HIV-1 Viral Load, COUNT: 618 |  | GX-IV 4d | 839236 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perfrom Preventive Maintenance | Only Two Modulea are Operational | Module B1 and B2 disabled due to H/W Failure, Data sent to Cepheid for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666780469024624751/5666780469024624751_signature_28.png | Macaturac, val | https://www.jotform.com/uploads/macarewecare/222927978372067/5666780469024624751/5666780469024624751_signature_31.png
| 2023-08-01 | 2023-08-01 13:20 | Reproductive Health and Wellness Center | Jacinto Street, Poblacion Davao City | 11 | Davao del Sur | RHIVDA Section | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25393 | DOH | rhwclab@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 135 ASSAY: Xpert HIV-1 Viral Load, COUNT: 871 |  | GX-IV 4D | 838262 | Due for Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666779909021051061/5666779909021051061_signature_28.png | Macaturac, val | https://www.jotform.com/uploads/macarewecare/222927978372067/5666779909021051061/5666779909021051061_signature_31.png
| 2023-08-01 | 2023-07-25 11:40 | Bayambang District Hospital | Bayambang, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25392 | PBSP | bevzneri@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1201 | 2023-07-25 | GX4 | 812830 | For Module Replacement | -Cleaned Cartridge Bays, Plunger Rods and I-CORE -Replaced Defective module 667602>>>210100620 -Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210100620 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666775870214783970/5666775870214783970_signature_28.png | Beverly A. Neri | https://www.jotform.com/uploads/macarewecare/222927978372067/5666775870214783970/5666775870214783970_signature_31.png
| 2023-08-01 | 2023-07-25 11:40 | Bayambang District Hospital | Bayambang, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25391 | PBSP | bevzneri@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1201 | 2023-07-25 | GX4 | 812830 | For Module Replacement | -Cleaned Cartridge Bays, Plunger Rods and I-CORE -Replaced Defective module 667602>>>210100620 -Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210100620 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666775470216575529/5666775470216575529_signature_28.png | Beverly A. Neri | https://www.jotform.com/uploads/macarewecare/222927978372067/5666775470216575529/5666775470216575529_signature_31.png
| 2023-08-01 | 2023-07-26 06:25 | Dagupan City Health Office | Dagupan City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25390 | DOH | thisisformyfamily@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1426 | 2023-07-19 | GX4 | 838286 | For Module Replacement | -Cleaned Cartridge bays, Plunger Rods,and I-CORE -Replace defective module 792263 >>> 645833 -Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 645833 | DR 10197 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666774160216824322/5666774160216824322_signature_28.png | Rudolph joseph guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5666774160216824322/5666774160216824322_signature_31.png
| 2023-08-01 | 2023-07-25 10:45 | Rosales Rural Health Unit | Resales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25389 | DOH | allan.vinluan2013@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 201 ASSAY: Xpert MTB/RIF Ultra, COUNT: 551 | 2023-07-19 | GX 4 | 838472 | Module B3 for Module Replacement | -Cleaned Cartridge bays, Plunger Rods and I-CORE -Performed Module Replacement on Module B3 SN 862036>>> 620197 -Generated Reports | Performed Module Replacement | Unit is Operational |  | Module 620197 | DR 10198 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666760380216924163/5666760380216924163_signature_28.png | Allan b. Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5666760380216924163/5666760380216924163_signature_31.png
| 2023-08-01 | 2023-07-25 10:45 | Rosales Rural Health Unit | Resales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25388 | DOH | allan.vinluan2013@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 201 ASSAY: Xpert MTB/RIF Ultra, COUNT: 551 | 2023-07-19 | GX 4 | 838472 | Module B3 for Module Replacement | -Cleaned Cartridge bays, Plunger Rods and I-CORE -Performed Module Replacement on Module B3 SN 862036>>> 620197 -Generated Reports | Performed Module Replacement | Unit is Operational |  | Module 620197 | DR 10198 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666760180219325645/5666760180219325645_signature_28.png | Allan b. Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5666760180219325645/5666760180219325645_signature_31.png
| 2023-08-01 | 2023-07-25 10:45 | Rosales Rural Health Unit | Resales, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25387 | DOH | allan.vinluan2013@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 201 ASSAY: Xpert MTB/RIF Ultra, COUNT: 551 | 2023-07-19 | GX 4 | 838472 | Module B3 for Module Replacement | -Cleaned Cartridge bays, Plunger Rods and I-CORE -Performed Module Replacement on Module B3 SN 862036>>> 620197 -Generated Reports | Performed Module Replacement | Unit is Operational |  | Module 620197 | DR 10198 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666759750214893373/5666759750214893373_signature_28.png | Allan b. Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5666759750214893373/5666759750214893373_signature_31.png
| 2023-08-01 | 2023-07-25 11:40 | Bayambang District Hospital | Bayambang, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25386 | PBSP | bevzneri@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1201 | 2023-07-25 | GX4 | 812830 | For Module Replacement | -Cleaned Cartridge Bays, Plunger Rods and I-CORE -Replaced Defective module 667602>>>210100620 -Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Module 210100620 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666756610212220609/5666756610212220609_signature_28.png | Beverly A. Neri | https://www.jotform.com/uploads/macarewecare/222927978372067/5666756610212220609/5666756610212220609_signature_31.png
| 2023-08-01 | 2023-07-25 14:36 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25385 | DOH | gbmm@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 | 2023-07-19 | GX4 | 838437 | For Module Replacement | -Cleaned Cartridge Bays, Plunger Rods and I-CORE -Replaced Defective modules SN 861278>>>>210097620 680419>>>>210059512 -Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Modules 210097620 and 210059512 | DR 10196 DR 10194 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666754910216482584/5666754910216482584_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5666754910216482584/5666754910216482584_signature_31.png
| 2023-08-01 | 2023-07-25 14:36 | San Carlos City Health Office | San Carlos City, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25384 | DOH | gbmm@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 | 2023-07-19 | GX4 | 838437 | For Module Replacement | -Cleaned Cartridge Bays, Plunger Rods and I-CORE -Replaced Defective modules SN 861278>>>>210097620 680419>>>>210059512 -Generated Reports | Performed Module Replacement | Unit is operational |  | GeneXpert Modules 210097620 and 210059512 | DR 10196 DR 10194 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666754510211726113/5666754510211726113_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5666754510211726113/5666754510211726113_signature_31.png
| 2023-08-01 | 2023-07-26 11:01 | Basista Rural Health Unit | Basista, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25383 | PBSP | coraflores14@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 71 ASSAY: Xpert MTB/RIF Ultra, COUNT: 152 |  | GX4 | 838860 | Due for Calibration | #NAME? | Performed XpertCheck | Unit is operational | All modules passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666753310217603221/5666753310217603221_signature_28.png | Cora S. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5666753310217603221/5666753310217603221_signature_31.png
| 2023-08-01 | 2023-07-26 11:01 | Basista Rural Health Unit | Basista, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25382 | PBSP | coraflores14@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 71 ASSAY: Xpert MTB/RIF Ultra, COUNT: 152 |  | GX4 | 838860 | Due for Calibration | #NAME? | Performed XpertCheck | Unit is operational | All modules passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666752840219978514/5666752840219978514_signature_28.png | Cora S. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5666752840219978514/5666752840219978514_signature_31.png
| 2023-08-01 | 2023-07-28 08:40 | Calasiao Rural Health Unit | Calasiao, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25381 | PBSP | calasiaoclinicallab@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 72 ASSAY: Xpert MTB/RIF Ultra, COUNT: 513 |  | GX4 | 834664 | Due for Calibration | #NAME? | Performed XpertCheck | Unit is operational | Module B4 failed for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666751620216295442/5666751620216295442_signature_28.png | John John Bauzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5666751620216295442/5666751620216295442_signature_31.png
| 2023-08-01 | 2023-07-28 08:40 | Calasiao Rural Health Unit | Calasiao, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25380 | PBSP | calasiaoclinicallab@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 72 ASSAY: Xpert MTB/RIF Ultra, COUNT: 513 |  | GX4 | 834664 | Due for Calibration | #NAME? | Performed XpertCheck | Unit is operational | Module B4 failed for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666751160217810801/5666751160217810801_signature_28.png | John John Bauzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5666751160217810801/5666751160217810801_signature_31.png
| 2023-08-01 | 2023-08-01 12:49 | Legazpi City Health Office | Legazpi City Hall Compound, Legazpi | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25379 | WHO | janzchel_neneh10@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1095q |  | GX4 | 846354 | B1, B2 for module replacement | Replaced module: 890880>>210101409, 907517>>210101426 Cleaned cartridge bays, plunger rod, icore,fan Generated reports | B1, B2 module replaced. Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666749518681994609/5666749518681994609_signature_28.png | Janzhel marie chan | https://www.jotform.com/uploads/macarewecare/222927978372067/5666749518681994609/5666749518681994609_signature_31.png
| 2023-08-01 | 2023-08-01 11:19 | Cotabato City Health Office | Cotabato City, Maguindanao | 12 | Maguindanao | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 25378 | DOH | onehope834@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF, COUNT: 51 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1580 |  | Gx-IV | 838528 | Module 2 and 4 for Service Required | Perform Xpertcheck at module 2 and 4 for troubleshooting | Perform Genexpert machine Preventive maintenance Installed new printer for Genexpert machine | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666716119113228568/5666716119113228568_signature_28.png | Emelyn Cadiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5666716119113228568/5666716119113228568_signature_31.png
| 2023-08-01 | 2023-08-01 11:40 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25377 | DOH | dorothygundaya@gmail.com | 26.35 | ASSAY: Xpert MTB/RIF, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1569 ASSAY: Xpert HIV-1 Viral Load, COUNT: 47 | 2023-07-21 | GX-4 | 838403 | None | None | Replacing Module B1 with SN: 860976 to SN: 210041238 Replacing Module B3 with SN: 790843 to SN: 905063 Replacing Module B4 with SN: 860982 to SN: 210041380 Preventive Maintenance | All modules are working |  | GeneXpert Module 6 Color | 900-0491R |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666713409214168279/5666713409214168279_signature_28.png | Dorothy Gundaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5666713409214168279/5666713409214168279_signature_31.png
| 2023-08-01 | 2023-08-01 11:24 | Giporlos Rural Health Unit | Giporlos Eastern Samar | 8 | Eastern Samar | Tb- dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25376 | PBSP | ningrediang@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 469 | 2023-08-01 | GXIV | 835060 | Unit for Calibration | #NAME? | #NAME? | Unit is Operational - Module B2 for service required |  | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666698533816462185/5666698533816462185_signature_28.png | Luningning Abuda | https://www.jotform.com/uploads/macarewecare/222927978372067/5666698533816462185/5666698533816462185_signature_31.png
| 2023-08-01 | 2023-08-01 10:50 | Sara RHU | Sara, Iloilo | 6 | Iloilo | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25375 | PBSP | candyparrenas@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 209 |  | GX4 | 838200 | Module B4 with error code 2014 and it failed in the Xpert check. | Run calibration for module B4 only and conduct PM | Run Xpert check for module B4 only | Module B1, B2 and B3 are in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666681686873528988/5666681686873528988_signature_28.png | Ma. Candelaria Parreñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5666681686873528988/5666681686873528988_signature_31.png
| 2023-08-01 | 2023-08-01 10:28 | Bicol South Luzon Sub National Reference Laboratory | Rizal St., Legazpi | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25374 | PBSP | rbalzaga.pbsp@gmail.com | 28 |  |  | GX4 | 808636 | For Preventive maintenence | Cleaned cartridge bay, plunger rod, icore Generated reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666672150502298959/5666672150502298959_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5666672150502298959/5666672150502298959_signature_31.png
| 2023-08-01 | 2023-08-01 10:20 | Bicol South Luzon Sub National Reference Laboratory | Rizal St., Legazpi | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25373 | PBSP | rbalzaga.pbsp@gmail.com | 26 |  |  | GX4 | 807579 | B1, B2 for Module Replacement | Replaced Modules: B1/784716>>210100698, B2/653834>>210100424 Cleaned cartridge bays, plunger rod, icore Generated reports | Replaced Modules B1, B2 Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666670960507144637/5666670960507144637_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5666670960507144637/5666670960507144637_signature_31.png
| 2023-08-01 | 2023-08-01 10:35 | Bicol South Luzon Sub National Reference Laboratory | Rizal St., Legazpi | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25372 | PBSP | rbalzaga.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 158 |  | GX4 10C | 110016954 | For Preventive maintenance | Cleaned cartridge bays, plunger rod, icore | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666668998686532494/5666668998686532494_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5666668998686532494/5666668998686532494_signature_31.png
| 2023-08-01 | 2023-08-01 09:06 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25371 | PBSP | crmcmolecularlab@gmail.com | 22.6 | ASSAY: Xpert SARS-CoV-2, COUNT: 4866 | 2023-08-01 | Gx-IV | 834442 | N/A | Perform Genexpert machine Xpertcheck/Calibration | Perform Genexpert machine Preventive maintenance | 4 operational modules | All modules passed the Xpertcheck/Calibration | Xpertcheck CE-5 | Lot #: 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666656699117052215/5666656699117052215_signature_28.png | CRMC Molecular Lab | https://www.jotform.com/uploads/macarewecare/222927978372067/5666656699117052215/5666656699117052215_signature_31.png
| 2023-08-01 | 2023-08-01 09:01 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25370 | PBSP | crmcmolecularlab@gmail.com | 23.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 4866 | 2023-08-01 | Gx-IV | 834441 | Module 1, 2, and 3 for Service Required | Perform Genexpert machine Xpertcheck/Calibration | Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Xpertcheck CE-5 | Lot #: 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666655579118079267/5666655579118079267_signature_28.png | CRMC Molecular Lab | https://www.jotform.com/uploads/macarewecare/222927978372067/5666655579118079267/5666655579118079267_signature_31.png
| 2023-08-01 | 2023-08-01 10:01 | Sta. Ignacia Rural Health Unit | Sta. Ignacia Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25369 | DOH | jdi0973@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 421 |  | GXIV-D | 831909 | A2 for troubleshooting | Troubleshooting | Performed preventive maintenance Performed cepheid troubkeshooting procedures Archived test results Gathered reports Performed functionality test | Unit is operational | A2 under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666647168374214941/5666647168374214941_signature_28.png | Jess D. Ibarra Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5666647168374214941/5666647168374214941_signature_31.png
| 2023-07-31 | 2023-07-31 16:48 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25368 | PBSP | crmctblab@gmail.com | 22.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 ASSAY: Xpert HBV VL, COUNT: 40 | 2023-07-31 | Gx-IV | 807463 | Module 3 for Service Required | Perform Genexpert machine Xpertcheck/Calibration | Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Xpertcheck CE-5 | Lot #: 1000725070 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666046869114854806/5666046869114854806_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5666046869114854806/5666046869114854806_signature_31.png
| 2023-07-31 | 2023-07-31 16:52 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25367 | PBSP | crmctblab@gmail.com | 22.3 | ASSAY: Xpert HBV VL, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 40 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 | 2023-07-31 | Gx-IV | 807607 | Module 3 for Service Required | Perform Genexpert machine Xpertcheck/Calibration | Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Xpertcheck CE-5 | Lot #: 1000725070 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666045659112374428/5666045659112374428_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5666045659112374428/5666045659112374428_signature_31.png
| 2023-07-31 | 2023-07-31 16:41 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Tb-Lab | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 25366 | PBSP | crmctblab@gmail.com | 24.3 | ASSAY: Xpert HBV VL, COUNT: 161 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2913 ASSAY: Xpert HIV-1 Viral Load, COUNT: 246 |  | Gx-XXVI | 839307 | Module 2 multiple temperature error | Report module B2 problem to Jotform reporting | Perform Genexpert machine Preventive maintenance | 15 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5666044619111230246/5666044619111230246_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5666044619111230246/5666044619111230246_signature_31.png
| 2023-07-31 | 2023-07-31 13:53 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25365 | PBSP | crmcmolecularlab@gmail.com | 26.2 | ASSAY: Xpert SARS-CoV-2, COUNT: 4866 |  | Gx-IV | 834441 | Module D4 multiple temperature error | Report module 4 problem to Jotform reporting | Perform Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665990909119958536/5665990909119958536_signature_28.png | crmcmolecularlab | https://www.jotform.com/uploads/macarewecare/222927978372067/5665990909119958536/5665990909119958536_signature_31.png
| 2023-07-31 | 2023-07-28 12:20 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25362 | PBSP | molecularlab@rmc.doh.gov.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3506 ASSAY: Xpert SARS-CoV-2, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 58 |  | GX-IV-D | 814374 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. Module B1/607982 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665982103611050153/5665982103611050153_signature_28.png | JUSTINE LORENZ V. FALZADO, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5665982103611050153/5665982103611050153_signature_31.png
| 2023-07-31 | 2023-07-28 15:48 | Pililla Rural Health Unit | M.A Roxas St., Pililla | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25364 | DOH | tbdotspililla@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 602 | 2023-07-25 | GX-IV-D | 821293 | Module B3 for replacement. Module B2/687232 has error 4012 | Replaced defective module. Run XpertCheck on module B2 | Performed maintenance Replaced defective module Run XpertCheck on module B2 Generated reports | Unit is operational. Module B2/687232 service required. |  | Old: B3/730900 New: B3/805270 | DR# 10287 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665982103611658472/5665982103611658472_signature_28.png | Karren R. Olaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5665982103611658472/5665982103611658472_signature_31.png
| 2023-07-31 | 2023-07-28 12:30 | Rizal Medical Center | Pasig Blvd. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 25363 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 26 | ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 9 ASSAY: Xpert SARS-CoV-2, COUNT: 14128 ASSAY: Others, COUNT: 9 |  | GX-IV-D | 838627 | Modules A1, B3, B4, C2, D2 and D4 has errors. | Checked modules. | Performed maintenance Cleaned cartridge bay, plunger rod and icore Checked modules Generated reports. | Unit is operational. Modules A1/860941 and B4/860923 not operational but reported already. Module B3/702534, C2/668748, D2/620837 and D4/661946 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665982103612511503/5665982103612511503_signature_28.png | JUSTINE LORENZ V. FALZADO, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5665982103612511503/5665982103612511503_signature_31.png
| 2023-07-31 | 2023-07-27 13:36 | Binangonan Rural Health Unit III | Osmeña St, Binangonan | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25361 | WHO | yanamanawari@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 551 | 2023-07-27 | GX-IV-D | 846045 | Module B4/904767 for replacement | Replaced module B4 | Performed maintenance Replaced defective module Generated reports | All modules operational. |  | Old: B4/904767 New: B4/210100341 | G 5948 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665982103615495818/5665982103615495818_signature_28.png | Dhayanara H. Zapanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5665982103615495818/5665982103615495818_signature_31.png
| 2023-07-31 | 2023-07-26 11:09 | West Rembo Health Center | West Rembo, Makati | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 25358 | DOH | mcabalosjr.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1215 | 2023-07-24 | GX-IV-D | 831712 | Module A4/757448 for replacement, Module A2/757446 error 1001 | Replaced defective module Run XpertCheck on module A2 | Performed maintenance Replaced module A4 Run XpertCheck on module A2 Generated reports | Unit is operational. Module A2 passed Xpertcheck but under observation. |  | Old: A4/757448 New: A4/210016336 | DR# 10252 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981973612836289/5665981973612836289_signature_28.png | Jun Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981973612836289/5665981973612836289_signature_31.png
| 2023-07-31 | 2023-07-27 11:39 | Binangonan Rural Health Unit I | Manila East Rd., Binangonan | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25359 | PBSP | jen_cerezola@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1497 |  | GX-IV-D | 838199 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. | Module B2/807335 has error 5007 under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981973616229418/5665981973616229418_signature_28.png | Jennifer Cerezola | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981973616229418/5665981973616229418_signature_31.png
| 2023-07-31 | 2023-07-26 11:03 | West Rembo Health Center | West Rembo, Makati | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25357 | PBSP | mcabalosjr.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1215 |  | GX-IV-D | 838875 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981973616438431/5665981973616438431_signature_28.png | Jun Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981973616438431/5665981973616438431_signature_31.png
| 2023-07-31 | 2023-07-27 11:42 | Binangonan Rural Health Unit I | Manila East Rd., Binangonan | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25360 | PBSP | jen_cerezola@yahoo.com | 25 |  |  | GX-IV-D | 807605 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981973618274190/5665981973618274190_signature_28.png | Jennifer Cerezola | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981973618274190/5665981973618274190_signature_31.png
| 2023-07-31 | 2023-07-31 13:55 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25356 | PBSP | crmcmolecularlab@gmail.com | 25.8 | ASSAY: Xpert SARS-CoV-2, COUNT: 4866 |  | Gx-IV | 834442 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981329116105356/5665981329116105356_signature_28.png | crmcmolecularlab | https://www.jotform.com/uploads/macarewecare/222927978372067/5665981329116105356/5665981329116105356_signature_31.png
| 2023-07-31 | 2023-07-31 13:46 | Cotabato Regional and Medical Center | Cotabato City, Maguindanao | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25355 | DOH | crmcmolecularlab@gmail.com | 21.3 | ASSAY: Xpert SARS-CoV-2, COUNT: 4866 | 2023-07-31 | Gx-IV | 837662 | Module 1 and 3 H/W failed | Replace module 1 SN:806117 with new module SN: 642942 Replace module 3 SN:674965 with new module SN: 643459 | Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 color | DR no: 10231 and 10232 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665979099118358463/5665979099118358463_signature_28.png | Crmc molecular lab | https://www.jotform.com/uploads/macarewecare/222927978372067/5665979099118358463/5665979099118358463_signature_31.png
| 2023-07-31 | 2023-07-27 11:34 | Madridejos RHU | Madridejos, Bantayan Island | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE Site visit | CEPHEID | Raymond Laurito |  |  |  |  | 25354 | PBSP | cabasevilma@gmail.com | 28 |  |  | Gx 4 | 835057 | Site visit clean fan filter | Site visit clean fan filter | Site visit clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665971583918677868/5665971583918677868_signature_28.png | Vilma cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5665971583918677868/5665971583918677868_signature_31.png
| 2023-07-31 | 2023-07-31 14:46 | NEGROS ORIENTAL PROVINCIAL HOSPITAL | Dumaguete City | 7 | Negros Oriental | TB Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25353 | DOH | rezelmarielagos@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 388 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2390 ASSAY: Xpert HIV-1 Viral Load, COUNT: 147 |  | GX-4 | 838525 | Defective module B1/806592. | - Preventive Maintenance. - Run XPC on module B3; B3 passed XPC Run. | Replaced module B1 with new module SN 896104. | modules 1, 2, 4 operational. module 3 for observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665967754914218073/5665967754914218073_signature_28.png | Rezel Marie Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5665967754914218073/5665967754914218073_signature_31.png
| 2023-07-31 | 2023-07-31 13:41 | Libacao RHU | Libacao, Aklan | 6 | Aklan | Laboratory Dept. | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 25352 | DOH | haladeeb18@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 106 |  | GX4 | 838540 | Module B1, B2 and B4 failed in calibration | Conduct PM and saves files. | Replaced the old module for B1 SN 862287 to new module SN 616665 Replaced the old module for B2 SN 862288 to new module 616415. Replaced the old module for B4 SN 862322 to new module SN 611883. | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665920169212298732/5665920169212298732_signature_28.png | Ma. Hala Deed Estanislao | https://www.jotform.com/uploads/macarewecare/222927978372067/5665920169212298732/5665920169212298732_signature_31.png
| 2023-07-31 | 2023-07-28 13:32 | Tinambac RHU | Zone-2, Binalay, Tinambac | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25351 | PBSP | maelagarde27@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1052 |  | GX4 | 808285 | For XpertCheck | Performed XpertCheck Cleaned Plunger rod, fan filter, cartridge bays, Icores | Performed XpertCheck, Preventive Maintenance | Unit is operational B1 - B2 - B3 - B4 - |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665896291928230669/5665896291928230669_signature_28.png | Mae Aguinillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5665896291928230669/5665896291928230669_signature_31.png
| 2023-07-31 | 2023-07-25 14:11 | National Tuberculosis Reference Laboratory | RITM Compound, Filinvest | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25350 | PBSP | mspogosa.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1961 | 2023-07-24 | GXIV-D | 707512 | Modules 2 and 3 with Temperature Problem | Module Replacement | Performed Maintenance Replaced Module 2 and 3 Generate Reports | All Modules are Operational |  | GX Module | G 5894 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665779003619750176/5665779003619750176_signature_28.png | Maria Guileane Pogosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5665779003619750176/5665779003619750176_signature_31.png
| 2023-07-31 | 2023-07-25 15:03 | Asian Hospital and Medical Center | 2205 Civic Drive, Filinvest | NCR | NCR 4th - Muntinlupa | Molecular | YES | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25349 | COMMERCIAL | agalang@asianhospital.com | 24 |  | 2023-07-24 | GXIV-D-10 C | 110014996 | Unit is due for Installation | Installation | Installed GeneXpert System Performed Functionality Test Generate Reports | System is Operational |  |  | SI 95990 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943611876570/5665757943611876570_signature_28.png | Alyssa Galang | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943611876570/5665757943611876570_signature_31.png
| 2023-07-31 | 2023-07-21 12:32 | Imus City Health Office II | Plaridel Subd, Imus | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25347 | DOH | amielynnmangulabnan1217@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1151 | 2023-07-19 | GXIV-D | 821284 | Module 1 with Temperature Problem UPs is Defective | Module Replacement UPS Replacement | Performed Maintenance Replaced UPS Replaced Module 1 Generate Reports | 3 Modules are Operational |  | GX Module 6-Colors | DR 10199 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943614267665/5665757943614267665_signature_28.png | Amielynn Mangulabnan | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943614267665/5665757943614267665_signature_31.png
| 2023-07-31 | 2023-07-19 12:08 | Noveleta Health Center | Municipal Compound, Noveleta | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25346 | PBSP | jacque_estinoso23@yahoo.com.ph | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 424 |  | GXIV-D-2 | 835186 | Unit is Due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943617292485/5665757943617292485_signature_28.png | Jacqueline Q Estinoso | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943617292485/5665757943617292485_signature_31.png
| 2023-07-31 | 2023-07-19 10:34 | Caridad Health Center | M. Dela Cruz St, Cavite City | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25345 | DOH | seyerlanie85@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1444 | 2023-07-17 | GXIV-D | 821227 | Module 2 Failed in XpertCheck Unit is Due for XpertCheck | Module Replacement XpertCheck | Performed Maintenance Replaced Module 2 Performed XpertCheck Generate Reports | All Modules are Operational | 730030->786729 | GX Module 6-Colors | DR 10150 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943617462949/5665757943617462949_signature_28.png | Melanie Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943617462949/5665757943617462949_signature_31.png
| 2023-07-31 | 2023-07-18 14:18 | Bagong Brgy Health Center | W. Zamora, Pandacan | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25348 | DOH | thea.peneyra@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1071 | 2023-07-17 | GXIV-D | 831694 | 2 Modules Failed in XpertCheck | Module Replacement | Performed Maintenance Replaced Modules 2 and 3 Generate Reports | 3 Modules are Operational, 1 Defective |  | GX Module | DR 10148 & 10149 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943618641720/5665757943618641720_signature_28.png | Maria Anthea Peneyra | https://www.jotform.com/uploads/macarewecare/222927978372067/5665757943618641720/5665757943618641720_signature_31.png
| 2023-07-30 | 2023-07-28 11:46 | Chinese General Hospital | Blumentriitt st, Manila | NCR | NCR 1st - Manila | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 25344 | COMMERCIAL | cchlaboratory2016@yahoo.com | 28 |  |  | MCS + | 08M463 | As per end user they encounter valve error | Checked unit Cheked pneumatic pump Checked valve and performed valve diagnostics Checked program parameters Performed functional test and simulation running | Troubleshooting | During testing and running no error encounter |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665308570519367188/5665308570519367188_signature_28.png | Jacqueline Zarzoso | https://www.jotform.com/uploads/macarewecare/222927978372067/5665308570519367188/5665308570519367188_signature_31.png
| 2023-07-30 | 2023-07-27 15:53 | National kidney and transplant institute | East avenue QC | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25343 | MACARE-HBDC | molecularlabnkti@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 677 ASSAY: Others, COUNT: 10 ASSAY: Xpert HBV VL, COUNT: 407 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1236 ASSAY: Xpert MTB/RIF Ultra, COUNT: 705 ASSAY: Xpert SARS-CoV-2, COUNT: 16910 ASSAY: Xpert HCV VL, COUNT: 34 ASSAY: Others, COUNT: 9 | 2023-07-24 | GXXVI | 110005587 | For module replacement and maintenance D2 633953 - 692814 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | C3 and D3 Ttemp error,B1 loss communication 13 module operational |  | 900-0491rs | 10273 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665303680514413949/5665303680514413949_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5665303680514413949/5665303680514413949_signature_31.png
| 2023-07-30 | 2023-07-27 11:16 | Payatas B Health Center | Bulacan st, Payatas QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25342 | DOH | racquelmrocero@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1477 | 2023-07-24 | GX4 | 831899 | For module replacement A2 758252 - 210026721 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | A3 module reported already, A1,A2,A4 operational |  | 900-0491r | 10255 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665302890517011004/5665302890517011004_signature_28.png | Racquel M Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5665302890517011004/5665302890517011004_signature_31.png
| 2023-07-30 | 2023-07-28 14:32 | National children's hospital | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25341 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3034 ASSAY: Xpert SARS-CoV-2, COUNT: 7226 ASSAY: Others, COUNT: 18 | 2023-07-26 | GXXVI | 110014171 | For module replacement and maintenance B3 210069245 - 898183 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit,performed module replacement | Module replacement and maintenance | Operational |  | 900-065 rs | 10296 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5665301420515475657/5665301420515475657_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5665301420515475657/5665301420515475657_signature_31.png
| 2023-07-28 | 2023-07-28 14:49 | Moalboal Rural Health Unit | Poblacion Moalboal Cebu | 7 | Cebu | Tb Laboratory | YES | TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25340 | PBSP | roavaul@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 514 |  | Gx4 | 838892 | Monitor LCD problem | Recommend for replacement of monitor Ask end user to report in jotform for proper channel of request for replacement | Troubleshoot monitor. Transfer to othe cpu and the monitor problem still exist Isolate problem | All modules are operational only the LCD monitor are defective |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663589013916737328/5663589013916737328_signature_28.png | Ava Gako | https://www.jotform.com/uploads/macarewecare/222927978372067/5663589013916737328/5663589013916737328_signature_31.png
| 2023-07-28 | 2023-07-28 15:00 | Pulilan Rural Health Unit I | Pulilan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25339 | DOH | celsptcruz19@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 611 | 2023-07-26 | GXIV | 831915 | Monitor for replacement | #NAME? | Replaced monitor | Unit is operational. |  | GeneXpert Monitor SN: CNK83224Q9 | 10293 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663385978313143649/5663385978313143649_signature_28.png | Cecile Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5663385978313143649/5663385978313143649_signature_31.png
| 2023-07-28 | 2023-07-28 14:18 | Mini Forest Health Center - District C | Quezon Boulevard, Brgy. 23-C Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25338 | DOH | miniforestlaboratory@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF, COUNT: 196 ASSAY: Xpert MTB/RIF Ultra, COUNT: 286 | 2023-07-21 | GX-IV 4D | 838507 | Module B4 encounter multiple temperature related error Due for Maintenance | Replaced the Defective Module SN: 861942 with new Module SN: 905958 | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | 900-0491rs | DR NO. 10221 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663345238517223469/5663345238517223469_signature_28.png | Joy N. Nio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5663345238517223469/5663345238517223469_signature_31.png
| 2023-07-28 | 2023-07-28 14:18 | Mini Forest Health Center - District C | Quezon Boulevard, Brgy. 23-C Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25337 | DOH | miniforestlaboratory@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF, COUNT: 196 ASSAY: Xpert MTB/RIF Ultra, COUNT: 286 | 2023-07-21 | GX-IV 4D | 838507 | Module B4 encounter multiple temperature related error Due for Maintenance | Replaced the Defective Module SN: 861942 with new Module SN: 905958 | Perform Preventive Maintenance and Module Replacement | All Modules are operational |  | 900-0491rs | DR NO. 10221 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663345228518071331/5663345228518071331_signature_28.png | Joy N. Nio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5663345228518071331/5663345228518071331_signature_31.png
| 2023-07-28 | 2023-07-28 13:46 | Sta. Ignacia Rural Health Unit | Ignacio st. Poblacion West Sta. Ignacia Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25336 | DOH | jdi0973@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 421 | 2023-07-24 | GXIV-D | 831909 | A4 for module replacement, UPS not Backing Up | Module replacement | Performed preventive maintenance Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is operational | UPS for reporting | 1 genexpert module | 10268 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663330823515088487/5663330823515088487_signature_28.png | Jess D. Ibarra Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5663330823515088487/5663330823515088487_signature_31.png
| 2023-07-28 | 2023-07-28 13:03 | Baguio General Hospital and Medical Center | BGH Compound, Baguio City, Benguet | CAR | Benguet | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25335 | PBSP | keithtagudar@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 724 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2120 | 2023-07-28 | Genexpert IV | 834685 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement - 786822 - 210068086 | Genexpert Module - 210068086 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663301864642514343/5663301864642514343_signature_28.png | Keith William Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5663301864642514343/5663301864642514343_signature_31.png
| 2023-07-28 | 2023-07-28 12:49 | Baguio General Hospital | Baguio City, Benguet | CAR | Benguet | TB Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25334 | PBSP | keithtagudar@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 35 ASSAY: Xpert MTB/RIF Ultra, COUNT: 308 ASSAY: Others, COUNT: 7 | 2023-07-28 | Genexpert IV | 120000171 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663293814647057804/5663293814647057804_signature_28.png | Keith William Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5663293814647057804/5663293814647057804_signature_31.png
| 2023-07-28 | 2023-07-28 12:23 | La Castellaña RHU | Municipality of La Castellaña | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25333 | DOH | milouepascual@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 818 |  | GX4 | 838483 | Module B3 failed in the calibration | Conduct PM and save files | Run Xpert check for 4 modules | Module B1, B2 and B4 are functional and passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663276120768185294/5663276120768185294_signature_28.png | Mira Luna Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/5663276120768185294/5663276120768185294_signature_31.png
| 2023-07-28 | 2023-07-28 11:40 | Bagong Barrio Health Center | Bagong Barrio, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25332 | DOH | cruzernitamay@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1337 | 2023-07-24 | GXIV | 831906 | Module A4 for replacement UPS for troubleshooting | - replaced A4 - checked and cleaned cartridge bay, plunger rod, filter - generated reports - checked UPS' back up mode (not working) | Replaced A4 | Unit is operational. UPS is recommended for repair/replacement. |  | GeneXpert Module SN: 210058431 | 10260 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663267338313361540/5663267338313361540_signature_28.png | Ernita May Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5663267338313361540/5663267338313361540_signature_31.png
| 2023-07-28 | 2023-07-28 11:36 | La Castellaña RHU | Municipality of La Castellaña | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25331 | DOH | milouepascual@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 818 |  | GX4 | 838483 | Module B3 failed in calibration | Conduct PM and save files | Run Xpert check for 4 modules | Module 1, 2 and 4 passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663266730764723260/5663266730764723260_signature_28.png | Mira Luna Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/5663266730764723260/5663266730764723260_signature_31.png
| 2023-07-28 | 2023-07-28 11:18 | Calabanga Rural Health Unit | San Pablo St., Calabanga | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25330 | DOH | martyarc11@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 849 |  | GX4 | 838398 | B1,B2 module report. For XpertCheck | Performed XpertCheck: Cleaned cartridge bays, plunger rod, fan filter Generate reports | Performed XpertCheck | Unit is operational B1 - Passed B2 - Service Required B3 - Service Required B4 - Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663250061423877326/5663250061423877326_signature_28.png | Marty R. Arcayera | https://www.jotform.com/uploads/macarewecare/222927978372067/5663250061423877326/5663250061423877326_signature_31.png
| 2023-07-28 | 2023-07-28 10:51 | Passi City Health Office | Poblacion, Passi City | 6 | Iloilo | Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 25329 | PBSP | passicholab@gmail.com | 23 |  | 2023-07-28 | GX4 | 804701 | Module B2-B3 with serial number 732595,656548 | Replace modules | Replace module B2,B3 with new serial number 210100430,210100442 | All modules ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663225634478546669/5663225634478546669_signature_28.png | John Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5663225634478546669/5663225634478546669_signature_31.png
| 2023-07-28 | 2023-07-28 10:41 | Passi City Health Office | Poblacion, Passi city | 6 | Iloilo | Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 25328 | PBSP | passicholab@gmail.com | 23 |  | 2023-07-28 | GX4 | 807610 | Module A2 Hardware Fail serial number 761282 | Replace module A2 | Replace Module A2 with new serial number 210100416 | Module A1, 2,3 ok to use module A4 Hardware fail. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663218374477186516/5663218374477186516_signature_28.png | John Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5663218374477186516/5663218374477186516_signature_31.png
| 2023-07-28 | 2023-07-28 10:02 | GOVERNOR CELESTINO GALLARES MEMORIAL HOSPITAL | TAGBILARAN BOHOL | 7 | Bohol | TB LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25327 | DOH | majieocmen@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 97 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1987 ASSAY: Xpert HIV-1 Viral Load, COUNT: 82 | 2023-07-17 | GX-IV | 820228 | FOR MODULE A2 AND MODULE A4 REPLACEMENT | COLECT IQR , SYSTEM LOG, AND ARCHIVE | PERFORM MODULE REPLACEMENT A2 DEF SN : 634601 WITH NEW SN:894275 A4 DEF SN: 722094 WITH NEW SN:210019612 | EQUIPMENT OPERATIONAL |  | GX MODULE 6-COLOR | 1011710132 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663201517343356652/5663201517343356652_signature_28.png | Ma. Jernai Sosoban | https://www.jotform.com/uploads/macarewecare/222927978372067/5663201517343356652/5663201517343356652_signature_31.png
| 2023-07-28 | 2023-07-28 10:17 | Victoria Rural Health Unit | Brgy. Bulo Victoria Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25326 | DOH | diazrmt@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 89 ASSAY: Xpert MTB/RIF Ultra, COUNT: 804 | 2023-07-24 | GXIV-D | 831669 | A3 for module replacement | Module replacement | Performed preventive maintenance Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 10269 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663200808316267512/5663200808316267512_signature_28.png | Alma G. Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5663200808316267512/5663200808316267512_signature_31.png
| 2023-07-28 | 2023-07-28 09:36 | Banlat Health Center | Daisy St, Banlat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25325 | PBSP | reyzelmariesantos@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 95 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1428 |  | GXIV 6C | 804688 | UPS not backing up | Checked UPS Checked line cord; ok UPS does not backup; for reporting Performed PM Generated report | Troubleshoot UPS | GX is operational. UPS for reporting. | Report UPS |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5663179008222084978/5663179008222084978_signature_28.png | Emma Ibarra | https://www.jotform.com/uploads/macarewecare/222927978372067/5663179008222084978/5663179008222084978_signature_31.png
| 2023-07-27 | 2023-07-27 15:34 | Baao MHO | Imperial st., Del Rosario, Poblacion | 5 | Camarines Sur | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25324 | PBSP | naschw.4u@gmail.com | 25 |  |  | GX4 | 808639 | CPU not booting | Checked power cord connections, cleaned ram and ram ports, | Performed troubleshooting steps | Non operational due to cpu not working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662536560523302085/5662536560523302085_signature_28.png | Ignacio Benosa Britanico | https://www.jotform.com/uploads/macarewecare/222927978372067/5662536560523302085/5662536560523302085_signature_31.png
| 2023-07-27 | 2023-07-27 14:24 | Libon Rural Health Unit | Zone-1, Libon | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25323 | DOH | marilynrazal@gmail.com | 29 |  | 2023-07-21 | GX4 | 838397 | B1,B4 for module Replacement | Replaced module 640264>>210039783, 861019>>210026807 Cleaned cartridge bays, fan, plunger rods, replaced filter Generated reports | Replaced module, Performed Preventive maintenence | Unit is operational |  |  | 10237, 10239 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662535980528936573/5662535980528936573_signature_28.png | Marilyn Razal | https://www.jotform.com/uploads/macarewecare/222927978372067/5662535980528936573/5662535980528936573_signature_31.png
| 2023-07-27 | 2023-07-27 14:03 | Tomas Claudio Health Center | Corner E. Quirino - Camus Street, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25322 | DOH | arleen_p_alconera@yahoo.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 166 ASSAY: Xpert MTB/RIF Ultra, COUNT: 566 | 2023-07-21 | GX-IV 4D | 838492 | Module B1 encounter multiple temperature related errors | Replaced the Defective Module SN:862214 with new Module SN:915997 Generate Pre and Post IQR | Perform Module Replacement | All Modules are operational |  | 900-0491r | DR NO. 10222 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662474745012474040/5662474745012474040_signature_28.png | Arleen P. Alconera RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5662474745012474040/5662474745012474040_signature_31.png
| 2023-07-27 | 2023-07-27 14:03 | Tomas Claudio Health Center | Corner E. Quirino - Camus Street, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25321 | DOH | arleen_p_alconera@yahoo.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 166 ASSAY: Xpert MTB/RIF Ultra, COUNT: 566 | 2023-07-21 | GX-IV 4D | 838492 | Module B1 encounter multiple temperature related errors | Replaced the Defective Module SN:862214 with new Module SN:915997 Generate Pre and Post IQR | Perform Module Replacement | All Modules are operational |  | 900-0491r | DR NO. 10222 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662474735012620347/5662474735012620347_signature_28.png | Arleen P. Alconera RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5662474735012620347/5662474735012620347_signature_31.png
| 2023-07-27 | 2023-07-27 13:57 | Toledo City Health Office | Ilihan Toledo City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25320 | PBSP | loriemaediapen@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 769 |  | Gx4 | 808640 | Module 1 SN # 667772 error 1001 and 1002 started July 13, 2023 | Reset module and condduct preventive maintenance | Reset Module B1, Clean plunger and Icore Clean Filter, cartridge bay, | All module are available but B1 encoutered run time termination error after 5 test |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662469555113969983/5662469555113969983_signature_28.png | Lorie Mae C. Diapen | https://www.jotform.com/uploads/macarewecare/222927978372067/5662469555113969983/5662469555113969983_signature_31.png
| 2023-07-27 | 2023-07-27 13:50 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25319 | DOH | laguitannataniel@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 | 2023-07-19 | GXIV | 831677 | Module A4 for replacement | #NAME? | Replaced A4 | Unit is operational. |  | GeneXpert Module SN: 210050153 | 10176 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662466080402833251/5662466080402833251_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5662466080402833251/5662466080402833251_signature_31.png
| 2023-07-27 | 2023-07-27 12:58 | Calamba City Health Office | Old Municipal Site, Calamba | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25318 | DOH | 2052456@aup.edu.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1305 | 2023-07-24 | GXIV | 838225 | Unit for PM. For module replacement of module B4. | Performed PM. Performed module replacement on B4. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module B4 (210043986) in replacement to defective B4 (807732). | DR#10258 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662433588215460086/5662433588215460086_signature_28.png | Mary Jane Lourdes R. Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5662433588215460086/5662433588215460086_signature_31.png
| 2023-07-27 | 2023-07-27 12:39 | San Jose Del Monte City Health Center VI | Muzon, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25317 | PBSP | theadelrosario1@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2879 ASSAY: Xpert HIV-1 Viral Load, COUNT: 139 | 2023-07-24 | GXIV | 807620 | Module C3 for replacement | #NAME? | Replaced C3 | Unit is operational. |  | GeneXpert Module SN: 210100319 | G 5897 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662426316266240603/5662426316266240603_signature_28.png | Krizette Marie Amorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5662426316266240603/5662426316266240603_signature_31.png
| 2023-07-27 | 2023-07-27 12:00 | Hermosa Rural Health Unit | Brgy. Palihan, Hermosa, Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25316 | DOH | kryztynangela@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 | 2023-07-17 | GXIV-D | 831645 | A2 Due for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is operational |  | 1 yenexpert module | 10137 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662407004496229305/5662407004496229305_signature_28.png | Mary Ann G. Patambang | https://www.jotform.com/uploads/macarewecare/222927978372067/5662407004496229305/5662407004496229305_signature_31.png
| 2023-07-27 | 2023-07-27 11:12 | CONG. SIMEON G. TORIBIO MEMORIAL HOSPITAL | CARMEN BOHOL | 7 | Bohol | TB DOTS | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 25315 | DOH | Himaya.vale@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 638 ASSAY: Xpert MTB/RIF Ultra, COUNT: 352 | 2023-07-17 | GX-IV | 838394 | MODULE B1 REPLACEMENT | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | PERFORM MODULE B1 REPLACEMENT DEF SN:861016 WITH NEW SN:210052368 | EQUIPMENT IS OPERATIONAL |  | GX MODULE 6-COLOR | 10118 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662403159214712060/5662403159214712060_signature_28.png | HIMAYA VALE,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5662403159214712060/5662403159214712060_signature_31.png
| 2023-07-27 | 2023-07-27 11:22 | Bagumbayan Rural Health unit | Bagumbayan, Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25314 | DOH | marizalenoid@yahoo.com | 28.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1174 |  | Gx-IV | 838261 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662377274183775967/5662377274183775967_signature_28.png | Ma. Rizal R. Dionela | https://www.jotform.com/uploads/macarewecare/222927978372067/5662377274183775967/5662377274183775967_signature_31.png
| 2023-07-27 | 2023-07-27 11:03 | Cuenca Rural Health Unit | Cuenca LGU | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25313 | DOH | ligayamijares50@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 629 | 2023-07-25 | GXIV | 821297 | Unit for PM. For module replacement of module A3. | Performed PM. Performed module replacement on A3. Generated reports. | PM and module replacement. | Unit is operational. |  |  | GX module A3 (800812) in replacement to defective A3 (749944). |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662366108215976992/5662366108215976992_signature_28.png | Ligaya R. Mijares | https://www.jotform.com/uploads/macarewecare/222927978372067/5662366108215976992/5662366108215976992_signature_31.png
| 2023-07-27 | 2023-07-27 10:46 | San pablo rural health unit | Poblacion san pablo, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25312 | PBSP | deasis.alfanta@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 220 | 2023-07-27 | GX 2 | 834829 | None | Replaced module 787229 with Module 210092867 | Preventive Maintenence, Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662355045016102361/5662355045016102361_signature_28.png | Elvie Alfanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5662355045016102361/5662355045016102361_signature_31.png
| 2023-07-27 | 2023-07-27 10:15 | Burauen District Hospital | Burauen leyte | 8 | Leyte | Tb-dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25311 | DOH | burauenlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 352 | 2023-07-27 | GXIV | 838551 | Unit for Preventive Maintenance | N/A | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662337023315977259/5662337023315977259_signature_28.png | Princess Geven | https://www.jotform.com/uploads/macarewecare/222927978372067/5662337023315977259/5662337023315977259_signature_31.png
| 2023-07-27 | 2023-07-27 08:56 | Bogo City Health Office | Poblacion, Bogo City | 7 | Cebu | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 25310 | PBSP | bogocholab@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 882 | 2023-07-26 | GX4 | 808630 | Due for xpertcheck | Perform Xpertcheck. | Perform Preventive maintenance for the machine and perform xpertcheck for all module's. | Module D2 and D4 fail for xpertcheck, Module D1 and D3 passed for xpertcheck. | UPS not charging need to repair. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5662289253916939959/5662289253916939959_signature_28.png | Reggie Dahili | https://www.jotform.com/uploads/macarewecare/222927978372067/5662289253916939959/5662289253916939959_signature_31.png
| 2023-07-26 | 2023-07-26 11:31 | Norala District Hospital | Poblacion, Norala, South Cotabato | 12 | South Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25309 | DOH | rezielrds0709@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1038 ASSAY: Xpert MTB/RIF, COUNT: 50 |  | Gx-IV | 838543 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661876916972833047/5661876916972833047_signature_28.png | Reziel D. Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/5661876916972833047/5661876916972833047_signature_31.png
| 2023-07-26 | 2023-07-26 15:35 | Iriga CHO | Sta. Cruz Sur, Iriga City | 5 | Camarines Sur | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25308 | DOH | lindareyesdelima1969@gmail.com | 28 |  |  | GX4 | 838391 | B1,B4 For re-Xpertcheck | Re-XpertCheck B1, & B4 Generate reports | Performed Re-XpertCheck | Unit is operational B1 - B2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661687743014435121/5661687743014435121_signature_28.png | Erlinda R. De Lima | https://www.jotform.com/uploads/macarewecare/222927978372067/5661687743014435121/5661687743014435121_signature_31.png
| 2023-07-26 | 2023-07-26 15:22 | Cordlife Medical Philippines | UP Ayala Technohub, Commonwealth Avenue | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CYTIVA | Patrick Los Baños |  |  |  |  | 25307 | COMMERCIAL | katrina.castro@cordlife.com | 25 |  |  | Sepax 2 | 6220.4122 | Unit intermittently powering off | #NAME? | #NAME? | Unit is not operational; requires parts replacement | Will coordinate with office for back-up unit availability |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661665502129656061/5661665502129656061_signature_28.png | Katrina Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5661665502129656061/5661665502129656061_signature_31.png
| 2023-07-26 | 2023-07-26 15:34 | Bantayan Rural Health Unit | Bantayan Island, Poblacion | 7 | Cebu | Laborattory | YES | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25306 | DOH | mcdelle916bausin@gmail.com | 26 |  |  | Gx 4 | 838509 | Defective module B1 862414 and B4 862356 | Clean fan filter | Install new modules B1 210093182 & B4 210093185 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661664298271523871/5661664298271523871_signature_28.png | Roxanne M. Bausin | https://www.jotform.com/uploads/macarewecare/222927978372067/5661664298271523871/5661664298271523871_signature_31.png
| 2023-07-26 | 2023-07-26 15:36 | Mexico Rural Health Unit I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25305 | DOH | jasperdizon20@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 468 | 2023-07-20 | Genexpert IV | 831877 | Monitor sometimes turn off | Troubleshoot Cables | Performed Maintenance Gathered Reports Performed Module Replacement (757998 - 753257) Performed Functionality Test | Unit is Operational | Monitor for monitoring | Genexpert Module 753257 | 10207 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661663788718136191/5661663788718136191_signature_28.png | Jonabel Lumba | https://www.jotform.com/uploads/macarewecare/222927978372067/5661663788718136191/5661663788718136191_signature_31.png
| 2023-07-26 | 2023-07-26 14:55 | Batis Health Center | Manalo St., San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25304 | DOH | almacenaio16@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1914 ASSAY: Xpert HIV-1 Viral Load, COUNT: 62 | 2023-07-25 | GXIV | 831598 | Unit for PM. For module replacement of module A2. | Performed PM. Performed module replacement on module A2. Generated reports | PM and module replacement. | Unit is operational. |  | GX module A2 (220001434) in replacement to defective A2 (604349). | DR#10277 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661648898217440782/5661648898217440782_signature_28.png | Rosario Almacen | https://www.jotform.com/uploads/macarewecare/222927978372067/5661648898217440782/5661648898217440782_signature_31.png
| 2023-07-26 | 2023-07-26 13:52 | Pili Rural Health Unit | Pili | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25303 | DOH | lfsanao724@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1274 |  | GX4 | 838408 | B3 For module replacement | Module replaced: B3 861110>>210029760 Cleaned Cartridge bays, plunger rod Generate reports | Performed module replacement, preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661603186014439686/5661603186014439686_signature_28.png | Liezl Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/5661603186014439686/5661603186014439686_signature_31.png
| 2023-07-26 | 2023-07-26 13:19 | Potrero Health Center | Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25302 | DOH | ailyn.santos19@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2095 | 2023-07-24 | GXIV | 804715 | Module B1 for replacement | #NAME? | Replaced module B1 | Unit is operational. |  | GeneXpert Module SN: 210100320 | G 5898 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661584706246568667/5661584706246568667_signature_28.png | Ailyn Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5661584706246568667/5661584706246568667_signature_31.png
| 2023-07-26 | 2023-07-26 13:20 | Domingo Santiago Health Center | Sampaloc, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25301 | DOH | fei_paas@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1148 | 2023-07-24 | GXIV 6C | 831527 | Module A1 for replacement | Performed PM Replaced Module A2 Generated report | Replaced Module A1: 6647645 -> 660595 | Unit is operational. All modules working |  | GX module | DR 10243 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661584444916352681/5661584444916352681_signature_28.png | Flory Fe A. Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5661584444916352681/5661584444916352681_signature_31.png
| 2023-07-26 | 2023-07-26 13:06 | GARCIA MEMORIAL PROVINCIAL HOSPITAL | TALIBON BOHOL | 7 | Bohol | TB DOTS | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 25300 | DOH | ybanezrosalina19@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 842 | 2023-07-17 | GX-IV | 838203 | FOR MODULE B2 AND B3 REPLACEMENT | COLLECT IQR , SYSTEM LOG, AND ARCHIVE | PERFORM MODULE B2 DEF SN:807020 WITH NEW SN:655947 MODULE B3 DEF SN:807025 NEW SN:653675 | EQUIPMENT IS OPERATIONAL |  | GX MODULE 6-COLOR | 1012410125 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661580092783182226/5661580092783182226_signature_28.png | ROSALINA A. YBANEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5661580092783182226/5661580092783182226_signature_31.png
| 2023-07-26 | 2023-07-21 15:43 | Mandaluyong City Health Office | Maysilo, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25299 | DOH | rsjalova1974@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2507 ASSAY: Xpert HIV-1 Viral Load, COUNT: 153 | 2023-07-19 | GXIV | 831596 | Modules C1, C2, C4 for replacement | #NAME? | Replaced C1, C2, C4 | Unit is operational. |  | GeneXpert Module SN: 619782, 210041087, 617065 | 10178 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661542886244723453/5661542886244723453_signature_28.png | Ann Patricia Abella | https://www.jotform.com/uploads/macarewecare/222927978372067/5661542886244723453/5661542886244723453_signature_31.png
| 2023-07-26 | 2023-07-21 15:43 | Mandaluyong City Health Office | Maysilo, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25298 | DOH | rsjalova1974@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2507 ASSAY: Xpert HIV-1 Viral Load, COUNT: 153 | 2023-07-19 | GXIV | 831596 | Modules C1, C2, C4 for replacement | #NAME? | Replaced C1, C2, C4 | Unit is operational. |  | GeneXpert Module SN: 619782, 210041087, 617065 | 10178 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661542546242713519/5661542546242713519_signature_28.png | Ann Patricia Abella | https://www.jotform.com/uploads/macarewecare/222927978372067/5661542546242713519/5661542546242713519_signature_31.png
| 2023-07-26 | 2023-07-21 12:51 | Punturin 3S Health Center | Punturin, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25297 | DOH | lukegonzaga09@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1074 | 2023-07-19 | GXIV | 831914 | Module B1, B4 for replacement | #NAME? | Replaced B1, B4 | Unit is operational. |  | GeneXpert Module SN: 220001462, 633420 | 10181 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661540566247225526/5661540566247225526_signature_28.png | Pauline Mae Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5661540566247225526/5661540566247225526_signature_31.png
| 2023-07-26 | 2023-07-21 12:51 | Punturin 3S Health Center | Punturin, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25296 | DOH | lukegonzaga09@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1074 | 2023-07-19 | GXIV | 831894 | Module B1, B4 for replacement | #NAME? | Replaced B1, B4 | Unit is operational. |  | GeneXpert Module SN: 220001462, 633420 | 10181 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661540296247839793/5661540296247839793_signature_28.png | Pauline Mae Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5661540296247839793/5661540296247839793_signature_31.png
| 2023-07-26 | 2023-07-26 11:50 | Marilao Rural Health Unit I | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 25295 | PBSP | jmjenilo27@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1537 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 236 |  | GXIV | 807466 | UPS problem | #NAME? | Checked UPS | Unit is operational. UPS for replacement/repair. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661539406244308059/5661539406244308059_signature_28.png | Jonathan Jenilo | https://www.jotform.com/uploads/macarewecare/222927978372067/5661539406244308059/5661539406244308059_signature_31.png
| 2023-07-26 | 2023-07-26 11:39 | Marilao Rural Healt Unit I | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25294 | DOH | jmjenilo27@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1660 ASSAY: Xpert HIV-1 Viral Load, COUNT: 82 | 2023-07-26 | GXIV | 831897 | Modules A1, A2 for replacement | #NAME? | Replaced A1, A2 | Unit is operational. |  | GeneXpert Module SN: 694358, 694368 | 10266, 10267 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661538436247979257/5661538436247979257_signature_28.png | Jonathan Jenilo | https://www.jotform.com/uploads/macarewecare/222927978372067/5661538436247979257/5661538436247979257_signature_31.png
| 2023-07-26 | 2023-07-26 11:41 | Esperanza Health Center | Sta. Mesa, Manila | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25293 | DOH | jong.agdamag@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1273 | 2023-07-24 | GXIV 6 Colors | 831891 | Module A1 and A2 for replacement | Performed PM Replaced module A1 and A2 Generated report | Replaced Modules A1: 784638 -> 641895 A2: 784635 -> 642924 | Unit is operational. All modules working |  | GX Module - 6 colors - 2mod | DR 10241, 10242 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661528634918016650/5661528634918016650_signature_28.png | Victor Agdamag Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5661528634918016650/5661528634918016650_signature_31.png
| 2023-07-26 | 2023-07-26 11:10 | Baybay City Health Office | Baybay city | 8 | Leyte | Tb-dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25292 | DOH | sajecopooh20@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 307 ASSAY: Xpert MTB/RIF Ultra, COUNT: 864 | 2023-07-26 | GXIV | 838555 | Mudule A3 H/W Failed | #NAME? | #NAME? | Module A3 - Service required Module A1 - multiple temperature error |  | Xpertcheck cartridge | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661505205313815429/5661505205313815429_signature_28.png | Safah Jean Ganieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5661505205313815429/5661505205313815429_signature_31.png
| 2023-07-26 | 2023-07-26 10:22 | Cagayan de Oro City Health Office | Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25291 | DOH | chocdotbdots@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3201 |  | GeneXpert | 838441 | H/W Fail on Module B1/806412 and B4/805415 | N/A | Preventive Maintenance Module Replacement Generate Report files Database Back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661497481819119695/5661497481819119695_signature_28.png | Piquero, Heide | https://www.jotform.com/uploads/macarewecare/222927978372067/5661497481819119695/5661497481819119695_signature_31.png
| 2023-07-26 | 2023-07-26 10:24 | Cagayan de Oro City Health Office | Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25290 | DOH | chocdotbdots@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3201 |  | GeneXpert | 838435 | H/W Fail Module C4/681968 | N/A | Module Replacement Preventive Maintenance Generate Report files Database Back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661496891811670541/5661496891811670541_signature_28.png | Piquero,Heide | https://www.jotform.com/uploads/macarewecare/222927978372067/5661496891811670541/5661496891811670541_signature_31.png
| 2023-07-26 | 2023-07-25 12:46 | Lung Center of the Philippines | Quaezon Avenue,QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25289 | MACARE-HBDC | lcpmicrolab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 716 ASSAY: Xpert SARS-CoV-2, COUNT: 684 ASSAY: Xpert HIV-1 Viral Load, COUNT: 7 |  | GX4 | 804205 | For maintenance d3 and d4 disabled encounter temp error | Cleaned unit,Fliter,plunger rod,cartridge bay Gathered reports | Maintenance | D3 and D4 disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661388333619069732/5661388333619069732_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5661388333619069732/5661388333619069732_signature_31.png
| 2023-07-26 | 2023-07-25 12:53 | Lung Center of the Philippines | Quezon Avenue,QC | NCR | NCR 2nd - Quezon City | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25288 | MACARE-HBDC | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 458 ASSAY: Xpert SARS-CoV-2, COUNT: 1604 ASSAY: Xpert HIV-1 Viral Load, COUNT: 76 ASSAY: Others, COUNT: 1 ASSAY: Others, COUNT: 9 | 2023-07-25 | GX4 | 838837 | For maintenance and due fro xpertcheck, b4 disabled h/w failed | Performed xpertcheck, Cleaned unit,filter,cartridge bay,pluger rod. Gathered reports Test run unit | Xpertcheck and Maintenance | B1,B2,B3, passed xpertcheck, B4 Failed xpertcheck | For module replacement B4 failed in xpertcheck | Xpertcheck | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661386813616796344/5661386813616796344_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5661386813616796344/5661386813616796344_signature_31.png
| 2023-07-26 | 2023-07-25 12:11 | Lung Center of the Philippines | Quezon Ave.,Q.C | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25287 | PBSP | lcpmicrolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4049 ASSAY: Xpert SARS-CoV-2, COUNT: 4426 | 2023-07-24 | GXXVI | 801080 | For Module replacement B1 628060 - 210100356 and B3 encounter loss communication error | Performed module replacement Cleaned unit,filter,plunger rod,cartridge bay Gathered reports. Replaced cable for module b3 Test run unit. | Module replacement and maintenance | 15 module operational b3 for observation |  | 900-0508 | G 5899 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5661385933617470089/5661385933617470089_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5661385933617470089/5661385933617470089_signature_31.png
| 2023-07-25 | 2023-07-19 14:47 | Tatalon health center | Victory Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25286 | MACARE-HBDC | jel472020@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 14 ASSAY: Xpert BCR-ABL Ultra, COUNT: 739 | 2023-07-17 | GX4 | 831610 | For module replacement A1 754328 - 639129,A2 754332 - 640485, A3 754333 - 210001864 | Performed module replacement,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900-0491rs | 10011016210161 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660819461904958569/5660819461904958569_signature_28.png | JSmin A Magtalas | https://www.jotform.com/uploads/macarewecare/222927978372067/5660819461904958569/5660819461904958569_signature_31.png
| 2023-07-25 | 2023-07-18 12:17 | East Avenue Medical Center | East avenue q.c | NCR | NCR 2nd - Quezon City | Tb dots | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25285 | PBSP | roseannparnes@yahoo.com | 31 | ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3656 |  | GX4 | 803087 | As per end user they encounter error 5007 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | For observation c2 and c3 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660818751903234751/5660818751903234751_signature_28.png | ROSE ANN PARNES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660818751903234751/5660818751903234751_signature_31.png
| 2023-07-25 | 2023-07-18 10:18 | East Avenue Medical Center | East avenue q.c | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25284 | MACARE-HBDC | eamcmolecularlab@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 139 ASSAY: Xpert HPV, COUNT: 540 ASSAY: Xpert SARS-CoV-2, COUNT: 2769 ASSAY: Others, COUNT: 2 |  | Gx4 10c | 110014085 | B2 module encounter continous 2014 error | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Troubleshooting and maintenance | B2 enable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660816261909050962/5660816261909050962_signature_28.png | JEV RYAN BONIFACIO | https://www.jotform.com/uploads/macarewecare/222927978372067/5660816261909050962/5660816261909050962_signature_31.png
| 2023-07-25 | 2023-07-18 10:00 | East avenue medical center | East avenue q.c | NCR | NCR 2nd - Quezon City | Molecular | NO | INSTALLATION | CEPHEID | Noel Floresca |  |  |  |  | 25283 | COMMERCIAL | eamcmolecularlab@gmail.com | 26 |  |  | GX4 10c | 110016781 | For installation | Performed installation | Installation and set up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660815081907605386/5660815081907605386_signature_28.png | Jev Ryan Bonifacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5660815081907605386/5660815081907605386_signature_31.png
| 2023-07-25 | 2023-07-25 14:40 | Masambong Health Center | Capoas, Masambong, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25282 | PBSP | michellebonrostro@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1135 | 2023-07-24 | GXIV 6C | 814160 | Module B2 for replacement | Performed PM Replaced Module B2 Generated report | Replaced module B2 814160 -> 210068088 | unit is operational. All modules working |  | GX Module 6 Colors | G 5890 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660786361426371725/5660786361426371725_signature_28.png | Michelle Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/5660786361426371725/5660786361426371725_signature_31.png
| 2023-07-25 | 2023-07-25 15:04 | Apalit Rural Health Unit | Apalit Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25281 | DOH | duenasblesilda@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 857 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 | 2023-07-19 | GXIV-D | 831999 | A1 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance 0erformed module replacement Archived test results Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 10192 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660785404643560810/5660785404643560810_signature_28.png | Bkesilda Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5660785404643560810/5660785404643560810_signature_31.png
| 2023-07-25 | 2023-07-25 14:45 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25280 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | VORTEX MIXER | K3CT82010446 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 200RPM AND 3000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660769361337694770/5660769361337694770_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660769361337694770/5660769361337694770_signature_31.png
| 2023-07-25 | 2023-07-25 14:32 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25279 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | VORTEX MIXER | K3CT82010445 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 200RPM AND 3000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660766831337541257/5660766831337541257_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660766831337541257/5660766831337541257_signature_31.png
| 2023-07-25 | 2023-07-25 14:30 | Bongabon District Hospital | Curva, Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25278 | DOH | tsantuajr0128@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 419 | 2023-07-20 | Genexpert IV | 831562 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Installed HIV assay Performed Functionality Test | Unit is Operational | Module Replacement = 754110 - 210014943 | Genexpert Module = 210014943 | 10208 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660762257512787244/5660762257512787244_signature_28.png | Joan Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5660762257512787244/5660762257512787244_signature_31.png
| 2023-07-25 | 2023-07-25 14:29 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAT | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  |  |  | IRIS | 25277 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | STATSPIN EXPRESS 2 | 731M50106146 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 1000RPM AND 3000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660759091332096162/5660759091332096162_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660759091332096162/5660759091332096162_signature_31.png
| 2023-07-25 | 2023-07-25 14:24 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25276 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | VORTEX MIXER | K3CT82010021 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEED AT 200RPM AND 3000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660757451331809849/5660757451331809849_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660757451331809849/5660757451331809849_signature_31.png
| 2023-07-25 | 2023-07-25 14:20 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25275 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | MEDIFUGE | 42358139 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 1000RPM AND 500RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660754631335940716/5660754631335940716_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660754631335940716/5660754631335940716_signature_31.png
| 2023-07-25 | 2023-07-25 14:18 | Bocaue Rural Health Unit | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25274 | DOH | zaychi20@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1123 | 2023-07-24 | GXIV | 831691 | Module B4 for replacement | #NAME? | Replaced B4 | Unit is operational. |  | GeneXpert Module SN: 624180 | 10263 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660752656324678729/5660752656324678729_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5660752656324678729/5660752656324678729_signature_31.png
| 2023-07-25 | 2023-07-25 14:15 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  |  |  | LW SCIENTIFIC | 25273 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | EZ SWING | 1101273 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 1000RPM AND 5000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660752191331997035/5660752191331997035_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660752191331997035/5660752191331997035_signature_31.png
| 2023-07-25 | 2023-07-25 14:10 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25272 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | MYSPIN 12 | HSH12116 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEED AT 500RPM AND 12500RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660748801337775686/5660748801337775686_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660748801337775686/5660748801337775686_signature_31.png
| 2023-07-25 | 2023-07-25 14:01 | Panabo City Health Office | Poblacion, Panabo City | 11 | Davao del Norte | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25271 | DOH | nadineeust67@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2430 |  | GX-IV 4D | 838295 | Module B1 and B2 encounter multiple temperature errors Due for Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Replaced the Defective Module SN : 808297 with new Module SN: 210057509 Replaced the Defective Module SN: 634050 with new Module SN: 635801 | Perfrom Preventive Maintenance and Module Replacement | Only Three Modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660746960217627200/5660746960217627200_signature_28.png | Louella Jane G. Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5660746960217627200/5660746960217627200_signature_31.png
| 2023-07-25 | 2023-07-25 14:06 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25270 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | MYSPIN 12 | HSH12119 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 500RPM AND 12500RPM | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660745351338572665/5660745351338572665_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660745351338572665/5660745351338572665_signature_31.png
| 2023-07-25 | 2023-07-25 14:03 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25269 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | MYSPIN 12 | HSH12117 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED SPEEDS AT 500RPM AND 12500RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660743341335474851/5660743341335474851_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660743341335474851/5660743341335474851_signature_31.png
| 2023-07-25 | 2023-07-25 13:59 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25268 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | THERMAL MIXER | 1014319110746 | NONE | NONE | PERFORMED MAINTENANCE VERIFIED TEMPERATURES AT 55°C AND 60°C VERIFIED SPEED AT 500RPM AND 2000RPM PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660741441334401692/5660741441334401692_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660741441334401692/5660741441334401692_signature_31.png
| 2023-07-25 | 2023-07-25 13:55 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | RBC BIOSCIENCE | Carlos Leniel Origenes |  |  |  |  | 25267 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | MAGCORE PLUS II | P26R021720135 | NONE | NONE | PERFORMED MAINTENANCE OF UNIT CHECKED AND VERIFIED ALIGNMENTS PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660739311338210330/5660739311338210330_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660739311338210330/5660739311338210330_signature_31.png
| 2023-07-25 | 2023-07-25 13:52 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | RBC BIOSCIENCE | Carlos Leniel Origenes |  |  |  |  | 25266 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | MAGCORE PLUS II | P26R022020135 | NONE | NONE | PERFORMED MAINTENANCE OF UNIT VERIFIED AND CHECKED ALIGNMENTS PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONALITY | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660737131336297033/5660737131336297033_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660737131336297033/5660737131336297033_signature_31.png
| 2023-07-25 | 2023-07-25 13:48 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25265 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | TSX2330FZ | 1161322301200520 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURE AT -20°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660734851334040335/5660734851334040335_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660734851334040335/5660734851334040335_signature_31.png
| 2023-07-25 | 2023-07-25 13:43 | Atang dela Rama Health Center | Pampanga St., Gagalangin | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  |  |  |  | 25264 | DOH | monskie_22@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1374 | 2023-07-19 | GXIV | 831644 | For installation of new AVR. | Performed installation of AVR. Generated reports. Pulled-out defective AVR. | Installation. | Unit is operational. |  | Stavol AVR-SVC-1 (SVC1K190320MX058) ins replacement to defective Stavol AVR-SVC-1 (SVC1K190320MX039). | DR#10203 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660733824886373624/5660733824886373624_signature_28.png | Maria Suzette Melencio | https://www.jotform.com/uploads/macarewecare/222927978372067/5660733824886373624/5660733824886373624_signature_31.png
| 2023-07-25 | 2023-07-25 13:13 | St. Elizabeth Hospital Inc. | General Santos City | 12 | South Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25263 | COMMERCIAL | laboratory@sehi.ph | 20.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 165 ASSAY: Xpert SARS-CoV-2, COUNT: 1530 |  | Gx-IV (10C) | 110003713 | All modules H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | All modules subject for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660733481871076744/5660733481871076744_signature_28.png | Roselle Jave Dignadice | https://www.jotform.com/uploads/macarewecare/222927978372067/5660733481871076744/5660733481871076744_signature_31.png
| 2023-07-25 | 2023-07-25 13:45 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | THERMO FISHER SCIENTIFIC | Carlos Leniel Origenes |  |  |  |  | 25262 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | 88400D | 160414501161220 | NONE | NONE | PERFORMED SURFACE CLEANING PERFORMED DEFROSTING VERIFIED TEMPERATURE AT -80°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660732961339067988/5660732961339067988_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660732961339067988/5660732961339067988_signature_31.png
| 2023-07-25 | 2023-07-25 13:37 | St. Elizabeth Hospital Inc. | General Santos City | 12 | South Cotabato | Laboratory | NO | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25260 | COMMERCIAL | laboratory@sehi.ph | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 41 ASSAY: Xpert SARS-CoV-2, COUNT: 3393 |  | Gx-IV | 831313 | Module 3 temperature error | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660731311872512167/5660731311872512167_signature_28.png | ROSELLE JAVE DIGNADICE | https://www.jotform.com/uploads/macarewecare/222927978372067/5660731311872512167/5660731311872512167_signature_31.png
| 2023-07-25 | 2023-07-21 10:30 | Social Hygiene Clinic - General Santos City | Brgy. Lagao, General Santos | 12 | South Cotabato | Laboratory | YES | INSTALLATION | CEPHEID | Jesson Tupas |  |  |  |  | 25261 | PBSP | dontingzon@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | Gx-IV | 120000147 | N/A | Genexpert machine Installation Installed Dell laptop SN: CPH0822F40058 Import HIV-Viral load Import MTB - Ultra | Genexpert machine Installation | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660731311872947617/5660731311872947617_signature_28.png | Jonnel Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5660731311872947617/5660731311872947617_signature_31.png
| 2023-07-25 | 2023-07-25 13:41 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | PERKIN ELMER | Carlos Leniel Origenes |  |  |  |  | 25259 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | PRE NAT II | 3310-2141 | NONE | NONE | PERFORMED MAINTENANCE CHECKED AND VERIFIED ALIGNMENTS PERFORMED WATER RUN PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660731081332385349/5660731081332385349_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660731081332385349/5660731081332385349_signature_31.png
| 2023-07-25 | 2023-07-25 13:37 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  |  |  | CONDURA | 25258 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | CTD311MNI | 18956716056 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 8°C, 2°C, AND -20°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660728731335115855/5660728731335115855_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660728731335115855/5660728731335115855_signature_31.png
| 2023-07-25 | 2023-07-25 13:33 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  |  |  | CONDURA | 25257 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | CBC283-NL | 688103125W | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURE AT 2°C AND 8°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660726381331520368/5660726381331520368_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660726381331520368/5660726381331520368_signature_31.png
| 2023-07-25 | 2023-07-25 13:29 | BIOPATH CLINICAL DIAGNOSTICS -EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  |  |  | CONDURA | 25256 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | CBC283-NL | 688103101W | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURE AT 2°C AND 8°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660723971336076203/5660723971336076203_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660723971336076203/5660723971336076203_signature_31.png
| 2023-07-25 | 2023-07-25 13:29 | Macabebe Rural Health Unit | Macabebe Rural Health Unit | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25255 | DOH | macabeberhutmlrtdl@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1083 |  | GXIV-D | 831893 | A1 and A3 due for module replacement | Module replacement Preventive maintenance | Performed Preventive Maintenance Performed module replacement Archived test results Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660722594647873733/5660722594647873733_signature_28.png | Salva Mari Sese | https://www.jotform.com/uploads/macarewecare/222927978372067/5660722594647873733/5660722594647873733_signature_31.png
| 2023-07-25 | 2023-07-25 12:38 | San Nicolas Health Center | Asuncion St. Binondo | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25254 | DOH | guadex7@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1336 | 2023-07-19 | GXIV | 831618 | Unit for PM. For module replacement of module A4. | Performed PM. Performed module replacement on A4. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module A4 (650710) in replacement to defective A4 (729715). | DR#10200 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660694704882096349/5660694704882096349_signature_28.png | Guadalupe L. Cube | https://www.jotform.com/uploads/macarewecare/222927978372067/5660694704882096349/5660694704882096349_signature_31.png
| 2023-07-25 | 2023-07-25 12:22 | Philippine Orthopedic Center | 78-A Maria Clara St., Sta. Mesa Heights, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25253 | PBSP | cmblanceta@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 993 | 2023-07-24 | GXIV 6C | 803088 | Module B2 for replacement | Performed PM Replaced Module B2 Generated report | Replaced Module B2 618609 -> 210068183 | Unit is operational. All modules working |  | GX Module 6 Color | G 5891 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660685601421775865/5660685601421775865_signature_28.png | Stella Maree Bernardo | https://www.jotform.com/uploads/macarewecare/222927978372067/5660685601421775865/5660685601421775865_signature_31.png
| 2023-07-25 | 2023-07-25 12:24 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  |  |  | ESCO | 25252 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | IFA-32-8 | 2020-T05163 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURE AT 56°C AND 60°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660684281335948843/5660684281335948843_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660684281335948843/5660684281335948843_signature_31.png
| 2023-07-25 | 2023-07-25 12:21 | BIOPATH CLINICAL DIAGNOSTICS - EROD | 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  |  |  | MEDX | 25251 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | LX-B50L | F5020078 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURE AT 129°C PERFORMED FUNTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONR |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660682301338130382/5660682301338130382_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660682301338130382/5660682301338130382_signature_31.png
| 2023-07-25 | 2023-07-25 12:11 | TACLOBAN CITY HEALTH OFFICE | Tacloban city | 8 | Leyte | Tb-dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25250 | DOH | marteetingzon@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 174 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1141 | 2023-07-25 | GXIV | 838421 | Module B1 H/W Failed | #NAME? | #NAME? | Module B1 - Service required Module B4 - multiple temperature error Module B2 and B3 are in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660681299199797625/5660681299199797625_signature_28.png | MARTEE TINGZON | https://www.jotform.com/uploads/macarewecare/222927978372067/5660681299199797625/5660681299199797625_signature_31.png
| 2023-07-25 | 2023-07-25 12:15 | BIOPATH CLINICAL DIAGNOSTICS - EROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  |  |  | ESCO | 25249 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | IFA-32-8 | 2020-T05006 | NONE | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURES AT 56°C AND 60°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONR |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660680401331349664/5660680401331349664_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660680401331349664/5660680401331349664_signature_31.png
| 2023-07-25 | 2023-07-25 12:04 | Kamuning Health Center | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25248 | DOH | vicvic.chua25@yahoo.com | 28 | ASSAY: Xpert HPV, COUNT: 648 ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 508 | 2023-07-19 | GXIV | 831612 | Module A1, A3 for replacement | #NAME? | #NAME? | Unit is operational. |  | GeneXpert Module SN: 210036556, 210029825 | 10182, 10183 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660675931346849095/5660675931346849095_signature_28.png | Jesusa Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5660675931346849095/5660675931346849095_signature_31.png
| 2023-07-25 | 2023-07-25 12:00 | BIOPATH CLINICAL DIAGNOSTICS - E. ROD | G14-15 GARDEN HEIGHTS CONDOMINIUM, 269 E. RODRIGUEZ SR. AVE., QUEZON CITY | NCR | NCR 2nd - Quezon City | MOLECULAR | NO | CALIBRATION MAINTENANCE | OTHERS | Carlos Leniel Origenes |  |  |  | BIOBASE | 25247 | COMMERCIAL | biopathqc@biopath.ph | 22 |  | 2023-07-25 | BKQ-B50 II | BKQ50II2007541 | None | NONE | PERFORMED SURFACE CLEANING VERIFIED TEMPERATURE AT 134°C PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL | NONE | NONE | NONE |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660673551333739721/5660673551333739721_signature_28.png | AMANDA INOCENTES | https://www.jotform.com/uploads/macarewecare/222927978372067/5660673551333739721/5660673551333739721_signature_31.png
| 2023-07-25 | 2023-07-25 11:58 | Minalin Rural Health Unit | Minalin Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25246 | PBSP | bernadetteann0318@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 663 |  | GXIV-D | 834688 | Error 2008 on all modules | Troubleshooting | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Run reliability test Gathered reports Performed functionality test | Unit is operational | Cartridge related error/issue |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660673234641353042/5660673234641353042_signature_28.png | Bernadette Lazaro | https://www.jotform.com/uploads/macarewecare/222927978372067/5660673234641353042/5660673234641353042_signature_31.png
| 2023-07-25 | 2023-07-25 10:49 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25245 | DOH | abigailgsula@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1400 | 2023-07-19 | GXIV-D | 831912 | A1 for module replacement A3 for re-xpertcheck error 1001 | Module replacement Re-xpertcheck | Performed surface cleaning Performed module replacement Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Performed calibration Gathered reports Performed functionality test | Unit is operational | Modules passed calibration A3 under observation | 1 genexpert module 1 xpertcheck kit | 10193 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660633050487025787/5660633050487025787_signature_28.png | Abigail Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5660633050487025787/5660633050487025787_signature_31.png
| 2023-07-25 | 2023-07-25 10:22 | Dr. Wenceslao Dela Paz Health Center | SB Park, Banlat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25244 | DOH | ameliafclemente2001@yahoo.com | 28 | ASSAY: Xpert HPV, COUNT: 492 ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2420 |  | GX IV 6C | 831876 | Module 4: Hardware Failed | Performed PM Xpertcheck Module 4 Passed XpertCheck Continue using Module 4 Will return when more error occured Generated Report | Performed XpertCheck on Module 4 | Unit is operational. Module 4 passed XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5660615481423061949/5660615481423061949_signature_28.png | amelia clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5660615481423061949/5660615481423061949_signature_31.png
| 2023-07-24 | 2023-07-24 16:04 | Cadiz City Health Office | Cadiz City, Negros Occ. | 6 | Negros Occidental | TB DOTS | NO | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 25243 | DOH | zhevebacomo@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1877 |  | GX4 | 838499 | Defective module for B4 with SN 862382 | Conduct PM and save files. | Installation of new module for B4 with SN 777925 | All modules are now functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659954876613006688/5659954876613006688_signature_28.png | Sherlyn Bacomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5659954876613006688/5659954876613006688_signature_31.png
| 2023-07-24 | 2023-07-18 12:34 | Muñoz City Health Office | West Poblacion, Science City of Muñoz, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25242 | DOH | scomrtdl@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 582 | 2023-07-17 | Genexpert IV | 831673 | Module 1 error | Performed Xpertcheck on A1 module (Xpertcheck Result = Passed) | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = (757027 & 761310 - 210005055 & 210001822) | Xpertcheck & Genexpert Module (210005055 & 210001822) | 10136 & 10135 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659942123619784187/5659942123619784187_signature_28.png | Catherine Celestino | https://www.jotform.com/uploads/macarewecare/222927978372067/5659942123619784187/5659942123619784187_signature_31.png
| 2023-07-24 | 2023-07-18 12:34 | Muñoz City Health Office | West Poblacion, Science City of Muñoz, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25241 | DOH | scomrtdl@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 582 | 2023-07-17 | Genexpert IV | 831673 | Module 1 error | Performed Xpertcheck on A1 module (Xpertcheck Result = Passed) | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = (757027 & 761310 - 210005055 & 210001822) | Xpertcheck & Genexpert Module (210005055 & 210001822) | 10136 & 10135 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659942043615676462/5659942043615676462_signature_28.png | Catherine Celestino | https://www.jotform.com/uploads/macarewecare/222927978372067/5659942043615676462/5659942043615676462_signature_31.png
| 2023-07-24 | 2023-07-21 17:11 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Nueva Ecija | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25240 | PBSP | jblmghmpl@gmail.com | 24 | ASSAY: Xpert CT/NG, COUNT: 0 ASSAY: Xpert HPV, COUNT: 0 | 2023-07-20 | Genexpert IV | 834445 | N/a | N/a | Performed Maintenance Gathered Reports Replaced Genexpert PC Performed Functionality Test | Unit is Operational | N/a | Genexpert PC - CPH0622F38701 | G 5826 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659940833616723294/5659940833616723294_signature_28.png | Ella Mae Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5659940833616723294/5659940833616723294_signature_31.png
| 2023-07-24 | 2023-07-21 19:15 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Immunology | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25240 | MACARE-HBDC | jblmgh.is@gmail.com | 24 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 200 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HBV VL, COUNT: 24 | 2023-07-20 | Genexpert IV | 812571 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 684570 & 684571 - 634842 & 635243 | Genexpert Module = 635243 & 634842 | 10142 & 10141 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659940833619341744/5659940833619341744_signature_28.png | Lerdie Boy Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5659940833619341744/5659940833619341744_signature_31.png
| 2023-07-24 | 2023-07-24 13:08 | Balamban Rural Health Unit | Poblacion, Balamban | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 25239 | DOH | climaco.mildred@yahoo.com | 22 |  | 2023-07-21 | Gx4 | 838231 | Module B1 Hardware Fail | Replace Module B1 | Replace Module B1 With new Serial number 879690 | All module's ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659848253916277594/5659848253916277594_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5659848253916277594/5659848253916277594_signature_31.png
| 2023-07-24 | 2023-07-24 12:23 | Pontevedra Rural Health Unit | Pontevedra, Negros Occidental | 6 | Negros Occidental | Laboratory Dept. | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 25238 | DOH | pontevedra.laboratory@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 755 |  | GX4 | 838491 | None | Install the new module for B1 with SN 682655 | Replacement of Old module for B1 with SN 862261 | All 4 modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659823675862891469/5659823675862891469_signature_28.png | Mae Ann Bonilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5659823675862891469/5659823675862891469_signature_31.png
| 2023-07-24 | 2023-07-24 12:12 | Minglanilla Rural Health Unit | Pider 1, Minglanilla Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 25237 | DOH | brendagonzagaparedes@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 998 ASSAY: Xpert MTB/RIF, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1423 |  | Gx4 | 838512 | Module 2 Sn 862343 encountered a run time error 1002 during running | Replaced new module for B2 | Replaced new module Sn 863260 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659819964216668610/5659819964216668610_signature_28.png | Brenda Paredes, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5659819964216668610/5659819964216668610_signature_31.png
| 2023-07-24 | 2023-02-28 13:57 | Siliman University Medical Center | Aldecoa Street Dumaguete City Negros Oriental | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 25236 | COMMERCIAL | ellisab.024@gmail.com | 21 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 4 | 2023-02-28 | GX4 | 821471 | Module Fail A3 with Serial number 210070889 | Install New Module | Install new Module with new Serial number 787982 | Module A3 ok to use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659782163912274718/5659782163912274718_signature_28.png | Basille Dipaling | https://www.jotform.com/uploads/macarewecare/222927978372067/5659782163912274718/5659782163912274718_signature_31.png
| 2023-07-24 | 2023-07-21 09:54 | Amai Pakpak Medical Center | Brgy. Datu Saber | BARMM | Lanao del Sur | TB DOTS | YES | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25235 | PBSP | rbarco2007@gmail.com | 28 |  |  | GeneXpert | 807457 | Defective UPS low power output | N/A | Replaced UPS SN/: 40208052003 Install additional AVR | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659741680955592805/5659741680955592805_signature_28.png | Raisah Ubarco | https://www.jotform.com/uploads/macarewecare/222927978372067/5659741680955592805/5659741680955592805_signature_31.png
| 2023-07-24 | 2023-07-21 14:42 | Talomo North Health Center | S.I.R Matina, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25234 | PBSP | alfredojrsantos@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1726 |  | GX-IV 4d | 815385 | Multiple error on B1 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Re arrange the GeneXpert Setup | Perform Preventive Maintenance and Troubleshooting | Module B1 under observation | High temps due to UPS was beside the GeneXpert Machine, and also UPS unit is low battery level |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5659426909919646855/5659426909919646855_signature_28.png | Alfredo H. Santos Jr. RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5659426909919646855/5659426909919646855_signature_31.png
| 2023-07-21 | 2023-07-21 15:56 | Pasay City Social Hygiene Clinic | Malibay Health Center, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25233 | MACARE-HBDC | marsha_dedios@yahoo.com | 26 |  |  | GXIV | 120000141 | For installation | Setup GX machine and accessories Generated report | Installed GeneXpert | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657413691919277748/5657413691919277748_signature_28.png | Marsha Catherine De Dios | https://www.jotform.com/uploads/macarewecare/222927978372067/5657413691919277748/5657413691919277748_signature_31.png
| 2023-07-21 | 2023-07-21 16:18 | Ospital ning Capas | Capas Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25232 | DOH | onclaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1651 | 2023-07-21 | GXIV-D | 831591 | UPS and Module A4 for replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed UPS and Module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module and 1 UPS | 10201 and 10202 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657371826015657201/5657371826015657201_signature_28.png | Hermes Frias | https://www.jotform.com/uploads/macarewecare/222927978372067/5657371826015657201/5657371826015657201_signature_31.png
| 2023-07-21 | 2023-07-21 15:20 | Rosario Reyes Health Center | San Andres, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25231 | DOH | moninabatil16.prime@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2241 | 2023-07-19 | GXIV6C | 831843 | Module A1, A2, and A4 for replacement | Performed PM Replaced 3 modules Grnerated report | Replaced 3 modules A1: 757666 -> 722232 A2: 757667 -> 725116 A4: 757669 -> 730662 | Unit is operational. All modules working |  | 3 GX Modules 6 colors | DR 10184, 10185, and 10186 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657338981918470485/5657338981918470485_signature_28.png | monina batil | https://www.jotform.com/uploads/macarewecare/222927978372067/5657338981918470485/5657338981918470485_signature_31.png
| 2023-07-21 | 2023-07-21 15:05 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25230 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 230 ASSAY: Xpert MTB/RIF, COUNT: 702 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2128 ASSAY: Xpert HIV-1 Viral Load, COUNT: 42 | 2023-07-17 | GXIV-4D | 838285 | Module A3 defective | -Replace defective module A3/676993 with new module SN: 880306 | #NAME? | All modules are functional. |  | 1 GeneXpert Module 6-Color 900-0491R | DR 10131 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657327566453197985/5657327566453197985_signature_28.png | James Ryan Galos | https://www.jotform.com/uploads/macarewecare/222927978372067/5657327566453197985/5657327566453197985_signature_31.png
| 2023-07-21 | 2023-07-21 15:01 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25229 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 228 ASSAY: Xpert MTB/RIF, COUNT: 672 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2231 ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 |  | GXIV-4D | 820206 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657323326458663476/5657323326458663476_signature_28.png | James Ryan Galos | https://www.jotform.com/uploads/macarewecare/222927978372067/5657323326458663476/5657323326458663476_signature_31.png
| 2023-07-21 | 2023-07-21 14:49 | LAPU-LAPU CITY HEALTH OFFICE | Lapu-lapu City, Cebu | 7 | Cebu | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 25228 | DOH | tanya_pax@yahoo.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 442 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1901 | 2023-07-17 | GX-IV | 837661 | FOR MODULE A1,A2, AND A3 REPLACEMENT | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | A1 DEF SN: 656994 ; NEW SN: 894967 A2 DEF SN: 806133 ; NEW SN: 860337 A3 DEF SN: 664865 ; NEW SN: 861244 | EQUIPMENT IS OPERATIONAL |  | GX MODULE 6-COLOR | 101331014510146 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657315483916724379/5657315483916724379_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5657315483916724379/5657315483916724379_signature_31.png
| 2023-07-21 | 2023-07-21 13:30 | LAPU LAPU CITY HEALTH OFFICE | LAPU-LAPU CITY | 7 | Cebu | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 25227 | DOH | tanya_pax@yahoo.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 442 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1901 | 2023-07-17 | GX-IV | 837661 | FOR MODULE A1,A2, AND A3 REPLACEMENT | COLLECT IQR , SYSTEM LOG, AND ARCHIVE | PERFORM MODULE REPLACEMENT A1 DEF SN: 656994 ; NEW SN: 894967 A2 DEF SN: 806133 ; NEW SN: 860337 A3 DEF SN: 664865 ; NEW SN: 861244 | EQUIPMENT IS OPERATIONAL |  | GX MODULE 6 | 101331014510146 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657307071324748592/5657307071324748592_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5657307071324748592/5657307071324748592_signature_31.png
| 2023-07-21 | 2023-07-21 14:07 | Murcia Rural Health Unit | Municipality of Murcia | 6 | Negros Occidental | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 25226 | PBSP | kambal.kapid@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1071 |  | GX2 | 835055 | None | Run Xpert check for 2 mofules | Xpert check calibration | The two modules passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657304036714042259/5657304036714042259_signature_28.png | Allan Benjamin | https://www.jotform.com/uploads/macarewecare/222927978372067/5657304036714042259/5657304036714042259_signature_31.png
| 2023-07-21 | 2023-07-21 13:24 | Belmonte Health Center | Don Quijote St. Sampaloc | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25225 | PBSP | spsurell@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1086 |  | GXIV | 814376 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. All modules passed. | Unit is located temporarily at Calabash Health Center. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657282406714962711/5657282406714962711_signature_28.png | Clarinda Sharon P.Surell | https://www.jotform.com/uploads/macarewecare/222927978372067/5657282406714962711/5657282406714962711_signature_31.png
| 2023-07-21 | 2023-07-21 13:21 | Belmonte Health Center | Don Quijote St. Sampaloc | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25224 | PBSP | spsurell@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1086 |  | GXIV | 808619 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Modules B1, B2 and B4 are service required. | Unit is located temporarily at Calabash Health Center. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657281206717067753/5657281206717067753_signature_28.png | Clarinda SharonP.Surell | https://www.jotform.com/uploads/macarewecare/222927978372067/5657281206717067753/5657281206717067753_signature_31.png
| 2023-07-21 | 2023-07-21 13:37 | Ospital ning Capas | Sto. Domingo 2nd, Municipal Compound, Capas, Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25223 | DOH | onclaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1621 | 2023-07-19 | GXIV-D | 831591 | Defective UPS and module A4, for replacemenet | Module replacement UPS replacement Preventive maintenance | Performed Preventive maintenance Performed module and UPS replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module and UPS | 10204 & 10202 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657275931118770379/5657275931118770379_signature_28.png | Hermes Frias | https://www.jotform.com/uploads/macarewecare/222927978372067/5657275931118770379/5657275931118770379_signature_31.png
| 2023-07-21 | 2023-07-21 12:46 | Pedro Gil Health Center | Pedro Gil St. Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25222 | PBSP | lorrainefabros@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 395 | 2023-07-20 | GX IV 6C | 812829 | Module B2 for replacement | Performed PM Replaced B2 Generated report | Replaced B2 635445 -> 210100154 | Unit is operational |  | GX module 6 colors | G 5825 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657244801913086199/5657244801913086199_signature_28.png | Lorraine Lucila L.Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5657244801913086199/5657244801913086199_signature_31.png
| 2023-07-21 | 2023-07-21 12:22 | M. Icasiano Health Center | Pedro Gil St., Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25221 | DOH | lorrainefabros@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1263 | 2023-07-19 | GXIV 6C | 831684 | Modules A1, A3, A4 for replacement Intermittent problem on CPU not booting up | Checked CPU Cleaned RAM again and switched position; problem solved CPU problem intermittent Performed PM Replaced 3 modules Generated report | Troubleshoot CPU Replaced 3 modules A1: 800869 -> 710480 A3: 757054 -> 711250 A4: 757055 -> 714937 | GX is operational. All modules working CPU for reporting |  | 3 GX modules 6 colors | DR 10187, 10188, 10189 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657233921911149883/5657233921911149883_signature_28.png | Lorraine Lucila L.Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5657233921911149883/5657233921911149883_signature_31.png
| 2023-07-21 | 2023-07-21 10:03 | Carcar City Health Office | Valladolid Carcar City Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 25220 | DOH | chaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1476 |  | Gx | 838283 | B2 failed during calibration and due to error 1002 during operation | Replaced new module Instruct medtech to maintain room temperature during running | Replaced new module with SN 861432 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657216485273192061/5657216485273192061_signature_28.png | Lucille Mae A. Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5657216485273192061/5657216485273192061_signature_31.png
| 2023-07-21 | 2023-07-21 09:54 | Amai Pakpak Medical Center | Brgy. Datu Saber | BARMM | Lanao del Sur | TB DOTS | YES | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25219 | PBSP | rbarco2007@gmail.com | 28 |  |  | GeneXpert | 807457 | Defective UPS low power output | N/A | Replaced UPS SN/: 40208052003 Install additional AVR | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657210874423532068/5657210874423532068_signature_28.png | Raisah Ubarco | https://www.jotform.com/uploads/macarewecare/222927978372067/5657210874423532068/5657210874423532068_signature_31.png
| 2023-07-21 | 2023-07-21 09:13 | Amai Pakpak Medical Center | Begy. Datu Saber, Marawi City | BARMM | Lanao del Sur | Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25218 | PBSP | apmcmollab@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 7937 |  | GeneXpert | 834444 | H/W Fail on Module B1 and B3 | N/A | Preventive Maintenance Module replacement Generate Report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657210494424482645/5657210494424482645_signature_28.png | Rayan S. Cosari, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5657210494424482645/5657210494424482645_signature_31.png
| 2023-07-21 | 2023-07-21 11:03 | Navotas Medical and Wellness Center | M. Naval St. Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25217 | DOH | alejandrinoknowell@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 25 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1460 | 2023-07-18 | GXIV | 831626 | Unit for PM. For module replacement of module A1 and A3. | Performed PM. Performed module replacement in A1 and A3. Generated reports. | PM and module replacement. | Unit is operational. Facility has defective aircon. |  | GX module A1 (743263) and A3 (210019525) in replacement to defective A1 (791798) and A3 (693447). | DR#10169 and DR#10170 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657207165919369196/5657207165919369196_signature_28.png | Noel C. Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5657207165919369196/5657207165919369196_signature_31.png
| 2023-07-21 | 2023-07-21 10:30 | Social Hygiene Clinic - General Santos City | Brgy. Lagao, General Santos | 12 | South Cotabato | Laboratory | YES | INSTALLATION | CEPHEID | Jesson Tupas |  |  |  |  | 25216 | PBSP | dontingzon@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | Gx-IV | 120000147 | N/A | Genexpert machine Installation Installed Dell laptop SN: CPH0822F40058 Import HIV-Viral load Import MTB - Ultra | Genexpert machine Installation | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657180533123838807/5657180533123838807_signature_28.png | Jonnel Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5657180533123838807/5657180533123838807_signature_31.png
| 2023-07-21 | 2023-07-21 08:50 | Eversly childs sanitarium | Mandaue City, Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 25215 | PBSP | alimurungjane@gmail.com | 29 |  |  | Gx 4 | 805270 | Defective module a3 sn 640129 | Clean fan filter, New module install module a3 sn | New module install module a3 sn | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657141412166383842/5657141412166383842_signature_28.png | Juanita V. Alimurung | https://www.jotform.com/uploads/macarewecare/222927978372067/5657141412166383842/5657141412166383842_signature_31.png
| 2023-07-21 | 2023-07-18 08:58 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 25214 | PBSP | cyrilcordova0515@gmail.com | 28 |  |  | Gx 4 | 804838 | Defective module C3 sn 665155 | Install module, clean fan filter | New module installed C3 sn 210092890 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657140232169352144/5657140232169352144_signature_28.png | Cyril cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5657140232169352144/5657140232169352144_signature_31.png
| 2023-07-21 | 2023-07-21 09:33 | Tacloban City Health Office | Tacloban | 8 | Leyte | Tb-dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25213 | DOH | marteetingzon20@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 180 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1133 | 2023-07-21 | GXIV | 838421 | Module B4 H/W Failed | #NAME? | #NAME? | Unit is Operational (Module B4 for monitoring temperature error) |  | 1 Xpertcheck cartridge | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657129659191086187/5657129659191086187_signature_28.png | MARTEE TINGZON | https://www.jotform.com/uploads/macarewecare/222927978372067/5657129659191086187/5657129659191086187_signature_31.png
| 2023-07-21 | 2023-07-21 08:47 | Tondo Medical Center | H. Lopez Blvd. Balut | NCR | NCR 1st - Manila | Covid laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25212 | DOH | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 244 ASSAY: Xpert SARS-CoV-2, COUNT: 1562 ASSAY: Xpert HIV-1 Viral Load, COUNT: 5 | 2023-07-18 | GXIV | 831872 | Unit for PM. For module replacement of module A2 and A3. | Performed PM. Performed module replacement of A2 and A3. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module A2 (210058805) and A3 (210058675) in replacement to defective A2 (612830) and A3 (633967) | DR#10166 and DR#10165 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5657106475918424181/5657106475918424181_signature_28.png | Analyn Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5657106475918424181/5657106475918424181_signature_31.png
| 2023-07-20 | 2023-07-20 11:00 | JARO RURAL HEALTH UNIT | Jaro Leyte | 8 | Leyte | TB-Dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25211 | PBSP | mendoza.anikks@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 639 | 2023-07-20 | GXIV-2D | 834817 | N/a | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5656719700512243551/5656719700512243551_signature_28.png | Anikka Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5656719700512243551/5656719700512243551_signature_31.png
| 2023-07-20 | 2023-07-20 13:32 | Malita Rural Health Unit | Poblacion, Malita Davao Occidental | 11 | Davao Occidental | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25210 | PBSP | mariastellarmt@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF, COUNT: 149 ASSAY: Xpert MTB/RIF Ultra, COUNT: 779 |  | GX-IV 4D | 805258 | Due For Calibration | Run XpertCheck Calibration on Four Modules | Perform XpertCheck Calibration | All Modules Passed the Calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5656704898402227830/5656704898402227830_signature_28.png | Maria Stella Solitana RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5656704898402227830/5656704898402227830_signature_31.png
| 2023-07-20 | 2023-07-20 14:07 | Agusan del Norte Provincial Health Office | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 25209 | PBSP | ronniebalansag74@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1611 |  | GX-4 | 807461 | Module A1 found Service Required from XPC Calibration | None | XPC Calibration Preventive Maintenance | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5656427637668399870/5656427637668399870_signature_28.png | Ronnie M. Balansag | https://www.jotform.com/uploads/macarewecare/222927978372067/5656427637668399870/5656427637668399870_signature_31.png
| 2023-07-20 | 2023-07-20 13:24 | Malapatan Rural Health Unit | Poblacion, Malapatan, Sarangani Province | 12 | Sarangani | Tb-lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25208 | DOH | rhumalapatan.12@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 315 |  | Gx-IV | 838419 | Module 1 stuck cartridge | Fixed/repair module 1 stuck cartridge | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5656405239216201863/5656405239216201863_signature_28.png | CERES MAE JIMENEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5656405239216201863/5656405239216201863_signature_31.png
| 2023-07-20 | 2023-07-20 12:14 | Basud Municipal Health Office | Maharlika Hwy., Poblacion II, Basud | 5 | Camarines Norte | RTDL LABORATORY | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25207 | PBSP | gbechano@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 559 |  | GX4 | 807600 | For Preventive Maintenance | Cleaned Cartridge bays, plunger rod, filter Generate reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5656359544519323039/5656359544519323039_signature_28.png | Grace B. Echano | https://www.jotform.com/uploads/macarewecare/222927978372067/5656359544519323039/5656359544519323039_signature_31.png
| 2023-07-20 | 2023-07-20 08:54 | Carcar City City Health Office | Valladollid Carcar City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25206 | DOH | chaska.lp14@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1590 |  | Gx4 | 838283 | Error 2037 B1 and B4 No air Filter damaged | Performed maintenance particular in Icore and plunger Need additional air filter | Performed plunger maintenance, clean Icore, valve, cartridge bay | All modules are operational except B2 waiting for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5656320820483211807/5656320820483211807_signature_28.png | Lucille Mae Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5656320820483211807/5656320820483211807_signature_31.png
| 2023-07-20 | 2023-07-20 09:39 | Vinzons Rural Health Unit | F. O. Balsa St., Brgy. II, Poblacion | 5 | Camarines Norte | RTDL LABORATORY | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25205 | DOH | bardonamabel@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 880 |  | GX4 | 838576 | 686319/B3 for Re-XpertCheck | Performed Re-XpertCheck: 686319/B3 - service required Cleaned Cartridge bays, fan filter, plunger rod, | Performed Re-XpertCheck, Preventive Maintenance | Unit is operational 686319/B3 - service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5656293874512249602/5656293874512249602_signature_28.png | Amabel Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/5656293874512249602/5656293874512249602_signature_31.png
| 2023-07-19 | 2023-07-19 15:26 | CEBU CITY HEALTH OFFICE | CEBU | 7 | Cebu | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 25204 | PBSP | litomanubag@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1033 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-IV | 838882 | FOR MODULE B1 REPLACEMENT | COLLECT IQR,SYSTEM LOG,AND ARCHIVE | PERFORM B1 DEF SN:861765 REPLACEMENT WITH NEW SN: 210100365 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655646772516492955/5655646772516492955_signature_28.png | JOSELITO MANUBAG | https://www.jotform.com/uploads/macarewecare/222927978372067/5655646772516492955/5655646772516492955_signature_31.png
| 2023-07-19 | 2023-07-19 15:15 | Camarin Health Center | Camarin, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25203 | PBSP | jasminfalala@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1674 | 2023-07-19 | GXIV | 815395 | Module B3 for replacement | #NAME? | Replaced module B3 | Unit is operational. |  | GeneXpert Module SN: 210100166 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655603179112694860/5655603179112694860_signature_28.png | Arianne Jasmin Roma | https://www.jotform.com/uploads/macarewecare/222927978372067/5655603179112694860/5655603179112694860_signature_31.png
| 2023-07-19 | 2023-07-19 14:45 | Talisay Municipal Health Office | Maharlika Hwy, Poblacion, Talisay | 5 | Camarines Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25202 | PBSP | savillaeditha@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 748 |  | GX4 | 835188 | For XpertCheck | Performed Xpertcheck Cleaned Cartridge bays, plunger rod, icore Generate Reports | Performed Xpertcheck, Preventive Maintenance | Unit is Operational B1 - passed B2 - passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655601810412869605/5655601810412869605_signature_28.png | Editha V. Savilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5655601810412869605/5655601810412869605_signature_31.png
| 2023-07-19 | 2023-07-19 14:59 | Sta. Lucia Health Center | Sta. Lucia, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25201 | DOH | josiedingle2004@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1296 | 2023-07-18 | GXIV 6C | 831901 | Module A1 and A3 for replacement | Performed PM Replaced modules A1 and A3 Generated report | Replaced modules A1 and A3 A1: 758240 -> 210043899 A3: 758243 -> 688716 | unit is operational. all modules working |  | GX modules 6C | DR 10160 and 10159 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655601763813809401/5655601763813809401_signature_28.png | Josephine L.Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5655601763813809401/5655601763813809401_signature_31.png
| 2023-07-19 | 2023-07-19 14:23 | Our Lady Of Mt. Carmel Medical Center | Saguin, City Of San Fernando, Pampanga | 3 | Pampanga | Blood Bank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Harold Toco |  |  |  |  | 25200 | COMMERCIAL | mt.carmel_laboratory@yahoo.com | 25 |  | 2023-07-12 | MCS+ | 15G043 | None, For PM and calibration. | Performed maintenance and calibration. Replaced pump rotor rubber roller. | Performed preventive maintenance and Calibration. | Unit is operational. | None | Rotor rubber roller. | DR num: 10103 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655576786217833705/5655576786217833705_signature_28.png | Rochelle Loren B. Guiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5655576786217833705/5655576786217833705_signature_31.png
| 2023-07-19 | 2023-07-19 13:48 | Dalaguete Rural Health Unit | Poblacion Dalaguete Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25199 | PBSP | ajreef5517@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1063 |  | Gx4 | 815464 | B1 SN: 716217 encountered a run timr error of 1002 several times and B4 2032 an improper tuning of ultrasonic horn | Performed maintenance and replacement in B1 and B4 conduct self test Run specimen in B1 and B4 | Replaced new module 210100651 and performed preventive maintnenance in B4 such as cleaning of Icore, Ultrasonic horn, plunger maintenance | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655553743148065004/5655553743148065004_signature_28.png | Alaizza Jane Rivera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5655553743148065004/5655553743148065004_signature_31.png
| 2023-07-19 | 2023-07-19 13:11 | Jose Panganiban RHU | Roy Padilla Blvd., South Poblacion, Jose Panganiban | 5 | Camarines Norte | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25198 | PBSP | materesitapenas1@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1671 |  | GX4 | 808641 | For Preventive Maintenance | Cleaned machine fan, fan filter, icore, plunger rod, cartridge bays Generate reports | Performed Preventive Maintenance | Unit is operational | B3 - sufficient report collected. To be reported |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655530411015589457/5655530411015589457_signature_28.png | Ma. Teresita Peñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5655530411015589457/5655530411015589457_signature_31.png
| 2023-07-19 | 2023-07-19 12:59 | Tondo Health Center | Gagalangin, Tondo | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25197 | DOH | malorinamanera@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1766 | 2023-07-18 | GXIV | 831620 | Unit for PM. For module replacement of module A1. | Performed PM. Performed module replacement of module A1. | PM and module replacement | Unit is operational. |  | GX module A1 (210012414) in replacement to defective A1 (638087). | DR#10164 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655521921527201260/5655521921527201260_signature_28.png | Ma. Lorina Lontoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5655521921527201260/5655521921527201260_signature_31.png
| 2023-07-19 | 2023-07-19 12:46 | Potrero Health Center | Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25196 | PBSP | ailyn.santos19@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2067 | 2023-07-19 | GXIV | 804715 | Module B4 for replacement | #NAME? | Replaced module B4 | Unit is operational. |  | GeneXpert Module SN: 210100169 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655513925229502788/5655513925229502788_signature_28.png | Ailyn Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5655513925229502788/5655513925229502788_signature_31.png
| 2023-07-19 | 2023-07-19 12:18 | Dalaguete Rural Health Unit | Poblacion Dalaguete, Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25195 | PBSP | ajreef5517@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1063 |  | Gx4 | 815464 | B1 SN# 716217 encountered run time error of 1002 due to temperature problem and B4 Sn 210073850 improper tuning of ultrasonic horn error 2032 | Replaced new module of B1 and conduct corrective maintenance in B4 Conduct specimen run in B1 and B4 | Replacement of new module in B1 SN 210100651 Clean B4 the ultrasonic horn and restart module | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655509443149892982/5655509443149892982_signature_28.png | Alaizza Jane Rivera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5655509443149892982/5655509443149892982_signature_31.png
| 2023-07-19 | 2023-07-19 11:58 | Parola Health Center | Parola, Tondo | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25194 | DOH | friaslynette@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 883 | 2023-07-18 | GXIV | 831871 | Unit for PM. For module replacement of B1 and B2. | Performed PM. Performed module replacement of B1 and B2. | PM and module replacement. | Unit is operational. |  | GX module B1 (701803) and B2 (701926) in replacement to defective B1 (660338) and B2 (673276). | DR#10167 and DR#10168 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655485891526538167/5655485891526538167_signature_28.png | Lynette Olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5655485891526538167/5655485891526538167_signature_31.png
| 2023-07-19 | 2023-07-19 11:23 | San Isidro Rural Health Unit | Poblacion, San Isidro, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25193 | PBSP | urian_marlyn@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 782 | 2023-07-19 | Genexpert IV | 805107 | Module 4 error | Performed Self Test | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 4 for Reporting | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655467341329069495/5655467341329069495_signature_28.png | Marlyn Urian | https://www.jotform.com/uploads/macarewecare/222927978372067/5655467341329069495/5655467341329069495_signature_31.png
| 2023-07-19 | 2023-07-19 10:18 | Arsenia J. Maximo Super Health Center | Dona Isaura, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25192 | PBSP | elyvette_macalalag@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 155 ASSAY: Xpert MTB/RIF, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1105 ASSAY: Xpert HIV-1 Viral Load, COUNT: 254 |  | GX IV 6C | 814162 | PC's time reverts back to the turned off time | Set windows time to automatic ;problem persists Replaced CMOS battery; problem solved Performed PM Generated report | Troubleshoot Windows Time and CMOS battery | Unit is operational. Date and time displays current time |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655433303815230496/5655433303815230496_signature_28.png | Elyvette M. Basa | https://www.jotform.com/uploads/macarewecare/222927978372067/5655433303815230496/5655433303815230496_signature_31.png
| 2023-07-19 | 2023-07-19 09:51 | Labo District Hospital | Talubatib, Labo | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25191 | DOH | analyn.consuelo@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1164 |  | GX4 | 838207 | B3 for Re-xpertcheck | Performed Re-XpertCheck Cleaned Cartridge bays, plunger rod, icore Generate reports | Performed Re-XpertCheck, Preventive Maintenance | Unit is operational Re-Xpertcheck B-3 passed For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655430661016496469/5655430661016496469_signature_28.png | Analyn C. Consuelo | https://www.jotform.com/uploads/macarewecare/222927978372067/5655430661016496469/5655430661016496469_signature_31.png
| 2023-07-19 | 2023-07-19 08:46 | Mayorga Rural Health Unit | Mayorgq leyte | 8 | Leyte | Tb-dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25190 | PBSP | irrnepsia@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 729 | 2023-07-19 | GXIV | 834819 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5655370607113801554/5655370607113801554_signature_28.png | Irene P. Sia | https://www.jotform.com/uploads/macarewecare/222927978372067/5655370607113801554/5655370607113801554_signature_31.png
| 2023-07-18 | 2023-07-12 15:03 | San Mateo Super Health Center | Guitnang bayan 1, San Mateo, | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25189 | PBSP | wilhelminanartatez@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1327 | 2023-07-10 | GX4 | 808616 | For xpertcheck and maintenance | Cleaned unit.cleaned Filter,plunger rod,cartridge bay,icore,gathered reports,performed xpertcheck, test run unit | Xpertcheck and maintenance | Operational |  | Xpertcheck | 41893 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654855411902111228/5654855411902111228_signature_28.png | Wilhelmina R. Nartatez | https://www.jotform.com/uploads/macarewecare/222927978372067/5654855411902111228/5654855411902111228_signature_31.png
| 2023-07-18 | 2023-07-18 16:20 | E. Rodriguez Health Center | 150 Pajo, Project 2, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25188 | DOH | clairebaptista32786@gmail.com | 31 | ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1073 ASSAY: Others, COUNT: 207 | 2023-07-18 | GXIV 6C | 831848 | Module A1 and A2 for replacement | Performed PM Replaced Module A1 & A2 Generated report | Replaced Module A1 and A2 A1: 757725 -> 210093175 A2: 793284 -> 210093174 | Unit is operational. All modules working |  | 2 GX Modules 6 Colors | DR 10157 & 10158 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654781416318785507/5654781416318785507_signature_28.png | mary claire p baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/5654781416318785507/5654781416318785507_signature_31.png
| 2023-07-18 | 2023-07-18 16:00 | Bagumbong Health Center | Bagumbong, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25187 | DOH | jenilyn_duenas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2412 | 2023-07-18 | GXIV | 831907 | A2 for replacement | #NAME? | Replaced A2 | Unit is operational. Monitor for request for replacement. |  | GeneXpert Module SN: 210045625 | 10163 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654770645227499431/5654770645227499431_signature_28.png | Jenilyn Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5654770645227499431/5654770645227499431_signature_31.png
| 2023-07-18 | 2023-07-18 16:00 | Bagumbong Health Center | Bagumbong, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25186 | DOH | jenilyn_duenas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2412 | 2023-07-18 | GXIV | 831907 | A2 for replacement | #NAME? | Replaced A2 | Unit is operational. Monitor for request for replacement. |  | GeneXpert Module SN: 210045625 | 10163 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654770465224200808/5654770465224200808_signature_28.png | Jenilyn Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5654770465224200808/5654770465224200808_signature_31.png
| 2023-07-18 | 2023-07-18 14:13 | Sta Rosa Community Hospital | Sta Rosa, Olango Island Lapulapu Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25185 | DOH | ariestumz@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 169 |  | Gx4 | 838534 | Due for Maintenance | Performed maintenance | Performed maintenance Clean Icore, plunger, and ultrasonic horn Clean filter and cartridge bay | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654757124714150474/5654757124714150474_signature_28.png | Ariston Tumulak, Rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5654757124714150474/5654757124714150474_signature_31.png
| 2023-07-18 | 2023-07-18 15:03 | Bukidnon Provincial Medical Center | Malaybalay City | 10 | Bukidnon | Laborato | YES | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25184 | PBSP | bpmclab.malaybalay@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2024 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 |  | GeneXpert | 804708 | N/A | N/A | Preventive Maintenance Module Replacement Generate Report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654739817221931434/5654739817221931434_signature_28.png | Maureen Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5654739817221931434/5654739817221931434_signature_31.png
| 2023-07-18 | 2023-07-18 14:51 | Velasquez Health Center | Varona St. Tondo | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25183 | PBSP | wcpatron11x4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1074 |  | GXIV | 812561 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck | Unit is operational. Module A1 (684560), service required. Modules A2, A3 and A4 passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654728621526130248/5654728621526130248_signature_28.png | Wilma C. Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5654728621526130248/5654728621526130248_signature_31.png
| 2023-07-18 | 2023-07-18 14:46 | Velasquez Health Center | Varona St. Tondo | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25182 | DOH | wcpatron11x4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1074 | 2023-07-17 | GXIV | 831534 | Unit for PM. For module replacement of module B1. | Performed PM. Performed module replacement of module B1. Generated report. | PM and module replacement. | Unit is operational. |  | GX module B1 (210057495) in replacement to defective B1 (761593). | DR#10143 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654722391528284454/5654722391528284454_signature_28.png | Wilma C. Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5654722391528284454/5654722391528284454_signature_31.png
| 2023-07-18 | 2023-07-18 13:48 | Atang dela Rama Health Center | Gagalangin, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25181 | DOH | monskie_22@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1376 | 2023-07-17 | GXIV | 831644 | Unit for PM. For module replacement of modules B2 and B4. | Performed PM. Performed module replacement of modules B2 and B4. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module B2 (210057654) and B4 (210044177) in replacement to defective B2 (756759) and B4 (756764). | DR#10144 and DR#10147 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654688251523600816/5654688251523600816_signature_28.png | Monaliza T. Balcos | https://www.jotform.com/uploads/macarewecare/222927978372067/5654688251523600816/5654688251523600816_signature_31.png
| 2023-07-18 | 2023-07-18 13:37 | San Fernando RHU | San Fernando, Camarines Sur | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25180 | DOH | josephus_garfin@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1282 |  | GX4 | 838496 | For Xpertcheck | Performed Xpertcheck: Cleaned Cartridge bays, plunger rod, icore Generate reports | Performed Xpertcheck, Preventive Maintenance | Unit is Operational B1 - Service Required B2 - Service Required B3 - Service Required B4 - Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654688026418322175/5654688026418322175_signature_28.png | Josephus Garfin | https://www.jotform.com/uploads/macarewecare/222927978372067/5654688026418322175/5654688026418322175_signature_31.png
| 2023-07-18 | 2023-07-18 13:17 | Lapulapu City Health Office | Lapulapu City, Cebu | 7 | Cebu | Tb Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25179 | DOH | tanya_pax@yahoo.com | 21.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1901 |  | Gx4 | 837661 | Error 1001, 1002 and 4009 | Performed maintenance and calibration in Module 4 | Calibrate module 4 Performed clean up in Icore, Plunger, and filter | All modules are available but due to frequent errors of all modules run time error of 1002. All modules are inoperational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654670572323504347/5654670572323504347_signature_28.png | Maria Lucille P. Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5654670572323504347/5654670572323504347_signature_31.png
| 2023-07-18 | 2023-07-18 12:04 | Passi City Health Office | Passi City, Iloilo | 6 | Iloilo | TB Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25178 | PBSP | jaguilar.passi@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3575 ASSAY: Xpert HIV-1 Viral Load, COUNT: 302 | 2023-07-18 | GXIV-4D | 807610 | Module A3 defective | -Replace defective module A3/663035 with new module SN: 210092868 | #NAME? | Module A2 pending for replacement. Other modules are functional. | Module A4 under observation. | GeneXpert Module 6-Colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654635744225016119/5654635744225016119_signature_28.png | Kaycee Rhee Pormilos | https://www.jotform.com/uploads/macarewecare/222927978372067/5654635744225016119/5654635744225016119_signature_31.png
| 2023-07-18 | 2023-07-18 11:16 | AFPMC-AFRIMS | V. Luna St., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25177 | MACARE-HBDC | mtvalderama@gmail.com | 23 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 66 ASSAY: Xpert MTB/RIF, COUNT: 86 ASSAY: Xpert MTB/RIF Ultra, COUNT: 534 ASSAY: Xpert SARS-CoV-2, COUNT: 2306 ASSAY: Xpert HCV VL, COUNT: 32 ASSAY: Xpert HIV-1 Viral Load, COUNT: 62 ASSAY: Others, COUNT: 30 | 2023-07-17 | GXIV 10C | 845462 | Module B3 for replacement | Perfomed PM Replaced module B3 Generated report | Replaced module B3 899121 -> 210022963 | unit is operational |  | GX Module 10C | G 5760 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654612956325835459/5654612956325835459_signature_28.png | Maria Theresa Vaderama | https://www.jotform.com/uploads/macarewecare/222927978372067/5654612956325835459/5654612956325835459_signature_31.png
| 2023-07-18 | 2023-07-18 11:18 | Mindanao Central Sanitarium | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25176 | DOH | ybioamie@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2045 | 2023-07-18 | GX 4 | 821214 | UPS and System unit unable to start up | For UPS checked the batteries voltage while the system unit reconnect all wires and memory | For UPS checked the batteries voltage while the system unit reconnect all wires and memory | System unit problem solved while the ups, batteries must be replaced and see if the problem persist. | UPS pulled out for battery replacement. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654601731911828431/5654601731911828431_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5654601731911828431/5654601731911828431_signature_31.png
| 2023-07-18 | 2023-07-18 10:49 | MINDANAO CENTRAL SANITARIUM | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25175 | DOH | ybioamie@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1865 | 2023-07-18 | GX 4 | 821214 | System unit and ups unable to start | Checked ups batteries and reconnect/ connect all available cables and memory inside the system unit. | Checked ups batteries and reconnect/ connect all available cables and memory inside the system unit. | System unit was able to start already and functional. While UPS batteries must be replaced to see if the problem persist | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654587301915936830/5654587301915936830_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5654587301915936830/5654587301915936830_signature_31.png
| 2023-07-18 | 2023-07-18 10:33 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | TB DOTS | NO | TRANSFER | CEPHEID | Joseph Acebedo |  |  |  |  | 25174 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 470 ASSAY: Xpert HCV VL, COUNT: 23 ASSAY: Xpert HBV VL, COUNT: 33 ASSAY: Xpert CT/NG, COUNT: 18 ASSAY: Xpert BCR-ABL Ultra, COUNT: 19 |  | GX-IV | 110008524 | FOR TRANSFER OF GX EQUIPMENT AND GX WORKSTATION FROM MOLECULAR LAB TO TB DOTS | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | PERFORM TRANSFER OF GENEXPERT EQUIPMENT, COMPUTER AND PERIPHERALS, MONITOR, UPS, AND AVR. | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654572333013049146/5654572333013049146_signature_28.png | Xavier C. Corraje | https://www.jotform.com/uploads/macarewecare/222927978372067/5654572333013049146/5654572333013049146_signature_31.png
| 2023-07-18 | 2023-07-18 10:05 | Bagong Pag-asa Health Center | Bagong Pag-asa, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25173 | DOH | erlindalorilla026@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1538 ASSAY: Xpert HPV, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1538 ASSAY: Others, COUNT: 79 |  | GXIV 6C | 831597 | Intermittent CPU problem not booting up | Checked CPU Cleaned RAM pins Performed PM Generated report | Troubleshoot CPU | Unit is working temporarily. CPU for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654555053118914473/5654555053118914473_signature_28.png | Erlinda Lorilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5654555053118914473/5654555053118914473_signature_31.png
| 2023-07-18 | 2023-07-18 09:54 | Bulakan Rural Health Unit II | Bulakan, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25172 | DOH | l_gemino@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 897 ASSAY: Xpert HIV-1 Viral Load, COUNT: 65 | 2023-07-03 | GXIV | 831641 | #NAME? | #NAME? | #NAME? | Unit is operational. Pulled out defective monitor. A2 and A3 requires service, for reporting. |  | GeneXpert Monitor SN: MMA1PSS002215079933W01 | 10073 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654550500294220378/5654550500294220378_signature_28.png | Lilibeth Gemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5654550500294220378/5654550500294220378_signature_31.png
| 2023-07-18 | 2023-07-14 09:28 | Philippine Red Cross QC | Malakas St.,Quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | REPAIR | HAEMONETICS | Noel Floresca |  |  |  |  | 25171 | MACARE-HBDC | lambert.cordero@redcross.org.ph | 28 |  | 2023-07-13 | MCS + | 02A227 | For replacement of UPS defective s.n 20051871211 - AT1200210729ST0246 | Replaced and installed UPS, performed functionality | Delivery and Installation | Operational |  | Accupower UPS 1.2KVA | 10107 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654474320175619366/5654474320175619366_signature_28.png | Lambert Cordero | https://www.jotform.com/uploads/macarewecare/222927978372067/5654474320175619366/5654474320175619366_signature_31.png
| 2023-07-18 | 2023-07-14 14:06 | Payatas B Health Center | Bulacan st.,Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION | CEPHEID | Noel Floresca |  |  |  |  | 25170 | DOH | racquelmrocero@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1470 | 2023-07-14 | GX4 | 831899 | Module A3 758253 encounter continous temp. Error 2014,1001,1002 | Performed rexpertcheck cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Rexpertcheck and maintenance | A2 reported already, A1,A4 operational,A3 failed in xpertcheck |  | Xpertcheck kits | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5654473380177346456/5654473380177346456_signature_28.png | Racquel M. Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5654473380177346456/5654473380177346456_signature_31.png
| 2023-07-17 | 2023-07-17 16:46 | Josefina Belmonte Duran Albay Provincial Hospital | Purok-2, Tuburan, Ligao City, Albay | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25169 | PBSP | salvadorlyndonrempillo@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1816 | 2023-07-12 | GX4 | 814371 | Module Replacement | Replaced Module: B1-890331>>210092884, B2-210068727>>210087159, B4-693979>>210092937 Cleaned Cartridge bays, plunger rod, icore Generated reports | For module replacement, preventive maintenance | Unit is operational |  |  | 3module, PL-011826, PL-011688 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653935778813668060/5653935778813668060_signature_28.png | SALVADOR LYNDON P. REMPILLO | https://www.jotform.com/uploads/macarewecare/222927978372067/5653935778813668060/5653935778813668060_signature_31.png
| 2023-07-17 | 2023-07-14 14:23 | Bangkal Health Center | Rodriguez Ave. | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 25168 | DOH | kapid5sepe@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1448 | 2023-07-03 | GX-IV-D | 831875 | For maintenance. Monitor frequent shutting down. | Run XpertCheck | Performed maintenance Cleaned cartridge bay, plunger rod, icore, filter Run XpertCheck Generated reports Checked monitor | All modules operational. All modules passed. |  | XpertCheck-CE-5 | DR# 10079 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849563612135063/5653849563612135063_signature_28.png | Eduardo O. Sepe V, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849563612135063/5653849563612135063_signature_31.png
| 2023-07-17 | 2023-07-14 11:14 | Department of Health Treatment and Rehabilitation Center - Bicutan | Camp Bagong Diwa | NCR | NCR 4th - Taguig | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 25167 | PBSP | dohtrcbicutan.lab@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1919 |  | GX-IV-D | 807452 | For XpertCheck | Performed maintenance Run XpertCheck | Performed maintenance Cleaned cartridge bay, plunger rod, icore. Run XpertCheck Generated reports | Module C3/801912 not operational. Modules C1, C2 and C4 operational. | Module C2/609195 and module C4/666170 service required. Module C3/ cannot calibrate. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849563616561713/5653849563616561713_signature_28.png | April Chakee Javier | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849563616561713/5653849563616561713_signature_31.png
| 2023-07-17 | 2023-07-13 15:18 | Cainta Rural Health Unit | Karangalan Dr. | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25165 | PBSP | rowenatec@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 803 ASSAY: Xpert HIV-1 Viral Load, COUNT: 67 |  | GX-IV-D | 838182 | For maintenance Module B1/807117 error 1002 Module B2/714059 continuous booting | Performed maintenance Checked modules | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Checked modules Generated reports | Modules B1/807117 and B2/714059 not operational, for reporting. Modules B3 ans B4 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473611669535/5653849473611669535_signature_28.png | Rowena Laxamana | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473611669535/5653849473611669535_signature_31.png
| 2023-07-17 | 2023-07-13 15:27 | Cainta Rural Health Unit | Karangalan Dr. | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25165 | PBSP | rowenatec@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1254 |  | GX-IV-D | 804687 | For XpertCheck | Run XpertCheck | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Run XpertCheck Generated reports | Module B1/682692 service required module B3/696028 has error 1002. Modules B2 and B4 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473616545163/5653849473616545163_signature_28.png | Rowena Laxamana | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473616545163/5653849473616545163_signature_31.png
| 2023-07-17 | 2023-07-12 14:54 | The Medical City | Ortigas Avenue, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25165 | COMMERCIAL | moleculardx@themedicalcity.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 3 ASSAY: Others, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 24 ASSAY: Xpert HPV, COUNT: 52 ASSAY: Xpert MTB/RIF, COUNT: 144 ASSAY: Xpert MTB/RIF Ultra, COUNT: 200 ASSAY: Xpert HCV VL, COUNT: 10 ASSAY: Xpert HIV-1 Viral Load, COUNT: 65 |  | GX-IV-L | 802630 | For maintenance. | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter. Generated reports. | Unit is operational. Module B4/210071314 has error 4012. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473616827491/5653849473616827491_signature_28.png | Kristine Parilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473616827491/5653849473616827491_signature_31.png
| 2023-07-17 | 2023-07-13 15:13 | Cainta Rural Health Unit | Karangalan Dr | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25165 | DOH | rowenatec@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 921 ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 |  | GX-IV-D | 831623 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports. | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473617668797/5653849473617668797_signature_28.png | Rowena Laxamana | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473617668797/5653849473617668797_signature_31.png
| 2023-07-17 | 2023-07-12 15:15 | The Medical City | Ortigas Ave, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25166 | COMMERCIAL | moleculardx@themedicalcity.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 131 ASSAY: Xpert HPV, COUNT: 108 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1015 ASSAY: Others, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 22 ASSAY: Xpert HIV-1 Viral Load, COUNT: 198 ASSAY: Xpert SARS-CoV-2, COUNT: 5 |  | GX-XVI-10C | 110015767 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports. LIS Setup | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473617833295/5653849473617833295_signature_28.png | Kristine Parilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5653849473617833295/5653849473617833295_signature_31.png
| 2023-07-17 | 2023-07-11 09:10 | San Lazaro Hospital | Quiricada St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25164 | PBSP | acong.pbsp@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1899 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1117 | 2023-07-10 | GXIV-D | 804763 | Module 1 with Temperature Error | Module Replacement | Replaced Module 1 Performed Maintenance Generate Reports | All Modules are Operational | 778497-> 210092899 | 900-0508 | G5726 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653740073617543530/5653740073617543530_signature_28.png | Ashley Cynna Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5653740073617543530/5653740073617543530_signature_31.png
| 2023-07-17 | 2023-07-11 09:10 | San Lazaro Hospital | Quiricada St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25163 | PBSP | acong.pbsp@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1899 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1117 | 2023-07-10 | GXIV-D | 804763 | Module 1 with Temperature Error | Module Replacement | Replaced Module 1 Performed Maintenance Generate Reports | All Modules are Operational | 778497-> 210092899 | 900-0508 | G5726 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653739543611149447/5653739543611149447_signature_28.png | Ashley Cynna Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5653739543611149447/5653739543611149447_signature_31.png
| 2023-07-17 | 2023-07-12 09:45 | AMOSUP Seamens Hospital | Intramuros, Manila | NCR | NCR 1st - Manila | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25162 | MACARE-HBDC | shmolecularlab@gmail.com | 26 | ASSAY: Xpert SARS-CoV-2, COUNT: 96 | 2023-07-12 | GXIV-D | 838836 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports Performed Self-Test | All Modules are Operational |  | XpertCheck Kit | DR 10099 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653738453614609237/5653738453614609237_signature_28.png | Camelle Marquina | https://www.jotform.com/uploads/macarewecare/222927978372067/5653738453614609237/5653738453614609237_signature_31.png
| 2023-07-17 | 2023-07-12 09:45 | AMOSUP Seamens Hospital | Intramuros, Manila | NCR | NCR 1st - Manila | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25161 | MACARE-HBDC | shmolecularlab@gmail.com | 26 | ASSAY: Xpert SARS-CoV-2, COUNT: 96 | 2023-07-12 | GXIV-D | 838836 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports Performed Self-Test | All Modules are Operational |  | XpertCheck Kit | DR 10099 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653737783611270432/5653737783611270432_signature_28.png | Camelle Marquina | https://www.jotform.com/uploads/macarewecare/222927978372067/5653737783611270432/5653737783611270432_signature_31.png
| 2023-07-17 | 2023-07-12 09:57 | AMOSUP Seamens Hospital | Intramuros, Manila | NCR | NCR 1st - Manila | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25160 | MACARE-HBDC | shmolecularlab@gmail.com | 26 | ASSAY: Xpert SARS-CoV-2, COUNT: 945 | 2023-07-12 | GX16-D | 838777 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Performed Self-Test Generate Reports | 1 Module is Defective, 15 Modules are Operational | D1 Failed | XpertCheck Kits | DR 10099 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653737223616833804/5653737223616833804_signature_28.png | Camelle Marquina | https://www.jotform.com/uploads/macarewecare/222927978372067/5653737223616833804/5653737223616833804_signature_31.png
| 2023-07-17 | 2023-07-12 09:57 | AMOSUP Seamens Hospital | Intramuros, Manila | NCR | NCR 1st - Manila | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25159 | MACARE-HBDC | shmolecularlab@gmail.com | 26 | ASSAY: Xpert SARS-CoV-2, COUNT: 945 | 2023-07-12 | GX16-D | 838777 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Performed Self-Test Generate Reports | 1 Module is Defective, 15 Modules are Operational | D1 Failed | XpertCheck Kits | DR 10099 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653736953611608974/5653736953611608974_signature_28.png | Camelle Marquina | https://www.jotform.com/uploads/macarewecare/222927978372067/5653736953611608974/5653736953611608974_signature_31.png
| 2023-07-17 | 2023-07-17 10:57 | Balasan Rural Health Unit | Balasan, Iloilo | 6 | Iloilo | Laboratory | YES | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 25158 | PBSP | charliemagnedk2@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1573 | 2023-07-17 | GXIV-4D | 807555 | Module B2 is defective | -Replace defective module B2/713796 with new module SN: 210092885 | #NAME? | All modules are functional. |  | GeneXpert Module 6-colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653724279402750309/5653724279402750309_signature_28.png | Charlie Magne Dequito | https://www.jotform.com/uploads/macarewecare/222927978372067/5653724279402750309/5653724279402750309_signature_31.png
| 2023-07-17 | 2023-07-12 13:14 | National Tuberculosis for Reference Laboratory | Filinvest | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25157 | PBSP | richarddelossantos10@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1969 |  | GXIV-D | 707512 | Modules have High Temperature and Temperature Error | Troubeshoot | Checked Fan if Working Checked Ambient Temperature Generate Reports | Fan is Not Working. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653673873611198157/5653673873611198157_signature_28.png | Richard Delos Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5653673873611198157/5653673873611198157_signature_31.png
| 2023-07-17 | 2023-07-13 09:38 | Imus City Health Office II | Plaridel Subd, Imus | 4A | Cavite | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25156 | DOH | amielynnmangulabnan1217@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1146 |  | GXIV-D | 821284 | UPS Problem | Troubleshoot | Checked Outlet Voltage Checked Voltage Output of the UPS Checked if Back-Up Function is Operational Generate Reports | Defective UPS. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653673873612524216/5653673873612524216_signature_28.png | Amielynn Mangulabnan | https://www.jotform.com/uploads/macarewecare/222927978372067/5653673873612524216/5653673873612524216_signature_31.png
| 2023-07-17 | 2023-07-17 09:05 | Danao City Health Office | Danao City, Cebu | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25155 | PBSP | chockenriquez@gmail.com | 23 |  |  | Gx 4 | 808286 | Error 1001, 2014 module 1 & due date for xpertcheck | Running xpertcheck, clean fan filter | Running xpertcheck, clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653662767378145190/5653662767378145190_signature_28.png | Nicanor Enriquez Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5653662767378145190/5653662767378145190_signature_31.png
| 2023-07-16 | 2023-07-11 11:35 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 25154 | DOH | jan19.cam@gmail.com | 27.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1436 |  | Gx-IV | 838504 | Module 1 H/W failed due multiple temperature error | Perform Xpertcheck at module 1 for troubleshooting | Perform Genexpert machine Preventive maintenance | Module 1 under observation due to temperature error | Module 1 passed the Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653308374429362921/5653308374429362921_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5653308374429362921/5653308374429362921_signature_31.png
| 2023-07-16 | 2023-07-14 10:21 | Davao Oriental Provincial Hospital - Cateel | National Highway, Cateel Davao Oriental | 11 | Davao Oriental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25153 | PBSP | mykadoysabas8@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 648 |  | GX-IV 4d | 807432 | For Preventive Maintenance and Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Perform XPERTCHECK Calibration | Perform Preventive Maintenance and XpertCheck Calibration | Module B1 for service Required B2, B3 and B4 Passed the Calibration | Module B1 for under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5653243395619403117/5653243395619403117_signature_28.png | Carol Myka L. Doysabas | https://www.jotform.com/uploads/macarewecare/222927978372067/5653243395619403117/5653243395619403117_signature_31.png
| 2023-07-14 | 2023-07-14 13:59 | Dimasalang Health Center | Tiago St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25152 | PBSP | kikokendz_060830@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1330 | 2023-07-11 | GXIV 6C | 805111 | Module B2 for replacement Unit for XpertCheck | Performed PM Replaced module B2 Performed XpertCheck Generated report | Replaced B2: 210008903 -> 210092905 Performed XpertCheck on B1, B3, and B4 | Unit is operational B1: Passed B2: New module B3: Service Required B4: Service Required |  | XpertChecknkit | Lot 1000725070 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5651255842749267334/5651255842749267334_signature_28.png | Majoy abenales | https://www.jotform.com/uploads/macarewecare/222927978372067/5651255842749267334/5651255842749267334_signature_31.png
| 2023-07-14 | 2023-07-14 12:29 | Badian RHU | Badian, Cebu | 7 | Cebu | TB Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 25151 | PBSP | avenzr@gmail.com | 32 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 674 |  | Gx4 | 807588 | Module 3 SN 662527 and Module 4 SN 662525 | Replaced defective modules with new module Generate data such as Gxx files, IQR, Sys log, Assay statistics Clean filter and icore, plunger | Replaced new module 3 witj SN 210093414 and module 4 with SN 210093378 | All modules are operational but recommended to install aircon in the aircon for genxepert alone as soon as possible due high temperature |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5651214810424893401/5651214810424893401_signature_28.png | Avenizer Absin, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5651214810424893401/5651214810424893401_signature_31.png
| 2023-07-14 | 2023-07-14 12:16 | Lagonoy Rural Health Unit | Lagonoy | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25150 | WHO | barbacenam128@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1464 |  | GX4 | 846355 | For Preventive Maintenance | Cleaned Cartridge bays, plunger rod, icore Generate Reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5651211997603615732/5651211997603615732_signature_28.png | Marietta Barbacena | https://www.jotform.com/uploads/macarewecare/222927978372067/5651211997603615732/5651211997603615732_signature_31.png
| 2023-07-14 | 2023-07-14 12:24 | Tigaon Municipal Health Office | Caraycayon, Tigaon | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25149 | DOH | maritesjudar@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1242 |  | GX4 | 838558 | For Xpertcheck | Performed Xpertcheck Cleaned Cartridge bays, plunger rod, icore Generate report | Performed Xpertcheck, Preventive Maintenance | Unit is operational B1 - Passed B2 - Service Required B3 - Passed B4 - Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5651209687606780607/5651209687606780607_signature_28.png | Marites Judar | https://www.jotform.com/uploads/macarewecare/222927978372067/5651209687606780607/5651209687606780607_signature_31.png
| 2023-07-14 | 2023-07-14 12:16 | Talavera Rural Health Unit II | San Ricardo, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25148 | PBSP | nelzon1966@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 606 | 2023-07-14 | Genexpert IV | 812831 | UPS not Working | Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | UPS for Replacement due to no Backup | Genexpert Module = 210092906 & 210092910 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5651181113992045889/5651181113992045889_signature_28.png | Nelson Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5651181113992045889/5651181113992045889_signature_31.png
| 2023-07-14 | 2023-07-14 08:48 | Region 1 Medical Center | Dagupan City, Pangasinan | 1 | Pangasinan | Molecular Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25147 | PBSP | efrenotariojr@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1547 ASSAY: Xpert SARS-CoV-2, COUNT: 2839 |  | GX 16 | 842230 | Module B1 High Temperature | - Let GX temperature drop to 30° C - Replace all filters - Generated Reports | Performed Troubleshooting | Unit is operational | Room Temperature too high for GeneXpert |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5651178118917966904/5651178118917966904_signature_28.png | Efren Notario Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5651178118917966904/5651178118917966904_signature_31.png
| 2023-07-14 | 2023-07-14 11:25 | Lubao Rural Health Unit | Lubao Pampanga | 3 | Pampanga | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25146 | PBSP | nathaliecoleencovar@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 928 |  | GXIV-D | 812834 | UPS not backing up | Troubleshooting | Performed UPS troubleshooting Gathered reports Performed functionality test | Unit is operational except UPS not backing up | UPS for reporting, for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5651148272511724171/5651148272511724171_signature_28.png | Nathalie Covar | https://www.jotform.com/uploads/macarewecare/222927978372067/5651148272511724171/5651148272511724171_signature_31.png
| 2023-07-14 | 2023-07-14 09:17 | Bagac Rural Health Unit | A. Bonifacio St. Brgy. Ibaba Bagac, Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25145 | PBSP | jonathanricalde_17@yahoo.com.ph | 30 | ASSAY: Xpert MTB/RIF, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 374 | 2023-07-06 | GXII-D | 835173 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | 2 modules passed calibration | 2 xpertcheck kits | 41818 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5651075382514403379/5651075382514403379_signature_28.png | Jonathan Ricalde | https://www.jotform.com/uploads/macarewecare/222927978372067/5651075382514403379/5651075382514403379_signature_31.png
| 2023-07-14 | 2023-07-14 07:00 | San Jose Municipal Health Unit | San Jose, Camarines Sur | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25144 | DOH | jhuneric2008@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1745 |  | GX4 | 838389 | For Xpertcheck | Performed Xpertcheck Cleaned Cartridge Bays, Plunger Rod, Icore Generate Reports | Performed Xpertcheck, Preventive Maintenance | Unit is operational B1 - service required B2 - Passed B3 - module unavailable B4 - Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5651009160983339466/5651009160983339466_signature_28.png | Rosmar Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5651009160983339466/5651009160983339466_signature_31.png
| 2023-07-13 | 2023-07-13 16:30 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25143 | PBSP | bmcgenexpertlab@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2787 | 2023-07-13 | GXIV | 834623 | Module C2 for replacement | #NAME? | Replaced C2 | Unit is operational. |  | GeneXpert Module SN: 210092873 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650465731011895016/5650465731011895016_signature_28.png | Michaella Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5650465731011895016/5650465731011895016_signature_31.png
| 2023-07-13 | 2023-07-13 14:32 | Cristino M. Paragas Memorial Hospital | Quiniput, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25142 | PBSP | fatimaabegailabdurahman@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 174 | 2023-07-13 | GX 4 | 815269 | None | All 4 modules undergone Xpertcheck Calibration | All 4 modules undergone Xpertcheck Calibration | All modules passed and available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650456299102443295/5650456299102443295_signature_28.png | Fatima Abegail Abdurahman | https://www.jotform.com/uploads/macarewecare/222927978372067/5650456299102443295/5650456299102443295_signature_31.png
| 2023-07-13 | 2023-07-13 14:53 | M. Icasiano Health Center | Pedro Gil St., Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25141 | DOH | lorrainefabros@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1264 |  | GXIV 6C | 831684 | CPU not booting up "OSD Lockout" stays on the monitor ; OK button broken | Tightened power cable of CPU; problem persists Test HDD; still works Replaced RAM; problem solved; CPU boots up Tightened monitor cables; problem persists Changed VGA cable; problem persists Generated report | Troubleshoot CPU Troubleshoot monitor | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650414754814618650/5650414754814618650_signature_28.png | Lorraine Lucila L Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5650414754814618650/5650414754814618650_signature_31.png
| 2023-07-13 | 2023-07-13 14:16 | Nasipit Rural Health Unit | Brgy. Talisay, Nasipit | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 25140 | DOH | peterdolotina@gmail.com | 25.63 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1604 |  | GX-4 | 837611 | Module B2 found Service Required thru XPC | None | Calibrating Module B2 Preventive Maintenance | 3 modules working 1 module found service required (for replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650385149025094153/5650385149025094153_signature_28.png | Peter C. Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/5650385149025094153/5650385149025094153_signature_31.png
| 2023-07-13 | 2023-07-13 13:56 | Dalaguete RHU | Poblacion Dalaguete Cebu | 7 | Cebu | Tb Lab | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 25139 | PBSP | ajreef5517@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1039 |  | Gx4 | 815464 | B2 SN 884841 Hardware failure due to temperature related problem | Replaced new module for B2 | Replaced new module SN 210092896 | All units modules are available except module 4 encounter an error of 2032 ultrasonic horn improper tuning |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650377455515933017/5650377455515933017_signature_28.png | Allaiza Jane Rivera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5650377455515933017/5650377455515933017_signature_31.png
| 2023-07-13 | 2023-07-06 09:46 | Urdaneta District Hospital | Urdaneta, Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25138 | PBSP | warlyacosta@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 35 ASSAY: , COUNT: 2259 |  | GX 4 | 805109 | For Calibration | #NAME? | Performed Calibration | Unit is operational | All modules pass |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650364898911405556/5650364898911405556_signature_28.png | Warlito Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5650364898911405556/5650364898911405556_signature_31.png
| 2023-07-13 | 2023-07-06 11:23 | Urdaneta Sacred Heart Hospital | Urdaneta, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25137 | COMMERCIAL | annabellelcatbagan@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 184 |  | GX 2 | 820303 | For Preventive Maintenance | #NAME? | Performed preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650363438919538903/5650363438919538903_signature_28.png | ANNABELLE CATBAGAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5650363438919538903/5650363438919538903_signature_31.png
| 2023-07-13 | 2023-07-07 08:08 | Basista RHU | Basista, Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25136 | PBSP | coraflores14@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 99 ASSAY: Xpert MTB/RIF Ultra, COUNT: 141 |  | GX 4 | 838860 | For Maintenance | #NAME? | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650361688916926156/5650361688916926156_signature_28.png | Cora s. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5650361688916926156/5650361688916926156_signature_31.png
| 2023-07-13 | 2023-07-12 09:52 | Candon City Health Office | Candon, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25134 | DOH | candoncholaboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Xpert BCR-ABL Ultra, COUNT: 460 | 2023-06-13 | GX4 | 838273 | Due for Calibration | #NAME? | Performed Calibration | Unit is operational | Monitor for reporting | XpertCheck Kits | DR 5344 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650358578912165396/5650358578912165396_signature_28.png | Michelle Bucalen | https://www.jotform.com/uploads/macarewecare/222927978372067/5650358578912165396/5650358578912165396_signature_31.png
| 2023-07-13 | 2023-07-12 11:47 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | La Union | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25135 | DOH | piaramuj@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 595 | 2023-06-13 | GX4 | 838270 | Due for Calibration | #NAME? | Performed Calibration | Unit is Operational | Modules B2 and B3 failed XpertCheck for reporting | XpertCheck Kits | DR 5342 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650358578914592833/5650358578914592833_signature_28.png | Pia Jumar T. Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/5650358578914592833/5650358578914592833_signature_31.png
| 2023-07-13 | 2023-07-11 08:31 | Bacarra RHU | Bacarra, Ilocos Norte | 1 | Ilocos Norte | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25129 | PBSP | davidangelogamidobarrozo@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 36 |  | GX 4 | 807430 | Due for Calibration | #NAME? | Performed Calibration | Unit is operational | Module B4 failed for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868914781891/5650357868914781891_signature_28.png | David Angelo Barrozo | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868914781891/5650357868914781891_signature_31.png
| 2023-07-13 | 2023-07-11 11:58 | Mariano Marcos Memorial Hospital and Medical Center | Batac, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25130 | DOH | mmmhmc.molecularlab@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 392 ASSAY: Xpert MTB/RIF Ultra, COUNT: 822 ASSAY: Xpert SARS-CoV-2, COUNT: 2992 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 | 2023-06-13 | GX 4 | 831908 | Due for Calibration | #NAME? | Performed Calibration | Unit is operational | Module A2 and A3 failed XpertCheck for reporting | XpertCheck Kits | DR 5339 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868915296235/5650357868915296235_signature_28.png | Allan Fellizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868915296235/5650357868915296235_signature_31.png
| 2023-07-13 | 2023-07-11 13:22 | Mariano Marcos Memorial Hospital and Medical Center | Batac, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25131 | PBSP | mmmhmc.molecularlab@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 200 ASSAY: Xpert MTB/RIF Ultra, COUNT: 512 ASSAY: Xpert SARS-CoV-2, COUNT: 2396 |  | GX4 | 838168 | For Troubleshooting, Prone to Lost Communication For Preventive Maintenance | #NAME? | Performed Troubleshooting and Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868915982229/5650357868915982229_signature_28.png | ALLAN FELLIZAR | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868915982229/5650357868915982229_signature_31.png
| 2023-07-13 | 2023-07-12 06:23 | Sinait District Hospital | Sinait, Ilocos Sur | 1 | Ilocos Sur | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25132 | DOH | isdhsinaitlaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 509 | 2023-06-13 | GX4 | 838277 | Due for Calibration | #NAME? | Performed Calibration | Unit is operational |  | XpertCheck Kits | DR 5343 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868916125173/5650357868916125173_signature_28.png | Aileen Sambrana | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868916125173/5650357868916125173_signature_31.png
| 2023-07-13 | 2023-07-11 10:06 | Laoag City Health Office | Laoag, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25133 | DOH | denver.mahor@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 389 | 2023-06-13 | GX4 | 838416 | Due for Calibration | #NAME? | Performed Calibration | Unit is operational |  | XpertCheck Kits | DR 541 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868919043820/5650357868919043820_signature_28.png | Rodenver Mahor | https://www.jotform.com/uploads/macarewecare/222927978372067/5650357868919043820/5650357868919043820_signature_31.png
| 2023-07-13 | 2023-07-13 12:57 | Pili Rural Health Unit | San Isidro, Pili, Camarines Sur | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25128 | DOH | lfsanao724@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1233 |  | GX4 | 838408 | For Xpertcheck | Performed Xpertcheck Cleaned Cartridge bays, plunger rod, icore Generate reports | Performed Xpertcheck, Preventive Maintenance | Unit is operational B1-passed B2-passed B3-service required (already reported) B4-passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650343766028029655/5650343766028029655_signature_28.png | Liezl Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/5650343766028029655/5650343766028029655_signature_31.png
| 2023-07-13 | 2023-07-13 11:41 | Dr. Elvira Lagrosa Health Center | R. Layug St., Brgy. 3, Pasay | NCR | NCR 4th - Pasay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25127 | PBSP | LBCANICULA@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2315 | 2023-07-06 | GX IV 6C | 805106 | Module 4 Insufficient kit last week New xpertchrck kit arrived Module 4 for xpertcheck | Performed PM Performed Xpertcheck on module 4 Module 4 (604721) failed xpertcheck Generated report | Performed XpertCheck on module 4 | Unut is operational |  | XpertCheck kit | DR #0041849 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650301220518694456/5650301220518694456_signature_28.png | Lorna Canicula | https://www.jotform.com/uploads/macarewecare/222927978372067/5650301220518694456/5650301220518694456_signature_31.png
| 2023-07-13 | 2023-07-13 10:11 | FAMILY PLANNING ORGANIZATION OF THE PHILIPPINES - ILOILO CITY | RIZAL ST. ILOILO CITY | 6 | Iloilo | LABORATORY | NO | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 25126 | PBSP | curdi1444@gmail.com | 25 |  |  | GX-IV | 120000143 | NEW INSTALLATION | COLLECT IQR | PERFORM NEW GENEXPERT INSTALLATION MERGE GX SN 120000143 AND 838869 USING NETWORK SWITH | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650248331127903738/5650248331127903738_signature_28.png | Yasmin P. Sabando | https://www.jotform.com/uploads/macarewecare/222927978372067/5650248331127903738/5650248331127903738_signature_31.png
| 2023-07-13 | 2023-07-13 09:39 | Giporlos Rural Health Unit | Giporlos | 8 | Eastern Samar | TB dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25125 | PBSP | ningrediang@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 57 ASSAY: Xpert MTB/RIF Ultra, COUNT: 433 | 2023-07-13 | GXIV - 2D | 835060 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650219675317832189/5650219675317832189_signature_28.png | Luningning R. Abuda | https://www.jotform.com/uploads/macarewecare/222927978372067/5650219675317832189/5650219675317832189_signature_31.png
| 2023-07-13 | 2023-07-13 09:15 | San Jose Public Health and Diagnostic Center | Poblacion 7, San Jose | 4B | Occidental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25124 | PBSP | naty_garcia1962@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1051 | 2023-07-10 | GXIV | 838197 | Unit for PM. For module replacement of module B4. | Performed PM. Performed module replacement on B4. Generated reports. | PM and module replacement. | Unit is operational |  | GX Module B4 (210092869) in replacement to defectivr B4 (807350). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5650204331466523253/5650204331466523253_signature_28.png | Natividad Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5650204331466523253/5650204331466523253_signature_31.png
| 2023-07-12 | 2023-07-12 15:08 | Zamboanga City Medical Centre | Dr. Evangelista Street, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25123 | PBSP | zcmctblab@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HPV, COUNT: 1212 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4968 ASSAY: Xpert SARS-CoV-2, COUNT: 14831 ASSAY: Xpert HIV-1 Viral Load, COUNT: 305 | 2023-07-12 | GX 4 | 839309 | None | Replaced module 210068702 with Module 210092892. And Preventive Maintenence | Replaced module 210068702 with Module 210092892. And Preventive Maintenence | Some modules are for replacement and some are available. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649603620275774221/5649603620275774221_signature_28.png | Aira Faith S. Runas | https://www.jotform.com/uploads/macarewecare/222927978372067/5649603620275774221/5649603620275774221_signature_31.png
| 2023-07-12 | 2023-07-12 15:11 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25122 | PBSP | laureanjo@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 976 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | Gx 4 | 838882 | Defective module 4 sn 210068078 | Clean fan filter | New module 4 installed sn 210085582 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649580838151308022/5649580838151308022_signature_28.png | Laurean Jo Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5649580838151308022/5649580838151308022_signature_31.png
| 2023-07-12 | 2023-07-12 15:41 | St Lukes Extension Clinic | J. Bocobo St., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Patrick Los Baños |  |  |  |  | 25121 | MACARE-HBDC | laboratory@slec.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 4696 ASSAY: Xpert MTB/RIF, COUNT: 317 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5262 |  | Infinity 48s | 809643 | Software stuck on waste container emptying screen | #NAME? | Performed basic troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649573866877623790/5649573866877623790_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5649573866877623790/5649573866877623790_signature_31.png
| 2023-07-12 | 2023-07-12 15:27 | Carcar CHO | Valladolid, Carcar City, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25120 | DOH | chaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1460 |  | Gx4 | 838283 | Error 2037 cartridge integrity test failed in B4 | Instruct medtech to check the cartridge before loading specimen Follow instruction for specimen handling And preparation | Clean Icore, Plunger maintenance and clean cartridge bay Clean filter Generate IQR, Sys log, Gxx files and Assay Stats | B1, B3, and B4 are operational except B2 due to frequent 1002 errors |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649566316429787436/5649566316429787436_signature_28.png | Lucille Mae A. Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5649566316429787436/5649566316429787436_signature_31.png
| 2023-07-12 | 2023-07-12 11:21 | Rosario Super Health Center | Dr. Sixto Antonio Ave., Rosario, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25119 | WHO | eden.cawayan@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1934 |  | GXIV 6C | 843924 | GeneXpert to transfer to better room | Moved and setup GX and all accessories to new room Performed PM Generated report | Transferred GX to new room | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649536693613890306/5649536693613890306_signature_28.png | Eden R. Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5649536693613890306/5649536693613890306_signature_31.png
| 2023-07-12 | 2023-07-12 13:16 | Palayan City Health Office | Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25118 | DOH | riazarizzamae@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 989 | 2023-07-12 | Genexpert IV | 831637 | N/a | N/a | Performed Maintenance Gathered Reports Installed HIV assay Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649512775417019565/5649512775417019565_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5649512775417019565/5649512775417019565_signature_31.png
| 2023-07-12 | 2023-07-12 13:26 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 25117 | DOH | ccmclaboratory01@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1398 |  | GXIV-4D | 838561 | Module A2 frequent termination error. | None | #NAME? | Waiting for module A1 allocation. All three other modules are operational. | Module A2 for observation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649491089615585339/5649491089615585339_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5649491089615585339/5649491089615585339_signature_31.png
| 2023-07-12 | 2023-07-12 12:49 | Atang dela Rama Health Center | Gagalangin, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25116 | DOH | monskie_22@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1374 |  | GXIV | 831644 | Unit for PM. AVR malfunctioning. | Performed PM. Performed troubleshooting of AVR. Generated reports. Generated photo and video of AVR malfunctioning. | Preventive maintenance and troubleshooting. | Unit is operational. AVR is defective and for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649474966626866507/5649474966626866507_signature_28.png | Monaliza Taruc | https://www.jotform.com/uploads/macarewecare/222927978372067/5649474966626866507/5649474966626866507_signature_31.png
| 2023-07-12 | 2023-07-12 12:42 | Tomas Claudio Health Center | Corner E. Quirino, Camus Street Davao city | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25115 | DOH | arleen_p_alconera@yahoo.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 191 ASSAY: Xpert MTB/RIF Ultra, COUNT: 520 |  | GX-IV 4d | 838492 | Modules B2 and B4 reported Module Error | Clean Optics | Troubleshoot the Modules | Modules reported is under observation. | Errors Encounter during Running are Code 5007 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649467896223484582/5649467896223484582_signature_28.png | Arleen Alconera RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5649467896223484582/5649467896223484582_signature_31.png
| 2023-07-12 | 2023-07-12 12:26 | Lapulapu cho | Lapulapu City, Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25114 | DOH | tanya_pax@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 466 |  | Gx4 | 837661 | Error 1002 and 4009 SN # 753535 | Reset module 4 | Reset module 4 Generate IQR, Sys log and assay stat | All modules are available but due to frequent Error 1002 and 4009 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649463632672094729/5649463632672094729_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5649463632672094729/5649463632672094729_signature_31.png
| 2023-07-12 | 2023-07-12 10:00 | Talisay CHU | Cansojong Talisay | 7 | Cebu | Tb Lab | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25113 | PBSP | marydelgella@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1400 |  | Gx4 | 815384 | Error 2014 in B2 | Clean filter, checking ambient temp in the area, schedule reset after running all specimen | Generate IQR, Sys log, IQR | Only 3 modules are running due terminnation error of module 2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649369386817613056/5649369386817613056_signature_28.png | Marydel Gella, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5649369386817613056/5649369386817613056_signature_31.png
| 2023-07-12 | 2023-01-18 14:57 | Cagayan Valley Medical Center | Tuguegarao City, Cagayan | 3 | Cagayan | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Matthew Roxas |  |  |  |  | 25112 | DOH | loleniamelad@ymail.com | 28 |  | 2023-01-18 | MCS+ | ISH0336 | N/a | Replaced Rollers on red Adjust the line sensor but not working Clean the machine | Performed Maintenance | Unit is Operational | Line sensor for Replacement due to low output | Rollers 2pcs | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5502686552261284914/5502686552261284914_signature_28.png | Lolenia D. Melaf | https://www.jotform.com/uploads/macarewecare/222927978372067/5502686552261284914/5502686552261284914_signature_31.png
| 2023-07-12 | 2023-07-12 09:41 | BALASAN RURAL HEALTH UNIT | BALASAN ILOILO | 6 | Iloilo | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 25111 | PBSP | charliemagnedk2@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1564 |  | GX-IV | 807555 | MODULE B3 STUCK CARTRIDGE | COLLECT IQR , STYSTEM LOG AND ASSAY STAT | #NAME? | STUCK CARTRIDGE ISSUE RESOLVED MOUDLE B2 H/W FAIL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649357936423722852/5649357936423722852_signature_28.png | Charlie Magne Dequito | https://www.jotform.com/uploads/macarewecare/222927978372067/5649357936423722852/5649357936423722852_signature_31.png
| 2023-07-12 | 2023-07-12 08:46 | Oras District Hospital | Oras E. Samar | 8 | Eastern Samar | Tb dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25110 | PBSP | espesopriscilla@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 45 ASSAY: Xpert MTB/RIF Ultra, COUNT: 792 | 2023-07-12 | GXIV - 2D | 835061 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649325679718184491/5649325679718184491_signature_28.png | Priscilla R. Espeso | https://www.jotform.com/uploads/macarewecare/222927978372067/5649325679718184491/5649325679718184491_signature_31.png
| 2023-07-12 | 2023-07-11 13:47 | Vitali lying-in Centre | Vitali Zamboanga City | 9 | Zamboanga del Sur | Labo | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25109 | PBSP | erianejanecardenas@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 | 2023-07-11 | GX 4 | 838140 | None | Xpertcheck calibration for all 4 modules and conducted preventive Maintenence | Xpertcheck calibration for all 4 modules and conducted preventive Maintenence | All modules are available. All modules passed the xpertcheck calibration. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649148188316861995/5649148188316861995_signature_28.png | Eriane Jane M. Cardenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5649148188316861995/5649148188316861995_signature_31.png
| 2023-07-12 | 2023-07-11 13:47 | Vitali lying-in Centre | Vitali Zamboanga City | 9 | Zamboanga del Sur | Labo | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25108 | PBSP | erianejanecardenas@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 | 2023-07-11 | GX 4 | 838140 | None | Xpertcheck calibration for all 4 modules and conducted preventive Maintenence | Xpertcheck calibration for all 4 modules and conducted preventive Maintenence | All modules are available. All modules passed the xpertcheck calibration. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5649148118311062160/5649148118311062160_signature_28.png | Eriane Jane M. Cardenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5649148118311062160/5649148118311062160_signature_31.png
| 2023-07-11 | 2023-07-11 14:08 | Arsenia Maximo Super Health Center | Quirino Hi-way, Novaliches, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25107 | PBSP | elyvette_macalalag@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 155 ASSAY: Xpert MTB/RIF, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1093 ASSAY: Xpert HIV-1 Viral Load, COUNT: 246 | 2023-07-07 | GXIV 6C | 814162 | Unit for XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. All modules passed. |  | XpertCheck Kit | DR #0041843 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5648674939479151602/5648674939479151602_signature_28.png | Elyvette Basa | https://www.jotform.com/uploads/macarewecare/222927978372067/5648674939479151602/5648674939479151602_signature_31.png
| 2023-07-11 | 2023-07-11 14:40 | Toledo CHO | Ilihan Toledo City, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25106 | PBSP | loriemaediapen@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 672 |  | Gx4 | 808640 | Due for Maintenance, | Performed maintenance | Performed maintenance, Back up database Clean Icore, Clean filter, plunger maintenance Generate IQR, Sys log, Gxx files and Assay Stat | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5648671086326304864/5648671086326304864_signature_28.png | Lorie Mae C. Diapen, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5648671086326304864/5648671086326304864_signature_31.png
| 2023-07-11 | 2023-07-11 14:32 | Bongabon District Hospital | Curva, Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25105 | DOH | tsantuajr0128@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 398 | 2023-07-11 | Genexpert IV | 831562 | Module 3 error 2008 | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 3 for observation | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5648666228118744696/5648666228118744696_signature_28.png | Teodorico Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5648666228118744696/5648666228118744696_signature_31.png
| 2023-07-11 | 2023-07-11 13:32 | Juan Posadas Health Center | Rodriguez St. Tondo | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25104 | PBSP | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1293 |  | GXIV | 805060 | CPU slow response and slow to archive. | Performed troubleshooting of CPU. Resetted ram, hdd and internal connections. Generated reports, photo and video. | Troubleshooting | Unit is operational. CPU is slow and recommended for replacement. | HP Z220SFF (sn: 2UA3520VNV) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5648634955429326313/5648634955429326313_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/5648634955429326313/5648634955429326313_signature_31.png
| 2023-07-11 | 2023-07-11 13:23 | Sta Cruz Rural Health Unit | Sta Cruz, Davao Del Sur | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25103 | DOH | chendalagan72@gmail.com | 25.6 |  |  | GX-IV 4D | 838505 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5648627546225672337/5648627546225672337_signature_28.png | Dorothy A. Dalagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5648627546225672337/5648627546225672337_signature_31.png
| 2023-07-11 | 2023-07-11 13:18 | SULAT RURAL HEALTH UNIT | SULAT | 8 | Eastern Samar | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25102 | PBSP | katherinelupos@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 85 ASSAY: Xpert MTB/RIF Ultra, COUNT: 650 | 2023-07-11 | GXIV | 835100 | Module B2 - stuck Cartridge Module B1 - received multiple temperature error | #NAME? | #NAME? | Unit is Operational ( module B1 waiting for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5648626026518424689/5648626026518424689_signature_28.png | Katherine Ann L. Alday | https://www.jotform.com/uploads/macarewecare/222927978372067/5648626026518424689/5648626026518424689_signature_31.png
| 2023-07-11 | 2023-07-11 10:21 | Hi-Precision Diagnostics | 674 N.S. Amoranto Sr. St, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25101 | CONSORTIUM | jose@hi-precision.com.ph | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 24 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Others, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4977 |  | GX IV 6 Colors | 835602 | 2 Mod Modules A1 and A3 for replacement | #NAME? | Replaced Modules A1 and A3 A1: 677953 -> 633074 A3: 671739 -> 904172 | Unit is operational. All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5648528682116641280/5648528682116641280_signature_28.png | Joselito T. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5648528682116641280/5648528682116641280_signature_31.png
| 2023-07-11 | 2023-07-06 16:35 | Caloocan Health Center | Sangandaan, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25100 | PBSP | caloocancentrallaboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3112 ASSAY: Xpert HIV-1 Viral Load, COUNT: 529 |  | GXIV 6C | 805271, 814154 | CPU 2 for troubleshooting | Tightened CPU parts connections CPU working properly Separated 2 GX Removed network switch Setup PC 1 and PC 2 Module B1 ERROR 2126; Tightened connections Generated report | Troubleshoot CPU2 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5648509812112295476/5648509812112295476_signature_28.png | ALBERT IGAMA | https://www.jotform.com/uploads/macarewecare/222927978372067/5648509812112295476/5648509812112295476_signature_31.png
| 2023-07-10 | 2023-07-10 16:27 | Tabaco CHU | Tabaco City | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25099 | PBSP | elenocbina@gmail.com | 28 |  | 2023-07-07 | GX4 | 804257 | For module replacement, D2 | Replaced Module: 617185>>210089677 Clean cartridge bays, plunger rod, fan filter | Module Replacement, Preventive Maintenance | Unit is operational |  | Module 210089677 | PL-011626 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5647873436816576511/5647873436816576511_signature_28.png | Eleanor Ocbina | https://www.jotform.com/uploads/macarewecare/222927978372067/5647873436816576511/5647873436816576511_signature_31.png
| 2023-07-10 | 2023-07-10 13:45 | Pilar Rural Health Unit | Pilar Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25098 | DOH | aimna_24@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 356 |  | GXIV-D | 831613 | Due for re-xpertcheck | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Performed calibration Gathered reports Performed functionality test | Unit is operation except module A3 | A3 failed xpertcheck for reporting, used 1 calibration kit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5647774167129105738/5647774167129105738_signature_28.png | Marciana Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5647774167129105738/5647774167129105738_signature_31.png
| 2023-07-10 | 2023-07-04 10:12 | Dasmariñas City Health Office II | DBB-E, Dasmariñas | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25096 | DOH | hazelmilan@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2422 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 |  | GXIV-D | 838535 | Stuck Cartridge in Module 4 | Self-Test | Removed Stuck Cartridge Performed Maintenance Generate Reports | 1 Modules is Operational, 3 Modules are Defective |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5647769203616385095/5647769203616385095_signature_28.png | Hazel Tabucol | https://www.jotform.com/uploads/macarewecare/222927978372067/5647769203616385095/5647769203616385095_signature_31.png
| 2023-07-10 | 2023-07-07 09:49 | Amadeo Rural Health Unit | Pangil Road, Amadeo | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25096 | PBSP | ginabautista789@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 876 ASSAY: Xpert HIV-1 Viral Load, COUNT: 76 |  | GXIV-D | 838198 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | All Modules are Operational and Passed the XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5647769203618359023/5647769203618359023_signature_28.png | Gina Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5647769203618359023/5647769203618359023_signature_31.png
| 2023-07-10 | 2023-07-05 15:13 | RAKKK Prophet Medical Center | Maharlika Highway, Gumaca | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25097 | PBSP | rpmcimoleclab20@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 248 ASSAY: Xpert SARS-CoV-2, COUNT: 493 |  | GXIV-D | 808268 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | All Modules are Operational and Passed the XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5647769203619624591/5647769203619624591_signature_28.png | Katrina Rose F. Aureaca | https://www.jotform.com/uploads/macarewecare/222927978372067/5647769203619624591/5647769203619624591_signature_31.png
| 2023-07-10 | 2023-07-10 11:34 | Bataan Genaeral Hospital and Medical Center | Balanga Bataan | 3 | Bataan | RHIVDA | NO | Genexpert machine relocation | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25095 | MACARE-HBDC | honeyeichel@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 19 ASSAY: Xpert HCV VL, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 39 |  | GXIV-D | 110017894 | Machine relocation | Machine relocation Installation and setup | Performed machine relocation from RHIVDA to PCR Laboratory Performed machine setup Performed preventive maintenance Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5647764597125602748/5647764597125602748_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5647764597125602748/5647764597125602748_signature_31.png
| 2023-07-10 | 2023-07-10 11:16 | Cebu South Medical Center | San Isidro Talisay City, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25094 | DOH | josephlorenzomercado@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 653 ASSAY: Xpert HIV-1 Viral Load, COUNT: 32 |  | Gx4 | 838559 | Error 2008 and 5007 occured rarely | Errors mostly are sample related problems. | Check ups and CPU, printer Check genxpert machine, bar code scanner, Clean Icore, plunger, and filter Generate database, gxx files, IQR and assay Stats | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5647686267031267611/5647686267031267611_signature_28.png | Josehp Lorenzo V. Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/5647686267031267611/5647686267031267611_signature_31.png
| 2023-07-10 | 2023-07-06 11:00 | Dr. Jorge P. Royeca Hospital | General Santos City, South Cotabato | 12 | South Cotabato | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25093 | PBSP | fatchoi12001@yahoo.com | 29.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 873 ASSAY: Xpert HIV-1 Viral Load, COUNT: 149 | 2023-07-06 | Gx-IV | 838130 | Module 4 H/W failed | Replace module 4 SN: 806718 with new module SN: 210092886 | Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 color SN: 210092886 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5647622267379133009/5647622267379133009_signature_28.png | Maria Ivy A. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5647622267379133009/5647622267379133009_signature_31.png
| 2023-07-10 | 2023-07-10 08:15 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25092 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 230 ASSAY: Xpert MTB/RIF, COUNT: 554 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2002 ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 |  | GXIV-4D | 838285 | Modules can't be detected | #NAME? | None | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5647575875819240849/5647575875819240849_signature_28.png | Leah Karen Noynay | https://www.jotform.com/uploads/macarewecare/222927978372067/5647575875819240849/5647575875819240849_signature_31.png
| 2023-07-08 | 2023-07-04 11:01 | NGC Super Health Center | MRB Compound,Pilor Drive, QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Noel Floresca |  |  |  |  | 25091 | DOH | tmcolisao@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1024 |  | GX4 | 831586 | For relocation and installation,unit from holy spirit health center to be transfer to NGC super health center | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, transfer and installed unit | Installation and maintenance | Operational | Complete accessories |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5646049358911538856/5646049358911538856_signature_28.png | TRECIA MARIE D. COLISAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5646049358911538856/5646049358911538856_signature_31.png
| 2023-07-08 | 2023-07-04 14:34 | Payatas B Health Center | Bulacan st., Payatas | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 25090 | DOH | racquelmrocero@gmail.com | 36 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1434 |  | GX4 | 831899 | A3 intermittent temp.error, A4 intermittent loss communication | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, replaced cable | Troubleshooting and maintenance | For rexpertcheck A3 module,A4 operational | A2 reported already |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5646048518919721693/5646048518919721693_signature_28.png | Racquel M. Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5646048518919721693/5646048518919721693_signature_31.png
| 2023-07-08 | 2023-07-05 16:09 | National kidney and transplant institute | East avenue q.c | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25089 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 596 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 350 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1237 ASSAY: Xpert MTB/RIF Ultra, COUNT: 575 ASSAY: Xpert SARS-CoV-2, COUNT: 17466 ASSAY: Xpert HCV VL, COUNT: 26 ASSAY: Others, COUNT: 9 | 2023-06-26 | GXXVI | 110005587 | For module replacement and maintenance A3 210028986 - 655287 C1 210028937 -210036872, C4 210028989 - 890075 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | 13 modules operational | D2 and D3 disabled | 900-0491rs | 10047 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5646047638911197784/5646047638911197784_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5646047638911197784/5646047638911197784_signature_31.png
| 2023-07-08 | 2023-07-07 11:50 | Lung Center of the Philippines | Quezon avenue q.c | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 25088 | PBSP | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3800 ASSAY: Xpert SARS-CoV-2, COUNT: 4299 | 2023-07-07 | GXXVI | 801080 | For module replacement and maintenance C1 753022 - 210092918 D4 870126 - 210093415 | leaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | 14 modules operational | B1,b3 disabled | 900-0508 | G5724 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5646047048912407140/5646047048912407140_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5646047048912407140/5646047048912407140_signature_31.png
| 2023-07-07 | 2023-07-07 12:32 | Davao del Norte Provincial Hospital Kapalong Zone | National Highway, Kapalong Davao del Norte | 11 | Davao del Norte | Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25087 | PBSP | lot2salcedo@yahoo.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3884 |  | GX-IV 4D | 815270 | Module C3 encounter hardware failure | Replaced the defective Module C3 SN:686037 with new Module SN: 210062831 Clean Surface area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log | Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645391255207122801/5645391255207122801_signature_28.png | Leona S. Olila | https://www.jotform.com/uploads/macarewecare/222927978372067/5645391255207122801/5645391255207122801_signature_31.png
| 2023-07-07 | 2023-07-05 14:47 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25086 | PBSP | vincentmilestone20@gmail.com | 22.5 |  |  | GX-IV 4d | 834440 | Module D2 hardware failure | Replaced the Defective Module SN: 783912 with new Module SN: 210092888 | Module Replacement | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645390095208537237/5645390095208537237_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5645390095208537237/5645390095208537237_signature_31.png
| 2023-07-07 | 2023-07-05 14:44 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25085 | PBSP | vincentmilestone20@gmail.com | 22.5 |  |  | GX-IV 4d | 834626 | Module B1 h/w failure | Replace the defective Module B1 SN: 893800 with new Module SN : 210093394 | Perform Module Replacement | All Module are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645389735207916422/5645389735207916422_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5645389735207916422/5645389735207916422_signature_31.png
| 2023-07-07 | 2023-07-07 16:19 | Minglanilla RHU | Poblacion Ward 1, Tiber Minglanilla Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25084 | DOH | brendagonzagaparedes@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1410 ASSAY: Xpert HPV, COUNT: 908 |  | Gx4 | 838512 | Error 2008 in Module 3 and Module 4 | Instruct medtech to review the samples before loading into the cartridge. | Performed clean up in plunger, Icore and Filter | All 3 modules are used due to frequent error of 1002 in Module 2. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645279404218626273/5645279404218626273_signature_28.png | Brenda Paredes, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5645279404218626273/5645279404218626273_signature_31.png
| 2023-07-07 | 2023-07-07 15:08 | Commonwealth Health Center | Katuparan St., Commonwealth, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25083 | DOH | rizzabrazil13@gmail.com | 31 | ASSAY: Xpert HPV, COUNT: 585 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2273 | 2023-06-27 | GXIV 6C | 831615 | Monitor for replacement | #NAME? | Replaced monitor | Unit is operational. |  | Monitor | DR #10060 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645238521969905194/5645238521969905194_signature_28.png | Rizza B. Clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5645238521969905194/5645238521969905194_signature_31.png
| 2023-07-07 | 2023-07-07 13:29 | Tarlac City Health Center 2 | Brgy. Matatalaib, Tarlac City | 3 | Tarlac | Labortory | NO | TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25082 | DOH | clarizannebeltran1996@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 0 |  | GXIV-D | 831917 | Genexpert Application not opening Genexpert undetected Slow CPU booting up Failed to connect to a windows service | Troubleshooting | Troubleshoot network connection Reset hard drive and RAM Checked network services | Unit is not fully operational, Genexpert Application not booting up | For reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645196554426428121/5645196554426428121_signature_28.png | Clariz Anne M. Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5645196554426428121/5645196554426428121_signature_31.png
| 2023-07-07 | 2023-07-07 13:38 | Pandan Rural Health Unit | Libod, Pandan | 5 | Catanduanes | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25081 | PBSP | mjnamata10@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 | 2023-03-13 | GX4 | 838863 | For XpertCheck | Performed XpertCheck Cleaned Cartridge bays, plunger rod, icore | Performed Xpertcheck, Preventive Maintenance | Unit is operational All modules passed |  | Xpertcheck kit | PL-010410 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645189903417077463/5645189903417077463_signature_28.png | Mary Joan I Namata | https://www.jotform.com/uploads/macarewecare/222927978372067/5645189903417077463/5645189903417077463_signature_31.png
| 2023-07-07 | 2023-07-07 13:44 | Calamba City Health Office | Calamba | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25080 | PBSP | 2052456@aup.edu.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1972 | 2023-07-03 | GXIV | 807570 | Unit for PM. For module replacement of B3. | Performed PM. Performed module replacement of B3. Generated reports. | PM and module replacement | Unit is operational. |  | GX module B3 (210092908) in replacement to defective B3 (684562). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645182983562947996/5645182983562947996_signature_28.png | Mary Jane Lourdes R. Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5645182983562947996/5645182983562947996_signature_31.png
| 2023-07-07 | 2023-07-07 11:57 | Sta. Rosa City Health Office 1 | Sta. Rosa | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25079 | PBSP | santarosacho1clinlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 904 |  | GXIV | 804839 | Unit for preventive maintenance. | Performed preventive maintenance. Generated reports. | Preventive maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645116193565773254/5645116193565773254_signature_28.png | ALONA NINA ALVARO | https://www.jotform.com/uploads/macarewecare/222927978372067/5645116193565773254/5645116193565773254_signature_31.png
| 2023-07-07 | 2023-07-07 11:51 | Sta. Rosa City Health Office 1 | Sta. Rosa | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25078 | MACARE-HBDC | santarosacho1clinlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 819 |  | GXIV | 838519 | Unit for preventive maintenance. For troubleshooting of modules. | Performed preventive maintenance | Preventive maintenance and troubleshooting. | Unit is operational. Disabled module C3 (temperature error). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645114433568845508/5645114433568845508_signature_28.png | ALONA NINA L. ALVARO | https://www.jotform.com/uploads/macarewecare/222927978372067/5645114433568845508/5645114433568845508_signature_31.png
| 2023-07-07 | 2023-07-07 11:49 | Talavera General Hospital | Maestrang Kikay, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  | roxasmatthew659@yahoo.com.ph |  |  | 25077 | MACARE-HBDC | pjgtehlab@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 389 ASSAY: Xpert HBV VL, COUNT: 39 | 2023-07-07 | Genexpert IV | 110013773 | Error 2008 | Performed Xpertcheck | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645113231318720063/5645113231318720063_signature_28.png | Fitz Gerald Manuzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5645113231318720063/5645113231318720063_signature_31.png
| 2023-07-07 | 2023-07-07 11:21 | Sta Rosa City Health Office 1 | Sta. Rosa City | 4A | Laguna | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25076 | MACARE-HBDC | csrlcho1covidlab@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 11 ASSAY: Xpert SARS-CoV-2, COUNT: 789 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GXIV | 838522 | Unit for preventive maintenance. For troubleshooting of modules. | Performed preventive maintenance. Performed troubleshooting of modules with errors. Generated reports. | Preventive maintenance and troubleshooting. | Unit is operational. Disabled module B4 (temperature errors). Subject for replacement. Re-enabled other modules that are not hardware related errors. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645101813563613436/5645101813563613436_signature_28.png | Mariza B. Garing | https://www.jotform.com/uploads/macarewecare/222927978372067/5645101813563613436/5645101813563613436_signature_31.png
| 2023-07-07 | 2023-07-07 11:26 | Sta. Rosa City Health Office 1 | Sta. Rosa | 4A | Laguna | Molecular laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 25075 | MACARE-HBDC | csrlcho1covidlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 47 ASSAY: Xpert SARS-CoV-2, COUNT: 4851 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX16 | 844981 | Unit for preventive maintenance. For troubleshooting of modules. | Performed preventive maintenance. Performed troubleshooting of modules with errors. Generated report. | Preventive maintenance and troubleshooting. | Unit is operational. Disabled module B4, C2, D1, D2 and D4 (temperature errors). Subject for replacement. Re-enabled other modules that are not hardware related errors. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645101813566769951/5645101813566769951_signature_28.png | Mariza B. Garing | https://www.jotform.com/uploads/macarewecare/222927978372067/5645101813566769951/5645101813566769951_signature_31.png
| 2023-07-07 | 2023-07-07 11:13 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25074 | DOH | laguitannataniel@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 859 | 2023-07-03 | GXIV | 831677 | Monitor for replacement | #NAME? | Replaced monitor | Unit is operational. |  | GeneXpert Monitor SN: MMA1PSS002215078AB3W01 | 10074 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5645090058649092160/5645090058649092160_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5645090058649092160/5645090058649092160_signature_31.png
| 2023-07-06 | 2023-07-06 15:23 | Baras Medicare Community Hospital | Bagong Sirang, Catanduanes | 5 | Catanduanes | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25073 | PBSP | malustesoreromt@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 463 |  | GX4 | 835215 | For Preventive Maintenance | Cleaned Cartridge Bays, plunger rod, icore Generated reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644383587898302553/5644383587898302553_signature_28.png | Teresita Vergara | https://www.jotform.com/uploads/macarewecare/222927978372067/5644383587898302553/5644383587898302553_signature_31.png
| 2023-07-06 | 2023-07-06 12:37 | Virac Municipal Health Office | San Isidro Village, Virac, Catanduanes | 5 | Catanduanes | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25072 | DOH | limarmariano50@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2143 |  | GX4 | 838563 | For Xpertcheck | Performed Xpertcheck Cleaned Cartridge Bays, Plunger rod, Icore Generate Reports | Performed Xpertcheck, Preventive Maintenance | Unit is Operational B1(616796)-Passed B2(862611)-Passed B3(736617)-Passed B4(863009)-Service Required | Xpertcheck: B4(863009)-Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644306847893229173/5644306847893229173_signature_28.png | Ramil Mariano | https://www.jotform.com/uploads/macarewecare/222927978372067/5644306847893229173/5644306847893229173_signature_31.png
| 2023-07-06 | 2023-07-06 13:09 | Carcar Cho | Valladolid, Carcar City Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25071 | DOH | chaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1370 |  | Gx4 | 838283 | Error 2008 in Module 3 and 4 during running | Instruct medtech to dilute the sampe before loading the samples to cartridge. | Performed troubleshooting in Module 3 and 4. Clean Icore, plunger and cartrdge bay Clean filter and replaced with new filter | Only 3 modules are available due to module 2 having frequent error of 1002 during running. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644299929773481544/5644299929773481544_signature_28.png | Lucille Mae A. Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5644299929773481544/5644299929773481544_signature_31.png
| 2023-07-06 | 2023-07-06 12:05 | Philippine Red Cross | Intramuros, Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  |  |  |  | 25070 | COMMERCIAL | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. | Preventive maintenance. | Performed PM. Verified calibration parameters. Performed simulation running. Generated report. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644269084912680281/5644269084912680281_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/5644269084912680281/5644269084912680281_signature_31.png
| 2023-07-06 | 2023-07-06 12:07 | Philippine Red Cross | Intramuros, Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  |  |  |  | 25069 | COMMERCIAL | dolly.banares@redcross.org.ph | 25 |  |  | MCS+ | 02A087 | Unit for preventive maintenance. | Performed PM. Verified calibration parameters. Performed simulation running. Generated report. | Preventive maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644269044918456062/5644269044918456062_signature_28.png | Dolly Banares | https://www.jotform.com/uploads/macarewecare/222927978372067/5644269044918456062/5644269044918456062_signature_31.png
| 2023-07-06 | 2023-07-06 11:55 | HINIGARAN RURAL HEALTH UNIT | Hinigaran, Negros Occidental | 6 | Negros Occidental | Tab Laboratory | YES | INSTALLATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25068 | PBSP | mfeliprada@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 64 ASSAY: Xpert MTB/RIF Ultra, COUNT: 868 |  | GX-IV | 815283 | No UPS and AVR; Processing problems during power interruptions and fluctuations. | #NAME? | #NAME? | -4 modules operational; - needs UPS and AVR. Defective UPS., at least 2KVA capacity of both UPS and AVR. - reported earlier by previous Gx medtech. | - Needs AVR and UPS due to frequent power fluctuations in the facility; |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644256435141751271/5644256435141751271_signature_28.png | May Flor M. Feliprada | https://www.jotform.com/uploads/macarewecare/222927978372067/5644256435141751271/5644256435141751271_signature_31.png
| 2023-07-06 | 2023-07-06 09:57 | Tarlac City Health Center 3 | Brgy. San Miguel, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25067 | DOH | daganlea572@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 41 ASSAY: Xpert MTB/RIF Ultra, COUNT: 448 |  | GXIV-F | 831697 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Generate reports Archived test results Performed functionality test | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644243200222815994/5644243200222815994_signature_28.png | Daisy Coles | https://www.jotform.com/uploads/macarewecare/222927978372067/5644243200222815994/5644243200222815994_signature_31.png
| 2023-07-06 | 2023-07-06 09:57 | Tarlac City Health Center 3 | Brgy. San Miguel, Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25066 | DOH | daganlea572@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 41 ASSAY: Xpert MTB/RIF Ultra, COUNT: 448 |  | GXIV-F | 831697 | Due for maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Generate reports Archived test results Performed functionality test | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644243150227342849/5644243150227342849_signature_28.png | Daisy Coles | https://www.jotform.com/uploads/macarewecare/222927978372067/5644243150227342849/5644243150227342849_signature_31.png
| 2023-07-06 | 2023-07-06 11:12 | Bureau of Quarantine | Intramuros Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25065 | COMMERCIAL | dklo.lab@quarantine.doh.gov.ph | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 543 |  | GX16 | 844418 | Unit for preventive maintenance. | Performed preventive maintenance. Generated reports. | Preventive maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644226424918976058/5644226424918976058_signature_28.png | Genesis Gabine Viole Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/5644226424918976058/5644226424918976058_signature_31.png
| 2023-07-06 | 2023-07-05 16:31 | Cordlife Medical Philippines Inc. | UP Ayala Technohub, Commonwealth Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | REPAIR CALIBRATION | CYTIVA | Gian Paolo Magnaye |  |  |  |  | 25063 | COMMERCIAL | katrina.castro@cordlife.com | 25 |  | 2023-07-05 | Smartmax AS-310 | 5131.3196 | Defective smartmax cover inside assembly | Replqved defective part Performed functionality test Performed unit calibration | Replaced defective part Performed unit calibration | Unit is operational |  |  | 95867 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644223532195072730/5644223532195072730_signature_28.png | Katrina Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5644223532195072730/5644223532195072730_signature_31.png
| 2023-07-06 | 2023-07-06 10:53 | Dr. Lagrosa Health Center | Layug St., Brgy 3, Pasaya City | NCR | NCR 4th - Pasay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25064 | PBSP | LBCANICULA@gmail.com | 35 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2293 | 2023-03-16 | GX IV 6C | 805106 | Unit for XpertCheck | Checked & cleaned cartridge bays, plunger rod, icore slot & replaced filter Performed unit calibration B2/644514 - Failed B4/604721 - Insufficient kit Generated report | Performed PM Performed XpertCheck | Unit is operational. |  | XpertCheck kit | 1000410191 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644223532198250730/5644223532198250730_signature_28.png | Lorna Canicula | https://www.jotform.com/uploads/macarewecare/222927978372067/5644223532198250730/5644223532198250730_signature_31.png
| 2023-07-06 | 2023-07-06 10:46 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | Tb dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25062 | DOH | sarah_yangzon@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 118 ASSAY: Xpert MTB/RIF Ultra, COUNT: 707 | 2023-07-06 | GXIV | 838406 | N/a | N/a | Preventive maintenance - Performed Functionality test - Generate reports | Unit is Operational (module B1 disabled- for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644212843015686905/5644212843015686905_signature_28.png | Sarah Yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5644212843015686905/5644212843015686905_signature_31.png
| 2023-07-06 | 2023-07-05 13:51 | South Cotabato Provincial Hospital | Koronadal City, South Cotabato | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25061 | PBSP | anamariermt22@gmail.com | 22.8 | ASSAY: Xpert MTB/RIF, COUNT: 413 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2406 |  | Gx-IV | 815380 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644062369028412038/5644062369028412038_signature_28.png | Ana Marie Grace C. Egasan | https://www.jotform.com/uploads/macarewecare/222927978372067/5644062369028412038/5644062369028412038_signature_31.png
| 2023-07-06 | 2023-07-03 10:04 | Sultan Kudarat Provincial Hospital | Isulan, Sultan Kudarat | 12 | Sultan Kudarat | Tb- Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25061 | PBSP | charmfrias1270@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2985 |  | Gx-IV | 807405 | Module 1 stuck cartridge | Fixed/Report module 1 stuck cartridge | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5644062369029635869/5644062369029635869_signature_28.png | Charma Grace E. Frias | https://www.jotform.com/uploads/macarewecare/222927978372067/5644062369029635869/5644062369029635869_signature_31.png
| 2023-07-05 | 2023-07-05 16:26 | Cordlife Medical Philippines Inc | UP Ayala Technohub, Commonwealth Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CYTIVA | Patrick Los Baños |  |  |  |  | 25060 | COMMERCIAL | katrina.castro@cordlife.com | 25 |  |  | Sepax 2 | 6220.4122 | None, for Preventive Maintenance | Performed Preventive Maintenance Performed functionality test | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643603090311295686/5643603090311295686_signature_28.png | Katrina Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5643603090311295686/5643603090311295686_signature_31.png
| 2023-07-05 | 2023-07-05 15:43 | Liloan Rural Health Unit | Liloan, Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 25059 | PBSP | liloanmholab@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 115 ASSAY: Xpert MTB/RIF Ultra, COUNT: 551 |  | Gx 4 | 835189 | Preventive maintenance | Clean fan filter/plunger rod maintenance | Clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643550585419948612/5643550585419948612_signature_28.png | Annabelle V. Buot | https://www.jotform.com/uploads/macarewecare/222927978372067/5643550585419948612/5643550585419948612_signature_31.png
| 2023-07-05 | 2023-07-05 14:25 | Dalaguete Rhu | Dalaguete, Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25058 | PBSP | ajreef5517@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1005 |  | Gx4 | 815464 | Module 3 and 4 Run time error 1002 | Replaced new module 3 and 4 | Module replacement of New module 3 SN 210060575 and New module 4 SN 210073850 Generate data IQR, Sys log, Assay Statistics and Gxx files | Only 3 modules are available, module 1 is excluded due to run time error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643480298616645552/5643480298616645552_signature_28.png | Alyzza Jane Rivera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5643480298616645552/5643480298616645552_signature_31.png
| 2023-07-05 | 2023-07-05 13:13 | LORENZO D. ZAYCO DISTRICT HOSPITAL | National Highway, Kabankalan City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 25057 | PBSP | beuagravante@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1782 |  | GX-IV | 807455 | Defective modules A1/682751 and A3/666201. | #NAME? | - Replaced modules A1 and A3 with new modules SN A1/210092460 and A3/210092893. | All 4 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643470077024125113/5643470077024125113_signature_28.png | BEULAH AGRAVANTE | https://www.jotform.com/uploads/macarewecare/222927978372067/5643470077024125113/5643470077024125113_signature_31.png
| 2023-07-05 | 2023-07-05 13:44 | Bugasong Rural Health Unit | Bugasong, Antique | 6 | Antique | Tb Dots | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 25056 | PBSP | jeanelyngabuyogimalay@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 914 |  | GXIV-4D | 815427 | #NAME? | #NAME? | #NAME? | Module B1-B3 are functional. Module B4 for reporting and further troubleshooting by Cepheid. CPU is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643457514228402149/5643457514228402149_signature_28.png | Jeanelyn G. Imalay | https://www.jotform.com/uploads/macarewecare/222927978372067/5643457514228402149/5643457514228402149_signature_31.png
| 2023-07-05 | 2023-07-05 13:22 | San Antonio District Hospital | San Mariano, San Antonio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25055 | PBSP | garciaadonisanthony@gmail.com | 29 |  | 2023-07-05 | Genexpert IV | 812832 | Monitor Problem | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit not Working | Monitor for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643442994028849219/5643442994028849219_signature_28.png | Adonis Anthony Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5643442994028849219/5643442994028849219_signature_31.png
| 2023-07-05 | 2023-07-05 13:22 | San Antonio District Hospital | San Mariano, San Antonio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 25054 | PBSP | garciaadonisanthony@gmail.com | 29 |  | 2023-07-05 | Genexpert IV | 812832 | Monitor Problem | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit not Working | Monitor for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643441264022339573/5643441264022339573_signature_28.png | Adonis Anthony Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5643441264022339573/5643441264022339573_signature_31.png
| 2023-07-05 | 2023-07-05 13:21 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 25053 | DOH | vanessadelmonte@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 578 | 2023-07-05 | GXIV | 838409 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643439695415656196/5643439695415656196_signature_28.png | Vanessa mabag | https://www.jotform.com/uploads/macarewecare/222927978372067/5643439695415656196/5643439695415656196_signature_31.png
| 2023-07-05 | 2023-07-05 13:01 | Tagbina Rural Health Unit | Poblacion, Tagbina | 13 | Surigao del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 25052 | PBSP | ellalaguros@yahoo.com | 24.18 | ASSAY: Xpert MTB/RIF, COUNT: 41 ASSAY: Xpert MTB/RIF Ultra, COUNT: 713 |  | GX-4 | 838877 | None | None | Replacing Module B1 with SN: 861257 to SN: 210092772 | All 4 modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643436925915265438/5643436925915265438_signature_28.png | Ellaquim L. Albarracin | https://www.jotform.com/uploads/macarewecare/222927978372067/5643436925915265438/5643436925915265438_signature_31.png
| 2023-07-05 | 2023-07-05 12:11 | Chinese General Hospital | Blumentritt St., Sta. Cruz Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 25051 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 21 ASSAY: Xpert CT/NG, COUNT: 37 ASSAY: Others, COUNT: HBV 83 ASSAY: Xpert HPV, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1481 ASSAY: Xpert HCV VL, COUNT: 15 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 | 2023-07-03 | GXIV | 835604 | Unit for PM. For module replacement of Module B4. | Performed PM. Performed module replacement on B4. Generated reports. | PM and module replacement. | Unit is operational |  | GX module B4 (682515) in replacement to defective B4 (651975). | DR 10072 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643399604913764230/5643399604913764230_signature_28.png | Mary Joyce Gaudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5643399604913764230/5643399604913764230_signature_31.png
| 2023-07-05 | 2023-06-29 10:34 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25050 | MACARE-HBDC | balbahazel@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 193 |  | GX 4 | 821632 | For Preventive Maintenance | Cleaned Cartridge bays, Plunger Rods, and I-CORE Generated Reports | Performed PREVENTIVE MAINTENANCE | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643269420518181402/5643269420518181402_signature_28.png | HAZEL GWYNETH BALBA | https://www.jotform.com/uploads/macarewecare/222927978372067/5643269420518181402/5643269420518181402_signature_31.png
| 2023-07-05 | 2023-06-29 12:04 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25049 | PBSP | keithtagudar@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 756 |  | GX 4 6C | 834685 | Module B2 | Cleaned Cartridge Bays, Plunger Rods and I-CORE Performed ReXpertCheck Generated Reports | Performed ReXpertCheck | Unit is operational | Module B2 failed. For reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643267560517917679/5643267560517917679_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5643267560517917679/5643267560517917679_signature_31.png
| 2023-07-05 | 2023-06-29 10:02 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | Molecular Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25048 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 22 |  | 2023-06-29 | GX 4 10C | 110016112 | For Preventive Maintenance | Cleaned Cartridge bays, Plunger Rods and I-CORE Generated Reports | Performed preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643265850513319659/5643265850513319659_signature_28.png | April Joy Gundran | https://www.jotform.com/uploads/macarewecare/222927978372067/5643265850513319659/5643265850513319659_signature_31.png
| 2023-07-05 | 2023-06-29 12:04 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 25047 | PBSP | keithtagudar@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 756 |  | GX 4 6C | 834685 | Module B2 | Cleaned Cartridge Bays, Plunger Rods and I-CORE Performed ReXpertCheck Generated Reports | Performed ReXpertCheck | Unit is operational | Module B2 failed. For reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643264070516232900/5643264070516232900_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5643264070516232900/5643264070516232900_signature_31.png
| 2023-07-05 | 2023-06-29 08:59 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 25046 | PBSP | keithtagudar@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 386 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1592 | 2023-06-29 | GX 4 6C | 805212 | Due for Module Replacement | -Cleaned Cartridge bays, I-CORE and Plunger Rods -Replaced filter -Replaced Module B3 875228 >>> 210059820 | Performed Module Replacement and Preventive Maintenance | Unit is operational |  | Module 21005982 | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643262290512408829/5643262290512408829_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5643262290512408829/5643262290512408829_signature_31.png
| 2023-07-05 | 2023-06-29 13:34 | Baguio Health Services Office | T. Alonzo, Baguio City | CAR | Benguet | Laboratory | NO | CALIBRATION | CEPHEID | Yasser Salih |  |  |  |  | 25045 | DOH | sherryannemarrero@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 993 | 2023-06-13 | GX 4 | 838281 | For Calibration | Cleaned Cartridge bays, Plunger Rods and I-CORE Generated Reports | Perform XpertCheck | Unit is operational | 3 passed 1 fail, Module B4 for reporting, B1 error after XpertCheck for reporting | XpertCheck Kits | DR 5335 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643261120512471035/5643261120512471035_signature_28.png | Sherry anne marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5643261120512471035/5643261120512471035_signature_31.png
| 2023-07-05 | 2023-06-30 14:12 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 25044 | PBSP | romarmtamayo@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1540 ASSAY: Xpert SARS-CoV-2, COUNT: 3222 |  | GX 16 | 842230 | Module B1 error prone | Performed reliability test Generated Reports | Troubleshoot module B1 | Unit is operational | Module B1 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5643259910514045080/5643259910514045080_signature_28.png | Romar M Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5643259910514045080/5643259910514045080_signature_31.png
| 2023-07-04 | 2023-07-04 16:45 | San Jose Rural Health Unit | San Jose, Antique | 6 | Antique | Tb Dots | YES | REPAIR CALIBRATION | CEPHEID | Wynnona Reoma |  |  |  |  | 25043 | PBSP | sanjoseantiquegxsite@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 102 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1370 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 | 2023-07-04 | GXIV-4D | 804714 | Module C1 defective Module C2- reported temp error | -Replace defective module C1/723403 with new module SN: 210067892 | #NAME? | Module C3 disabled - ultrasonic horn error; Module C2 - xpertcheck service required; Module C1 and C4 - functional. |  | GeneXpert Module 6 colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642954249713736287/5642954249713736287_signature_28.png | Oscar Delano Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5642954249713736287/5642954249713736287_signature_31.png
| 2023-07-04 | 2023-07-04 16:33 | San Jose Rural Health Unit | San Jose, Antique | 6 | Antique | TB Dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25042 | DOH | sanjoseantiquegxsite@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 102 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1369 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 |  | GXIV-4D | 838412 | Module B2 and B4 - h/w failed | None | #NAME? | Module B1 and B3 are functional. Module B2 and B4 waiting for allocation - xpertcheck service required. | 2 xpc cartridge used. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642954249712914150/5642954249712914150_signature_28.png | Oscar Delano Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5642954249712914150/5642954249712914150_signature_31.png
| 2023-07-04 | 2023-07-04 16:55 | Sumilang Super Health Center | Sumilang, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25041 | PBSP | myrlising@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 150 ASSAY: Xpert MTB/RIF Ultra, COUNT: 287 |  | GXIV 6C | 814163 | Asked assistance in retrieving file from the hard drive of reported cpu | #NAME? | troubleshoot hard drive | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642705959613344220/5642705959613344220_signature_28.png | Myrna P. Lising | https://www.jotform.com/uploads/macarewecare/222927978372067/5642705959613344220/5642705959613344220_signature_31.png
| 2023-07-04 | 2023-07-04 15:29 | St Lukes Extension Clinic | J Bocobo St, Manila | NCR | NCR 1st - Manila | TB Laboratory | YES | INSTALLATION | CEPHEID | Patrick Los Baños |  |  |  |  | 25040 | MACARE-HBDC | laboratory@slec.ph | 30 | ASSAY: Others, COUNT: |  | GX-IV 10C | 110016786 | None | Installed unit and workstation Performed functionality test Generated reports | Installed unit and workstation | Unit is operational | Room intake and AC temporarily not functioning, recommended to be operational before start of processing samples |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642661425691259117/5642661425691259117_signature_28.png | Khariz Cabiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5642661425691259117/5642661425691259117_signature_31.png
| 2023-07-04 | 2023-07-04 13:34 | Guiguinto Rural Health Unit II | Guiguinto, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25039 | PBSP | joie_calica@yahoo.com | 27 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 796 ASSAY: Xpert HIV-1 Viral Load, COUNT: 178 | 2023-07-04 | GXIV | 814165 | Module B4 (784627) for replacement. | #NAME? | Replaced B4 | Unit is operational. |  | GeneXpert Module SN: 210091145 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642586400711227224/5642586400711227224_signature_28.png | Ligaya Calica | https://www.jotform.com/uploads/macarewecare/222927978372067/5642586400711227224/5642586400711227224_signature_31.png
| 2023-07-04 | 2023-07-04 12:24 | Pilar Rural Health Unit | Local and National Government Complex, Brgy., Calongay, Pilar, Sorsogon | 5 | Sorsogon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25038 | DOH | mariarramae@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1542 |  | GX4 | 838266 | For XpertCheck | Performed XpertCheck Cleaned Cartridge bays, plunger rod, fan filter Generated reports | Performed XpertCheck Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642577589111970320/5642577589111970320_signature_28.png | Arra A. Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5642577589111970320/5642577589111970320_signature_31.png
| 2023-07-04 | 2023-07-04 11:50 | Gerona Rural Health Unit | Brgy. Pinasling Geron Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25037 | DOH | huwakie15@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 457 |  | GXIV-D | 831693 | Module A3 for re-xpertcheck | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except A1, A3 passed calibration | Used 1 xpertcheck kit, under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642530138018132716/5642530138018132716_signature_28.png | Ana Marie Nool | https://www.jotform.com/uploads/macarewecare/222927978372067/5642530138018132716/5642530138018132716_signature_31.png
| 2023-07-04 | 2023-07-04 11:31 | Mandaue CHO | Mandaue City, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25036 | DOH | elma.bragat@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3287 |  | Gx4 | 838510 | Error 4009 Hardware Failure in Module 4 Sn 861023 | Clean Filter, adjust aircon temp | Generate data, IQR, Sys log and Assay statistics, gxx files | C4, B1, B2 are unusable due to temperature related problem |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642510817616190632/5642510817616190632_signature_28.png | Elma G. Bragat, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5642510817616190632/5642510817616190632_signature_31.png
| 2023-07-04 | 2023-07-04 10:47 | Ramos Rural Health Unit | Ramos Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 25035 | DOH | jadeeh8143@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 71 ASSAY: Xpert MTB/RIF Ultra, COUNT: 334 |  | GXIV-D | 831721 | UPS not backing up for more than 1 minute | UPS Troubleshooting Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Performed UPS troubleshooting Gathered reports Performed functionality test | Units is operational except for module 2 and UPS | UPS Battery for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642494128016868898/5642494128016868898_signature_28.png | Emily Jade Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/5642494128016868898/5642494128016868898_signature_31.png
| 2023-07-04 | 2023-07-04 10:39 | Cagayan de Oro City Health Office | Brgy. 40, Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25034 | DOH | hbpiquero.pbsp@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3226 |  | GeneXpert | 838441 | None | None | Preventive Maintenance Generate Report files Database back up | B1/806412- Error 4009 , B2/731414-H/W Fail, B4/806415-H/W Failand B3/806414 some results are Indeterminate |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5642479388022270442/5642479388022270442_signature_28.png | Piquero, Heide | https://www.jotform.com/uploads/macarewecare/222927978372067/5642479388022270442/5642479388022270442_signature_31.png
| 2023-07-03 | 2023-06-30 13:59 | Solano Rural Health Unit | Roxas, Solano, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25033 | DOH | rhusolanolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 556 | 2023-06-30 | Genexpert IV | 838201 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641798293615987831/5641798293615987831_signature_28.png | Roderick Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5641798293615987831/5641798293615987831_signature_31.png
| 2023-07-03 | 2023-06-30 13:59 | Solano Rural Health Unit | Roxas, Solano, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 25032 | DOH | rhusolanolab@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 556 | 2023-06-30 | Genexpert IV | 838201 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641798223611987926/5641798223611987926_signature_28.png | Roderick Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5641798223611987926/5641798223611987926_signature_31.png
| 2023-07-03 | 2023-07-03 13:05 | Sagay City Health Office | Poblacion, Sagay City | 6 | Negros Occidental | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 25031 | PBSP | maravelesdebra@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 134 | 2023-06-30 | GX4 | 838144 | Module B2 error code 2014 with serial number 691053 | Replace Module B2 with new serial number 210092919 | Replace Module B2 | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641767583914961327/5641767583914961327_signature_28.png | Debra Maraveles | https://www.jotform.com/uploads/macarewecare/222927978372067/5641767583914961327/5641767583914961327_signature_31.png
| 2023-07-03 | 2023-06-27 15:56 | Philippine heart center | East ave. Qc | NCR | NCR 2nd - Quezon City | Bloodbank | NO | INSTALLATION | HAEMONETICS | Noel Floresca |  |  |  |  | 25030 | COMMERCIAL | baluyotjeric@gmail.com | 28 |  |  | MCS + | 11H204 | For delivery | Performed installation,and calibration Checked all parameters Test run unit | Installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641747153614663453/5641747153614663453_signature_28.png | Amelia L. Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5641747153614663453/5641747153614663453_signature_31.png
| 2023-07-03 | 2023-06-29 16:58 | Oriental Mindoro Provincial Hospital | Calapan | 4B | Oriental Mindoro | Molecular | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25029 | PBSP | omphlabsarscov2@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2379 ASSAY: Xpert SARS-CoV-2, COUNT: 2314 ASSAY: Xpert HIV-1 Viral Load, COUNT: 44 |  | GX16-D | 842229 | Module C3 with Temperarure Error Stuck Cartridge in A3 | Module Replacement Self-Test | Performed Maintenance Performed Self-Test and Removed Stuck Cartridge Replaced Module C3 Generate Reports | All Modules are Operational | C3: 883404-> 210091311 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641721293613208833/5641721293613208833_signature_28.png | Maria Theresa Yaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5641721293613208833/5641721293613208833_signature_31.png
| 2023-07-03 | 2023-06-27 13:09 | GMA Rural Health Unit II | Brgy. Memije, General Mariano Alvarez | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 25028 | PBSP | ldlconti@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1680 | 2023-06-27 | GXIV -D | 814152 | Module 2 with Temperature Errors | Module Replacement | Performed Maintenance Replaced Module 2 Generate Reports Performed Functionality Test | 3 Modules are Operational, 1 Module for Replacement |  | 900-0508 | G 5700 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641721293617838506/5641721293617838506_signature_28.png | Lorena L. Conti | https://www.jotform.com/uploads/macarewecare/222927978372067/5641721293617838506/5641721293617838506_signature_31.png
| 2023-07-03 | 2023-07-03 13:17 | Balasan Municipal Health Office | Balasan, Iloilo | 6 | Iloilo | Laboratory | YES | REPAIR ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25027 | PBSP | charliemagnedk2@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1578 |  | GXIV-4D | 807555 | Module B1 Stuck Cartridge | #NAME? | #NAME? | Module B2 waiting for allocation. Three other modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641717522694032855/5641717522694032855_signature_28.png | Charlie Magne Dequito | https://www.jotform.com/uploads/macarewecare/222927978372067/5641717522694032855/5641717522694032855_signature_31.png
| 2023-07-03 | 2023-06-30 13:09 | Libmanan Rural Health Unit | Libmanan, Camarines Sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25026 | PBSP | amyquiapo1972@gmail.com | 26 |  | 2023-06-26 | GX4 | 805260 | For module replacement | Replaced module: 677428>>210092909 Cleaned cartridge bays, plunger rod Generated reports | Performed module replacement, preventive maintenence | Unit is operational |  | Module 210092909 | PL-011613 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641645670328997146/5641645670328997146_signature_28.png | Amy Quiapo | https://www.jotform.com/uploads/macarewecare/222927978372067/5641645670328997146/5641645670328997146_signature_31.png
| 2023-07-03 | 2023-07-03 11:11 | San Francisco RHU | San Francisco Camotes, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25025 | PBSP | ejoey22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 |  | Gx2 | 835058 | Due for prevebntive maintenance | Performed preventive maintenance | Performed machine check up Check ups, monitor, bar code scanner cpu Generate datta, IQR, sys log and assay statistics Back up database and archieve test results. Clean plunger, Icore and clean filter | The machines are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641635813913650219/5641635813913650219_signature_28.png | Josiecorra Albaraccin | https://www.jotform.com/uploads/macarewecare/222927978372067/5641635813913650219/5641635813913650219_signature_31.png
| 2023-07-03 | 2023-07-03 11:11 | San Francisco RHU | San Francisco Camotes, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25024 | PBSP | ejoey22@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 |  | Gx2 | 835058 | Due for prevebntive maintenance | Performed preventive maintenance | Performed machine check up Check ups, monitor, bar code scanner cpu Generate datta, IQR, sys log and assay statistics Back up database and archieve test results. Clean plunger, Icore and clean filter | The machines are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641635603912536136/5641635603912536136_signature_28.png | Josiecorra Albaraccin | https://www.jotform.com/uploads/macarewecare/222927978372067/5641635603912536136/5641635603912536136_signature_31.png
| 2023-07-03 | 2023-07-03 11:11 | San Francisco RHU | San Francisco Camotes, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 25023 | PBSP | ejoey22@gmail.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 489 |  | Gx2 | 835058 | Due for prevebntive maintenance | Performed preventive maintenance | Performed machine check up Check ups, monitor, bar code scanner cpu Generate datta, IQR, sys log and assay statistics Back up database and archieve test results. Clean plunger, Icore and clean filter | The machines are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5641634733918827272/5641634733918827272_signature_28.png | Josiecorra Albaraccin | https://www.jotform.com/uploads/macarewecare/222927978372067/5641634733918827272/5641634733918827272_signature_31.png
| 2023-07-01 | 2023-06-29 13:52 | Malungon Rural Health Unit | Malungon, Sarangani Province | 12 | Sarangani | Tb-Lab | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 25022 | DOH | madameerrmie@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 226 |  | Gx-IV | 838546 | Module 2 for Service Required Module 3 h/w failed due to multiple temperature error | Perform Xpertcheck at module 3 for troubleshooting | Perform Genexpert machine Preventive maintenance | 3 operational modules | Module 3 the Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5640293115811184553/5640293115811184553_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5640293115811184553/5640293115811184553_signature_31.png
| 2023-07-01 | 2023-06-30 10:32 | Polomolok Rural Health Unit | Poblacion, Polomolok, South Cotabato | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 25021 | DOH | define.awkward90@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1707 |  | Gx-IV | 838265 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5640293115813957506/5640293115813957506_signature_28.png | Jethro Camaganacan | https://www.jotform.com/uploads/macarewecare/222927978372067/5640293115813957506/5640293115813957506_signature_31.png
| 2023-07-01 | 2023-06-30 12:26 | University of Santo tomas hospital | España, Manila | NCR | NCR 1st - Manila | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 25020 | COMMERCIAL | usth_bloodbank@yahoo.com | 28 |  |  | MCS + | 02F069 | As per end user they encounter unusual sound coming from centrifuge | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,cleaned centrifuge,performed functionality test.remove debris in centrifuge | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5640085603812890140/5640085603812890140_signature_28.png | Mylene Balb0a | https://www.jotform.com/uploads/macarewecare/222927978372067/5640085603812890140/5640085603812890140_signature_31.png
| 2023-06-30 | 2023-06-29 16:12 | Davao Oriental Provincial Medical Center | National Highway, Mati City | 11 | Davao Oriental | Microbiology | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 25019 | PBSP | ofeliamendiola3@gmail.com | 28.5 |  |  | GX-IV 4D | 804693 | Stuck Cartridge, For Preventive Maintenance and XpertCheck Calibration | Tuning the module cartridge bay for enable to remove the stuck cartridge Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform Xpertcheck Calibration | Trouble the Modulea that encounter stuck cartridge, perform Preventive Maintenance and XpertCheck Calibration | Successfully removed the stucked cartridges and all Modules Passed the Calibration | High Room Temp |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639324661818155617/5639324661818155617_signature_28.png | Ofelia Mendiola RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5639324661818155617/5639324661818155617_signature_31.png
| 2023-06-30 | 2023-06-30 16:04 | San Miguel District Hospital | San Miguel, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 25018 | PBSP | smdh_laboratory@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 942 | 2023-06-30 | GXIV | 807604 | Modules A2 (629504), A3 (662463) for replacement | #NAME? | Replaced A2, A3 | Unit is operational. |  | GeneXpert Module SN: 210060794, 210092891 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639223665313166904/5639223665313166904_signature_28.png | Florencia Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5639223665313166904/5639223665313166904_signature_31.png
| 2023-06-30 | 2023-06-30 14:45 | Pedro Gil Health Center | Pedro Gil, Sta. Ana | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25017 | PBSP | acg_medtech@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 313 |  | GXIV 6C | 812829 | For XpertCheck | #NAME? | Performed XpertCheck | Unit is operational B1: New B2: Failed B3: New B4: Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639180007424925740/5639180007424925740_signature_28.png | Lorraine Lucila L Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5639180007424925740/5639180007424925740_signature_31.png
| 2023-06-30 | 2023-06-30 14:48 | San Fernando Rural Health Unit | San Fernando | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25016 | DOH | josephus_garfin@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1278 |  | GX4 | 838496 | For Re-XpertCheck | Performed re-XpertCheck Cleaned cartridge bays, plunger rod Generated reports | Performed Re-XpertCheck | Unit is operational | B2 re-XpertCheck for retry suggested |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639176160611471755/5639176160611471755_signature_28.png | Josephus garfin | https://www.jotform.com/uploads/macarewecare/222927978372067/5639176160611471755/5639176160611471755_signature_31.png
| 2023-06-30 | 2023-06-30 14:27 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25015 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 228 ASSAY: Xpert MTB/RIF, COUNT: 562 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2002 ASSAY: Xpert HIV-1 Viral Load, COUNT: 26 |  | GXIV-4D | 820206 | None | None | #NAME? | All four modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639162140324053043/5639162140324053043_signature_28.png | Myra Riccil Estose,Rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5639162140324053043/5639162140324053043_signature_31.png
| 2023-06-30 | 2023-06-30 14:28 | TALISAY CITY HEALTH UNIT | CANSOJONG TALISAY CITY | 7 | Cebu | RTDL | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 25014 | PBSP | marydelgella@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1344 ASSAY: Xpert HIV-1 Viral Load, COUNT: 54 |  | GX-IV | 815384 | FOR MODULE B3 AND MODULE B4 REPLACEMENT | #NAME? | -PERFORM MODULE B3 DEF SN: 210019110 REPLACEMENT WITH NEW SN: 210092127 -PERFORM MODULE B4 DEF SN: 210018733 WITH NEW SN: 210092942 -PERFORM GX EQUIPMENT CLEANING | EQUIPMENT IS OPERATIONAL | Module B2 FOR OBSERVATION |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639161136937039762/5639161136937039762_signature_28.png | Marydel Gella | https://www.jotform.com/uploads/macarewecare/222927978372067/5639161136937039762/5639161136937039762_signature_31.png
| 2023-06-30 | 2023-06-30 14:22 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 25013 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 230 ASSAY: Xpert MTB/RIF, COUNT: 601 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1946 ASSAY: Xpert HIV-1 Viral Load, COUNT: 27 |  | GXIV-4D | 838285 | None | None | #NAME? | All modules are functional (w/ service module A3) | Waiting for module A3 allocation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639161100327420394/5639161100327420394_signature_28.png | Myra Riccil Estose, Rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5639161100327420394/5639161100327420394_signature_31.png
| 2023-06-30 | 2023-06-30 12:18 | Zamboanga City Medical Centre | Dr. Evangelista Street, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25012 | WHO | jbcua.pbsp@gmail.com | 28 | ASSAY: Others, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 356 ASSAY: Others, COUNT: 329 ASSAY: Xpert SARS-CoV-2, COUNT: 2462 ASSAY: Xpert HIV-1 Viral Load, COUNT: 28 | 2023-06-30 | GX 4 | 848250 | None | All 4 modules undergone Xpertcheck. While module 1 and 2 service required , module 3 and 4 passed. | Xpertcheck calibration and Preventive Maintenence | All modules can be used | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639088115917690386/5639088115917690386_signature_28.png | James peniel b. Cua | https://www.jotform.com/uploads/macarewecare/222927978372067/5639088115917690386/5639088115917690386_signature_31.png
| 2023-06-30 | 2023-06-30 12:12 | Zamboanga City Medical Centre | Dr. Evangelists Street, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 25011 | PBSP | jbcua.pbsp@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Others, COUNT: 1211 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4944 ASSAY: Xpert SARS-CoV-2, COUNT: 15477 ASSAY: Xpert HIV-1 Viral Load, COUNT: 303 | 2023-06-30 | GX 16 | 839309 | Module A1, A2 and C3 h/w failed | All 16 modules undergone Xpertcheck calibration and Preventive Maintenence. | All 16 modules undergone Xpertcheck calibration and Preventive Maintenence. | Module A1, A2 and C3 h/w failed | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639088115918624962/5639088115918624962_signature_28.png | James peniel b. Cua | https://www.jotform.com/uploads/macarewecare/222927978372067/5639088115918624962/5639088115918624962_signature_31.png
| 2023-06-30 | 2023-06-30 11:20 | Rosario Reyes Health Center | 627 San Andres St., Malate | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25010 | DOH | moninabatil16.prime@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2369 |  | GXIV 6C | 831843 | Module 2 frequent error | -Performed PM -Module 1, 2 and 4 already for replacement -Assisted end user in using jotform -Generated report | Checked module 2 Jotform assistance | Unit is operational but only module 3 works perfectly. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639065307019822308/5639065307019822308_signature_28.png | monina batil | https://www.jotform.com/uploads/macarewecare/222927978372067/5639065307019822308/5639065307019822308_signature_31.png
| 2023-06-30 | 2023-06-30 09:16 | Misamis Occidental Provincial Hospital | Oroquieta City | 10 | Misamis Occidental | Laboratot | YES | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25009 | PBSP | mophlaboratory2017@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1458 |  | GeneXpert | 807618 | Module A3/653705 and A4/695639 fail the calibration | Replaced Module new module A3/210067861 A4/210059819 | Preventive Maintenance Module replacement Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639004737187939812/5639004737187939812_signature_28.png | Mary Auxilium Omandam | https://www.jotform.com/uploads/macarewecare/222927978372067/5639004737187939812/5639004737187939812_signature_31.png
| 2023-06-30 | 2023-06-30 09:16 | Misamis Occidental Provincial Hospital | Oroquieta City | 10 | Misamis Occidental | Laboratot | YES | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 25008 | PBSP | mophlaboratory2017@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1458 |  | GeneXpert | 807618 | Module A3/653705 and A4/695639 fail the calibration | Replaced Module new module A3/210067861 A4/210059819 | Preventive Maintenance Module replacement Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5639004162212842816/5639004162212842816_signature_28.png | Mary Auxilium Omandam | https://www.jotform.com/uploads/macarewecare/222927978372067/5639004162212842816/5639004162212842816_signature_31.png
| 2023-06-30 | 2023-06-30 09:42 | Kalinga Health Hub Marikina | JP Rizal St. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 25007 | CONSORTIUM | jasmin.catral@innovationsch.org | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 235 |  | GX-IV | 110008418 | For maintenance/check-up | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Functionality test Generated reports | All modules operational and ready to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638990490937368484/5638990490937368484_signature_28.png | Jasmin Catral RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5638990490937368484/5638990490937368484_signature_31.png
| 2023-06-30 | 2023-06-30 07:52 | Pilar RHU | Poblacion Pilar Camotes | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 25006 | PBSP | steveyong3384@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 75 |  | Gx2 | 834813 | Due for Preventive Maintenance | Performed preventive maintenance, and Instruct medtech to maintain ambient temperature 23 to 28 Celcius | Performed maintenance, clean Icore, Plunger, and cartridge bay Generate data IQR, Sys log, and Gxx files Check ups and Avr | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638985241389776984/5638985241389776984_signature_28.png | Steve Cureg, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5638985241389776984/5638985241389776984_signature_31.png
| 2023-06-30 | 2023-06-30 09:21 | Veterans Memorial Medical Center | North Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory - Bacteriology | NO | Software Upgrade | CEPHEID | Carlos Leniel Origenes |  |  |  |  | 25005 | MACARE-HBDC | carlos.leniel.origenes@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 142 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1044 |  | GXIV-D 6C | 839150 | Unable to transmit results to LIS | Performed software upgrade from GX Dx v5.1 to 6.4 | Troubleshooting of unit Testing using service laptop with GX Dx v6.4 Performed functionality tests | Unit is operational | For USB to Lan and LIS configuration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638976534814620576/5638976534814620576_signature_28.png | Veterans Memorial Medical Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5638976534814620576/5638976534814620576_signature_31.png
| 2023-06-29 | 2023-06-29 14:44 | National children's hospital | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 25004 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2970 ASSAY: Xpert SARS-CoV-2, COUNT: 7703 ASSAY: Others, COUNT: 18 | 2023-06-26 | GXXVI (10c) | 110014171 | For module replacement A1 210069229 - 887323, B4 210069244 - 210062093 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit. Performed module replacement | Module replacement | Operational |  | 900-0650rs | 10046 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638509376266965207/5638509376266965207_signature_28.png | Jessica Faye Basilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5638509376266965207/5638509376266965207_signature_31.png
| 2023-06-29 | 2023-06-29 17:23 | Region II Trauma and Medical Center | Bayombong, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 25003 | PBSP | cecillepabance@gmail.com | 29 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1816 ASSAY: Xpert HIV-1 Viral Load, COUNT: 54 | 2023-06-29 | Genexpert IV | 807458 | Module 3 and 4 error, Module 3 stuck cartridge | Performed Xpertcheck on B3 and B4 module | Performed Maintenance Gathered Reports Reset Gateway Board Performed Functionality Test | Unit is Operational | Xpertcheck Result = B3, B4 - Passed | Xpertcheck 3pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638405021991966542/5638405021991966542_signature_28.png | Maricris N Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5638405021991966542/5638405021991966542_signature_31.png
| 2023-06-29 | 2023-06-29 16:48 | Bicol Medical Center | Concepcion, Naga City | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 25002 | PBSP | bmc.bacteriology@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 1575 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6280 ASSAY: Xpert SARS-CoV-2, COUNT: 3094 ASSAY: Xpert HIV-1 Viral Load, COUNT: 326 | 2023-06-27 | GX-16 | 842225 | D2,D3,D4 for module replacement | Replaced module: D2 886348>>210092932, D3 886347>>210092889, D4 878345>>210092875 Cleaned cartridge bays, plunger rods Generated reports | Performed module replacement & preventive maintenence | Unit is operational D2 886348>>210092932 D3 886347>>210092889 D4 878345>>210092875 |  | Module: 210092932, 210092889, 210092875 | PL-011612 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638404156718425458/5638404156718425458_signature_28.png | Angelika Mae P. Bichara | https://www.jotform.com/uploads/macarewecare/222927978372067/5638404156718425458/5638404156718425458_signature_31.png
| 2023-06-29 | 2023-06-29 17:20 | Banlat Health Center | Brgy. Banlat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 25001 | PBSP | reyzelmariesantos@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 95 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1440 |  | GXIV 6C | 804688 | Module A1 not detected | -Performed PM -Reset GX; problem persists -Switch lan cables of A1 and A2 and tighten gateway; problem solved -Generated report | Troubleshoot module 1 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638400342297302692/5638400342297302692_signature_28.png | Reyzel marie S. Dulay | https://www.jotform.com/uploads/macarewecare/222927978372067/5638400342297302692/5638400342297302692_signature_31.png
| 2023-06-29 | 2023-06-29 16:30 | Bicol Medical Center | Concepcion, Naga City | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 25000 | MACARE-HBDC | bmc.bacteriology@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 913 ASSAY: Xpert MTB/RIF Ultra, COUNT: 315 ASSAY: Xpert SARS-CoV-2, COUNT: 515 |  | GX-4 | 110008458 | For re-XpertCheck | Performed re-XpertCheck Cleaned plunger rod and cartridge bays Generated reports | Performed re-XpertCheck, Preventive Maintenance | Unit is operational B3 failed re-XpertCheck | B3 failed on re-XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638392996714960978/5638392996714960978_signature_28.png | MARIANNE BOJADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5638392996714960978/5638392996714960978_signature_31.png
| 2023-06-29 | 2023-06-29 14:41 | Hi-Precision Diagnostics | N.S. Amroanto St., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24999 | CONSORTIUM | jose@hi-precision.com.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 23 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4899 | 2023-03-13 | GX IV 6C | 835602 | For XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. A1: Failed A2: New A3: Waiting for Replacement A4: Passed |  | XpertCheck kit | PL 010369 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638319719411880156/5638319719411880156_signature_28.png | Joselito Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5638319719411880156/5638319719411880156_signature_31.png
| 2023-06-29 | 2023-06-29 14:13 | Bureau of Plant Industry | 692 San Andres Street, Malate, Manila | NCR | NCR 1st - Manila | Central Laboratory | YES | INSTALLATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 24998 | COMMERCIAL | laurie_asis@yahoo.com | 25 |  | 2023-06-26 | Stemi 305 | 3943010040 | None | None | Cleaning of Eyepieces and surfaces Installed purchased Axiocam 208 Testing of Axiocam 208 with the router Performed functionality testing | Unit is operational | Left gooseneck light is loose | Axiocam 208 SN:5321004523 | SI 95796 and DR 10044 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638295586918847057/5638295586918847057_signature_28.png | Laurie Salumbides | https://www.jotform.com/uploads/macarewecare/222927978372067/5638295586918847057/5638295586918847057_signature_31.png
| 2023-06-29 | 2023-06-29 13:40 | Bureau of Plant Industry | 692 San Andres Street, Malate, Manila | NCR | NCR 1st - Manila | Central Laboratory | YES | INSTALLATION MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 24997 | COMMERCIAL | laurie_asis@yahoo.com | 25 |  | 2023-06-26 | Primo Star | 3116003729 | None | Installed Blue Filter | Cleaning of Eyepieces, Objectives, and Surfaces Performed Preventive Maintenance Tested Axiocam 208 Color with the Router Installed new power adapter Performed functionality tests | Unit is Operational | Missing one rubber stand underneath the base | Blue Filter from PN: 415500-1804-000 Power Adapter PN: 000000-0670-104 | SI 95796 and DR 10044 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638293366913685874/5638293366913685874_signature_28.png | Laurie Salumbides | https://www.jotform.com/uploads/macarewecare/222927978372067/5638293366913685874/5638293366913685874_signature_31.png
| 2023-06-29 | 2023-06-29 12:49 | Kasibu Rural Health Unit | Poblacion, Kasibu, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24996 | PBSP | mhokasibulaboratory@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 416 | 2023-06-29 | Genexpert II | 834802 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Archived Results Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 3pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638258661484717506/5638258661484717506_signature_28.png | Irene Dario | https://www.jotform.com/uploads/macarewecare/222927978372067/5638258661484717506/5638258661484717506_signature_31.png
| 2023-06-29 | 2023-06-29 12:23 | Philippine Orthopedic Center | 78-A Maria Clara St., Sta. Mesa Heights | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24995 | PBSP | cmblanceta@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 799 | 2023-06-29 | GXIV - 6C | 803088 | Module B1 for replacement Modules B2, B3, B4 for XpertCheck | -Checked and cleaned cartridge bays, plunger rod, Icore slots and filter -Replaced module 1 (724964 -> 210092883) -Performed XpertCheck on Modules B2, B3,B4) -Generated report | Replaced Module B1 Performed XpertCheck | Unit is operational. B1: Newly Replaced B2: Failed B3: Passed B4: Passed |  | GX Module 6C | G 5706 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638239389414318140/5638239389414318140_signature_28.png | CLAIRE MARIE B. LANCETA | https://www.jotform.com/uploads/macarewecare/222927978372067/5638239389414318140/5638239389414318140_signature_31.png
| 2023-06-29 | 2023-06-29 12:00 | Poro Rhu | Poblacion Poro, Camotes | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 24994 | DOH | avilacharmaine6@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 376 |  | Gx4 | 838279 | 1. UPS low output voltage 181 to 191 Vac 2. Module 2 SN # 808037 Error 4012 last March 2023 and Error 1001 run time error last May 2023, May 2023 Module 3 encounter error 2008 | Observed UPS 24 hours charging, direct avr to power outlet. Instruct end-user to observed the samples quality to avoid viscous samples | Performed troubleshooting of UPS low output voltage 181 Vac Clean Icore, plunger maintenance, clean filter and cartridge bay Generate data IQR, Sys log and Assay, archive and back up database | All modules are operational but no ups attched to the unit. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638211548107375255/5638211548107375255_signature_28.png | Charmaine Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/5638211548107375255/5638211548107375255_signature_31.png
| 2023-06-29 | 2023-06-29 11:36 | San Sebastian District Hospital | Tuban St., Sablayan | 4B | Occidental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24993 | PBSP | anthonycastrormt@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2256 | 2023-06-29 | GXIV | 812791 | Unit for PM. For module replacement of module B3. | Performed PM. Performed module replacement on B3. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B3 (210091102) in replacement to defective B3 (652387) | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638195550521337920/5638195550521337920_signature_28.png | Anthony D. Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5638195550521337920/5638195550521337920_signature_31.png
| 2023-06-29 | 2023-06-29 11:37 | Talisay Rural Health Unit | Talisay, Batangas | 4A | Batangas | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24992 | DOH | Adettlourodriguez@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 434 |  | GXIV | 820201 | Module A4 for check up (temp. Errors) | #NAME? | #NAME? | Unit is operational. A4 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638194763229874970/5638194763229874970_signature_28.png | Adett Lou Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5638194763229874970/5638194763229874970_signature_31.png
| 2023-06-29 | 2023-06-29 10:40 | Santa Ignacia Rural Health Unit | Santa Ignacia Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24991 | DOH | jdi0973@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 373 |  | GXIV-D | 831909 | A2 error 2035 | Re-xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Performed calibration Gathered reports Performed functionality test | Unit is operational except A4, A2 passed calibration | Used 1 xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638180335011175743/5638180335011175743_signature_28.png | Jess D. Ibarra Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5638180335011175743/5638180335011175743_signature_31.png
| 2023-06-29 | 2023-06-29 10:42 | Abuyog District Hospital | Abuyog Leyte | 8 | Leyte | Tb-dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24990 | PBSP | adh.leyte@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 357 ASSAY: Xpert MTB/RIF Ultra, COUNT: 137 | 2023-06-29 | GXIV | 815274 | unit for Calibration | #NAME? | #NAME? | Unit is Operational |  | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638161204494384282/5638161204494384282_signature_28.png | Maryln C. Redoblado | https://www.jotform.com/uploads/macarewecare/222927978372067/5638161204494384282/5638161204494384282_signature_31.png
| 2023-06-29 | 2023-06-29 10:10 | Philippine Red Cross Zamboanga | Petit Baracks, Zamboanga City, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | Ups pulled out for replacement | HAEMONETICS | Gian Carlo Aizon |  |  |  |  | 24989 | COMMERCIAL | zamboangacity@redcross.org.ph | 0 |  | 2023-06-29 | MCS + | 01M256 | Ups malfunction | Replaced battery but still problem occurs. Ups pulled out for replacement | Replaced battery but still problem occurs. Ups pulled out for replacement | Ups is broken and must be replaced | For replacement, Serial Number of UPS JB0320055820 model Back-UPS RS 1500 brand is APC | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638151995917786037/5638151995917786037_signature_28.png | Nimfa A. Linao | https://www.jotform.com/uploads/macarewecare/222927978372067/5638151995917786037/5638151995917786037_signature_31.png
| 2023-06-29 | 2023-06-29 09:37 | Eversly child sanitarium | Jagobiao, mandaue city | 7 | Cebu | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24988 | PBSP | alimurungjane@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 267 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2813 |  | Gx 4 | 805270 | Module a3 640129, temperature error 1002 | Clean fan filter | Clean fan filter, soft copy of IQR and system log | Module a3 temperature error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638125495668501829/5638125495668501829_signature_28.png | Juanita v. Almurung | https://www.jotform.com/uploads/macarewecare/222927978372067/5638125495668501829/5638125495668501829_signature_31.png
| 2023-06-29 | 2023-06-29 09:31 | Bicol SLSNRL | BRTTH Compound, Rizal St., Legazpi City | 5 | Albay | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24987 | PBSP | rbalzaga.pbsp@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 585 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2052 ASSAY: Xpert HIV-1 Viral Load, COUNT: 186 | 2023-06-13 | GX4 | 807579 | For replacement of gateway board | Replaced gateway board. 807579>>751351 Retained 807579 SN Cleaned plunger rod, cartridge bays Generated reporteds | Performed gateway board replacement, preventive maintenance | Unit is operational |  | 751351 Gateway Board SN | PL-011515 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638124571411269252/5638124571411269252_signature_28.png | Richard B. Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5638124571411269252/5638124571411269252_signature_31.png
| 2023-06-29 | 2023-06-29 09:32 | Baguio General Hospital and Medical Center | Baguio City | CAR | Benguet | TB Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24986 | PBSP | keithtagudar@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 756 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2156 | 2023-06-29 | GX4 6C | 834685 | Due for Preventive Maintenance | Cleaned Cartridge bays, Plunger Rods and I-CORE Generated Reports | Performed preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5638120753819767588/5638120753819767588_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5638120753819767588/5638120753819767588_signature_31.png
| 2023-06-28 | 2023-06-28 12:59 | Kalinga Provincial Hospital | Bulanao, Tabuk City, Kalinga | CAR | Kalinga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 24985 | DOH | leombelgica139@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1288 | 2023-06-28 | Genexpert IV | 838466 | Module 2 and 4 for xpertcheck | Performed Xpertcheck on Module 2 and 4 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck Result = B2 - Passed & B1 - Service Required | Xpertcheck 2pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5637381489125019528/5637381489125019528_signature_28.png | Leo Mark Belgica | https://www.jotform.com/uploads/macarewecare/222927978372067/5637381489125019528/5637381489125019528_signature_31.png
| 2023-06-27 | 2023-06-26 14:16 | Kabacan Rural Health Unit | Kabacan, North Cotabato | 12 | Cotabato | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24984 | PBSP | yayo_kmsi@yahoo.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 834 |  | Gx-IV (2 modules) | 835182 | Genexpert CPU/Computer problem | Replace defective CPU SN: *CPH1119F04268* with new CPU SN:*CPH0622F38156* Import MTB-rif ultra | Perform Genexpert machine Preventive maintenance | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636898700314821099/5636898700314821099_signature_28.png | Joan A. Sangines | https://www.jotform.com/uploads/macarewecare/222927978372067/5636898700314821099/5636898700314821099_signature_31.png
| 2023-06-27 | 2023-06-26 14:16 | Kabacan Rural Health Unit | Kabacan, North Cotabato | 12 | Cotabato | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24983 | PBSP | yayo_kmsi@yahoo.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 834 |  | Gx-IV (2 modules) | 835182 | Genexpert CPU/Computer problem | Replace defective CPU SN: *CPH1119F04268* with new CPU SN:*CPH0622F38156* Import MTB-rif ultra | Perform Genexpert machine Preventive maintenance | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636898560318504524/5636898560318504524_signature_28.png | Joan A. Sangines | https://www.jotform.com/uploads/macarewecare/222927978372067/5636898560318504524/5636898560318504524_signature_31.png
| 2023-06-27 | 2023-06-26 14:16 | Kabacan Rural Health Unit | Kabacan, North Cotabato | 12 | Cotabato | Tb-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24982 | PBSP | yayo_kmsi@yahoo.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 834 |  | Gx-IV (2 modules) | 835182 | Genexpert CPU/Computer problem | Replace defective CPU SN: *CPH1119F04268* with new CPU SN:*CPH0622F38156* Import MTB-rif ultra | Perform Genexpert machine Preventive maintenance | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636898370313087259/5636898370313087259_signature_28.png | Joan A. Sangines | https://www.jotform.com/uploads/macarewecare/222927978372067/5636898370313087259/5636898370313087259_signature_31.png
| 2023-06-27 | 2023-06-27 15:34 | Medellin RHU | Poblacion, Medellin Cebu | 7 | Cebu | Tb Laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 24981 | DOH | medellindtc@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1149 |  | Gx4 | 838506 | Module 2 Sn 862245 encounter run time error 1002 several times and Module 4 Sn 633038 encounter error 1002 and 4009 multiple times. | Performed calibration in Module 2 and Module 4 Conduct PMS in genxpert machine Module 2 and 4 required for module replacement | Performed calibration in Module 2 and 4 Install expertcheck ver 1.5a Clean Icore, plunger and filter Generate data IQR, sys log, Assay Stats | Only 1 and module 3 are available due to run time error of Module 2 and 4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636622307113938229/5636622307113938229_signature_28.png | Raul M.. Galano | https://www.jotform.com/uploads/macarewecare/222927978372067/5636622307113938229/5636622307113938229_signature_31.png
| 2023-06-27 | 2023-06-27 15:08 | Wenceslao Dela Paz Health Center | SB Park, Tandang Sora, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24980 | DOH | ameliafclemente2001@yahoo.com | 29 |  |  | GXIV 6C | 831876 | Some modules experiencing error | -Checked and cleaned cartridge bays, plunger rod, icore slot and filter -Checked system log -Module 4 experienced error 2005 twice this week already -Will ask update if same error happens by friday -Generate report | Checked the status of modules | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636617533121077675/5636617533121077675_signature_28.png | Camille Joy Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5636617533121077675/5636617533121077675_signature_31.png
| 2023-06-27 | 2023-06-27 15:37 | Cebu City health office | Cebu city | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24979 | COMMERCIAL | jonnifercodoyceniza@gmail.com | 29 |  |  | Gx 4 | 829419 | Xpertcheck for module 3 | Xpertcheck for module 3 | Xpertcheck for module 3 | Systems opeartional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636608971125995974/5636608971125995974_signature_28.png | Jonnifer C. Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5636608971125995974/5636608971125995974_signature_31.png
| 2023-06-27 | 2023-06-27 15:12 | Cebu City Health Office | Cebu city | 7 | Cebu | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Raymond Laurito |  |  |  |  | 24978 | PBSP | laureanjo@yahoo.com | 28 |  |  | Gx 4 | 838882 | Defective module B1 & B4 | Self test B1 & B4, | Gathering data system log, IQR | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636593715812952035/5636593715812952035_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5636593715812952035/5636593715812952035_signature_31.png
| 2023-06-27 | 2023-06-27 13:25 | East Avenue Medical Center | East avenue,QC | NCR | NCR 2nd - Quezon City | Bloodbank | YES | REPAIR CALIBRATION | HAEMONETICS | Noel Floresca |  |  |  |  | 24977 | COMMERCIAL | bloodbankeamc@gmail.com | 29 |  | 2023-06-15 | MCS + | 15H044G | Defective CPU Board | Performed repair Delivered and replaced cpu board Check and adjust Calibration ,Centrifuge,DPM,SPM,Weigher,Line sensor,Checked parameters. Test run unit | Repair and calibration | Operational |  | CPU Board | 95730 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636570209227965048/5636570209227965048_signature_28.png | Jocelyn Rote | https://www.jotform.com/uploads/macarewecare/222927978372067/5636570209227965048/5636570209227965048_signature_31.png
| 2023-06-27 | 2023-06-27 14:14 | Abucay Rural Health Unit | Abucay Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24976 | DOH | parreraledelina22176@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 75 ASSAY: Xpert MTB/RIF Ultra, COUNT: 402 |  | GXIV-D | 831878 | A2 Temperature error | Re-xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Performed calibration Gathered reports Performed functionality test | Unit is operational except A1 and A3 | Used 1 xpertcheck kit, A2 failed calibration for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636563150691453724/5636563150691453724_signature_28.png | Mylene M. Forbes | https://www.jotform.com/uploads/macarewecare/222927978372067/5636563150691453724/5636563150691453724_signature_31.png
| 2023-06-27 | 2023-06-27 13:57 | Bagong Pag-Asa Health Center | Bagong Pag-asa, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24975 | DOH | erlindalorilla026@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1520 |  | GXIV 6C | 831597 | CPU not booting up | -Checked CPU -Replaced HDD; problem persist -Replaced RAM; problem solved -Checked and cleaned cartridge bays, plunger rod, icore slot and filter -reversed cable connection; problem persists -Generated report | Troubleshoot CPU | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636559793127436603/5636559793127436603_signature_28.png | ERLINDA LORILLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5636559793127436603/5636559793127436603_signature_31.png
| 2023-06-27 | 2023-06-27 14:07 | Bagong Pag-asa Health Center | Bagong Pag-asa, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24974 | DOH | erlindalorilla026@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1520 |  | GXIV 6C | 831597 | White line in monitor | -Checked and cleaned cartridge bays, plunger rod, icore slot and filter -Troubleshoot monitor -reversed VGA cable connection; problem persist -replaced VGA; problem solved -Generated report | Troubleshoot monitor | Unit is operational. Module 2 under monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636557973121107676/5636557973121107676_signature_28.png | ERLINDA LORILLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5636557973121107676/5636557973121107676_signature_31.png
| 2023-06-27 | 2023-06-27 14:11 | Mandaluyong City Health Office | Maysilo, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24973 | DOH | rsjalova1974@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2475 ASSAY: Xpert HIV-1 Viral Load, COUNT: 130 |  | GXIV | 831596 | Unit for maintenance | #NAME? | Performed maintenance | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636556378511061123/5636556378511061123_signature_28.png | Revena Sibayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5636556378511061123/5636556378511061123_signature_31.png
| 2023-06-27 | 2023-06-27 14:03 | Mandaluyong City Health Office | Maysilo, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24972 | PBSP | rsjalova1974@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2794 ASSAY: Xpert HIV-1 Viral Load, COUNT: 149 |  | GXIV | 805062 | Module A3 stuck cartridge | #NAME? | Performed troubleshooting | Unit is operational except A3. Module stuck on booting. Unable to generate IQR. Module A3 (698345) for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636554188519818065/5636554188519818065_signature_28.png | Revena Sibayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5636554188519818065/5636554188519818065_signature_31.png
| 2023-06-27 | 2023-06-27 12:16 | Bataan General Hospital | Balanga Bataan | 3 | Bataan | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24971 | PBSP | honeyeichel@gmail.com | 28 | ASSAY: Xpert SARS-CoV-2, COUNT: 13510 |  | GXIV-D | 829390/834433 | B1 error 4009,1002 B2 error 4009 B3 error 4009 D3 error 1001,1002 UPS Battery error ( not charging, drained battery) | Troubleshooting Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Gathered reports Performed functionality test | Unit is operational excep B1,B2,B3, D3 and UPS | For reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636552230697892364/5636552230697892364_signature_28.png | Honey Eichel M. Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5636552230697892364/5636552230697892364_signature_31.png
| 2023-06-27 | 2023-06-27 13:21 | Tacloban City Health Office | Tacloban City | 8 | Leyte | Tb dots | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24970 | DOH | ammegzedd@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 240 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1065 | 2023-06-27 | GXIV | 838421 | Defective LCD monitor (SN: CNK9430B0R) | #NAME? | #NAME? | Unit is Operational |  | LCD monitor (Model: AOC sn/ROXNB1A004636) | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636532874119506607/5636532874119506607_signature_28.png | Gemma malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5636532874119506607/5636532874119506607_signature_31.png
| 2023-06-27 | 2023-06-27 10:51 | Masambong Health Center | Capoas, Masambong, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24969 | PBSP | michellebonrostro@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 973 | 2023-06-27 | GXIV 6C | 814160 | Module B3 and B4 for replacement | -Checked and cleaned cartridge bays, plunger rod, icore slot and filter -Replaced module B3 and B4 -Checked module 2 system log, error 4009 -Helped move GX machine to bigger room -Generated report | Replaced Module B3 and B4 B3: 690602 -> 210059649 B4: 690604 -> 210063218 | Unit is operational |  | GX Modules | G 5697 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636466796612954240/5636466796612954240_signature_28.png | Michelle D. Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/5636466796612954240/5636466796612954240_signature_31.png
| 2023-06-27 | 2023-06-27 11:04 | Davao del Norte Provincial Hospital-Igacos Samal | Datu Taganiog St.,Brgy. Peñaplata Island Garden City of Samal | 11 | Davao del Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24968 | PBSP | labddnhigacoszone@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 960 |  | GX-IV 4D | 815383 | For Preventive Maintenance and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform Xpertcheck Calibration | Perform Preventive Maintenance and XpertCheck Calibration | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636462516227957865/5636462516227957865_signature_28.png | Janice Fe Pastor | https://www.jotform.com/uploads/macarewecare/222927978372067/5636462516227957865/5636462516227957865_signature_31.png
| 2023-06-27 | 2023-06-27 11:27 | Lianga Rural Health Unit | Purok 3, Poblacion, Lianga | 13 | Surigao del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 24967 | PBSP | gyvxs@yahoo.com | 28.15 | ASSAY: Xpert MTB/RIF, COUNT: 59 ASSAY: Xpert MTB/RIF Ultra, COUNT: 511 |  | GX-4 | 835174 | None | None | Replacing Module B1 with SN: 669053 to SN: 210068498 | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636460964813497012/5636460964813497012_signature_28.png | Gyver I. Sistona | https://www.jotform.com/uploads/macarewecare/222927978372067/5636460964813497012/5636460964813497012_signature_31.png
| 2023-06-27 | 2023-06-27 11:28 | Lasam District Hospital | Ignacio B Jorado, Lasam, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24966 | DOH | lasamdh.laboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 51 ASSAY: Xpert MTB/RIF Ultra, COUNT: 311 | 2023-02-14 | Genexpert IV | 838530 | Module 4 Error 4012 | Performed Xpertcheck | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = B1, B2, B3 - Passed & B4 - Service Required | Xpertcheck 4pc | 9624 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636460427184502586/5636460427184502586_signature_28.png | Verbell Manuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5636460427184502586/5636460427184502586_signature_31.png
| 2023-06-27 | 2023-06-27 11:19 | Potrero Health Center | Durian St. Malabon | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24965 | PBSP | ailyn.santos19@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2029 | 2023-06-27 | GXIV | 804715 | Unit for PM. For module replacement of B2. | Performed PM. Performed module replacment on B2. Generated reports. | PM and module replacment. | Unit is operational. For repirting of modules B1 and B4. |  | Gx module B2 (210059647) in replacement to defective B2 (672907). | G5698 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636454897214036606/5636454897214036606_signature_28.png | Ailyn M.Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5636454897214036606/5636454897214036606_signature_31.png
| 2023-06-27 | 2023-06-27 10:57 | Lubao Rural Health Unit | Lubao Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24964 | PBSP | nathaliecoleencovar@gmail.com | 31 | ASSAY: Xpert MTB/RIF, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 922 |  | GXIV-D | 812834 | B1 stuck cartridge | Preventive maintenance Cartridge removal | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Removed stuck cartridge Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5636442173697859374/5636442173697859374_signature_28.png | Nathalie Covar | https://www.jotform.com/uploads/macarewecare/222927978372067/5636442173697859374/5636442173697859374_signature_31.png
| 2023-06-26 | 2023-06-26 17:47 | Cagayan Valley Center for Health Development (DOH R02) | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24963 | PBSP | mctorio.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1111 ASSAY: Xpert SARS-CoV-2, COUNT: 2576 | 2023-06-26 | Genexpert IV | 834434 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635822339312472908/5635822339312472908_signature_28.png | Mikeala Anne C. Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5635822339312472908/5635822339312472908_signature_31.png
| 2023-06-26 | 2023-06-26 17:42 | Cagayan Valley Center for Health Development (DOH R02) | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24962 | PBSP | mctorio.pbsp@gmail.com | 25 |  | 2023-06-26 | Genexpert IV 10c | 120000178 | N/a | N/a | Performed Maintenance Gathered Reports Installed Drivers Installed Barcode Scanner Installed Genexpert Installed PC Installed Printer Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635819939314848721/5635819939314848721_signature_28.png | Mikeala Anne C. Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5635819939314848721/5635819939314848721_signature_31.png
| 2023-06-26 | 2023-06-26 17:32 | Cagayan Valley Center for Health Development (DOH R02) | Carig Norte, Tuguegarao City Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24961 | PBSP | mctorio.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1111 ASSAY: Xpert SARS-CoV-2, COUNT: 2576 | 2023-06-26 | Genexpert IV | 808628 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 665783 & 628109 - 210063032 & 210063113 | Genexpert Module = 210063032 & 210063113 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635817049318468707/5635817049318468707_signature_28.png | Mikeala Anne C. Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5635817049318468707/5635817049318468707_signature_31.png
| 2023-06-26 | 2023-06-26 17:27 | Cagayan Valley Center for Health Development(DOH R02) | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE Replaced Module Door | CEPHEID | Matthew Roxas |  |  |  |  | 24960 | PBSP | mctorio.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2217 ASSAY: Xpert SARS-CoV-2, COUNT: 6457 | 2023-06-26 | Genexpert XVI | 842894 | Module D3 Hardware Failure, D2 Module Door | Replaced Module Door of D2 from Defective Module | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 888973 - 210061145 | Genexpert Module - 210061145 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635811389315954518/5635811389315954518_signature_28.png | Mikeala Anne Torio | https://www.jotform.com/uploads/macarewecare/222927978372067/5635811389315954518/5635811389315954518_signature_31.png
| 2023-06-26 | 2023-06-26 16:46 | Nqbua Rural Health Unit | Zone 1, San Roque, Poblacion, Nabua, Camarines Sur | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 24959 | PBSP | fevalenzuela1965@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1050 |  | GX4 | 834620 | CPU delayed powered on | Cleaned ram, checked power cord, checked connections Cleaned cartridge bays, plunger rod, fan filter Collected reports | Troubleshooting CPU, Performed Preventive maintenance | Unit is operational | CPU For replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635789821987953942/5635789821987953942_signature_28.png | Valenzuela, Fe L. | https://www.jotform.com/uploads/macarewecare/222927978372067/5635789821987953942/5635789821987953942_signature_31.png
| 2023-06-26 | 2023-06-26 14:46 | Department of Health - Northern Mindanao TB Reference Laboratory | JB Serina St., Barangay Carmen, Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24958 | DOH | rfreneefleur184@gmail.com | 26 |  | 2023-06-09 | GeneXpert | 120000170 | N/A | N/A | GeneXpert Machine Installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635717800954290945/5635717800954290945_signature_28.png | Renee Fleur Clutario | https://www.jotform.com/uploads/macarewecare/222927978372067/5635717800954290945/5635717800954290945_signature_31.png
| 2023-06-26 | 2023-06-26 13:31 | Alcala Municipal Health Office | Centro Norte | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24957 | DOH | rhualcala@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 472 | 2023-02-14 | Genexpert IV | 838531 | Module 2 Hardware Failure | Performed Xpertcheck | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = B1, B3, B4 - Passed & B2 - Service Required | Xpertcheck 4pc | 9628 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635694889319020900/5635694889319020900_signature_28.png | Clarisa Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/5635694889319020900/5635694889319020900_signature_31.png
| 2023-06-26 | 2023-06-23 12:49 | Dasmariñas City Health Office II | DBB-E, Dasmariñas | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24956 | DOH | hazelmilan0777@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2441 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 |  | GXIV-D | 838535 | Modules 1 and 4 with Temperature and Valve Error | XpertCheck | Performed Maintenance Performed XpertCheck on Modules 1 and 4 Performed Self-Test and Functionality Test Generate Reports | 2 Modules are Operational, 2 Modules for Replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684183614843390/5635684183614843390_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684183614843390/5635684183614843390_signature_31.png
| 2023-06-26 | 2023-06-21 10:03 | Bagong Brgy Health Center | Zamora, Pandacan | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24954 | DOH | thea.peneyra@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1028 |  | GXIV-D | 831694 | Module 1 with Temperature Error | XpertCheck | Performed Maintenance Perform XpertCheck in Module 1 Generate Reports | 1 Module is Operational, 3 Modules are Defective |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063611897215/5635684063611897215_signature_28.png | Maria Anthea S Peneyra | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063611897215/5635684063611897215_signature_31.png
| 2023-06-26 | 2023-06-21 07:54 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24952 | MACARE-HBDC | fabella.mollabgene@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 3303 | 2023-06-20 | GXIV-D-10C | 110015482 | Modules 1 and 3 have Temperature Related Errors | Module Replacement | Replaced Modules 1 and 3 Performed Maintenance Performed Self-Test and Functionality Generate Reports | All Modules are Operational | 210078318-> 210010209, 210078487-> 210031286 | GX Module Kit 10-C | DR 10030 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063612188961/5635684063612188961_signature_28.png | CHRISTIANNE BERNADETTE BAUTISTA | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063612188961/5635684063612188961_signature_31.png
| 2023-06-26 | 2023-06-23 10:45 | Imus City Health Office II | Bayan Luma, Imus | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24951 | DOH | amielynnmangulabnan1217@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1163 |  | GXIV-D | 821284 | Modules 3 with 5000 Series Error | XpertCheck | Performed Maintenance Performed XpertCheck on Module 3 Generate Reports Performed Functionality Test | 2 Modules are Operational, 2 for Replacement | Module 3 Failed in XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063612724426/5635684063612724426_signature_28.png | Amielynn B Mangulabnan | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063612724426/5635684063612724426_signature_31.png
| 2023-06-26 | 2023-06-22 12:51 | Molino III City Health Unit | Molino, Bacoor | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24953 | DOH | joanrnd@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1228 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 |  | GXIV-D | 821221 | UPS has No Backup Function | Battery Replacement | Replaced UPS Battery Replaced BIOS Battery Performed Maintenance Generate Reports | 3 Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063614159978/5635684063614159978_signature_28.png | Jonalyn Aranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063614159978/5635684063614159978_signature_31.png
| 2023-06-26 | 2023-06-21 08:16 | Dr. Jose Fabella Memorial Hospital | Tayuman St., Manila | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24955 | MACARE-HBDC | fabella.mollabgene@gmail.com | 27 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 5414 ASSAY: Others, COUNT: 10 |  | GXIV-L | 834953 | Module 1 with Temperature Error | XpertCheck Maintenance | Performed Maintenance Performed XpertCheck on Reported Error Generate Reports | 3 Modules are Operational, 1 Defective | Used Spare Cartridge |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063616609613/5635684063616609613_signature_28.png | Christianne Bernadette Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5635684063616609613/5635684063616609613_signature_31.png
| 2023-06-26 | 2023-06-26 13:16 | Baao Rural Health Unit | Imerial St., Del Rosario, Baao, Camarines Sur | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24950 | PBSP | briciarandy@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1516 | 2023-03-15 | GX4 | 808639 | For XpertCheck | Performed XpertCheck Cleaned cartridge bays, plunger rod, fan filter Generated reports | Performed XpertCheck & Preventive Maintenance | Unit is Operational | All passed for XpertCheck | Xpertcheck kit | PL-010397 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635672841985507736/5635672841985507736_signature_28.png | Randy Bricia | https://www.jotform.com/uploads/macarewecare/222927978372067/5635672841985507736/5635672841985507736_signature_31.png
| 2023-06-26 | 2023-04-20 09:17 | Jose Dalman Rural Health Unit | Poblacion Jose Dalman, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24949 | PBSP | ferranbryder@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 345 | 2023-04-20 | GX 4 | 834827 | Module 2 h/w failed | Module 787281 replaced with Module 606719 | Module Replacement and Preventive Maintenence | All 2 modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635622738411108819/5635622738411108819_signature_28.png | Ferrari Bryder Tomoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5635622738411108819/5635622738411108819_signature_31.png
| 2023-06-26 | 2023-06-26 11:48 | Carcar City Health Office | Valladolid Carcar Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 24948 | DOH | chaska.lp14@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1314 |  | Gx4 | 838283 | Error 2008 in B3 and Error 5006 in B4 | Clean Icore, Plunger and clean filter Instruct medtech to add extra time for viscous sputum | Performed troubleshooting and analysis | Only B1, B3, and B4 are used due to recurrent Run time error for module 2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5635608353455175825/5635608353455175825_signature_28.png | Lucille Mae A. Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5635608353455175825/5635608353455175825_signature_31.png
| 2023-06-25 | 2023-06-22 09:58 | Umingan Community Hospital | Burgos St. Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24947 | DOH | jeffreyvrumbaoa@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 349 ASSAY: Xpert MTB/RIF Ultra, COUNT: 921 |  | GX 4 | 838473 | Module B3 | #NAME? | Perform ReXpertCheck and Maintenance | Unit is operational | Passed ReXpertCheck. Module B3 for Monitoring |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5634886957464539861/5634886957464539861_signature_28.png | Jeffrey V. Rumbaoa | https://www.jotform.com/uploads/macarewecare/222927978372067/5634886957464539861/5634886957464539861_signature_31.png
| 2023-06-25 | 2023-06-22 09:58 | Umingan Community Hospital | Burgos St. Umingan, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24946 | DOH | jeffreyvrumbaoa@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 349 ASSAY: Xpert MTB/RIF Ultra, COUNT: 921 |  | GX 4 | 838473 | Module B3 | #NAME? | Perform ReXpertCheck and Maintenance | Unit is operational | Passed ReXpertCheck. Module B3 for Monitoring |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5634886597464304498/5634886597464304498_signature_28.png | Jeffrey V. Rumbaoa | https://www.jotform.com/uploads/macarewecare/222927978372067/5634886597464304498/5634886597464304498_signature_31.png
| 2023-06-25 | 2023-06-22 13:15 | Rosales RHU | Rosales | 1 | Pangasinan | TB DOTS | NO | REPAIR TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24945 | DOH | allan.vinluan2013@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 247 ASSAY: Xpert MTB/RIF Ultra, COUNT: 446 |  | GX4 | 838472 | Module B2 for troubleshooting | #NAME? | For ReXpertCheck | Unit is operational | Module B2 passed ReXpertCheck. For Monitoring |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5634885417461943102/5634885417461943102_signature_28.png | Allan Vinlluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5634885417461943102/5634885417461943102_signature_31.png
| 2023-06-25 | 2023-06-22 13:15 | Rosales RHU | Rosales | 1 | Pangasinan | TB DOTS | NO | REPAIR TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24944 | DOH | allan.vinluan2013@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 247 ASSAY: Xpert MTB/RIF Ultra, COUNT: 446 |  | GX4 | 838472 | Module B2 for troubleshooting | #NAME? | For ReXpertCheck | Unit is operational | Module B2 passed ReXpertCheck. For Monitoring |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5634884997466707617/5634884997466707617_signature_28.png | Allan Vinlluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5634884997466707617/5634884997466707617_signature_31.png
| 2023-06-25 | 2023-06-23 09:28 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24943 | PBSP | r1mc.molecularlab@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1529 ASSAY: Xpert SARS-CoV-2, COUNT: 3406 |  | GX 16 | 842230 | Module B1 for Troubleshooting | Performed Reliability test on module B1 | Troubleshoot module B1 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5634876697463011352/5634876697463011352_signature_28.png | Mark Noni Pimentel | https://www.jotform.com/uploads/macarewecare/222927978372067/5634876697463011352/5634876697463011352_signature_31.png
| 2023-06-24 | 2023-06-24 09:48 | St.lukes medical center QC | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 24942 | COMMERCIAL | biomed.qc@stlukes.com.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 260 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 171 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2934 ASSAY: Xpert HCV VL, COUNT: 5 |  | GX4 D | 807155 | Due for maintenance and module 4 210072753 encounter continous loss communication error | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit. Checked and reset gateway board Replace all cable Checked powersupply voltage Swap module position module 4 and module 3 Performed reliability | Maintenance and Troubleshooting | For observation | Swa0 module location module 4 and module 3 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5634188370028944903/5634188370028944903_signature_28.png | Ma. Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5634188370028944903/5634188370028944903_signature_31.png
| 2023-06-23 | 2023-06-23 15:44 | Sumilang Super Health Center | Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24941 | PBSP | maricor.obnamia05@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 150 ASSAY: Xpert MTB/RIF Ultra, COUNT: 240 | 2023-06-23 | GX IV 6C | 814163 | CPU not booting up | #NAME? | Installed bak up laptop temporarily Troubleshoot printer | Unit is operational. |  | Lenovo Thinkpad CPH1019F03493 | G 5692 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633162835381529850/5633162835381529850_signature_28.png | Maria Corazon obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/5633162835381529850/5633162835381529850_signature_31.png
| 2023-06-23 | 2023-06-20 13:55 | Talavera Rural Health Unit II | San Ricardo, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24940 | PBSP | nelzon1966@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 545 | 2023-06-20 | Genexpert IV | 812831 | Module 4, HW Failure UPS not Working | Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Calibration (B2, B3 = Passed & B1 = Service Required) Performed Functionality Test | Unit is Operational | UPS for Replacement | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633156623614600946/5633156623614600946_signature_28.png | Nelson Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5633156623614600946/5633156623614600946_signature_31.png
| 2023-06-23 | 2023-06-16 15:11 | Cabanatuan City Health Office | Kapt. Pepe, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24938 | DOH | cholab2021@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 200 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2371 | 2023-06-16 | Genexpert IV | 831592 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633156623615903250/5633156623615903250_signature_28.png | Lea Samson | https://www.jotform.com/uploads/macarewecare/222927978372067/5633156623615903250/5633156623615903250_signature_31.png
| 2023-06-23 | 2023-06-23 13:35 | Dr. Emigdio C. Cruz Sr Memorial Hospital | San Agustin, Arayat, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24939 | PBSP | leilani101683@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 469 | 2023-06-23 | Genexpert IV | 812790 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 5pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633156623616703056/5633156623616703056_signature_28.png | Leilani Marie Apostol | https://www.jotform.com/uploads/macarewecare/222927978372067/5633156623616703056/5633156623616703056_signature_31.png
| 2023-06-23 | 2023-06-23 14:17 | Escalante City Health Office | Poblacion Escalante City, Negros Occidental | 6 | Negros Occidental | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24937 | DOH | sheila.negrido@yahoo.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1007 |  | Gx4 | 838523 | Due for Calibration and Maintenance Ambient Temperature in the area too high 35 degrees. | Performed maintenance and clean up Instruct enduser to provide aircon ASAP ( as soon as possible) to prevent overheating of all modules during operation | Performed maintenance, clean Icore and plunger, clean filter Check ups voltage input and output | All modules are beyond 30 degrees celcius Recommended for additional aircon due to high temp in the module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633107031413222166/5633107031413222166_signature_28.png | Sheila Mae Negrido | https://www.jotform.com/uploads/macarewecare/222927978372067/5633107031413222166/5633107031413222166_signature_31.png
| 2023-06-23 | 2023-06-21 13:50 | Panabo City Health Office | Poblacion, Panabo City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24936 | DOH | nadineeust67@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2302 |  | GX-IV 4D | 838295 | For Preventive Maintenance Barcode Reader error | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Reset the Barcode Reader Connection and try to change the port | Perform Preventive Maintenance Troubleshoot the barcode reader | Unit has only 1 Module operational Barcode Reader error is fixed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633087408113504872/5633087408113504872_signature_28.png | Nadine Gemma Eustaquio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5633087408113504872/5633087408113504872_signature_31.png
| 2023-06-23 | 2023-06-23 13:02 | Davao Chest Center | Villa Abrille Street, Davao City | 11 | Davao del Sur | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24935 | PBSP | dcclabdept@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2262 |  | GX-IV 4D | 804835 | For Calibration and Preventive | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration | Perform Preventive Maintenance and XpertCheck Calibration | Module A1 retry and for service required Module A2 for service required Modules A3 and A4 retry and skipped | Data will be sent to Cepheid to further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633087008116351701/5633087008116351701_signature_28.png | CRIZZ VALEREY B. BAUTISTA | https://www.jotform.com/uploads/macarewecare/222927978372067/5633087008116351701/5633087008116351701_signature_31.png
| 2023-06-23 | 2023-06-23 13:15 | Kolambugan Rural Health Unit | Kolambugan | 10 | Lanao del Norte | Laboratory | YES | REPAIR TROUBLESHOOTING | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24934 | PBSP | yojecinaj@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 607 |  | GeneXpert | 815272 | System Unit Booting issue | Replaced CMOS battery Adjust Booting options | Preventive Maintenance Generate Report files Database back up | Operatiknal |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633067827402376305/5633067827402376305_signature_28.png | Janice Joy R. Siso | https://www.jotform.com/uploads/macarewecare/222927978372067/5633067827402376305/5633067827402376305_signature_31.png
| 2023-06-23 | 2023-06-23 12:43 | BAYAWAN CITY HEALTH OFFICE | T. Diao St., Bayawan City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 24933 | PBSP | nymphar07@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 469 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1629 |  | GX-4 | 808282 | Defective module 619532/B3 | #NAME? | - Replaced module B3 with new SN 210068643 - Generate Pre and Post IQR - Replaced printer toner | All 4 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633049956444399084/5633049956444399084_signature_28.png | NYMPHA RIDAD | https://www.jotform.com/uploads/macarewecare/222927978372067/5633049956444399084/5633049956444399084_signature_31.png
| 2023-06-23 | 2023-06-23 12:12 | Central Luzon Doctors Hospital | Hospital Drive San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24931 | COMMERCIAL | rohmhardameg@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 107 ASSAY: Xpert SARS-CoV-2, COUNT: 2747 |  | GXIV-D | 839232 | Due for maintenance | Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Gathered reports Performed functionality test | Unit is operational | B4 under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633034583231701576/5633034583231701576_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/5633034583231701576/5633034583231701576_signature_31.png
| 2023-06-23 | 2023-06-23 12:16 | Central Luzon Doctors Hospital | Hospital Drive San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24932 | COMMERCIAL | rohmhardameg@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 7 ASSAY: Xpert SARS-CoV-2, COUNT: 336 |  | GXII-D | 830841 | Due for maintenance | Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633034583233331165/5633034583233331165_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/5633034583233331165/5633034583233331165_signature_31.png
| 2023-06-23 | 2023-06-23 10:42 | Gerona Rural Health Umit | Brgy. Pinasling Gerona Tarlac | 3 | Tarlac | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24930 | DOH | huwakie15@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 436 |  | GXIV-D | 831693 | A1 error 4004 | Troubleshooting | Check valve position Performed self test Performed valbe adjustment Gathered reports Performed functionality test | Unit is operationap except A1 and A3 | For reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633034583236402928/5633034583236402928_signature_28.png | Artemio Nool III | https://www.jotform.com/uploads/macarewecare/222927978372067/5633034583236402928/5633034583236402928_signature_31.png
| 2023-06-23 | 2023-06-23 11:11 | Manila Doctors Hospital | 667 U.N. Ave., Ermita, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24929 | COMMERCIAL | lab.bacteriology@maniladoctors.com.ph | 24 | ASSAY: Xpert CT/NG, COUNT: 90 ASSAY: Xpert HBV VL, COUNT: 55 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1013 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 ASSAY: Others, COUNT: 9 | 2023-06-20 | GX IV 10C | 110001648 | For XpertCheck | #NAME? | Performed XpertCheck | Unit is operational | All Modilules Passed | Xpertcheck kit | 1000397715 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5633022657118608565/5633022657118608565_signature_28.png | Dyan Valois | https://www.jotform.com/uploads/macarewecare/222927978372067/5633022657118608565/5633022657118608565_signature_31.png
| 2023-06-23 | 2023-06-23 10:33 | Bulacan Medical Center | Mojon St. Malolos | 3 | Bulacan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24928 | PBSP | bmcmolecularlab@yahoo.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 3576 | 2023-06-22 | GXIV | 884623 | Unit for PM. For module replacement of module C4. | Performed PM. Performed module replacement on C4. Generated reports. | PM and module replacement. | Unit is operational. |  | GX Module C4 (210091108) in replacement to defective C4 (784871). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632971509071788925/5632971509071788925_signature_28.png | Roxan Resare | https://www.jotform.com/uploads/macarewecare/222927978372067/5632971509071788925/5632971509071788925_signature_31.png
| 2023-06-23 | 2023-06-23 10:21 | Sagay City Health Office | Sagay, Negros Occidental | 6 | Negros Occidental | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24927 | PBSP | maravelesdebra@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 675 |  | Gx4 | 838144 | Module 1 Hardware Failure | Replaced new module Generate data IQR, Sys log and Gxx files Recommended enduser to maintain the ambient temperature in the area | Replaced module 1 with new module SN 210091070 | The units are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632965009118863363/5632965009118863363_signature_28.png | Debora Abellar, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5632965009118863363/5632965009118863363_signature_31.png
| 2023-06-23 | 2023-06-23 09:49 | Talisay City Health Office | Talisay City Negros Occidental | 6 | Negros Occidental | Tb Dots | YES | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 24926 | PBSP | hadmeoffguard@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | Gx4 | 847222 | UPS discharged automatically Liebert GXT MTS. UPS Battery defective. | Troubleshoot ups | Performed troubleshooting check input voltage 230 and output voltge zero volts | The UPS cannot be charge. The machine is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632945673342704141/5632945673342704141_signature_28.png | Hadney Besa | https://www.jotform.com/uploads/macarewecare/222927978372067/5632945673342704141/5632945673342704141_signature_31.png
| 2023-06-23 | 2023-06-23 00:42 | Philippine Heart Center | East Ave., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR | HAEMONETICS | Harold Toco |  |  |  |  | 24925 | COMMERCIAL | phc_bb@yahoo.com | 25 |  | 2023-06-23 | MCS+ | 02D168 | Valve out of position | Checked valve functionality.. | Checked pinch valve connections.. Performed valve toggle functionality.. Run simulatoin test.. | Unit is operational.. |  | None | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632620738118704371/5632620738118704371_signature_28.png | PHC Blood Bank | https://www.jotform.com/uploads/macarewecare/222927978372067/5632620738118704371/5632620738118704371_signature_31.png
| 2023-06-22 | 2023-06-21 09:10 | Rodriguez rhu | J.P Rizal st,Montalban | 4A | Rizal | Tb dots | NO | Checked up | CEPHEID | Noel Floresca |  |  |  |  | 24922 | PBSP | karlbanal25@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2485 ASSAY: Xpert HIV-1 Viral Load, COUNT: 95 |  | GX4 | 838894 | B2 encountered 2008 error | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Gathered reports | B2 encounter 2008 error, B1,B2,B4 operati9nal |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632512378923138557/5632512378923138557_signature_28.png | Karl Michael Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5632512378923138557/5632512378923138557_signature_31.png
| 2023-06-22 | 2023-06-22 15:24 | National children's hospital | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24924 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2983 ASSAY: Xpert SARS-CoV-2, COUNT: 7856 ASSAY: Others, COUNT: 18 | 2023-06-20 | GXXVI 10c | 110014171 | For module replacement D3 210069281 - 210025172 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Maintenance and module replacement | A1,b4 reported, 14 modules operational |  | 900-0650rs | 10029 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632512378923659533/5632512378923659533_signature_28.png | National children's hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/5632512378923659533/5632512378923659533_signature_31.png
| 2023-06-22 | 2023-06-21 09:31 | Rodriguez rhu | J.P Rizal, Montalban | 4A | Rizal | Tb dots | NO | Checked up | CEPHEID | Noel Floresca |  |  |  |  | 24923 | PBSP | karlbanal25@gmail.com | 35 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2249 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 |  | GX4 | 807580 | UPS unable to backup power supply | Checked ups | Check up | Unit operational, ups unable backup power supply | UPS for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632512378924191149/5632512378924191149_signature_28.png | Karl Michael Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5632512378924191149/5632512378924191149_signature_31.png
| 2023-06-22 | 2023-06-22 14:34 | Josefine Belmonte Duran Memorial District Hospital | Purok 2, Tuburan, Ligao City, Albay | 5 | Albay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24921 | PBSP | johnrayds.22@gmail.com | 27 |  |  | GX4 | 814371 | For xpertcheck | Performed xpertcheck Cleaned cartridge bays, plunger rods, fan filter Generated reports | Performed XpertCheck, Preventive Maintenence | Unit is operational B1(890331) -service required B2(210068727)-service required B4(693979)-service required | Xpertcheck B1(890331) -service required B2(210068727)-service required B4(693979)-service required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632325990323877876/5632325990323877876_signature_28.png | JOHNRAY DE LOS SANTOS | https://www.jotform.com/uploads/macarewecare/222927978372067/5632325990323877876/5632325990323877876_signature_31.png
| 2023-06-22 | 2023-06-22 15:38 | M. Icasiano Health Center | Pedro Gil, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24920 | DOH | acg_medtech@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1301 |  | GXIV 6C | 831684 | CPU not booting up | -Checked and cleaned cartridge bays, icore slot, plunger tod and filter -CPU not booting up -Replaced HDD; problem persists -Replaced RAM; problem solved -Generated report | Troubleshoot CPU | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632298430513462213/5632298430513462213_signature_28.png | Agnes G. Ranches | https://www.jotform.com/uploads/macarewecare/222927978372067/5632298430513462213/5632298430513462213_signature_31.png
| 2023-06-22 | 2023-06-22 14:13 | Cawayan Rural Health Unit | Cawayan, Masbate | 5 | Masbate | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24919 | DOH | rhucawayan@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 216 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1179 | 2023-06-19 | GXIV | 838267 | Module C3(663614) for replacement. Unit for calibration. | #NAME? | #NAME? | Unit is operational. C1: C2: C4: |  | GeneXpert Module SN: 784120 XpertCheck Kit | 10024 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632253993912047447/5632253993912047447_signature_28.png | Celedonio Manuel Condor | https://www.jotform.com/uploads/macarewecare/222927978372067/5632253993912047447/5632253993912047447_signature_31.png
| 2023-06-22 | 2023-06-22 12:07 | Lambunao Rural Health Unit | Lambunao, Iloilo | 6 | Iloilo | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24918 | PBSP | gailsantocildes@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1691 |  | GXIV-4D | 815279 | Module B2 Stuck Cartridge/Broken Door | #NAME? | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632164746257380869/5632164746257380869_signature_28.png | Novie Gail Pardilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5632164746257380869/5632164746257380869_signature_31.png
| 2023-06-22 | 2023-06-22 11:47 | Philippine Red Cross Zamboanga City Chapter | Petit Baracks, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR | HAEMONETICS | Gian Carlo Aizon |  |  |  |  | 24917 | COMMERCIAL | zamboangacity@redcross.org.ph | 0 |  | 2023-06-22 | MCS + | 01M256 | Cuff motor malfunction | Replaced cuff motor | Replaced cuff motor | Functional and ready to use | N/A | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632160750419360219/5632160750419360219_signature_28.png | Philippine Red Cross Zamboanga City Chapter | https://www.jotform.com/uploads/macarewecare/222927978372067/5632160750419360219/5632160750419360219_signature_31.png
| 2023-06-22 | 2023-06-22 11:44 | Schistosomiasis Research and Training Center | Palo leyte | 8 | Leyte | Tb dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24916 | PBSP | alramos.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 108 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1255 | 2023-06-22 | GXIV | 707521 | for Module replacement (B1,B2,B3) | #NAME? | -Module replacement * B1/ 608682 into 210063219 * B2/ 652378 into 210060753 * B3/ 665794 into 210068543 -Performed Functionality test -Preventive maintenance -Generate reports -Pullout defective modules | Unit is Operational |  | Genexpert Module 6 color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632154603919751406/5632154603919751406_signature_28.png | Arabelle Joy L. Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5632154603919751406/5632154603919751406_signature_31.png
| 2023-06-22 | 2023-06-22 10:49 | BACOLOD CITY HEALTH OFFICE | Bacolod City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 24915 | PBSP | cyrilcordova0515@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5279 |  | GX-4 | 804267 | Defective Module 677816/E1. | #NAME? | - Replaced module E1 with new module SN 210068557. - Generate and saved IQR and System Log. | - Module 666155/C3 with temperature-related errors. - 15 modules operational, 1 module for report to Cepheid. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632150094119759347/5632150094119759347_signature_28.png | CYRIL CORDOVA | https://www.jotform.com/uploads/macarewecare/222927978372067/5632150094119759347/5632150094119759347_signature_31.png
| 2023-06-22 | 2023-06-22 10:56 | Cataingan Rural Health Unit | Cataingan, Masbate | 5 | Masbate | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24914 | PBSP | catainganrhu@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 143 | 2023-06-22 | GXIV | 812787 | Modules B3(679862) and B4(662447) for replacement. Unit for calibration. | #NAME? | #NAME? | Unit is operational. B1: passed B2: passed |  | GeneXpert Module SN: 765744, 775243 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632129822905170237/5632129822905170237_signature_28.png | Rowena Alicante | https://www.jotform.com/uploads/macarewecare/222927978372067/5632129822905170237/5632129822905170237_signature_31.png
| 2023-06-22 | 2023-06-22 11:00 | Bicol South Luzon Sub National Reference Laboratory | BRTTH Compound, Rizal St., Legazpi City | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24913 | PBSP | dbbalane.pbsp@gmail.com | 20 | ASSAY: Others, COUNT: 79 |  | GX4 | 110016954 | For preventive maintenence | Cleaned cartridge bays, plunger rod,fan filter Generated reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632124007717037675/5632124007717037675_signature_28.png | Darryl Balane | https://www.jotform.com/uploads/macarewecare/222927978372067/5632124007717037675/5632124007717037675_signature_31.png
| 2023-06-22 | 2023-06-22 10:33 | Bicol South Luzon Sub National Reference | BRTTH Compound, Rizal St., Legazpi City | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24912 | PBSP | dbbalane.pbsp@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 6 ASSAY: Others, COUNT: 337 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1094 ASSAY: Xpert HIV-1 Viral Load, COUNT: 149 |  | GX4 | 808636 | For preventive maintenence | Cleaned cartridge bays, fan filter, plunger rod Generated reports | Performed Preventive Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632116907716933944/5632116907716933944_signature_28.png | Darryl Balane | https://www.jotform.com/uploads/macarewecare/222927978372067/5632116907716933944/5632116907716933944_signature_31.png
| 2023-06-22 | 2023-06-22 10:41 | Bicol South Luzon Sun National Reference Laboratory | BRTTH Compound, Rizal St., Legazpi City | 5 | Albay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 24911 | PBSP | ddbalane.pbsp@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 603 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2086 ASSAY: Xpert HIV-1 Viral Load, COUNT: 181 |  | GX4 | 807579 | For preventive maintenance | Cleaned cartridge bays, plunger rod, fan filter Generated reports | Performed Preventive Maintenance | Unit is operational | 784716(D1) & 653834(D2) temperature error code: 1002, 4009 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632113537711792267/5632113537711792267_signature_28.png | Darryl Balane | https://www.jotform.com/uploads/macarewecare/222927978372067/5632113537711792267/5632113537711792267_signature_31.png
| 2023-06-22 | 2023-06-22 10:05 | Commonwealth Health Center | Katuparan, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24910 | DOH | commonwealthclinical@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 586 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2262 |  | GXIV - 6C | 831615 | Module A3 = Hardware Failed Module A4 = Temp problem Monitor flickers | #NAME? | Re-XpertCheck module A3 and A4 Troubleshoot monitor | Unit is operational but A3 and A4 failed XpertCheck and monitor flickers |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632099220684580240/5632099220684580240_signature_28.png | BLUE ROB PACURE | https://www.jotform.com/uploads/macarewecare/222927978372067/5632099220684580240/5632099220684580240_signature_31.png
| 2023-06-22 | 2023-06-22 09:29 | Dinalupihan Rural Health Unit | Brgy. Colo Dinalupihan Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24909 | PBSP | dereggyjuswa@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 132 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1673 ASSAY: Xpert HIV-1 Viral Load, COUNT: 57 | 2023-06-19 | GXIV-D | 807561 | A2 and A3 for module replacement | Module replacement | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 2 genexpert modules | 41579 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5632067348616399060/5632067348616399060_signature_28.png | Joshua Deregla | https://www.jotform.com/uploads/macarewecare/222927978372067/5632067348616399060/5632067348616399060_signature_31.png
| 2023-06-21 | 2023-06-21 14:23 | Kabacan Rural Health Unit | Kabacan, North Cotabato | 12 | Cotabato | Tb - Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24908 | PBSP | yayo_kmsi@yahoo.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 823 |  | Gx-IV ( 2 modules ) | 835182 | Genexpert CPU problem | Cancelled CPU replacement due to HDMI to VGA cable Requirements for new CPU | Genexpert machine Preventive maintenance | 2 operational modules and Genexpert CPU subject for replacement | Reschedule CPU replacement and need to buy HDMI to VGA Cable for installation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631646915017806320/5631646915017806320_signature_28.png | Joan A. Sangines | https://www.jotform.com/uploads/macarewecare/222927978372067/5631646915017806320/5631646915017806320_signature_31.png
| 2023-06-21 | 2023-06-21 18:31 | Casiguran Rural Health Unit | Laboratory | 3 | Aurora | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24907 | DOH | cezalen@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 295 ASSAY: Xpert MTB/RIF Ultra, COUNT: 367 | 2023-02-14 | Genexpert IV | 831844 | Module 1 and 2 error | Performed Xpertcheck on Module A1 & A2 | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 5pc | 9631 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631529961802336659/5631529961802336659_signature_28.png | Princess Aguilar | https://www.jotform.com/uploads/macarewecare/222927978372067/5631529961802336659/5631529961802336659_signature_31.png
| 2023-06-21 | 2023-06-21 14:59 | Santolan Super Health Center | Santolan, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24906 | PBSP | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3995 |  | GXIV 6C | 838898 | Module B2 hardware failed | -Checked and cleaned cartridge bays, icore slot, plunger rod an filter -Replaced gateway board -Module B2 hardware problem has been solved -Replaced the serial of the gateway board to 883898 -Assigned instrument letter to B -Performed self-test -Generated report | Returned to troubleshoot problem on Module B2 | Unit is operational. All modules working properly |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631407160686443687/5631407160686443687_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5631407160686443687/5631407160686443687_signature_31.png
| 2023-06-21 | 2023-06-21 15:07 | TUBIGON RURAL HEALTH UNIT | TUBIGON BOHOL | 7 | Bohol | LABORATORY | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  |  |  |  | 24905 | PBSP | josephraytan@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 825 ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 |  | GX-IV | 815431 | MODULE B1 TEMPERATURE ERROR | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | PERFORM MODULE B1 CALIBRATION, PASSED | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631407015913250792/5631407015913250792_signature_28.png | Joseph Ray Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5631407015913250792/5631407015913250792_signature_31.png
| 2023-06-21 | 2023-06-21 13:20 | Dalaguete Rural Health Unit | Poblacion Dalaguete Cebu | 7 | Cebu | Tb Laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 24904 | PBSP | ajreef5517@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 967 |  | Gx4 | 815464 | Error in Calibration results for Module 4 | Delete Cepheid folder and programmed data in Reinstall expertcheck and run Expertcheck in module 4 | Performed calibration in module 4 | Only 1 modules is used by enduser to avoid cartridge wastage |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631352000227909327/5631352000227909327_signature_28.png | Alaizza Rivera, Rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5631352000227909327/5631352000227909327_signature_31.png
| 2023-06-21 | 2023-06-21 12:45 | West Visayas State University Medical Center | Iloilo City, Iloilo | 6 | Iloilo | Tb Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24902 | PBSP | tvanrlujan@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 728 | 2023-06-21 | GXIV-4D | 804699 | Module B1 & B4 - XpertCheck Service Required | -Replace defective module B2/618510 with new module SN: 210068567 -Replace defective module B4/210019708 with new module SN: 210068572 | #NAME? | All modules are functional. |  | 2 units GeneXpert Module 6-Colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631322220519089507/5631322220519089507_signature_28.png | Van Ryan Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5631322220519089507/5631322220519089507_signature_31.png
| 2023-06-21 | 2023-06-21 12:10 | Sumilang Super Health Center | Sumilang, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24901 | PBSP | maricor.obnamia05@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1220 |  | GX IV 6C | 814163 | PC not booting up | #NAME? | Returned today to continue troubleshooting | Unit is operational. Temporary fix only, will wait for the new cpu to arrive |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631303818517485993/5631303818517485993_signature_28.png | maria corazon obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/5631303818517485993/5631303818517485993_signature_31.png
| 2023-06-21 | 2023-06-21 10:08 | DON EMILIO DEL VALLE MEMORIAL HOSPITAL | BOOD, UBAY, BOHOL | 7 | Bohol | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24900 | PBSP | dedvmhlaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 858 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1751 |  | GX-IV | 807603 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631247130827766889/5631247130827766889_signature_28.png | JAN ARKID A. DIACOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5631247130827766889/5631247130827766889_signature_31.png
| 2023-06-21 | 2023-06-21 10:09 | LORENZO D. ZAYCO DISTRICT HOSPITAL | National Highway, Kabankalan City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 24899 | PBSP | beuagravante@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1696 |  | GX-4 | 807455 | E4009 and E2014 in modules | - Cleabed Gx filter. - Adjust ACU settings to affect lower temp at Genexpert modules. - Run dummy test - | - Checked room temperature - room temp around 22 C - Checked Genexpert IQR for temperature readings - temp at Gx at around 32 C on all modules - Checked ACU airflow to Gx machine, adjusted louvers, checked dilter, adjusted fan speed | Module A1 HW failed duribn dummy test- for replacement. Module A3 - completed specimen run without errors vut with intermittent temp-related errors in Sytem Log - recommended for replacement. - 3 modules usable as of today | A1 and A3 for replacement, reported thru Jotform by the end user |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631231977463218594/5631231977463218594_signature_28.png | BEULAH AGRAVANTE | https://www.jotform.com/uploads/macarewecare/222927978372067/5631231977463218594/5631231977463218594_signature_31.png
| 2023-06-21 | 2023-06-21 10:05 | Anahawan District Hospital | Anahawan S. leyte | 8 | Southern Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24898 | DOH | claritomirasol@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 212 ASSAY: Xpert MTB/RIF Ultra, COUNT: 372 | 2023-06-21 | GXIV | 838548 | N/a | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631226971476530987/5631226971476530987_signature_28.png | Mirasol Clarito | https://www.jotform.com/uploads/macarewecare/222927978372067/5631226971476530987/5631226971476530987_signature_31.png
| 2023-06-21 | 2023-06-21 09:34 | Aroroy Rural Health Unit | Aroroy, Masbate | 5 | Masbate | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24897 | DOH | morenorhenaluz@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 197 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1017 | 2023-06-19 | GXIV | 838268 | #NAME? | #NAME? | #NAME? | Unit is operational. B1: passed B2: failed B3: passed |  | GeneXpert Module SN: 668421 XpertCheck Kit | 10025 & 10017 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5631220685429520696/5631220685429520696_signature_28.png | Rhena Luz | https://www.jotform.com/uploads/macarewecare/222927978372067/5631220685429520696/5631220685429520696_signature_31.png
| 2023-06-20 | 2023-06-20 16:53 | GARCIA MEMORIAL PROVINCIAL HOSPITAL | TALIBON BOHOL | 7 | Bohol | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24896 | DOH | ybanezrosalina19@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 153 ASSAY: Xpert MTB/RIF Ultra, COUNT: 761 |  | GX-IV | 838203 | FOR MODULE B1 CALIBRATION | COLLECT IQR, SYSTEM LOG,AND ARCHIVE | -PERFORM EQUIPMENT MAINTENANCE -PERFORM MODULE B1 SN:730958 CALIBRATION, SERVICE REQUIRED | EQUIPMENT OPERATIONAL, MODULE B1,B2, AND B3 DISABLED DUE TO MULTIPLE ERRORS |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630608169741856627/5630608169741856627_signature_28.png | Rosalina Ybanez | https://www.jotform.com/uploads/macarewecare/222927978372067/5630608169741856627/5630608169741856627_signature_31.png
| 2023-06-20 | 2023-06-20 15:09 | Sumilang Super Health Center | Sumilang, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24895 | PBSP | maricor.obnamia05@gmail.com | 25 |  |  | GX IV 6C | 814163 | PC not booting up | -Replaced CMOS battery -Updated date & time -Troubleshoot in BIOS -Still not booting up, "Start PXE over IPV4" and "BootDevice not found" appears on screen -Tried troubleshooting in BIOS menu still the problem persists -Did preventive maintenance on GX -will return tomorrow for further troubleshooting | #NAME? | GX machine is working but the PC is not booting up |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630583443331558981/5630583443331558981_signature_28.png | Maria Corazon D. Obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/5630583443331558981/5630583443331558981_signature_31.png
| 2023-06-20 | 2023-06-20 16:03 | Moalboal RHU | Poblacion Moalboal, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 24894 | PBSP | culturahoneylournrmt@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 443 |  | Gx4 | 838892 | Monitor problem | Transfer monitor to other cpu and removed ground.. Clean Icore, filter and cartrdge bay | Performed troubleshooting of the Monitor Generate data IQR, Sys log and Gxx files | All modules are operational and monitor are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630579150214793706/5630579150214793706_signature_28.png | Honeylou C. Cultura, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5630579150214793706/5630579150214793706_signature_31.png
| 2023-06-20 | 2023-06-20 14:30 | Badian RHU | Poblacion, Badian Cebu | 7 | Cebu | Tb Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24893 | PBSP | avenzr@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 640 |  | Gx4 | 807588 | For calibration and Maintenance | Performed calibration and maintenance Advice medtech to install airconditioning unit intended for genxepert Calibration results are the ff: B1 - Retry and Passed B2 - Passed B3 - Retry and Failed B4 - Retry and Failed | Performed calibration and Conduct preventive maintenance Generate data IQR and Sys Log, back up database | All modules are operation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630541909187711680/5630541909187711680_signature_28.png | Avenizer P. Absin | https://www.jotform.com/uploads/macarewecare/222927978372067/5630541909187711680/5630541909187711680_signature_31.png
| 2023-06-20 | 2023-06-20 14:12 | PULUPANDAN RURAL HEALTH UNIT | Pulupandan, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 24892 | WHO | jackieorcajada@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 92 ASSAY: Xpert MTB/RIF Ultra, COUNT: 320 |  | GX-IV | 847255 | UPS Problem. Smoke was observed coming out from the UPS unit. | #NAME? | - Checked UPS power cord for any sign of burns. - Checked UPS input for shorts - none detected. - Opened UPS unit and checked for burnt components - none observed. - Checked individual output of each of the 6 12-V batteries - all batteries around 12 Vdc output. - Checked output of Charging circuit - output 88 Vdc. - Power ON the UPS unit, tried to configure back to Online mode. However, UPS mode can only be selected to Bypass Mode. | Possible defective UPS control board. UPS non-functional. GeneXpert machine, CPU, LCD monitor and printer are functional. | For request of backup UPS - at least 2 KVA or above. Also recommending installation of AVR (at least 2 KVA, Servo-type).Temporarily, GeneXpert mahcine, CPU, monitor and printer are connected directly to line voltage. Possible problem if fluctuations or blackout occur during specimen runs. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630517432123155282/5630517432123155282_signature_28.png | Jackie Lou O. Orcajada | https://www.jotform.com/uploads/macarewecare/222927978372067/5630517432123155282/5630517432123155282_signature_31.png
| 2023-06-20 | 2023-06-20 14:07 | Balamban Rural Health Unit | Poblacion, Balamban Cebu | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 24891 | DOH | climaco.mildred@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-06-19 | Gx4 | 838564 | GeneXpert Gx 4 Installation | Install GeneXpert Machine only. | Install GeneXpert Machine only. | GeneXpert Machine ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630512765412665668/5630512765412665668_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5630512765412665668/5630512765412665668_signature_31.png
| 2023-06-20 | 2023-06-20 13:58 | Maysan 3S Health Center | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24890 | PBSP | neldimatera@gmail.com | 35 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 815381 | Additional troubleshooting from Cepheid | #NAME? | Performed various tests and troubleshooting | Unit is not operational. GX system for follow up report. Room temperature exceeds 30 deg. Celsius threshold (avg. module temp. 35 deg). Advised end user to request for repair/maintenance of aircon to prevent module temp. errors on other GX unit. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630505055359100489/5630505055359100489_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/5630505055359100489/5630505055359100489_signature_31.png
| 2023-06-20 | 2023-06-20 12:44 | Tomas Claudio Health Center | Corner E. Quirino, Camus St. Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24889 | DOH | arleen_p_alconera@yahoo.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 222 ASSAY: Xpert MTB/RIF Ultra, COUNT: 472 |  | GX-IV 4D | 838492 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Unit has 3 Functional Modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630489378423241312/5630489378423241312_signature_28.png | GAELMARE M. RELLITA | https://www.jotform.com/uploads/macarewecare/222927978372067/5630489378423241312/5630489378423241312_signature_31.png
| 2023-06-20 | 2023-06-20 11:38 | Calaanan Health Center | Nadurata St. Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24888 | DOH | rosaleslorina0917@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1222 |  | GXIV | 831892 | Unit for PM. For Xpertcheck of Module A4. | Performed PM. Performed Xpertcheck on module A4. Generated reports. | PM and Xpertcheck. | Unit is operational. Module A4 passed. | Used 1 spare Xpertcheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630432558717587003/5630432558717587003_signature_28.png | Lorina Rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/5630432558717587003/5630432558717587003_signature_31.png
| 2023-06-20 | 2023-06-20 11:48 | Pasig City General Hospital | M. Eusebio, Pasig, Metro Manila | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24887 | PBSP | maniquizederlina@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3174 |  | GXIV - 6C | 805010 | Module B1 and B4 for replacement | #NAME? | Replaced Defective Modules B1: 210091374 -> 210059582 B4: 612401 -> 210063233 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630427704426034581/5630427704426034581_signature_28.png | Ederlina Maniquiz Sagales | https://www.jotform.com/uploads/macarewecare/222927978372067/5630427704426034581/5630427704426034581_signature_31.png
| 2023-06-20 | 2023-06-20 11:17 | Koronadal City Health Office | Koronadal City, South Cotabato | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24886 | PBSP | calvokenneth12@gmail.com | 24.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1914 | 2023-06-20 | Gx-IV | 707529 | N/A | Perform Genexpert machine Xpertcheck/Calibration | Perform Genexpert machine Preventive maintenance | 4 Operational modules | All modules passed the Xpertcheck | Xpertcheck -CE-5 | Lot# 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630419821015314258/5630419821015314258_signature_28.png | Kenneth Paolo T. Calvo | https://www.jotform.com/uploads/macarewecare/222927978372067/5630419821015314258/5630419821015314258_signature_31.png
| 2023-06-20 | 2023-06-15 14:43 | General Trias Social Hygiene Clinic | Hospital Road, Brgy., Pinagtipunan, General Trias | 4A | Cavite | Social Hygiene | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24885 | MACARE-HBDC | sittieasyahpangandaman@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: |  | GXIV -D-10C | 110016785 | Unit is due for Installation | Installation | Installed GeneXpert System Generate Report | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630330253614244414/5630330253614244414_signature_28.png | Sittie Asyah Pangandaman | https://www.jotform.com/uploads/macarewecare/222927978372067/5630330253614244414/5630330253614244414_signature_31.png
| 2023-06-20 | 2023-06-20 08:38 | Bohol Provincial Health Office | Dao, Tagbilaran City, Bohol | 7 | Bohol | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24884 | PBSP | aries2cris@yahoo.com | 28 |  |  | Gx 4 | 804286 | Defective modules 1,3 & 4 | Clean fan filter/ replace modules | New module installed in b1, b3 & b4 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5630321231422051088/5630321231422051088_signature_28.png | Ma Crisanta T. Estomago | https://www.jotform.com/uploads/macarewecare/222927978372067/5630321231422051088/5630321231422051088_signature_31.png
| 2023-06-19 | 2023-06-19 11:48 | Dalaguete RHU | Poblacion Dalaguete Cebu | 7 | Cebu | TB Laboratory | YES | REPAIR CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 24883 | PBSP | ajreef5517@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 777 |  | Gx4 | 815464 | Stucked cartridge in Module 2 due to Error 1018 | Removed stucked cartridge in Module 2 and calibrate Module 3 and 4 failed during calibration Module 1 and 2 Passed Module 3 and 4 Failed | Removed strucked cartridge and Calibrate Genexpert machine | All modules are available but only module 1 and 2 are used by the enduser due to frequent errors Module 2 and 3. Module 1 and 2 - Passed Module 3 and 4 - Failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5629573043416688099/5629573043416688099_signature_28.png | Alaizza Rivera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5629573043416688099/5629573043416688099_signature_31.png
| 2023-06-19 | 2023-06-16 15:11 | Cabanatuan City Health Office | Kapt. Pepe, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24882 | DOH | cholab2021@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 200 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2371 | 2023-06-16 | Genexpert IV | 831592 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5629560933611131716/5629560933611131716_signature_28.png | Lea Samson | https://www.jotform.com/uploads/macarewecare/222927978372067/5629560933611131716/5629560933611131716_signature_31.png
| 2023-06-19 | 2023-06-15 09:06 | Philippine heart center | East avenue q.c | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24881 | COMMERCIAL | baluyotjeric@gmail.com | 28 |  | 2023-04-19 | MCS + | O20168 | Due for Maintenance | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance | Operational |  | Urethane rollers | 9910 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5629440593619871720/5629440593619871720_signature_28.png | Jeric Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5629440593619871720/5629440593619871720_signature_31.png
| 2023-06-19 | 2023-06-15 10:15 | Philippine heart center | East avenue q.c | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24880 | COMMERCIAL | baluyotjeric@gmail.com | 28 |  | 2023-04-19 | MCS + | 01M201 | Due for maintenance | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance | Operational |  | Urethane rollers | 9910 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5629439713619620548/5629439713619620548_signature_28.png | JERIC BALUYOT | https://www.jotform.com/uploads/macarewecare/222927978372067/5629439713619620548/5629439713619620548_signature_31.png
| 2023-06-19 | 2023-06-15 14:26 | Philippine Red Cross qc chapter | Malakas st., Quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24879 | COMMERCIAL | prcqc.bloodservices@gmail.com | 28 |  |  | MCS + | 02A227 | Due for maintenance , UPS Unable to back up defective battery | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance | Machine operational, | For replacement of UPS |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5629438883618092890/5629438883618092890_signature_28.png | Lambert Cordero | https://www.jotform.com/uploads/macarewecare/222927978372067/5629438883618092890/5629438883618092890_signature_31.png
| 2023-06-19 | 2023-06-18 11:00 | Philippine heart center | East avenue q.c | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 24878 | COMMERCIAL | baluyotjeric@gmail.com | 27 |  |  | MCS + | 20168 | Malfunction of display panel button and unable to press | Performed troubleshooting Performed reset of CPU board Check cable and ribbon Swapping of display board assembly Cheked cable connection | Troubleshooting | Unit unoperational due to malfunction of display panel button | For replacement of display board assembly |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5629438343615985365/5629438343615985365_signature_28.png | Jeric Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5629438343615985365/5629438343615985365_signature_31.png
| 2023-06-18 | 2023-06-15 12:51 | Dr. Jorge P. Royeca Hospital | General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24877 | PBSP | fatchoi12001@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 839 ASSAY: Xpert HIV-1 Viral Load, COUNT: 131 | 2023-06-15 | Gx-IV | 838130 | Module 4 H/W failed | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Xpertcheck / Calibration | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Xpertcheck / Calibration | 3 operational modules | All module passed the Calibration | Xpertcheck CE -5 | Lot 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5628970988652868756/5628970988652868756_signature_28.png | Maria Ivy A. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5628970988652868756/5628970988652868756_signature_31.png
| 2023-06-16 | 2023-06-16 08:00 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24876 | PBSP | r1mc.molecularlab@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1488 ASSAY: Xpert SARS-CoV-2, COUNT: 3569 |  | Gx 16 | 842230 | Module B1 | #NAME? | Collect System Log Report and Troubleshooting | Unit is operational | Module B1 was disabled upon booting. No Errors Logged |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5627280897122925494/5627280897122925494_signature_28.png | Efren notario jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5627280897122925494/5627280897122925494_signature_31.png
| 2023-06-16 | 2023-06-16 08:00 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24875 | PBSP | r1mc.molecularlab@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1488 ASSAY: Xpert SARS-CoV-2, COUNT: 3569 |  | Gx 16 | 842230 | Module B1 | #NAME? | Collect System Log Report and Troubleshooting | Unit is operational | Module B1 was disabled upon booting. No Errors Logged |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5627280627129045492/5627280627129045492_signature_28.png | Efren notario jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5627280627129045492/5627280627129045492_signature_31.png
| 2023-06-16 | 2023-06-16 15:09 | Philippine Institute of Volcanology and Seismology | C.P. Garcia Ave., UP Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | MAINTENANCE | CARL ZEISS | Carlos Leniel Origenes |  |  |  |  | 24874 | COMMERCIAL | vgggrds.laboratory@gmail.com | 25 |  |  | Stemi 508 | 3942010146 | None | None | Cleaning of eyepiece, camera lens, lens, and body Cleaning, checking, and alignment of mechanical parts Functionality Testing of Power Supply Functionality Testing of the Camera with the software Functionality Testing of the System | Unit is operational. Defective Workstation, recommended to be replaced. | There is a prism for the camera, a small knob below the camera. Controls the blocking of the right eyepiece and camera. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5627095586712322524/5627095586712322524_signature_28.png | DOST-PHIVOLCS VGGGRDS PETROGRAPHY AND PETROLOGY LABORATORY | https://www.jotform.com/uploads/macarewecare/222927978372067/5627095586712322524/5627095586712322524_signature_31.png
| 2023-06-16 | 2023-06-16 14:43 | Tabaco City Health Unit | Tabaco City | 5 | Albay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24873 | PBSP | elenocbina@gmail.com | 33 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3138 | 2023-06-16 | GX4 | 804257 | For XpertCheck | XpertCheck Cleaned fan filter, cartridge bays, plunger ror Generated reports | Performed XpertCheck, Preventive Maintenance | Unit is operational All passed |  | XpertCheck kit | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5627085741185943215/5627085741185943215_signature_28.png | Eleanor Ocbina | https://www.jotform.com/uploads/macarewecare/222927978372067/5627085741185943215/5627085741185943215_signature_31.png
| 2023-06-16 | 2023-06-16 14:00 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24872 | PBSP | atumbaga1079@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert MTB/RIF, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 31 ASSAY: Others, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 190 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 |  | GXIV | 847726 | Unit for maintenance and calibration | #NAME? | Performed maintenance and calibration | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5627079302112283262/5627079302112283262_signature_28.png | Teodoro, Karl-Wynne M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5627079302112283262/5627079302112283262_signature_31.png
| 2023-06-16 | 2023-06-16 13:34 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24871 | PBSP | spmcrhivda@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1464 ASSAY: Xpert HIV-1 Viral Load, COUNT: 498 |  | GX-IV 4D | 808298 | Module B1 for replacement | Replace the defective Module B1 SN: 671900 with new Module SN: 210068532 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses | Perform Module Replacement | All Modulea are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5627041966916599751/5627041966916599751_signature_28.png | Christelle Sienes RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5627041966916599751/5627041966916599751_signature_31.png
| 2023-06-16 | 2023-06-16 13:25 | Barili Rural Health Unit | Poblacion Barili, Cebu | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24870 | PBSP | barilirhumain@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 627 |  | Gx4 | 834693 | Module 3 SN 786755 Hardware failure Error 1001 and Module 4 SN 662771 Error 2014 and Error 1001 | Replaced new module 4 SN # 210068556 | Replaced Module 4 Sn # 662771 | Only 3 modules are operational, Module 3 Hardware Failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5627027251338304245/5627027251338304245_signature_28.png | Mary Joy B. Bajarias | https://www.jotform.com/uploads/macarewecare/222927978372067/5627027251338304245/5627027251338304245_signature_31.png
| 2023-06-16 | 2023-06-16 13:21 | San Pedro City Health Office 1 | San Pedro | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24869 | DOH | venersantos0425@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 259 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1548 |  | GXIV | 821224 | Unit for PM. For Xpertcheck of module A3. | Performed PM. Performed Xpertcheck on module A3. Generated reports. | PM and Xpertcheck. | Unit is operational. Module A3, service required. | Used 1 spare Xpertcheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5627025154951755582/5627025154951755582_signature_28.png | Vener R. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5627025154951755582/5627025154951755582_signature_31.png
| 2023-06-16 | 2023-06-16 10:58 | Manila Doctors Hospital | United Nations Ave., Manila | NCR | NCR 1st - Manila | Microbiology Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24868 | COMMERCIAL | rfbuentiempo@maniladoctors.com.ph | 19 | ASSAY: Others, COUNT: 10 ASSAY: Xpert CT/NG, COUNT: 92 ASSAY: Xpert HBV VL, COUNT: 55 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 986 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 |  | GXIV 10C | 110001648 | Module B3 Temp Problem | -Checked and cleaned cartridge bays, plunger rod, icore slot and filter -Checked system log -B3 have many error 1002 -Disabled Module B3 -Module B3 for replacement -Generated report | Checked system logs and PM | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626981248222651272/5626981248222651272_signature_28.png | DK Valois | https://www.jotform.com/uploads/macarewecare/222927978372067/5626981248222651272/5626981248222651272_signature_31.png
| 2023-06-16 | 2023-06-16 11:40 | Sta Rosa City Health Office 2 | Kabesang Moldes St. Sta. Rosa | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24867 | DOH | annie.soliman@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: |  | GXIV | 821156 | Unit for PM. For Xpertcheck of module A3. | Performed PM. Performed Xpertcheck on module A3. Generated reports | PM and Xpertcheck. | Unit is operational. Module A3, service required. | Used 1 spare Xpertcheck. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626980914952664288/5626980914952664288_signature_28.png | Annie Soliman | https://www.jotform.com/uploads/macarewecare/222927978372067/5626980914952664288/5626980914952664288_signature_31.png
| 2023-06-16 | 2023-06-15 16:22 | Toledo Cho | Ilihan Toledo City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24866 | PBSP | loriemaediapen@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 783 |  | Gx4 | 808640 | Due for Maintenance and clean filter, clean plunger | Peformed maintenance and clean, back up data | Conduct maintenance, replaced filter, clean Plunger, Icore, Back up database, generate IQR, sys log, assay statistics, gxx results | All modules are oprational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626926863926616372/5626926863926616372_signature_28.png | Lorie Mae Diapen, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5626926863926616372/5626926863926616372_signature_31.png
| 2023-06-16 | 2023-06-16 10:15 | Los Banos Rural Health Unit | Los Banos,Laguna | 4A | Laguna | Laboratory | NO | REPAIR TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24865 | PBSP | rmbtaguic@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1074 |  | GXIV | 804716 | Modules not detected. | Performed repair and troubleshooting. Generated reports. | Repair and troubleshooting | Unit is operational. Endorsed additional Lan cable for backup. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626912474954569029/5626912474954569029_signature_28.png | Rhea Marie Dorado | https://www.jotform.com/uploads/macarewecare/222927978372067/5626912474954569029/5626912474954569029_signature_31.png
| 2023-06-16 | 2023-06-16 09:59 | Tuburan Rural Health Unit | Poblacion,Tuburan,Cebu | 7 | Cebu | Laboratory | YES | Pull out GeneXpert Machine system. | CEPHEID | Jovy Noel |  |  |  |  | 24864 | DOH | marielgracia20@gmail.com | 23 |  | 2023-06-15 | Gx4 | 838564 | GeneXpert system transfer to Balamban Rural Health Unit | Pull out GeneXpert Machine S/N.838564, Cpu S/N.MJ0C9VF6, Voltage Regulator S/N.20PX100104(SVC), UPS S/N.AT1000190402ST0391, LCD S/N.CNK9430322 | Pull out GeneXpert Machine S/N.838564, Cpu S/N.MJ0C9VF6, Voltage Regulator S/N.20PX100104(SVC), UPS S/N.AT1000190402ST0391, LCD S/N.CNK9430322 | All are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626906725417449868/5626906725417449868_signature_28.png | Mariel R. Gracia | https://www.jotform.com/uploads/macarewecare/222927978372067/5626906725417449868/5626906725417449868_signature_31.png
| 2023-06-16 | 2023-06-16 10:03 | Gerona Rural Health Unit | Brgy. Pinasling Gerona Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24863 | DOH | huwakie15@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 94 ASSAY: Xpert MTB/RIF Ultra, COUNT: 423 |  | GXIV-D | 831693 | Error 4004, 2016 For re-xpertcheck | Troubleshooting Re-xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | Used 1 xpertcheck kit, under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626904694871366850/5626904694871366850_signature_28.png | Artemio I. Nool III | https://www.jotform.com/uploads/macarewecare/222927978372067/5626904694871366850/5626904694871366850_signature_31.png
| 2023-06-16 | 2023-06-16 09:44 | Balamban Rural Health Unit | Poblacion, Balamban Cebu | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  |  |  |  | 24862 | DOH | climaco.mildred@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 51 | 2023-06-09 | Gx4 | 838231 | Module B3 with serial number 807124 Invalid resolt 2x | Perform XPert Check for Module B3 | Perform Xpertcheck for Module B3 | Module B3 Pass for Calibration Ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626896615414014096/5626896615414014096_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5626896615414014096/5626896615414014096_signature_31.png
| 2023-06-15 | 2023-06-15 15:16 | Wenceslao Dela Paz Health Center | Culiat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24861 | DOH | ameliafclemente2001@yahoo.com | 28 | ASSAY: Xpert HPV, COUNT: 492 ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2253 |  | GXIV 6C | 831876 | Health Center for renovation | #NAME? | Packed all equipment and transferred to another location | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626365693688962381/5626365693688962381_signature_28.png | amelia clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5626365693688962381/5626365693688962381_signature_31.png
| 2023-06-15 | 2023-06-15 13:15 | Santa Ignacia Rural Health Unit | Santa Ignacia Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24860 | DOH | jdi0973@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 48 ASSAY: Xpert MTB/RIF Ultra, COUNT: 323 |  | GXIV-D | 831909 | CPU Error 3F0 | Troubleshooting and repair | Performed hardware reset Checked Hard drive and run harddrive test Perdormed surface cleaning Gathered reports Performed functionality test | Unit is operational | Under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626270344281991838/5626270344281991838_signature_28.png | Jess D. Ibarra Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5626270344281991838/5626270344281991838_signature_31.png
| 2023-06-15 | 2023-06-15 14:14 | San Luis Rural Health Unit | Poblacion, San Luis, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24859 | DOH | mohan_santos@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 265 | 2023-06-15 | Genexpert IV | 831714 | N/a | N/a | Performed Maintenance Gathered Reports Archived Results (1 Year) Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626242563616391264/5626242563616391264_signature_28.png | Mohan S. Pangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5626242563616391264/5626242563616391264_signature_31.png
| 2023-06-15 | 2023-06-15 13:29 | Santolan Super Health Center | Santolan, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24858 | PBSP | samandlet17@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3973 |  | GXIV 6C | 838898 | Stuck Cartridge - Module B2 | -Perfomer prevenetive maintenance -Removed stuck cartridge -Hardware Failed module b2 fue to error 4008 -Will comeback tomorrow to test the supply voltage -generated reports | Removed stuck cartridge | unit is operational. B2 disabled temporarily |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626165229463252689/5626165229463252689_signature_28.png | Lolita Ong Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5626165229463252689/5626165229463252689_signature_31.png
| 2023-06-15 | 2023-06-15 12:55 | Talisay CHU | Cansojong, Talisay City | 7 | Cebu | Tb Laboratory | YES | CALIBRATION TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 24857 | PBSP | marydelgella@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1237 |  | Gx4 | 815384 | Error 2014 in Module 4 SN 210018733 and module 3 210019110 due 5011 cartridge wasted | Performed calibration and generate data | Performed Expert Check calibration | Only 2 modules are available |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626147194421097325/5626147194421097325_signature_28.png | Marydel Gella | https://www.jotform.com/uploads/macarewecare/222927978372067/5626147194421097325/5626147194421097325_signature_31.png
| 2023-06-15 | 2023-06-15 12:30 | LAPU-LAPU CITY HEALTH OFFICE | LAPU-LAPU CITY | 7 | Cebu | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 24856 | PBSP | tanya_pax@yahoo.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 650 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1625 |  | GX-IV | 808637 | FOR MODULE B3 REPLACEMENT | #NAME? | -PERFORM MODULE B3 DEF SN: 724674 REPLACEMENT WITH NEW SN: 210060756 | EQUIPMENT IS OPERATIONAL | MODULE B4 FOR OBSERVATION |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626135542326440204/5626135542326440204_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5626135542326440204/5626135542326440204_signature_31.png
| 2023-06-15 | 2023-06-15 12:29 | Cebu TB Reference Laboratory | SRP Cebu City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24855 | PBSP | asmantiquilla.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | Gx4 | 120000180 | For monitoring after installation | Performed check up | Performed machine check up and power check | Units are operational, but the site dont have enough exhaust pressure. For the meantime, CTRL paused its laboratory opration while awaiting the materials for renovation of negative pressure machine. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626134164429064064/5626134164429064064_signature_28.png | Allan Mantiquilla, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5626134164429064064/5626134164429064064_signature_31.png
| 2023-06-15 | 2023-06-15 09:42 | Eastern Samar Provincial Hospital | Borongan City | 8 | Eastern Samar | Tb dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24854 | PBSP | nathanycarmel@gmai.com | 34 |  | 2023-06-15 | GXIV | 804256 | N/a | Preventive maintenance | Preventive maintenance | Unit is Operational (Air Conditioning is not working properly- room temperature above 30°C) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626105824293202782/5626105824293202782_signature_28.png | Glynnie Operario | https://www.jotform.com/uploads/macarewecare/222927978372067/5626105824293202782/5626105824293202782_signature_31.png
| 2023-06-15 | 2023-06-15 09:42 | Eastern Samar Provincial Hospital | Borongan City | 8 | Eastern Samar | Tb dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24853 | PBSP | nathanycarmel@gmai.com | 34 |  | 2023-06-15 | GXIV | 804256 | N/a | Preventive maintenance | Preventive maintenance | Unit is Operational (Air Conditioning is not working properly- room temperature above 30°C) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626105614298608858/5626105614298608858_signature_28.png | Glynnie Operario | https://www.jotform.com/uploads/macarewecare/222927978372067/5626105614298608858/5626105614298608858_signature_31.png
| 2023-06-15 | 2023-06-15 11:27 | Presentacion RHU | Brgy. Pili, Presentacion, Camarines Sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24852 | PBSP | dianacorreisrael@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 622 | 2023-06-01 | GX-4 | 834808 | For Module Replacement | Replaced Module: 643064>>210059921 Cleaned fan, fan filter, plunger rod, cartridge bays Generated reports | Replaced Defective Module, Performed Preventive Maintenance | Unit is Operational |  | 210059921 | 40974 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626096231015785545/5626096231015785545_signature_28.png | Diana israel | https://www.jotform.com/uploads/macarewecare/222927978372067/5626096231015785545/5626096231015785545_signature_31.png
| 2023-06-15 | 2023-06-15 11:14 | Parola Health Center | Parola, Tondo | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24851 | DOH | friaslynette@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 824 |  | GXIV | 831871 | Unit for PM. For Re-Xpertcheck of moeule B4. | Performed PM. Performed Re-Xpertcheck on module B4. Generated reports. | PM and Re-Xpertcheck. | Unit is operational. | Used 1 spare Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626095102442159823/5626095102442159823_signature_28.png | Lynette Olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5626095102442159823/5626095102442159823_signature_31.png
| 2023-06-15 | 2023-06-14 11:43 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24850 | PBSP | cdh.laboratorydept@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 19 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1111 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 |  | GXIV-D | 807562 | Module C2 not detected | Troubleshooting and repair | Performed surface cleaning Performed preventive maintenance Reset Gateway board Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626070984289998356/5626070984289998356_signature_28.png | Rowena Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5626070984289998356/5626070984289998356_signature_31.png
| 2023-06-15 | 2023-06-14 11:43 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24849 | PBSP | cdh.laboratorydept@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 19 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1111 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 |  | GXIV-D | 807562 | Module C2 not detected | Troubleshooting and repair | Performed surface cleaning Performed preventive maintenance Reset Gateway board Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5626070744283150908/5626070744283150908_signature_28.png | Rowena Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5626070744283150908/5626070744283150908_signature_31.png
| 2023-06-14 | 2023-06-07 16:24 | Libmanan RHU II | San Isidro, Libmanan | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24848 | PBSP | amyquiapo1972@gmail.com | 25 |  | 2026-04-27 | GX4 | 805260 | For Module Replacement, For XpertCheck | Replaced Module B4(705303>>736878) B1(616978>>744281) Performed XpertCheck Cleaned cartridge bays,plunger rod, fan,replaced fan filter Generated reports | Replace Module Performed XpertCheck Performed Preventive Maintenance | Unit is operational B3 failed | B2 677428 failed xpertcheck | Module 744281 736878 XpertCheck kit | Delivery report 0040635, 0040661 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625534381127997559/5625534381127997559_signature_28.png | Amy Quiapo | https://www.jotform.com/uploads/macarewecare/222927978372067/5625534381127997559/5625534381127997559_signature_31.png
| 2023-06-14 | 2023-06-07 16:24 | Libmanan RHU II | San Isidro, Libmanan | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24847 | PBSP | amyquiapo1972@gmail.com | 25 |  | 2026-04-27 | GX4 | 805260 | For Module Replacement, For XpertCheck | Replaced Module B4(705303>>736878) B1(616978>>744281) Performed XpertCheck Cleaned cartridge bays,plunger rod, fan,replaced fan filter Generated reports | Replace Module Performed XpertCheck Performed Preventive Maintenance | Unit is operational B3 failed | B2 677428 failed xpertcheck | Module 744281 736878 XpertCheck kit | Delivery report 0040635, 0040661 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625534011122172709/5625534011122172709_signature_28.png | Amy Quiapo | https://www.jotform.com/uploads/macarewecare/222927978372067/5625534011122172709/5625534011122172709_signature_31.png
| 2023-06-14 | 2023-06-14 18:02 | Chong Hua Hospital | Fuente Osmena, Cebu City | 7 | Cebu | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 24846 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | Mcs+ | 02F035 | Due for calibration and preventive maintenanance | Performed maintenance and calibration | Performed calibration and Preventive maintenance | Mcs+ is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625463736517422006/5625463736517422006_signature_28.png | Dave Gabini | https://www.jotform.com/uploads/macarewecare/222927978372067/5625463736517422006/5625463736517422006_signature_31.png
| 2023-06-14 | 2023-06-14 17:56 | Chong Hua Hospital | Fuente Osmena Cebu City | 7 | Cebu | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 24845 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  |  | Mcs+ | 11J219 | Due for Calibration | Calibrate and Performed maintenance | Performed calibration and Preventive maintenance | MCS+ is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625461046512744059/5625461046512744059_signature_28.png | Dave Gabini, Biomedical Supervisor | https://www.jotform.com/uploads/macarewecare/222927978372067/5625461046512744059/5625461046512744059_signature_31.png
| 2023-06-14 | 2023-06-14 17:01 | Adela Serra Ty Memorial Medical Center | Capitol Hills, Telaje, Tandag City | 13 | Surigao del Sur | RTDL | NO | INSTALLATION | CEPHEID | Faizar Avenido |  |  |  |  | 24844 | PBSP | astmmclaboratory17@gmail.com | 25.22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4849 ASSAY: Xpert HIV-1 Viral Load, COUNT: 26 |  | GX-4 | 805279 | 3 modules found service required from the last XPC | Reported to Cepheid all service required modules | Installation of new back-up Cepheid GX-4 Machine with SN: 805279 | 6 modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625428433113128333/5625428433113128333_signature_28.png | Alberto Plaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5625428433113128333/5625428433113128333_signature_31.png
| 2023-06-14 | 2023-06-14 14:17 | Masambong Health Center | Capoas, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24843 | PBSP | cocristy987@gmail.com | 34 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 914 |  | GXIV 6C | 814160 | Horizontal line in monitor | #NAME? | Troubleshoot monitor | Monitor is operational but with horizontal line |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625388121394934767/5625388121394934767_signature_28.png | cristy cortezano | https://www.jotform.com/uploads/macarewecare/222927978372067/5625388121394934767/5625388121394934767_signature_31.png
| 2023-06-14 | 2023-06-14 15:06 | Talisay City Health Office | Talisay City proper | 6 | Negros Occidental | TB DOTS | NO | INSTALLATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24842 | WHO | hadmeoffguard@gmail.com | 4 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX4 | 847222 | Printer Problem | #NAME? | #NAME? | Printer software installation completed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625368960499406457/5625368960499406457_signature_28.png | Hadney Besa | https://www.jotform.com/uploads/macarewecare/222927978372067/5625368960499406457/5625368960499406457_signature_31.png
| 2023-06-14 | 2023-06-08 14:05 | Mexico Rural Health Unit IV | Pandacaqui, Mexico, Pampanga | 3 | Pampanga | Laborarory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24840 | DOH | annbathan79@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 514 | 2023-06-08 | Genexpert IV | 831902 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625335294744416177/5625335294744416177_signature_28.png | Mary Ann Carreon | https://www.jotform.com/uploads/macarewecare/222927978372067/5625335294744416177/5625335294744416177_signature_31.png
| 2023-06-14 | 2023-06-07 13:58 | Mexico Rural Health Unit I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 24839 | DOH | jasperdizon20@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 73 ASSAY: Xpert MTB/RIF Ultra, COUNT: 407 | 2023-06-07 | Genexpert IV | 831877 | N/a | N/a | Performed Maintenance Gathered Reports Performed Xpertcheck on A4 Module Performed Functionality Test | Unit is Operational | Xpertcheck Result = Service Required | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625335294749793375/5625335294749793375_signature_28.png | Jasper Dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5625335294749793375/5625335294749793375_signature_31.png
| 2023-06-14 | 2023-06-14 14:20 | Candaba Rural Health Unit | Poblacion, Candaba, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24841 | DOH | rochelannlontoc@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 564 | 2023-06-14 | Genexpert IV | 831873 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625335304742138339/5625335304742138339_signature_28.png | Al-JANE MAGPAYO | https://www.jotform.com/uploads/macarewecare/222927978372067/5625335304742138339/5625335304742138339_signature_31.png
| 2023-06-14 | 2023-06-14 13:16 | Sulat Rural Health Unit | Sulat Eastern Samar | 8 | Eastern Samar | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24838 | PBSP | rhusulat@gmail.com | 31 | ASSAY: Xpert MTB/RIF, COUNT: 152 ASSAY: Xpert MTB/RIF Ultra, COUNT: 647 | 2023-06-14 | GXIV- 2D | 835100 | Defective Module B2/787997 for replacement | #NAME? | -Install new module on B2 sn: 210067846 - Performed Preventive maintenance - Generate reports | Unit is Operational | Room air conditioning is not working properly | Genexpert Module 6 color sn: 210067846 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625295303612189488/5625295303612189488_signature_28.png | Edel Abogaa | https://www.jotform.com/uploads/macarewecare/222927978372067/5625295303612189488/5625295303612189488_signature_31.png
| 2023-06-14 | 2023-06-14 12:50 | Kalinga Health Hub - Mandaluyong | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24837 | CONSORTIUM | shad.famisaran@innovationsch.org | 27 | ASSAY: Xpert MTB/RIF, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 611 | 2023-06-13 | GXIV | 110008415 | Module B4 (210035637) for replacement | #NAME? | Replaced B4 | Unit is operational. |  | GeneXpert Module SN: 729705 | 5333 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625280233612957201/5625280233612957201_signature_28.png | Shad Famisaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5625280233612957201/5625280233612957201_signature_31.png
| 2023-06-14 | 2023-06-14 11:24 | Tacurong City Health Office | Tacurong City, Sultan Kudarat | 12 | Sultan Kudarat | Tb - Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24836 | PBSP | torrecampoalma8@gmail.com | 22.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1863 | 2023-06-21 | Gx-IV | 815393 | Module 2 H/W failed due to multiple temperature error | Replace module 2 SN:792515 with new module SN:210059883 | Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 color SN: 210059883 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625234590427591234/5625234590427591234_signature_28.png | Alma Torrecampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5625234590427591234/5625234590427591234_signature_31.png
| 2023-06-14 | 2023-06-08 10:14 | Molino III City Health Unit | Malinis Compound, Molino, Bacoor | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24835 | DOH | joanrnd@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1167 |  | GXIV-D | 821221 | Module 1 with Temperature Error UPS Problem | XpertCheck | Performed Maintenance Performed XpertCheck on Module 1 UPS Functionality Test Generate Reports | UPS has Defective Battery, 3 Modules are Operational, 1 Defective | B1: Failed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625195033611619770/5625195033611619770_signature_28.png | Jonalyn Aranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5625195033611619770/5625195033611619770_signature_31.png
| 2023-06-14 | 2023-06-09 11:19 | General Emilio Aguinaldo Memorial Hospital | Brgy. Luciano, Trece Martires | 4A | Cavite | Laboratory | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24835 | PBSP | djaemiel@gmail.com | 28 |  |  | GXIV-D-10C | 120000175 | Unit is due for Installation | Installation | Installed GeneXpert System Generate Reports | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5625195033615963547/5625195033615963547_signature_28.png | JAEMIEL JUVIRENS A. DIMARANAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5625195033615963547/5625195033615963547_signature_31.png
| 2023-06-13 | 2023-06-13 14:03 | Philippine Red Cross Zamboanga | Pettit baracks, Zamboanga ci6 | 9 | Surigao del Sur | Laboratory | NO | TROUBLESHOOTING | HAEMONETICS | Gian Carlo Aizon |  |  |  |  | 24834 | COMMERCIAL | zamboangacity@redcross.org.ph | 0 |  | 2023-06-13 | MCS+ | 01M256 | Cuff motor malfunction and ups malfunction | Undergone troubleshooting using UPP Technician Configuration and Diagnostic card. Regarding ups all batteries have been checked and all passed still ups gives overload output. | Undergone troubleshooting using UPP Technician Configuration and Diagnostic card. Regarding ups all batteries have been checked and all passed still ups gives overload output. | Ups not working and cuff motor malfunction needs to be replaced | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5624468805414256123/5624468805414256123_signature_28.png | Joan A. Luna | https://www.jotform.com/uploads/macarewecare/222927978372067/5624468805414256123/5624468805414256123_signature_31.png
| 2023-06-13 | 2023-06-13 09:04 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Tb - Lab | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24833 | DOH | jan19.cam@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF, COUNT: 99 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1378 | 2023-06-08 | Gx-IV | 838504 | Module 4 H/W failed due to multiple temperature error | Perform Xpertcheck at module 4 for troubleshooting | Perform Genexpert machine Preventive maintenance | 3 operational modules | Module 4 for Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5624281017811584966/5624281017811584966_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5624281017811584966/5624281017811584966_signature_31.png
| 2023-06-13 | 2023-06-13 08:52 | Misamis Occidental Provincial Hospital | Oroquieta City | 10 | Misamis Occidental | Laboratory | YES | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24832 | PBSP | mophlaboratory2017@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1345 |  | GeneXper | 807618 | Module A1/662352 Service Required after calibration | Replaced module new module A1/210091310 | Preventive Maintenance Module Replacement Xpertcheck calibration Generate Report Files | Operational | A3/653705 and A4/695639 Service Required during calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5624275022227100655/5624275022227100655_signature_28.png | Mary Auxilium Omandam | https://www.jotform.com/uploads/macarewecare/222927978372067/5624275022227100655/5624275022227100655_signature_31.png
| 2023-06-12 | 2023-06-09 10:26 | Malungon Rural Health Unit | Poblacion Malungon, Sarangani Province | 12 | Sarangani | Tb - Lab | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24831 | DOH | madameerrmie@gmail.com | 28.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 179 |  | Gx-IV | 838546 | Module 2 multiple temperature error | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 2 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 2 for troubleshooting | 3 operational modules | Module 2 for Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623851421411484625/5623851421411484625_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5623851421411484625/5623851421411484625_signature_31.png
| 2023-06-12 | 2023-06-07 14:52 | Lung center of the Philippines | Quezon avenue, Quezon City | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24830 | MACARE-HBDC | lcpmicrolab@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 303 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1180 ASSAY: Xpert SARS-CoV-2, COUNT: 1263 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GX4 | 804205 | Due for maintenance, d4 disable loss communication,d3 disable temp error | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | D1,d2 operational, d3 and d4 disable | Last xpert check 6.30.2021 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623792108111496053/5623792108111496053_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5623792108111496053/5623792108111496053_signature_31.png
| 2023-06-12 | 2023-06-07 14:58 | Lung center of the Philippines | Quezon avenue, Quezon City | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24829 | PBSP | lcpmicrolab@gmail.com | 32 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3391 ASSAY: Xpert SARS-CoV-2, COUNT: 3881 |  | GXXVI | 801080 | Due for maintenance b1,b3,c1,d4 disable | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | 12 module operational, 4 module disable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623791528113636287/5623791528113636287_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5623791528113636287/5623791528113636287_signature_31.png
| 2023-06-12 | 2023-06-07 14:58 | Lung center of the Philippines | Quezon avenue, Quezon City | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24828 | PBSP | lcpmicrolab@gmail.com | 32 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3391 ASSAY: Xpert SARS-CoV-2, COUNT: 3881 |  | GXXVI | 801080 | Due for maintenance b1,b3,c1,d4 disable | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | 12 module operational, 4 module disable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623789878114857534/5623789878114857534_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5623789878114857534/5623789878114857534_signature_31.png
| 2023-06-12 | 2023-06-07 14:46 | Lung center of the Philippines | Quezon ave,Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24827 | MACARE-HBDC | lcpmicrolab@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 40 ASSAY: , COUNT: 426 ASSAY: Xpert SARS-CoV-2, COUNT: 1346 ASSAY: Xpert HIV-1 Viral Load, COUNT: 65 ASSAY: Others, COUNT: 9 ASSAY: Others, COUNT: 1 |  | GX4 | 838837 | Due for maintenance, b4 hw failed, | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | B4 disabled, 3 module operational | Last xpertcheck july 2022 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623788628115357493/5623788628115357493_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5623788628115357493/5623788628115357493_signature_31.png
| 2023-06-12 | 2023-06-07 14:46 | Lung center of the Philippines | Quezon ave,Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24826 | MACARE-HBDC | lcpmicrolab@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 40 ASSAY: , COUNT: 426 ASSAY: Xpert SARS-CoV-2, COUNT: 1346 ASSAY: Xpert HIV-1 Viral Load, COUNT: 65 ASSAY: Others, COUNT: 9 ASSAY: Others, COUNT: 1 |  | GX4 | 838837 | Due for maintenance, b4 hw failed, | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | B4 disabled, 3 module operational | Last xpertcheck july 2022 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623787948118602038/5623787948118602038_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5623787948118602038/5623787948118602038_signature_31.png
| 2023-06-12 | 2023-06-07 15:33 | Lung center of the Philippines | Quezon avenue, Quezon City | NCR | NCR 2nd - Quezon City | Koica | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24825 | PBSP | nilabis.pbsp@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1330 |  | GX4 | 707450 | Due for maintenance | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623787288111287597/5623787288111287597_signature_28.png | Norman Labis | https://www.jotform.com/uploads/macarewecare/222927978372067/5623787288111287597/5623787288111287597_signature_31.png
| 2023-06-12 | 2023-06-07 15:33 | Lung center of the Philippines | Quezon avenue, Quezon City | NCR | NCR 2nd - Quezon City | Koica | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24824 | PBSP | nilabis.pbsp@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1330 |  | GX4 | 707450 | Due for maintenance | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623786788112053147/5623786788112053147_signature_28.png | Norman Labis | https://www.jotform.com/uploads/macarewecare/222927978372067/5623786788112053147/5623786788112053147_signature_31.png
| 2023-06-12 | 2023-06-07 15:37 | Lung center of the Philippines | Quezon avenue, Quezon City | NCR | NCR 2nd - Quezon City | Koica | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24823 | PBSP | nilabis.pbsp@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 66 |  | GX4 | 704584 | Due for maintenance | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623786388112099704/5623786388112099704_signature_28.png | Norman Charles Labis | https://www.jotform.com/uploads/macarewecare/222927978372067/5623786388112099704/5623786388112099704_signature_31.png
| 2023-06-12 | 2023-06-09 09:30 | Jose B. Lingad Memorial General Hospital | San fernando,Pampanga | 3 | Pampanga | Laboratory | NO | Laptop replacement | CEPHEID | Noel Floresca |  |  |  |  | 24823 | COMMERCIAL | jblmgh.is@gmail.com | 26 |  | 2023-06-08 | GX4 | 812571 | Malfunction of laptop display | Delivered and installation laptop Upload ADF files Set up laptop Test run unit | Laptop replacement | Operational |  | Laptop | 5327 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623786388113652819/5623786388113652819_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5623786388113652819/5623786388113652819_signature_31.png
| 2023-06-12 | 2023-06-09 14:23 | Philippine heart center | East avenue q.c | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 24823 | PBSP | phcmolecularlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 308 ASSAY: Others, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 2212 | 2023-06-09 | GX4 | 834443 | For module replacement B1 603916 - 210063239 | Performed module, replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | Operational |  | 900-0508 | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5623786388115525450/5623786388115525450_signature_28.png | Arjay Abichuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5623786388115525450/5623786388115525450_signature_31.png
| 2023-06-09 | 2023-06-07 13:23 | Midsalip Rural health unit | Midsalip, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24821 | DOH | genaliner@rocketmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 334 | 2023-06-07 | GX 4 | 837664 | Module 4 error 2032 | Module 4 undergone Xpertcheck calibration, results is Service required. | Preventive Maintenence and Module 4 undergone Xpertcheck calibration | Module 1, 2 and 3 are available. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5621037310317585263/5621037310317585263_signature_28.png | Genaline Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5621037310317585263/5621037310317585263_signature_31.png
| 2023-06-09 | 2023-06-09 13:56 | Mindanao central sanitarium | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24822 | DOH | ybioamie@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2019 | 2023-06-09 | GX 4 | 821214 | Module 2 error 2014 and 1001 | Module 2 undergone Xpertcheck calibration and the result is Service required. | Preventive Maintenence and Module 2 undergone Xpertcheck calibration | Module 3 and 4 are available | Module 1 for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5621037310319773636/5621037310319773636_signature_28.png | Amie b ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5621037310319773636/5621037310319773636_signature_31.png
| 2023-06-09 | 2023-06-09 14:26 | Samboan RHU | Poblacion, Samboan RHU | 7 | Cebu | Tb Labodatory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24820 | DOH | sheneannesenagan@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 133 |  | Gx4 | 838513 | B1 862517 encouner Error 4012 during Selft test on April 5, 23. | Advice medtech to maintain the ambient temperature between 21 to 28 degrees Celcius | Perfomed maintenance, clean Filter, clean Icore and plunger Generate data IQR, Sys log and Assay stat, Gxx files, Back up database | All modules are operational and annual calibration was due on May 31, 2023 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5621016732636581567/5621016732636581567_signature_28.png | Shene Anne A. Senagan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5621016732636581567/5621016732636581567_signature_31.png
| 2023-06-09 | 2023-06-09 13:58 | Taguig City Health Office | Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24819 | PBSP | dumawalgina20@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2688 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-06-09 | GXIV | 814372 | Module B4 (743460) for replacement | #NAME? | Replaced B4 | Unit is operational. |  | GeneXpert Module SN: 210059802 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620998619223900851/5620998619223900851_signature_28.png | Gina Dumawal | https://www.jotform.com/uploads/macarewecare/222927978372067/5620998619223900851/5620998619223900851_signature_31.png
| 2023-06-09 | 2023-06-09 11:08 | Victoria Rural Health Unit | Brgy. Bulo Victoria Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24818 | DOH | diazrmt@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 178 ASSAY: Xpert MTB/RIF Ultra, COUNT: 725 |  | GXIV-D | 831669 | A1 and A2 temperature error For Preventive maintenance | Preventive maintenance | Performed surface cleaning Performed cleaning of cartridge bay,icore,plunger rod and filter Generate reports Archived test results Performed functionality test | Unit is operational except for modules A1,A2 and A3 | A1 and A2 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620895547649111566/5620895547649111566_signature_28.png | Alma D. Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5620895547649111566/5620895547649111566_signature_31.png
| 2023-06-09 | 2023-06-09 10:30 | Los Banos Rural Health Unit | Los Banos | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24817 | PBSP | rmbtaguic@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1059 | 2023-06-05 | GXIV | 804716 | Unit for PM. For module replacement of module B1. | Performed PM. Performed module replacement of module B1. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B1 (210091372) in replacement to defective B1 (664587). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620881913422403274/5620881913422403274_signature_28.png | Maria Josalee Cabrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5620881913422403274/5620881913422403274_signature_31.png
| 2023-06-09 | 2023-06-09 09:51 | Atimonan Rural Health Unit | Leon Guinto Blvd., Atimonan | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24816 | DOH | AMCL2008@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 354 | 2023-04-25 | GXIV - 6C | 838206 | For XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. All modules passed | All Passed | XpertCheck kit | 9933 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620873224523326986/5620873224523326986_signature_28.png | Marife M.Villamiel | https://www.jotform.com/uploads/macarewecare/222927978372067/5620873224523326986/5620873224523326986_signature_31.png
| 2023-06-09 | 2023-06-09 09:25 | Tarlac City Health Center 2 | Brgy. Matatalaib, Tarlac City | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24815 | DOH | clarizannebeltran1996@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 747 |  | GXIV-D | 831917 | Temperature error 1001 on module A4:758602 | Re-xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test reaults Performed calibration Gathered reports Performed functionality test | Unit is operational | Used 2 xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620846084511441852/5620846084511441852_signature_28.png | Clariz Anne M. Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5620846084511441852/5620846084511441852_signature_31.png
| 2023-06-08 | 2023-06-08 14:25 | Dalaguete RHU | Poblacion Dalaguete | 7 | Cebu | Tb Laboratoy | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 24814 | PBSP | ajreef5517@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 938 |  | Gx4 | 815464 | Module 1 Sn 716217 Run time error 1002 and Module B4 SN 714498 error run time error 1002 and 4009 temperatire Verification error | Adjust aircon setting, performed plunger maintenance | Reset module, Performed self test, genearate data, sys log, IQR and Gxx files | Only 1 module is used by the end user due to this run time error of Module 1, 3 and 4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620203383181975324/5620203383181975324_signature_28.png | Allyza Valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/5620203383181975324/5620203383181975324_signature_31.png
| 2023-06-08 | 2023-06-08 15:19 | Alabat Island District Hospital | Alabat, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24813 | DOH | letpiquero@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 660 |  | GXIV 6C | 831614 | Module C4 stuck cartridge | #NAME? | troubleshooting stuck cartridge | unit is operational. disadbled C4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620191144129634350/5620191144129634350_signature_28.png | Carmelita Piquero | https://www.jotform.com/uploads/macarewecare/222927978372067/5620191144129634350/5620191144129634350_signature_31.png
| 2023-06-08 | 2023-06-07 13:35 | Tampakan Rural Health Unit | Poblacion, Tampakan South Cotabato | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24811 | DOH | jmedtech_lab@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 713 |  | Gx-IV | 838544 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620173952375073653/5620173952375073653_signature_28.png | Jovi G. GELIO | https://www.jotform.com/uploads/macarewecare/222927978372067/5620173952375073653/5620173952375073653_signature_31.png
| 2023-06-08 | 2023-06-08 14:19 | Philippine Red Cross - General Santos City Blood Center | General Santos City | 12 | South Cotabato | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Jesson Tupas |  |  |  |  | 24812 | COMMERCIAL | j_fuscablo@yahoo.com | 23.2 |  |  | MCS+ | 03J030 | N/A | Perform MCS+ Preventive maintenance Perform MCS+ Calibration Perform cycle test | Perform MCS+ Preventive maintenance Perform MCS+ Calibration Perform cycle test | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620173952377247859/5620173952377247859_signature_28.png | Johnny Fuscablo,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5620173952377247859/5620173952377247859_signature_31.png
| 2023-06-08 | 2023-06-08 14:51 | Northern Mindanao Medical Center | Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24810 | PBSP | nmmcgenexpress@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 737 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7078 ASSAY: Xpert SARS-CoV-2, COUNT: 11171 ASSAY: Xpert HIV-1 Viral Load, COUNT: 546 |  | GeneXpert | 844118 | Service Required modules A1/713054, A3/625215, and C3/686629 | Replaced module new module SN: A1/210067859, A3/210059879 and C3/210063174 | Preventive maintenance Module Replacement Generate report files Database backup | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620167759641913309/5620167759641913309_signature_28.png | NMMC | https://www.jotform.com/uploads/macarewecare/222927978372067/5620167759641913309/5620167759641913309_signature_31.png
| 2023-06-08 | 2023-06-08 13:30 | Tondo Health Center | Gagalangin Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24809 | DOH | malorinamanera@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1852 |  | GXIV | 831620 | Unit for PM. For removal of stucked cartridge. | Performed PM. Performed removal of stucked cartridge. Generated reports | PM and repair. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620118359489816399/5620118359489816399_signature_28.png | Lontoc, Ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5620118359489816399/5620118359489816399_signature_31.png
| 2023-06-08 | 2023-06-08 11:07 | Talisay CHU | Cansojong Poblacion Talisay City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24808 | PBSP | marydelgella@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 350 |  | Gx4 | 815384 | Error 5011 Module 3 recurring error and Hardware failure on B4 210018733 due to Error 2014 Temperature offshot | Observed module 4 within the week | Performed self test, clean Icore, clean Restart module | Only 2 modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620044248262546072/5620044248262546072_signature_28.png | Marydel Gella | https://www.jotform.com/uploads/macarewecare/222927978372067/5620044248262546072/5620044248262546072_signature_31.png
| 2023-06-08 | 2023-06-08 12:52 | Mariveles Rural Health Unit | Municipal Compound Mariveles, Bataan | 3 | Bataan | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24807 | DOH | rhu1lab@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1072 |  | GXIV-D | 831894 | Due for calibration Error 2034 Laptop BIOS defective( start up date error) Slow PC Unlicensed Operating system | Xpertcheck, troubleshooting and Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Reset Gateway Board Performed calibration Gathered reports Performed functionality test | Unit is operational, all modules passed calibration | For PC replacement, used 4 xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620099804716103511/5620099804716103511_signature_28.png | Jossa Mae Monroe | https://www.jotform.com/uploads/macarewecare/222927978372067/5620099804716103511/5620099804716103511_signature_31.png
| 2023-06-08 | 2023-06-08 12:15 | Pontevedra RHU | Pontevedra, Negros Occidental | 6 | Negros Occidental | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24806 | DOH | pontevedra.laboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 692 |  | GX4 | 838491 | Error 2008 | Conduct PM and saves files(GXX File, IQR, System Log & Archive) | Run Xpert check calibration for module 2 only with SN 862258 | Module 1 already reported for replacement. Module 3 & 4 are in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620073794546463490/5620073794546463490_signature_28.png | Mae Ann Bonilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5620073794546463490/5620073794546463490_signature_31.png
| 2023-06-08 | 2023-06-08 11:28 | Eastern Visayas Regional Medical Center | Bgry. 93 Bagacay, Tacloban City | 8 | Leyte | Laboratory | NO | INSTALLATION | CEPHEID | John Paul Labutap |  |  |  |  | 24805 | PBSP | merl_agner62@yahoo.com.sg | 24 |  | 2023-06-08 | GXIV - 10C | 120000177 | N/a | N/a | Installation for GXIV 10C Setup all devices CPU - SN: CPH0722F39014 LCD Monitor - SN: V5GAH854 UPS - SN: 40208632105 | Unit is Operational |  | GXIV 10C, CPU, lcd monitor, Ups, | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620051175422774266/5620051175422774266_signature_28.png | Ma Merlina Vistal | https://www.jotform.com/uploads/macarewecare/222927978372067/5620051175422774266/5620051175422774266_signature_31.png
| 2023-06-08 | 2023-06-08 11:07 | Talisay CHU | Cansojong Poblacion Talisay City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24804 | PBSP | marydelgella@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 350 |  | Gx4 | 815384 | Error 5011 Module 3 recurring error and Hardware failure on B4 210018733 due to Error 2014 Temperature offshot | Observed module 4 within the week | Performed self test, clean Icore, clean Restart module | Only 2 modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5620044248262546072/5620044248262546072_signature_28.png | Marydel Gella | https://www.jotform.com/uploads/macarewecare/222927978372067/5620044248262546072/5620044248262546072_signature_31.png
| 2023-06-07 | 2023-06-07 12:43 | Medellin Rhu | Poblacion, Medellin Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24803 | DOH | rhumedellin@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1090 |  | Gx4 | 838506 | Due for Maintenance Module 2 Sn 862245 and Module 4 Sn 633038 run time error 1002 | Performed maintenance and generate datA, | Perfomed maintenance, Clean filter, clean Icore, clean plunger and clean cartridge bay | O nly 2 units are available due to run time error of modue 2 and 4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619371018816272901/5619371018816272901_signature_28.png | Raul Galano, Rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5619371018816272901/5619371018816272901_signature_31.png
| 2023-06-07 | 2023-06-07 15:16 | Quezon Medical Center | Lucena, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24802 | PBSP | qmcrtdl2015@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 508 ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2498 ASSAY: Xpert HIV-1 Viral Load, COUNT: 130 | 2023-06-01 | GXIV - 6C | 807462 | Module 4 recurring error | #NAME? | Replaced module D4: 677265 -> 210059928 | Unit is operational. UPS not giving back up |  | GX module 6C | DR# 0040973 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619318977914598057/5619318977914598057_signature_28.png | Jo-Ann Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5619318977914598057/5619318977914598057_signature_31.png
| 2023-06-07 | 2023-06-07 14:38 | Barili Rural Health Unit | Poblacion Barili,Cebu | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 24801 | PBSP | barilimainrhu@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 324 | 2023-06-07 | GX4 | 834693 | Due for Preventive Maintenance. | Perform Preventive Maintenance | Perform Preventive Maintenance | B1,B2 modules are functional ok to use. B3,B4 modules not ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619299614673742199/5619299614673742199_signature_28.png | Grace D. Cui | https://www.jotform.com/uploads/macarewecare/222927978372067/5619299614673742199/5619299614673742199_signature_31.png
| 2023-06-07 | 2023-06-07 14:35 | Rogaciano M. Mercado Memorial Hospital | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24800 | PBSP | rhia.aragon@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1986 | 2023-06-07 | GXIV | 805252 | Module A2 (643729) for replacement | #NAME? | Replaced module A2 | Unit is operational. |  | GeneXpert Modules SN: 210059813 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619296602324890731/5619296602324890731_signature_28.png | Rhia Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/5619296602324890731/5619296602324890731_signature_31.png
| 2023-06-07 | 2023-06-07 14:06 | Davao TB Reference Laboratory | DOH Compound JP Laurel Bajada Davao City | 11 | Davao del Sur | TB DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24799 | WHO | dtrlmark.pbsp@gmail.com | 29.5 |  |  | GX IV 4D 10C | 848248 | Module B3 for Replacement | Replaced the Defective Module SN 210011400 with new Module SN 210010841 | Perform Module Replacement | All Modules are operational | High Room Temperature, User is advise to adhere the recommended room temperature before operating the GeneXpert System |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619283001714363575/5619283001714363575_signature_28.png | IRENE M. ABEJUELA | https://www.jotform.com/uploads/macarewecare/222927978372067/5619283001714363575/5619283001714363575_signature_31.png
| 2023-06-07 | 2023-06-07 13:58 | Mexico Rural Health Unit I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 24798 | DOH | jasperdizon20@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 73 ASSAY: Xpert MTB/RIF Ultra, COUNT: 407 | 2023-06-07 | Genexpert IV | 831877 | N/a | N/a | Performed Maintenance Gathered Reports Performed Xpertcheck on A4 Module Performed Functionality Test | Unit is Operational | Xpertcheck Result = Service Required | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619276758228909534/5619276758228909534_signature_28.png | Jasper Dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5619276758228909534/5619276758228909534_signature_31.png
| 2023-06-07 | 2023-06-07 08:44 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24797 | DOH | ccmclaboratory01@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 179 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1221 |  | GXIV-4D | 838561 | Due for monthly PM | None | #NAME? | Waiting for module B1 allocation. All other modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619251517714095984/5619251517714095984_signature_28.png | Fatimah B. Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5619251517714095984/5619251517714095984_signature_31.png
| 2023-06-07 | 2023-06-07 13:22 | Tarlac City Health Center 2 | Matatalaib Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24796 | DOH | clarizannebeltran1996@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 17 ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 |  | GXIV-D | 831917 | For Preventive maintenance | Preventive maintenance | Genexpert maintenance -Performed surface cleaning -Performed cleaning of cartridge bay,icore,plunger rod and filter -Generate reports -Archived test results -Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619248231028256233/5619248231028256233_signature_28.png | Clariz Anne M. Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5619248231028256233/5619248231028256233_signature_31.png
| 2023-06-07 | 2023-06-07 09:12 | Dagupan CHO | Herero-Perez | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24795 | DOH | thisisformyfamily@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 85 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1236 |  | GX 4 | 838286 | Due for Preventive Maintenance | #NAME? | Performed Preventive Maintenance and Troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619198427126732311/5619198427126732311_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5619198427126732311/5619198427126732311_signature_31.png
| 2023-06-07 | 2023-06-07 09:12 | Dagupan CHO | Herero-Perez | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24794 | DOH | thisisformyfamily@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 85 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1236 |  | GX 4 | 838286 | Due for Preventive Maintenance | #NAME? | Performed Preventive Maintenance and Troubleshooting | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619198177124250827/5619198177124250827_signature_28.png | Rudolph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5619198177124250827/5619198177124250827_signature_31.png
| 2023-06-07 | 2023-06-07 08:05 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24793 | PBSP | romarmtamayo@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1457 ASSAY: Xpert SARS-CoV-2, COUNT: 3775 |  | GX 16 | 842230 | Module B1 | #NAME? | Troubleshoot module B1 and other modules | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619179327125751179/5619179327125751179_signature_28.png | Romar Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5619179327125751179/5619179327125751179_signature_31.png
| 2023-06-07 | 2023-06-07 08:05 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24792 | PBSP | romarmtamayo@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1457 ASSAY: Xpert SARS-CoV-2, COUNT: 3775 |  | GX 16 | 842230 | Module B1 | #NAME? | Troubleshoot module B1 and other modules | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619179047128051687/5619179047128051687_signature_28.png | Romar Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5619179047128051687/5619179047128051687_signature_31.png
| 2023-06-07 | 2023-06-07 09:50 | Bicol Medical Center | Concepcion Pequeñe, Naga City | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24791 | PBSP | bmc.bacteriology@gmail.com | 33 | ASSAY: Xpert HPV, COUNT: 1066 ASSAY: Xpert MTB/RIF Ultra, COUNT: 182 ASSAY: Xpert SARS-CoV-2, COUNT: 469 |  | GX4 | 110008458 | For Preventive Maintenance | Cleaned Cartridge Bays,plunger rod, filter Generated reports | Performed Preventive Maintenance | Unit is Operational | B3 with temperature error codes, for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619152571819881708/5619152571819881708_signature_28.png | Patricia Loise C. BACHILLER | https://www.jotform.com/uploads/macarewecare/222927978372067/5619152571819881708/5619152571819881708_signature_31.png
| 2023-06-07 | 2023-06-07 09:45 | Bicol Medical Center | Concepcion Pequeña, Naga City | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24790 | PBSP | bmc.bacteriology@gmail.com | 34 | ASSAY: Xpert HPV, COUNT: 1570 ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6338 ASSAY: Xpert SARS-CoV-2, COUNT: 3149 ASSAY: Xpert HIV-1 Viral Load, COUNT: 306 | 2023-06-07 | GX 16 | 842225 | For XpertCheck | XpertCheck Cleaned Cartridge Bays, Plunger Rod, Fan filter Generated Report | Performed XpertCheck & Preventive Maintenance | Unit is operational D2 failed D3 failed D4 failed | High temperature statuses of modules as shown in IQR | XpertCheck Kit | PL 010400 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619150731811012998/5619150731811012998_signature_28.png | Patricia Loise C. Bachiller | https://www.jotform.com/uploads/macarewecare/222927978372067/5619150731811012998/5619150731811012998_signature_31.png
| 2023-06-07 | 2023-06-07 09:56 | Aloguinsan Rural Health Unit | Poblacion, Aloguinsan, Cebu | 7 | Cebu | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 24789 | PBSP | limjennifer82@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 422 | 2023-06-07 | Gx2 | 834812 | Due for Xpertcheck | Perform Xpertcheck and Preventive Maintenance. | Perform Xpertcheck and Preventive Maintenance. | All modules pass for Xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5619145697274739759/5619145697274739759_signature_28.png | Jennifer L. Sayson | https://www.jotform.com/uploads/macarewecare/222927978372067/5619145697274739759/5619145697274739759_signature_31.png
| 2023-06-06 | 2023-06-06 11:16 | Eastern Pangasinan District Hospital | Brgy. Magallanes Tayug, Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24788 | PBSP | epdhlaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 557 | 2023-06-01 | GX4 | 815394 | Due for Module Replacement and Maintenance | -Cleaned Cartridge bays, Plunger Rods, I-CORE -Replaced defective module B3 864409 >>> 210059893 -Generated Reports | Performed Module Replacement and Maintenance | Unit is operational |  | GeneXpert Module 210059893 | DR No. 0040967 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618576577127032214/5618576577127032214_signature_28.png | Fergyl Escano | https://www.jotform.com/uploads/macarewecare/222927978372067/5618576577127032214/5618576577127032214_signature_31.png
| 2023-06-06 | 2023-06-06 09:09 | Rosales Rural Health Unit | Rosales Pangasinan | 1 | Pangasinan | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24787 | DOH | allan.vinluan2013@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 266 ASSAY: Xpert MTB/RIF Ultra, COUNT: 400 |  | GX 4 | 83472 | For troubleshooting of module B2 | Cleaned Cartridge Bays, I-CORE, Plunger Rods Generated reports | Performed Troubleshooting and Maintenance | Unit is operational | Module B2 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618547997122885457/5618547997122885457_signature_28.png | Allan B.Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5618547997122885457/5618547997122885457_signature_31.png
| 2023-06-06 | 2023-06-06 16:41 | Caloocan City Health Office | Caloocan, Metro Manila | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24786 | PBSP | crissyname@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3119 ASSAY: Xpert HIV-1 Viral Load, COUNT: 492 |  | GXIV 6C | 805271 | Module A4 for replacement | #NAME? | Replaced A4: 695309 -> 210059851 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618511569128808103/5618511569128808103_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5618511569128808103/5618511569128808103_signature_31.png
| 2023-06-06 | 2023-06-06 16:47 | Caloocan City Health Office | Caloocan, Metro Manila | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24785 | PBSP | crissyname@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3119 ASSAY: Xpert HIV-1 Viral Load, COUNT: 492 |  | GX IV 6C | 814154 | Module B3 for replacement | #NAME? | Replaced Module B3: 690081 -> 210063155 | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618509779127189434/5618509779127189434_signature_28.png | Maricris C. Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5618509779127189434/5618509779127189434_signature_31.png
| 2023-06-06 | 2023-06-06 13:43 | Iriga City Health Office | Sta. Cruz Sur, Iriga, Camarines Sur | 5 | Camarines Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24784 | DOH | kerov_xxx@yahoo.com | 26 |  | 2023-06-06 | GX4 | 838391 | For xpertcheck | Performed XpertCheck cleaned cartridge bays,plunger rod, filter Generated reports | xpertcheck & preventive maintenence | Unit is operational B1 - passed B2 - failed B3 - failed B4 - passed | All 4modules with temperature error codes | XpertCheck kit | 5271 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618450315719839658/5618450315719839658_signature_28.png | KEROV N. CORTAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5618450315719839658/5618450315719839658_signature_31.png
| 2023-06-06 | 2023-06-06 13:32 | Davao del Norte Provincial Hospital-Kapalong Zone | Kapalong, Davao del Norte | 11 | Davao del Norte | Laboratory | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24783 | PBSP | lot2salcedo@yahoo.com | 27.5 |  |  | GX-IV 4D | 815270 | For Calibration | Run XPC on the four Modules Generate IQR, Assay Stat and System Log | Perform XpertCheck Calibration | Module 3 for service required, other Modules Passed the Calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618448693298234591/5618448693298234591_signature_28.png | LEONA S. OLILA | https://www.jotform.com/uploads/macarewecare/222927978372067/5618448693298234591/5618448693298234591_signature_31.png
| 2023-06-06 | 2023-06-06 13:57 | Amparo Health Center | Amparo, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24782 | PBSP | jjpsjasperjoyce@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1447 | 2023-06-06 | GXIV | 838873 | Module B1 (694240) for replacement | #NAME? | Replaced B1 | Unit is operational. |  | GeneXpert Module SN: 210059825 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618404962617613095/5618404962617613095_signature_28.png | Jas Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/5618404962617613095/5618404962617613095_signature_31.png
| 2023-06-06 | 2023-06-06 13:48 | Sta Rosa Community Hospital | Sta Rosa, Olango Island Lapulapu City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24781 | DOH | ariestumz@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 150 |  | Gx4 | 838534 | Due for Monthly Maintenance | Performed maintenance, generate data, Sys Log, Assay Statistics | Performed maintenance, clean Filter, clean Icore, clean plunger | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618401252781100999/5618401252781100999_signature_28.png | Aries Tumulak, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5618401252781100999/5618401252781100999_signature_31.png
| 2023-06-06 | 2023-06-06 12:11 | Democrito O. Plaza Memorial Hospital | Brgy. Patin-ay, Prosperidad | 13 | Agusan del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 24780 | PBSP | lermabpabillan@yahoo.com | 20.4 | ASSAY: Xpert MTB/RIF, COUNT: 143 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1419 |  | GX-4 | 805335 | CPU not booting up. | CPU already reported to Cepheid. Waiting for replacement parts. | Replacing Module D3 with SN : 610085 to SN: 210067890 | All modules are working and operational Temporary Laptop was used. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618373137285612492/5618373137285612492_signature_28.png | Lerma B. Pabillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5618373137285612492/5618373137285612492_signature_31.png
| 2023-06-06 | 2023-06-06 11:46 | Santolan Super Health Center | Evangelista Ave., Pasig, Metro Manila | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24779 | PBSP | samandlet17@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3968 |  | GX IV - 6C | 838898 | Module 2 for replacement | -Checked and cleaned cartridge bays, icore slot, plunger rod and filter -Replaced module 2 -Generated reports | Replaced Module 2 - 890193 -> 210059845 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618371093419728226/5618371093419728226_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5618371093419728226/5618371093419728226_signature_31.png
| 2023-06-06 | 2023-06-06 12:02 | Cebu City Health Center | Cebu City | 7 | Cebu | Laboratory | NO | REPAIR CALIBRATION ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24778 | DOH | laureanjo@yahoo.com | 25 |  |  | Gx 4 | 838527 | 005 real time clock power loss | Replace new battery | Replace new battery adust systems settings | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618339454894385860/5618339454894385860_signature_28.png | Laurean jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5618339454894385860/5618339454894385860_signature_31.png
| 2023-06-06 | 2023-06-06 11:09 | Cebu City Health Department | Cebu City | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 24777 | DOH | jonnifercodoyceniza@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 507 | 2023-06-06 | GX4 | 829419 | Maintenance | Perform Preventive maintenance | Perform Preventive maintenance | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618334234521658464/5618334234521658464_signature_28.png | Jonnifer C. Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5618334234521658464/5618334234521658464_signature_31.png
| 2023-06-06 | 2023-06-06 10:27 | Biopath Clinical Diagnostic Afrims | Cap Building, Cebu City | 7 | Cebu | Tb Laboratory | YES | REPAIR | RBC BIOSCIENCE | Mark Efren Ramas |  |  |  |  | 24776 | MACARE-HBDC | mry_kee@yahop.com | 24 |  |  | Magcore Plus | P26R021920135 | Motor X axis problem | Clean up slide guide of the motor | Repair motor | Units are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5618279994519696811/5618279994519696811_signature_28.png | Mary Joy Kee, | https://www.jotform.com/uploads/macarewecare/222927978372067/5618279994519696811/5618279994519696811_signature_31.png
| 2023-06-05 | 2023-06-05 11:12 | Love YourSelf Cebu | 6 Avelino Morales St, Cebu City, 6000 Cebu | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 24775 | COMMERCIAL | rmission@loveyourself.ph | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 | 2023-06-01 | Gx4 | 814150 | Due for Preventive Maintenance | Perform Preventive Maintenance | Perform Preventive Maintenance | All modules functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617446095838926942/5617446095838926942_signature_28.png | RALPH PEARVIN G. MISSION, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5617446095838926942/5617446095838926942_signature_31.png
| 2023-06-05 | 2023-06-02 21:47 | Balamban Rural Health Unit | Poblacion,Tuburan,Cebu | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  |  |  |  | 24774 | DOH | climaco.mildred@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-05-30 | Gx4 | 838231 | Module B4 error code 5007 3x,2008 | Calibrate Module B4 | Calibrate Module B4 | Module B4 Pass for calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617407025834747740/5617407025834747740_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5617407025834747740/5617407025834747740_signature_31.png
| 2023-06-05 | 2023-06-01 09:07 | Romblon Provincial Hospital | M.L. Quezon St, Odiongan | 4B | Romblon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24773 | PBSP | mjfaderogaya.rmt@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 623 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GXIV-D | 807459 | Module 1 Failed in XpertCheck | Module Replacement | Performed Maintenance Replaced Module 1 Generate Reports | All Modules are Operational | 626529 -> 210091141 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361773618836663/5617361773618836663_signature_28.png | Mary Jane S. Faderogaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361773618836663/5617361773618836663_signature_31.png
| 2023-06-05 | 2023-06-01 09:26 | Romblon Provincial Hospital | M.L. Quezon St, Odiongan | 4B | Romblon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24773 | PBSP | mjfaderogaya.rmt@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 |  | GXIV-D | 847304 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | All Modules are Oprational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361773619341411/5617361773619341411_signature_28.png | Mary Jane Faderogaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361773619341411/5617361773619341411_signature_31.png
| 2023-06-05 | 2023-05-24 13:02 | RAKKK Prophet Medical Center | Maharlika Highway, Gumaca | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24771 | PBSP | rpmcimoleclab20@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 168 ASSAY: Xpert SARS-CoV-2, COUNT: 151 |  | GXIV-D | 834625 | Module Failed in XpertCheck | Module Replacement | Performed Maintenance Replaced 1 Module Generate Reports | All Modules are Operational | 784829 -> 210092154 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361663613456712/5617361663613456712_signature_28.png | KATRINA ROSE F. AUREADA | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361663613456712/5617361663613456712_signature_31.png
| 2023-06-05 | 2023-05-25 14:14 | Dasmariñas (Paliparan) City Chealth Office IV | Phase III, Paliparan | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24769 | PBSP | krlandaya25@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1871 ASSAY: Xpert HIV-1 Viral Load, COUNT: 75 ASSAY: Xpert HPV, COUNT: 42 | 2023-05-16 | GXIV-D | 838139 | Module 2 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 2 Generate Reports | All Modules are Operational | 210068735-> 210091142 | Gx Module Kit | G 5574 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361663618966921/5617361663618966921_signature_28.png | Khristofer Ralph L. Andaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361663618966921/5617361663618966921_signature_31.png
| 2023-06-05 | 2023-05-23 13:05 | Putatan Health Center | National Road, Putatan | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24768 | PBSP | apmolera@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 84 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1350 ASSAY: Xpert HIV-1 Viral Load, COUNT: 81 |  | GXIV-D | 834690 | Module 1 with Temperature Error | Module Replacement | Replaced Module 1 Gemerate Reports Performed Functionality Test | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361653613952376/5617361653613952376_signature_28.png | Aleli P. Molera | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361653613952376/5617361653613952376_signature_31.png
| 2023-06-05 | 2023-05-24 13:05 | RAKKK Prophet Medical Center | Maharlika Highway, Gumaca | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24772 | PBSP | rpmcimoleclab20@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 211 ASSAY: Xpert SARS-CoV-2, COUNT: 562 |  | GXIV-D | 838871 | 1 Module Failed in XpertCheck | Module Replacement | Performed Maintenance Replaced 1 Module Generate Reports | All Modules are Operational | 863445-> 210092155 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361663619142216/5617361663619142216_signature_28.png | Katrina Rose F. Aureada | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361663619142216/5617361663619142216_signature_31.png
| 2023-06-05 | 2023-05-26 08:06 | General Trias City Health Office | Hospital Road, General Trias | 4A | Cavite | Laboratory | NO | MAINTENANCE Install Network Switch | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24770 | PBSP | lore_montoya1169@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 198 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1728 |  | GXIV-D | 804691 | For Installation of Network Switch | Installation | Performed Maintenance Install Switch betweeon PBSP and DOH units | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361663619806634/5617361663619806634_signature_28.png | Lorena R Montoya | https://www.jotform.com/uploads/macarewecare/222927978372067/5617361663619806634/5617361663619806634_signature_31.png
| 2023-06-05 | 2023-05-30 10:30 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24767 | DOH | wcpatron11x4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2057 | 2023-05-30 | GXIV | 831534 | Unit for PM. For Xpertcheck of modules B1 and B3. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Module B3 (876732), service required. |  | 3 spare Xpertcheck | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617338702144904026/5617338702144904026_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5617338702144904026/5617338702144904026_signature_31.png
| 2023-06-05 | 2023-05-30 10:30 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24766 | DOH | wcpatron11x4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2057 | 2023-05-30 | GXIV | 831534 | Unit for PM. For Xpertcheck of modules B1 and B3. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Module B3 (876732), service required. |  | 3 spare Xpertcheck | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617338662148748580/5617338662148748580_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5617338662148748580/5617338662148748580_signature_31.png
| 2023-06-05 | 2023-06-02 13:53 | Lung Center of the Philippines | Quezon ave.,Quezon city | NCR | NCR 1st - Manila | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24765 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | MCS + | 08H188 | Due for maintenance | Cleaned unit, Checked rollers,dpm,spm,line sensor,valve,centrifuge,cuff, Checked parameters, Performed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617329103611774607/5617329103611774607_signature_28.png | Kettle Unlayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5617329103611774607/5617329103611774607_signature_31.png
| 2023-06-05 | 2023-06-02 13:53 | Lung Center of the Philippines | Quezon ave.,Quezon city | NCR | NCR 1st - Manila | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24764 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | MCS + | 08H188 | Due for maintenance | Cleaned unit, Checked rollers,dpm,spm,line sensor,valve,centrifuge,cuff, Checked parameters, Performed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617328963619116365/5617328963619116365_signature_28.png | Kettle Unlayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5617328963619116365/5617328963619116365_signature_31.png
| 2023-06-05 | 2023-06-02 14:57 | Lung Center of the Philippines | Quezon ave.,Quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24763 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | MCS + | 02D166 | Due for maintenance | Performed maintenance Cleaned unit Checked rollers,spm,dpm,line sensor,centrifuge,valve,cuff, Perfomed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617327833611497039/5617327833611497039_signature_28.png | Kettle Unlayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5617327833611497039/5617327833611497039_signature_31.png
| 2023-06-05 | 2023-06-02 14:57 | Lung Center of the Philippines | Quezon ave.,Quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24762 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | MCS + | 02D166 | Due for maintenance | Performed maintenance Cleaned unit Checked rollers,spm,dpm,line sensor,centrifuge,valve,cuff, Perfomed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5617327763614726753/5617327763614726753_signature_28.png | Kettle Unlayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5617327763614726753/5617327763614726753_signature_31.png
| 2023-06-04 | 2023-06-04 10:31 | Burauen District Hospital | Burauen leyte | 8 | Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24761 | DOH | burauenlab@gmail.com | 29 |  | 2023-05-31 | GXIV | 838551 | N/a | N/a | Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5616555698489657294/5616555698489657294_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5616555698489657294/5616555698489657294_signature_31.png
| 2023-06-04 | 2023-05-31 11:05 | SACCL | Rizal ave.,Manila | NCR | NCR 1st - Manila | Saccl | NO | ON-SITE ASSISTANCE | CEPHEID | Noel Floresca |  |  |  |  | 24760 | MACARE-HBDC | estrellesb@yahoo.com | 28 |  | 2023-05-31 | GX4 | 821257 | For replacement of workstation to laptop | Delivered and replaced workstation | Workstation replacement | Operational |  | Laptop | Hp laptop |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5616506492713620363/5616506492713620363_signature_28.png | Blessie Estrelles | https://www.jotform.com/uploads/macarewecare/222927978372067/5616506492713620363/5616506492713620363_signature_31.png
| 2023-06-02 | 2023-06-01 11:10 | Pangasinan Provincial Hospital | Bulingit San Carlos City Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24759 | PBSP | sweetsellveras08@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 145 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3047 | 2023-05-31 | GX 4 | 808283 | Due for Module Replacement | -Replaced Module 730175 >>> 210067929 and Filter -Cleaned Cartridge Bays, Plunger Rods, I-CORE -Generated Reports | Module Replacement and Maintenance | Unit is Operational |  | GeneXpert Module SN 210067929 | G 5614 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5615057840336062013/5615057840336062013_signature_28.png | Sweet Sell C. Veras | https://www.jotform.com/uploads/macarewecare/222927978372067/5615057840336062013/5615057840336062013_signature_31.png
| 2023-06-02 | 2023-06-01 11:10 | Pangasinan Provincial Hospital | Bulingit San Carlos City Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24758 | PBSP | sweetsellveras08@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 145 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3047 | 2023-05-31 | GX 4 | 808283 | Due for Module Replacement | -Replaced Module 730175 >>> 210067929 and Filter -Cleaned Cartridge Bays, Plunger Rods, I-CORE -Generated Reports | Module Replacement and Maintenance | Unit is Operational |  | GeneXpert Module SN 210067929 | G 5614 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5615057550335247445/5615057550335247445_signature_28.png | Sweet Sell C. Veras | https://www.jotform.com/uploads/macarewecare/222927978372067/5615057550335247445/5615057550335247445_signature_31.png
| 2023-06-02 | 2023-06-02 14:25 | Commonwealth Health Center | Katuparan, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24757 | DOH | commonwealthclinical@gmail.com | 31 | ASSAY: Xpert HPV, COUNT: 586 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2129 |  | GXIV - 6C | 831615 | Module 3 - frequent error Monitor flickering | -Checked and cleaned cartridge bays, plunger rod, icore slot and filter -checked monitor cables -checked module 3 system log -generated reports | #NAME? | Unit is operational Monitor (CNK832248W) flickers Module 3 under monitoring |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614994723453331516/5614994723453331516_signature_28.png | Rizza Clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5614994723453331516/5614994723453331516_signature_31.png
| 2023-06-02 | 2023-06-02 08:31 | Region 1 Medical Center | Dagupan City | 1 | Pangasinan | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Yasser Salih |  |  |  |  | 24756 | PBSP | romarmtamayo@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 71 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1440 ASSAY: Xpert SARS-CoV-2, COUNT: 3870 |  | GX 16 | 842230 | Due for Preventive Maintenance | #NAME? | Performed preventive maintenance and Troubleshooting | Unit is operational | Module B1 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614985720334714304/5614985720334714304_signature_28.png | Romar M. Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5614985720334714304/5614985720334714304_signature_31.png
| 2023-06-02 | 2023-06-02 14:46 | Bolinao Rural Health Unit | Bolinao Pqngasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24755 | DOH | nylemae14@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV-D | 838278 | CPU for replacement | Repair | Performed CPU Replacement Gathered reports Performed functionality test | Unit is operational | Used backup CPU SN:2UA5201PHX |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614979613638660674/5614979613638660674_signature_28.png | Emylyn Celzo | https://www.jotform.com/uploads/macarewecare/222927978372067/5614979613638660674/5614979613638660674_signature_31.png
| 2023-06-02 | 2023-06-02 08:29 | Region 1 Medical Center | Dagupan City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24754 | PBSP | romarmtamayo@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 58 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1701 ASSAY: Xpert SARS-CoV-2, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 220 |  | GX 4 | 804726 | For Preventive Maintenance | #NAME? | Performed Preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614979140338141166/5614979140338141166_signature_28.png | Romar M. Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5614979140338141166/5614979140338141166_signature_31.png
| 2023-06-02 | 2023-06-02 14:10 | Tabaco City Health Unit | Tabaco, Albay | 5 | Albay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 24753 | PBSP | elyjoy17rmt@gmail.com | 31 |  |  | Gx4 | 804257 | Monitor not working | Checked monitor Cleaned fan filter, fan, plunger rod, cartridge bays Generated report | Performed troubleshooting and preventive maintenance | Unit is operational D2 & D3 disabled Monitor not working | Monitor for replacement. D2(617185) temperature error code and D3(600725) temperature error code. For replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614961601717637031/5614961601717637031_signature_28.png | Ely Joy Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5614961601717637031/5614961601717637031_signature_31.png
| 2023-06-02 | 2023-06-02 13:00 | San Carlos City Health Office | Palaris St., San Carlos City, Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24752 | DOH | gbmmbg@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 70 ASSAY: Xpert MTB/RIF Ultra, COUNT: 468 |  | GX 4 | 838437 | Due for Preventive Maintenance | #NAME? | Performed preventive maintenance | Unit is Operational | Modules B1 and B2 already Reported |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614955001529145455/5614955001529145455_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5614955001529145455/5614955001529145455_signature_31.png
| 2023-06-02 | 2023-06-02 13:31 | Negros Oriental Provincial Hospital | Dumaguete City | 6 | Negros Oriental | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24751 | PBSP | rezelmarielagos@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 594 |  | Gx4 | 805275 | Module 4 SN 679876 Error 4008 | Replaced new module and generate data | Replaced new module SN 210059887 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614934663122962462/5614934663122962462_signature_28.png | Rezel Marie Lagos, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5614934663122962462/5614934663122962462_signature_31.png
| 2023-06-02 | 2023-06-02 11:57 | Soccsksargen General Hospital | Surallah | 12 | South Cotabato | Tb-Lab | YES | INSTALLATION | CEPHEID | Jesson Tupas |  |  |  |  | 24750 | PBSP | johnellamarisbautista@gmail.com | 24.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Others, COUNT: 0 |  | Gx-IV (10-C) | 120000172 | N/A | Install Genexpert machine SN: 120000172 (10-C) Install Genexpert CPU Install Genexpert monitor Install Genexpert UPS Install Genexpert wireless mouse w/ mouse pad and Keyboard Install Genexpert Printer Epson M1100 Install/import assay MTB-XDR Install/import assay MTB-ULTRA | Genexpert machine SN: 120000172 (10-C) installation | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614890556318492549/5614890556318492549_signature_28.png | JOHNELLA MARIS R. BAUTISTA | https://www.jotform.com/uploads/macarewecare/222927978372067/5614890556318492549/5614890556318492549_signature_31.png
| 2023-06-02 | 2023-06-02 12:00 | Malilipot Rural Health Unit | Brgy. 5, Malilipot, Albay | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24749 | PBSP | faith.buban2124@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2197 | 2023-04-27 | GX4 | 808280 | For module replacement | Replaced module:B4 655206>>713836 Cleaned fan,fan filter,cartridge bays,Icores, plunger rod | Replaced defective module and preventive maintenance | Unit is operational |  | Module 713836 | 40630 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614882948826747101/5614882948826747101_signature_28.png | Fatima Ana C. Buban | https://www.jotform.com/uploads/macarewecare/222927978372067/5614882948826747101/5614882948826747101_signature_31.png
| 2023-06-02 | 2023-06-02 11:46 | Weatern Pangasinan District Hospital | Alaminos Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24748 | PBSP | wpdhlab@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 272 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1460 |  | GXIV-D | 807598 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational, all modules passed calibration | Used 4 kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614880307172707052/5614880307172707052_signature_28.png | Gemarie Rabadon | https://www.jotform.com/uploads/macarewecare/222927978372067/5614880307172707052/5614880307172707052_signature_31.png
| 2023-06-02 | 2023-06-02 11:21 | San Jose Del Monte City Health Center VI | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24747 | PBSP | krizetteamorio@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2885 ASSAY: Xpert HIV-1 Viral Load, COUNT: 120 |  | GXIV | 807620 | Printer for installation | #NAME? | Installed printer | Unit is operational. Advised end user to avoid using after market gx fan filter. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614855765428341032/5614855765428341032_signature_28.png | Krizette Marie Amorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5614855765428341032/5614855765428341032_signature_31.png
| 2023-06-02 | 2023-06-02 09:05 | Felipe Abrigo Memorial Hospital | Brgy. Cantahay, Guiuan E. Samar | 8 | Eastern Samar | Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24746 | PBSP | dagantejas@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 174 ASSAY: Xpert MTB/RIF Ultra, COUNT: 722 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2023-06-02 | GXIV | 815271 | #NAME? | #NAME? | #NAME? | Unit is Operational |  | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614777101399338579/5614777101399338579_signature_28.png | Jasmin Joy D. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5614777101399338579/5614777101399338579_signature_31.png
| 2023-06-01 | 2023-06-01 14:47 | Roxas City Health Office | Bang-Bang St., Roxas City | 6 | Capiz | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 24745 | PBSP | rudy_camacho@ymail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2080 ASSAY: Xpert HIV-1 Viral Load, COUNT: 89 |  | GX4 | 805333 | Module B3 with SN 687242 need service | Conduct PM and save files | Replacement of new module for B3 with SN 210091312 | All modules are now functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614132504124607671/5614132504124607671_signature_28.png | Ruby Camacho | https://www.jotform.com/uploads/macarewecare/222927978372067/5614132504124607671/5614132504124607671_signature_31.png
| 2023-06-01 | 2023-06-01 14:19 | Mindanao central sanitarium | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24744 | DOH | ybiuamie@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1936 | 2023-06-01 | GX 4 | 821214 | Module 1 and 2 temp error 4012 and 2014 | Preventive Maintenence, data gathering, site inspection | Preventive Maintenence | Module 3 and 4 are available. Module 1 and 2 unavailable due to excessive errors | Module 1 and 2 already undergone Xpertcheck calibration last May 18, 2023 . Module 1 is waiting for replacement while 2 passed the calibration. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614103560466777706/5614103560466777706_signature_28.png | Amie b. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5614103560466777706/5614103560466777706_signature_31.png
| 2023-06-01 | 2023-06-01 14:13 | Sultan kudarat Provincial Hospital | Isulan, Sultan kudarat | 12 | Sultan Kudarat | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24743 | PBSP | charmfrias1270@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2931 | 2023-06-01 | Gx-IV | 807405 | Module 3 H/W failed | Replace module 3 SN:688969 with new module SN: 210091313 Perform Genexpert machine Preventive maintenance | Replace module 3 SN:688969 with new module SN: 210091313 Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 color SN:210091313 | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614098056118804139/5614098056118804139_signature_28.png | Charma Grace E. Frias | https://www.jotform.com/uploads/macarewecare/222927978372067/5614098056118804139/5614098056118804139_signature_31.png
| 2023-06-01 | 2023-06-01 12:55 | Sto. Domingo District Hospital | Poblacion, Sto. Domimgo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 24742 | DOH | lhaydeesel@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 980 | 2023-06-01 | Genexpert IV | 831582 | Module 2 HW Failure | Performed Xpertcheck on module 2 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational, Module 2 for Monitoring | Xpertcheck Result = Passed | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614096942529031537/5614096942529031537_signature_28.png | Myrish Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/5614096942529031537/5614096942529031537_signature_31.png
| 2023-06-01 | 2023-05-25 12:19 | Amma Jadsac District Hospital | Poblacion, Pudtol, Apayao | CAR | Apayao | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24739 | PBSP | dlgblacar.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2449 | 2023-05-25 | Genexpert IV | 807609 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 647183 - 210091324 | Genexpert Module = 210091324 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614095352523715575/5614095352523715575_signature_28.png | Daisy Lynn Grace B. Lacar | https://www.jotform.com/uploads/macarewecare/222927978372067/5614095352523715575/5614095352523715575_signature_31.png
| 2023-06-01 | 2023-05-26 11:34 | Kalinga Provincial Hospital | Bolanao, Tabuk City Kalinga | CAR | Kalinga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24740 | DOH | leombelgica139@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1339 | 2023-05-26 | Genexpert IV | 838466 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614095352525471133/5614095352525471133_signature_28.png | Micah Janelle B. Tuguinay | https://www.jotform.com/uploads/macarewecare/222927978372067/5614095352525471133/5614095352525471133_signature_31.png
| 2023-06-01 | 2023-05-26 13:02 | Kalinga Provincial Hospital | Bulanao, Tabuk City, Kalinga | CAR | Kalinga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24741 | PBSP | tuguinaymicahjanelle@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1316 | 2023-05-26 | Genexpert IV | 807622 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 794446 & 613842 - 210091167 & 210091169 | Genexpert Module =210091167 & 210091169 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614095352527974187/5614095352527974187_signature_28.png | Micah Janelle Tuguinay | https://www.jotform.com/uploads/macarewecare/222927978372067/5614095352527974187/5614095352527974187_signature_31.png
| 2023-06-01 | 2023-05-26 08:05 | Cagayan Valley Medical Center | Carig, Tuguegarao City, Cagayan | 2 | Cagayan | Bloodbank | NO | REPAIR MAINTENANCE | HAEMONETICS | Matthew Roxas |  |  |  |  | 24739 | MACARE-HBDC | ba.gomez007@gmail.com | 26 |  | 2023-01-20 | MCS+ | ISHO336 | N/a | N/a | Clean the MCS+ Replaced Rollers | Unit is Operational | N/a | 2pc Rollers | 9522 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614095352528696082/5614095352528696082_signature_28.png | Benj Angelo M. Gomez | https://www.jotform.com/uploads/macarewecare/222927978372067/5614095352528696082/5614095352528696082_signature_31.png
| 2023-06-01 | 2023-06-01 13:33 | Masambong Health Center | Capoas, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24738 | PBSP | michellebonrostro@gmail.com | 26 |  | 2023-05-24 | GXIV - 10C | 120000173 | For installation | Installation and setup of GeneXpert Setup GeneXpert Users Import MTB XDR assay Installed Printer Gathered reports Performed functionality test | Installation | Unit is operational |  | GeneXpert Machine Set | 40927 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614076762021747265/5614076762021747265_signature_28.png | Michelle Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/5614076762021747265/5614076762021747265_signature_31.png
| 2023-06-01 | 2023-06-01 13:13 | Bagumbong Health Center | Bagumbong, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24737 | DOH | jenilyn_duenas@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2432 |  | GXIV | 831907 | Module A4 for re XpertCheck | #NAME? | Performed XpertCheck | Unit is operational. A4 passed re XpertCheck. A4 under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614072015429119256/5614072015429119256_signature_28.png | Jenilyn Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5614072015429119256/5614072015429119256_signature_31.png
| 2023-06-01 | 2023-06-01 13:18 | Masambong Health Center | Capoas, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24736 | PBSP | michellebonrostro@gmail.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 856 | 2023-01-09 | GXIV 6C | 814160 | Unit for Calibration | #NAME? | Performed XpertCheck | Unit is operational B1 & B2 Passed B3 & B4 Failed |  | XpertCheck Kit | 37979 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614064482026354426/5614064482026354426_signature_28.png | Michelle D. Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/5614064482026354426/5614064482026354426_signature_31.png
| 2023-06-01 | 2023-06-01 11:17 | West Visayas State University Medical Center | Iloilo City, Iloilo | 6 | Iloilo | TB Dots | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24735 | PBSP | tvanrlujan@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 | 2023-06-01 | GXIV-4D R1 | 707531 | #NAME? | -Replace defective module C1/805098 with new module SN: 210091373 | -Module Replacement -XpertCheck Calibration on module C2, C3, & C4 -Preventive Maintenance | All four modules are functional. Module C2, C3, & C4 passed XPC. |  | GeneXpert Module 6-colors, XpertCheck-CE-5 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614024024327161551/5614024024327161551_signature_28.png | Van Ryan Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5614024024327161551/5614024024327161551_signature_31.png
| 2023-06-01 | 2023-06-01 11:21 | West Visayas State University Medical Center | Iloilo City, Iloilo | 6 | Iloilo | TB Dots | YES | CALIBRATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24734 | PBSP | tvanrlujan@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 504 |  | GXIV-4D | 804699 | None | None | #NAME? | Module B1 & B3 passed XPC; Module B2 & B4 - retry then service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614023214326446025/5614023214326446025_signature_28.png | Van Ryan Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5614023214326446025/5614023214326446025_signature_31.png
| 2023-06-01 | 2023-06-01 11:39 | Cebu TB reference lab | SRP near DPWH, cebu city | 7 | Cebu | Laboratory | NO | INSTALLATION | CEPHEID | Raymond Laurito |  |  |  |  | 24733 | PBSP | asmantiquilla.pbsp@gmail.com | 28 |  |  | Gx 4 | 120000180 | New installation | New installation | New installation | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5614002180127880555/5614002180127880555_signature_28.png | ALLAN S. MANTIQUILLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5614002180127880555/5614002180127880555_signature_31.png
| 2023-06-01 | 2023-06-01 09:00 | Bulan Rural Health Unit | Brgy. Aquino, Bulan, Sorsogon | 5 | Sorsogon | Lab | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24732 | PBSP | bing.gonzales33@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2724 |  | Gx4 | 807983 | For preventive maintenance | Cleaned cartridge bays, plunger rod, fan and fan filter Generated reports | Perfo,and preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613915719713430412/5613915719713430412_signature_28.png | Claribel G. Nacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5613915719713430412/5613915719713430412_signature_31.png
| 2023-06-01 | 2023-06-01 08:57 | Carmen Health Center | Brgy. Carmen Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24731 | DOH | murielmelchor11@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 593 |  | GeneXpert | 838442 | N/A | Preventive Maintenance Generate Report files Database back up | N/A | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613912580019301131/5613912580019301131_signature_28.png | Muriel melchor | https://www.jotform.com/uploads/macarewecare/222927978372067/5613912580019301131/5613912580019301131_signature_31.png
| 2023-06-01 | 2023-05-30 10:30 | Amai Pakpak Medical Center | Brgy. Datu Saber, Marawi City | BARMM | Lanao del Sur | Molecular Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24730 | PBSP | mohhannifa@gmail.com | 22 | ASSAY: Xpert SARS-CoV-2, COUNT: 8378 |  | GeneXpert | 834444 | H/W Fail on module B2/806377 and B4/693683 | Replaced module new module B2/210091158 and B4/210092141 | Preventive maintenance Module Replacement Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613901048905194019/5613901048905194019_signature_28.png | Moh Hannif Abdussalam | https://www.jotform.com/uploads/macarewecare/222927978372067/5613901048905194019/5613901048905194019_signature_31.png
| 2023-05-31 | 2023-05-31 15:08 | Philippine Fiber Industry Development Authority | Visayas Ave, Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | MAINTENANCE | CARL ZEISS | Patrick Los Baños |  |  |  |  | 24729 | COMMERCIAL | mksapida@philfida.da.gov.ph | 26 |  |  | STEMI 305 | 3943002734 | For Preventive Maintenance cleaning | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613275469699689287/5613275469699689287_signature_28.png | Mark Kevin C. Sapida | https://www.jotform.com/uploads/macarewecare/222927978372067/5613275469699689287/5613275469699689287_signature_31.png
| 2023-05-31 | 2023-05-31 15:17 | Philippine Fiber Industry Development Authority | Visayas Ave., Diliman, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | MAINTENANCE | CARL ZEISS | Patrick Los Baños |  |  |  |  | 24728 | COMMERCIAL | mksapida@philfida.da.gov.ph | 26 |  |  | Primo Star HD | 3153001421 | For Preventive Maintenance Cleaning | #NAME? | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613269019321597303/5613269019321597303_signature_28.png | Mark Kevin C. Sapida | https://www.jotform.com/uploads/macarewecare/222927978372067/5613269019321597303/5613269019321597303_signature_31.png
| 2023-05-31 | 2023-05-31 14:18 | Minglanilla RHU | Tiber 1, Minglanilla City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24727 | DOH | brendagonzagaparedes@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 851 |  | Gx4 | 838512 | Error 2008 Module and Error 2037 | Performed maintenance | Performed self test and performed maintenance, generate IQR, sys log and Assay stats | Only 3 modules, except module 2 waiting for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613235854214942599/5613235854214942599_signature_28.png | Brenda Paredes, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5613235854214942599/5613235854214942599_signature_31.png
| 2023-05-31 | 2023-05-31 14:00 | Juban Rural Health Unit | North Poblacion, Juban, Sorsogon | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24726 | PBSP | gaytarean@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 783 | 2023-05-31 | GX4 | 835214 | For Module Replacement | Replaced module: 789101>>210091002 Cleaned cartridge bays, plunger rods, fan Replace fan filter Generated reports | Replaced Module, Preventive Maintenance | Unit is operational |  | 210091002 | 40822 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613225303804068337/5613225303804068337_signature_28.png | Rean G. Gracilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5613225303804068337/5613225303804068337_signature_31.png
| 2023-05-31 | 2023-05-31 13:46 | Sasa District Health Center | Sasa, Davao city | 11 | Davao del Sur | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24725 | PBSP | decebendavid@gmail.com | 26.5 |  |  | GX-IV | 838889 | For Preventive maintenance User always encounter auto mail popuplogin in windows | Clean Surface Area, Cartridge Bays, Plunger Rod and Fan Filters Uninstall Windows Mail,check the keyboard for ghost touching Remove the keycap to avoid ghost touching | Perform Preventive Maintenance, Windows Troubleshooting | All Modules are operational Windows System Desktop under observation if ghosting still occurs | Removed keycap turnover to enduser |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613221308116070539/5613221308116070539_signature_28.png | Deceben V. David | https://www.jotform.com/uploads/macarewecare/222927978372067/5613221308116070539/5613221308116070539_signature_31.png
| 2023-05-31 | 2023-05-31 13:46 | Sasa District Health Center | Sasa, Davao city | 11 | Davao del Sur | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24724 | PBSP | decebendavid@gmail.com | 26.5 |  |  | GX-IV | 838889 | For Preventive maintenance User always encounter auto mail popuplogin in windows | Clean Surface Area, Cartridge Bays, Plunger Rod and Fan Filters Uninstall Windows Mail,check the keyboard for ghost touching Remove the keycap to avoid ghost touching | Perform Preventive Maintenance, Windows Troubleshooting | All Modules are operational Windows System Desktop under observation if ghosting still occurs | Removed keycap turnover to enduser |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613221038117189660/5613221038117189660_signature_28.png | Deceben V. David | https://www.jotform.com/uploads/macarewecare/222927978372067/5613221038117189660/5613221038117189660_signature_31.png
| 2023-05-31 | 2023-05-31 11:34 | Baguio General Hospital | BGH Compound Baguio City | CAR | Benguet | HACT | YES | CALIBRATION MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 24723 | MACARE-HBDC | balbahazel@gmail.com | 23 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 148 | 2023-05-17 | Gx4 | 821632 | For xpertcheck calibration | Performed xpertcheck | Performed xpertcheck Performed maintenance | 3 modules operational..1 module disable status.. | 3 modules PASSED..1 module FAILED | Xpertcheck kit | SI 95549 DR5245 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613208618836411798/5613208618836411798_signature_28.png | Balba, Hazel Gwyneth T. | https://www.jotform.com/uploads/macarewecare/222927978372067/5613208618836411798/5613208618836411798_signature_31.png
| 2023-05-31 | 2023-05-31 12:59 | Baguio General Hospital | BGH Compound Baguio City | CAR | Benguet | TB Laboratory | NO | INSTALLATION | CEPHEID | Harold Toco |  |  |  |  | 24722 | PBSP | dmgonzales.pbsp@gmail.com | 24 |  | 2023-05-31 | Gx4 10C | 120000171 | For unit installation | Performed unit installation | Performed unit installation | All 4 modules operational | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613208548832587434/5613208548832587434_signature_28.png | Donnel M. Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5613208548832587434/5613208548832587434_signature_31.png
| 2023-05-31 | 2023-05-31 13:14 | Gerona Rural Health Unit | Gerona Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24721 | DOH | huwakie15@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 114 ASSAY: Xpert MTB/RIF Ultra, COUNT: 404 |  | GXIV-D | 831693 | A1: 757222, error 4004 | Troubleshooting | Performed preventibe maintenance Troubleshoot module A1 valve position Performed valve alignment Checked module board and valve connection Performed functionality test | Unit is operational except modules A1 and A3 | A1 for reporting, valve error |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613196901689941169/5613196901689941169_signature_28.png | Artemio Nool III | https://www.jotform.com/uploads/macarewecare/222927978372067/5613196901689941169/5613196901689941169_signature_31.png
| 2023-05-31 | 2023-05-31 12:55 | Carcar City Health Office | Valladollid Carcar City | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24720 | DOH | chaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 453 |  | Gx4 | 838283 | Due for Maintenance and Runtime error 1002, air filter damage | Performed maintenance, clean filter, Install assay HIV VL | Performed maintenance, generate data IQR, Sys log, assay stat | All modules are operational, new avr installed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613187046029271271/5613187046029271271_signature_28.png | Lucille A. Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5613187046029271271/5613187046029271271_signature_31.png
| 2023-05-31 | 2023-05-31 11:37 | Muñoz City Health Office | Poblacion West, Sciene City of Muñoz, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24719 | DOH | scomrtdl@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 625 | 2023-05-31 | Genexpert IV | 831673 | Module 1 error | N/a, Valve at the center position | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module 1 for monitoring | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613139594619326191/5613139594619326191_signature_28.png | Catherine Celestino | https://www.jotform.com/uploads/macarewecare/222927978372067/5613139594619326191/5613139594619326191_signature_31.png
| 2023-05-31 | 2023-05-26 12:32 | San Jose Rural Health Office | San Jose | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24718 | DOH | sanjoserhulab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 845 | 2023-05-23 | GXIV | 821420 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck | Unit is operational. Module A1 (731600), service required. |  | Xpertcheck Kit | DR 9483 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613100837968600077/5613100837968600077_signature_28.png | Dolly Jeane Nitro | https://www.jotform.com/uploads/macarewecare/222927978372067/5613100837968600077/5613100837968600077_signature_31.png
| 2023-05-31 | 2023-05-26 12:32 | San Jose Rural Health Office | San Jose | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24717 | DOH | sanjoserhulab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 845 | 2023-05-23 | GXIV | 821420 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck | Unit is operational. Module A1 (731600), service required. |  | Xpertcheck Kit | DR 9483 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613100797966199441/5613100797966199441_signature_28.png | Dolly Jeane Nitro | https://www.jotform.com/uploads/macarewecare/222927978372067/5613100797966199441/5613100797966199441_signature_31.png
| 2023-05-31 | 2023-05-31 09:55 | Dr. Jose Fabella Memorial Hospital | Tayuman St. | NCR | NCR 1st - Manila | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24716 | COMMERCIAL | fabella.mollabgene@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 5156 ASSAY: Others, COUNT: SarsCov2 Flu RSV plus 10 | 2023-05-26 | GXIV | 834953 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module A3 (789441), Service required. |  | Xpertcheck | DR 5281 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613098597966814103/5613098597966814103_signature_28.png | Christianne Bernadette Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5613098597966814103/5613098597966814103_signature_31.png
| 2023-05-31 | 2023-05-31 10:04 | Matnog Medicare Community Hospital | Purok-3, Gadgaron, Matnog, Sorsogon | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24715 | PBSP | mernadithg@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 798 | 2023-05-31 | Gx4 | 835200 | For Module replacement | Replaced module: 789043>>210091134 Cleaned cartridge bays, plunger rod, fan filter Generated reports | Preventive maintenance and module replacement | Unit is operational |  | 210091134 | 40825 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613086673808967798/5613086673808967798_signature_28.png | Mernadith A. Garcera,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5613086673808967798/5613086673808967798_signature_31.png
| 2023-05-31 | 2023-05-31 08:29 | Benguet Provincial Health Office | La Trinidad, Brnguet | CAR | Benguet | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 24714 | PBSP | benguetphortdl@gmail.com | 24 |  | 2023-05-31 | Gx4 | 807621 | Module B1 for module replacement SN: 766856 | Performed module replacement Performed maintenance | Module replacement on B1 | All 4 modules operational | New mod SN: 210091156 | New mod 210091156 | Monarch/PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613027764812152475/5613027764812152475_signature_28.png | TILLBER T. BANIAGA JR. | https://www.jotform.com/uploads/macarewecare/222927978372067/5613027764812152475/5613027764812152475_signature_31.png
| 2023-05-31 | 2023-05-31 08:29 | Benguet Provincial Health Office | La Trinidad, Brnguet | CAR | Benguet | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 24713 | PBSP | benguetphortdl@gmail.com | 24 |  | 2023-05-31 | Gx4 | 807621 | Module B1 for module replacement SN: 766856 | Performed module replacement Performed maintenance | Module replacement on B1 | All 4 modules operational | New mod SN: 210091156 | New mod 210091156 | Monarch/PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5613027384812302243/5613027384812302243_signature_28.png | TILLBER T. BANIAGA JR. | https://www.jotform.com/uploads/macarewecare/222927978372067/5613027384812302243/5613027384812302243_signature_31.png
| 2023-05-30 | 2023-05-30 15:50 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24712 | PBSP | vincentmilestone20@gmail.com | 25.4 |  |  | GX-IV | 834440 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Module d2 encounter error, data will be sent to Cepheid for further Assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612555253315260736/5612555253315260736_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5612555253315260736/5612555253315260736_signature_31.png
| 2023-05-30 | 2023-05-30 15:43 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24711 | PBSP | vincentmilestone20@gmail.com | 24.6 |  |  | GX16 | 838756 | Module Replacement | Replaced the Defective Module SN :795584 with new Module SN 2100911500 Replaced the Defective Module SN: 866583 with new Module SN: 210091151 Replaced the Defective Module SN: 711301 with new Module SN 210091152 Replaced the Defective Module SN: 866396 with new Module SN: 210091153 Replaced the Defective Module SN: 866442 with new Module SN:210091154 Replaced the Defective Module SN: 657188 with new Module SN: 210091155 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Replacement | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612554793318408507/5612554793318408507_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5612554793318408507/5612554793318408507_signature_31.png
| 2023-05-30 | 2023-05-30 11:11 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | OPD OB Department | NO | MAINTENANCE | MEDGYN | Roi Kenneth Fagutao |  |  |  |  | 24710 | COMMERCIAL | benedict.alconera@drmc.doh.gov.ph | 25.8 |  |  | MGC-200 Cryotherapy | 11263 | For Preventive Maintenance | Attempt Preventive Maintenance | Attempt to Perform Preventive Maintenance | Unit was unable to test the functionality due to previously reported misplace / missing Colposcope and Cryotherapy parts ; Laptop with its Software and Nitrogen Oxide tank |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612554013314899523/5612554013314899523_signature_28.png | Benedict John M. Alconera | https://www.jotform.com/uploads/macarewecare/222927978372067/5612554013314899523/5612554013314899523_signature_31.png
| 2023-05-30 | 2023-05-30 11:02 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | OPD OB | NO | MAINTENANCE | MEDGYN | Roi Kenneth Fagutao |  |  |  |  | 24709 | COMMERCIAL | benedict.alconera@drmc.doh.gov.ph | 25.6 |  |  | Digital Colposcope - AL 106 | 111AEQ0004 | For Preventive Maintenance | Attempt to Perfrom Preventive Maintenance | Attempt Perfrom Preventive Maintenance | Unit was unable to test its functionality due to missing/lost parts of the Colposcope and Cryotherapy Gun System. The missing parts are the Laptop with Colposcope Software and the Nitrogen Oxide Gas Tank |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612553323312698835/5612553323312698835_signature_28.png | Benedict John M. Alconera | https://www.jotform.com/uploads/macarewecare/222927978372067/5612553323312698835/5612553323312698835_signature_31.png
| 2023-05-30 | 2023-05-30 16:33 | Sorsogon Provincial Hospital | Macabog, Sorsogon City | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24708 | PBSP | wizard682000@yahoo.com | 27 |  | 2023-05-30 | Gx4 | 835213 | Gateway replacement | Replaced gateway, reset serial number to old serial number(835213) Cleaned cartridge bays, plunger rod, fan filter Generated reports | Performed Preventive Maintenance. Replace gateway | Unit is operational |  | Gateway 755474 | PL-010497 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612528082153475873/5612528082153475873_signature_28.png | Reynaldo B. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5612528082153475873/5612528082153475873_signature_31.png
| 2023-05-30 | 2023-05-30 16:43 | Madrid District Hospital | Madrid | 13 | Surigao del Sur | RTDL | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 24707 | DOH | madridlaboratory@gmail.com | 26.9 | ASSAY: Xpert MTB/RIF, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 768 |  | GX-4 | 838439 | Temperature-related error found at Module B2 (E4017) | None | Troubleshooting and XPC at module B2 Preventive Maintenance | 3 modules working, 1 module for replacement (Service Required) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612457682913865609/5612457682913865609_signature_28.png | Rochelle J. Pueblas | https://www.jotform.com/uploads/macarewecare/222927978372067/5612457682913865609/5612457682913865609_signature_31.png
| 2023-05-30 | 2023-05-30 15:40 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24706 | PBSP | vincentmilestone20@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3002 ASSAY: Xpert HIV-1 Viral Load, COUNT: 301 ASSAY: Xpert SARS-CoV-2, COUNT: 7945 |  | Gx-IV | 805063 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612431240224749666/5612431240224749666_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5612431240224749666/5612431240224749666_signature_31.png
| 2023-05-30 | 2023-05-30 15:34 | Davao Regional Medical Center | Apokon, Tagum City | 12 | Davao del Norte | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24705 | PBSP | vincentmilestone20@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3002 ASSAY: Xpert SARS-CoV-2, COUNT: 7945 ASSAY: Xpert HIV-1 Viral Load, COUNT: 301 |  | Gx-IV | 834626 | Module 1 Multiple temperature error | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612430180222752225/5612430180222752225_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5612430180222752225/5612430180222752225_signature_31.png
| 2023-05-30 | 2023-05-23 11:16 | East Avenue Medical Center | East Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24704 | MACARE-HBDC | eamcmolecularlab@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 560 ASSAY: Xpert SARS-CoV-2, COUNT: 2088 ASSAY: Xpert SARS-CoV-2, COUNT: 2 | 2023-05-16 | GX4- 10C | 110014085 | Due for Calibration and Maintenance | #NAME? | XpertCheck and Maintenance | Unit is operational All modules passed XpertCheck | Work With Noel Floresca | XpertCheck Kits | G5564 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606935970415750657/5606935970415750657_signature_28.png | ANN GYLENE I. SISANTE | https://www.jotform.com/uploads/macarewecare/222927978372067/5606935970415750657/5606935970415750657_signature_31.png
| 2023-05-30 | 2023-05-30 15:04 | Macare In-house | Aurora Grand | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24703 | COMMERCIAL | carlos.leniel.origenes@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 40 ASSAY: Xpert CT/NG, COUNT: 384 ASSAY: Xpert HPV, COUNT: 415 ASSAY: Xpert MTB/RIF, COUNT: 60 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1228 ASSAY: Xpert HCV VL, COUNT: 37 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1141 |  | GX 4 | 807128 | Due for XpertCheck | #NAME? | XpertCheck and Maintenance | Unit Is Operational All modules passed XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612399403617849805/5612399403617849805_signature_28.png | Carlos Origenes | https://www.jotform.com/uploads/macarewecare/222927978372067/5612399403617849805/5612399403617849805_signature_31.png
| 2023-05-30 | 2023-05-30 15:04 | Santa Rosa Rural Health Unit | Rizal, Santa Rosa, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24702 | DOH | addc1995@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 23 ASSAY: Xpert MTB/RIF Ultra, COUNT: 135 | 2023-05-30 | Genexpert IV | 831599 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612397512301225045/5612397512301225045_signature_28.png | Angelo Espino | https://www.jotform.com/uploads/macarewecare/222927978372067/5612397512301225045/5612397512301225045_signature_31.png
| 2023-05-30 | 2023-05-30 13:19 | Schistosomiasis Research and Training Center | Palo leyte | 8 | Leyte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24701 | PBSP | mbesplago.pbsp@gmail.com | 28 |  | 2023-05-30 | GXIV | 707521 | Unit for Calibration | #NAME? | #NAME? | Module B1,B2,B3- service required results on calibration |  | Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612340480375268200/5612340480375268200_signature_28.png | Maria Chrissa E. Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5612340480375268200/5612340480375268200_signature_31.png
| 2023-05-30 | 2023-05-30 13:08 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24700 | DOH | laguitannataniel@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1230 |  | GXIV | 831845 | Various temperature and ultrasonic horn error on modules A1 (757678), A2 (759993), A3 (786592) | #NAME? | Performed re XpertCheck | Unit is operational. All modules passed re XpertCheck. All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612330923525592044/5612330923525592044_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5612330923525592044/5612330923525592044_signature_31.png
| 2023-05-30 | 2023-05-30 12:59 | Bagong Silang Ph. 2 Health Center | Bagong Silang, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24699 | DOH | laguitannataniel@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 813 |  | GXIV | 831677 | - Monitor problem (SN: CNK83224PZ) - A1 (784136) for re XpertCheck | #NAME? | Performed troubleshooting on GeneXpert System | Unit is operational. Monitor is for reporting. Module A1 is for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612324733529192195/5612324733529192195_signature_28.png | Nataniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5612324733529192195/5612324733529192195_signature_31.png
| 2023-05-30 | 2023-05-30 12:28 | Bagong Silang Ph. 1 Health Center | Bagong Silang, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24698 | PBSP | alcaydeirene07@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1714 | 2023-05-30 | GXIV | 838856 | Module B1 (863470) for replacement | #NAME? | Replaced module B1 | Unit is operational. |  | GeneXpert Module SN: 210091375 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612303023527508791/5612303023527508791_signature_28.png | Irene Alcaide | https://www.jotform.com/uploads/macarewecare/222927978372067/5612303023527508791/5612303023527508791_signature_31.png
| 2023-05-30 | 2023-05-30 11:09 | Tarlac City Health Center X | Tibag Tarlac City | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24697 | DOH | rodelie220@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 81 ASSAY: Xpert MTB/RIF Ultra, COUNT: 547 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 |  | GXIV-D | 831587 | CPU not booting up | Troubleshooting and repair | Troubleshoot CPU Reset wire connections and Ram Gathered reports Performed functionality test | Unit is operational | CPU under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612256793529978681/5612256793529978681_signature_28.png | Rodelie C. Verayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5612256793529978681/5612256793529978681_signature_31.png
| 2023-05-30 | 2023-05-30 10:30 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24696 | DOH | wcpatron11x4@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2057 | 2023-05-30 | GXIV | 831534 | Unit for PM. For Xpertcheck of modules B1 and B3. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Module B3 (876732), service required. |  | 3 spare Xpertcheck | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5612246550322221794/5612246550322221794_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5612246550322221794/5612246550322221794_signature_31.png
| 2023-05-29 | 2023-05-25 10:45 | Baras RHU | Baras rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24695 | PBSP | mitchz0323@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 685 | 2023-05-25 | GX4 | 815389 | For module replacement and maintenance B2 696133 - 210091157 | Performed module replacement,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900-0508 | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5611774238911277465/5611774238911277465_signature_28.png | Michelle Avela | https://www.jotform.com/uploads/macarewecare/222927978372067/5611774238911277465/5611774238911277465_signature_31.png
| 2023-05-29 | 2023-05-29 11:59 | Olongapo City Health Office | West Bajac-Bajac Olongapo City Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24694 | PBSP | roselynmaganomilan@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 582 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1126 | 2023-05-15 | GXIV-D | 804700 | D1 for module replacement | Troubleshooting Module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Troubleshoot UPS Gathered reports Performed functionality test | Unit is operational | Ups not charging for reporting | 1 genexpert module | 40827 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5611425821987339415/5611425821987339415_signature_28.png | Roselyn M. Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5611425821987339415/5611425821987339415_signature_31.png
| 2023-05-29 | 2023-05-29 10:22 | Dinalupihan Rural Health Unit 2 | Brgy Colo, Dinalupihan Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24693 | PBSP | dereggyjuswa@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 236 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1555 ASSAY: Xpert HIV-1 Viral Load, COUNT: 33 | 2023-04-27 | GXIV-D | 807561 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except modules B2 and B3 | B2 and B3 for reporting, failed xpertcheck | 6 xpertcheck kits | 40654 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5611372428025038293/5611372428025038293_signature_28.png | Joshua Deregla | https://www.jotform.com/uploads/macarewecare/222927978372067/5611372428025038293/5611372428025038293_signature_31.png
| 2023-05-29 | 2023-05-26 14:41 | Sudipen Rural Health Unit | Sudipen, La Union | 1 | La Union | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24692 | DOH | piaramuj@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 529 |  | GX 4 | 838270 | Due for ReXpertCheck and Maintenance | #NAME? | Performed maintenance and ReXpertCheck | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5611262080635733154/5611262080635733154_signature_28.png | Pia Jumar T. Ramirez | https://www.jotform.com/uploads/macarewecare/222927978372067/5611262080635733154/5611262080635733154_signature_31.png
| 2023-05-29 | 2023-05-24 11:38 | Carmen Rural Health Unit | Carmen, North Cotabato | 12 | Cotabato | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24691 | PBSP | lumawagrosalyn1@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 303 ASSAY: Xpert MTB/RIF Ultra, COUNT: 423 | 2023-05-24 | Gx-IV | 815284 | Module 4 stuck cartridge | Fixed/repair module 4 stuck cartridge Perform Genexpert machine Xpertcheck/Calibration Perform Genexpert machine Preventive maintenance | Fixed/repair module 4 stuck cartridge Perform Genexpert machine Xpertcheck/Calibration Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Xpertcheck CE -5 | Lot: 1000378669 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5611155724269949286/5611155724269949286_signature_28.png | Rosalyn L. Lumawag | https://www.jotform.com/uploads/macarewecare/222927978372067/5611155724269949286/5611155724269949286_signature_31.png
| 2023-05-28 | 2023-05-24 14:04 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | RHIVDA | NO | INSTALLATION | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24690 | MACARE-HBDC | mariaana.hizon@gmail.com | 31 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HPV, COUNT: 0 ASSAY: Xpert CT/NG, COUNT: 0 ASSAY: Xpert HCV VL, COUNT: 0 ASSAY: Xpert HBV VL, COUNT: 0 | 2023-05-19 | GXIV-D | 110017894 | Machine setup and installation | Installation | Machine installation and setup Gathered reports Performed functionality test | Unit is operational |  | Genexpert machine | 5256 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5610680264329997443/5610680264329997443_signature_28.png | Maria Ana Hizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5610680264329997443/5610680264329997443_signature_31.png
| 2023-05-28 | 2023-05-24 14:04 | Bataan General Hospital and Medical Center | Balanga Bataan | 3 | Bataan | RHIVDA | NO | INSTALLATION | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24689 | MACARE-HBDC | mariaana.hizon@gmail.com | 31 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HPV, COUNT: 0 ASSAY: Xpert CT/NG, COUNT: 0 ASSAY: Xpert HCV VL, COUNT: 0 ASSAY: Xpert HBV VL, COUNT: 0 | 2023-05-19 | GXIV-D | 110017894 | Machine setup and installation | Installation | Machine installation and setup Gathered reports Performed functionality test | Unit is operational |  | Genexpert machine | 5256 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5610680054321988312/5610680054321988312_signature_28.png | Maria Ana Hizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5610680054321988312/5610680054321988312_signature_31.png
| 2023-05-26 | 2023-05-23 15:30 | Philippine blood center | Quezon avenue., Quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24688 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15H029G | For maintenance and calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | Maintenance and Calibration | Operational | Work with yasser |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5609107597404372754/5609107597404372754_signature_28.png | Ivan | https://www.jotform.com/uploads/macarewecare/222927978372067/5609107597404372754/5609107597404372754_signature_31.png
| 2023-05-26 | 2023-05-23 15:33 | Philippine blood center | Quezon avenue., Quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24687 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15H028G | For maintenance and calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | Maintenance and Calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5609106947407783795/5609106947407783795_signature_28.png | Ivan | https://www.jotform.com/uploads/macarewecare/222927978372067/5609106947407783795/5609106947407783795_signature_31.png
| 2023-05-26 | 2023-05-23 15:37 | Philippine blood center | Quezon avenue.,Quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24686 | MACARE-HBDC | strobecatcher@gmail.com | 28 |  |  | MCS + | 15030G | For maintenance and calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | Maintenance and Calibration | Operational | Work with yasser |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5609106297401581839/5609106297401581839_signature_28.png | Ivan | https://www.jotform.com/uploads/macarewecare/222927978372067/5609106297401581839/5609106297401581839_signature_31.png
| 2023-05-26 | 2023-05-26 08:58 | Polillo RHU | San jose st.,Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24685 | PBSP | villaverde.thea@gmail.com | 31 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 666 | 2023-05-26 | GX4 | 815388 | For module replacement and maintenance | Performed module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | Genexpert 6 color module | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5609105177405748504/5609105177405748504_signature_28.png | Thea Villaverde | https://www.jotform.com/uploads/macarewecare/222927978372067/5609105177405748504/5609105177405748504_signature_31.png
| 2023-05-26 | 2023-05-26 08:58 | Polillo RHU | San jose st.,Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24684 | PBSP | villaverde.thea@gmail.com | 31 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 666 | 2023-05-26 | GX4 | 815388 | For module replacement and maintenance | Performed module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | Genexpert 6 color module | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5609104507406532026/5609104507406532026_signature_28.png | Thea Villaverde | https://www.jotform.com/uploads/macarewecare/222927978372067/5609104507406532026/5609104507406532026_signature_31.png
| 2023-05-26 | 2023-05-24 15:04 | University of Perpetual Help Dalta Medical Center | Alabang Zapote rd.,Pamplona, Las Piñas City | NCR | NCR 4th - Las Piñas | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24683 | COMMERCIAL | uphdmc_lab@yahoo.com | 29 |  | 2023-05-09 | MCS + | 11J220 | Due for Calibration and Maintenance | cleaned unit, checked valve,rollers,adjust line sensor value ,checked display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked program parameters, checked and adjust pressure to 300mmhg | Calibration and Maintenance | Operational | Work with Yasser Salih | Urethane rollers | 9983 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5609103827408335628/5609103827408335628_signature_28.png | Therese Omega | https://www.jotform.com/uploads/macarewecare/222927978372067/5609103827408335628/5609103827408335628_signature_31.png
| 2023-05-26 | 2023-05-26 11:20 | Cristino Paragas Memorial Hospital | Quiniput, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24682 | PBSP | fatimaabegailabdurahman@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 53 | 2023-05-26 | GX 4 | 815269 | System unit keep on restarting | Replaced System unit | Installed System unit with serial number CPH0622F37113 | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5609006674029630655/5609006674029630655_signature_28.png | Fatima Abegail F. Abdurahman | https://www.jotform.com/uploads/macarewecare/222927978372067/5609006674029630655/5609006674029630655_signature_31.png
| 2023-05-26 | 2023-05-26 11:20 | Cristino Paragas Memorial Hospital | Quiniput, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24681 | PBSP | fatimaabegailabdurahman@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 53 | 2023-05-26 | GX 4 | 815269 | System unit keep on restarting | Replaced System unit | Installed System unit with serial number CPH0622F37113 | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5609006544024586531/5609006544024586531_signature_28.png | Fatima Abegail F. Abdurahman | https://www.jotform.com/uploads/macarewecare/222927978372067/5609006544024586531/5609006544024586531_signature_31.png
| 2023-05-26 | 2023-05-26 16:08 | Camarin Health Center | Camarin, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24680 | PBSP | cindy.carpo06@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1461 |  | GXIV | 815395 | UPS not powering on (SN: 40223411704) | - powered on UPS (UPS' switch is hard to press) * back up mode: 150V * output: 230V - generated reports - checked and cleaned cartridge bay, plunger rod, filter | Checked UPS' functionality | Unit is operational. UPS seems operational but difficult to power on. UPS' switch is recommended for repair/replacement. B3 for reporting (temp. Errors found). |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608985719429311166/5608985719429311166_signature_28.png | Cindy Carpo | https://www.jotform.com/uploads/macarewecare/222927978372067/5608985719429311166/5608985719429311166_signature_31.png
| 2023-05-26 | 2023-05-26 11:20 | Cristino Paragas Memorial Hospital | Quiniput, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24679 | PBSP | fatimaabegailabdurahman@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 53 | 2023-05-26 | GX 4 | 815269 | System unit keep on restarting | Replaced System unit | Installed System unit | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608887754028660301/5608887754028660301_signature_28.png | Fatima Abegail F. Abdurahman | https://www.jotform.com/uploads/macarewecare/222927978372067/5608887754028660301/5608887754028660301_signature_31.png
| 2023-05-26 | 2023-05-26 13:04 | Reproductive Health and Wellness Center | Jacinto St. Poblacion Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24678 | MACARE-HBDC | rhwclab@gmail.com | 24.5 |  |  | GX-IV 4D | 839236 | Module B2 Encounter Error | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Perform Functionality Test Run XPC for error verification | Perform Module Troubleshooting and Preventive maintenance | Module B2 (SN:676723) for service required | Only 2 Modules are operational |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608876296667504249/5608876296667504249_signature_28.png | Elaind malano | https://www.jotform.com/uploads/macarewecare/222927978372067/5608876296667504249/5608876296667504249_signature_31.png
| 2023-05-26 | 2023-05-26 12:31 | University of Santo Tomas Hospital | Sampaloc, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Cienne Mae Valerio |  |  |  |  | 24677 | CONSORTIUM | ust_imp@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 798 | 2023-05-24 | GX IV | 110000711 | Errors on module 2 and3 | Replaced Module 2 and 3 Cleaned Filter Cleaned bay Plunger Rod Maintenance Generated Reports | Performed Module Replacement Performed Preventive Maintenance | Unit is Operational | 21004322 > 682156; 21000477 > 715805 | 900-GX | 5269 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608863644819322936/5608863644819322936_signature_28.png | Lourdes fajardo | https://www.jotform.com/uploads/macarewecare/222927978372067/5608863644819322936/5608863644819322936_signature_31.png
| 2023-05-26 | 2023-05-26 12:32 | San Jose Rural Health Office | San Jose | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24676 | DOH | sanjoserhulab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 845 | 2023-05-23 | GXIV | 821420 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck | Unit is operational. Module A1 (731600), service required. |  | Xpertcheck Kit | DR 9483 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608859305804878742/5608859305804878742_signature_28.png | Dolly Jeane Nitro | https://www.jotform.com/uploads/macarewecare/222927978372067/5608859305804878742/5608859305804878742_signature_31.png
| 2023-05-26 | 2023-05-26 12:25 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24675 | MACARE-HBDC | laboratory@loveyourself.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 192 ASSAY: Xpert MTB/RIF Ultra, COUNT: 253 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1372 |  | GXIV | 807956 | Troubleshooting of modules | #NAME? | Checked System Log | Unit is operational. A1: temperature errors found B3: assay specific error |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608851519427081675/5608851519427081675_signature_28.png | Ivan Rhyan B. Tee | https://www.jotform.com/uploads/macarewecare/222927978372067/5608851519427081675/5608851519427081675_signature_31.png
| 2023-05-26 | 2023-05-26 12:19 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24674 | PBSP | cdh.laboratorydept@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 56 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1053 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 |  | GXIV-D | 807562 | Gateway board for replacement | Repair Preventive maintenance | Performed surface cleaning Performed preventive maintenance Performed gateway board replacement Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608843413017425072/5608843413017425072_signature_28.png | Rhesa Bucu | https://www.jotform.com/uploads/macarewecare/222927978372067/5608843413017425072/5608843413017425072_signature_31.png
| 2023-05-26 | 2023-05-26 12:06 | Torrijos Rural Health Unit | Torrijos, Marinduque | 4B | Marinduque | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24673 | MACARE-HBDC | jm.constantino20@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 239 |  | GX IV 6C | 847724 | For preventive maintenance | Checked and cleaned cartridge bays, plunger rod, icore slot, and filters | performed preventive maintenance | Ubit is operational. 4 modules working properly |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608834860819990861/5608834860819990861_signature_28.png | JOANNA MAY CONSTANTINO | https://www.jotform.com/uploads/macarewecare/222927978372067/5608834860819990861/5608834860819990861_signature_31.png
| 2023-05-26 | 2023-05-26 11:56 | Manuel J. Santos Hospital | 554 Montilla Blvd. Butuan City | 13 | Agusan del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 24672 | COMMERCIAL | molecularlab@mjsantos.com.ph | 21.27 | ASSAY: Xpert SARS-CoV-2, COUNT: 4215 |  | GX-4 | 831292 | None | None | Replacing Module A3 with SN: 210068313 to SN: 210091161 Replacing Module A4 with SN: 210070803 to SN: 210091164 | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608829639424319190/5608829639424319190_signature_28.png | Maria Kycel Dela Torre | https://www.jotform.com/uploads/macarewecare/222927978372067/5608829639424319190/5608829639424319190_signature_31.png
| 2023-05-26 | 2023-03-31 09:11 | University of Santo Tomas | Sampaloc, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Cienne Mae Valerio |  |  |  |  | 24671 | CONSORTIUM | ustimp@yahoo.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 720 | 2023-03-20 | GXIV | 110000711 | Unit due for Xpertcheck | Performed Xpertcheck and Preventive Maintenance | Performed Maintenance Cleaned filter Cleaned Icore Cleaned Cartridge Bay Plunger Maintenance Gathered Reports | Unit is Operational. All Modules Passed. |  | XPERTCHECK-CE-5 | 9721 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608814824814542702/5608814824814542702_signature_28.png | Sherleen calaunan | https://www.jotform.com/uploads/macarewecare/222927978372067/5608814824814542702/5608814824814542702_signature_31.png
| 2023-05-26 | 2023-04-20 09:55 | Siargao Island Medical Center | Dapa, Surigao del Norte | 10 | Surigao del Norte | TB DOTS | NO | REPAIR | CEPHEID | Cienne Mae Valerio |  |  |  |  | 24670 | DOH | neilvpcongreso17@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 56 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1285 ASSAY: Xpert SARS-CoV-2, COUNT: 1798 | 2023-03-28 | GX IV | 838444 | Module B1 and B2 not operational. | Performed Preventive Maintenanace Cleaned Cartridge Bay Cleaned Icore Cleaned Filter Gathered Reports Performed Functionality Test | Performed Module Replacement on B1 and B2 | UNIT IS OPERATIONAL 4 MODULES WORKING | B1 665231 - 679741 B2 777837 - 729817 | 900-GX | 9840, 9842 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608814824815898317/5608814824815898317_signature_28.png | Neil vesper P. Congreso | https://www.jotform.com/uploads/macarewecare/222927978372067/5608814824815898317/5608814824815898317_signature_31.png
| 2023-05-26 | 2023-05-26 10:31 | Batangas Medical Center | Kumintang Ibaba, Batangas | 4A | Batangas | Covid Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24669 | PBSP | batmc.tbculturelab@gmail.comm | 25 |  | 2023-05-23 | GXIV | 814153 | Unit for PM. For installation of new CPU. | Performed PM. Performed installation of new CPU. Generated reports. | Preventive maintenance and installation. | Unit is operational. |  | Lenovo Thinkcentre CPU (CPH0622F38432) in replacement to defective HP Z240 SFF CPU (2UA703297X). | G5597 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608796036268139173/5608796036268139173_signature_28.png | Marienan Marasigan | https://www.jotform.com/uploads/macarewecare/222927978372067/5608796036268139173/5608796036268139173_signature_31.png
| 2023-05-26 | 2023-05-25 12:50 | Bolina Rural Health Unit | Bolinao Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24667 | DOH | nylemae14@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 360 |  | GXIV-D | 838278 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | 6 xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608783993012687684/5608783993012687684_signature_28.png | Emylyn Celzo | https://www.jotform.com/uploads/macarewecare/222927978372067/5608783993012687684/5608783993012687684_signature_31.png
| 2023-05-26 | 2023-05-26 10:14 | Tarlac Provincial Hospital | Hospital Drive Brgy. San Vicente Tarlac City | 3 | Tarlac | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24668 | PBSP | molecularlab.tph@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2023-05-24 | GXIV-D 10C | 120000179 | Genexpert machine installation | Installation | Installation and setup of Genexpert machine Gathered reports Performed functionality test | Unit is operational |  | Genexpert machine | 40926 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608784003013547911/5608784003013547911_signature_28.png | Camille Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5608784003013547911/5608784003013547911_signature_31.png
| 2023-05-26 | 2023-05-26 09:39 | Potrero Health Center | Potrero, Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24666 | PBSP | ailyn.santos19@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2055 |  | GXIV | 804715 | B2 for further troubleshooting | #NAME? | Performed troubleshooting steps from Cepheid | Unit is operational. B2 is under observation. B4 is also under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608751279423473097/5608751279423473097_signature_28.png | Marilyn Laureño Damayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5608751279423473097/5608751279423473097_signature_31.png
| 2023-05-26 | 2023-05-23 09:04 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta. Cruz, Manila | NCR | NCR 1st - Manila | COVID Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24665 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 214 ASSAY: Xpert MTB/RIF Ultra, COUNT: 106 ASSAY: Xpert SARS-CoV-2, COUNT: 312 | 2023-05-22 | GX IV | 110014099 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. All modules passed. |  | XpertCheck Kit | G 5565 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608749109422768479/5608749109422768479_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5608749109422768479/5608749109422768479_signature_31.png
| 2023-05-26 | 2023-05-23 09:04 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta. Cruz, Manila | NCR | NCR 1st - Manila | COVID Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24664 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 214 ASSAY: Xpert MTB/RIF Ultra, COUNT: 106 ASSAY: Xpert SARS-CoV-2, COUNT: 312 | 2023-05-22 | GX IV | 110014099 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. All modules passed. |  | XpertCheck Kit | G 5565 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608748829429163466/5608748829429163466_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5608748829429163466/5608748829429163466_signature_31.png
| 2023-05-26 | 2023-05-25 15:42 | Angono Rural Health Unit | Angono, Rizal | 4A | Rizal | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24663 | PBSP | siennablanco@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 857 |  | GXIV | 819864 | UPS for troubleshooting | - tested UPS's output - tested back up mode - generated reports | Checked UPS | Unit is operational. Unit is directly plugged on outlet. UPS is not powering on. UPS for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608748119426572858/5608748119426572858_signature_28.png | Sienna Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5608748119426572858/5608748119426572858_signature_31.png
| 2023-05-26 | 2023-05-24 12:35 | Maysan 3S Health Station | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24662 | PBSP | neldimatera@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 814151 | Power supply error on all modules (4008) | - checked power input on power cord (227 V) - checked power input on transformer (225 V) - checked fuse if busted (fuse is ok) - checked DC Output (0 V) *busted capacitor | Checked power supply | Unit is not operational. PSU for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608741339421872490/5608741339421872490_signature_28.png | Cecilia Guardiano | https://www.jotform.com/uploads/macarewecare/222927978372067/5608741339421872490/5608741339421872490_signature_31.png
| 2023-05-26 | 2023-05-24 12:35 | Maysan 3S Health Station | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24661 | PBSP | cesguardiano@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 814151 | Power supply error on all modules (4008) | - checked power input on power cord (227 V) - checked power input on transformer (225 V) - checked fuse if busted (fuse is ok) - checked DC Output (0 V) *busted capacitor | Checked power supply | Unit is not operational. PSU for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608741049429719233/5608741049429719233_signature_28.png | Cecilia Guardiano | https://www.jotform.com/uploads/macarewecare/222927978372067/5608741049429719233/5608741049429719233_signature_31.png
| 2023-05-26 | 2023-05-25 12:06 | Binangonan Rural Health Unit | Binangonan, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24660 | PBSP | jen_cerezola@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1411 | 2023-05-23 | GXIV | 838199 | Modules B1 (807322), B3 (662510), C3 (784721) for replacement | #NAME? | Replaced B1, B3, C3 | Unit is operational. |  | GeneXpert Module SN: 210091186, 210091187, 210091188 | G 5600 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608739889425181375/5608739889425181375_signature_28.png | Jennifer Cerezola | https://www.jotform.com/uploads/macarewecare/222927978372067/5608739889425181375/5608739889425181375_signature_31.png
| 2023-05-26 | 2023-05-25 12:06 | Binangonan Rural Health Unit | Binangonan, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24659 | PBSP | jen_cerezola@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1411 | 2023-05-23 | GXIV | 838199 | Modules B1 (807322), B3 (662510), C3 (784721) for replacement | #NAME? | Replaced B1, B3, C3 | Unit is operational. |  | GeneXpert Module SN: 210091186, 210091187, 210091188 | G 5600 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608739639429814312/5608739639429814312_signature_28.png | Jennifer Cerezola | https://www.jotform.com/uploads/macarewecare/222927978372067/5608739639429814312/5608739639429814312_signature_31.png
| 2023-05-26 | 2023-05-26 09:07 | Lapulapu City Health Office | Lapulapu City Health Office | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24658 | PBSP | tanya_pax@yahoo.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 828 |  | Gx4 | 808637 | Temperature related problem Error 1002 | Performed maintenance, generate data Install intenet set up, install chrome | Performed maintenance, generate IQR, Sys Log, Gxx files and Assay statistics Install internet in the cpu unit | All modules are operationals |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608730841967586876/5608730841967586876_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5608730841967586876/5608730841967586876_signature_31.png
| 2023-05-26 | 2023-05-26 08:32 | Tacloban City Health Office | Tacloban City | 8 | Leyte | Tb dots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  |  |  |  | 24657 | DOH | ammegzedd@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 307 ASSAY: Xpert MTB/RIF Ultra, COUNT: 981 | 2023-05-26 | GXIV | 838421 | PC monitor problem SN: CNK9430B0R | #NAME? | #NAME? | Unit is Operational | PC monitor defective - Dead pixels on Liquid Crystal Display (LCD) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608710630016630611/5608710630016630611_signature_28.png | Gemma malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5608710630016630611/5608710630016630611_signature_31.png
| 2023-05-26 | 2023-05-24 08:35 | Laoang Rural Health Unit | Laoang Northern Samar | 8 | Northern Samar | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24656 | DOH | aliceuytan@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 322 ASSAY: Xpert MTB/RIF Ultra, COUNT: 707 | 2023-05-24 | GXIV | 838202 | Unit for Preventive maintenance | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608700740011162202/5608700740011162202_signature_28.png | Alice tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5608700740011162202/5608700740011162202_signature_31.png
| 2023-05-26 | 2023-05-24 08:35 | Laoang Rural Health Unit | Laoang Northern Samar | 8 | Northern Samar | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24655 | DOH | aliceuytan@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 322 ASSAY: Xpert MTB/RIF Ultra, COUNT: 707 | 2023-05-24 | GXIV | 838202 | Unit for Preventive maintenance | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608700470012171505/5608700470012171505_signature_28.png | Alice tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5608700470012171505/5608700470012171505_signature_31.png
| 2023-05-26 | 2023-05-24 10:08 | San Roque Rural Health Unit | San Roque Northern Samar | 8 | Northern Samar | TB dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24654 | DOH | vfmorales1970@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 145 ASSAY: Xpert MTB/RIF Ultra, COUNT: 583 | 2023-05-24 | GXIV | 838498 | Unit for Preventive Maintenance | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608699280015281380/5608699280015281380_signature_28.png | Vivian Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/5608699280015281380/5608699280015281380_signature_31.png
| 2023-05-26 | 2023-05-24 10:08 | San Roque Rural Health Unit | San Roque Northern Samar | 8 | Northern Samar | TB dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24653 | DOH | vfmorales1970@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 145 ASSAY: Xpert MTB/RIF Ultra, COUNT: 583 | 2023-05-24 | GXIV | 838498 | Unit for Preventive Maintenance | N/a | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608698910019565527/5608698910019565527_signature_28.png | Vivian Morales | https://www.jotform.com/uploads/macarewecare/222927978372067/5608698910019565527/5608698910019565527_signature_31.png
| 2023-05-26 | 2023-05-25 11:08 | Samar Provincial Hospital | Catbalogan Samar | 8 | Samar | Tb dots | NO | INSTALLATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24652 | PBSP | leslieleones36@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 836 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1301 | 2023-05-25 | GXIV | 808291 | Defective LCD monitor SN: 509NTKFGR189 | #NAME? | #NAME? | Unit is Operational |  | LCD monitor (ThinkVision sn: V5GTL380) | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608698510012111364/5608698510012111364_signature_28.png | Leslie daryl leones | https://www.jotform.com/uploads/macarewecare/222927978372067/5608698510012111364/5608698510012111364_signature_31.png
| 2023-05-26 | 2023-05-25 11:08 | Samar Provincial Hospital | Catbalogan Samar | 8 | Samar | Tb dots | NO | INSTALLATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24651 | PBSP | leslieleones36@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 836 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1301 | 2023-05-25 | GXIV | 808291 | Defective LCD monitor SN: 509NTKFGR189 | #NAME? | #NAME? | Unit is Operational |  | LCD monitor (ThinkVision sn: V5GTL380) | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608698260019667181/5608698260019667181_signature_28.png | Leslie daryl leones | https://www.jotform.com/uploads/macarewecare/222927978372067/5608698260019667181/5608698260019667181_signature_31.png
| 2023-05-25 | 2023-05-25 19:28 | Marinduque Provincial Hospital | Boac, Marinduque | 4B | Marinduque | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24650 | PBSP | anarvas73@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 305 ASSAY: Xpert SARS-CoV-2, COUNT: 1181 ASSAY: Xpert HIV-1 Viral Load, COUNT: 7 |  | GX IV 6C | 835087 | For preventive maintenance | Checked and cleaned cartridge bays, plunger rod, icore slot and replaced filter Generated report | Perforemed preventive maintenance | 2 unit still operational. 2 units empty |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608242652418657866/5608242652418657866_signature_28.png | Michael Lamac | https://www.jotform.com/uploads/macarewecare/222927978372067/5608242652418657866/5608242652418657866_signature_31.png
| 2023-05-25 | 2023-05-25 19:16 | Marinduque Provincial Hospital | Boac, Marinduque | 4B | Marinduque | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24649 | PBSP | anarvas73@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 198 ASSAY: Xpert SARS-CoV-2, COUNT: 1250 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-05-25 | GX IV 6C | 835084 | 2 Modules for replacement | -Checkedncartridge bays, plunger rod, icore slot and replaced filters -Replaced B1: 787972 -> 647486 -Replaced B2: 787987 ->210091205 -Generated reports | Replaced Modules B1 &B2 | 2 modules operational. 2 modules empty |  | 2 GX modules 6C | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608242212416545442/5608242212416545442_signature_28.png | Michael Lamac | https://www.jotform.com/uploads/macarewecare/222927978372067/5608242212416545442/5608242212416545442_signature_31.png
| 2023-05-25 | 2023-05-25 16:33 | Pasig City General Hospital | Maybunga, Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 24648 | PBSP | vilmarumadhay@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 92 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3022 | 2023-05-25 | GX4 | 805010 | B1 & B4 module troubleshooting | Checked gateway connections, module connections & tested GX power supply & UPS voltage Cleaned cartridge bays, plunger rod, filter Generated reports | Performed preventive maintenence & troubleshooting steps | Unit is operational B1 & B4 are disabled |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608137077122704841/5608137077122704841_signature_28.png | Vilma Umadhay | https://www.jotform.com/uploads/macarewecare/222927978372067/5608137077122704841/5608137077122704841_signature_31.png
| 2023-05-25 | 2023-05-25 16:22 | GOV. CELESTINO GALLARES MEMORIAL HOSPITA | TAGBILARAN BOHOL | 7 | Bohol | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24647 | PBSP | majieocmen@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 309 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1761 ASSAY: Xpert HIV-1 Viral Load, COUNT: 87 |  | GXIV | 838878 | MODULE B3 TEMP ERROR 1002 | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | PERFORM XPC-MODULE B3 PASSED | FOR OBSERVATION AND REPORTING |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608125006322813846/5608125006322813846_signature_28.png | Ma. Jernai Sosoban | https://www.jotform.com/uploads/macarewecare/222927978372067/5608125006322813846/5608125006322813846_signature_31.png
| 2023-05-25 | 2023-05-25 15:42 | Angono Rural Health Unit | Angono, Rizal | 4A | Rizal | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24646 | PBSP | siennablanco@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 857 |  | GXIV | 819864 | UPS for troubleshooting | - tested UPS's output - tested back up mode - generated reports | Checked UPS | Unit is operational. Unit is directly plugged on outlet. UPS is not powering on. UPS for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608100127223241255/5608100127223241255_signature_28.png | Sienna Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5608100127223241255/5608100127223241255_signature_31.png
| 2023-05-25 | 2023-05-18 16:25 | Bago City Health Office | Bago City | 6 | Negros Occidental | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24645 | DOH | limsiacoroselyn@yahoo.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 205 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1098 |  | GX4 | 83848 | Module B1 with error code 1002 | Conduct Preventive Maintenance | Run Xpert check calibration for module B1 with SN 862117 | Module B2, B3 & B4 are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608089337328756957/5608089337328756957_signature_28.png | Roselyn Rufon | https://www.jotform.com/uploads/macarewecare/222927978372067/5608089337328756957/5608089337328756957_signature_31.png
| 2023-05-25 | 2023-05-25 15:16 | BOHOL PROVINCIAL HEALTH OFFICE | TAGBILARAN | 7 | Bohol | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24644 | PBSP | aries2cris@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 953 ASSAY: Xpert MTB/RIF Ultra, COUNT: 809 |  | GXIV | 804286 | FOR MAINTENANCE AND CALIBRATION | COLLECT IQR,SYSTEM LOG,ARCHIVE,AND CALIBRATION DATA | PERFORM GX MAINTENANCE AND CLEANING PERFORM XPC: B1,B3,AND B4 FOR SERVICE B2-PASSED | EQUIPMENT OPERATIONAL, B4 IS DISABLED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608085055118677014/5608085055118677014_signature_28.png | Ma.Crisanta T.Estomago | https://www.jotform.com/uploads/macarewecare/222927978372067/5608085055118677014/5608085055118677014_signature_31.png
| 2023-05-25 | 2023-05-25 14:36 | Taytay Rural Health Unit | Taytay, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24643 | PBSP | dannelli11@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 265 | 2023-05-25 | GXIV | 812836 | Monitor for replacement | #NAME? | Replaced monitor | Unit is operational. |  | Monitor SN: V5GAC610 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608062077227793060/5608062077227793060_signature_28.png | Dino Dannelli De Los Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5608062077227793060/5608062077227793060_signature_31.png
| 2023-05-25 | 2023-05-25 14:12 | Tuburan Rural Health Unit | Poblacion,Tuburan,Cebu | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 24642 | DOH | marielgracia20@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 164 | 2023-05-23 | Gx4 | 838564 | Maintenance | Clean all modules in side and out side of the machine. | Perform Preventive Maintenance | All modules are clean and functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608046785413699888/5608046785413699888_signature_28.png | Mariel R. Gracia | https://www.jotform.com/uploads/macarewecare/222927978372067/5608046785413699888/5608046785413699888_signature_31.png
| 2023-05-25 | 2023-05-25 13:33 | Escalante City Health Office | Escalante City , Negros Occidental | 6 | Negros Occidental | Tb laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24641 | DOH | sheila.negrido@yahoo.com | 34 | ASSAY: Xpert MTB/RIF, COUNT: 278 |  | GX4 | 838523 | No reported problems | Conduct Maintenance , perfomed plunger clean up and clean filter, instruct the medtech never performed test if temperature exceeds 35 degrees C Provide additional ventilation like electric fan if no additional aircon is provided | Performed Maintenance , clean filter , generate data IQR , sys log, assay stat, plunger maintenance, check ambient temperature | Units are operational but need additional aircon |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608024976918096901/5608024976918096901_signature_28.png | Sheila Mae Negrido | https://www.jotform.com/uploads/macarewecare/222927978372067/5608024976918096901/5608024976918096901_signature_31.png
| 2023-05-25 | 2023-05-25 13:20 | Sagay City Health Office | Sagay City , Negros Occidental | 6 | Negros Occidental | Tb laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24640 | PBSP | maravelesdebra@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 231 |  | GX4 | 838144 | Error 4012 during start up of the genexpert machine 806305 | Performed maintenance, clean filter, clean plunger and adjust airconditioning | Performed self test , Get IQR , generate data 3 months , sys log , assay statistics | Module encounter module failure during start up SN # 806305 , after the reset and restart the module 806305 turn into normal condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608019656914289506/5608019656914289506_signature_28.png | Debra M. Abellar | https://www.jotform.com/uploads/macarewecare/222927978372067/5608019656914289506/5608019656914289506_signature_31.png
| 2023-05-25 | 2023-05-25 12:58 | Panabo City Health Office | Poblacion Panabo | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24639 | DOH | nadineeust67@gmail.com | 23.7 |  |  | GX-IV 4D | 838295 | Module B3 encounter temperature error | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Perform Functionality Test Run XPC for error verification | Perform Preventive Maintenance and Module Troubleshooting | Module B3 for service required | Only 1 Module is operational |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5608008591187022087/5608008591187022087_signature_28.png | Nadine Gemma Eustaquio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5608008591187022087/5608008591187022087_signature_31.png
| 2023-05-25 | 2023-05-25 12:49 | Chinese General Hospital | Blumentritt, Sta. Cruz | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  |  |  |  | 24638 | COMMERCIAL | cghmcbloodbank2016@gmail.com | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. | Performed preventive maintenance. Verified calibration parameters. Performed simulation running. | Preventive maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607995710263746577/5607995710263746577_signature_28.png | Jacqueline Zarzoso | https://www.jotform.com/uploads/macarewecare/222927978372067/5607995710263746577/5607995710263746577_signature_31.png
| 2023-05-25 | 2023-05-24 13:17 | Tondo Health Center | Gagalangin, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24637 | COMMERCIAL | ma.lorinamanera@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1663 |  | GXIV | 831620 | Unit for PM. | Performed Preventive Maintenance. Generated reports. | Preventive Maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607969020269524119/5607969020269524119_signature_28.png | Lontoc, Ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5607969020269524119/5607969020269524119_signature_31.png
| 2023-05-25 | 2023-05-24 13:17 | Tondo Health Center | Gagalangin, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24636 | COMMERCIAL | ma.lorinamanera@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1663 |  | GXIV | 831620 | Unit for PM. | Performed Preventive Maintenance. Generated reports. | Preventive Maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607968980267049104/5607968980267049104_signature_28.png | Lontoc, Ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5607968980267049104/5607968980267049104_signature_31.png
| 2023-05-25 | 2023-05-25 11:47 | Butuan Medical Center | Baan KM 5, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24635 | PBSP | bmccrcl2022@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1149 ASSAY: Xpert SARS-CoV-2, COUNT: 864 ASSAY: Xpert HIV-1 Viral Load, COUNT: 82 |  | GX-4 | 805099 | None | None | Replacing Module B3 with SN: 719843 to SN: 210092142 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607961516518459013/5607961516518459013_signature_28.png | Joechita L. Doysabas | https://www.jotform.com/uploads/macarewecare/222927978372067/5607961516518459013/5607961516518459013_signature_31.png
| 2023-05-25 | 2023-05-25 10:53 | CANDIJAY COMMUNITY HOSPITAL | CANDIJAY BOHOL | 7 | Bohol | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24634 | DOH | senenmilo24@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 475 ASSAY: Xpert MTB/RIF Ultra, COUNT: 281 |  | GX-IV | 838396 | MODULE B4 TEMPERATURE ERROR | #NAME? | #NAME? | MODULE B3&B4 DISABLED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607934179019658989/5607934179019658989_signature_28.png | Lian Rose Milo | https://www.jotform.com/uploads/macarewecare/222927978372067/5607934179019658989/5607934179019658989_signature_31.png
| 2023-05-25 | 2023-05-25 09:30 | Teresita Jalandoni Provincial Hospital | Silay, negros occidental | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 24633 | PBSP | feljplz@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 127 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2678 |  | Gx 4 | 805093 | Defective module 4 sn791669 | Clean fan filter, New module 4 installed sn210091149 | New module 4 installed sn210091149 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607886990328167721/5607886990328167721_signature_28.png | Feljelyn P. Bela-ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5607886990328167721/5607886990328167721_signature_31.png
| 2023-05-25 | 2023-05-25 08:25 | Bacarra Rural Health Unit | Bacarra, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24632 | PBSP | davidangelogamidobarrozo@gmail.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX 4 | 807430 | Due for CPU Replacement and Maintenance | Replaced Cpu Cleaned Cartridge bays, Filter, Plunger Rods and I-CORE Generated reports | Replaced CPU and Performed Maintenance | Unit is operational | Modules B1 and B4 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607857732115911075/5607857732115911075_signature_28.png | David Angelo Barrozo | https://www.jotform.com/uploads/macarewecare/222927978372067/5607857732115911075/5607857732115911075_signature_31.png
| 2023-05-25 | 2023-05-24 12:30 | San Isidro Health Center | Dr. A. Santos Ave., San Isidro, Parañaque | NCR | NCR 4th - Parañaque | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24631 | PBSP | herminadoa@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1678 ASSAY: Xpert HIV-1 Viral Load, COUNT: 110 | 2023-05-23 | GX 4 | 838192 | Due for Module Replacement and Maintenance | -Cleaned Cartridge Bays, Plunger Rods and I-CORE -Replaced Module 806897 >>>210091236 and Module 806898>>> 210091268 -Generated Reports | Performed module replacement and Maintenance | Unit is Operational | Work with Noel Floresca | GeneXpert Modules SN: 210091236 and 210091268 | G 5598 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607854802115257093/5607854802115257093_signature_28.png | Amelia G. Herminado | https://www.jotform.com/uploads/macarewecare/222927978372067/5607854802115257093/5607854802115257093_signature_31.png
| 2023-05-24 | 2023-05-24 14:34 | Tuao Rural Health Unit | Centro 2, Tuao, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24630 | DOH | annrollaine@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 79 ASSAY: Xpert MTB/RIF Ultra, COUNT: 429 | 2023-02-14 | Genexpert | 838465 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | 9625 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607286081111604322/5607286081111604322_signature_28.png | Rollaine Ann Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5607286081111604322/5607286081111604322_signature_31.png
| 2023-05-24 | 2023-05-24 10:25 | TUBIGON RURAL HEALTH UNIT | TUBIGON BOHOL | 7 | Bohol | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24629 | PBSP | josephraytan@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 931 ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 |  | GX-IV | 815431 | FOR MAINTENANCE AND CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607212485917983094/5607212485917983094_signature_28.png | Joseph Ray F. Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5607212485917983094/5607212485917983094_signature_31.png
| 2023-05-24 | 2023-05-24 14:55 | University of The East Ramon Magsaysay Memorial Hospital | Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24628 | CONSORTIUM | uerm.laboratory@gmail.com | 25 | ASSAY: Xpert C. difficile BT, COUNT: 44 ASSAY: Xpert CT/NG, COUNT: 60 ASSAY: Xpert MTB/RIF, COUNT: 345 ASSAY: Xpert MTB/RIF Ultra, COUNT: 256 ASSAY: Others, COUNT: 1 |  | GXIV | 847578 | 4plex assay definition file for importing | - imported assays (Xpert Xpress_SARS-CoV-2_Flu_RSV_5.gxa and Xpress SARS-CoV-2_Flu_RSV plus_1.gxa) - checked and cleaned cartridge bay, plunger rod, filter, icore slot - generated reports | #NAME? | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607209155555144602/5607209155555144602_signature_28.png | Gilbert R. Claveria | https://www.jotform.com/uploads/macarewecare/222927978372067/5607209155555144602/5607209155555144602_signature_31.png
| 2023-05-24 | 2023-05-24 14:40 | CONG. SIMEON TORIBIO MEMORIAL HOSPITAL | CARMEN BOHOL | 7 | Bohol | TB DOTS LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24627 | DOH | himaya.vale@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 798 ASSAY: Xpert MTB/RIF Ultra, COUNT: 167 |  | GX-IV | 838394 | MODULE B1 TEMP ERROR 1002 and 4012, MODULE B2 ERROR 4011 and 2034, MODULE B3 ERROR 1002 and 4017 | COLLECT IQR, SYSTEM LOG,AND ARCHIVE | #NAME? | #VALUE! |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607200645214174789/5607200645214174789_signature_28.png | Himaya Vale | https://www.jotform.com/uploads/macarewecare/222927978372067/5607200645214174789/5607200645214174789_signature_31.png
| 2023-05-24 | 2023-05-24 10:25 | TUBIGON RURAL HEALTH UNIT | TUBIGON BOHOL | 7 | Bohol | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24626 | PBSP | josephraytan@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 931 ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 |  | GX-IV | 815431 | FOR MAINTENANCE AND CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607195695214017797/5607195695214017797_signature_28.png | Joseph Ray F. Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5607195695214017797/5607195695214017797_signature_31.png
| 2023-05-24 | 2023-05-24 12:07 | Braulio E. Dujali Rural Health Unit | Braulio E. Dujali, Davao Del Norte | 11 | Davao del Norte | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24625 | PBSP | sblacorda@yahoo.com | 27.4 |  |  | GX-IV 4D | 835216 | Module B2 for replacement | Replaced the Defective Module B2 SN: 788796 with new Module SN : 210091166 Generate IQR, Assay Stat and System Log | Perform Module Replacement | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607158764457710588/5607158764457710588_signature_28.png | SHARON ROSE B. LACORDA | https://www.jotform.com/uploads/macarewecare/222927978372067/5607158764457710588/5607158764457710588_signature_31.png
| 2023-05-24 | 2023-05-24 11:46 | Bayawan City Health Office | T. Diao St., Bayawan City, Negros Oriental | 7 | Negros Oriental | Laboratory | YES | REPAIR | CEPHEID | Gerald Fritz Selga |  |  |  |  | 24624 | PBSP | nymphar07@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 620 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1083 |  | GX-4 | 808282 | Modules B2/701412 and B4/662360 resulted in Requiring Service during last Xpertcheck. | Preventive Maintenance, Plunger rod maintenance, filter cleaning, module bay cleaning, archive tests. | Replaced module B2 and B4 with new modules with SN 210091172 and 210091185 respectively. | Module B3 with HW Fail; for report to Cepheid. Modules B1, B2 and B4 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607147149615928297/5607147149615928297_signature_28.png | Nympha Ridad | https://www.jotform.com/uploads/macarewecare/222927978372067/5607147149615928297/5607147149615928297_signature_31.png
| 2023-05-24 | 2023-05-24 12:37 | Surigao del Norte Provincial Hospital | Bad-as, Placer | 13 | Surigao del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24623 | PBSP | biyoyojayl@gmail.com | 23.45 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1372 |  | GX-4 | 815435 | None | None | Replacing Module B1 with SN: 805830 to SN: 210091072 Replacing Module B4 with SN: 688216 to SN: 210091081 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607147018856111004/5607147018856111004_signature_28.png | Jay Biyoyo | https://www.jotform.com/uploads/macarewecare/222927978372067/5607147018856111004/5607147018856111004_signature_31.png
| 2023-05-24 | 2023-05-24 08:46 | Piat Municipal Health Office | Poblacion, Piat, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24621 | DOH | marieatalipsalas@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 76 ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 | 2023-02-14 | Genexpert IV | 838532 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibratiom Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | 9626 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607136704411005127/5607136704411005127_signature_28.png | Marie A. Salas | https://www.jotform.com/uploads/macarewecare/222927978372067/5607136704411005127/5607136704411005127_signature_31.png
| 2023-05-24 | 2023-05-24 09:59 | Nuestra Sra. De Piat District Hospital | Maguilling, Piat, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24622 | PBSP | mariagenevaligas@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 48 ASSAY: Xpert MTB/RIF Ultra, COUNT: 252 | 2023-05-24 | Genexpert II | 835091 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 787869 - 210091136 | Genexpert Module = 210091136 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607136704414556128/5607136704414556128_signature_28.png | Maria Geneva M. Ligas | https://www.jotform.com/uploads/macarewecare/222927978372067/5607136704414556128/5607136704414556128_signature_31.png
| 2023-05-24 | 2023-05-24 12:33 | Conner District Hospital | Ripang, Conner, Apayao | CAR | Apayao | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24620 | MACARE-HBDC | cdhlaboratory2019@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 0 | 2023-05-17 | Genexpert IV | 847727 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | 5244 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5607136704415021706/5607136704415021706_signature_28.png | MARK G.TALAY | https://www.jotform.com/uploads/macarewecare/222927978372067/5607136704415021706/5607136704415021706_signature_31.png
| 2023-05-24 | 2023-05-24 08:54 | Lubao Rural Health Unit 3 | Sitio Barin, San Isidro Lubao Pampanga | 3 | Pampanga | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24619 | PBSP | nathaliecoleencovar@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 42 ASSAY: Xpert MTB/RIF Ultra, COUNT: 833 | 2053-05-15 | GXIV-D | 812834 | B3 for module replacement | Module replacemwnt Preventive maintenance | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 40823 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606997028729148313/5606997028729148313_signature_28.png | Nathalie Covar | https://www.jotform.com/uploads/macarewecare/222927978372067/5606997028729148313/5606997028729148313_signature_31.png
| 2023-05-24 | 2023-05-23 11:16 | East Avenue Medical Center | East Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24618 | MACARE-HBDC | eamcmolecularlab@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 560 ASSAY: Xpert SARS-CoV-2, COUNT: 2088 ASSAY: Xpert SARS-CoV-2, COUNT: 2 | 2023-05-16 | GX4- 10C | 110014085 | Due for Calibration and Maintenance | #NAME? | XpertCheck and Maintenance | Unit is operational All modules passed XpertCheck | Work With Noel Floresca | XpertCheck Kits | G5564 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606935970415750657/5606935970415750657_signature_28.png | ANN GYLENE I. SISANTE | https://www.jotform.com/uploads/macarewecare/222927978372067/5606935970415750657/5606935970415750657_signature_31.png
| 2023-05-24 | 2023-05-23 10:29 | East Avenue Medical Center | East Avenue, Quezon City | NCR | NCR 2nd - Quezon City | TB DOTS | NO | MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24618 | PBSP | roseannparnes@yahoo.com | 32 | ASSAY: Xpert HPV, COUNT: 65 ASSAY: Xpert MTB/RIF, COUNT: 181 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3774 |  | GX4 | 803087 | Due for preventive maintenance | #NAME? | Performed preventive maintenance | Unit Is Operational | Work with Noel Floresca |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606935970418947124/5606935970418947124_signature_28.png | Rose Ann Parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5606935970418947124/5606935970418947124_signature_31.png
| 2023-05-23 | 2023-05-23 18:41 | Cadiz City Health Office | Cadiz City | 6 | Negros Occidental | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24617 | PBSP | zhevebacomo@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 164 |  | Gx4 | 807614 | Error ultrasonic horn failure and gateway board damage | Replaced gateway board | Replaced gateway board | The units are now operational. All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606479781189761175/5606479781189761175_signature_28.png | Zherlyn Bacumo | https://www.jotform.com/uploads/macarewecare/222927978372067/5606479781189761175/5606479781189761175_signature_31.png
| 2023-05-23 | 2023-05-23 16:53 | Roxas City Health Office | Roxas City, Capiz | 6 | Capiz | Tb Dots | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 24616 | PBSP | ruby_camacho@ymail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2010 |  | GXIV-4D | 805333 | Module B4 - Error 2008 H/W fail | None | #NAME? | Module B3 pending for replacement. Module B1, B2, B4 are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606414490117435310/5606414490117435310_signature_28.png | Ruby C. Lut | https://www.jotform.com/uploads/macarewecare/222927978372067/5606414490117435310/5606414490117435310_signature_31.png
| 2023-05-23 | 2023-05-17 09:30 | Dr. Cornelio T. Martinez Sr. Memorial Hospital | Kiamba, Sarangani | 12 | Sarangani | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24615 | PBSP | mamartinez.pbsp@gmail.com | 28.3 | ASSAY: Xpert MTB/RIF, COUNT: 112 ASSAY: Xpert MTB/RIF Ultra, COUNT: 365 | 2023-05-17 | Gx-IV | 812017 | N/A | Perform Genexpert machine Xpertcheck/ Calibration Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/ Calibration Perform Genexpert machine Preventive maintenance | 4 operational modules | All modules passed the Xpertcheck/Calibration | Xpertcheck CE-5 | Lot no: 1000330696 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606376629302058864/5606376629302058864_signature_28.png | Madelene R. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5606376629302058864/5606376629302058864_signature_31.png
| 2023-05-23 | 2023-05-23 15:34 | Dr. Amado Diaz Provincial Foundation Hospital | Midsayap, North Cotabato | 12 | Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24614 | PBSP | cadornaangelica0314@gmail.com | 26.9 | ASSAY: Xpert MTB/RIF, COUNT: 216 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1022 | 2023-05-23 | Gx-IV | 807623 | Module 3 and 4 for Service Required Genexpert machine modules temperature not within the acceptable range of 15-30°C | Replace module 3 SN: 662097 with new module SN: 210092161 Replace module 3 SN: 662098 with new module SN: 210085010 Perform Genexpert machine Preventive maintenance | Replace module 3 SN: 662097 with new module SN: 210092161 Replace module 3 SN: 662098 with new module SN: 210085010 Perform Genexpert machine Preventive maintenance | 4 Operational modules | Genexpert machine need to install air-conditioning unit for required temperature of all module | Genexpert module 6 color | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606376629302274185/5606376629302274185_signature_28.png | Angelica cadorna | https://www.jotform.com/uploads/macarewecare/222927978372067/5606376629302274185/5606376629302274185_signature_31.png
| 2023-05-23 | 2023-05-23 15:32 | Danao City Health Office | Danao city, cebu | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 24613 | PBSP | faiha2@gmail.com | 23 |  |  | Gx 4 | 808286 | Dedective module b3 sn | New module b3 installed sn210091171, clean gan filter | New module b3 installed sn210091171 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606370555914384302/5606370555914384302_signature_28.png | Janice Ferrariz | https://www.jotform.com/uploads/macarewecare/222927978372067/5606370555914384302/5606370555914384302_signature_31.png
| 2023-05-23 | 2023-05-23 14:48 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24612 | PBSP | spmcrhivda@gmail.com | 23.8 |  |  | GX-IV 4D | 808298 | Module B2 for module replacement | Replaced the Defective Module B2 SN:603064 with new Module SN: 210091170 | Perform Module Replacement | 3 modules are operational | Module B1 encounter multiple error, data sent to Cepheid for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606341287808814625/5606341287808814625_signature_28.png | Flordeliza Go, RMT MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/5606341287808814625/5606341287808814625_signature_31.png
| 2023-05-23 | 2023-05-23 13:54 | Sen. Gerardo M. Roxas Memorial District Hospital | Dao, Capiz | 6 | Capiz | TB Dots | YES | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 24611 | PBSP | bermendirowena@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1459 | 2023-05-23 | GXIV-4D | 815280 | Module B1 & B4 - Xpert Check service required | -Replace defective module B1/695246 with new module SN: 210088787 -Replace defective module B4/670432 with new module SN: 210089518 | #NAME? | All modules are operational. |  | GeneXpert Module 6-Colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606310296117441422/5606310296117441422_signature_28.png | Rowena L. Bermendi | https://www.jotform.com/uploads/macarewecare/222927978372067/5606310296117441422/5606310296117441422_signature_31.png
| 2023-05-23 | 2023-05-23 13:26 | Northern Samar Provincial Hospital | Catarman Northern Samar | 8 | Northern Samar | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24610 | PBSP | dinahada70@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 640 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1744 ASSAY: Xpert HIV-1 Viral Load, COUNT: 29 | 2023-05-23 | GXIV | 807566 | #NAME? | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606296408856176772/5606296408856176772_signature_28.png | Dinah F. Ada | https://www.jotform.com/uploads/macarewecare/222927978372067/5606296408856176772/5606296408856176772_signature_31.png
| 2023-05-23 | 2023-05-23 11:02 | University of The Philippines - National Institutes of Health | Pedro Gil St. Ermita, Manila | NCR | NCR 1st - Manila | Central Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  |  |  |  | 24609 | MACARE-HBDC | vvluzentales@up.edu.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GX XVI | 110014186 | Unit for installation | #NAME? | Installed unit | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606219913122714853/5606219913122714853_signature_28.png | Vivienne Luzentales | https://www.jotform.com/uploads/macarewecare/222927978372067/5606219913122714853/5606219913122714853_signature_31.png
| 2023-05-23 | 2023-05-23 09:03 | Jose Reyes Memorial Medical Center | Rizal Ave., Sta. Cruz Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24608 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 1069 ASSAY: Xpert SARS-CoV-2, COUNT: 4573 | 2023-05-22 | GX16 | 110013856 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. |  | Xpertcheck 5 box | G5565 and G5593 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606143943416915513/5606143943416915513_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5606143943416915513/5606143943416915513_signature_31.png
| 2023-05-23 | 2023-05-17 14:24 | Nasugbu Rural Health Unit | J.P.Laurel St., Nasugbu | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24605 | PBSP | lucenavicente15@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 931 | 2023-05-16 | GXIV-D | 807581 | Module 1 with Lost Communication Error | Gateway Replacement | Performed Maintenance Replaced Gateway Board Archived Tests Generate Reports | Unit is Operational |  | Gateway Board | G 5567 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122670412993606/5606122670412993606_signature_28.png | Lucena i. Vicente | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122670412993606/5606122670412993606_signature_31.png
| 2023-05-23 | 2023-05-16 10:22 | GMA Rural Health Unit II | Brgy. Memije, General Mariano Alvarez | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24604 | PBSP | ldlconti@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1523 |  | GXIV-D | 814152 | Module 1 with Temperature Error | Module Replacement | Performed Maintenance Replaced Module 1 Generate Reports | 3 Modules are Operational, 1 Defective | 606336-> 210091143 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122670418833030/5606122670418833030_signature_28.png | Lorena L. Conti | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122670418833030/5606122670418833030_signature_31.png
| 2023-05-23 | 2023-05-18 11:10 | Tanza Rural Health Unit | Municipal Compound, Tanza | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24606 | PBSP | gutierrezrhian@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 619 | 2023-05-16 | GXIV-D | 812835 | CPU with Hard Drive Problem | CPu Replacement | Replaced CPU Pulled-out Back-up Laptop Generate Reports | All Modules are Operational |  | Lenovo Desktop | G 5576 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122680415794637/5606122680415794637_signature_28.png | Mary Ann C. Gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122680415794637/5606122680415794637_signature_31.png
| 2023-05-23 | 2023-05-18 12:06 | General Trias City Health Office | Prial Road, General Trias | 4A | Cavite | Laboratory | NO | MAINTENANCE Installation of Network Switch | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24607 | PBSP | loremontoya1169@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1694 ASSAY: Xpert HIV-1 Viral Load, COUNT: 194 |  | GXIV-D | 804691 | Unit is due for Maintenance | Maintenance | Performed Maintenance Generate Reports | 4 Modules are Operational. | Wrong Network Switch |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122680417714877/5606122680417714877_signature_28.png | Lorena R Montoya | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122680417714877/5606122680417714877_signature_31.png
| 2023-05-23 | 2023-05-09 10:00 | Carmona Rural Health Unit | Carmona | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24600 | DOH | irishzenreyes2197@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1116 |  | GXIV-D | 821216 | Module 2 with syringe problem | XpertCheck | Performed Maintenance Performed XpertCheck on reported Module Performed Self-Test Generate Reports | All Modules are Operational | Used Spare Kit, A2 for Observatiom |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122230416304111/5606122230416304111_signature_28.png | Irishzen reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122230416304111/5606122230416304111_signature_31.png
| 2023-05-23 | 2023-05-12 12:19 | San Lazaro Hospital | Quiricada St., Sta.Cruz | NCR | NCR 1st - Manila | HIV Training Center | NO | INSTALLATION | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24601 | MACARE-HBDC | sacclkiteval@gmail.com | 26 |  | 2023-05-10 | GXIV-D | 821257 | Unit is due for Installatiom | Installation | Installed GeneXpert Machine Generate IQR | Unit is Operational | With Yasser Salih | GeneXpert System | 9992 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122230417104614/5606122230417104614_signature_28.png | MELANIE PEREZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122230417104614/5606122230417104614_signature_31.png
| 2023-05-23 | 2023-05-10 10:18 | Rosario Rural Health Unit | Scout Torillo, Rosario | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24603 | DOH | arlenepenaarmijo@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1051 | 2023-04-25 | GXIV-D | 838537 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 1 Module is Operational, 2 Modules for Replacement, 1 Module Not Detected in XpertCheck | Module 3 Passed, Module 2 Not Detected | XpertCheck Kit | 9930 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122240416467337/5606122240416467337_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122240416467337/5606122240416467337_signature_31.png
| 2023-05-23 | 2023-05-10 11:51 | Tanza Rural Health Unit | Municipal Compound, Tanza | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24603 | PBSP | gutierrezrhian@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 525 |  | GXIV-D | 838162 | Module 1 with Temperature Error | Module Replacement | Replaced Module 1 Performed Maintenance Generate Reports | All Modules are Operational | 807003-> 677119 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122240416879116/5606122240416879116_signature_28.png | Robert torres | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122240416879116/5606122240416879116_signature_31.png
| 2023-05-23 | 2023-05-09 11:32 | GMA Rural Health Unit II | Brgy. Memije, General Mariano Alvarez | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24602 | PBSP | ldlconti@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1474 |  | GXIV-D | 814152 | Monitor for Replacement Module 4 with Temperature Errors | Module Replacement Monitor Replacement | Performed Maintenance Replaced Module 4 Replaced Monitor Generate Reports | 2 Modules are Defective, 2 Modules are Operational | 619639 -> 891932 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122240418092890/5606122240418092890_signature_28.png | Lorena L. Conti | https://www.jotform.com/uploads/macarewecare/222927978372067/5606122240418092890/5606122240418092890_signature_31.png
| 2023-05-23 | 2023-05-17 10:12 | Alaminos Rural Health Unit | D. Fandiño St. Alaminos | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24599 | PBSP | irahrobles@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 260 | 2023-05-15 | GXIV | 838447 | Unit for PM. For module replacement of module B2. | Performed PM. Performed module replacement on B2. | PM and module replacement. | Unit is operational. |  | GX Module B2 (210089549) in replacement to defective B2 (861446). | Monarch DR 0040814 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606071849488525074/5606071849488525074_signature_28.png | Iralyn M. Robles | https://www.jotform.com/uploads/macarewecare/222927978372067/5606071849488525074/5606071849488525074_signature_31.png
| 2023-05-23 | 2023-05-18 13:16 | Occidental Mindoro Provincial Hospital | Tayamaan, Mamburao | 4B | Occidental Mindoro | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24598 | PBSP | jaysonlopez12@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 328 ASSAY: Xpert SARS-CoV-2, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 | 2023-05-16 | GXIV | 807987 | Unit for PM. For module replacement of modules B1 and B3. | Performed PM. Performed module replacement on B1 and B3. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module B1 (626470) and B3 (647203) in replacement to defective B1 (685248) and B3 (662114). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606071559488774762/5606071559488774762_signature_28.png | Jayson Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5606071559488774762/5606071559488774762_signature_31.png
| 2023-05-23 | 2023-05-18 13:16 | Occidental Mindoro Provincial Hospital | Tayamaan, Mamburao | 4B | Occidental Mindoro | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24597 | PBSP | jaysonlopez12@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 328 ASSAY: Xpert SARS-CoV-2, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 | 2023-05-16 | GXIV | 807987 | Unit for PM. For module replacement of modules B1 and B3. | Performed PM. Performed module replacement on B1 and B3. Generated reports. | PM and module replacement. | Unit is operational. |  | GX module B1 (626470) and B3 (647203) in replacement to defective B1 (685248) and B3 (662114). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5606071539483943612/5606071539483943612_signature_28.png | Jayson Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5606071539483943612/5606071539483943612_signature_31.png
| 2023-05-22 | 2023-05-19 08:51 | Philippine Heart Center | East avenue.,Quezon city | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24596 | MACARE-HBDC | phcmolecularlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1219 ASSAY: Others, COUNT: 16 ASSAY: Xpert SARS-CoV-2, COUNT: 10728 |  | GXXVI | 838655 | For maintenance c1 error 4011,C3 error 4009,C4 error 1002,D2 error 4017. disabled | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | 4 modules disable, 12 modules available |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605672717407161555/5605672717407161555_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5605672717407161555/5605672717407161555_signature_31.png
| 2023-05-22 | 2023-05-19 09:48 | Philippine heart center | East avenue, Quezon city | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24595 | PBSP | phcmolecularlab@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 241 ASSAY: Others, COUNT: 5 ASSAY: Xpert SARS-CoV-2, COUNT: 1755 |  | GX4 | 834443 | B1 unable to detect | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed troubleshooting | Maintenance and troubleshooting | B1 unable to use 3 modules available |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605670677405654780/5605670677405654780_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5605670677405654780/5605670677405654780_signature_31.png
| 2023-05-22 | 2023-05-19 09:52 | Philippine heart center | East avenue, Quezon City | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24594 | MACARE-HBDC | phcmolecularlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 241 ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Others, COUNT: 5 ASSAY: Xpert SARS-CoV-2, COUNT: 1755 |  | GX4 | 804260 | C4 disabled encounter temp error | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | C4 d8sabled, 3 modules available |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605670127404893036/5605670127404893036_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5605670127404893036/5605670127404893036_signature_31.png
| 2023-05-22 | 2023-05-19 09:52 | Philippine heart center | East avenue, Quezon City | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24593 | MACARE-HBDC | phcmolecularlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 241 ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Others, COUNT: 5 ASSAY: Xpert SARS-CoV-2, COUNT: 1755 |  | GX4 | 804260 | C4 disabled encounter temp error | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | C4 d8sabled, 3 modules available |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605669607402718965/5605669607402718965_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5605669607402718965/5605669607402718965_signature_31.png
| 2023-05-22 | 2023-05-19 12:21 | Tatalon health center | Victory ave.Quezon CIty | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 24592 | DOH | jel472020@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 127 ASSAY: Xpert MTB/RIF Ultra, COUNT: 608 |  | GX4 | 831610 | A1 encounter ng continous 5007 error | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and troubleshooting | A4 only available,2 module disabled, 1 module not in use due to continous error 5007 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605668887407956614/5605668887407956614_signature_28.png | Jasmin Magtalas | https://www.jotform.com/uploads/macarewecare/222927978372067/5605668887407956614/5605668887407956614_signature_31.png
| 2023-05-22 | 2023-05-19 14:19 | San Mateo RHU 5 | Sunflower st.,Modesta VIl.age , San Mateo | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24591 | PBSP | mmacahis@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 436 | 2023-05-16 | GX4 | 838161 | For module replacement B4 807010 - 210091113 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | Operational |  | 900-0508 | G5584 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605667597402441472/5605667597402441472_signature_28.png | Monina Rica Macahis- De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5605667597402441472/5605667597402441472_signature_31.png
| 2023-05-22 | 2023-05-18 09:13 | Ramon Rural Health Unit | Bugallion Proper, Ramon, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24590 | DOH | rhuramonlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1096 | 2023-02-14 | Genexpert IV | 838480 | Module 3 error | Performed Xpertcheck = All Modules Passed | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Air-Conditioner Broken, Unit is not Operational | Do not use until Air-Conditioner is fixed | Xpertcheck 4pc | 9619 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605624316914219707/5605624316914219707_signature_28.png | Janette Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5605624316914219707/5605624316914219707_signature_31.png
| 2023-05-22 | 2023-05-18 09:13 | Ramon Rural Health Unit | Bugallion Proper, Ramon, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24589 | DOH | rhuramonlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1096 | 2023-02-14 | Genexpert IV | 838480 | Module 3 error | Performed Xpertcheck = All Modules Passed | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Air-Conditioner Broken, Unit is not Operational | Do not use until Air-Conditioner is fixed | Xpertcheck 4pc | 9619 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605624196913799802/5605624196913799802_signature_28.png | Janette Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5605624196913799802/5605624196913799802_signature_31.png
| 2023-05-22 | 2023-05-18 09:13 | Ramon Rural Health Unit | Bugallion Proper, Ramon, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24588 | DOH | rhuramonlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1096 | 2023-02-14 | Genexpert IV | 838480 | Module 3 error | Performed Xpertcheck = All Modules Passed | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Air-Conditioner Broken, Unit is not Operational | Do not use until Air-Conditioner is fixed | Xpertcheck 4pc | 9619 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605623936912716825/5605623936912716825_signature_28.png | Janette Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5605623936912716825/5605623936912716825_signature_31.png
| 2023-05-22 | 2023-05-18 11:10 | Alicia Rural Health Unit | Magsaysay, Alicia, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24587 | DOH | labalicia.rhu@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 60 ASSAY: Xpert MTB/RIF Ultra, COUNT: 590 | 2023-05-18 | Genexpert IV | 838521 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | 9620 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605622826915009977/5605622826915009977_signature_28.png | Jeslie Raymunda Valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/5605622826915009977/5605622826915009977_signature_31.png
| 2023-05-22 | 2023-05-18 11:10 | Alicia Rural Health Unit | Magsaysay, Alicia, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24586 | DOH | labalicia.rhu@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 60 ASSAY: Xpert MTB/RIF Ultra, COUNT: 590 | 2023-05-18 | Genexpert IV | 838521 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | 9620 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605622406919222647/5605622406919222647_signature_28.png | Jeslie Raymunda Valdez | https://www.jotform.com/uploads/macarewecare/222927978372067/5605622406919222647/5605622406919222647_signature_31.png
| 2023-05-22 | 2023-05-22 14:54 | Gapan Rural Health Unit | San Nicolas, Gapan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24585 | PBSP | daisymaygarcia14@gmail.com | 29 |  | 2023-05-22 | Genexpert IV | 814182 | Printer not working | Troubleshoot printer Installed Driver | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Printer defective | Genexpert Module - 210091148 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605481916923821921/5605481916923821921_signature_28.png | Daisy May Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5605481916923821921/5605481916923821921_signature_31.png
| 2023-05-22 | 2023-05-22 11:18 | Northern Mindanao Medical Center | Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24584 | PBSP | nmmcgenexpress@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 690 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7071 ASSAY: Xpert SARS-CoV-2, COUNT: 11576 ASSAY: Xpert HIV-1 Viral Load, COUNT: 525 |  | GeneXpert | 844118 | N/A | N/A | Preventive Maintenance Xpertcheck Calibration Generate report files Database back up | operational | Module A1/A3/C3 requiring service and has a high temperature log, all other modules pass the calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605453013223879134/5605453013223879134_signature_28.png | Rione Frances Lomongo | https://www.jotform.com/uploads/macarewecare/222927978372067/5605453013223879134/5605453013223879134_signature_31.png
| 2023-05-22 | 2023-05-12 09:13 | Metro Antipolo Hospital and Medical Center Inc. | 1348 Marcos Highway, Brgy. Mayamot, Antipolo City | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24583 | COMMERCIAL | mahmci_labcmt@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 44 | 2023-05-09 | GXIV | 110005265 | Unit for maintenance | #NAME? | Performed maintenance | Unit is operational. |  | XpertCheck Kit | 9984 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605446560415907982/5605446560415907982_signature_28.png | Grazie Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/5605446560415907982/5605446560415907982_signature_31.png
| 2023-05-22 | 2023-05-12 13:57 | San Miguel District Hospital | San Miguel, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24582 | PBSP | smdh_laboratory@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 91 ASSAY: Xpert MTB/RIF Ultra, COUNT: 823 | 2023-05-12 | GXIV | 807604 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. A1: passed A2: service required A3: retry A4: passed |  | Xpertcheck kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5605446560417777597/5605446560417777597_signature_28.png | Florencia Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5605446560417777597/5605446560417777597_signature_31.png
| 2023-05-19 | 2023-05-19 14:40 | Sariaya Rural Health Unit | Sariaya, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24581 | PBSP | marielinnetabril@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1890 | 2023-05-16 | GX IV 6C | 831890 | Module A4 (757623) for replacement | #NAME? | Replaced A4 | Unit is operational. A4: 757623 -> 210091100 |  | GeneXpert Module 6C SN:210091100 | G5582 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5603121653311363408/5603121653311363408_signature_28.png | Marie linnet Abril | https://www.jotform.com/uploads/macarewecare/222927978372067/5603121653311363408/5603121653311363408_signature_31.png
| 2023-05-19 | 2023-05-17 09:30 | Dr. Cornelio T. Martinez Sr. Memorial Hospital | Kiamba, Sarangani | 12 | Sarangani | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24580 | PBSP | mamartinez.pbsp@gmail.com | 28.3 | ASSAY: Xpert MTB/RIF, COUNT: 112 ASSAY: Xpert MTB/RIF Ultra, COUNT: 365 | 2023-05-17 | Gx-IV | 812017 | N/A | Perform Genexpert machine Xpertcheck/ Calibration Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Xpertcheck/ Calibration Perform Genexpert machine Preventive maintenance | 4 operational modules | All modules passed the Xpertcheck/Calibration | Xpertcheck CE-5 | Lot no: 1000330696 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602877720328305031/5602877720328305031_signature_28.png | Madelene R. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5602877720328305031/5602877720328305031_signature_31.png
| 2023-05-19 | 2023-05-19 14:27 | Tupi Rural Health Unit | Tupi | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24579 | PBSP | pedzella@gmail.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 701 | 2023-05-19 | Gx-IV (2 modules) | 835184 | Module 2 H/W failed | Replace module 2 SN:788668 with new module SN:210091139 Perform Genexpert machine Preventive maintenance | Replace module 2 SN:788668 with new module SN:210091139 Perform Genexpert machine Preventive maintenance | 2 operational modules |  | Genexpert module 6 color | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602876800322821916/5602876800322821916_signature_28.png | Ella Esmailene S. Pedregosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5602876800322821916/5602876800322821916_signature_31.png
| 2023-05-19 | 2023-05-19 13:50 | Western Visayas Sanitarium | Brgy. Inangayan, Sta. Barbara | 6 | Iloilo | TB-DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24578 | DOH | migzko@yahoo.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 79 ASSAY: Xpert MTB/RIF Ultra, COUNT: 215 |  | GX4 | 838393 | Error 2124 & 2008 | Conduct Preventive Maintenance and saves files(IQR, System Log, Assay Stats & GXC File) | Run Xpert check calibration for 4 modules | All modules passed the calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602871544115929349/5602871544115929349_signature_28.png | Ruth Dequilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5602871544115929349/5602871544115929349_signature_31.png
| 2023-05-19 | 2023-05-19 14:21 | Davao Doctors Hospital | Quirino Avenue, Davao City | 11 | Davao del Sur | Molecular | YES | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24577 | COMMERCIAL | arpson@ddh.com.ph | 24.3 |  | 2023-05-12 | GX4D | 839151 | Module A3 and A4 replacement | Replace Module A3 SN: 866248 with new Module SN: 210092430 Replaced the Defective Module A4 SN: 210009883 With new Module SN : 210092567 | Module Replacement | Unit is operational |  | 900-0508 | Sales Invoice #95540 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602870687919742159/5602870687919742159_signature_28.png | Nicole Andre Benedict A. Romarate | https://www.jotform.com/uploads/macarewecare/222927978372067/5602870687919742159/5602870687919742159_signature_31.png
| 2023-05-19 | 2023-05-19 12:15 | Nueva Vizcaya Provincial Hospital | Almaguer, Bambang, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24576 | DOH | dumlaojoy4@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1152 ASSAY: Xpert MTB/RIF, COUNT: 85 | 2023-02-14 | Genexpert IV | 838272 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | 9622 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602795357122073659/5602795357122073659_signature_28.png | Bernadette Dumlao | https://www.jotform.com/uploads/macarewecare/222927978372067/5602795357122073659/5602795357122073659_signature_31.png
| 2023-05-19 | 2023-05-19 10:58 | Talakag Rural Health Unit | Talakag | 10 | Bukidnon | TB DOTS | NO | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24575 | DOH | mhotalakaglaboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 207 |  | GeneXpert | 838566 | N/A | N/A | Preventive Maintenance Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602745819819168369/5602745819819168369_signature_28.png | Khiza tinoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5602745819819168369/5602745819819168369_signature_31.png
| 2023-05-19 | 2023-05-15 13:12 | Concepcion District Hospital | L. Cortez st. Concepcion, Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24574 | PBSP | cdh.laboratorydept@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1025 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 | 2023-04-27 | GXIV-D | 807562 | B2: 654652 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 40618 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602719879181508255/5602719879181508255_signature_28.png | Maria Rowena T. Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5602719879181508255/5602719879181508255_signature_31.png
| 2023-05-19 | 2023-05-15 13:12 | Concepcion District Hospital | L. Cortez st. Concepcion, Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24573 | PBSP | cdh.laboratorydept@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1025 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 | 2023-04-27 | GXIV-D | 807562 | B2: 654652 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 40618 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602719659184685865/5602719659184685865_signature_28.png | Maria Rowena T. Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5602719659184685865/5602719659184685865_signature_31.png
| 2023-05-19 | 2023-05-19 09:15 | Albuera Rural Health Unit | Albuera, Leyte | 8 | Leyte | Tb Dots | YES | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24572 | PBSP | mariaana.barte@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 167 | 2023-05-19 | GXIV-2D | 834818 | Due for XPC | #NAME? | #NAME? | All modules passed Xpertcheck and is operational. | 3 extra XPC collected from site. | XpertCheck-CE-5 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602684353226389360/5602684353226389360_signature_28.png | Maria Ana S. Barte | https://www.jotform.com/uploads/macarewecare/222927978372067/5602684353226389360/5602684353226389360_signature_31.png
| 2023-05-18 | 2023-05-17 14:30 | National kidney and transplant institute | East ave,Quezon city | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 24570 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 441 ASSAY: Others, COUNT: 9 ASSAY: Xpert HBV VL, COUNT: 234 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1243 ASSAY: Xpert MTB/RIF Ultra, COUNT: 285 ASSAY: Xpert SARS-CoV-2, COUNT: 18697 ASSAY: Xpert HCV VL, COUNT: 23 ASSAY: Xpert SARS-CoV-2, COUNT: 9 |  | GXXVI | 110005587 | C1 for troubleshooting loss communication | Performed troubleshooting, checked power output,checked gatewaycleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Troubleshooting and maintenance | 13 modules operational, 2 module reported, 1 module for reporting | Workwith Yasser Salih |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602173138791097504/5602173138791097504_signature_28.png | Ma. Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5602173138791097504/5602173138791097504_signature_31.png
| 2023-05-18 | 2023-05-18 09:49 | Philippine Children Medical Center | Quezon ave, Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 24571 | PBSP | kate_july27@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1777 | 2023-05-16 | GX4 | 807672 | For module replacement b2 645843 - 714190 b3 625055 - 714957 | Performed module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | Operational all module functioning |  | 900-0491rs | G5585 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602173138796569945/5602173138796569945_signature_28.png | Katherine Tabano | https://www.jotform.com/uploads/macarewecare/222927978372067/5602173138796569945/5602173138796569945_signature_31.png
| 2023-05-18 | 2023-05-18 14:20 | Calbayog City Health Office | Talahib, Calbayog City | 8 | Samar | TB-Dots | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24569 | PBSP | mardelen_l@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 465 ASSAY: Xpert MTB/RIF Ultra, COUNT: 821 | 2023-05-18 | GXIV | 804265 | #NAME? | #NAME? | #NAME? | Unit is Operational |  | XpertCheck Cartridges | PBSP |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5602143804121416046/5602143804121416046_signature_28.png | Vina Baile | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5602143804121416046/5602143804121416046_signature_31.png
| 2023-05-18 | 2023-05-18 14:20 | Calbayog City Health Office | Talking, Calbayog City | 8 | Samar | TB-Dots | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24568 | PBSP | mardelen_l@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 465 ASSAY: Xpert MTB/RIF Ultra, COUNT: 821 | 2023-05-18 | GXIV | 804265 | #NAME? | #NAME? | #NAME? | Unit is Operational |  | XpertCheck Cartridges | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602143804121416046/5602143804121416046_signature_28.png | Vina Baile | https://www.jotform.com/uploads/macarewecare/222927978372067/5602143804121416046/5602143804121416046_signature_31.png
| 2023-05-18 | 2023-05-18 16:15 | Philippine Red Cross - Davao Chapter | Roxas Boulevard, Davao City | 11 | Davao del Sur | Blood Center | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Roi Kenneth Fagutao |  |  |  |  | 24567 | COMMERCIAL | davaocitybloodservices@gmail.com | 24.7 |  |  | MCS+ | 02J021 | For Calibration and Preventive Maintenance | Collect Diagnostic Data | Perfrom Preventive Maintenance and Calibration, and Perfrom Cycle Tests | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602104485193023320/5602104485193023320_signature_28.png | Mars Panes | https://www.jotform.com/uploads/macarewecare/222927978372067/5602104485193023320/5602104485193023320_signature_31.png
| 2023-05-18 | 2023-05-18 16:15 | Philippine Red Cross - Davao Chapter | Roxas Boulevard, Davao City | 11 | Davao del Sur | Blood Center | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Roi Kenneth Fagutao |  |  |  |  | 24566 | COMMERCIAL | davaocitybloodservices@gmail.com | 24.7 |  |  | MCS+ | 02J021 | For Calibration and Preventive Maintenance | Collect Diagnostic Data | Perfrom Preventive Maintenance and Calibration, and Perfrom Cycle Tests | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602104445197812207/5602104445197812207_signature_28.png | Mars Panes | https://www.jotform.com/uploads/macarewecare/222927978372067/5602104445197812207/5602104445197812207_signature_31.png
| 2023-05-18 | 2023-05-18 11:46 | Lingayen Rural Health Unit 2 | Domalandan Lingayen Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24565 | PBSP | michelledoriarmt@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 147 ASSAY: Xpert MTB/RIF Ultra, COUNT: 875 |  | GXIV-D | 834629 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602078959188798578/5602078959188798578_signature_28.png | Melanie Sison | https://www.jotform.com/uploads/macarewecare/222927978372067/5602078959188798578/5602078959188798578_signature_31.png
| 2023-05-18 | 2023-05-18 14:24 | Patnanungan Rural Health Unit | Patnanungan, Quezon Province | 4A | Quezon | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24564 | DOH | navagemuelcarlp17@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 34 ASSAY: Xpert MTB/RIF Ultra, COUNT: 186 | 2023-04-27 | GXIV | 831611 | #NAME? | #NAME? | #NAME? | Unit is operational. Desktop under observation. |  | Barcode Scanner SN: 22236010556963 | 9958 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602012475419115250/5602012475419115250_signature_28.png | Gemuel Carl P. Nava | https://www.jotform.com/uploads/macarewecare/222927978372067/5602012475419115250/5602012475419115250_signature_31.png
| 2023-05-18 | 2023-05-18 14:22 | Pagsanjan Rural Helath Unit | Pagsanjan, Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24563 | PBSP | dotspagsanjan@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2023-05-16 | GX IV 6C | 807568 | CPU failed to boot up instantly | #NAME? | replaced cpu, mouse & keyboard new SN: CPH0622F38169 | Unit is operational |  | Lenovo configured CPU, Windows 10 | G5577 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5602007993707308303/5602007993707308303_signature_28.png | Lei Arnie Abrigo | https://www.jotform.com/uploads/macarewecare/222927978372067/5602007993707308303/5602007993707308303_signature_31.png
| 2023-05-18 | 2023-05-18 13:44 | Mindanao central sanitarium | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24562 | DOH | ybioamie@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1936 | 2023-05-18 | GX 4 | 821214 | Module 1 and 2 temperature error | Module 1 and 2 undergone Xpertcheck | Preventive Maintenence, module 1 and 2 undergone Xpertcheck. | Module 3 and 4 are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601997630214587958/5601997630214587958_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5601997630214587958/5601997630214587958_signature_31.png
| 2023-05-18 | 2023-05-18 11:19 | Tacurong City Health Office | Tacurong City | 12 | Sultan Kudarat | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24561 | PBSP | cho_clinlab@yahoo.com | 24.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1861 | 2023-05-18 | Gx-IV | 815393 | Module 2 H/W failed Module 3 H/W failed | Replace module 3 SN: 893610 with new module SN: 210091132 Perform Genexpert machine Preventive maintenance | Replace module 3 SN: 893610 with new module SN: 210091132 Perform Genexpert machine Preventive maintenance | 3 operational modules |  | Genexpert module 6 color | 900-0508RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601949403429148806/5601949403429148806_signature_28.png | Alma Torrecampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5601949403429148806/5601949403429148806_signature_31.png
| 2023-05-18 | 2023-05-18 11:28 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24560 | DOH | cauayanmho@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 270 ASSAY: Xpert MTB/RIF Ultra, COUNT: 970 |  | GX4 | 838477 | Module B1 error code 4012 & module B2 error code 1002 | Conduct preventive maintenance and save files (IQR, System Log, Assay Stats & GXC File) | Run Xpert check for module B1 & B2. | Module B1 w/ SN no. 862002 failed in Xpert check. Module B2 w/ SN no. 862003 passed in Xpert check |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601915443593018538/5601915443593018538_signature_28.png | Jose Ali C. Reyes, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5601915443593018538/5601915443593018538_signature_31.png
| 2023-05-18 | 2023-05-18 11:30 | Biliran Provincial Hospital | Naval, Biliran | 8 | Biliran | Tb Dots | YES | INSTALLATION ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24559 | PBSP | ghelmission13@yahoo.com | 24.5 | ASSAY: Xpert MTB/RIF, COUNT: 489 ASSAY: Xpert MTB/RIF Ultra, COUNT: 854 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 | 2023-05-18 | GXIV-4D | 815273 | Defective Monitor | -Install new LCD ThinkVision S22e-20 Flat Panel Monitor with SN: V5GTL350 | #NAME? | All modules are operational. |  | Lenovo ThinkVision S22e-20 Flat Panel Monitor | 800-0600 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601905903716722248/5601905903716722248_signature_28.png | Gheland Corto | https://www.jotform.com/uploads/macarewecare/222927978372067/5601905903716722248/5601905903716722248_signature_31.png
| 2023-05-18 | 2023-05-18 11:16 | Tagoloan Rural Health Unit | Tagoloan | 10 | Misamis Oriental | TB DOTS | NO | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24558 | DOH | dianekate22@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1211 |  | GeneXpert | 838574 | N/A | N/A | Preventive Maintenance Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601892508323211571/5601892508323211571_signature_28.png | Kate Diane Liquit | https://www.jotform.com/uploads/macarewecare/222927978372067/5601892508323211571/5601892508323211571_signature_31.png
| 2023-05-18 | 2023-05-18 08:58 | Siquijor IPHO | Caipilan Siquijor | 7 | Siquijor | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24557 | PBSP | glinniericamor73@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1198 |  | Gx4 | 808638 | Stucked up cartridge in Module 667651 | Removed stucked cartridge | Pull out stucked cartridge Module 1 SN 667651 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601858182616983353/5601858182616983353_signature_28.png | Glinnie Ricamor | https://www.jotform.com/uploads/macarewecare/222927978372067/5601858182616983353/5601858182616983353_signature_31.png
| 2023-05-17 | 2023-05-17 16:19 | Southern Isabela Medical Center | Rosario, Santiago City, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24556 | PBSP | simcmicrobiology@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 203 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3404 ASSAY: Xpert HIV-1 Viral Load, COUNT: 101 | 2023-05-17 | Genexpert IV | 805104 | Module A4 Booting Problem | Troubleshoot the module A4 Plug and unplug the module wire Release wire from tie | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601232519977880313/5601232519977880313_signature_28.png | Cassandra Cates Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5601232519977880313/5601232519977880313_signature_31.png
| 2023-05-17 | 2023-05-17 15:48 | AFPC - AFRIMS | V. Luna Road, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24555 | MACARE-HBDC | mtvalderama@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert SARS-CoV-2, COUNT: 5082 |  | GX IV 6C | 805294 | Module B3 error | -checked and cleaned cartridge bays, plunger rod, icore slot -replaced filter -checked B3 system log -error 4012 in B3 for 3 days -generate reports | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601199157222668983/5601199157222668983_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/5601199157222668983/5601199157222668983_signature_31.png
| 2023-05-17 | 2023-05-17 14:51 | AFPMC-AFRIMS | V. Luna Road, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24554 | CULMEN | mtvalderama@gmail.com | 27 | ASSAY: Xpert HBV VL, COUNT: 48 ASSAY: Xpert MTB/RIF, COUNT: 175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 377 ASSAY: Xpert SARS-CoV-2, COUNT: 2319 ASSAY: Xpert HCV VL, COUNT: 36 ASSAY: Xpert HIV-1 Viral Load, COUNT: 51 ASSAY: Others, COUNT: 18 | 2023-05-16 | GXIV - 10C | 845462 | Module B2 (898809) for replacement | #NAME? | Replaced B2 from SN 898809 to 911309 | unit is operational. |  | GeneXpert module SN: 911309 | G 5569 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601189007223687004/5601189007223687004_signature_28.png | Maria Theresa Valderama | https://www.jotform.com/uploads/macarewecare/222927978372067/5601189007223687004/5601189007223687004_signature_31.png
| 2023-05-17 | 2023-05-17 15:23 | San Jose Del Monte City Health Center VI | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24553 | PBSP | theadelrosario1@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1760 ASSAY: Xpert HIV-1 Viral Load, COUNT: 62 | 2023-05-17 | GXIV | 814161 | Module B2 (666021) for replacement | #NAME? | Replaced B2 | Unit is operational. |  | GeneXpert Module SN: 210091288 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601176836614786088/5601176836614786088_signature_28.png | Thea Del Rosario | https://www.jotform.com/uploads/macarewecare/222927978372067/5601176836614786088/5601176836614786088_signature_31.png
| 2023-05-17 | 2023-05-17 15:16 | Negros Oriental Provincial Hospital | Dumaguete City | 6 | Negros Oriental | Tb Laboratory | YES | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 24552 | PBSP | rezelmarielagos@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 615 |  | Gx4 | 805275 | Stucked up cartridge and Error 4008 | Remove cartridge | Remove cartridge, Get IQR and Sys log | Module 4 is Hardware fail |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601172876126306532/5601172876126306532_signature_28.png | Rezel Marie S. Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5601172876126306532/5601172876126306532_signature_31.png
| 2023-05-17 | 2023-05-17 15:00 | CEBU DOCTOR’S UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | MOLECULAR LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24551 | COMMERCIAL | katrinamariebalcera@yahoo.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 10 ASSAY: Others, COUNT: CT - 2 ; ASSAY: Xpert CT/NG, COUNT: 8 ASSAY: Xpert HBV VL, COUNT: 19 ASSAY: Xpert MTB/RIF Ultra, COUNT: 305 ASSAY: Others, COUNT: NG - 1 ASSAY: Xpert HCV VL, COUNT: 12 |  | GX-IV | 110008524 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601164909531187040/5601164909531187040_signature_28.png | Katrina Marie R. Balcera | https://www.jotform.com/uploads/macarewecare/222927978372067/5601164909531187040/5601164909531187040_signature_31.png
| 2023-05-17 | 2023-05-17 14:26 | Ormoc City Health Office | Ormoc City, Leyte | 8 | Leyte | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24550 | PBSP | theresesaladaga@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1232 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1582 ASSAY: Xpert HIV-1 Viral Load, COUNT: 59 | 2023-05-17 | GXIV-4D | 804253 | Module B2 defective | -Replace defective module B2/892169 with new module SN: 210089520 | #NAME? | All modules are operational. |  | 900-GX GeneXpert Module 6 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601146873497704912/5601146873497704912_signature_28.png | Ma. Therese A. Saladaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5601146873497704912/5601146873497704912_signature_31.png
| 2023-05-17 | 2023-05-17 12:35 | Pasig City General Hospital | F. Legazpi Extension, Maybunga, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Dan Emil Bodino |  |  |  |  | 24549 | PBSP | maniquizederlina@yahoo.com | 25 |  | 2023-05-17 | GX4 | 805010 | H/W Failed for B1, stuck cartridge | Plunger rod adjustment for stuck cartridge Performed self test & reliability test Generated report Checked cable connections inside the machine Generates reports | Performed self test | Unit is operational | B1 210091374 H/W failed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601135774731867659/5601135774731867659_signature_28.png | Ederlina Sagales | https://www.jotform.com/uploads/macarewecare/222927978372067/5601135774731867659/5601135774731867659_signature_31.png
| 2023-05-17 | 2023-05-17 13:23 | Davao TB Reference Laboratory | DOH Compound, JP Laurel Avenue, Davao City | 11 | Davao del Sur | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24548 | WHO | dtrlmark.pbsp@gmail.com | 27.8 |  |  | GX-IV 10C | 848248 | Module B3 encounter error | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform functionality test of the Module B3 | Perfrom Preventive Maintenance and troubleshooting | Module B3 encounter multiple Temperature error, data willbe sent to Cepheid for further Assessment | Module B3 is Disabled, modules only available for testing |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601106059435628039/5601106059435628039_signature_28.png | Joseph Mark Anque | https://www.jotform.com/uploads/macarewecare/222927978372067/5601106059435628039/5601106059435628039_signature_31.png
| 2023-05-17 | 2023-05-17 13:16 | Doña Remedios Trinidad Rural Health Unit | DRT, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24547 | PBSP | maricel.rmt24@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA | 2023-05-17 | GXIV | 838872 | #NAME? | #NAME? | #NAME? | Unit is operational. Laptop is for reporting. Laptop and defective monitor for pull-out. |  | Lenovo Monitor SN: V5GTL340 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601102177851360408/5601102177851360408_signature_28.png | Cecilia Angeles | https://www.jotform.com/uploads/macarewecare/222927978372067/5601102177851360408/5601102177851360408_signature_31.png
| 2023-05-17 | 2023-05-17 12:54 | Pangasinan Provincial Hospital | Bulingit San Carlos City Pangasinan | 1 | Pangasinan | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24546 | PBSP | judzfm@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 253 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2923 | 2023-05-15 | GXIV-D | 808283 | A3:666225 for module replacement | Module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 40828 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601087581349900435/5601087581349900435_signature_28.png | Judy Ann F. Mabunga | https://www.jotform.com/uploads/macarewecare/222927978372067/5601087581349900435/5601087581349900435_signature_31.png
| 2023-05-17 | 2023-05-17 08:30 | Victoria Rural health Unit | Brgy. Bulo Victoria Tarlac | 3 | Tarlac | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24545 | DOH | diazrmt@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 223 ASSAY: Xpert MTB/RIF Ultra, COUNT: 686 |  | GXIV-D | 831669 | A1 and A2 for Re-Xpertcheck temperature error | Re-Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Performed calibration Gathered reports Performed functionality test | Unit is operational | Under observation. Used 2 xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601083651348189545/5601083651348189545_signature_28.png | Alma Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5601083651348189545/5601083651348189545_signature_31.png
| 2023-05-17 | 2023-05-17 12:17 | Calamba City Health Office | Brgy.7, Calamba | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24544 | PBSP | 2052456@aup.edu.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1973 | 2023-05-16 | GXIV | 807570 | Unit for PM. For module replacement of module B1. | Performed PM. Performed module replacement on B1. Generated reports. | PM and module replacement. | Unit is operational. |  | GX Module B1 (210090905) in replacement to defective B1 (624926). | G 5573 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601067280525412929/5601067280525412929_signature_28.png | Mary Jane Lourdes R. Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5601067280525412929/5601067280525412929_signature_31.png
| 2023-05-17 | 2023-05-17 11:36 | Baybay City Health Office | Baybay City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24543 | DOH | sajecopooh20@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 512 ASSAY: Xpert MTB/RIF Ultra, COUNT: 512 | 2023-05-17 | GXIV | 838555 | #NAME? | #NAME? | #NAME? | Unit is Operational |  | Xpertcheck Cartridge | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601044557912896027/5601044557912896027_signature_28.png | sarah jean r gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5601044557912896027/5601044557912896027_signature_31.png
| 2023-05-17 | 2023-05-17 11:12 | Batis Health Center | 428 F. Manalo St., San Juan City, NCR | NCR | NCR 2nd - San Juan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24542 | DOH | almacenaio16@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 286 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1645 ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 | 2023-05-16 | GX 4 | 831598 | For Module Replacement Stuck Cartridge | -Removed stuck Cartridge in module A1 -Performed module replacement for A1 630159 >>> 210091071 - Cleaned Cartridge Bays, Plunger rods and I-core - Replaced dirty filter with new filter - Generated reports | #NAME? | Unit is operational | Monitor shuts down randomly | GeneXpert Module 6-COLOR SN: 210091071 | G 5583 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601038128812494707/5601038128812494707_signature_28.png | Rosario T. Almacen | https://www.jotform.com/uploads/macarewecare/222927978372067/5601038128812494707/5601038128812494707_signature_31.png
| 2023-05-17 | 2023-05-17 10:38 | Banlat Health Center | Tandang Sora, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24541 | PBSP | ibarra.emma1227@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 95 ASSAY: Xpert MTB/RIF, COUNT: 142 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1363 | 2023-05-16 | GX IV | 804688 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational A1 - passed A2 - passed march 2023 A3 - newly replaced A4 - passed |  | XpertCheck Kit | G 5579 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5601010513817956174/5601010513817956174_signature_28.png | Emma Ibarra | https://www.jotform.com/uploads/macarewecare/222927978372067/5601010513817956174/5601010513817956174_signature_31.png
| 2023-05-17 | 2023-05-17 10:12 | Alaminos Rural Health Unit | D. Fandiño St. Alaminos | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24540 | PBSP | irahrobles@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 260 | 2023-05-15 | GXIV | 838447 | Unit for PM. For module replacement of module B2. | Performed PM. Performed module replacement on B2. | PM and module replacement. | Unit is operational. |  | GX Module B2 (210089549) in replacement to defective B2 (861446). | Monarch DR 0040814 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600992900523897892/5600992900523897892_signature_28.png | Iralyn M. Robles | https://www.jotform.com/uploads/macarewecare/222927978372067/5600992900523897892/5600992900523897892_signature_31.png
| 2023-05-17 | 2023-05-16 15:23 | Palimbang Rural Health Unit | Palimbang | 12 | Sultan Kudarat | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24539 | WHO | tina_villegas2319@yahoo.com | 25.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 108 ASSAY: Xpert MTB/RIF, COUNT: 187 | 2023-05-16 | Gx-IV | 847257 | Module 2 H/W Failed | Replace module 2 SN:910142 with new module SN:682714 Perform Genexpert machine Preventive maintenance | Replace module 2 SN:910142 with new module SN:682714 Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 color | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600893709313598989/5600893709313598989_signature_28.png | Cristina E. Villegas | https://www.jotform.com/uploads/macarewecare/222927978372067/5600893709313598989/5600893709313598989_signature_31.png
| 2023-05-16 | 2023-05-16 12:46 | Eastern Pangasinan District Hospital | Magallanes Tayug Pangasinan | 1 | Pangasinan | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24538 | PBSP | epdhlaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 481 |  | GXIV-D | 815349 | Due for calibration B3 temperature error | Preventive maintenance Xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except B3 | All modules passed, B3 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600370169185951004/5600370169185951004_signature_28.png | Fergyl Escaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5600370169185951004/5600370169185951004_signature_31.png
| 2023-05-16 | 2023-05-11 10:44 | Kalinga Provincial Hospital | Bulanao, Tabuk City, Kalinga | CAR | Kalinga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24535 | PBSP | leombelgica139@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1274 | 2023-05-11 | Genexpert IV | 807622 | Module 4 door problem Module 1 error 1002 | Recorded Video of problem | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 693422 - 677543 | Genexpert Module = 677453 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600343938514817728/5600343938514817728_signature_28.png | Leo Mark Belgica | https://www.jotform.com/uploads/macarewecare/222927978372067/5600343938514817728/5600343938514817728_signature_31.png
| 2023-05-16 | 2023-05-11 10:58 | Kalinga Provincial Hospital | Bulanao, Tabuk City, Kalinga | CAR | Kalinga | Laborarory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24536 | DOH | leombelgica139@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1344 | 2023-02-14 | Genexpert IV | 838466 | Module 2, 3, 4 errors | Performed Xpertcheck | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = B1, B2, B4 - Passed = B3 - Service Required | Xpertcheck 5pc | 9627 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600343948516678080/5600343948516678080_signature_28.png | Leo Mark Belgica | https://www.jotform.com/uploads/macarewecare/222927978372067/5600343948516678080/5600343948516678080_signature_31.png
| 2023-05-16 | 2023-05-16 15:50 | Munoz City Health Office | Poblacion West, Science City of Munoz, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 24537 | DOH | scomrtdl@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 61 ASSAY: Xpert MTB/RIF Ultra, COUNT: 612 | 2023-05-16 | Genexpert IV | 831673 | Module 1 error 5002 | Performed Xpertcheck on module A1 | Performed Maintenance Gathered Reports Performed Xpertcheck on Module A1 Enabled Module A1 Performed Functionality Test | Unit is Operational | Xpertcheck result = Passed | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600343948519516793/5600343948519516793_signature_28.png | Ronalyn Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5600343948519516793/5600343948519516793_signature_31.png
| 2023-05-16 | 2023-05-16 15:50 | Pasig City Health Office | F. Legazpi Extension, Maybunga, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24534 | PBSP | pcghlaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 151 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2979 | 2023-05-16 | GX4 | 805010 | For Module Replacement | Module Replacement(644019>>210091374) Cleaned module bay, plunger rod, filter Generated reports | Performed Preventive Maintenance | Unit is operational |  | GeneXpert module 210091374 | G 5566 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600339204028153037/5600339204028153037_signature_28.png | Chermarie Cabang | https://www.jotform.com/uploads/macarewecare/222927978372067/5600339204028153037/5600339204028153037_signature_31.png
| 2023-05-16 | 2023-05-16 14:49 | Labuan General hospital | Labuan, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24533 | DOH | jenne060213@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2055 | 2023-05-16 | GX 4 | 838568 | Module 1 h/w failed | Preventive Maintenence and Module 1 undergone Xpertcheck calibration. Result is module 2 service required | Module 1 undergone Xpertcheck calibration | All 3 modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600296458116988974/5600296458116988974_signature_28.png | Jennerose Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5600296458116988974/5600296458116988974_signature_31.png
| 2023-05-16 | 2023-05-16 13:48 | Padada Rural Health Unit | Poblacion, Padada, Davao Del Sur | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24532 | DOH | torresrmt2016@gmail.com | 25.8 |  |  | GX-IV | 838511 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600275786913537546/5600275786913537546_signature_28.png | TORRES, MARL DAVE DWIGHT T. | https://www.jotform.com/uploads/macarewecare/222927978372067/5600275786913537546/5600275786913537546_signature_31.png
| 2023-05-16 | 2023-05-16 14:15 | A.H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24531 | PBSP | dinajaen1990@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 120 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1947 | 2023-05-16 | GXIV | 805067 | Module A2 (634968) for replacement | #NAME? | Replaced A2 | Unit is operational. |  | GeneXpert Module SN: 210091321 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600273152135549901/5600273152135549901_signature_28.png | Dina Jaen | https://www.jotform.com/uploads/macarewecare/222927978372067/5600273152135549901/5600273152135549901_signature_31.png
| 2023-05-16 | 2023-05-16 11:22 | Marilao Rural Health Unit I | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24530 | DOH | lab.rhu1marilao@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 108 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1898 ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 | 2023-05-16 | GXIV | 831897 | Temperature error on A3 (644075) | #NAME? | Performed XpertCheck on A3 | Unit is operational. A3: passed A3 is under observation. |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600196353418321188/5600196353418321188_signature_28.png | Jonathan Jenilo | https://www.jotform.com/uploads/macarewecare/222927978372067/5600196353418321188/5600196353418321188_signature_31.png
| 2023-05-16 | 2023-05-16 11:32 | Chinese General Hospital | Blumentritt, Sta Cruz | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24529 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 13 ASSAY: Xpert CT/NG, COUNT: 35 ASSAY: Xpert HBV VL, COUNT: 57 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1318 ASSAY: Xpert HCV VL, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 |  | GXIV | 835604 | Unit for Preventive Maintenance. For checking of modules B4. | Performed Preventive Maintenance. Checked module B4. Generated reports. | Preventive maintenance and troubleshooting. | Unit is operational. Module B4 load error probably due to uprised valve. 3 modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5600178560529960502/5600178560529960502_signature_28.png | Maria Odessa Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/5600178560529960502/5600178560529960502_signature_31.png
| 2023-05-15 | 2023-05-15 17:38 | Cebu City Health Office | Cebu City, Maboll | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24528 | PBSP | laureanjo@yahoo.com | 26 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 491 |  | GX4 | 808620 | Defective Module B2 SN 631758 | Replaced new module | Replaced New Module SN 210092152 | All Modules are opertional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599533866917310750/5599533866917310750_signature_28.png | Laureen Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5599533866917310750/5599533866917310750_signature_31.png
| 2023-05-15 | 2023-05-15 14:35 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 24527 | PBSP | laureanjo@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 933 ASSAY: Xpert MTB/RIF Ultra, COUNT: 479 ASSAY: Xpert SARS-CoV-2, COUNT: 16 |  | Gx 4 | 804840 | Defective modules 1 sn 730761 & 2 sn698742 | New modules installed module 1 sn210092143 & 2 sn210092147 | New modules installed module 1 sn210092143 & 2 sn210092147 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599423220653491366/5599423220653491366_signature_28.png | Laurean Jo. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5599423220653491366/5599423220653491366_signature_31.png
| 2023-05-15 | 2023-05-15 10:13 | Gilbert O. Teodoro Memorial Hospital | Malacam0a Camiling, Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24526 | PBSP | almabellapdancel3@gmail.com | 31 | ASSAY: Xpert MTB/RIF, COUNT: 110 ASSAY: Xpert MTB/RIF Ultra, COUNT: 571 | 2023-04-27 | GXIV-D | 812838 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports HIV VL assay installation Performed functionality test | Unit is operational | All modules passed calibration | 4 xoertcheck kit | 40649 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599264135512007292/5599264135512007292_signature_28.png | Alma Bella P. Dancel | https://www.jotform.com/uploads/macarewecare/222927978372067/5599264135512007292/5599264135512007292_signature_31.png
| 2023-05-15 | 2023-05-10 17:03 | Murcia Rural Health Unit | Rizal st, Brgy zone 3 murcia | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 24525 | PBSP | kambal.kapid@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1118 |  | Gx 4 | 835055 | Defective module 2 sn 787916 | New module installed, clean fan filter and fan | New module installed 2 sn 713807 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599251636915313507/5599251636915313507_signature_28.png | Allan M.Benjamin | https://www.jotform.com/uploads/macarewecare/222927978372067/5599251636915313507/5599251636915313507_signature_31.png
| 2023-05-15 | 2023-05-12 14:22 | National children's hospital | E.rodriguez sr. Ave, Quezon city | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION | CEPHEID | Noel Floresca |  |  |  |  | 24524 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 7 ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2734 ASSAY: Xpert SARS-CoV-2, COUNT: 7545 ASSAY: Xpert SARS-CoV-2, COUNT: 18 | 2023-05-10 | GXXVI (10c) | 110014171 | For xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed xpertcheck | Xpertcheck | 14 modules operational | A1 , B4 failed xpertcheck | Xpertcheck | 9988 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599218160411608951/5599218160411608951_signature_28.png | National Children's Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/5599218160411608951/5599218160411608951_signature_31.png
| 2023-05-15 | 2023-05-11 12:29 | Philippine Red Cross (Port area) | Bonifacio drive,Port area,Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 24523 | MACARE-HBDC | dolly.banares@redcross.org.ph | 26 |  |  | MCS + | 02AO87 | As per end user they encounter error system problem | Performed troubleshooting,reset cpu boardchecked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure,test run unit | Troubleshooting and Calibration | Operational | Work with Yasser Salih |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599217530416920564/5599217530416920564_signature_28.png | Dolly Bañares | https://www.jotform.com/uploads/macarewecare/222927978372067/5599217530416920564/5599217530416920564_signature_31.png
| 2023-05-15 | 2023-05-11 14:27 | Philippine Red Cross (Port area) | Bonifacio drive, Port area. Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 24522 | MACARE-HBDC | dolly.banares@redcross.org.ph | 26 |  |  | MCS + | 12D063 | For maintenance and calibration checking | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance and Calibration | Operational | Work with Yasser Salih |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599216140412771126/5599216140412771126_signature_28.png | Dolly Bañares | https://www.jotform.com/uploads/macarewecare/222927978372067/5599216140412771126/5599216140412771126_signature_31.png
| 2023-05-15 | 2023-05-12 | South Cotabato Provincial Hospital | Koronadal City | 12 | South Cotabato | Molecular laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  | jessontupas.macaremed@gmail.com | dappmc.tblab@gmail.com |  | 24521 | COMMERCIAL | jessontupas.macaremed@gmail.com | 27 | ASSAY: Xpert SARS-CoV-2, COUNT: 7376 |  | Gx-XVI | 839314 | Module B1 C4 and D1 multiple temperature error Module A3, C2 and D4 waiting for replacement module | Perform Xpertcheck at module B1, C4, and D1 for troubleshooting Perform Genexpert machine Preventive maintenance | Perform Xpertcheck at module B1, C4, and D1 Perform Genexpert machine Preventive maintenance | Module B1,C4 and D1 Passed the calibration 15 operational modules | Module A3, C2, and D4 waiting for replacement module |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599213666902436023/5599213666902436023_signature_28.png | Noel Fuscablo | https://www.jotform.com/uploads/macarewecare/222927978372067/5599213666902436023/5599213666902436023_signature_31.png
| 2023-05-15 | 2023-05-05 11:36 | Batangas Medical Center | Batangas City | 4A | Batangas | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24520 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 180 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2100 ASSAY: Xpert SARS-CoV-2, COUNT: 1997 ASSAY: Xpert HIV-1 Viral Load, COUNT: 602 | 2023-05-02 | GX16 | 844120 | Unit for PM. For module replacement of modules A3, B4 and D4. | Performed PM. Performed module replacement on A3, B4 and D4. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module A3 (680897), B4 (682115) and D4 (680439) in replacement to defective A3 (892245) B4 (892247) and D4 (892447). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599209510416189565/5599209510416189565_signature_28.png | Patricia Louisse Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5599209510416189565/5599209510416189565_signature_31.png
| 2023-05-15 | 2023-05-05 11:36 | Batangas Medical Center | Batangas City | 4A | Batangas | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24519 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 180 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2100 ASSAY: Xpert SARS-CoV-2, COUNT: 1997 ASSAY: Xpert HIV-1 Viral Load, COUNT: 602 | 2023-05-02 | GX16 | 844120 | Unit for PM. For module replacement of modules A3, B4 and D4. | Performed PM. Performed module replacement on A3, B4 and D4. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module A3 (680897), B4 (682115) and D4 (680439) in replacement to defective A3 (892245) B4 (892247) and D4 (892447). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599209470418061183/5599209470418061183_signature_28.png | Patricia Louisse Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5599209470418061183/5599209470418061183_signature_31.png
| 2023-05-15 | 2023-05-05 11:29 | Batangas Medical Center | Batangas City | 4A | Batangas | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24518 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2023-05-02 | GXIV | 834627 | Unit for PM. For module replacement of module B2. | Performed PM. Performed module replacement on B3. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B2 (677854) in replacement to defective B2 (784301) | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599209100414422997/5599209100414422997_signature_28.png | Patricia Louisse Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5599209100414422997/5599209100414422997_signature_31.png
| 2023-05-15 | 2023-05-05 11:29 | Batangas Medical Center | Batangas City | 4A | Batangas | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24517 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2023-05-02 | GXIV | 834627 | Unit for PM. For module replacement of module B2. | Performed PM. Performed module replacement on B3. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B2 (677854) in replacement to defective B2 (784301) | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599209070415264240/5599209070415264240_signature_28.png | Patricia Louisse Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5599209070415264240/5599209070415264240_signature_31.png
| 2023-05-15 | 2023-05-12 14:28 | Calauan Rural Health Unit | Albela St. Calauan | 4A | Laguna | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24516 | DOH | mesucalauan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 490 | 2023-05-08 | GXIV | 821164 | Unit for PM. Unit for Xpertcheck. For installation of new UPS. | Performed installation of UPS. Performed PM. Performed Xpertcheck. Generated reports. | PM, Xpertcheck and installation of UPS. | Unit is operational. Module A3 (783229), service required. |  | Xpertcheck. Accupower UPS (AT1000211227ST0130) in replacement to defective UPS (AT1000180910ST0078) | Xpertcheck DR 9612, UPS DR 9979 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599208580419358821/5599208580419358821_signature_28.png | Sharina jean yu | https://www.jotform.com/uploads/macarewecare/222927978372067/5599208580419358821/5599208580419358821_signature_31.png
| 2023-05-15 | 2023-05-12 14:28 | Calauan Rural Health Unit | Albela St. Calauan | 4A | Laguna | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24515 | DOH | mesucalauan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 490 | 2023-05-08 | GXIV | 821164 | Unit for PM. Unit for Xpertcheck. For installation of new UPS. | Performed installation of UPS. Performed PM. Performed Xpertcheck. Generated reports. | PM, Xpertcheck and installation of UPS. | Unit is operational. Module A3 (783229), service required. |  | Xpertcheck. Accupower UPS (AT1000211227ST0130) in replacement to defective UPS (AT1000180910ST0078) | Xpertcheck DR 9612, UPS DR 9979 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5599208560413589551/5599208560413589551_signature_28.png | Sharina jean yu | https://www.jotform.com/uploads/macarewecare/222927978372067/5599208560413589551/5599208560413589551_signature_31.png
| 2023-05-14 | 2023-05-11 14:12 | Mini Forest Health Center | Quezon Boulevard, Barangay 23-C Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24514 | DOH | miniforestlaboratory@gmail.com | 24.7 |  |  | GX-IV 4D | 838507 | Error at Module B4 | Run XpertCheck Calibration for Verificationol of Error | Perform Module Troubleshooting | Module B4 for service required | Data will be sent to Cepheid for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5598789024276678756/5598789024276678756_signature_28.png | Joy L. Nio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5598789024276678756/5598789024276678756_signature_31.png
| 2023-05-14 | 2023-05-11 14:12 | Mini Forest Health Center | Quezon Boulevard, Barangay 23-C Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24513 | DOH | miniforestlaboratory@gmail.com | 24.7 |  |  | GX-IV 4D | 838507 | Error at Module B4 | Run XpertCheck Calibration for Verificationol of Error | Perform Module Troubleshooting | Module B4 for service required | Data will be sent to Cepheid for further Assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5598788984271462452/5598788984271462452_signature_28.png | Joy L. Nio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5598788984271462452/5598788984271462452_signature_31.png
| 2023-05-12 | 2023-05-12 14:28 | Calauan Rural Health Unit | Albela St. Calauan | 4A | Laguna | Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24512 | DOH | mesucalauan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 490 | 2023-05-08 | GXIV | 821164 | Unit for PM. Unit for Xpertcheck. For installation of new UPS. | Performed installation of UPS. Performed PM. Performed Xpertcheck. Generated reports. | PM, Xpertcheck and installatiom of UPS. | Unit is operational. Module A3 (783229), service required. |  | Xpertcheck. Accupower UPS (AT1000211227ST0130) in replacement to defective UPS (AT1000180910ST0078) | Xpertcheck DR 9612, UPS DR 9979 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5596832609682738335/5596832609682738335_signature_28.png | Sharina jean yu | https://www.jotform.com/uploads/macarewecare/222927978372067/5596832609682738335/5596832609682738335_signature_31.png
| 2023-05-12 | 2023-05-12 13:13 | Nagcarlan Rural Health Unit | Brgy. Poblacion II, Nagcarlan | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24511 | DOH | zekcoronado@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 816 | 2023-02-13 | GXIV | 821163 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. | PM and Xpertcheck. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck | DR 9611 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5596787740113345105/5596787740113345105_signature_28.png | Jojo B Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5596787740113345105/5596787740113345105_signature_31.png
| 2023-05-12 | 2023-05-12 11:23 | Dr. Jose Rizal Memorial Hospital | Lawaan, Dapitan city | 9 | Zamboanga del Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24510 | PBSP | lovinoantonieta@gmail.com | 28 |  | 2023-05-12 | GX 4 | 804698 | Module 2 error 1002 | Undergone Xpertcheck, All modules passed/ Preventive Maintenence | Undergone Xpertcheck/ Preventive Maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5596734894367170038/5596734894367170038_signature_28.png | Antonieta B. Lovino | https://www.jotform.com/uploads/macarewecare/222927978372067/5596734894367170038/5596734894367170038_signature_31.png
| 2023-05-12 | 2023-05-11 14:44 | Sultan kudarat Provincial Hospital | Isulan | 12 | Sultan Kudarat | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24509 | PBSP | charmfrias1270@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2997 |  | Gx-IV | 807405 | Module 3 multiple temperature error | Troubleshoot module 3 due to multiple temperature error Use or run dummy cartridge for troubleshooting Reported problem to jotform | Troubleshoot module 3 due to multiple temperature error Use or run dummy cartridge for troubleshooting Reported problem to jotform | 4 operational modules | Troubleshooting result for Module 3 Error 5017 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5596722822582064253/5596722822582064253_signature_28.png | Charma Grace Frias | https://www.jotform.com/uploads/macarewecare/222927978372067/5596722822582064253/5596722822582064253_signature_31.png
| 2023-05-12 | 2023-05-12 10:31 | Amang Rodriguez Memorial Medical Center | Sumulong Hi-way, Sto. Niño, Marikina | NCR | NCR 2nd - Marikina | Bacteriology | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24508 | PBSP | rheddtimoteo27@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 271 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2697 | 2023-04-28 | GX IV - 6C | 807982 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. 1 new, 3 passed |  | XpertCheck kit | DR#40734 PL-011238 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5596687619018674429/5596687619018674429_signature_28.png | Christine Rose Timoteo | https://www.jotform.com/uploads/macarewecare/222927978372067/5596687619018674429/5596687619018674429_signature_31.png
| 2023-05-12 | 2023-05-12 09:18 | Zamboanga del norte Medical Centre | Sicayag dipolog city, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24507 | PBSP | thelmamorgia@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 471 | 2023-05-12 | GX 4 | 815285 | Module 4 error 4017 | Replaced module 731313 with Module 658574 | Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5596672309536474855/5596672309536474855_signature_28.png | Thelma Wabe | https://www.jotform.com/uploads/macarewecare/222927978372067/5596672309536474855/5596672309536474855_signature_31.png
| 2023-05-11 | 2023-05-11 15:19 | The Black Nazarene Hospital Inc. | Brgy 2. San Nicolas Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24506 | CONSORTIUM | tbnhi.lab@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 19 | 2023-05-08 | GXIV-D | 835542 | B1:789943 B4: 790381 for moduke replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 2 genexpert modules | 9981 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595994211944978710/5595994211944978710_signature_28.png | Joanna Marie Salud | https://www.jotform.com/uploads/macarewecare/222927978372067/5595994211944978710/5595994211944978710_signature_31.png
| 2023-05-11 | 2023-05-11 11:51 | Bantay Municipal Health office | Zone 5 Bantay Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24505 | PBSP | lornarabara_23@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 61Lorna |  | GXIV-D | 807585 | Date and time not updating automatically | Repair/troubleshooting Oreventive maintenance | Performed preventive maintenance Replaced BIOS battery Update date and time Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595987141941903736/5595987141941903736_signature_28.png | Lorna Rabara | https://www.jotform.com/uploads/macarewecare/222927978372067/5595987141941903736/5595987141941903736_signature_31.png
| 2023-05-11 | 2023-05-11 15:08 | Maysan 3S Health Center | F. Alarcon St., Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24504 | PBSP | cesguardiano@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 505 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 | 2023-05-11 | GXIV | 814151 | Module A3 (688222) for replacement | - replaced A3 - generated reports - modules A1, A2, A3, A4 generated power supply error 4008 - reset gateway board and other cables - swapped new module position A3 with A1 - performed self test on A1, A2, A3, A4 (problem persist) | Replaced module A3 | Unit is not operational. Needs further troubleshooting. For reporting. |  | GeneXpert Module SN: 677431 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595986800014076721/5595986800014076721_signature_28.png | Cecilia Guardiano | https://www.jotform.com/uploads/macarewecare/222927978372067/5595986800014076721/5595986800014076721_signature_31.png
| 2023-05-11 | 2023-05-11 13:49 | Guimbal Rural Health Unit | Guimbal, Iloilo | 6 | Iloilo | Laboratory | YES | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24503 | PBSP | toto.davis.td@gmail.com | 33 | ASSAY: Xpert MTB/RIF, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1617 |  | GXIV-4D | 807587 | None | None | #NAME? | All modules are functional. | Room AC is not functional, ambient temp is greater than 30C and module temp is ranging 34-37C |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595966452814439642/5595966452814439642_signature_28.png | Davis James Secapuri | https://www.jotform.com/uploads/macarewecare/222927978372067/5595966452814439642/5595966452814439642_signature_31.png
| 2023-05-11 | 2023-05-11 14:18 | PHILIPPINE RED CROSS-ILOILO CITY CHAPTER | BONIFACIO DRIVE,ILOILO CITY | 6 | Iloilo | BLOOD CENTER | NO | REPAIR | HAEMONETICS | Joseph Acebedo |  |  |  |  | 24502 | COMMERCIAL | rlcerbas@gmail.com | 20 |  |  | MCS+ | 98G049 | SELF TEST PROBLEM ERROR 479-WEIGHER TEST FAILED | #NAME? | -PERFORM INITIAL TROUBLESHOOTING -PERFORM WEIGHER REPLACEMENT USING SERVICE PART -DEFECTIVE WEIGHER SN:929984 | EQUIPMENT IS OPERATIONAL AFTER WEIGHER LOAD CELL REPLACEMENT |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595954576327116320/5595954576327116320_signature_28.png | Ryan L Cerbas | https://www.jotform.com/uploads/macarewecare/222927978372067/5595954576327116320/5595954576327116320_signature_31.png
| 2023-05-11 | 2023-05-11 13:49 | Guimbal Rural Health Unit | Guimbal, Iloilo | 6 | Iloilo | Laboratory | YES | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24501 | PBSP | toto.davis.td@gmail.com | 33 | ASSAY: Xpert MTB/RIF, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1617 |  | GXIV-4D | 807587 | None | None | #NAME? | All modules are functional. | Room AC is not functional, ambient temp is greater than 30C and module temp is ranging 34-37C |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595936219015460791/5595936219015460791_signature_28.png | Davis James Secapuri | https://www.jotform.com/uploads/macarewecare/222927978372067/5595936219015460791/5595936219015460791_signature_31.png
| 2023-05-11 | 2023-05-11 12:50 | Meycauayan City Health Center VII | Pandayan, Meycauayan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24500 | DOH | ronamarieiris_25@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 70 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1272 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 | 2023-05-11 | GXIV | 831904 | #NAME? | #NAME? | For re-XpertCheck | Unit is operational. A1: passed A4: failed A1 under observation. |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595925330019624653/5595925330019624653_signature_28.png | Rona Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5595925330019624653/5595925330019624653_signature_31.png
| 2023-05-11 | 2023-05-09 13:18 | Buluan District Hospital | Buluan, Maguindanao | BARMM | Maguindanao | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24498 | PBSP | rowenasinomandigesmael@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 301 ASSAY: Xpert MTB/RIF Ultra, COUNT: 846 | 2023-05-09 | Gx-IV | 815490 | Module 3 and 4 for Service Required | Replace module 3 SN: 694735 with new module SN: 791953 Replace module 4 SN: 694870 with new module SN: 803313 Perform Genexpert machine Preventive maintenance | Replace module 3 SN: 694735 with new module SN: 791953 Replace module 4 SN: 694870 with new module SN: 803313 Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 color SN:791953 SN:803313 | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595895078415662182/5595895078415662182_signature_28.png | Rowena Esmael | https://www.jotform.com/uploads/macarewecare/222927978372067/5595895078415662182/5595895078415662182_signature_31.png
| 2023-05-11 | 2023-05-11 12:14 | Tacurong City Health Office | Tacurong City | 12 | South Cotabato | Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24499 | PBSP | cho_clinlab@yahoo.com | 27.6 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1857 | 2023-05-11 | Gx-IV | 815393 | Module 3 subject for Module replacement | Perform Genexpert machine Preventive maintenance Perform Genexpert Xpertcheck/Calibration Report module 2 problem due to temperature error | Perform Genexpert machine Preventive maintenance Perform Genexpert Xpertcheck/Calibration Report module 2 problem due to temperature error | 3 operational modules | Module 1,2 and 4 passed the Calibration | Xpertcheck -Ce-5 | 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595895088412789347/5595895088412789347_signature_28.png | John Lorenz Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5595895088412789347/5595895088412789347_signature_31.png
| 2023-05-11 | 2023-05-10 14:23 | San Jose Rural Health Unit | San Jose de Buenavista, Antique | 6 | Antique | Laboratory | YES | REPAIR TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 24496 | PBSP | sanjoseantiquegxsite@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 192 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2244 |  | GXIV-4D | 804714 | Module C3 encountered E2032 | #NAME? | #NAME? | All modules are available. Module C3 to be reported and observed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595870837672853922/5595870837672853922_signature_28.png | Oscar Delano Salazar IV | https://www.jotform.com/uploads/macarewecare/222927978372067/5595870837672853922/5595870837672853922_signature_31.png
| 2023-05-11 | 2023-05-11 11:53 | San Joaquin Rural Health Unit | San Joaquin, Iloilo | 6 | Iloilo | Laboratory | YES | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24497 | PBSP | rhusjlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 95 ASSAY: Xpert MTB/RIF Ultra, COUNT: 250 |  | GXIV-4D | 838887 | None | None | #NAME? | All modules are functional. | Collected one box XPC |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595870837679734170/5595870837679734170_signature_28.png | Romelita Sevillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5595870837679734170/5595870837679734170_signature_31.png
| 2023-05-11 | 2023-05-11 11:45 | Taguig City Health Office | Bay Breeze, Taguig City | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24495 | PBSP | zel091177@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2769 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-05-11 | GX4 6C | 814372 | For module replacement | Replaced module: 864132>>891216 and 766239>>664140 Cleaned module cartridge bay, fan filter, plunger rod Generated reports | Performed module replacement | Unit is operational |  |  | G 5556 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595868559989665299/5595868559989665299_signature_28.png | Liezl dela cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5595868559989665299/5595868559989665299_signature_31.png
| 2023-05-11 | 2023-05-11 11:03 | Baybay City Health Office | Baybay City Health Office | 8 | Leyte | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24494 | DOH | sajecopooh20@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 519 |  | Gx4 | 838555 | 4009 and 1002 | Performed maintenance | Conduct maintanenance, clean Icore, clean plunger Get gxx files, IQR and Sys log Back up database | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595838483422769797/5595838483422769797_signature_28.png | Sarah Jean R. Gabieta, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5595838483422769797/5595838483422769797_signature_31.png
| 2023-05-11 | 2023-05-11 10:59 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24493 | DOH | meggysantiago@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 304 ASSAY: Xpert MTB/RIF Ultra, COUNT: 338 | 2023-05-11 | GXIV | 838548 | N/A | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595834757264399084/5595834757264399084_signature_28.png | MARY GRACE SANTIAGO | https://www.jotform.com/uploads/macarewecare/222927978372067/5595834757264399084/5595834757264399084_signature_31.png
| 2023-05-11 | 2023-05-11 10:38 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 24492 | DOH | dorothygundaya@gmail.com | 27.67 | ASSAY: Xpert MTB/RIF, COUNT: 137 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1818 ASSAY: Xpert HIV-1 Viral Load, COUNT: 47 |  | GX-4 | 838403 | Module B1 and B4 reported Module B3 for XPC error checking | None | XPC Calibration at Module B3 | Only 1 module working 3 modules are prior for replacements |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595824433903302538/5595824433903302538_signature_28.png | Dorothy Gundaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5595824433903302538/5595824433903302538_signature_31.png
| 2023-05-11 | 2023-05-11 10:28 | FAMILY PLANNING ORGANIZATION OF THE PHILIPPINES-ILOILO | RIZAL ST. ILOILO CITY | 6 | Iloilo | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  |  |  |  | 24491 | PBSP | fpopiloilo.lab@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 |  | GX-IV | 838869 | FOR INSTALLATION | COLLECT IQR, ALL MODULES AVAILABLE | PERFORM GENEXPERT INSTALLATION | EQUIPMENT IS OPERATIONAL | #NAME? |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595815768839653992/5595815768839653992_signature_28.png | FPOP ILOILO | https://www.jotform.com/uploads/macarewecare/222927978372067/5595815768839653992/5595815768839653992_signature_31.png
| 2023-05-10 | 2023-05-10 13:38 | Philippine General Hospital | Taft Ave., Ermita | NCR | NCR 1st - Manila | MRL | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24490 | PBSP | mmcajucom@up.edu.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 143 ASSAY: Xpert MTB/RIF Ultra, COUNT: 103 ASSAY: Xpert HIV-1 Viral Load, COUNT: 242 |  | GXIV | 802511 | Unit for PM. | Performed PM. Generated reports. | PM. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595167450161969143/5595167450161969143_signature_28.png | MARC AGNEW M. CAJUCOM | https://www.jotform.com/uploads/macarewecare/222927978372067/5595167450161969143/5595167450161969143_signature_31.png
| 2023-05-10 | 2023-05-10 13:34 | Philippine General Hospital | Taft Ave., Ermita | NCR | NCR 1st - Manila | MRL | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24490 | PBSP | mmcajucom@up.edu.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 143 ASSAY: Xpert MTB/RIF Ultra, COUNT: 103 ASSAY: Xpert HIV-1 Viral Load, COUNT: 242 | 2023-05-08 | GXIV | 838145 | Unit for PM. For module replacement of module B4. | Performed PM. Performed module replacement on B4. Generated reports | PM and module replacement. | Unit is operational. |  | GX Module B4 (604800) in replacement to defective B4 (806707). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595167450169640392/5595167450169640392_signature_28.png | MARC AGNEW M. CAJUCOM | https://www.jotform.com/uploads/macarewecare/222927978372067/5595167450169640392/5595167450169640392_signature_31.png
| 2023-05-10 | 2023-05-09 15:55 | Balanga Rural Health Unit | Sampaguita St. San Jose Balanga,Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24487 | PBSP | cho.melaniealonzo@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 160 ASSAY: Xpert MTB/RIF Ultra, COUNT: 605 | 2023-04-27 | GXII-D | 835102 | Due for calibration | Preventive maintenance Xpertcheck | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration | 2 xpertcheck kits | 40648 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595146239184586586/5595146239184586586_signature_28.png | Melanie Alonzo | https://www.jotform.com/uploads/macarewecare/222927978372067/5595146239184586586/5595146239184586586_signature_31.png
| 2023-05-10 | 2023-05-08 14:27 | Bolinao Rural Health Unit | Bolinao Pangasinan | 1 | Pangasinan | Lqboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24488 | DOH | nylemae14@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: N/a |  | GXIV-D | 838278 | Cpu not booting up | Troubleshooting | Troubleshoot CPU | UPS and CPU not operational for reporting | Possible error defective power supply |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595146239185254170/5595146239185254170_signature_28.png | Emylyn Celzo | https://www.jotform.com/uploads/macarewecare/222927978372067/5595146239185254170/5595146239185254170_signature_31.png
| 2023-05-10 | 2023-05-08 12:35 | Agno Rural Health Office | Poblacion West Agno Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24489 | PBSP | renahumildearenas@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 37 ASSAY: Xpert MTB/RIF Ultra, COUNT: 145 | 2023-04-27 | GXII-D | 834837 | Due for calibration | Calibration Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed | 2 xpertcheck kits | 40639 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595146249186740984/5595146249186740984_signature_28.png | Maria Cris Rena H. Arenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5595146249186740984/5595146249186740984_signature_31.png
| 2023-05-10 | 2023-05-10 15:06 | Riverside Medical Center | BS Aquino Drive, Bacolod City | 6 | Negros Occidental | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24486 | COMMERCIAL | liza.paltriguera@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 6 ASSAY: Xpert HBV VL, COUNT: 62 ASSAY: Xpert HPV, COUNT: 41 ASSAY: Xpert MTB/RIF Ultra, COUNT: 73 ASSAY: Xpert HCV VL, COUNT: 78 |  | Gx 4 | 110005268 | Duedate Xpertcheck Calibration | Cleaning fan filter, XpertCheck calibration | XpertCheck Calibration | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595125915651028362/5595125915651028362_signature_28.png | Liza Mae Paltriguera | https://www.jotform.com/uploads/macarewecare/222927978372067/5595125915651028362/5595125915651028362_signature_31.png
| 2023-05-10 | 2023-05-10 15:12 | Hilongos Rural Health Unit | Hilongos, Poblacion Leyte | 8 | Leyte | Tb Laboratory | YES | INSTALLATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 24485 | PBSP | leovillaflor@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 166 |  | Gx4 | 838143 | Module 1 Serial Number 806782 Error 4009 | Install new module in Module 1 | Replaced module 1 with new module 860586 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595122972718572881/5595122972718572881_signature_28.png | LEONIDES VILLAFLOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5595122972718572881/5595122972718572881_signature_31.png
| 2023-05-10 | 2023-05-10 13:43 | Makati Medical Center | Makati City | NCR | NCR 4th - Makati | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24484 | COMMERCIAL | ejmagante.medilinx@outlook.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 49 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 239 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Others, COUNT: 248 ASSAY: Xpert MTB/RIF Ultra, COUNT: 961 ASSAY: Xpert HCV VL, COUNT: 2 | 2023-05-10 | GXIV | 110001647 | Unit for maintenance. | #NAME? | Performed maintenance | Unit is operational. |  | XpertCheck Kit | 9977 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595081803228534441/5595081803228534441_signature_28.png | Ej Magante | https://www.jotform.com/uploads/macarewecare/222927978372067/5595081803228534441/5595081803228534441_signature_31.png
| 2023-05-10 | 2023-05-10 13:39 | Department of Health, Regional Office II | Maharlika Highway, Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24483 | PBSP | mctorio.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 80 ASSAY: Xpert MTB/RIF Ultra, COUNT: 898 ASSAY: Xpert SARS-CoV-2, COUNT: 1715 | 2023-05-10 | Genexpert IV | 808628 | Module C1 and C2 Hardware Failure | Gathered System Log and IQR | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module C1 and C2 for Reporting | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595068015418830759/5595068015418830759_signature_28.png | Mika Caliguiran | https://www.jotform.com/uploads/macarewecare/222927978372067/5595068015418830759/5595068015418830759_signature_31.png
| 2023-05-10 | 2023-05-10 13:33 | Department of Health, Regional Office II | Maharlika Highway, Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24482 | PBSP | mctorio.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 80 ASSAY: Xpert MTB/RIF Ultra, COUNT: 898 ASSAY: Xpert SARS-CoV-2, COUNT: 1715 | 2023-05-10 | Genexpert IV | 834434 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 805485 - 673003 | Genexpert Module 673003 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595064555414990767/5595064555414990767_signature_28.png | Mika Caliguiran | https://www.jotform.com/uploads/macarewecare/222927978372067/5595064555414990767/5595064555414990767_signature_31.png
| 2023-05-10 | 2023-05-10 13:20 | Department of Health, Regional Office II | Maharlika Highway, Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24481 | PBSP | mctorio.pbsp@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 203 ASSAY: Xpert MTB/RIF, COUNT: 1989 ASSAY: Xpert SARS-CoV-2, COUNT: 4428 | 2023-05-10 | Genexpert XVI | 842894 | Module D3 Stuck Door | Manual Opened the module door | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational, Module D2 Broken Module Door | Module Replacement = 888630 & 888734 - 662108 & 667599 | Genexpert Module = 662108 & 667599 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595061085419177285/5595061085419177285_signature_28.png | Mika Caliguiran | https://www.jotform.com/uploads/macarewecare/222927978372067/5595061085419177285/5595061085419177285_signature_31.png
| 2023-05-10 | 2023-05-10 12:15 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24480 | DOH | lornsabaigar@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 124 ASSAY: Xpert MTB/RIF Ultra, COUNT: 453 | 2023-05-10 | GXIV | 838409 | N/a | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5595017370426595498/5595017370426595498_signature_28.png | lorna abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/5595017370426595498/5595017370426595498_signature_31.png
| 2023-05-10 | 2023-05-10 11:26 | BALASAN RURAL HEALTH UNIT | BALASAN ILOILO | 6 | Iloilo | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24479 | PBSP | charliemagnedk2@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 124 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1412 |  | GX-IV | 807555 | FOR MODULE B2 REPLACEMENT | #NAME? | -PERFORM MODULE B2 DEF SN 635442 REPLACEMENT WITH NEW SN 713796 -PERFORM PREVENTIVE MAINTENANCE | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594988793312292583/5594988793312292583_signature_28.png | Charlie Magne C. Dequito | https://www.jotform.com/uploads/macarewecare/222927978372067/5594988793312292583/5594988793312292583_signature_31.png
| 2023-05-10 | 2023-05-10 11:18 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 24478 | PBSP | cyrilcordova0515@gmail.com | 26 |  |  | GX 4 | 804267 | Defective module 4 sn 660932 | New module installed sn 688723 | New module installed sn 688723 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594981569116558480/5594981569116558480_signature_28.png | Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5594981569116558480/5594981569116558480_signature_31.png
| 2023-05-10 | 2023-05-10 10:48 | Kalinga Health Hub - Mandaluyong | Brgu. Malamig, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24477 | CONSORTIUM | shad.famisaran@innovationsch.org | 28 | ASSAY: Xpert MTB/RIF, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 596 | 2023-05-08 | GX4 | 110008415 | Need for calibration | Checked and cleaned cartridge bay, plunger rod, filter Performed xpertcheck Generated reports | Performed XpertCheck | Unit is operational B1 - passed B2 - passed B3 - passed B4 - failed |  | XpertCheck kit | 9973 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594976239019789921/5594976239019789921_signature_28.png | Shad Famisaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5594976239019789921/5594976239019789921_signature_31.png
| 2023-05-10 | 2023-05-10 10:45 | Kalinga Helath Hub Marikina | J.P. Rizal St., Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24476 | CONSORTIUM | krisna.morante@innovationsch.org | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 653 | 2023-05-08 | GX IV 6C | 110008418 | Module 2 and 4 defective | Module replacement B2 and B4 B2: 220000754 -> 897578 B4: 210035674 -> 210026764 | Performed maintenance Replaced module 2 and 4 Generate reports | Unit is operational, 2 modules newly replaced |  | GX Module 6-Colors | 9977 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594976119011160661/5594976119011160661_signature_28.png | Krisna Morante | https://www.jotform.com/uploads/macarewecare/222927978372067/5594976119011160661/5594976119011160661_signature_31.png
| 2023-05-10 | 2023-05-10 10:48 | Kalinga Health Hub - Mandaluyong | Brgy. Malamig, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24475 | CONSORTIUM | shad.famisaran@innovationsch.org | 28 | ASSAY: Xpert MTB/RIF, COUNT: 53 ASSAY: Xpert MTB/RIF Ultra, COUNT: 655 | 2023-05-08 | GXIV | 110008420 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. B1: passed B2: passed B3: passed B4: passed |  | XpertCheck Kit | 9973 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594975953229837594/5594975953229837594_signature_28.png | Shad Famisaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5594975953229837594/5594975953229837594_signature_31.png
| 2023-05-10 | 2023-05-10 10:59 | Philippine General Hospital | Taft Ave., Ermita | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 24474 | MACARE-HBDC | pghmolecularlab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 479 ASSAY: Xpert SARS-CoV-2, COUNT: 14229 ASSAY: Others, COUNT: SARS-CoV-2_Flu_RSV plus = 10 |  | GX16 | 844975 | Unit for Preventive Maintenance. | Performed Preventive Maintenance. Generated reports. | Preventive Maintenance. | Unit is operational. Generated reports for modules C1, C2, C4 and D4 with temperature errors. Modules C1, C2, C4 and D4 are confirmed for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594972130164863720/5594972130164863720_signature_28.png | Joanne Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5594972130164863720/5594972130164863720_signature_31.png
| 2023-05-10 | 2023-05-10 10:56 | Philippine General Hospital | Taft Ave., Ermita | NCR | NCR 1st - Manila | Molecular laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24473 | PBSP | pghmolecularlab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 248 ASSAY: Xpert SARS-CoV-2, COUNT: 4186 |  | GXIV | 834624 | Unit for Preventive Maintenance | Performe preventive maintenance. Generated reports | Preventive Maintenance | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594971730161869432/5594971730161869432_signature_28.png | Joanne Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5594971730161869432/5594971730161869432_signature_31.png
| 2023-05-10 | 2023-05-09 13:27 | Tacloban City Health Office | Tacloban City | 8 | Leyte | Tb dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24472 | DOH | marteetingzon20@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 343 ASSAY: Xpert MTB/RIF Ultra, COUNT: 918 | 2023-05-09 | GXIV | 838421 | Module B4 (temperature error) | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594923847515932569/5594923847515932569_signature_28.png | Martee Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5594923847515932569/5594923847515932569_signature_31.png
| 2023-05-10 | 2023-05-10 08:27 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24469 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 131 ASSAY: Xpert MTB/RIF Ultra, COUNT: 241 ASSAY: Xpert SARS-CoV-2, COUNT: 1922 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GXIV | 834430 | Unit for PM. | Performed PM. Generated reports. | PM. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594880580168723854/5594880580168723854_signature_28.png | Vernelee Panlaqui | https://www.jotform.com/uploads/macarewecare/222927978372067/5594880580168723854/5594880580168723854_signature_31.png
| 2023-05-10 | 2023-05-10 08:24 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24470 | DOH | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 131 ASSAY: Xpert MTB/RIF Ultra, COUNT: 241 ASSAY: Xpert SARS-CoV-2, COUNT: 1922 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GXIV | 831872 | Unit for PM. | Performed PM. Generated reports. | PM. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594880590162621990/5594880590162621990_signature_28.png | Vernelee Panlaqui | https://www.jotform.com/uploads/macarewecare/222927978372067/5594880590162621990/5594880590162621990_signature_31.png
| 2023-05-10 | 2023-05-10 08:22 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 24471 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 366 ASSAY: Xpert MTB/RIF, COUNT: 459 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1312 ASSAY: Xpert SARS-CoV-2, COUNT: 10882 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 |  | GX16 | 844121 | Unit for PM. For installation of new lan cable. | Performed PM. Performed lan cable installation. Generated reports. | PM and installation. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594880590165683541/5594880590165683541_signature_28.png | Vernelee Panlaqui | https://www.jotform.com/uploads/macarewecare/222927978372067/5594880590165683541/5594880590165683541_signature_31.png
| 2023-05-09 | 2023-05-09 15:39 | Sebaste Rural Health Unit | Sebaste, Antique | 6 | Antique | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24468 | PBSP | francisconaciongayo@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 303 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1812 | 2023-05-09 | GXIV-4D | 808287 | Module B2, B3, and B4 - XPC service required | -Replace defective module B2/673580 with new module SN: 684390 -Replace defective module B3/629736 with new module SN: 688184 -Replace defective module B4/666125 with new module SN: 690576 | #NAME? | All modules are operational. |  | 900-GX GeneXpert Module 6, 3 units | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594276779123397053/5594276779123397053_signature_28.png | FRANCISCO S. NACIONGAYO | https://www.jotform.com/uploads/macarewecare/222927978372067/5594276779123397053/5594276779123397053_signature_31.png
| 2023-05-09 | 2023-05-09 11:41 | Alfonso Ponce Enrile Memorial District Hospital | Mabini st. Smart, Gonzaga, Cagayan | 2 | Cagayan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24467 | DOH | shellyvee1993@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 96 ASSAY: Xpert MTB/RIF Ultra, COUNT: 657 | 2023-02-14 | Genexpert IV | 838474 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = B1, B2, B3 - Passed = B4 - Service Required | Xpertcheck 5pc | 9630 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594274741024197380/5594274741024197380_signature_28.png | SHELLY VEE MAOANA G. GONZALES | https://www.jotform.com/uploads/macarewecare/222927978372067/5594274741024197380/5594274741024197380_signature_31.png
| 2023-05-09 | 2023-05-08 12:41 | Amma Jadsac District Hospital | Poblacion, Pudtol, Apayao | CAR | Apayao | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24466 | PBSP | dlgblacar.pbsp@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2399 | 2023-05-08 | Genexpert IV | 807609 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 681803 - 691108 | Genexpert Module - 691108 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594274741026986724/5594274741026986724_signature_28.png | Daisy Lynn Grace B. Lacar | https://www.jotform.com/uploads/macarewecare/222927978372067/5594274741026986724/5594274741026986724_signature_31.png
| 2023-05-09 | 2023-05-09 14:40 | Nabunturan Rural Health Unit | Poblacion Nabunturan, Davao De Oro | 11 | Davao de Oro | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24465 | WHO | danneth_pcanastra@yahoomail.com | 26.8 |  |  | GX-IV 4d | 847253 | Module B3 for Replacement | Replace the defective Module B3 SN: 910202 with new Module SN 684339 Generate IQR, Assay Stat and System Log | Perform Module Replacement | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594247769963179635/5594247769963179635_signature_28.png | DANNETH P. CANASTRA | https://www.jotform.com/uploads/macarewecare/222927978372067/5594247769963179635/5594247769963179635_signature_31.png
| 2023-05-09 | 2023-05-09 14:24 | Mayor Hilarion A. Ramiro Sr. Medica Center | Ozamiz City | 10 | Misamis Occidental | Molecular Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24464 | PBSP | mharsmc.covidlabgenex@gmail.com | 28 | ASSAY: Xpert SARS-CoV-2, COUNT: 5692 |  | GeneXpert | 834437 | Failed the xpertcheck calibration module B2/784112 | Replace module new module SN/720964 | Preventive maintenance Module Replacement Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594232695919906189/5594232695919906189_signature_28.png | GRETCHEN LANAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5594232695919906189/5594232695919906189_signature_31.png
| 2023-05-09 | 2023-01-27 14:43 | Don Emilio Valle Memorial Hospital | Ubay, Bohol | 7 | Bohol | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24463 | PBSP | dedvmhlaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 567 | 2023-01-19 | Gx4 | 807603 | 4008 | Replaced module | Replace new module | Operational and 4 modules available |  | Module | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594214596915351753/5594214596915351753_signature_28.png | Sharmaine Pagolong | https://www.jotform.com/uploads/macarewecare/222927978372067/5594214596915351753/5594214596915351753_signature_31.png
| 2023-05-09 | 2023-05-09 13:30 | Caraga Regional Hospital | Rizal St. Brgy, Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24462 | PBSP | abpabellan.pbsp@gmail.com | 26.6 | ASSAY: Xpert MTB/RIF, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4351 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 |  | GX-4 | 807602 | None | None | Replacing Module C1 with SN: 642566 to SN: 713394 | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594212490273647650/5594212490273647650_signature_28.png | Andrew B. Pabellan | https://www.jotform.com/uploads/macarewecare/222927978372067/5594212490273647650/5594212490273647650_signature_31.png
| 2023-05-09 | 2023-05-08 09:05 | Sagonsongan Rural Health Unit | Sagonsongan, Marawi City | BARMM | Lanao del Sur | TB DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24461 | DOH | rbarco2007@gmail.com | 29 |  |  | GeneXpert | 838260 | Failed module 3 SN/869703 | Replaced Module | Preventive Maintenance Module Replacement Generate report files Data base back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594179294614871234/5594179294614871234_signature_28.png | Inshira Macabando | https://www.jotform.com/uploads/macarewecare/222927978372067/5594179294614871234/5594179294614871234_signature_31.png
| 2023-05-09 | 2023-05-09 12:25 | Doña Remedios Trinidad Rural Health Unit | DRT, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24460 | PBSP | maricel.rmt24@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 6 ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 594 ASSAY: Xpert HIV-1 Viral Load, COUNT: 37 | 2023-05-09 | GXIV | 838872 | Module D2 (863200) for replacement | #NAME? | Replaced D2 | Unit is operational. |  | GeneXpert Module SN: 838872 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594162622112071087/5594162622112071087_signature_28.png | Cecilia Angeles | https://www.jotform.com/uploads/macarewecare/222927978372067/5594162622112071087/5594162622112071087_signature_31.png
| 2023-05-09 | 2023-04-28 14:12 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  |  |  |  | 24459 | COMMERCIAL | sniperdxcero@gmail.com | 26 |  |  | MCS+ | 15H045G | For calibration | Performed calibration | Performed calibration Checked rotor pumps Checked calibration parameters Simulation run | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594152637815048151/5594152637815048151_signature_28.png | Ferdinand Barroga, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5594152637815048151/5594152637815048151_signature_31.png
| 2023-05-09 | 2023-04-28 14:12 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  |  |  |  | 24458 | COMMERCIAL | sniperdxcero@gmail.com | 26 |  |  | MCS+ | 15H045G | For calibration | Performed calibration | Performed calibration Checked rotor pumps Checked calibration parameters Simulation run | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594152467814817236/5594152467814817236_signature_28.png | Ferdinand Barroga, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5594152467814817236/5594152467814817236_signature_31.png
| 2023-05-09 | 2023-04-28 16:06 | St. Luke's Medical Center - Global City | 5th Avenue | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24456 | COMMERCIAL | ci.bgc@stlukes.com.ph | 25 |  |  | GX-IV-L | 808411 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod, filter Generated reports | Unit is operational. UPS has no backup. Recommend to replace. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594144967815676861/5594144967815676861_signature_28.png | Jane Katherina Navora | https://www.jotform.com/uploads/macarewecare/222927978372067/5594144967815676861/5594144967815676861_signature_31.png
| 2023-05-09 | 2023-04-28 15:59 | St. Luke's Medical Center - Global City | 5th Avenue | NCR | NCR 4th - Taguig | Cellular Immunology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24457 | COMMERCIAL | ci.bgc@stlukes.com.ph | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 153 ASSAY: Xpert CT/NG, COUNT: 1278 ASSAY: Others, COUNT: 11 ASSAY: Xpert HBV VL, COUNT: 77 ASSAY: Xpert HPV, COUNT: 909 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1958 ASSAY: Xpert HCV VL, COUNT: 10 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-IV-D | 839231 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod, filter Generated reports | Unit is operational. UPS has no backup. Recommend to replace. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594144967816327772/5594144967816327772_signature_28.png | Jane Katherina Navora | https://www.jotform.com/uploads/macarewecare/222927978372067/5594144967816327772/5594144967816327772_signature_31.png
| 2023-05-09 | 2023-04-25 14:35 | Marikina City Health Office | Shoe Avenue | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24455 | PBSP | amymitra22@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 142 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GX-IV-D | 805211 | For setup | Setup GeneXpert system | Setup GeneXpert system Performed maintenance Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594143217815619631/5594143217815619631_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5594143217815619631/5594143217815619631_signature_31.png
| 2023-05-09 | 2023-04-28 13:38 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24453 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 15438 ASSAY: Others, COUNT: 9 |  | GX-XVI-D | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod, filter Generated reports | Unit is operational. Modules A1/860941 and B4/860923 not operational, reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594143217811027381/5594143217811027381_signature_28.png | Justine Lorenz V. Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5594143217811027381/5594143217811027381_signature_31.png
| 2023-05-09 | 2023-04-28 13:34 | Rizal Medical Center | Pasig Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24454 | PBSP | molecularlab@rmc.doh.gov.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3206 ASSAY: Xpert SARS-CoV-2, COUNT: 8 ASSAY: Xpert HIV-1 Viral Load, COUNT: 37 |  | GX-IV-D | 814374 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger, filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594143217814502264/5594143217814502264_signature_28.png | Justine Lorenz V. Falzado | https://www.jotform.com/uploads/macarewecare/222927978372067/5594143217814502264/5594143217814502264_signature_31.png
| 2023-05-09 | 2023-04-25 14:51 | Marikina City Health Office | Shove Avenue | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24452 | PBSP | amymitra22@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 213 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1403 ASSAY: Xpert HIV-1 Viral Load, COUNT: 137 |  | GX-IV-D | 804254 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore Generated reports | Unit is operational. Modules D2 and D3 not operational. Reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594143217818545954/5594143217818545954_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5594143217818545954/5594143217818545954_signature_31.png
| 2023-05-09 | 2023-05-09 11:08 | Lapulapu CHO | Pajo Lapulapu City | 7 | Cebu | TB laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 24451 | DOH | tanya_pax@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 741 |  | Gx4 | 837661 | Error 1001, 2014 in Module 1 SN: 656994 | Peformed calibration | Conduct calibration for Module 1 | Modules are available but error occured in Module 1-3. Only module 4 is error free. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5594112478427935174/5594112478427935174_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5594112478427935174/5594112478427935174_signature_31.png
| 2023-05-08 | 2023-05-08 15:53 | Negros Oriental Provincial Hospital | Dumaguete City | 7 | Negros Oriental | Micro Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24450 | DOH | rezelmarielagos@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 699 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1515 ASSAY: Xpert HIV-1 Viral Load, COUNT: 101 |  | GX4 | 838525 | Module B1 lost communication 2120 with SN 806592 | Conduct PM and save files( IQR, System Log, GXC File and Assay Stat) | Run Xpert check calibration | Module B2, B3 & B4 are in good condition |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593435540945309904/5593435540945309904_signature_28.png | Rezel Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5593435540945309904/5593435540945309904_signature_31.png
| 2023-05-08 | 2023-05-05 14:34 | Asian hospital and medical center | Civic drive, Alabang Muntinlupa | NCR | NCR 4th - Muntinlupa | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24449 | COMMERCIAL | jmebres@asianhospital.com | 28 |  | 2023-05-04 | MCS + | 01F116 | For Maintenance and Calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance and Calibration | Operational |  | Urethane roller | 9971 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593425968798237528/5593425968798237528_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5593425968798237528/5593425968798237528_signature_31.png
| 2023-05-08 | 2023-05-05 14:34 | Asian hospital and medical center | Civic drive, Alabang Muntinlupa | NCR | NCR 4th - Muntinlupa | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24448 | COMMERCIAL | jmebres@asianhospital.com | 28 |  | 2023-05-04 | MCS + | 01F116 | For Maintenance and Calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance and Calibration | Operational |  | Urethane roller | 9971 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593425668793431028/5593425668793431028_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5593425668793431028/5593425668793431028_signature_31.png
| 2023-05-08 | 2023-05-05 15:18 | Asian hospital and medical center | Civic dr.,Alabang muntinlupa | NCR | NCR 4th - Muntinlupa | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24447 | COMMERCIAL | jmebres@asianhospital.com | 28 |  | 2023-05-04 | MCS + | 1OB129 | For maintenance and calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance and Calibration | Operational |  | Urethane rollers | 9971 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593425078798022903/5593425078798022903_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5593425078798022903/5593425078798022903_signature_31.png
| 2023-05-08 | 2023-05-05 15:18 | Asian hospital and medical center | Civic dr.,Alabang muntinlupa | NCR | NCR 4th - Muntinlupa | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24446 | COMMERCIAL | jmebres@asianhospital.com | 28 |  | 2023-05-04 | MCS + | 1OB129 | For maintenance and calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance and Calibration | Operational |  | Urethane rollers | 9971 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593424878793785108/5593424878793785108_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5593424878793785108/5593424878793785108_signature_31.png
| 2023-05-08 | 2023-05-08 13:53 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 24445 | DOH | ccmclaboratory01@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 349 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1081 | 2023-05-08 | GXIV-4D | 838561 | Module B1 hardware failed | None | #NAME? | Module B2, B3, B4 - operational; Module B1 - under observation (frequent temp error) | Module B1 encountering frequent temp error, for observation. | 1 XPC cartridge | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593347534716420130/5593347534716420130_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5593347534716420130/5593347534716420130_signature_31.png
| 2023-05-08 | 2023-05-08 13:02 | Manuel J. Santos Hospital | 554 Montilla Blvd. Butuan City | 13 | Agusan del Norte | Molecular Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24444 | COMMERCIAL | molecularlab@mjsantos.com.ph | 18.55 | ASSAY: Xpert SARS-CoV-2, COUNT: 4572 |  | GX-4 | 831292 | Defective Monitor | None | Replacing defective monitor HP ProDisplay 19A to Provision Monitor Preventive Maintenance XpertCheck of all modules | 2 modules passed XPC 2 module service required (for replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593329887581455230/5593329887581455230_signature_28.png | Jubily Dyanne A. Yap | https://www.jotform.com/uploads/macarewecare/222927978372067/5593329887581455230/5593329887581455230_signature_31.png
| 2023-05-08 | 2023-05-05 14:37 | Baliwag Rural Health Unit | Baliwag, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  | rcb.macare@gmail.com |  |  | 24443 | PBSP | baliwagrhulab.003@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 64 ASSAY: Xpert MTB/RIF Ultra, COUNT: 521 | 2023-05-05 | GXIV | 838891 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. B1: passed B2: passed B3: passed B4: passed |  | Xpertcheck kit | NA |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5593320630416340929/5593320630416340929_signature_28.png | Erwin Valenzuela | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5593320630416340929/5593320630416340929_signature_31.png
| 2023-05-08 | 2023-05-05 11:59 | Norzagaray Rural Health Unit I | Norzagaray, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24442 | PBSP | rowenacruz@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 77 ASSAY: Xpert MTB/RIF Ultra, COUNT: 849 | 2023-05-05 | GXIV | 838167 | Module B2 (807097) for replacement. | #NAME? | Replaced B2 | Unit is operational |  | GeneXpert Module SN: 790979 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593320600419707417/5593320600419707417_signature_28.png | Rowena Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5593320600419707417/5593320600419707417_signature_31.png
| 2023-05-08 | 2023-05-08 09:21 | Saguiaran Rural Health Unit | Saguiaran | 10 | Lanao del Sur | TB DOTS | NO | CALIBRATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24441 | DOH | saharaalimaradial.muti@gmail.com | 22 |  |  | GeneXpert | 838298 | Temperature error on module B4/793555 | N/A | Preventive Maintenance Xpertcheck Calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593182709429971343/5593182709429971343_signature_28.png | Sahara Muti | https://www.jotform.com/uploads/macarewecare/222927978372067/5593182709429971343/5593182709429971343_signature_31.png
| 2023-05-08 | 2023-05-08 09:16 | Tuburan Rural Health Unit | Poblacion,Tuburan,Cebu | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 24440 | PBSP | marielgracia20@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 75 | 2023-05-05 | Gx4 | 815485 | Due for Calibration | Perform Xpertcheck | Perform Xpertcheck | All modules pass for Calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593181956916053272/5593181956916053272_signature_28.png | Mariel R. Gracia | https://www.jotform.com/uploads/macarewecare/222927978372067/5593181956916053272/5593181956916053272_signature_31.png
| 2023-05-08 | 2023-05-04 09:07 | Balayan Rural Health Unit | Brgy Gumamela, Balayan | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24439 | PBSP | joanmanguerra@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 665 | 2023-05-03 | GXIV-D | 838451 | Module 3 with Temperature Error | Module Replacement | Replaced Module 3 Performed Maintenance Generate Reports | All Modules are Operational | 861231-> 717833 | 900-0491RS | G 5528 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593167980413293441/5593167980413293441_signature_28.png | Joan Manguerra | https://www.jotform.com/uploads/macarewecare/222927978372067/5593167980413293441/5593167980413293441_signature_31.png
| 2023-05-08 | 2023-05-04 12:01 | Lemery Rural Health Unit | Illustre Ave., Lemery | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24438 | PBSP | maritonie1912@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1013 | 2023-05-03 | GXIV-D | 838179 | 2 Modules have Temperature Error | Module Replacement | Replaced 2 Modules Performed Maintenance Generate Reports | All Modules are Operational | 807150, 807147 -> 636198, 640299 | 900-0491RS | G 5529 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593167980414024858/5593167980414024858_signature_28.png | Maritonie Gardiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5593167980414024858/5593167980414024858_signature_31.png
| 2023-05-08 | 2023-05-05 12:07 | Dasmariñas City Health Office II | Brgy. Sta. Cruz I, Dasmariñas | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24437 | DOH | hazelmilan0777@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 313 | 2023-04-12 | GXIV-D | 838535 | Unit is Due for XpertCheck | XpertCheck | Performed XpertCheck Performed Maintenance Generate Reports | 2 Modules are Newly Replaced, 1 Module Reported, 1 Passed |  | XpertCheck Kit | 9883 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5593167980415273473/5593167980415273473_signature_28.png | Hazel Tabucol | https://www.jotform.com/uploads/macarewecare/222927978372067/5593167980415273473/5593167980415273473_signature_31.png
| 2023-05-06 | 2023-05-06 11:00 | Amai Pakpak Medical Center | Marawi City | BARMM | Lanao del Sur | TB DOTS | YES | INSTALLATION REPAIR | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24436 | PBSP | rbarco2007@gmail.com | 28 |  |  | GeneXpert | 807457 | Stuck cartridge module 3 Communication error module 1 and 3 | Installed new CPU Valve drive adjustment Gateway board reset to fix communication error | New CPU installation Pull out stuck cartridge Reset gateway board Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5591631019421159234/5591631019421159234_signature_28.png | Raisah Barco | https://www.jotform.com/uploads/macarewecare/222927978372067/5591631019421159234/5591631019421159234_signature_31.png
| 2023-05-06 | 2023-05-06 10:48 | Jose B. Lingad Memorial General Hospital | Dolores San Fernando City, Pampamga | 3 | Pampanga | Molecular Lab | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Harold Toco |  |  |  |  | 24435 | PBSP | jblmghmpl@gmail.com | 28 |  |  | Gx4 | 834445 | CPU not booting.No power | Checked CPU connections and performed booting.. | Checked connections.. Power on several times | CPU not functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5591566837582374576/5591566837582374576_signature_28.png | Vince Daryll Navarro | https://www.jotform.com/uploads/macarewecare/222927978372067/5591566837582374576/5591566837582374576_signature_31.png
| 2023-05-06 | 2023-05-06 09:01 | Our Lady Of Mount Carmel Medical Center | San Fernando, Pampanga | 3 | Pampanga | Blood Bank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Harold Toco |  |  |  |  | 24434 | MACARE-HBDC | mt.carmel_laboratory@yahoo.com | 30 |  | 2023-05-06 | MCS+ | 15G043 | Noisy Centrifuge rotation.. | Removed bowl debris. | Checked centrifuge assembly. Removed bowl debris. Run functionality. Run actual UPP procedure. | Unit is operational. |  | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5591480813144666876/5591480813144666876_signature_28.png | Rochelle Loren Guiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5591480813144666876/5591480813144666876_signature_31.png
| 2023-05-06 | 2023-05-06 09:01 | Our Lady Of Mount Carmel Medical Center | San Fernando, Pampanga | 3 | Pampanga | Blood Bank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Harold Toco |  |  |  |  | 24433 | MACARE-HBDC | mt.carmel_laboratory@yahoo.com | 30 |  | 2023-05-06 | MCS+ | 15G043 | Noisy Centrifuge rotation.. | Removed bowl debris. | Checked centrifuge assembly. Removed bowl debris. Run functionality. Run actual UPP procedure. | Unit is operational. |  | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5591480713145507818/5591480713145507818_signature_28.png | Rochelle Loren Guiao | https://www.jotform.com/uploads/macarewecare/222927978372067/5591480713145507818/5591480713145507818_signature_31.png
| 2023-05-06 | 2023-05-06 09:30 | Eastern Visayas Regional Medical Center | Bagacay, Tacloban City | 8 | Leyte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24432 | PBSP | shairamethystcantos@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1579 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3869 ASSAY: Xpert HIV-1 Viral Load, COUNT: 179 | 2023-05-06 | GXIV | 808631 | Module B2 for Replacement | - Module replacement B2/642453 into SN:730362 - Performed preventive maintenance - Performed functionality test - Generate Reports | #NAME? | Unit is Operational |  | GeneXpert Module 6 Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5591465806421779055/5591465806421779055_signature_28.png | Shaira Amethyst C. Nicolas | https://www.jotform.com/uploads/macarewecare/222927978372067/5591465806421779055/5591465806421779055_signature_31.png
| 2023-05-05 | 2023-05-05 14:34 | Asian hospital and medical center | Civic drive, Alabang Muntinlupa | NCR | NCR 4th - Muntinlupa | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24431 | COMMERCIAL | jmebres@asianhospital.com | 28 |  | 2023-05-04 | MCS + | 01F116 | For Maintenance and Calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance and Calibration | Operational |  | Urethane roller | 9971 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5591061485127732237/5591061485127732237_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5591061485127732237/5591061485127732237_signature_31.png
| 2023-05-05 | 2023-05-05 15:18 | Asian hospital and medical center | Civic dr.,Alabang muntinlupa | NCR | NCR 4th - Muntinlupa | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24430 | COMMERCIAL | agimang@asianhospital.com | 28 |  | 2023-05-04 | MCS + | 1OB129 | For maintenance and calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality, replaced rollers,checked parameters, checked pressure | Maintenance and Calibration | Operational |  | Urethane rollers | 9971 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5591060935126387962/5591060935126387962_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5591060935126387962/5591060935126387962_signature_31.png
| 2023-05-05 | 2023-05-05 17:05 | Caloocan City Health Office | Sangandaan, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24429 | PBSP | crissyname@gmail.com | 32 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3027 ASSAY: Xpert HIV-1 Viral Load, COUNT: 401 | 2023-04-28 | GX IV | 814154 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. Module b3 failed XpertCheck |  | XpertCheck kit | DR# 40738 PL-011242 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590888040558772013/5590888040558772013_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5590888040558772013/5590888040558772013_signature_31.png
| 2023-05-05 | 2023-05-05 16:01 | Caloocan City Health Office | Sangandaan, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24428 | PBSP | crissyname@gmail.com | 31 | ASSAY: Xpert MTB/RIF, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3027 ASSAY: Xpert HIV-1 Viral Load, COUNT: 401 | 2023-04-28 | GX IV | 805271 | For XpertCheck | Run xpertcheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. A4 failed XpertCheck |  | XpertCheck kit | DR# 40730 PL-011234 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590882820559782583/5590882820559782583_signature_28.png | Maricris C. Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5590882820559782583/5590882820559782583_signature_31.png
| 2023-05-05 | 2023-05-05 15:37 | Cebu TB Reference Lab | Cebu City | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE Installation | CEPHEID | Raymond Laurito |  |  |  |  | 24427 | PBSP | cabello_clarissa@yahoo.con | 28 |  |  | Gx 4 | 838882 | Transfered in Cebu CHO | Re-installation genexpert machine to Cebu CHO | Re-installation genexpert machine to Cebu CHO | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590817866212574217/5590817866212574217_signature_28.png | Clarissa Cabello | https://www.jotform.com/uploads/macarewecare/222927978372067/5590817866212574217/5590817866212574217_signature_31.png
| 2023-05-05 | 2023-05-05 14:57 | Rosario Rural Health Unit | Poblacion B, Rosario | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24426 | DOH | liberty_yutuc@yahoo.com | 25 |  | 2023-01-17 | GXIV | 821421 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Module B4 (761196) , service required. |  | Xpertcheck Kit | DR 9482 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590811092029734255/5590811092029734255_signature_28.png | Liberty I. Yutuc | https://www.jotform.com/uploads/macarewecare/222927978372067/5590811092029734255/5590811092029734255_signature_31.png
| 2023-05-05 | 2023-05-05 14:58 | San Manuel Rural Health Unit | San Manuel Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24425 | DOH | ganapinhazel22@gmail.com | 33 | ASSAY: Xpert MTB/RIF, COUNT: 76 ASSAY: Xpert MTB/RIF Ultra, COUNT: 384 | 2023-03-23 | GXIV-D | 838516 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational exceot B4:862532 | 3 modules passed calibration, B4 for reporting | 4 xpertcheck kits | 9775 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590808561519752556/5590808561519752556_signature_28.png | Hazel Ganapin | https://www.jotform.com/uploads/macarewecare/222927978372067/5590808561519752556/5590808561519752556_signature_31.png
| 2023-05-05 | 2023-05-05 14:18 | Julian Felipe Health Center | Julian Felipe St., Brgy 8, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24424 | PBSP | maoi.lloyd@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1342 | 2023-05-03 | GX IV 6C | 838884 | Module 4 defective | Module 4 replacement from 863606 to 210046118 | Performed maintenance Cleaned cartridge bays, plunger rod, icores, replaced filter Generate reports | Unit is operational 1 newly replaced | B4: 863606 to 210046118 | GX module 6C | G 5538 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590773820559689287/5590773820559689287_signature_28.png | Lloyd Maoi De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5590773820559689287/5590773820559689287_signature_31.png
| 2023-05-05 | 2023-05-05 13:41 | Lipa City Health Office | Morada St. Lipa | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24423 | PBSP | mylaolayao@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1028 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 | 2023-05-02 | GXIV | 805329 | Unit for PM. For module replacement of module B3. | Performed PM. Performed module replacement on B3. | PM and module replacement. | Unit is operational. |  | Gx module B3 (210041477) in replacement to defective B3 (621002). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590752807822112008/5590752807822112008_signature_28.png | Myla Olayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5590752807822112008/5590752807822112008_signature_31.png
| 2023-05-05 | 2023-05-05 13:41 | Lipa City Health Office | Morada St. Lipa | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24422 | PBSP | mylaolayao@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1028 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 | 2023-05-02 | GXIV | 805329 | Unit for PM. For module replacement of module B3. | Performed PM. Performed module replacement on B3. | PM and module replacement. | Unit is operational. |  | Gx module B3 (210041477) in replacement to defective B3 (621002). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590752577823419296/5590752577823419296_signature_28.png | Myla Olayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5590752577823419296/5590752577823419296_signature_31.png
| 2023-05-05 | 2023-05-05 13:47 | Lipa City Health Office | Morada St. Lipa | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24421 | PBSP | mylaolayao@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1028 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 | 2023-05-02 | GXIV | 808642 | Unit for PM. For module replacement of modules C2, C3 and C4. | Performed PM. Performed module replacement of modules C2, C3 and C4. Generated reports. | PM and module replacement | Unit is operational. |  | Gx module C2 (600607), C3 (608356) and C4 (625061) in replacement to defective C2 (667796), C3 (667798) and C4 (667799). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590751807829386391/5590751807829386391_signature_28.png | Myla M. Olayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5590751807829386391/5590751807829386391_signature_31.png
| 2023-05-05 | 2023-05-03 15:08 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | TB-DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24420 | PBSP | vincentmilestone20@gmail.com | 22.6 |  |  | GX-XVI | 838756 | Powersupply is faulty | Replace the powersupply of the GeneXpert XVI, checking the connection and continuity of circuit Return the Service Module SN: 870474 Perform XpertCheck Calibration of the 16 Modules | Replacement of Powersupply | Modules A4, B2, B3, C1, C2, D4 for service required | Data will be sent to Cepheid |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590735423311397592/5590735423311397592_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5590735423311397592/5590735423311397592_signature_31.png
| 2023-05-05 | 2023-05-03 15:08 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | TB-DOTS | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24419 | PBSP | vincentmilestone20@gmail.com | 22.6 |  |  | GX-XVI | 838756 | Powersupply is faulty | Replace the powersupply of the GeneXpert XVI, checking the connection and continuity of circuit Return the Service Module SN: 870474 Perform XpertCheck Calibration of the 16 Modules | Replacement of Powersupply | Modules A4, B2, B3, C1, C2, D4 for service required | Data will be sent to Cepheid |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590735383311769523/5590735383311769523_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5590735383311769523/5590735383311769523_signature_31.png
| 2023-05-05 | 2023-05-05 12:57 | Sisin Rural Health Unit | Poblacion Central Sison Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24418 | PBSP | sisonrhu@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 80 ASSAY: Xpert MTB/RIF Ultra, COUNT: 286 | 2023-04-27 | GXIV-D | 834684 | B2:784440 for module replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 Genexpert module | 40626 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590721601512697322/5590721601512697322_signature_28.png | Jeannette O. Alcaide | https://www.jotform.com/uploads/macarewecare/222927978372067/5590721601512697322/5590721601512697322_signature_31.png
| 2023-05-05 | 2023-05-05 09:55 | San Jacinto Rural Health Unit | Capaoay Poblacion San Jacinto Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24416 | DOH | allen.indong@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 269 | 2023-03-23 | GXIV-D | 838518 | Due for Calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration | 4 xpertcheck kits | 9778 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590714811511597295/5590714811511597295_signature_28.png | Alexander F. Indong | https://www.jotform.com/uploads/macarewecare/222927978372067/5590714811511597295/5590714811511597295_signature_31.png
| 2023-05-05 | 2023-05-04 15:41 | Dasol Community Hospital | Bobonot Dasol Pangasinan | 1 | Pangasinan | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24415 | DOH | pamelajeanbejar@yahoo.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 104 ASSAY: Xpert MTB/RIF Ultra, COUNT: 453 | 2023-02-02 | GXIV-D | 838287 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modukes passed calibration | 4 xpertcheck kits | 9576 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590714811518980755/5590714811518980755_signature_28.png | Pamela Jean B. Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5590714811518980755/5590714811518980755_signature_31.png
| 2023-05-05 | 2023-05-04 09:04 | Subic Rural Health Unit | Subic, Zambales | 3 | Zambales | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24417 | PBSP | sanchezmacristina313@gmail.com | 31 | ASSAY: Xpert MTB/RIF, COUNT: 761 ASSAY: , COUNT: 2099 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 | 2023-04-27 | GXIV-D | 812788 | B1 for module replacement, due for calibration | Calibration Module replacement Preventive maintenance | Performed preventive maintenance Xpertcheck application update to 1.5a Performed calibration Performed module replacement Gathered reports Performed functionality test | Unit is operational | All modules passed calibration | 1 Genexpert Module, 3 xpertcheck kits | 0040627 and 0040651 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590714821513788296/5590714821513788296_signature_28.png | Ma. Cristina Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/5590714821513788296/5590714821513788296_signature_31.png
| 2023-05-05 | 2023-05-05 11:00 | Sumilang Super Health Center | Dr. Garcia St Cor Lopez Jaena Brgy Sumilang Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24414 | PBSP | maricor.obnamia05@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1271 | 2023-05-03 | GX 4 | 814163 | For Module replacement | -performerd module replacement for Module B2 676490 >>> 896529 and Module B3 696434 >>> 902304 -Cleaned Cartridge Bays, Plunger Rods, Filter and I-CORE -Generated Reports | Perform preventive maintenance | Unit is operational | Module B4 for reporting | -GeneXpert Module 6-color SN: 896529 -GeneXpert Module 6-color SN: 902304 | G 5534 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590706520419662904/5590706520419662904_signature_28.png | Maria Corazon Obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/5590706520419662904/5590706520419662904_signature_31.png
| 2023-05-05 | 2023-05-05 11:00 | Sumilang Super Health Center | Dr. Garcia St Cor Lopez Jaena Brgy Sumilang Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24413 | PBSP | maricor.obnamia05@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1271 | 2023-05-03 | GX 4 | 814163 | For Module replacement | -performerd module replacement for Module B2 676490 >>> 896529 and Module B3 696434 >>> 902304 -Cleaned Cartridge Bays, Plunger Rods, Filter and I-CORE -Generated Reports | Perform preventive maintenance | Unit is operational | Module B4 for reporting | -GeneXpert Module 6-color SN: 896529 -GeneXpert Module 6-color SN: 902304 | G 5534 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590706350417813126/5590706350417813126_signature_28.png | Maria Corazon Obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/5590706350417813126/5590706350417813126_signature_31.png
| 2023-05-05 | 2023-05-05 10:52 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Quezon District, Mabini st, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24412 | PBSP | bryanmichaelbalunes@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 1742 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3057 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 205 | 2023-05-05 | Genexpert IV | 844111 | Module 1 Lost Communication error | Troubleshoot the Lan Cable | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 892012 - 780938 | Genexpert Module = 780938 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590698922972379346/5590698922972379346_signature_28.png | Bryan Michael Balunes | https://www.jotform.com/uploads/macarewecare/222927978372067/5590698922972379346/5590698922972379346_signature_31.png
| 2023-05-05 | 2023-05-04 13:12 | Guimba Community Hospital | Guimba, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24411 | PBSP | marquezandrelyn18@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 58 ASSAY: Xpert MTB/RIF Ultra, COUNT: 961 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 | 2023-05-04 | Genexpert IV | 807560 | N/a | N/a | Performed Maintenanc3 Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 633463 - 758156 | Genexpert Module = 758156 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590698922977926153/5590698922977926153_signature_28.png | Andrelyn Marquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5590698922977926153/5590698922977926153_signature_31.png
| 2023-05-05 | 2023-05-05 11:06 | Pedro Gil Health Center | Pedro Gil St., Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24410 | PBSP | lorrainefabros@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 116 | 2023-05-03 | GXIV 6C | 812829 | Module 1 and 3 for replacement | Module replacement | Performed maintenance Replaced module 1 from 724845 to 868124 Replaced module 3 from 658655 to 885647 Generate Reports | Unit is operational 2 nrwly replaced | B1: 724845 to 868124 B3: 658655 to 885647 | GX Module 6C - 2mod | G 5533 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590680610556146660/5590680610556146660_signature_28.png | Lorraine L Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5590680610556146660/5590680610556146660_signature_31.png
| 2023-05-05 | 2023-05-05 11:29 | Batangas Medical Center | Batangas City | 4A | Batangas | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24409 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2023-05-02 | GXIV | 834627 | Unit for PM. For module replacement of module B2. | Performed PM. Performed module replacement on B3. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B2 (677854) in replacement to defective B2 (784301) | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590678417822095985/5590678417822095985_signature_28.png | Patricia Louisse Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5590678417822095985/5590678417822095985_signature_31.png
| 2023-05-05 | 2023-05-05 11:32 | Batangas Medical Center | Batangas City | 4A | Batangas | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24408 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2023-05-02 | GXIV | 804683 | Unit for PM. For module replacement of modules A2 and A3. | Performed PM. Performed module replacement on modules A2 and A3. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module A2 (685157) and A3 (686792) in replacement to defective A2 (613959) and A3 (602011). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590677947825969801/5590677947825969801_signature_28.png | Patricia Louisse Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5590677947825969801/5590677947825969801_signature_31.png
| 2023-05-05 | 2023-05-05 09:41 | Samar Provincial Hospital | Catbalogan Samar | 8 | Samar | TB dots | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24407 | PBSP | leslieleones36@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 942 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1181 | 2023-05-05 | GXIV | 808291 | #NAME? | - Module Replacement C3/804720 into 714095 - Performed Xpertcheck on Module C1 and C4 - Performed preventive medicine - Performed functionality test - Generate Reports | #NAME? | Unit is Operational |  | GeneXpert Module 6 Color and Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590606179468609885/5590606179468609885_signature_28.png | Leslie Daryl Leones | https://www.jotform.com/uploads/macarewecare/222927978372067/5590606179468609885/5590606179468609885_signature_31.png
| 2023-05-05 | 2023-05-05 09:34 | Rosario Super Health Center | Dr. Sixto Antonio Ave., Rosario, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24406 | PBSP | eden.cawayan@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 23 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1885 | 2023-05-03 | GX-4 | 843924 | For module replacement and preventive maintenance | Performed module replacement(638232>>684120) Cleaned cartridge bays, plunger rod Collected report | Performed preventive maintenance | Unit is operational |  | Gene Xpert module 684120 | G 5535 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590605229618897402/5590605229618897402_signature_28.png | Eden R. Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5590605229618897402/5590605229618897402_signature_31.png
| 2023-05-05 | 2023-05-04 13:36 | National Children's Hospital | E. Rodriguez Avenue, Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24405 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 22 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 7 ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2659 ASSAY: Xpert SARS-CoV-2, COUNT: 7371 ASSAY: Xpert SARS-CoV-2, COUNT: 18 |  | GX XVI 10C | 110014171 | For preventive maintenanace | Cleaned cartridge bays, plunger rod, icore and filters Generate report | Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5590566790555803559/5590566790555803559_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5590566790555803559/5590566790555803559_signature_31.png
| 2023-05-04 | 2023-05-04 15:00 | National.kidney and transplant institute | East ave,Quezon city | NCR | NCR 2nd - Quezon City | Mole | NO | Checkup | CEPHEID | Noel Floresca |  |  |  |  | 24404 | MACARE-HBDC | molecularlabnkti@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 386 ASSAY: Others, COUNT: 8 ASSAY: Xpert HBV VL, COUNT: 195 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1243 ASSAY: Xpert MTB/RIF Ultra, COUNT: 201 ASSAY: Xpert SARS-CoV-2, COUNT: 18920 ASSAY: Xpert HCV VL, COUNT: 21 ASSAY: Xpert SARS-CoV-2, COUNT: 9 |  | GXXVI | 110005587 | Disable A3,C4,temp error,loss comm.C1 | Gathered data | Checked up | A3,C4,C1 disable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589958909106548833/5589958909106548833_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5589958909106548833/5589958909106548833_signature_31.png
| 2023-05-04 | 2023-05-04 15:31 | Rogaciano M. Mercado Memorial General Hospital | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24403 | PBSP | rhia.aragon@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 121 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1797 |  | GXIV | 805252 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. A1: passed A2: failed A3: passed A4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589950504293120706/5589950504293120706_signature_28.png | Rhia Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/5589950504293120706/5589950504293120706_signature_31.png
| 2023-05-04 | 2023-05-04 14:35 | Cebu TB reference lab | Cebu City | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE Packing of genexpert unit | CEPHEID | Raymond Laurito |  |  |  |  | 24402 | PBSP | asmantiquilla.pbsp@gmail.com | 28 |  |  | GX IV | 806201 | Need to transfer to Cebu CHO | Packing of genexpert unit | Packing of genexpert unit | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589920017113588984/5589920017113588984_signature_28.png | ALLAN S. MANTIQUILLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5589920017113588984/5589920017113588984_signature_31.png
| 2023-05-04 | 2023-05-04 10:06 | San Lazaro Hospital | Quiricada St. | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24401 | PBSP | acong.pbsp@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2060 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1083 | 2023-05-03 | GX 4 | 804766 | For module replacement | -Performed module replacement 606103 >>718638 -Cleaned cartridge bays and Plunger rods -Generated reports | Perform preventive maintenance | Unit is operational | Work with Carlos Origenes and Dan Bodino | GeneXpert Module SN 718638 | G 5536 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589915270413192250/5589915270413192250_signature_28.png | Ashley Cynna Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5589915270413192250/5589915270413192250_signature_31.png
| 2023-05-04 | 2023-05-04 14:01 | Toledo cho | Ilihan Toledo City | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24400 | PBSP | loriemaediapen@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 790 |  | Gx4 | 808640 | Module B2 Serial Number 691154 failed during calibration | Replaced module B2 with new module 791542 | Replacement of defective module B2 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589895943916168887/5589895943916168887_signature_28.png | Lore Mae C. Diapen | https://www.jotform.com/uploads/macarewecare/222927978372067/5589895943916168887/5589895943916168887_signature_31.png
| 2023-05-04 | 2023-05-04 13:42 | Dr. Rafael S. Tumbokon Memorial Hospital | Kalibo, Aklan | 6 | Aklan | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24399 | PBSP | gonzagaevarmt@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 299 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3186 ASSAY: Xpert HIV-1 Viral Load, COUNT: 81 |  | GX 4 | 804680 | Due to calibration schedule | Conduct PM and save files(IQR, System Log, Assay Stat and GXC file) | Run Xpert check calibration for module D1, D3 and D4 | Modules D1, D3 & D4 passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589895930317960672/5589895930317960672_signature_28.png | Marniel C. Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5589895930317960672/5589895930317960672_signature_31.png
| 2023-05-04 | 2023-05-04 13:46 | San Francisco Rural Health Unit | Brgy. 4, San Francisco | 13 | Agusan del Sur | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24398 | PBSP | periabrasangelica@gmail.com | 29.2 | ASSAY: Xpert MTB/RIF, COUNT: 121 ASSAY: Xpert MTB/RIF Ultra, COUNT: 786 |  | GX-4 | 835078 | None | None | Replacing Module B2 with SN: 677485 to SN: 696050 | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589889874415882984/5589889874415882984_signature_28.png | Ma. Angelica O. Periabras | https://www.jotform.com/uploads/macarewecare/222927978372067/5589889874415882984/5589889874415882984_signature_31.png
| 2023-05-04 | 2023-05-04 13:29 | National Childrens Hospital | E.rodriguez ave,Quezon city | NCR | NCR 2nd - Quezon City | Molecular lab | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24397 | PBSP | nchmolecularlab2021@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 131 ASSAY: Xpert MTB/RIF Ultra, COUNT: 174 ASSAY: Xpert SARS-CoV-2, COUNT: 1050 |  | GX4 | 803109 | For maintenance | Cleaned unit,filter,icore,cartridge bay,plunger rod,back up gxx,gathered reports | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589883299105378179/5589883299105378179_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5589883299105378179/5589883299105378179_signature_31.png
| 2023-05-04 | 2023-05-04 11:32 | Northern Samar Provincial Hospital | Catarman Northern Samar | 8 | Northern Samar | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24396 | PBSP | dinahada70@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 719 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1612 ASSAY: Xpert HIV-1 Viral Load, COUNT: 29 | 2023-05-04 | GXIV | 807566 | Module B4 for Replacement | -Module replacement B4/676759 into 657532 - Performed Xpertcheck Calibration -Performed preventive maintenance - Generate Reports | #NAME? | Unit is Operational ( Module B1 for monitoring temperature error) |  | GeneXpert Module 6 Color, Xpertcheck kit | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589809308856554074/5589809308856554074_signature_28.png | Dinah F. Ada | https://www.jotform.com/uploads/macarewecare/222927978372067/5589809308856554074/5589809308856554074_signature_31.png
| 2023-05-04 | 2023-05-04 10:58 | National childrens hospital | E.rodriguez,Quezon city | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24395 | MACARE-HBDC | zarabellehilario@yahoo.com | 27 |  |  | MCs + | 17H632SPG | For maintenance and calibration | Cleaned unit, Checked roller, checked centrifuge,checked dpm,spm,cuff,checked parameters,performed functionality test,test run unit | Maintenance and calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589804157111875494/5589804157111875494_signature_28.png | Zarabelle Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5589804157111875494/5589804157111875494_signature_31.png
| 2023-05-04 | 2023-05-04 10:05 | San Lazaro Hospital | Quiricada St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24394 | PBSP | acong.pbsp@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2060 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1083 | 2023-05-03 | GX-4 R1 | 707843 | For module replacement and preventive maintenance | Performed module replacement(641703>>674311) cleaned cartridge bay, plunger rod. Collected reports | Performed preventive maintenance | Unit is operational | Worked with sir Carlos | Gene xpert module 674311 | G 5536 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589758343823563858/5589758343823563858_signature_28.png | Ashley Cynna Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5589758343823563858/5589758343823563858_signature_31.png
| 2023-05-04 | 2023-05-04 09:20 | Molave Rural health unit | Molave, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24393 | DOH | aldrinacain@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 835 | 2023-05-04 | GX 4 | 838233 | Module 1, 2, 3 and 4 error 1002, 1001, 4008 and | Preventive Maintenence. Module 1 and 4 service required Module 2 passed Module 3 retry | All modules undergone Xpertcheck and Preventive Maintenence | Module 2, 3 and 4 are available, module 1 h/w failed | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589729315564500696/5589729315564500696_signature_28.png | Aldrin Acain | https://www.jotform.com/uploads/macarewecare/222927978372067/5589729315564500696/5589729315564500696_signature_31.png
| 2023-05-03 | 2023-05-03 16:40 | Cebu TB reference lab | Cebu City | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24392 | PBSP | asmantiquilla.pbsp@gmail.com | 25 |  |  | GX 4 | 838263 | Packing of genexpert machine | Packing of genexpert machine | Packing of genexpert machine | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589128009156822298/5589128009156822298_signature_28.png | ALLAN S. MANTIQUILLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5589128009156822298/5589128009156822298_signature_31.png
| 2023-05-03 | 2023-05-03 13:48 | Juan Posadas Health Center | Rodriguez St. | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24391 | PBSP | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 15 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1200 | 2023-05-02 | GXIV | 805060 | Unit for PM. For module replacement of module B4. | Performed PM. Performef module replacement on module B4. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B4 (632590) in replacement to defective B4 (619323) | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589029147826588092/5589029147826588092_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/5589029147826588092/5589029147826588092_signature_31.png
| 2023-05-03 | 2023-05-03 13:28 | Democrito O. Plaza Memorial Hospital | National Highway, Brgy. Patin-ay, Prosperidad | 13 | Agusan del Sur | RTDL | NO | REPAIR TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 24390 | PBSP | lermabpabillan@yahoo.com | 20.63 | ASSAY: Xpert MTB/RIF, COUNT: 341 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3017 |  | GX-4 | 805335 | Intermittent Errors occurs at Module C3 | Cleaning and Reseating RAM slots Changing CMOS battery for Real Time Clock | Replacing Module C1 with SN: 885698 to SN: 803321 Troubleshooting of CPU Desktop | 3 modules working and 1 module under observation. CPU is functional (to be observed when power surges occurs) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589025446419139679/5589025446419139679_signature_28.png | Lerma B. Pabillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5589025446419139679/5589025446419139679_signature_31.png
| 2023-05-03 | 2023-05-03 13:28 | Democrito O. Plaza Memorial Hospital | National Highway, Brgy. Patin-ay, Prosperidad | 13 | Agusan del Sur | RTDL | NO | REPAIR TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 24389 | PBSP | lermabpabillan@yahoo.com | 20.63 | ASSAY: Xpert MTB/RIF, COUNT: 341 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3017 |  | GX-4 | 805335 | Intermittent Errors occurs at Module C3 | Cleaning and Reseating RAM slots Changing CMOS battery for Real Time Clock | Replacing Module C1 with SN: 885698 to SN: 803321 Troubleshooting of CPU Desktop | 3 modules working and 1 module under observation. CPU is functional (to be observed when power surges occurs) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5589024346417904766/5589024346417904766_signature_28.png | Lerma B. Pabillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5589024346417904766/5589024346417904766_signature_31.png
| 2023-05-03 | 2023-05-03 12:31 | Parola Health Center | Parola, Tondo | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24388 | DOH | friaslynette@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 686 | 2023-05-03 | GXIV | 831871 | Unit for PM. For re-Xpertcheck of module B4. | Performed PM. Performed re-Xpertcheck on module B4. Generated reports. | PM and re-Xpertcheck. | Unit is operational. Module B4 passed re-Xpertcheck. | Used 2 spare xpertcheck. | Xpertcheck | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588989117828085122/5588989117828085122_signature_28.png | Lynette Olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5588989117828085122/5588989117828085122_signature_31.png
| 2023-05-03 | 2023-05-03 12:10 | Kamuning Health Center | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24387 | DOH | vicvic.chua25@yahoo.com | 25 | ASSAY: Others, COUNT: 687 ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 536 |  | GXIV | 831612 | Temp. errors on Modules A2 and A4 | #NAME? | Performed XpertCheck to further diagnose | Unit is operational. B2: failed B4: failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588980277616011900/5588980277616011900_signature_28.png | Myrna De Lara | https://www.jotform.com/uploads/macarewecare/222927978372067/5588980277616011900/5588980277616011900_signature_31.png
| 2023-05-03 | 2023-05-03 09:26 | Macare In-house | Aurora Grand | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24386 | COMMERCIAL | florescanoel77@gmail.com | 28 |  | 2023-05-03 | GX4 | 821257 | For xpert check & maintenance | Cleaned cartridge bay, plunger rod, icore Checked cpu, monitor, barcode scanner performed xpert check Collected reports | Performed xpert check & preventive maintenance | Unit is operational |  | Xpert check kit | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588955040415025983/5588955040415025983_signature_28.png | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5588955040415025983/5588955040415025983_signature_31.png
| 2023-05-03 | 2023-05-03 10:04 | Tanza Rural Health Unit | Municipal Compound, Tanza | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24385 | PBSP | gutierrezrhian@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 476 | 2023-05-03 | GXIV-D | 838162 | 2 Modules Failed in XpertCheck | Module Replacement | Replaced 2 Modules Performed Maintenance Generate Reports | 3 Modules are Operational, 1 For Replacement | 807005, 807006 -> 714809, 717149 | 900-0491RS | G 5527 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588918922018512521/5588918922018512521_signature_28.png | MARY ANN C. GUTIERREZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5588918922018512521/5588918922018512521_signature_31.png
| 2023-05-03 | 2023-05-03 09:24 | Macare Medicals, Inc. | 1005 Aurora Blvd, Project 3, Quezon City, 1102 | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24384 | COMMERCIAL | florescanoel77@gmail.com | 28 |  | 2023-05-03 | GX - IV | 821257 | For XpertCheck & Maintenance | Run Xpertcheck | Performed maintenance Run xpertcheck Generated reports | Unit is operational. All modules passed xpertcheck |  | Xpertcheck kit | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588868130414496849/5588868130414496849_signature_28.png | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5588868130414496849/5588868130414496849_signature_31.png
| 2023-05-03 | 2023-05-03 09:20 | Macare In-house | 1005 Aurora Grand | NCR | NCR 2nd - Quezon City | Macare In-house | NO | CALIBRATION MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24383 | COMMERCIAL | florescanoel77@gmail.com | 28 |  | 2023-05-03 | GX 4 | 821257 | For Xpert Check and Maintenance | #NAME? | Xpert Check and Maintenance | Unit is operational | All modules passed Xpert Check | XpertCheck | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588866080417513561/5588866080417513561_signature_28.png | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5588866080417513561/5588866080417513561_signature_31.png
| 2023-05-02 | 2023-04-27 14:04 | Philippine Children's Medical Center | Quezon ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24382 | PBSP | mjcordova02@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1681 | 2023-04-27 | GX4 | 807672 | For xpertcheck and maintenance | Performed xpertcheck and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Xpertcheck and maintenance | Moduleb1,b4 passed xpertcheck | B2 and b3 reported already | Xpertcheck-ce-5 | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588416175128473456/5588416175128473456_signature_28.png | Mark Joseph Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5588416175128473456/5588416175128473456_signature_31.png
| 2023-05-02 | 2023-04-27 09:08 | Datu Halun Sakilan Memorial Hospital | Datu halun, bonggao Tawi - Tawi. | 9 | Tawi-Tawi | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24381 | PBSP | alnijarelias@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 14 | 2023-04-27 | GX 4 | 808643 | Module 4 h/w failed | Replaced module 731450 with Module 892988 | Module Replacement | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588214973519631569/5588214973519631569_signature_28.png | Alnijar s. Elias | https://www.jotform.com/uploads/macarewecare/222927978372067/5588214973519631569/5588214973519631569_signature_31.png
| 2023-05-02 | 2023-05-02 14:17 | Tomas Claudio Health Center | Corner E. Quirino, Camus St. Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24380 | DOH | arleen_p_alconera@yahoo.com | 27 |  |  | GX-IV 4d | 838492 | Module B1 encounter temperature related error | Run XPC on Module B1 | Perform Module Troubleshooting | Module B1 SN:862214 for service required Reports sent to Cepheid for further Assessment | Module B1 temporarily disabled |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588175676321805343/5588175676321805343_signature_28.png | Arleen P. Alconera | https://www.jotform.com/uploads/macarewecare/222927978372067/5588175676321805343/5588175676321805343_signature_31.png
| 2023-05-02 | 2023-05-02 14:12 | Tomas Claudio Health Center (District A) | Corner E. Quirino, Camus St. Davao City | 11 | Davao del Sur | TB-DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24379 | DOH | arleen_p_alconera@yahoo.com | 27.8 |  |  | GX-IV 4D | 838492 | Module B1 encounter temperature related error | Perform XPC at Module B1 | Perform Module Troubleshooting | Module B1 SN:862214 For Service Required | Module B1 temporarily disable |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588173246322536136/5588173246322536136_signature_28.png | Arleen Alconera RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5588173246322536136/5588173246322536136_signature_31.png
| 2023-05-02 | 2023-05-02 13:51 | Mandaue City Health Office | Mandaue City Cebu | 7 | Cebu | Tb laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 24378 | DOH | elma.bragat@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 790 |  | Gx4 | 838508 | Error 2014 | Performed calibration | Conduct and perform calibration of Genxpert 838508 Module Serial number 629957 | Module 1 passed, modulse 3 and 4 are operational. B2 is no longer used waiting for module replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588165836913222559/5588165836913222559_signature_28.png | Elma Bragat Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5588165836913222559/5588165836913222559_signature_31.png
| 2023-05-02 | 2023-05-02 11:09 | Bagumbong Health Center | Bagumbong, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 24377 | DOH | jenilyn_duenas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2352 |  | GXIV | 831907 | Desktop is not booting up properly. | #NAME? | Reset RAM | Unit is operational. Desktop is under observation. Found temp. error on A4. Advised end user to monitor A4. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588090842576231145/5588090842576231145_signature_28.png | Jenilyn Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5588090842576231145/5588090842576231145_signature_31.png
| 2023-05-02 | 2023-04-27 12:33 | Medical Center Manila Inc. | 850 UN Ave., Paco, Manila | NCR | NCR 1st - Manila | Microbiology Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24376 | CONSORTIUM | pathology.laboratories@manilamed.com.ph | 29 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 881 | 2023-04-25 | GXIV - 6C | 835540 | B4 Failed Xpertcheck due for module replacement | Performed module replacement on B4 old module: 791638 new: 916540 Cleaned cartridge bays, icore, plunger rod maintenance, filter Generated reports Performed self-test Advised end-user about ambient temp of laboratory | Module replacement Preventive maintenance | Unit is operational |  | Gene xpert module 6C | 9928 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588007540417651410/5588007540417651410_signature_28.png | Frances Velarde | https://www.jotform.com/uploads/macarewecare/222927978372067/5588007540417651410/5588007540417651410_signature_31.png
| 2023-05-02 | 2023-04-27 12:33 | Medical Center Manila Inc. | 850 UN Ave., Paco, Manila | NCR | NCR 1st - Manila | Microbiology Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24375 | CONSORTIUM | pathology.laboratories@manilamed.com.ph | 29 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 881 | 2023-04-25 | GXIV - 6C | 835540 | B4 Failed Xpertcheck due for module replacement | Performed module replacement on B4 old module: 791638 new: 916540 Cleaned cartridge bays, icore, plunger rod maintenance, filter Generated reports Performed self-test Advised end-user about ambient temp of laboratory | Module replacement Preventive maintenance | Unit is operational |  | Gene xpert module 6C | 9928 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5588007470418993758/5588007470418993758_signature_28.png | Frances Velarde | https://www.jotform.com/uploads/macarewecare/222927978372067/5588007470418993758/5588007470418993758_signature_31.png
| 2023-05-02 | 2023-04-25 13:19 | Alabang Health Center | Alabang-Zapote Road | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24372 | DOH | cathycarasig@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1238 | 2023-04-25 | GXIV-D | 831913 | Monitor with Lines in Display | Monitor Replacement | Replaced Monitor Performed Maintenance Generate Reports | All Modules are Operational | CNK83107VS->MMA1PSS00221507C8C3W01 | FC-E-Monitor | 9924 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5587971420414819207/5587971420414819207_signature_28.png | Catherine carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/5587971420414819207/5587971420414819207_signature_31.png
| 2023-05-02 | 2023-04-27 08:40 | Romblon Provincial Hospital | Brgy. Liwanag, Odiongan, Romblon | 4B | Romblon | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24374 | PBSP | mjfaderogaya.rmt@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 628 |  | GXIV-D | 807459 | Modules 2 and 4 H/W Failed Unit is due for XpertCheck | Module Replacement | Replaced Modules 2 and 4 Performed Maintenance Performed XpertCheck Generate Reports | 2 Modules are Newly Replaced, 1 Module Passed, 1 Module Failed | 210014984, 661376 -> 651738, 653984 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5587971430411166146/5587971430411166146_signature_28.png | Fivi F. Manliguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5587971430411166146/5587971430411166146_signature_31.png
| 2023-05-02 | 2023-04-26 10:13 | General Emilio Aguinaldo Memorial Hospital | Trece-Indang Rd., Trece Martires | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24373 | PBSP | djaemiel@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1723 ASSAY: Xpert HIV-1 Viral Load, COUNT: 103 |  | GXIV-L | 802860 | Unit is Due for XpertCheck | XpertCheck | Performed XpertCheck Performed Maintenance Generate Reports | All Modules are Operational | Xpertcheck: 1 and 2 Passed, 3 and 4 Newly Replaced |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5587971430416936123/5587971430416936123_signature_28.png | JAEMIEL JUVIRENS A. DIMARANAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5587971430416936123/5587971430416936123_signature_31.png
| 2023-05-02 | 2023-04-28 15:50 | The Medical City | Origami Ave., Paris City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24371 | COMMERCIAL | moleculardx@themedicalcity.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 29 ASSAY: Xpert HPV, COUNT: 37 ASSAY: Xpert MTB/RIF Ultra, COUNT: 296 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 | 2023-04-29 | GX16 10C | 110015767 | For module replacement B3 210082269 - 210022184 | cleaned unit, cartridge bay, plunger rod, filter, icore. Gathered reports, performed visual check, test run unit. Replaced module | Module replacement, maintenance | Unit is operational |  | 900-0650RS | G-5516 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5587960750411807509/5587960750411807509_signature_28.png | Xchosa Kay Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/5587960750411807509/5587960750411807509_signature_31.png
| 2023-04-30 | 2023-04-26 11:44 | Hilongos Rural Health Unit | Hilongos leyte | 8 | Leyte | TB dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24370 | PBSP | leovillaflor@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 207 ASSAY: Xpert MTB/RIF Ultra, COUNT: 694 | 2023-04-26 | GXIV | 838143 | #NAME? | #NAME? | #NAME? | Unit is Operational (Module B1 disabled - waiting for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5586607481322385927/5586607481322385927_signature_28.png | Leonides Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5586607481322385927/5586607481322385927_signature_31.png
| 2023-04-30 | 2023-04-26 11:44 | Hilongos Rural Health Unit | Hilongos leyte | 8 | Leyte | TB dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24369 | PBSP | leovillaflor@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 207 ASSAY: Xpert MTB/RIF Ultra, COUNT: 694 | 2023-04-26 | GXIV | 838143 | #NAME? | #NAME? | #NAME? | Unit is Operational (Module B1 disabled - waiting for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5586607251322742890/5586607251322742890_signature_28.png | Leonides Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5586607251322742890/5586607251322742890_signature_31.png
| 2023-04-30 | 2023-04-30 10:30 | Makati Medical Center | Salcedo St. Makati City | NCR | NCR 4th - Makati | Laboratory (Medi Linx) | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 24367 | COMMERCIAL | reymirsanapo.medilinx@outlook.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 43 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 227 ASSAY: Others, COUNT: GBS G3 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 897 ASSAY: Xpert CT/NG, COUNT: NG 2 ASSAY: Xpert HCV VL, COUNT: 19 ASSAY: Xpert HIV-1 Viral Load, COUNT: 571 |  | GXIV-10C | 110001647 | Unit for PM. For troubleshooting of Loss Communication problem. | Performed PM. Performed troutbleshooting. Replaced all module Lan cables. Checked ultrasonic horns on all modules and gateway board. Generated reports. | PM and troubleshooting. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5586350434811766334/5586350434811766334_signature_28.png | Reymir Sanapo | https://www.jotform.com/uploads/macarewecare/222927978372067/5586350434811766334/5586350434811766334_signature_31.png
| 2023-04-30 | 2023-04-28 13:47 | Dimasalang Health Center | Sta Cruz | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24368 | PBSP | kikokendz_060830@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 14 ASSAY: Xpert MTB/RIF, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 841 | 2023-04-27 | GXIV | 707072 | Unit for PM. For module replacement of modules B2 and B4. | Performed PM. Performed module replacement on B2 and B4. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B2 (629868) and B4 (631438) in replacement to defective B2 (693886) and B4 (634018). | G 5514 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5586350444811096690/5586350444811096690_signature_28.png | Maria joy cleofas m. Abenales | https://www.jotform.com/uploads/macarewecare/222927978372067/5586350444811096690/5586350444811096690_signature_31.png
| 2023-04-29 | 2023-04-27 10:56 | Lake Sebu Rural Health Unit | Lake Sebu | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24366 | DOH | jennypingkian23@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 149 ASSAY: Xpert MTB/RIF Ultra, COUNT: 352 |  | Gx-IV | 838569 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5585894441411577358/5585894441411577358_signature_28.png | Jenny A. Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/5585894441411577358/5585894441411577358_signature_31.png
| 2023-04-28 | 2023-04-28 17:44 | St Lukes Extension Clinic | J Bocobo Street Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24365 | MACARE-HBDC | laboratory@slec.ph | 24 | ASSAY: Xpert CT/NG, COUNT: 4707 ASSAY: Xpert MTB/RIF, COUNT: 520 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4659 | 2023-04-25 | Infinity 48s | 809643 | Module C1 for replacement | Replace module C1 (877673) with module (677392) | Replace module C1 Performed maintenance Cleaning | Unit is operational | Modules A2 and F2 for reporting, Work with Patrik Los Banos, Gian Paolo Magnaye and Cienne Valerio | GeneXpert Module 6-color SN: 677392 | DR 9927 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584853044114538247/5584853044114538247_signature_28.png | Maria Teresita D. Baclig | https://www.jotform.com/uploads/macarewecare/222927978372067/5584853044114538247/5584853044114538247_signature_31.png
| 2023-04-28 | 2023-04-28 15:24 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24364 | COMMERCIAL | jblmgh.is@gmail.com | 29 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 ASSAY: Xpert HBV VL, COUNT: 5 | 2023-04-27 | Genexpert IV | 812571 | Cannot Start Test HCV Viral Load | Update HCV Viral Load Assay | Performed Maintenance Gathered Reports Replaced UPS Replaced Barcode Scanner Performed Functionality Test | Unit is Operational | N/a | Barcode Scanner, UPS | 9959, 9916 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584850226777749716/5584850226777749716_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5584850226777749716/5584850226777749716_signature_31.png
| 2023-04-28 | 2023-04-28 15:50 | The Medical City | Origami Ave., Paris City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24363 | COMMERCIAL | moleculardx@themedicalcity.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 29 ASSAY: Xpert HPV, COUNT: 37 ASSAY: Xpert MTB/RIF Ultra, COUNT: 296 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 5 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 | 2023-04-29 | GX16 10C | 110015767 | For module replacement B3 210082269 - 210022184 | cleaned unit, cartridge bay, plunger rod, filter, icore. Gathered reports, performed visual check, test run unit. Replaced module | Module replacement, maintenance | Unit is operational |  | 900-0650RS | G-5516 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584803195517585659/5584803195517585659_signature_28.png | Xchosa Kay Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/5584803195517585659/5584803195517585659_signature_31.png
| 2023-04-28 | 2023-04-28 15:41 | Cagayan de Oro City Health Office | Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24362 | DOH | chocdotbdots@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3099 |  | GeneXpert | 838441 | Error 1001 on Module B1/806412 | None | Preventive Maintenance Xpertcheck Calibration Generate report files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584775340271719886/5584775340271719886_signature_28.png | Kate Diane Liquit | https://www.jotform.com/uploads/macarewecare/222927978372067/5584775340271719886/5584775340271719886_signature_31.png
| 2023-04-28 | 2023-04-28 14:33 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 24361 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 197 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1392 ASSAY: Xpert MTB/RIF, COUNT: 965 |  | GXIV-4D | 838285 | Module A3 - frequent temp error | None | #NAME? | Module A3 - Xpertcheck service required, for replacement. All other three modules are operational. | 2 XPC cartridge used. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584733435411761047/5584733435411761047_signature_28.png | Jogeveve Alforque | https://www.jotform.com/uploads/macarewecare/222927978372067/5584733435411761047/5584733435411761047_signature_31.png
| 2023-04-28 | 2023-04-28 13:44 | Monkayo Rural Health Unit | Poblacion, Monkayo Davao De Oro | 11 | Davao de Oro | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24360 | DOH | kendralois730@gmail.com | 29.5 |  |  | GX-IV | 838487 | Stuck cartridge and for Preventive Maintenance | Troubleshoot the module that has stuck cartridge in it Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Module Troubleshooting and Perform Preventive Maintenance | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584700290886710832/5584700290886710832_signature_28.png | Karen S Anuta | https://www.jotform.com/uploads/macarewecare/222927978372067/5584700290886710832/5584700290886710832_signature_31.png
| 2023-04-28 | 2023-04-28 11:04 | Macare In-house | Aurora Grand | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Yasser Salih |  |  |  | UPS | 24359 | DOH | carlos.Ieniel.origenes@gmaiI.com | 25 |  |  | 1 KVA | AT1000181128ST0032 | -Won't power on -Won't backup | None | #NAME? | Unit is operational | Removed one good battery and removed one bad battery |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584611820415173767/5584611820415173767_signature_28.png | Macare In-house | https://www.jotform.com/uploads/macarewecare/222927978372067/5584611820415173767/5584611820415173767_signature_31.png
| 2023-04-28 | 2023-04-28 11:09 | Macare In-house | Aurora Grand | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Dan Emil Bodino |  |  |  | UPS | 24358 | DOH | carlos.leniel.origenes@gmail.com | 25 |  |  | 1KVA | AT1000181128ST0100 | Won't power on, Won't backup | None | Checked fuse, batteries and motherboard | Unit is not operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584610660419367874/5584610660419367874_signature_28.png | Macare In-house | https://www.jotform.com/uploads/macarewecare/222927978372067/5584610660419367874/5584610660419367874_signature_31.png
| 2023-04-28 | 2023-04-28 10:57 | Macare In-house | Aurora Grand | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | OTHERS | Yasser Salih |  |  |  | UPS | 24357 | DOH | carlos.leniel.origenes@gmail.com | 25 |  |  | 1 KVA | AT1000180910ST0226 | Won't power on Won't Back up | None | #NAME? | Unit is not operational | Removed working batteries |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584602240418432425/5584602240418432425_signature_28.png | Macare In-House | https://www.jotform.com/uploads/macarewecare/222927978372067/5584602240418432425/5584602240418432425_signature_31.png
| 2023-04-27 | 2023-04-27 11:30 | Mangaldan Rural Health Unit | Poblacion Mangaldan,Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24356 | DOH | arleneabad32@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 414 | 2023-04-27 | GXIV-D | 838288 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration | 4 xpertcheck kits | 9779 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584051269411883962/5584051269411883962_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5584051269411883962/5584051269411883962_signature_31.png
| 2023-04-27 | 2023-04-27 10:15 | Sta. Barbara Rural health Unit | Poblacion Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24355 | DOH | dianeruiz147@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 191 ASSAY: Xpert MTB/RIF Ultra, COUNT: 198 | 2023-03-23 | GXIV-D | 838438 | Due for Calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational |  | 4 xpertcheck kits | 9777 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584051269418987548/5584051269418987548_signature_28.png | Diane Joy Q. Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5584051269418987548/5584051269418987548_signature_31.png
| 2023-04-27 | 2023-04-27 10:15 | Sta. Barbara Rural health Unit | Poblacion Sta. Barbara, Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24354 | DOH | dianeruiz147@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 191 ASSAY: Xpert MTB/RIF Ultra, COUNT: 198 | 2023-03-23 | GXIV-D | 838438 | Due for Calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational |  | 4 xpertcheck kits | 9777 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584036699413111585/5584036699413111585_signature_28.png | Diane Joy Q. Ruiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5584036699413111585/5584036699413111585_signature_31.png
| 2023-04-27 | 2023-04-27 11:30 | Mangaldan Rural Health Unit | Poblacion Mangaldan,Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24353 | DOH | arleneabad32@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 87 ASSAY: Xpert MTB/RIF Ultra, COUNT: 414 | 2023-04-27 | GXIV-D | 838288 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration | 4 xpertcheck kits | 9779 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5584036699415882463/5584036699415882463_signature_28.png | Arlene Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5584036699415882463/5584036699415882463_signature_31.png
| 2023-04-27 | 2023-04-27 15:15 | General Tinio Rural Health Unit | Poblacion Central, General Tinio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 24352 | DOH | jadelyn_pajarillaga@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 71 ASSAY: Xpert MTB/RIF Ultra, COUNT: 690 | 2023-04-27 | Genexpert IV | 831695 | Module 3 Error | Performed Xpertcheck in Module 3 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck Result = Service Required | Xpertcheck 2pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583933958425744680/5583933958425744680_signature_28.png | Jadelyn Pajarillaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5583933958425744680/5583933958425744680_signature_31.png
| 2023-04-27 | 2023-04-27 15:45 | Samboan Rural Health Unit | Poblacion Samboan Cebu | 7 | Cebu | Tb laboratory | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24351 | DOH | sheneannesenagan@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 199 |  | Gx4 | 838513 | Due for maintenance | Performed maintenance | Performed maintenance clean Icore, clean filter and plunger clean up | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583910607713576321/5583910607713576321_signature_28.png | Shene Ann A. Senagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5583910607713576321/5583910607713576321_signature_31.png
| 2023-04-27 | 2023-04-27 13:35 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 24350 | DOH | ccmclaboratory01@gmail.com | 23 |  |  | GXIV-4D | 838561 | Module A1 - frequent temp error | None | #NAME? | Module A1 is disabled. Module A2, A3, & A4 are operational. | Module A1 for further observation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583832548587982379/5583832548587982379_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5583832548587982379/5583832548587982379_signature_31.png
| 2023-04-27 | 2023-03-07 10:27 | National kidney and transplant institute | East ave qc | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 24349 | MACARE-HBDC | molecularlabnkti@gmail.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 880 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Others, COUNT: 51 ASSAY: Xpert HBV VL, COUNT: 442 ASSAY: Xpert HPV, COUNT: 19 ASSAY: Xpert MTB/RIF, COUNT: 540 ASSAY: Xpert SARS-CoV-2, COUNT: 450 ASSAY: Xpert HCV VL, COUNT: 74 |  | GX4 | 838524 | B1 862548 and B4 862557 encounter continous temperature error 1001,1002 | Checked b1 and b4 module,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Troubleshooting and maintenance | For replacement b1 862548 and b4 862557 due to continous temperature error encountered |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583793337253442452/5583793337253442452_signature_28.png | Ma. Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5583793337253442452/5583793337253442452_signature_31.png
| 2023-04-27 | 2023-04-25 14:29 | D.santiago health center | D.santiago st | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 24348 | DOH | fei_paas@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 65 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1057 |  | GX4 | 831527 | Ups unable to backup defective battery | Delivered and installed 2 pcs battery,test run unit | Repair | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583790107254877356/5583790107254877356_signature_28.png | Flory Fe Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5583790107254877356/5583790107254877356_signature_31.png
| 2023-04-27 | 2023-04-27 10:45 | Teresita L. Jalandoni Provincial Hospital | Silay City | 6 | Negros Occidental | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24347 | PBSP | mrvillan.pbsp@gmail.com | 4 |  |  | GX IV | 805093 | Module A4 with SN No. 791669 and error code 2126 | Conduct PM and save files for IQR, System Log and GXC File | Run Xpert check calibration for module A4 only. | Module A1, A2 and A3 is in good conditions. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583748913114474084/5583748913114474084_signature_28.png | Ms. Arna Villan | https://www.jotform.com/uploads/macarewecare/222927978372067/5583748913114474084/5583748913114474084_signature_31.png
| 2023-04-27 | 2023-04-27 08:54 | Salvacion Oppus - Yniguez Memorial Provincial Hospital | Maasin City | 8 | Southern Leyte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24346 | PBSP | loreensalva@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 682 ASSAY: Xpert MTB/RIF Ultra, COUNT: 882 | 2023-04-27 | GXIV | 804262 | - Module B1 error (1002) | #NAME? | #NAME? | Unit is Operational | Module B2 and B3 passed, Module B1 and B4 service required | Xpertcheck Cartridges | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583665395117485163/5583665395117485163_signature_28.png | Agnes Jill M Obligado | https://www.jotform.com/uploads/macarewecare/222927978372067/5583665395117485163/5583665395117485163_signature_31.png
| 2023-04-27 | 2023-04-26 12:41 | Umingan Community Hospital | Burgos St. Poblacion west Umingan Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24345 | DOH | jeffreyvrumbaoa@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 632 ASSAY: Xpert MTB/RIF Ultra, COUNT: 824 | 2023-02-02 | GXIV-D | 838473 | Due for calibration | Xpertcheck Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except B1 and B4, failed xpertcheck | For reporting | 6 xpertcheck kits | 9577 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583545409413434354/5583545409413434354_signature_28.png | Jeffrey Rumbaoa | https://www.jotform.com/uploads/macarewecare/222927978372067/5583545409413434354/5583545409413434354_signature_31.png
| 2023-04-27 | 2023-04-25 12:07 | Bayambang District Hospital | Bayambang Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24344 | PBSP | bevzneri@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 898 | 2023-04-11 | GXIV-D | 812830 | Gateway board replacement | Gateway board replacement Preventive maintenance | Performed surface cleaning Performed preventive maintenance Performed gateway board replacement Performed Serial number configuration Gathered reports Performed functionality test | Unit is operational |  | 1 Gateway board | 40021 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583544019417651023/5583544019417651023_signature_28.png | Beverly Neri | https://www.jotform.com/uploads/macarewecare/222927978372067/5583544019417651023/5583544019417651023_signature_31.png
| 2023-04-26 | 2023-04-26 12:52 | Jose B. Lingad Memorial Regional Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24343 | COMMERCIAL | jblmgh.is@gmail.com | 29 |  | 2023-04-26 | Genexpert IV | 812571 | SQL not Detected | Reinstalled Genexpert Application Troubleshoot the scanner Calibrate the barcode scannner | Disable built in keyboard Installed HCV Installed HBV Installed HIV Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Barcode scanner not functioning, for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5583003980512564288/5583003980512564288_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5583003980512564288/5583003980512564288_signature_31.png
| 2023-04-26 | 2023-04-26 14:13 | Philippine Red-Cross zamboanga chapter | Petit barracks zamboanga city | 9 | Zamboanga del Sur | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  |  |  |  | 24342 | COMMERCIAL | zamboangacity@redcross.org.ph | 25 |  |  | Mcs + | 01M256 | Preventive maintenance | Preventive maintenance | Preventive maintenance | Systems operational | PM with gian carlo aizon |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582993197995025994/5582993197995025994_signature_28.png | Maria Sandra Anitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5582993197995025994/5582993197995025994_signature_31.png
| 2023-04-26 | 2023-04-26 13:22 | Nabas Rural Health Unit | Nabas Aklan | 6 | Aklan | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24341 | PBSP | lambito.godelyn61@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 0 |  | GX-IV | 815387 | INSTALL NEW SYSTEM UNIT | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582960961625277524/5582960961625277524_signature_28.png | Godelyn A Lambito | https://www.jotform.com/uploads/macarewecare/222927978372067/5582960961625277524/5582960961625277524_signature_31.png
| 2023-04-26 | 2023-04-26 10:11 | Minglanilla RHU | Ward 1, Pider, Minglanilla Cebu | 7 | Cebu | Tb Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24340 | DOH | brendagonzagaparedes@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 694 |  | Gx 4 | 838512 | Error 4009 | Clean Icore, Plunger and clean filter Restart module Performed calibration of module 2 due to error Error 4009 | Perfomed calibration of Module 2 Serial number 862343 | All modules are operational except module 2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582847494214569122/5582847494214569122_signature_28.png | Brenda Gonzaga Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/5582847494214569122/5582847494214569122_signature_31.png
| 2023-04-26 | 2023-04-26 10:01 | DR. RAFAEL S. TUMBOKON MEMORIAL HOSPITAL | KALIBO AKLAN | 6 | Aklan | MICROBIOLOGY LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 24339 | PBSP | gonzagaevarmt@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 361 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3186 ASSAY: Xpert HIV-1 Viral Load, COUNT: 81 |  | GX-IV | 804680 | FOR GATEWAY BOARD REPLACEMENT | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | MODULE D1 FOR OBSERVATION AFTER GATEWAY BOARD REPLACEMENT |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582845907332797135/5582845907332797135_signature_28.png | Eva Raquel D. Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5582845907332797135/5582845907332797135_signature_31.png
| 2023-04-25 | 2023-04-25 14:43 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24338 | PBSP | tindocvelche@gmail.com | 25.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 5411 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-04-25 | Gx-IV | 834442 | Module 2 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Replace module 2 SN:636313 with new module SN: 867625 | Perform Genexpert machine Preventive maintenance Replace module 2 SN:636313 with new module SN: 867625 | 4 operational modules |  | Genexpert module 6 colors | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582215182323066038/5582215182323066038_signature_28.png | Velche B. Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5582215182323066038/5582215182323066038_signature_31.png
| 2023-04-25 | 2023-04-25 14:43 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24337 | PBSP | tindocvelche@gmail.com | 25.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 5411 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-04-25 | Gx-IV | 834442 | Module 2 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Replace module 2 SN:636313 with new module SN: 867625 | Perform Genexpert machine Preventive maintenance Replace module 2 SN:636313 with new module SN: 867625 | 4 operational modules |  | Genexpert module 6 colors | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582215152329441032/5582215152329441032_signature_28.png | Velche B. Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5582215152329441032/5582215152329441032_signature_31.png
| 2023-04-25 | 2023-04-25 14:38 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24336 | PBSP | tindocvelche@gmail.com | 25.3 | ASSAY: Xpert SARS-CoV-2, COUNT: 5411 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-04-25 | Gx-IV | 834441 | Module 4 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Replace module 4 SN:640800 with new module SN: 805746 | Perform Genexpert machine Preventive maintenance Replace module 4 SN:640800 with new module SN: 805746 | 4 operational modules |  | Genexpert module 6 colors | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582214362321008512/5582214362321008512_signature_28.png | Velche B. Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5582214362321008512/5582214362321008512_signature_31.png
| 2023-04-25 | 2023-04-25 14:38 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24335 | PBSP | tindocvelche@gmail.com | 25.3 | ASSAY: Xpert SARS-CoV-2, COUNT: 5411 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-04-25 | Gx-IV | 834441 | Module 4 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Replace module 4 SN:640800 with new module SN: 805746 | Perform Genexpert machine Preventive maintenance Replace module 4 SN:640800 with new module SN: 805746 | 4 operational modules |  | Genexpert module 6 colors | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582214272324348200/5582214272324348200_signature_28.png | Velche B. Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5582214272324348200/5582214272324348200_signature_31.png
| 2023-04-25 | 2023-04-25 14:38 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24334 | PBSP | tindocvelche@gmail.com | 25.3 | ASSAY: Xpert SARS-CoV-2, COUNT: 5411 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-04-25 | Gx-IV | 834441 | Module 4 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Replace module 4 SN:640800 with new module SN: 805746 | Perform Genexpert machine Preventive maintenance Replace module 4 SN:640800 with new module SN: 805746 | 4 operational modules |  | Genexpert module 6 colors | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582214242326509063/5582214242326509063_signature_28.png | Velche B. Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5582214242326509063/5582214242326509063_signature_31.png
| 2023-04-25 | 2023-04-25 15:24 | San Leonardo Rural Health Unit | Diversion, San Leonardo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24333 | PBSP | ligayadomingo20@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 136 ASSAY: Xpert MTB/RIF Ultra, COUNT: 778 | 2023-04-25 | Genexpert IV | 846042 | Cannot perform Test | Troubleshoot the Ethernet cable Configure the IP Address | Performed Maintenance Gathered Reports Installed Printer Drivers Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582208022991745611/5582208022991745611_signature_28.png | Ligaya Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5582208022991745611/5582208022991745611_signature_31.png
| 2023-04-25 | 2023-04-24 14:05 | Gapan Rural Health Unit II | San Nicolas, Gapan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24332 | PBSP | emmaliwaggutierrez@gmail.com | 29 |  | 2023-04-24 | Genexpert Iv | 814182 | N/a | Performed Xpertcheck on B4 module | Installed Genexpert Installed the Computer Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertchech Result - B1, B2, B3 = Passed, B4 = Service Required | Xpertcheck 5pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582208022995837188/5582208022995837188_signature_28.png | Emma Gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5582208022995837188/5582208022995837188_signature_31.png
| 2023-04-25 | 2023-04-25 15:10 | Markina City Health Office | Shoe Ave. | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | Dan Emil Bodino |  |  |  |  | 24331 | DOH | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 286 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1165 ASSAY: Xpert HIV-1 Viral Load, COUNT: 109 |  | GX-4 6C | 831588 | For maintenance | Performed maintenance | Performed maintenance: Cleaned plunger rod, icore lens, cartridge bays. Generated reports. | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582165664612348404/5582165664612348404_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5582165664612348404/5582165664612348404_signature_31.png
| 2023-04-25 | 2023-04-25 13:57 | A.H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24330 | PBSP | dinajaen1990@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 417 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1934 ASSAY: Xpert SARS-CoV-2, COUNT: 4644 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX XVI | 839292 | Unit for installation | #NAME? | Installed unit | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582141433757483180/5582141433757483180_signature_28.png | Dina Jaen | https://www.jotform.com/uploads/macarewecare/222927978372067/5582141433757483180/5582141433757483180_signature_31.png
| 2023-04-25 | 2023-04-25 10:28 | Butuan Medical Center | KM 5. Baan, Butuan City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24329 | PBSP | bmccrcl2022@gmail.com | 28.77 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1027 ASSAY: Xpert SARS-CoV-2, COUNT: 1208 ASSAY: Xpert HIV-1 Viral Load, COUNT: 72 |  | GX-4 | 805099 | Module B3 found service required from XPC | None | XPC calibration and Preventive Maintenance | 3 modules working, 1 modules service required (intermittent errors, use when needed) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5582002054262187967/5582002054262187967_signature_28.png | Joechita L. Doysabas | https://www.jotform.com/uploads/macarewecare/222927978372067/5582002054262187967/5582002054262187967_signature_31.png
| 2023-04-25 | 2023-04-25 10:11 | Jose Reyes Memorial Medical Center | Rizal Ave. Sta Cruz Manilla | NCR | NCR 1st - Manila | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24328 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 1069 ASSAY: Xpert SARS-CoV-2, COUNT: 4152 | 2023-04-24 | Gx 16 | 110013856 | Unit for PM. For module replacement of module C3 | Performed PM Performed module replacement Generated reports | PM and module replacement. | Unit is operational | Work with Shed Daroya | Gx modue C3(210007212) in replacement to defective module C3(210069240) | DR 9923 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581989432867981478/5581989432867981478_signature_28.png | NELSON MARK SALVADOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5581989432867981478/5581989432867981478_signature_31.png
| 2023-04-25 | 2023-04-25 10:16 | Carcar Cityhealth Office | Valladolid Carcar City Cebu | 7 | Cebu | Tb Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24327 | DOH | chaska.lp14@gmail.co | 25 | ASSAY: Xpert MTB/RIF, COUNT: 492 |  | Gx 4 | 838283 | Error 1001 | Conduct and performed calibration | Conduct Calibration of Module 2 Serial 802769 Clean Filter, Icore and plunger Get data gxx files, IQR, Sys log and Assag Statistics | Module 2 Serial Number 838283 failed during caibration. The other modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581985172779568761/5581985172779568761_signature_28.png | Lucille Mae Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5581985172779568761/5581985172779568761_signature_31.png
| 2023-04-25 | 2023-04-25 08:55 | E. RODRIGUEZ HEALTH CENTER | 150 Pajo, Project 2, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Paolo Magnaye |  |  |  |  | 24326 | DOH | clairebaptista32786@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 998 ASSAY: Xpert HPV, COUNT: 211 | 2023-04-24 | GX4 6C | 831848 | Module 1, 2, 3 Faulty | Replaced module 757727 with 632957 | Performed preventive maintenance Generate Reports Replaced module 3 Cleaning cartridge bays, icores, pluungers and filters Performed functionality test | Unit is operational with 2 out of 4 modules |  | Genexpert moduke 6C 632957 | DR# 9922 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581960300414394718/5581960300414394718_signature_28.png | Mary claire baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/5581960300414394718/5581960300414394718_signature_31.png
| 2023-04-24 | 2023-04-19 13:48 | Antipolo CHO | M.Santos,Antipolo city | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24325 | PBSP | acho.laboratory@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 8642 ASSAY: Xpert HIV-1 Viral Load, COUNT: 638 | 2023-04-17 | GXXVI | 842226 | For module replacement and maintenance B4 886430 - 881319 | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900-0491rs | G5452 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581482743471001312/5581482743471001312_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5581482743471001312/5581482743471001312_signature_31.png
| 2023-04-24 | 2023-04-19 09:40 | Taytay rhu 5 | Kadalagahan st,taytay rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24324 | PBSP | dannelli11@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 195 | 2023-04-17 | GX4 | 812836 | For module replacement b2 639463 - 893033 | Performed module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900-0491rs | G5453 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581482743476969318/5581482743476969318_signature_28.png | Dino Delos Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5581482743476969318/5581482743476969318_signature_31.png
| 2023-04-24 | 2023-04-24 15:58 | Coron Rural Health Unit | Poblacion District, Coron, Palawan | 4B | Palawan | TB-DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24323 | PBSP | mecaarmateo04@gmail.com | 25.7 |  |  | GX-IV 4D | 807612 | For Preventive Maintenance and Module Replacement | Replace the defective Module B4 SN:662451 with new module SN:656889 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance and Module Replacement | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581332237223448912/5581332237223448912_signature_28.png | Marie Eloise Armateo | https://www.jotform.com/uploads/macarewecare/222927978372067/5581332237223448912/5581332237223448912_signature_31.png
| 2023-04-24 | 2023-04-13 11:39 | The Black Nazarene Hospital Inc. | San Nicolas, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | Ronnel Belizar |  |  |  |  | 24321 | CONSORTIUM | tbnhi.lab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 19 | 2023-04-11 | GXIV | 835542 | Unit for Xpertcheck. | #NAME? | Performed Xpertcheck | Unit is operational. B1: failed B2: passed B3: passed B4: failed |  | Xpertcheck kit | 9876 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560412470784/5581241560412470784_signature_28.png | Joanna Marie Salud | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560412470784/5581241560412470784_signature_31.png
| 2023-04-24 | 2023-04-12 16:18 | Ilocos Training and Regional Medical Center | San Fernando, La Union | 1 | La Union | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24322 | PBSP | itrmc.tblaboratory@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 702 ASSAY: Xpert MTB/RIF, COUNT: 555 ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 ASSAY: Others, COUNT: 623 ASSAY: Xpert SARS-CoV-2, COUNT: 322 | 2023-04-12 | GXIV | 848249 | Module B4 (210002727) for replacement | #NAME? | Replaced B4 | Unit is operational. |  | GeneXpert Module SN: 210002902 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560414710245/5581241560414710245_signature_28.png | Jennylen Torres | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560414710245/5581241560414710245_signature_31.png
| 2023-04-24 | 2023-04-12 11:38 | La Union Medical Center | Agoo, La Union | 1 | La Union | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 24320 | PBSP | lumcrtdl@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 168 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1462 | 2023-04-12 | GXIV | 707400 | Modules A1 (647587), A2 (682801), A4 (784357) for replacement | #NAME? | Replaced modules | Unit is operational. |  | GeneXpert Module SN: 678987, 690638, 693507 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560414266613/5581241560414266613_signature_28.png | Josephine Paglingayen | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560414266613/5581241560414266613_signature_31.png
| 2023-04-24 | 2023-04-13 15:30 | Gabriela Silang General Hospital | Vigan, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24320 | PBSP | isphgsstc@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 132 ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 | 2023-04-13 | GXIV | 807578 | Module D4 (793228) for replacement | #NAME? | #NAME? | Unit is operational. D1: passed D2: passed D3: passed |  | GeneXpert Module SN: 865209 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560416024168/5581241560416024168_signature_28.png | Rachelle Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560416024168/5581241560416024168_signature_31.png
| 2023-04-24 | 2023-04-12 16:14 | Ilocos Training and Regional Medical Center | San Fernando, La Union | 1 | La Union | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24321 | PBSP | itrmc.tblaboratory@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 21 ASSAY: Xpert HPV, COUNT: 303 ASSAY: Xpert MTB/RIF Ultra, COUNT: 993 ASSAY: Xpert SARS-CoV-2, COUNT: 171 ASSAY: Xpert HIV-1 Viral Load, COUNT: 174 | 2023-04-21 | GXIV | 707524 | Module B4 (611737) for replacement | #NAME? | Replaced B4 | Unit is operational. |  | GeneXpert Module SN: 677556 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560419667702/5581241560419667702_signature_28.png | Jennylen Torres | https://www.jotform.com/uploads/macarewecare/222927978372067/5581241560419667702/5581241560419667702_signature_31.png
| 2023-04-24 | 2023-04-19 15:32 | Ospital ng Palawan | Puerto Prinsesa | 4B | Palawan | Laboratory | YES | REPAIR TROUBLESHOOTING | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24319 | PBSP | Jblim.pbsp@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 370 ASSAY: Xpert SARS-CoV-2, COUNT: 126 ASSAY: Xpert HIV-1 Viral Load, COUNT: 162 |  | GeneXpert | 804282 | Gx DX software wont open due to an error with the SQL server service manager that has stopped | Back up corrupted files on the MSSQL folder then replaced with the same file (.mdf and .ldf) from the "Template Data" folder and Restart SQL server Reinstall GxDX v4.8 32bit software on the laptop | Troubleshoot Generate reporting files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581165059118095441/5581165059118095441_signature_28.png | Jill Diala Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5581165059118095441/5581165059118095441_signature_31.png
| 2023-04-24 | 2023-04-24 09:36 | Mangatarem Rural Health Unit | Osmeña st. Mangatarem Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24318 | DOH | rhumangatarem.lab@gmail.com | 32 | ASSAY: Xpert MTB/RIF, COUNT: 132 ASSAY: Xpert MTB/RIF Ultra, COUNT: 530 | 2023-02-02 | GXIV-D | 838284 | Due for Calibration | Xpertcheck Preventive maintenance | #NAME? | Unit is operational except B3 | All modukes passed calibration | 3 xpertcheck kits | 9575 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581104179658278715/5581104179658278715_signature_28.png | Keziah T. Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5581104179658278715/5581104179658278715_signature_31.png
| 2023-04-24 | 2023-04-20 09:17 | Jose Dalman Rural Health Unit | Poblacion Jose Dalman, Zamboanga del norte | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24317 | PBSP | ferranbryder@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 345 | 2023-04-20 | GX 4 | 834827 | Module 2 h/w failed | Module 787281 replaced with Module 606719 | Module Replacement and Preventive Maintenence | All 2 modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581064039315859362/5581064039315859362_signature_28.png | Ferrari Bryder Tomoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5581064039315859362/5581064039315859362_signature_31.png
| 2023-04-24 | 2023-04-20 11:50 | Gumaca District Hospital | Gumaca | 4A | Quezon | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24316 | PBSP | gdhpathology@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 398 |  | GX-IV-D | 804684 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod, filter. | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063520414981650/5581063520414981650_signature_28.png | Josephine Nada | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063520414981650/5581063520414981650_signature_31.png
| 2023-04-24 | 2023-04-20 10:20 | Rakk Prophet Medical Center | Gumaca | 4A | Quezon | Laboratory | NO | TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 24315 | PBSP | rpmcimoleclab20@gmail.com | 27 |  |  | GX-IV-D | 808268 | CPU unable to power on. | Checked CPU connections | Checked CPU internal connections Connect the GeneXpert unit to network switch | Unit is operational. Currently connected to network switch. CPU not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063520412682122/5581063520412682122_signature_28.png | Katrina Rose F. Aureada | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063520412682122/5581063520412682122_signature_31.png
| 2023-04-24 | 2023-04-20 10:12 | Rakk Prophet Medical Center Inc. | Gumaca | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24315 | PBSP | rpmcimoleclab20@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 179 ASSAY: Xpert SARS-CoV-2, COUNT: 161 |  | GX-IV-D | 834625 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. Module B1/784829 service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063520412739248/5581063520412739248_signature_28.png | Katrina Rose F. Aureada | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063520412739248/5581063520412739248_signature_31.png
| 2023-04-24 | 2023-04-17 13:41 | Candelaria Rural Health Unit | Candelaria Municipal hall Compound | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24314 | PBSP | shirleyzrubz@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1478 |  | GX-IV-D | 814155 | For XpertCheck | Performed XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. | All modules passed. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370411752422/5581063370411752422_signature_28.png | Shirley R. Dollente | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370411752422/5581063370411752422_signature_31.png
| 2023-04-24 | 2023-04-18 11:46 | Labo District Hospital | Labo | 5 | Camarines Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24315 | DOH | analyn.consuelo@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 135 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1139 | 2023-04-12 | GX-IV-D | 838207 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. | Module B2/807309 service required | XpertCheck kit | DR# 9882 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370413437095/5581063370413437095_signature_28.png | Analyn C. Consuelo | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370413437095/5581063370413437095_signature_31.png
| 2023-04-24 | 2023-04-19 08:01 | Lopez Rural Health Unit | Lopez | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24313 | DOH | sgdegala@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 920 | 2023-04-12 | GX-IV-D | 831889 | Module C3 and C4 for module replacement. | Replaced defective module | Performed maintenance Replaced defective modules Generated reports | Unit is operational. |  | Old: C3/666758, C4/690591 New: C3/688921, C4/862408 | DR# 9884 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370415456759/5581063370415456759_signature_28.png | Shekinah Glare Degala | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370415456759/5581063370415456759_signature_31.png
| 2023-04-24 | 2023-04-18 13:00 | Vinzons Rural Health Unit | Vinzons | 5 | Camarines Norte | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24313 | DOH | bardonamabel@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 793 | 2023-04-12 | GX-IV-D | 838576 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. | All modules passed. | XpertCheck kit | DR# 9880 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370417779360/5581063370417779360_signature_28.png | Amabel A. Bardon | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370417779360/5581063370417779360_signature_31.png
| 2023-04-24 | 2023-04-20 10:06 | Rakk Prophet Medical Center Inc. | Gumaca | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24314 | PBSP | rpmcimoleclab20@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 ASSAY: Xpert SARS-CoV-2, COUNT: 601 |  | GX-IV-D | 838871 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. Module B3 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370418735867/5581063370418735867_signature_28.png | Katrina Rose F. Aureada | https://www.jotform.com/uploads/macarewecare/222927978372067/5581063370418735867/5581063370418735867_signature_31.png
| 2023-04-24 | 2023-04-20 14:54 | National Kidney Transplant Institute | East Ave., Quezon City | NCR | NCR 2nd - Quezon City | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24312 | COMMERCIAL | molecularlabnkti@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 346 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 162 ASSAY: Others, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1252 ASSAY: Xpert MTB/RIF Ultra, COUNT: 105 ASSAY: Xpert SARS-CoV-2, COUNT: 19272 ASSAY: Xpert HCV VL, COUNT: 19 ASSAY: Others, COUNT: 9 | 2023-04-18 | GX XVI | 110005587 | Modules A2(210028754), D1(210028990), D4(210028994) for replacement | #NAME? | Replaced modules | Unit is operational. |  | GeneXpert Module SN: 633806, 674723, 677552 | 9907, 9906 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581044880415409550/5581044880415409550_signature_28.png | Odessa Barreto | https://www.jotform.com/uploads/macarewecare/222927978372067/5581044880415409550/5581044880415409550_signature_31.png
| 2023-04-24 | 2023-04-20 10:23 | Alabat District Hospital | National Rd., Alabat | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24311 | DOH | letpiquero@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 583 | 2023-04-12 | GXIV-D | 831614 | Module 2 Failed in XpertCheck | Module Replacement | Performed Maintenance Replaced Module 2 Generate Reports | All Modules are Operational | C2: 754442 -> 716709 | GeneXpert Module 6-Colors | 9888 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581023040411417475/5581023040411417475_signature_28.png | Carmelita Piquero | https://www.jotform.com/uploads/macarewecare/222927978372067/5581023040411417475/5581023040411417475_signature_31.png
| 2023-04-24 | 2023-04-19 14:14 | San Francisco Rural Health Unit | Brgy. Poblacion, San Francisco | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24310 | DOH | micovelarde13@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 340 | 2023-04-12 | GXIV-D | 831698 | 2 Modulea Failed in XpertCheck | Module Replacement | Replaced Modules 1 and 4 Performed Maintenance Generate Reports | All Modules are Operational | 757169, 757209 -> 643530, 635021 | GeneXpert Module 6-Colors | 9886 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581023040419524119/5581023040419524119_signature_28.png | Michael Francis Velarde | https://www.jotform.com/uploads/macarewecare/222927978372067/5581023040419524119/5581023040419524119_signature_31.png
| 2023-04-24 | 2023-04-17 17:09 | Quezon Medical Center | Lucena | 4A | Quezon | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24309 | PBSP | joannpalermo13@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 438 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2485 ASSAY: Xpert HIV-1 Viral Load, COUNT: 109 | 2023-04-17 | GX-IV-D | 807462 | Module D1 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | Unit is operational. | All Modules Passed | Old: D1/766357 New: D1/210003676 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170414272388/5581021170414272388_signature_28.png | Jo-Ann Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170414272388/5581021170414272388_signature_31.png
| 2023-04-24 | 2023-04-18 17:14 | Tagkawayan Rural Health Unit | Brgy., Poblacion, Tagkawayan | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE ReXpertCheck | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24308 | DOH | rvteopy@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 376 | 2023-04-12 | GXIV-D | 831916 | Module 4 Has Temperature Errors Modules 2 and 3 with Temperature Errors | Module Replacement XpertCheck | Replaced Module 4 Performed Maintenance XpertCheck Modules 2 and 3 Generate Reports | 3 Modules are Operational, 1 Defective | XpertCheck: B2-Passed, B3-Failed | GeneXpert Module 6-Colors | 9885 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170415523839/5581021170415523839_signature_28.png | Rodelo v. Teopy | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170415523839/5581021170415523839_signature_31.png
| 2023-04-24 | 2023-04-17 16:58 | Quezon Medical Center | Quezon Ave., Lucena | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24309 | PBSP | joannpalermo13@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2394 ASSAY: Xpert HPV, COUNT: 403 ASSAY: Xpert HIV-1 Viral Load, COUNT: 81 |  | GXIV-D | 804692 | Module 1 has Temperature Error | Module Replacement | Performed Maintenamce Replaced Module 1 Generate Reports | All Modules are Operational | 638879->621976 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170415532589/5581021170415532589_signature_28.png | JO-ANN YAMBALLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170415532589/5581021170415532589_signature_31.png
| 2023-04-24 | 2023-04-17 15:08 | Sariaya Rural Health Unit | Gomez St., Sariaya | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24308 | DOH | rhu_sariaya@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1894 | 2023-04-12 | GXIV-D | 831890 | Module 4 Failed in XpertCheck | Module Replacement | Replaced Module 4 Performed Maintenance Generate Reports | All Modules are Operational | A4: 733097 -> 757623 | GeneXpert Module 6-Colors | 9887 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170417759165/5581021170417759165_signature_28.png | Nerissa asamli | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170417759165/5581021170417759165_signature_31.png
| 2023-04-24 | 2023-04-19 10:35 | San Narciso Rural Health Unit | San Jose Street, San Narciso | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24309 | PBSP | morena_veluz@yahoo.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 522 |  | GXIV-D | 838163 | Module 3 Failed in XpertCheck | Module Replacement | Replaced Module 3 Performed Maintenance Generate Reports | All Modules are Operational. Keyboard is Not Fully Functional | B3:806998 -> 655327 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170418399870/5581021170418399870_signature_28.png | MA. MORENA V. DIMACULANGAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5581021170418399870/5581021170418399870_signature_31.png
| 2023-04-22 | 2023-04-19 13:20 | Victoria Rural Health Unit | Poblacion Uno., Victoria | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24306 | PBSP | lractrebor24@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 675 | 2023-03-31 | GXIV- 2 placer | 835178 | Unit for PM. For module replacement of module B2. | Performed PM. Performed module replacement on B2. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B2 (657364) im replacement to defective B2 (788665). | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5579442735426555533/5579442735426555533_signature_28.png | CARL ROBERT RACELIS | https://www.jotform.com/uploads/macarewecare/222927978372067/5579442735426555533/5579442735426555533_signature_31.png
| 2023-04-22 | 2023-04-19 08:09 | Bongabong Municipal Health Office | Brgy. Aplaya, Lapu Lapu St., Bongabong | 4B | Occidental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24307 | PBSP | cortezmariel30@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 282 | 2023-03-31 | GXIV | 834687 | Unit for PM. For module replacement of modules A1 and A4. | Performed PM. Performed module replacement on A1 and A4. Generated reports. | PM and module replacement, | Unit is operational. |  | Gx module A1 (697878) and A4 (745368) in replacement to defective A1 (786829) and A4 (786837). | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5579442735428936565/5579442735428936565_signature_28.png | Mariel A. Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/5579442735428936565/5579442735428936565_signature_31.png
| 2023-04-22 | 2023-04-18 12:38 | Bulalacao Municipal Health Office | Poblacion, Bulalacao | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24307 | PBSP | queence.macaraig@gmail.com | 25 | ASSAY: Others, COUNT: 0 | 2023-03-31 | GXIV | 835185 | For module replacement of module B2. Unit for PM. | Performed module replacement on B2. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B2 (753679) im replacement to defective B2 (788911). | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5579442735429736382/5579442735429736382_signature_28.png | Marilyn C.Contreras | https://www.jotform.com/uploads/macarewecare/222927978372067/5579442735429736382/5579442735429736382_signature_31.png
| 2023-04-20 | 2023-04-20 12:56 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24305 | PBSP | cdh.laboratorydept@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 124 ASSAY: Xpert MTB/RIF Ultra, COUNT: 959 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GXIV-D | 807562 | B1:789715 for replacement | Module replacement Preventive maintenance | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational except B2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577899499412502946/5577899499412502946_signature_28.png | Kristina Cassandra Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5577899499412502946/5577899499412502946_signature_31.png
| 2023-04-20 | 2023-04-19 12:02 | Senator Ninoy Aquino Rural Health Unit | Kulaman, Senator Ninoy Aquino | 12 | Sultan Kudarat | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24304 | PBSP | benjieramos1975@gmail.com.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 97 ASSAY: Xpert MTB/RIF Ultra, COUNT: 371 | 2023-04-19 | Gx-IV | 838129 | Module 4 for Service Required | Perform Genexpert machine Preventive maintenance Replace module 4 SN:806700 with new module SN: 632395 | Perform Genexpert machine Preventive maintenance Replace module 4 SN:806700 with new module SN: 632395 | 4 operational modules |  | Genexpert module 6 color SN: 632395 | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577876349324556514/5577876349324556514_signature_28.png | Biengie Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5577876349324556514/5577876349324556514_signature_31.png
| 2023-04-20 | 2023-04-20 14:54 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24303 | PBSP | laureanjo@yahoo.com | 25 |  |  | GXIV | 808620 | Due date for xpertcheck | Xpertcheck all modules | Xpertcheck all modules | Module 2 service require, 1,3,4passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577832221226034581/5577832221226034581_signature_28.png | Laurean jo cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5577832221226034581/5577832221226034581_signature_31.png
| 2023-04-20 | 2023-04-20 11:55 | Banlat Health Center | Banlat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24302 | PBSP | lorinemendiola@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 96 ASSAY: Xpert MTB/RIF, COUNT: 176 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1383 | 2023-04-19 | GXIV | 804688 | Module A3 (624630) for replacement | #NAME? | Replaced A3 | Unit is operational. |  | GeneXpert Module SN: 664557 | G 5469 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577828396927646880/5577828396927646880_signature_28.png | Lorine Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5577828396927646880/5577828396927646880_signature_31.png
| 2023-04-20 | 2023-04-20 13:50 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24301 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 203 ASSAY: Xpert MTB/RIF, COUNT: 891 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1403 |  | GXIV-4D | 820206 | None | None | Preventive Maintenance | All four modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577795227511667329/5577795227511667329_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5577795227511667329/5577795227511667329_signature_31.png
| 2023-04-20 | 2023-04-20 13:48 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24300 | DOH | vsmmcdeptpatho.microbioloy@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 195 ASSAY: Xpert MTB/RIF, COUNT: 1015 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1334 |  | GXIV-4D | 838285 | None | None | #NAME? | All four modules are functional (w/ service module A3) | Original module A3 for XPC verification |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577795227514247272/5577795227514247272_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5577795227514247272/5577795227514247272_signature_31.png
| 2023-04-20 | 2023-04-20 13:50 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24299 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 203 ASSAY: Xpert MTB/RIF, COUNT: 891 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1403 |  | GXIV-4D | 820206 | None | None | Preventive Maintenance | All four modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577794827512991415/5577794827512991415_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5577794827512991415/5577794827512991415_signature_31.png
| 2023-04-20 | 2023-04-20 13:48 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24298 | DOH | vsmmcdeptpatho.microbioloy@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 195 ASSAY: Xpert MTB/RIF, COUNT: 1015 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1334 |  | GXIV-4D | 838285 | None | None | #NAME? | All four modules are functional (w/ service module A3) | Original module A3 for XPC verification |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577794827515316532/5577794827515316532_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5577794827515316532/5577794827515316532_signature_31.png
| 2023-04-20 | 2023-04-20 13:32 | Daram Rural Health Unit | Daram Samar | 8 | Samar | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24297 | PBSP | riobarras27@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 164 ASSAY: Xpert MTB/RIF Ultra, COUNT: 302 | 2023-04-20 | GXIV | 838868 | Module B1 for replacement | - Module Replacement B1/861539 into 722433 - Performed functionality test - Performed preventive maintenance - Generate Reports | #NAME? | Unit is Operational (Room air condition not working) |  | GeneXpert Module 6 Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577785806379157786/5577785806379157786_signature_28.png | Rioreta Barras | https://www.jotform.com/uploads/macarewecare/222927978372067/5577785806379157786/5577785806379157786_signature_31.png
| 2023-04-20 | 2023-04-20 12:38 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24296 | DOH | dyrosemarie71@gmail.com | 28.7 |  |  | GX-IV 4D | 838401 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Fan Filters, Plunger Rod, and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577761531418337900/5577761531418337900_signature_28.png | Rosemarie B. Dy | https://www.jotform.com/uploads/macarewecare/222927978372067/5577761531418337900/5577761531418337900_signature_31.png
| 2023-04-20 | 2023-04-20 12:46 | Ace Medical Center Bohol | Mansasa, Tagbilaran Bohol | 7 | Cebu | Tb Laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 24295 | CONSORTIUM | laboratory@acemcbohol.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 80 |  | Gx 4 | 840969 | Due for Calibration | Performed calibration | Clean Filter, clean ICore and plunger Get IQR, Gxx files, sys log and assay Stat | All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577755021304567227/5577755021304567227_signature_28.png | Dianne Dawn Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5577755021304567227/5577755021304567227_signature_31.png
| 2023-04-20 | 2023-04-20 11:55 | Banlat Health Center | Banlat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24294 | PBSP | lorinemendiola@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 96 ASSAY: Xpert MTB/RIF, COUNT: 176 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1383 | 2023-04-19 | GXIV | 804688 | Module A3 (624630) for replacement | #NAME? | Replaced A3 | Unit is operational. |  | GeneXpert Module SN: 664557 | G 5469 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577727427028045522/5577727427028045522_signature_28.png | Lorine Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5577727427028045522/5577727427028045522_signature_31.png
| 2023-04-20 | 2023-04-20 09:53 | Tacurong City Health Office | Tacurong City | 12 | Sultan Kudarat | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24293 | PBSP | mjgppalabrica@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1837 |  | Gx-IV | 835396 | Module 2 and 3 H/w failed due to multiple temperature errors | Perform Genexpert machine Preventive maintenance Report Genexpert module problems to jotform reporting | Perform Genexpert machine Preventive maintenance Report Genexpert module problems to jotform reporting | 2 operational modules | Module Problem reported to Jotform |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577654877421079278/5577654877421079278_signature_28.png | Mary June Grace Palabrica-Villacruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5577654877421079278/5577654877421079278_signature_31.png
| 2023-04-20 | 2023-04-18 09:30 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24292 | DOH | jan19.cam@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 473 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1200 |  | Gx-IV | 838504 | Module 4 H/w failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 4 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 4 for troubleshooting | 4 operational modules | Module 4 Passed for Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577654877422802206/5577654877422802206_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5577654877422802206/5577654877422802206_signature_31.png
| 2023-04-20 | 2023-04-18 09:30 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24290 | DOH | jan19.cam@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 473 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1200 |  | Gx-IV | 838504 | Module 4 H/w failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 4 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 4 for troubleshooting | 4 operational modules | Module 4 Passed for Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577654067423646959/5577654067423646959_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5577654067423646959/5577654067423646959_signature_31.png
| 2023-04-20 | 2023-04-20 09:53 | Tacurong City Health Office | Tacurong City | 12 | Sultan Kudarat | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24291 | PBSP | mjgppalabrica@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 8 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1837 |  | Gx-IV | 835396 | Module 2 and 3 H/w failed due to multiple temperature errors | Perform Genexpert machine Preventive maintenance Report Genexpert module problems to jotform reporting | Perform Genexpert machine Preventive maintenance Report Genexpert module problems to jotform reporting | 2 operational modules | Module Problem reported to Jotform |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577654067427092528/5577654067427092528_signature_28.png | Mary June Grace Palabrica-Villacruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5577654067427092528/5577654067427092528_signature_31.png
| 2023-04-19 | 2023-04-19 15:40 | Cebu TB Reference Lab | Cebu City | 7 | Cebu | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24289 | PBSP | titusianno@gmail.com | 26 |  |  | GXIV | 838882 | Due for xpercheck calibration module 2 and 4 | Xpertcheck calibration for module 2&4 | Xpertcheck calibration for module 2&4 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5577029350432842816/5577029350432842816_signature_28.png | TITUS IAN VALERIO | https://www.jotform.com/uploads/macarewecare/222927978372067/5577029350432842816/5577029350432842816_signature_31.png
| 2023-04-19 | 2023-04-19 14:33 | Pangasinan Provincial Hospital | Bolingit San Carlos City, Pangasinan | 1 | Pangasinan | Tb dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24288 | PBSP | judzfm@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 433 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2676 | 2023-01-09 | GXIV-D | 808283 | Due for calibration | Xpertcheck calibration Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except | For reporting | 5 xpertcheck kits | 3800 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576979302214002976/5576979302214002976_signature_28.png | Judy Ann F. Mabunga | https://www.jotform.com/uploads/macarewecare/222927978372067/5576979302214002976/5576979302214002976_signature_31.png
| 2023-04-19 | 2023-04-19 15:00 | Philippine Red Cross Tower | Boni. Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Yasser Salih |  |  |  |  | 24287 | COMMERCIAL | bloodcollection.tower@gmail.com | 25 |  |  | MCS + | 08B049 | Unit for maintenance | #NAME? | Perform maintenance | Unitis operational | With Ronnel Belizar and Gian Paolo Magnaye |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576978865423875465/5576978865423875465_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/5576978865423875465/5576978865423875465_signature_31.png
| 2023-04-19 | 2023-04-19 15:01 | Philippine Red Cross Tower | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Gian Paolo Magnaye |  |  |  |  | 24286 | COMMERCIAL | bloodcollection.tower@gmail.com | 25 |  |  | MCS+ | 01J038 | Unit for Maintenance | #NAME? | Performed maintenance | Operational | With Ronnel Belizar and Yasser Salih |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576978770917436781/5576978770917436781_signature_28.png | Marnelli Mari Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/5576978770917436781/5576978770917436781_signature_31.png
| 2023-04-19 | 2023-04-19 14:26 | Research Institute for Tropical Medicine | 9002 Research Drive, Filinvest Corporate City, Alabang Muntinlupa City, Metro Manila | NCR | NCR 4th - Muntinlupa | Laboratory | NO | INSTALLATION | CEPHEID | Harold Toco |  |  |  |  | 24285 | MACARE-HBDC | arglaboratory@gmail.com | 25 |  | 2023-04-19 | GX4 10-C | 110017870 | For unit installation | Performed unit installation | Performed Unit installation | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576953481617985429/5576953481617985429_signature_28.png | Marianne Co | https://www.jotform.com/uploads/macarewecare/222927978372067/5576953481617985429/5576953481617985429_signature_31.png
| 2023-04-19 | 2023-04-19 14:26 | Research Institute for Tropical Medicine | 9002 Research Drive, Filinvest Corporate City, Alabang Muntinlupa City, Metro Manila | NCR | NCR 4th - Muntinlupa | Laboratory | NO | INSTALLATION | CEPHEID | Harold Toco |  |  |  |  | 24284 | MACARE-HBDC | arglaboratory@gmail.com | 25 |  | 2023-04-19 | GX4 10-C | 110017870 | For unit installation | Performed unit installation | Performed Unit installation | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576953211618280549/5576953211618280549_signature_28.png | Marianne Co | https://www.jotform.com/uploads/macarewecare/222927978372067/5576953211618280549/5576953211618280549_signature_31.png
| 2023-04-19 | 2023-04-19 13:56 | Research Institute for Tropical Medicine | 9002 Research Drive, Filinvest Corporate City, Alabang, Muntinlupa City, Metro Manila | NCR | NCR 4th - Muntinlupa | Laboratory | NO | INSTALLATION | CEPHEID | Harold Toco |  |  |  |  | 24283 | MACARE-HBDC | arglaboratory@gmail.com | 25 |  | 2023-04-19 | GX4-10C | 110017869 | Unit Installation | Performed Unit Installation | Performed Unit installation | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576951761619101958/5576951761619101958_signature_28.png | Marianne Co | https://www.jotform.com/uploads/macarewecare/222927978372067/5576951761619101958/5576951761619101958_signature_31.png
| 2023-04-19 | 2023-04-19 13:01 | Talavera General Hospital | Poblacion, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24282 | MACARE-HBDC | talaverahope@gmail.com | 28 | ASSAY: Xpert HBV VL, COUNT: 7 ASSAY: Xpert MTB/RIF Ultra, COUNT: 246 ASSAY: Xpert HIV-1 Viral Load, COUNT: 65 | 2023-04-19 | Genexpert IV | 110013773 | Error 2008 | For Monitoring | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576926004211940639/5576926004211940639_signature_28.png | Fitz Manuzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5576926004211940639/5576926004211940639_signature_31.png
| 2023-04-19 | 2023-04-19 13:06 | Tomas Claudio Health Center (District A) | Corner E. Quirino, Camus Street, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24281 | DOH | arleen_p_alconera@yahoo.com | 28.7 |  |  | GX-IV 4D | 838492 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Fan Filters, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Unit is operational | Module B1 encounter temperature, data will be sent to Cepheid for further assessment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576906973427751860/5576906973427751860_signature_28.png | Arleen Alconera | https://www.jotform.com/uploads/macarewecare/222927978372067/5576906973427751860/5576906973427751860_signature_31.png
| 2023-04-19 | 2023-04-19 13:05 | San Carlos City Health Office | Palaris st. san Carlos City Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24280 | DOH | gbmmbg@yahoo.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 112 ASSAY: Xpert MTB/RIF Ultra, COUNT: 365 | 2023-02-03 | GXIV-D | 838437 | Due for Calibration | Calibration Preventive maintenance | #NAME? | Unit is operational except B1 and B2 |  | 3 xpertcheck kits | 9589 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576904050423610947/5576904050423610947_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5576904050423610947/5576904050423610947_signature_31.png
| 2023-04-19 | 2023-04-18 13:05 | Reproductive Health and Wellness Center | Jacinto Street, Poblacion Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24279 | DOH | rhwclab@gmail.com | 23.7 |  |  | GX-IV 4D | 838262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576899633423846577/5576899633423846577_signature_28.png | Valencio B. Macatutac Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5576899633423846577/5576899633423846577_signature_31.png
| 2023-04-19 | 2023-04-18 13:07 | Reproductive Health and Wellness Center | Jacinto Street, Poblacion Davao City | 11 | Davao del Sur | TB HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24278 | MACARE-HBDC | rhwclab@gmail.com | 23.8 |  |  | GX-IV 4D | 839236 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Only Three Modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576899633423746347/5576899633423746347_signature_28.png | Valencio B. Macatutac Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5576899633423746347/5576899633423746347_signature_31.png
| 2023-04-19 | 2023-04-19 11:56 | Magsaysay Rural Health Unit | Brgy. Poblacion, Magsaysay | 10 | Misamis Oriental | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24277 | DOH | jkakilgan12@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 131 |  | GX-4 | 838556 | Module 1 and 4 for error checking | None | XpertCheck Calibration of all Modules Quarterly Preventive Maintenance Data Collection and Site Visit | 3 modules passed in XPC, 1 module for replacement (Service Required) | None |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576889123516176134/5576889123516176134_signature_28.png | Jesusa N. Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/5576889123516176134/5576889123516176134_signature_31.png
| 2023-04-19 | 2023-04-19 12:25 | DR. VICTORIA R. POTENCIANO MEDICAL CENTER | 163 EDSA Mandaluyong City | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Yasser Salih |  |  |  |  | 24276 | CONSORTIUM | lab@vrp.com.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 239 ASSAY: Xpert MTB/RIF Ultra, COUNT: 205 | 2023-04-18 | GX 4 | 835600 | Modules B2 (696025) and B3 (791611) for replacement | #NAME? | Module replacement of B2 and B3 | Unit is operational | With Ronnel Belizar and Gian Paolo Magnaye | Genexpert modules SN 607775, SN 210057342 | 9908 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576881835424096292/5576881835424096292_signature_28.png | Charissa Torres | https://www.jotform.com/uploads/macarewecare/222927978372067/5576881835424096292/5576881835424096292_signature_31.png
| 2023-04-19 | 2023-04-18 10:13 | Bagong Silang Ph. 1 Health Center | Bagong Silang, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24275 | PBSP | alcaideirene07@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1583 | 2023-04-18 | GXIV | 838856 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. B1: failed B2: passed B3: passed B4: passed |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576880404519121115/5576880404519121115_signature_28.png | Airene Alcaide | https://www.jotform.com/uploads/macarewecare/222927978372067/5576880404519121115/5576880404519121115_signature_31.png
| 2023-04-19 | 2023-04-19 09:21 | Ospital ng Palawan | Puerto Prinsesa | 4B | Palawan | Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24274 | PBSP | jblim.pbsp@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 241 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2642 ASSAY: Xpert SARS-CoV-2, COUNT: 2695 ASSAY: Xpert HIV-1 Viral Load, COUNT: 66 |  | GeneXpert | 842893 | Defective module A1/693393 A4/885696 B2/888430 | Replaced module New module: A1/ 210045619 A4/210046004 B2/210045893 | Module Replacement Preventive Maintenance Generate reporting files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576774381509862161/5576774381509862161_signature_28.png | Jill Diala Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5576774381509862161/5576774381509862161_signature_31.png
| 2023-04-19 | 2023-04-19 09:21 | Balamban Rural Health Unit | Poblacion, Balamban Cebu | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  |  |  |  | 24273 | DOH | climaco.mildred@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 316 | 2023-04-19 | Gx4 | 838231 | B1 Module with serial number 806851 Hardware Fail. | Perform Xpertcheck | Perform XpertCheck Module B1 | Module B1 Pass for Calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576773185415139197/5576773185415139197_signature_28.png | Mildred Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5576773185415139197/5576773185415139197_signature_31.png
| 2023-04-18 | 2023-04-14 13:05 | Mexico Rural Health Unit I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24270 | DOH | jasperdizon20@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 115 ASSAY: Xpert BCR-ABL Ultra, COUNT: 348 | 2023-04-14 | Genexpert IV | 831877 | Module A4 error | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576190472026456047/5576190472026456047_signature_28.png | Von Romeo Pacheco | https://www.jotform.com/uploads/macarewecare/222927978372067/5576190472026456047/5576190472026456047_signature_31.png
| 2023-04-18 | 2023-04-14 11:50 | Jose B. Lingad Memorial General Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24272 | COMMERCIAL | jblmgh.is@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 79 ASSAY: Xpert HCV VL, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1169 | 2023-04-14 | Genexpert IV | 812571 | Genexpert not seen on Application UPS not working | Troubleshoot the Etherenet Cable Replaced the Ethernet Cable Troubleshoot the UPS | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | UPS Battery for Replacement | Ethernet Cable | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576190482023051231/5576190482023051231_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5576190482023051231/5576190482023051231_signature_31.png
| 2023-04-18 | 2023-04-18 15:07 | Sto Domingo District Hospital | Pulung Bulig, Sto Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24271 | DOH | lhaydeesel@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-04-18 | Genexpert IV | 831582 | Module 3 error 1002 | Module 3 for monitoring | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576190482026087170/5576190482026087170_signature_28.png | Ma. Czarina Anne F. Panahon | https://www.jotform.com/uploads/macarewecare/222927978372067/5576190482026087170/5576190482026087170_signature_31.png
| 2023-04-18 | 2023-04-18 15:04 | LAPU LAPU CITY HEALTH OFFICE | LAPULAPU CITY | 7 | Cebu | LABORATORY | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  |  |  |  | 24269 | DOH | tanya_pax@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 824 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1610 |  | GX-IV | 837661 | MODULE A2 & A3 FOR XPERTCHECK | COLLECT IQR AND XPERTCHECK DATA | -PERFORM CALIBRATION: A2 SN 806133 - RETRY THEN FOR SERVICE A3 SN 664865 - FOR SERVICE | EQUIPMENT OPERATIONAL, MODULE A2 AND A3 DISABLED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576109529427691741/5576109529427691741_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5576109529427691741/5576109529427691741_signature_31.png
| 2023-04-18 | 2023-04-18 13:12 | Olongapo City Health Office | Rizal St. West Bajac-bajac Olongapo City, Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24268 | PBSP | roselynmaganomilan@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 570 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1170 ASSAY: Xpert HIV-1 Viral Load, COUNT: 112 | 2023-03-31 | GXIV-D | 834807 | C1:787163 for Module replacement | Module replacement Preventive maintenance | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 39996 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576048249246831511/5576048249246831511_signature_28.png | Roselyn M. Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5576048249246831511/5576048249246831511_signature_31.png
| 2023-04-18 | 2023-04-18 12:08 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 24267 | DOH | suzzettekho34@gmail.com | 26.12 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1537 | 2023-03-28 | GX-4 | 838554 | None | None | Replacing Module B2 with SN: 862237 to SN: 609246 | All modules passed and functional | None | GeneXpert Module 6-Color | 9848 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5576006463911496481/5576006463911496481_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/5576006463911496481/5576006463911496481_signature_31.png
| 2023-04-18 | 2023-04-18 10:06 | Velez General Hospital | Ramos St,. Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 24266 | CONSORTIUM | emaglaoy@velezhospital.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 0 | 2023-04-18 | GX2 | 110017599 | Install the machines | Install genexpert machines, cpu and monitor, bar code scanner | Installation of genexpert machine, Cpu and monitor bar code scanner | Machines are operational | Install the machines with Engr Raymond Laurito | Cpu Serial numberCPH0822F40946 | DR 9871 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575974177223619166/5575974177223619166_signature_28.png | Elaine Maglaoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5575974177223619166/5575974177223619166_signature_31.png
| 2023-04-18 | 2023-04-18 10:13 | Bagong Silang Ph. 1 Health Center | Bagong Silang, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24265 | PBSP | alcaideirene07@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1583 | 2023-04-18 | GXIV | 838856 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. B1: failed B2: passed B3: passed B4: passed |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575956837841834254/5575956837841834254_signature_28.png | Airene Alcaide | https://www.jotform.com/uploads/macarewecare/222927978372067/5575956837841834254/5575956837841834254_signature_31.png
| 2023-04-18 | 2023-04-18 10:00 | Leyte Provincial Hospital | Pawing, Palo Leyte | 8 | Leyte | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  |  |  |  | 24264 | PBSP | kareenchavero@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 402 ASSAY: Xpert MTB/RIF Ultra, COUNT: 251 | 2023-04-18 | GXIV | 815267 | - CPU error Booting SN: 2AU7021M0Z | #NAME? | #NAME? | Unit is Operational |  | 3V lithium battery (CR2032) | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575935803228920113/5575935803228920113_signature_28.png | Kareen T. Chavero | https://www.jotform.com/uploads/macarewecare/222927978372067/5575935803228920113/5575935803228920113_signature_31.png
| 2023-04-18 | 2023-04-18 10:06 | Velez General Hospital | Ramos St,. Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 24263 | CONSORTIUM | emaglaoy@velezhospital.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 0 | 2023-04-18 | GX2 | 110017599 | Install the machines | Install genexpert machines, cpu and monitor, bar code scanner | Installation of genexpert machine, Cpu and monitor bar code scanner | Machines are operational | Install the machines with Engr Raymond Laurito | Cpu Serial numberCPH0822F40946 | DR 9871 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575934495125111080/5575934495125111080_signature_28.png | Elaine Maglaoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5575934495125111080/5575934495125111080_signature_31.png
| 2023-04-18 | 2023-04-18 09:30 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24262 | DOH | jan19.cam@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 473 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1200 |  | Gx-IV | 838504 | Module 4 H/w failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 4 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 4 for troubleshooting | 4 operational modules | Module 4 Passed for Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575926736719374922/5575926736719374922_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5575926736719374922/5575926736719374922_signature_31.png
| 2023-04-18 | 2023-04-18 09:30 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24261 | DOH | jan19.cam@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 473 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1200 |  | Gx-IV | 838504 | Module 4 H/w failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 4 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 4 for troubleshooting | 4 operational modules | Module 4 Passed for Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575926476717403179/5575926476717403179_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5575926476717403179/5575926476717403179_signature_31.png
| 2023-04-18 | 2023-04-18 09:50 | Isabela City Health Office | East Terminal sunrise isabela city basilan | 9 | Basilan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24260 | PBSP | reynantemontano@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 768 | 2023-04-18 | GX 4 | 807595 | Module 1 h/w failed | Module 676313 replaced with Module 905109 | Module Replacement | All modules are available, UPS must be replaced immediately. | UPS must be replaced immediately. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575923364421039435/5575923364421039435_signature_28.png | Reynante L. Montaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5575923364421039435/5575923364421039435_signature_31.png
| 2023-04-18 | 2023-04-18 08:58 | Negros Oriental Provincial Hospital RTDL | Dumaguete City | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 24259 | DOH | rezelmarielagos@gmail.com | 23 |  | 2023-04-14 | Gx4 | 838525 | Module B2,B4 | Replace new module. | Replace new module with serial number B2-688622, B4-690680. | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575892417227664058/5575892417227664058_signature_28.png | Rezelmarie Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5575892417227664058/5575892417227664058_signature_31.png
| 2023-04-18 | 2023-04-18 08:58 | Oriental Mindoro Southern District Hospital | Brgy. Odiong, Roxas | 4A | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24258 | PBSP | markedisonmartinez@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1960 | 2023-03-31 | GXIV | 807590 | For module replacement of module B1. | Performed module replacement on B1. Generated reports. | Module replacement. | Unit is operational. |  | Gx module B1 (614394) in replacement to defective B1(708246) | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575891718117256420/5575891718117256420_signature_28.png | Mark Edison G. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5575891718117256420/5575891718117256420_signature_31.png
| 2023-04-17 | 2023-04-17 15:34 | Malapatan Rural Health Unit | Malapatan, Sarangani | 12 | Sarangani | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24257 | DOH | rhu.malapatan12@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 9 ASSAY: Xpert MTB/RIF Ultra, COUNT: 169 | 2023-04-17 | Gx-IV | 838419 | Module 3 for Service Required | Perform Genexpert machine Preventive maintenance Replace module 3 SN: 861214 with new module SN: 912975 | Perform Genexpert machine Preventive maintenance Replace module 3 SN: 861214 with new module SN: 912975 | 4 operational modules |  | Genexpert module 6 color SN: 912975 | DR number: 9847 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575482923166982043/5575482923166982043_signature_28.png | CERES MAE JIMENEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5575482923166982043/5575482923166982043_signature_31.png
| 2023-04-17 | 2023-04-17 15:20 | Macare In-House | Aurora Grand, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | OTHERS | Gian Paolo Magnaye |  |  |  | UPS | 24256 | DOH | carlos.leniel.origenes@gmail.com | 25 |  |  | 1KVA | 20057560904 | Does not back up | N/A | Tested voltage ouput in line mode and back up mode Surface cleaning | For the commissioning | Burnt Board |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575273630416068863/5575273630416068863_signature_28.png | Carlos Leniel Origenes | https://www.jotform.com/uploads/macarewecare/222927978372067/5575273630416068863/5575273630416068863_signature_31.png
| 2023-04-17 | 2023-04-17 15:44 | Macare In-house | Aurora Grand, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | OTHERS | Yasser Salih |  |  |  | UPS | 24255 | DOH | carlos.leniel.origenes@gmail.com | 23 |  | 2023-04-17 | 1 KVA | 20110481208 | Does not backup | Replace Batteries | Tested voltage output in line mode and back up mode Surface cleaning | Unit is operational |  | 2 batteries | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575273600411279080/5575273600411279080_signature_28.png | Macare In-house | https://www.jotform.com/uploads/macarewecare/222927978372067/5575273600411279080/5575273600411279080_signature_31.png
| 2023-04-17 | 2023-04-17 15:44 | Macare In-house | Aurora Grand, Queen City | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | OTHERS | Dan Emil Bodino |  |  |  | UPS | 24254 | DOH | carlos.leniel.origenes@gmail.com | 23 |  | 2023-04-17 | 1KVA | AT1000180709ST0077 | Does not backup | N/A | Tested voltage output in line mode and backup mode,surface cleaning | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575273530412444464/5575273530412444464_signature_28.png | Macare In-house | https://www.jotform.com/uploads/macarewecare/222927978372067/5575273530412444464/5575273530412444464_signature_31.png
| 2023-04-17 | 2023-04-17 14:38 | Macare In-House | Aurora Grand, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Carlos Leniel Origenes |  |  |  |  | 24253 | PBSP | carlos.leniel.origenes@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 0 | 2023-04-14 | GX4 6C | 838869 | Modules not detected | Replaced gateway board | Archive test results Cleaned cartridge bays, icores, plungers, and filter Generate IQR, Assay Stats, and System Logs Perform functionality test | Unit is operational |  | Gateway board SN6140S22040325 | G5450 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575253050416283789/5575253050416283789_signature_28.png | Macare In-house | https://www.jotform.com/uploads/macarewecare/222927978372067/5575253050416283789/5575253050416283789_signature_31.png
| 2023-04-17 | 2023-04-17 13:59 | Sarangani Health Care Facility | Alabel, Sarangani Province | 12 | Sarangani | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24252 | DOH | ronayadao@yahoo.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 73 ASSAY: Xpert MTB/RIF Ultra, COUNT: 861 | 2023-04-17 | Gx-IV | 838542 | Module 3 and 4 for Service Required | Perform Genexpert machine Preventive maintenance Replace module 3 SN:862139 with new Module SN:714753 Replace module 4 SN:862137 with new Module SN:620030 | Perform Genexpert machine Preventive maintenance Replace module 3 SN:862139 with new Module SN:714753 Replace module 4 SN:862137 with new Module SN:620030 | 4 operational modules |  | Genexpert module 6 Color | DR number: 9791 and 9792 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575238319378482077/5575238319378482077_signature_28.png | Ronora Yadao | https://www.jotform.com/uploads/macarewecare/222927978372067/5575238319378482077/5575238319378482077_signature_31.png
| 2023-04-17 | 2023-04-17 13:25 | Quezon City Jail | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | Disassembling of GeneXpert for Pull out | CEPHEID | Ronnel Belizar |  |  |  |  | 24251 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 838881 | Unit for transfer | #NAME? | Disassembled Unit | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575187011524041011/5575187011524041011_signature_28.png | Carlee Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5575187011524041011/5575187011524041011_signature_31.png
| 2023-04-17 | 2023-04-17 12:33 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB dots | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24250 | DOH | marteetingzon20@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 379 ASSAY: Xpert MTB/RIF Ultra, COUNT: 790 | 2023-04-17 | GXIV | 838421 | Module B4 (Error 1002) | #NAME? | #NAME? | Unit is Operational (Module B4 observe for module error) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575158475119157635/5575158475119157635_signature_28.png | Martee H. Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5575158475119157635/5575158475119157635_signature_31.png
| 2023-04-17 | 2023-04-17 10:14 | Buenavista Rural Health Unit | Buenavista, Guimaras | 6 | Guimaras | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24249 | PBSP | clys0910@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 29 | 2023-04-17 | GXIV-4D | 838861 | Module B3 - XpertCheck service required | -Replace defective module B3/863021 with new module SN: 914644 | #NAME? | All modules are operational. |  | 900-0491RS GeneXpert Module 6-Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575157249825899773/5575157249825899773_signature_28.png | Raymund A. Salapio | https://www.jotform.com/uploads/macarewecare/222927978372067/5575157249825899773/5575157249825899773_signature_31.png
| 2023-04-17 | 2023-04-17 11:11 | A.H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | Uninstallation of GeneXpert | CEPHEID | Ronnel Belizar |  |  |  |  | 24248 | PBSP | dinajaen1990@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 432 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1884 ASSAY: Xpert SARS-CoV-2, COUNT: 4728 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GXIV | 839292 | Packing of unit for transfer to be picked up by PBSP | #NAME? | Packed unit | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575106211529998419/5575106211529998419_signature_28.png | Dina Jaen | https://www.jotform.com/uploads/macarewecare/222927978372067/5575106211529998419/5575106211529998419_signature_31.png
| 2023-04-17 | 2023-04-13 14:20 | International Organization for Migration | H.V. Dela Costa | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | John Monsour Doctor |  |  |  |  | 24248 | COMMERCIAL | jsabalboro@iom.int | 25 |  |  | Primostar | 3134004145 | For maintenance | Performed maintenance | Performed maintenance Cleaned eyepiece, objectives and microscope body Check focus | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810015247183/5575035810015247183_signature_28.png | Joy B. Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810015247183/5575035810015247183_signature_31.png
| 2023-04-17 | 2023-04-14 15:15 | Kalinga Health Hub | JP Rizal St. | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24247 | CONSORTIUM | jasmin.catral@innovationsch.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 836 | 2023-03-20 | GX-IV-D | 110008418 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. Module 2 and 4 service required |  | XpertCheck kit | DR# 9718 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810015498206/5575035810015498206_signature_28.png | Jasmin catral | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810015498206/5575035810015498206_signature_31.png
| 2023-04-17 | 2023-04-14 11:33 | Pililla Rural Health Unit | M.A Roxas St., Pililla | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 24246 | DOH | tbdotspililla@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 101 ASSAY: Xpert MTB/RIF Ultra, COUNT: 484 |  | GX-IV-D | 821293 | Module B3/730900 has error 4015 | Run XpertCheck on module B3 | Performed maintenance Run XpertCheck on module B3 Generated reports | Unit is operational. Module B3/730900 failed in XpertCheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810016469519/5575035810016469519_signature_28.png | Karren Abegail R. Olaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810016469519/5575035810016469519_signature_31.png
| 2023-04-17 | 2023-04-13 16:20 | West Rembo Health Center | JP Rizal, West Rembo | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24247 | DOH | danise_er_19@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 247 ASSAY: Xpert MTB/RIF Ultra, COUNT: 910 | 2023-03-27 | GX-IV-D | 831712 | Module B3 for replacement For XpertCheck | Replaced defective module Run XpertCheck | Performed maintenance Replaced defective module Run XpertCheck Generated reports | Unit is operational. Module B4 service required. |  | Old: B3/7457447 New: B3/685568 XpertCheck kot | DR# 9794/9513 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810018289269/5575035810018289269_signature_28.png | Danis Maebel V. Calimag | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810018289269/5575035810018289269_signature_31.png
| 2023-04-17 | 2023-04-14 10:26 | Camp Capinpin Army Station Hospital | Tanay | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24246 | DOH | cathialpha@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 144 | 2023-03-27 | GX-IV-D | 831896 | Modules A1 and A4 for replacement | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. |  | Old: A1/757276, A4/758194 New: A1/730925, A4/860127 | DR# 9809 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810019398625/5575035810019398625_signature_28.png | Anne Catherine Alpha | https://www.jotform.com/uploads/macarewecare/222927978372067/5575035810019398625/5575035810019398625_signature_31.png
| 2023-04-17 | 2023-04-13 10:39 | GMA Rural Health Unit II | Brgy Memije, General Mariano Alvarez | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24242 | PBSP | ldlconti@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1365 |  | GXIV-D | 814152 | Unit is due for XpertCheck | XpertCheck | Performed XpertCheck Performed Maintenance Generate Reports | 3 Modules are Operational, 1 Module Reported | XpertCheck: 2 modules passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575026241616418260/5575026241616418260_signature_28.png | Lorena L. Conti | https://www.jotform.com/uploads/macarewecare/222927978372067/5575026241616418260/5575026241616418260_signature_31.png
| 2023-04-17 | 2023-04-14 15:28 | Research Institute for Tropical Medicine | RITM Compound, Filinvest | NCR | NCR 4th - Muntinlupa | PCAV | NO | TROUBLESHOOTING Inspection | OTHERS | Brenguiler Bagbaguen |  |  |  | RHEA | 24243 | MACARE-HBDC | PerezRoehl1968@gmail.com | 25 |  |  | Compact | C-0321 | Unit has Problem in Screen | Troubleshoot | Reset LCD and Cable Connections Checked Functionality Used Old LCD to Isolate Problem | New LCD has No Touch Response. Backlight Output Only. | Used Old Screen for the Machine to Use Temporarily |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575026241619305474/5575026241619305474_signature_28.png | Roehl eric perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5575026241619305474/5575026241619305474_signature_31.png
| 2023-04-17 | 2023-04-14 14:03 | Asian Hospital and Medical Center | Civic Drive, Filinvest | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | Inspection and Assessment | HAEMONETICS | Brenguiler Bagbaguen |  |  |  |  | 24245 | MACARE-HBDC | agimang@asianhospital.com | 24 |  |  | MCS+ | 1F116 | Unit has Pump Rotor Roller Problem | Cleaning | Checked and Cleaned Pump Rotor Rollers Cleaned Valves Checked Functionality | Pump Rotor Roller for Replacement. For Quotation of Part Replacement. | Pump Rotor Rollers are Worn-out |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575026251612935600/5575026251612935600_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5575026251612935600/5575026251612935600_signature_31.png
| 2023-04-17 | 2023-04-14 13:49 | Asian Hospital and Medical Center | Civic Drive, Filinvest | NCR | NCR 4th - Muntinlupa | Blood Bank | NO | Inspection and Assessment | HAEMONETICS | Brenguiler Bagbaguen |  |  |  |  | 24244 | MACARE-HBDC | agimang@asianhospital.com | 24 |  |  | MCS+ | 10B129 | Unit has Pump Rotor Roller Problem | Cleaning | Checked Pump Rotor Rollers Checked Machine Parameters Cleaned Pumps and Valves Performed Functionality Test | Pump Rotor Rollers for Replacement. For Quotation of Replacement Parts | Pump Rotor Rollers are Worn-out |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575026251614722643/5575026251614722643_signature_28.png | Angel Danica B. Gimang | https://www.jotform.com/uploads/macarewecare/222927978372067/5575026251614722643/5575026251614722643_signature_31.png
| 2023-04-17 | 2023-04-12 12:21 | Panapaan Health Center | Floraville Subd., Panapaan, Bacoor | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24237 | PBSP | cecille_tobias@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2245 ASSAY: Xpert HIV-1 Viral Load, COUNT: 230 | 2023-04-11 | GXIV-D | 807584 | Modules 3 and 4 with Temperature Errors | Module Replacement | Performed Maintenance Replaced 2 Modules Generate Reports | All Modules are Operational | 807917, 657198 -> 876167, 877122 | GeneXpert Module Kit 6- Colors | G 5428 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025011614553822/5575025011614553822_signature_28.png | Maria Cecila V. Tobias | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025011614553822/5575025011614553822_signature_31.png
| 2023-04-17 | 2023-04-13 10:39 | GMA Rural Health Unit II | Brgy Memije, General Mariano Alvarez | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24238 | PBSP | ldlconti@yahoo.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1365 |  | GXIV-D | 814152 | Unit is due for XpertCheck | XpertCheck | Performed XpertCheck Performed Maintenance Generate Reports | 3 Modules are Operational, 1 Module Reported | XpertCheck: 2 modules passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025011619961962/5575025011619961962_signature_28.png | Lorena L. Conti | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025011619961962/5575025011619961962_signature_31.png
| 2023-04-17 | 2023-04-12 10:41 | Kawit Rural Health Unit | Magdiwang Hwhy, Kawit | 4A | Cavite | Laboratory | NO | MAINTENANCE ReXpertCheck | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24241 | DOH | alambraaljhon0@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 999 |  | GXIV-D | 820705 | Module 3 and 4 with Temperature Errors | XpertCheck | Performed Maintenance Performed XpertCheck in Modules 3 and 4 Generate Reports | 2 Modules are Operational, 2 Module are Defective | Used Spare Kits. 3 and 4 Failed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025021614275728/5575025021614275728_signature_28.png | Aljhon G, Alambra | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025021614275728/5575025021614275728_signature_31.png
| 2023-04-17 | 2023-04-13 12:05 | Bulihan Rural Health Unit | AFP Housing, Bulihan, Silang | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24240 | PBSP | alei22775@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 790 |  | GXIV-D | 838446 | Module 3 with Temperature Error | Module Replacement | Performed Maintenance Performed Module Replacement Generate Reports | All Modules are Operational | B3: 861472 -> 730137 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025021617293334/5575025021617293334_signature_28.png | Alejandra Alegre | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025021617293334/5575025021617293334_signature_31.png
| 2023-04-17 | 2023-04-12 09:44 | Caridad Health Center | Dela Cruz St, Cavite City | 4A | Cavite | Laboratory | NO | MAINTENANCE ReXpertCheck | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24239 | DOH | roblesjm1230@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1482 |  | GXIV-D | 821227 | Module 2 with Temperature Error | ReXpertCheck | Performed Maintenance Performed ReXpertCheck on Module 2 Generate Reports | 3 Modules are Operational, 1 Module is Defective | Used Spare Kit |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025021617690053/5575025021617690053_signature_28.png | JAN MARIE KIRSTEN ROBLES | https://www.jotform.com/uploads/macarewecare/222927978372067/5575025021617690053/5575025021617690053_signature_31.png
| 2023-04-15 | 2023-04-15 21:30 | St.lukes medical center qc | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 24236 | COMMERCIAL | biomed.qc@stlukes.com.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 250 ASSAY: Others, COUNT: 3 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 71 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2981 |  | GX4 | 807155 | As per end user they encounter error h/w failed in C1,dusty filter | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed restart 9f workstation | Troubleshooting and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5573748400719081949/5573748400719081949_signature_28.png | Phillip Jordan M. Salvador RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5573748400719081949/5573748400719081949_signature_31.png
| 2023-04-15 | 2023-04-14 14:36 | Chinese general hospital | Blumentritt road sta cruz,manila | NCR | NCR 1st - Manila | Bloodbank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 24235 | COMMERCIAL | cghmcbloodbank2016@gmail.com | 27 |  |  | MCS + | 08M463 | As per end user they encounter error in initialization red valve error | Performed self test,checked cpu board ,pneumatic pump,checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test,performed isolation running | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5573733970717709515/5573733970717709515_signature_28.png | Matt Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5573733970717709515/5573733970717709515_signature_31.png
| 2023-04-15 | 2023-04-14 13:29 | Buhangin District Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24234 | DOH | buhanginlaboratory@gmail.com | 24.8 |  |  | GX-IV 4D | 838232 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5573342640016354065/5573342640016354065_signature_28.png | Catherine Deocampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5573342640016354065/5573342640016354065_signature_31.png
| 2023-04-14 | 2023-04-14 11:50 | Jose B. Lingad Memorial General Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24232 | COMMERCIAL | jblmgh.is@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 79 ASSAY: Xpert HCV VL, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1169 | 2023-04-14 | Genexpert IV | 812571 | Genexpert not seen on Application | Troubleshoot the Etherenet Cable Replaced the Ethernet Cable | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Ethernet Cable | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572836538407593902/5572836538407593902_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5572836538407593902/5572836538407593902_signature_31.png
| 2023-04-14 | 2023-04-14 13:05 | Mexico Rural Health Unit I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24233 | DOH | jasperdizon20@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 115 ASSAY: Xpert BCR-ABL Ultra, COUNT: 348 | 2023-04-14 | Genexpert IV | 831877 | Module A4 error | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572836538408863721/5572836538408863721_signature_28.png | Von Romeo Pacheco | https://www.jotform.com/uploads/macarewecare/222927978372067/5572836538408863721/5572836538408863721_signature_31.png
| 2023-04-14 | 2023-04-14 13:05 | Mexico Rural Health Unit I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24231 | DOH | jasperdizon20@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 115 ASSAY: Xpert BCR-ABL Ultra, COUNT: 348 | 2023-04-14 | Genexpert IV | 831877 | Module A4 error | Performed Troubleshooting | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572827468402041910/5572827468402041910_signature_28.png | Von Romeo Pacheco | https://www.jotform.com/uploads/macarewecare/222927978372067/5572827468402041910/5572827468402041910_signature_31.png
| 2023-04-14 | 2023-04-14 11:50 | Jose B. Lingad Memorial General Hospital | Dolores, City of San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24231 | COMMERCIAL | jblmgh.is@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 79 ASSAY: Xpert HCV VL, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1169 | 2023-04-14 | Genexpert IV | 812571 | Genexpert not seen on Application | Troubleshoot the Etherenet Cable Replaced the Ethernet Cable | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Ethernet Cable | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572827468409499740/5572827468409499740_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5572827468409499740/5572827468409499740_signature_31.png
| 2023-04-14 | 2023-04-14 15:45 | Bohol Provincial Health Office | Dao, Tagbilaran | 7 | Cebu | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24230 | PBSP | elyngdumaran@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 1129 ASSAY: Xpert MTB/RIF Ultra, COUNT: 655 |  | Gx IV | 804286 | Detected computer virus | Run kapersky anti virus | Run kapersky anti virus | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572709814416827010/5572709814416827010_signature_28.png | Elyn G. Dumaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5572709814416827010/5572709814416827010_signature_31.png
| 2023-04-14 | 2023-04-14 13:27 | Dr. Jorge P. Royeca Hospital | General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24229 | PBSP | fatchoi12001@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 163 ASSAY: Xpert MTB/RIF Ultra, COUNT: 530 ASSAY: Xpert HIV-1 Viral Load, COUNT: 53 | 2023-04-14 | Gx-IV | 805007 | Error 3005 | Resolved Error 3005 Perform Genexpert Xpertcheck at module 2,3, and 4 Perform Genexpert machine Preventive maintenance | Resolved Error 3005 Perform Genexpert Xpertcheck at module 2,3, and 4 Perform Genexpert machine Preventive maintenance | 4 operational modules | Module 2,3, and 4 Passed for Xpertcheck/Calibration | Xpertcheck CE-5 | Lot no. 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572610541812427679/5572610541812427679_signature_28.png | Maria Ivy A. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5572610541812427679/5572610541812427679_signature_31.png
| 2023-04-14 | 2023-04-14 13:32 | Medellin Rhu | Poblacion Medellin Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24228 | DOH | rhumedellin@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 440 |  | Gx4 | 838506 | Due for monthly maintenance | Performed preventive maintenance the genexpert machine | Clean machine, fan filter, Icore and cartridge bay Gather data IQR, Sys log, Gxx files and Assay stats Performed self test, plunger maintenance test Test the output voltage of the Avr and Ups Backup the database | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572603059554268665/5572603059554268665_signature_28.png | Raul Galano, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5572603059554268665/5572603059554268665_signature_31.png
| 2023-04-14 | 2023-04-14 12:06 | Tanauan Rural Health Unit | Mayapis, Brgy. 3, Tanauan | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 24227 | DOH | ramilomher27@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1676 | 2023-03-27 | GXIV | 821299 | Unit for PM. For module replacement of module A2. For troubleshooting of module A3 (H/W failure). | Performed PM. Performed module replacement on A2. Performed troubleshooting on A3. Generated reports. | PM, module replacement, troubleshooting. | Unit is operational. Unable to Xpertcheck module A3 (731305) with H/W failure. Xpertcheck software shows unable to Xpertcheck even after reseting gateway board and connections. Module A3 for reporting |  | Gx module A2 (639512) in replacement to defective A2 (793145). | DR 9802 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572561737489212122/5572561737489212122_signature_28.png | Edita Ramilo | https://www.jotform.com/uploads/macarewecare/222927978372067/5572561737489212122/5572561737489212122_signature_31.png
| 2023-04-14 | 2023-04-14 11:23 | SIPALAY CITY HEALTH OFFICE | Sipalay City, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 24226 | PBSP | olivesabrine27@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 363 ASSAY: Xpert MTB/RIF Ultra, COUNT: 859 |  | GX-IV | 815281 | Defective module B3/714096. | Preventive Maintenance - module bay cleaning, optic slots cleaning, filter cleaning, plunger rod maintenance, archive tests. | Replaced module B3 with new SN 629047. | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572523302526047617/5572523302526047617_signature_28.png | oliva decena | https://www.jotform.com/uploads/macarewecare/222927978372067/5572523302526047617/5572523302526047617_signature_31.png
| 2023-04-14 | 2023-04-14 10:31 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24225 | DOH | dulzna@gmail.com | 23.3 | ASSAY: Xpert MTB/RIF, COUNT: 191 ASSAY: Xpert MTB/RIF Ultra, COUNT: 481 ASSAY: Xpert HIV-1 Viral Load, COUNT: 15 | 2023-04-14 | Gx-IV | 838570 | Module 4 for Service Required Module 3 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Replace module 4 SN: 903024 with new module SN: 617089 Perform Xpertcheck at module 3 for troubleshooting | Perform Genexpert machine Preventive maintenance Replace module 4 SN: 903024 with new module SN: 617089 Perform Xpertcheck at module 3 for troubleshooting | 4 operational modules | Module 3 passed the Module Xpertcheck | Genexpert module 6 color SN:617089 | DR number: 9822 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572504181812173000/5572504181812173000_signature_28.png | Dulce Aurora Ariston | https://www.jotform.com/uploads/macarewecare/222927978372067/5572504181812173000/5572504181812173000_signature_31.png
| 2023-04-14 | 2023-04-14 10:26 | President Roxas Rural Health Unit | President Roxas, Capiz | 6 | Capiz | TB Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24224 | PBSP | mlaadelacruz@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 195 ASSAY: Xpert MTB/RIF Ultra, COUNT: 726 |  | GXIV-4D | 815282 | CPU error - IPv4 PXE | #NAME? | #NAME? | CPU and all modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572488721224207976/5572488721224207976_signature_28.png | Ma Louella Ann dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5572488721224207976/5572488721224207976_signature_31.png
| 2023-04-14 | 2023-04-14 10:08 | SAN JOSE RURAL HEALTH UNIT - RTDL SITE | TRADE TOWN DALIPE, SAN JOSE | 6 | Antique | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24222 | DOH | sanjoseantiquegxsite@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 395 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2211 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 | 2023-03-28 | GX-IV | 838412 | FOR MODULE B1 AND B2 REPLACEMENT | #NAME? | -PERFORM MODULE B1 DEFECTIVE SN:86116 REPLACEMENT WITH NEW SN:776202 AND MODULE B3 DEFECTIVE SN:861118 REPLACEMENT WITH NEW SN:717580 -PERFORM PREVENTIVE MAINTENANCE | EQUIPMENT IS OPERATIONAL | MODULE B4 H/W FAIL | 900-0491RS | 98339834 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572485881718621297/5572485881718621297_signature_28.png | Oscar Delano Salazar IV | https://www.jotform.com/uploads/macarewecare/222927978372067/5572485881718621297/5572485881718621297_signature_31.png
| 2023-04-14 | 2023-04-14 10:17 | SAN JOSE RURAL HEALTH UNIT - RTDL SITE | TRADE TOWN DALIPE, SAN JOSE | 6 | Antique | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  |  |  |  | 24223 | PBSP | sanjoseantiquegxsite@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 395 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2211 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 |  | GX-IV | 804714 | MODULE C3 SN 886646 ERROR 2032 | COLLECT IQR, SYSTEM LOG, AND ARCHIVE | -PERFORM GX EQUIPMENT CLEANING AND MAINTENANCE -PERFORM MODULE C3 ERROR 2032 TROUBLESHOOTING AND INSPECT WIRES | EQUIPMENT IS OPERATIONAL | MODULE C3 FOR OBSERVATION |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572485891713687934/5572485891713687934_signature_28.png | Oscar Delano Salazar IV | https://www.jotform.com/uploads/macarewecare/222927978372067/5572485891713687934/5572485891713687934_signature_31.png
| 2023-04-14 | 2023-04-14 10:11 | Talisay Rural Health Unit | Poblacion 5, Talisay | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24221 | DOH | medcertrhutalisay@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 399 | 2023-04-11 | GXIV | 820201 | Unit for PM. For module replacement of module A2. | Performed PM. Performed module replacement on A2. Generated reports | PM and module replacement. | Unit is operational. |  | Gx module A2 (694110) in replacement to defective A2 (721636). | DR 9805 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572477072528895684/5572477072528895684_signature_28.png | Adettlou A. Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5572477072528895684/5572477072528895684_signature_31.png
| 2023-04-14 | 2023-04-14 09:56 | Eastern Visayas Medical Regional Center | Bagacay, Tacloban City | 8 | Leyte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24220 | PBSP | merl_agner62@yahoo.com.sg | 26 | ASSAY: Xpert MTB/RIF, COUNT: 1860 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3499 ASSAY: Xpert HIV-1 Viral Load, COUNT: 158 | 2023-04-14 | GXIV | 808631 | #NAME? | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572476514361261078/5572476514361261078_signature_28.png | ma. merlina a. vistal | https://www.jotform.com/uploads/macarewecare/222927978372067/5572476514361261078/5572476514361261078_signature_31.png
| 2023-04-14 | 2023-04-14 01:51 | Panabo City Health Office | Poblacion, Panabo City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24219 | DOH | nadineeust67@gmail.com | 23.8 |  |  | GX-IV 4D | 838295 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Only 2 modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572432904615231872/5572432904615231872_signature_28.png | Nadine Gemma Eustaquio RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5572432904615231872/5572432904615231872_signature_31.png
| 2023-04-14 | 2023-04-13 15:14 | Lung center of the Philippines | Quezon avenue qc | NCR | NCR 2nd - Quezon City | Koica laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 24218 | PBSP | ccsinaguinan.pbsp@gmail.com | 28 |  | 2023-04-13 | GX4 r1 | 704584 | For laptop replacement CNU412BVFK - CPH1221F30044 DELL5420 | Installed laptop, upload ADF files,gathered reports | Laptop installation | Operational |  | Laptop | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572409461218221988/5572409461218221988_signature_28.png | Cecilia Sinaguinan | https://www.jotform.com/uploads/macarewecare/222927978372067/5572409461218221988/5572409461218221988_signature_31.png
| 2023-04-13 | 2023-04-13 21:08 | Cong. Simeon Toribio Memorial Hospital | Carmen, Bohol | 7 | Cebu | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 24217 | DOH | Himaya.vale@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 901 ASSAY: Xpert MTB/RIF Ultra, COUNT: 79 |  | GXIV | 838394 | Defective module sn861075 | New module installed sn676721 | Replace module, clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5572011553518829357/5572011553518829357_signature_28.png | Himaya Vale | https://www.jotform.com/uploads/macarewecare/222927978372067/5572011553518829357/5572011553518829357_signature_31.png
| 2023-04-13 | 2023-04-13 16:37 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24216 | DOH | riazarizzamae@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 128 ASSAY: Xpert MTB/RIF Ultra, COUNT: 817 | 2023-05-22 | Genexpert IV | 831637 | UPS not working | Performed Troubleshooting on UPS | Performed Maintenance Gathered Reports Performed Module Replacement = 784354 - 694493 Performed Functionality Test | Unit is Operational | UPS for Replacement | Genexpert Module - 694493 | 9773 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571935052024024618/5571935052024024618_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5571935052024024618/5571935052024024618_signature_31.png
| 2023-04-13 | 2023-04-13 16:37 | Palayan City Health Office | Singalat, Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24215 | DOH | riazarizzamae@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 128 ASSAY: Xpert MTB/RIF Ultra, COUNT: 817 | 2023-05-22 | Genexpert IV | 831637 | UPS not working | Performed Troubleshooting on UPS | Performed Maintenance Gathered Reports Performed Module Replacement = 784354 - 694493 Performed Functionality Test | Unit is Operational | UPS for Replacement | Genexpert Module - 694493 | 9773 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571934992023000740/5571934992023000740_signature_28.png | Rizza Mae Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5571934992023000740/5571934992023000740_signature_31.png
| 2023-04-13 | 2023-04-13 16:36 | Balasan Rural Health Unit | Balasan, Iloilo | 6 | Iloilo | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24214 | PBSP | charliemagnedk2@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 231 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1386 | 2023-04-13 | GX4 | 807555 | Module B1 - hardware failure | -Replace defective module B1/669140 with new module SN: 711361 | #NAME? | All modules are operational, module B2 for module replacement. | Module B2 - failed XPC; Module B3 & B4 - passed; pulled out one extra XPC | 900-0491RS GeneXpert Module 6 Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571857171918515230/5571857171918515230_signature_28.png | Charlie Magne Dequito | https://www.jotform.com/uploads/macarewecare/222927978372067/5571857171918515230/5571857171918515230_signature_31.png
| 2023-04-13 | 2023-04-13 16:40 | Mariveles Rural Health Unit | Mariveles Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24213 | DOH | angelicaroseconos@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 103 ASSAY: Xpert MTB/RIF Ultra, COUNT: 931 ASSAY: Xpert SARS-CoV-2, COUNT: 196 | 2023-02-07 | GXIV-D | 831894 | Due for calibration | Preventeive maintenance Calibration | Performed surface cleaning Performed preventive maintenance Xpertcheck applivcation update Performed calibration Gathered reports Performed functionality test | Unit is operational | Error 2034 encountered during calibration | Xpertcheck kit | 9605 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571849052691184595/5571849052691184595_signature_28.png | Angelica Conos | https://www.jotform.com/uploads/macarewecare/222927978372067/5571849052691184595/5571849052691184595_signature_31.png
| 2023-04-13 | 2023-04-13 14:20 | International Organization for Migration | H.V. Dela Costa | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | John Monsour Doctor |  |  |  |  | 24212 | COMMERCIAL | jsabalboro@iom.int | 25 |  |  | Primostar | 3134004145 | For maintenance | Performed maintenance | Performed maintenance Cleaned eyepiece, objectives and microscope body Check focus | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571831344316414317/5571831344316414317_signature_28.png | Joy B. Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5571831344316414317/5571831344316414317_signature_31.png
| 2023-04-13 | 2023-04-12 10:03 | Pasig City General Hospital | M. Eusebio | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24211 | PBSP | maniquizederlina@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 363 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2683 |  | GX-IV-D | 805010 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational. Module B1/644019 service required |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830454312206394/5571830454312206394_signature_28.png | Ederlina Sagales | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830454312206394/5571830454312206394_signature_31.png
| 2023-04-13 | 2023-04-13 12:38 | International Organization for Migration | H.V. Dela Costa | NCR | NCR 4th - Makati | Laboratory | NO | REPAIR | CEPHEID | John Monsour Doctor |  |  |  |  | 24209 | IOM | jsabalboro@iom.int | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 579 ASSAY: Others, COUNT: 73 ASSAY: Xpert SARS-CoV-2, COUNT: 34 | 2023-04-10 | GX-IV-D-10C | 110004816 | Modules B2 and B4 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | All modules operational. |  | Old: B2/210024923, B4/210024322 New: B2/210017005, B4/210021760 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830454312536901/5571830454312536901_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830454312536901/5571830454312536901_signature_31.png
| 2023-04-13 | 2023-04-13 12:46 | International Organization for Migration | H.V. Dela Costa | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24208 | IOM | jsabalboro@iom.int | 25 |  |  | GX-IV-D | 805288 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830454314585675/5571830454314585675_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830454314585675/5571830454314585675_signature_31.png
| 2023-04-13 | 2023-04-12 12:04 | Nagpayong Super Health Center | Jacob St., Nagpayong | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24210 | DOH | michelindg23@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 176 ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 | 2023-02-07 | GX-IV-D | 831905 | For UPS replacement | Replaced UPS | Performed maintenance Replaced UPS Run XpertCheck on module B2 Generated reports | Unit is operational. Modules B2, B3 and B4 not operational. Reported already. |  | UPS | DR# 9599 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830454315115770/5571830454315115770_signature_28.png | Michelle De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830454315115770/5571830454315115770_signature_31.png
| 2023-04-13 | 2023-04-13 13:17 | International Organization for Migration | H.V Dela Costa | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24211 | IOM | jsabalboro@iom.int | 25 |  |  | GX-IV-D | 110002868 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore Generated reports | All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830464316374052/5571830464316374052_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5571830464316374052/5571830464316374052_signature_31.png
| 2023-04-13 | 2023-04-13 15:53 | Mariveles District Hospital | Mariveles Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24207 | MACARE-HBDC | mdh.mdl2020@gmail.com | 30 | ASSAY: Xpert SARS-CoV-2, COUNT: 1746 |  | GXXVI-D | 724561 | Stuck cartridge on D3 | Repair and troubleshooting | Performed maintenance Performed surface cleaning Remove stuck cartridge Gathered reports Performed functionality test | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571821242695894161/5571821242695894161_signature_28.png | John Nathaniel C. Sallao | https://www.jotform.com/uploads/macarewecare/222927978372067/5571821242695894161/5571821242695894161_signature_31.png
| 2023-04-13 | 2023-04-13 12:03 | Sta Rosa District Hospital | Sta Rosa Olango Lapulapu city | 7 | Cebu | Tb laboratory | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24206 | DOH | ariestumz@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 30 |  | Gx 4 | 838534 | Due for Monthly calibration | Conduct preventive maintenance | Clean filter, clean Icore, and cartridge bay, check ups voltage | All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571815037221819276/5571815037221819276_signature_28.png | Ariston Tumulak | https://www.jotform.com/uploads/macarewecare/222927978372067/5571815037221819276/5571815037221819276_signature_31.png
| 2023-04-13 | 2023-04-13 14:17 | Polanco Rural Health Unit | Poblacion South polanco, Zamboanga del Norte | 9 | Zamboanga del Norte | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24205 | PBSP | rhupolancocl@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 674 | 2023-04-13 | GX 4 | 838539 | None | Replaced Module 862313 with Module 647451 | Module Replacement and Preventive Maintenence | All Modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571788183521672579/5571788183521672579_signature_28.png | Gergean Lura | https://www.jotform.com/uploads/macarewecare/222927978372067/5571788183521672579/5571788183521672579_signature_31.png
| 2023-04-13 | 2023-04-13 14:19 | Calamba City Health Office | Calamba | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24204 | PBSP | 2052456@aup.edu.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1918 | 2023-04-11 | GXIV | 807570 | Unit for PM. For re-Xpertcheck of module B1. | Performed PM. Performed re-Xpertcheck on module B1. Generated reports. | PM and Xpertcheck. | Unit is operational. Module B1 (624926), service required. |  | Xpertcheck kit | G 5429 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571780795017058409/5571780795017058409_signature_28.png | Mary Jane Lourdes R. Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5571780795017058409/5571780795017058409_signature_31.png
| 2023-04-13 | 2023-04-12 15:43 | Ipil Rural Health Unit | Poblacion Ipil, Zamboanga Sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24203 | PBSP | rheasolana@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1553 | 2023-04-12 | GX 4 | 808615 | Module 2 hardware failed | Module Replacement and Preventive maintenence | Module Replacement and Preventive maintenence | All 4 modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571758522525970939/5571758522525970939_signature_28.png | RHEAMAY SOLANA | https://www.jotform.com/uploads/macarewecare/222927978372067/5571758522525970939/5571758522525970939_signature_31.png
| 2023-04-13 | 2023-04-13 13:14 | Malungon Rural Health Unit | Malungon | 12 | Sarangani | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24202 | DOH | ulbenarioe@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 45 |  | Gx-IV | 838546 | Module 2 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 2 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 2 for troubleshooting | 3 operational modules | Module 2 passed for Xpertcheck/calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571750335003751256/5571750335003751256_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5571750335003751256/5571750335003751256_signature_31.png
| 2023-04-13 | 2023-04-13 13:14 | Malungon Rural Health Unit | Malungon | 12 | Sarangani | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24201 | DOH | ulbenarioe@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 45 |  | Gx-IV | 838546 | Module 2 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 2 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 2 for troubleshooting | 3 operational modules | Module 2 passed for Xpertcheck/calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571750265002002514/5571750265002002514_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5571750265002002514/5571750265002002514_signature_31.png
| 2023-04-13 | 2023-04-13 12:15 | Ajuy Rural Health Unit | Ajuy, Iloilo | 6 | Iloilo | TB Dots | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24200 | PBSP | jaaziahjesse@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 24 | 2023-04-13 | GX4 | 815429 | Module B1 & B3 - hardware failure | -Replace defective module B1/695903 with new module SN: 718278 -Replace defective module B3/695905 with new module SN: 793567 | #NAME? | All four modules are operational. | Module B2 & B4 passed Xpertcheck calibration. Pulled out 2 extra XPC | 2 900-0491RS GeneXpert Module 6 | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571715516617794656/5571715516617794656_signature_28.png | Razhel Anne Libo-on | https://www.jotform.com/uploads/macarewecare/222927978372067/5571715516617794656/5571715516617794656_signature_31.png
| 2023-04-13 | 2023-04-13 12:46 | Samar Provincial Hospital | Catbalogan City | 8 | Samar | TB dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24199 | PBSP | leslieleones36@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 1047 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1036 | 2023-04-13 | GXIV | 808291 | #NAME? | -Replace defective module C2/ 695897 into 676713 - Performed functionality test -Performed preventive maintenance - Generate reports | #NAME? | Unit is Operational (Module C3 H/W Failed) |  | GeneXpert Module 6 Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571708041711597033/5571708041711597033_signature_28.png | leslie daryl leones | https://www.jotform.com/uploads/macarewecare/222927978372067/5571708041711597033/5571708041711597033_signature_31.png
| 2023-04-13 | 2023-04-13 12:18 | Nasipit Rural Health Unit | Brgy. Talisay, Nasipit | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24198 | DOH | peterdolotina@gmail.com | 28.2 | ASSAY: Xpert MTB/RIF, COUNT: 162 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1323 | 2023-03-24 | GX-4 | 837611 | None | None | Replacing Module B2 with SN: 806518 to SN: 698547 | All modules are passed and functional |  | GenXpert Module 6-Color | 9784 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571694579127651362/5571694579127651362_signature_28.png | Peter Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/5571694579127651362/5571694579127651362_signature_31.png
| 2023-04-13 | 2023-04-13 12:16 | Lipa City Health Office | Morada Ave. Lipa | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24197 | PBSP | mylaolayao@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1010 | 2023-04-11 | GXIV | 808642 | Unit for PM. For module replacement of module C1. | Performed PM. Performed module replacement on C1. | PM and module replacement. | Unit is operational. |  | Gx module C1 (693460) in replacement to defective C1 (667760). | G 5430 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571690331111463655/5571690331111463655_signature_28.png | Myla Olayao | https://www.jotform.com/uploads/macarewecare/222927978372067/5571690331111463655/5571690331111463655_signature_31.png
| 2023-04-13 | 2023-04-13 12:14 | SEBASTE RURAL HEALTH UNIT | SEBASTE ANTIQUE | 6 | Antique | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24197 | PBSP | francisconaciongayo@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 203 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1812 |  | GX-IV | 808287 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL, MODULE B3 DISABLED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571690334529862669/5571690334529862669_signature_28.png | Francisco S. Naciongayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5571690334529862669/5571690334529862669_signature_31.png
| 2023-04-13 | 2023-04-13 11:13 | Tondo foreshore | Pacheco st, tondo manila | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION | CEPHEID | Noel Floresca |  |  |  |  | 24196 | COMMERCIAL | synch28@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXXVI | 110016343 | For installation | Installed and set up unit, gathered reports, test run unit | Installation | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571680015314868564/5571680015314868564_signature_28.png | Dennis B. Aguilon | https://www.jotform.com/uploads/macarewecare/222927978372067/5571680015314868564/5571680015314868564_signature_31.png
| 2023-04-13 | 2023-04-13 10:58 | Cadiz CHO | Cadiz City | 6 | Negros Occidental | TB DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 24195 | DOH | zhevebacomo@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 153 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1361 |  | GX4 | 838499 | Defective module for B1/862379, B2/862380 and B3/862381 | Preventive Maintenance | New Modules for B1/760112, B2/606933 and B3/715690 | Module B4 for replacement due failed in Xpert Check |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571657537971299346/5571657537971299346_signature_28.png | Zherlyn Bacomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5571657537971299346/5571657537971299346_signature_31.png
| 2023-04-13 | 2023-04-13 10:31 | Cuenca Rural Health Unit | Poblacion 2, Cuenca | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24194 | DOH | LIGAYAMIJARES50@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 583 | 2023-03-27 | GXIV | 821297 | Unit for PM. Unit for Xpertcheck. For module replacement of module A1. | Performed PM. Performed module replacement on module A1. Performed Xpertcheck. Generate reports. | PM, Xpertcheck, module replacement. | Unit is operational. Module A3, service required. |  | Gx module A1 (750974) in replacement to defective A1 (731311). Xpertcheck. | Gx module DR 9808, Xpertcheck DR 9485 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571626611112953102/5571626611112953102_signature_28.png | Ligaya R. Mijares | https://www.jotform.com/uploads/macarewecare/222927978372067/5571626611112953102/5571626611112953102_signature_31.png
| 2023-04-13 | 2023-04-12 13:06 | Dimasalang Health Center (Lanuza Health Center) | Blumentritt, Sta. Cruz | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24193 | PBSP | joneevamarie@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1341 | 2023-04-11 | GXIV | 805111 | Unit for PM. For module replacement of module B1. | Performed PM. Performed module replacement on B1. Generated reports. | PM and module replacement | Unit is operational. |  | Gx module B1 (874866) in replacement to defective B1 (630369) | G 5431 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571610818413235457/5571610818413235457_signature_28.png | Eva Marie Jone | https://www.jotform.com/uploads/macarewecare/222927978372067/5571610818413235457/5571610818413235457_signature_31.png
| 2023-04-13 | 2023-04-12 14:30 | Calbayog City Health Office | Calbayog City | 8 | Samar | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24192 | PBSP | maganteelena@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 562 ASSAY: Xpert MTB/RIF Ultra, COUNT: 679 | 2023-04-12 | GXIV | 804265 | #NAME? | - Replace defective Module B4/721870 - SN: 725420 - Performed preventive maintenance -Performed functionality test - Generate Reports | #NAME? | Unit is Operational |  | GeneXpert Module 6 Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571602945643841009/5571602945643841009_signature_28.png | Elena M. Villa | https://www.jotform.com/uploads/macarewecare/222927978372067/5571602945643841009/5571602945643841009_signature_31.png
| 2023-04-13 | 2023-04-13 09:51 | Gandara RHU | Gandara Samar | 8 | Samar | TB dots | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24191 | DOH | materesaaliman78@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 269 ASSAY: Xpert MTB/RIF Ultra, COUNT: 511 | 2023-04-13 | GXIV | 838406 | #NAME? | -Replace defective module B3/754026 into 617444 -Run XPC on module B1 -Performed preventive maintenance -Performed functionality test - Generate Reports | #NAME? | Unit is Operational | Module B1 service required on XPC | - GeneXpert Module 6 Color - 1 XPC cartridge | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571602945645582858/5571602945645582858_signature_28.png | Ma.teresa E.Aliman | https://www.jotform.com/uploads/macarewecare/222927978372067/5571602945645582858/5571602945645582858_signature_31.png
| 2023-04-13 | 2023-04-13 08:37 | Mayor Hilarion A. Ramiro Sr. Medical Center | Ozamiz City | 10 | Misamis Occidental | Microbiology Laboratory | YES | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24190 | PBSP | mharsmicrobio@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1136 ASSAY: Xpert HIV-1 Viral Load, COUNT: 34 |  | GeneXpert | 804765 | Module B4/618492 failed during calibration | Replace module defective B4/618492 New Module replacement SN/617234 | Module Replacement Generate reporting files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571594307648829285/5571594307648829285_signature_28.png | Felipe D Saren Jr - RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5571594307648829285/5571594307648829285_signature_31.png
| 2023-04-13 | 2023-04-13 09:20 | Mayor Hilarion A. Ramiro Sr. Medical Center | Ozamiz City | 10 | Misamis Occidental | Molecular Laboratory | YES | CALIBRATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24189 | PBSP | mharsmc.covidlabgenex@gmail.com | 29 | ASSAY: Xpert SARS-CoV-2, COUNT: 5638 |  | GeneXpert | 834437 | N/A | N/A | Xpertcheck calibration Generate Reporting files Database back up | Operational | B1: Passed , B2:Retry then Fail , B3/B4 : Retry calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571593557648124137/5571593557648124137_signature_28.png | Felipe D Saren Jr - RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5571593557648124137/5571593557648124137_signature_31.png
| 2023-04-13 | 2023-04-13 08:37 | Mayor Hilarion A. Ramiro Sr. Medical Center | Ozamiz City | 10 | Misamis Occidental | Microbiology Laboratory | YES | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24188 | PBSP | mharsmicrobio@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1136 ASSAY: Xpert HIV-1 Viral Load, COUNT: 34 |  | GeneXpert | 804765 | Module B4/618492 failed during calibration | Replace module defective B4/618492 New Module replacement SN/617234 | Module Replacement Generate reporting files Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571577275716592201/5571577275716592201_signature_28.png | Felipe D Saren Jr - RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5571577275716592201/5571577275716592201_signature_31.png
| 2023-04-13 | 2023-03-31 12:12 | San Pedro City Health Office 1 | San Pedro | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24187 | DOH | venersantos0425@gmail.com | 25 |  | 2023-03-27 | GXIV | 821224 | Unit for Pm. For module replacement of module B2. | Performed PM. Performed module replacement on B2. Generated reports. | PM and module replacement | Unit is operational. Unable to replace module B3 due to mistaken allocation. |  | Gx module B2 (678381) in replacement to defective B2 (709892). | DR 9799 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571553680706750052/5571553680706750052_signature_28.png | Vener R. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5571553680706750052/5571553680706750052_signature_31.png
| 2023-04-13 | 2023-04-12 08:07 | Amma Jadsac District Hospital | Poblacion Pudyol Apayao | CAR | Apayao | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24186 | PBSP | dlgblacar.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 19 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2300 | 2023-04-12 | GXIV-D | 807609 | A1 and A3 for module replacement, UPS not functional | Preventive maintenance Calibration Module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Performed calibration Gathered reports Performed functionality test | Unit is operational except A2 and A4. UPS for request | A2 failed calibration | 2 genexpert module, 1 xpertcheck kit | 400083939039353 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571313799416075454/5571313799416075454_signature_28.png | Daisy Lynn Lacar | https://www.jotform.com/uploads/macarewecare/222927978372067/5571313799416075454/5571313799416075454_signature_31.png
| 2023-04-12 | 2023-04-12 17:18 | Carcar RTDL | Valladolid Carcar Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24185 | DOH | chaska.lp14@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 668 |  | Gx4 | 838283 | Module 2 Serial Number 802769, Error 1001 | Adjust temperature of the aircon, clean filter and reset module 2 | Clean Filter, Reset module, Get data IQR and Sys log, Gxx files | All modules are available however during few more test error 1001 occured |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5571007357222373705/5571007357222373705_signature_28.png | Lucille Mae Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5571007357222373705/5571007357222373705_signature_31.png
| 2023-04-12 | 2023-04-12 11:53 | Mexico Rural Health Unit IV | Pandacaqui, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24183 | DOH | annbathan79@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 431 | 2023-04-22 | Genexpert IV | 831902 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 758257 & 758259 - 639028 & 908350 | Genexpert Module - 639028 & 908350 | 9769 & 9770 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570975645212231692/5570975645212231692_signature_28.png | Mary Ann Carreon | https://www.jotform.com/uploads/macarewecare/222927978372067/5570975645212231692/5570975645212231692_signature_31.png
| 2023-04-12 | 2023-04-12 10:50 | Magalang Rural Health Unit | San Nicolas 1, Magalang, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24184 | WHO | bevlozano29@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 145 ASSAY: Xpert MTB/RIF Ultra, COUNT: 440 | 2023-04-12 | Genexpert IV | 846044 | Module 2 stuck cartridge UPS not working | Open Module door manually | Performed Maintenance Gathered Reports Performed Troubleshooting Performed Functionality Test | Unit is Operational | UPS for Replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570975645213713891/5570975645213713891_signature_28.png | Beverly Rose D. Lozano | https://www.jotform.com/uploads/macarewecare/222927978372067/5570975645213713891/5570975645213713891_signature_31.png
| 2023-04-12 | 2023-04-12 13:25 | LA CARLOTA CITY HEALTH OFFICE | Jereza St., La Carlota City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 24182 | PBSP | shane.gipanago@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 278 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1506 |  | GX-IV | 807594 | Defective module 690194/A1 | #NAME? | - Replaced module 690194/A1 with new module SN 678090 | All 4 modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570871114717882457/5570871114717882457_signature_28.png | shane m. gipanago | https://www.jotform.com/uploads/macarewecare/222927978372067/5570871114717882457/5570871114717882457_signature_31.png
| 2023-04-12 | 2023-04-12 12:24 | Valentina G. Plaza Memorial Hospital | San Teodoro, Bunawan | 13 | Agusan del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24181 | PBSP | bdhclinlab@gmail.com | 25.15 | ASSAY: Xpert MTB/RIF, COUNT: 802 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4066 |  | GX-4 | 815436 | None | None | Replacing Module A2 with SN: 769073 to SN: 725400 | All modules passed and functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570849436916048802/5570849436916048802_signature_28.png | Rosephine Toyong | https://www.jotform.com/uploads/macarewecare/222927978372067/5570849436916048802/5570849436916048802_signature_31.png
| 2023-04-12 | 2023-04-12 10:58 | Northern Samar Provincial Hospital | Catarman Northern Samar | 8 | Northern Samar | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24180 | PBSP | dinahada70@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 793 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1495 ASSAY: Xpert HIV-1 Viral Load, COUNT: 29 | 2023-04-12 | GXIV | 807566 | Stuck cartridge on module B3 (SN:730410) | #NAME? | #NAME? | #NAME? | Module B4 disabled (multiple temperature error) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570780028853545303/5570780028853545303_signature_28.png | Dinah Ada | https://www.jotform.com/uploads/macarewecare/222927978372067/5570780028853545303/5570780028853545303_signature_31.png
| 2023-04-12 | 2023-04-12 10:55 | CALUYA MUNICIPAL HEALTH OFFICE | CALUYA, ANTIQUE | 6 | Antique | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  |  |  |  | 24179 | PBSP | pcb1.caluya@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: N/A |  | GX-II | 835103 | NEW INSTALLATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570776341019489034/5570776341019489034_signature_28.png | Mary Grace P. Francisco,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5570776341019489034/5570776341019489034_signature_31.png
| 2023-04-12 | 2023-04-12 09:47 | Tubigon Rural Health Unit | Potohan, Tubigon | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 24178 | PBSP | josephraytan@gmail.com | 21 |  |  | Gx IV | 815431 | Defective module 4 sn695938 | New module 4 sn913043 | Replace module | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570743275511702296/5570743275511702296_signature_28.png | Joseph Ray F. Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5570743275511702296/5570743275511702296_signature_31.png
| 2023-04-12 | 2023-04-12 09:36 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Shiedrick Daroya |  |  |  |  | 24177 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 02A087 | For checking of MCS+ due to manifold issue with Sir Hawie. | Performed troubleshooting of MCS+. Checked manifold sensor using manifold from new kits and service manifold. Tried diagnostic and UPP protocol cards for isolation of problem. | Troubleshooting and repair. | Unit is operational. Manifold sensor is operational. Kit issue with specific lot numbers. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570738711113443577/5570738711113443577_signature_28.png | Lalaine Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5570738711113443577/5570738711113443577_signature_31.png
| 2023-04-12 | 2023-04-12 09:31 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Bloodbank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Shiedrick Daroya |  |  |  |  | 24176 | MACARE-HBDC | lalaine.ilagan@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | For checking of MCS+ due to manifold issue with Sir Hawie. | Performed troubleshooting of MCS+. Checked manifold sensor using manifold from new kits and service manifold. Tried diagnostic and UPP protocol cards for isolation of problem. | Troubleshooting and repair. | Unit is operational. Manifold sensor is operational. Kit issue with specific lot numbers. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570738711119610673/5570738711119610673_signature_28.png | Lalaine Ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5570738711119610673/5570738711119610673_signature_31.png
| 2023-04-12 | 2023-04-12 09:23 | Cebu City Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24175 | DOH | ccmclaboratory01@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 456 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1052 | 2023-03-28 | GX4 | 838561 | Module A2 & A4 - hardware failure | -Replace defective module A2/752968 with new module SN: 638250 -Replace defective module A4/862900 with new module SN: 802146 | #NAME? | Module A1 disabled- frequent temp error; Module A2, A3, A4 are operational. | Module A1 for XPC verification | 2 - 900-0491RS GeneXpert Module 6-Color | 9832/9838 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570730273911987211/5570730273911987211_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5570730273911987211/5570730273911987211_signature_31.png
| 2023-04-12 | 2023-04-12 09:10 | Kidapawan City Health Office | Kidapawan City, North Cotabato | 12 | Cotabato | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24174 | DOH | sg93rmt@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF, COUNT: 267 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2076 | 2023-04-12 | Gx-IV | 838545 | Module 3 and 4 H/W failed | Perform Genexpert machine Preventive maintenance Replace module 4 SN:697836 with new module SN:901953 Perform Xpertcheck at module 3 for troubleshooting | Perform Genexpert machine Preventive maintenance Replace module 4 SN:697836 with new module SN:901953 Perform Xpertcheck at module 3 for troubleshooting | 3 operational modules |  | Genexpert module 6 color | DR no. 9823 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570725849965324298/5570725849965324298_signature_28.png | TESSA MAE E. SAGUINDAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5570725849965324298/5570725849965324298_signature_31.png
| 2023-04-12 | 2023-04-11 15:43 | Antipas Rural Health Unit | Antipas, North Cotabato | 12 | Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24174 | DOH | rhuantipas@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 157 ASSAY: Xpert MTB/RIF Ultra, COUNT: 545 | 2023-04-11 | Gx-IV | 838259 | Module 1 H/W failed | Replaced module 1 SN: 807148 with new module SN: 688478 Perform Genexpert machine preventive maintenance | Replaced module 1 SN: 807148 with new module SN: 688478 Perform Genexpert machine preventive maintenance | 4 operational modules |  | Genexpert module 6 color | DR number: 9783 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570725849965456476/5570725849965456476_signature_28.png | Alvin A. Codilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5570725849965456476/5570725849965456476_signature_31.png
| 2023-04-11 | 2023-04-11 15:49 | East avenue medical center | East avenue q.c | NCR | NCR 2nd - Quezon City | Bloodbank | YES | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 24173 | MACARE-HBDC | eamc.biomed@gmail.com | 27 |  | 2023-03-02 | MCS + | 15H044 | No display defective cpu board | Performed troubleshooting, delivered and installed for stand by cpu board ,checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | Troubleshooting | Unit operational | Returned defective cpu board for safe keeping | CPU Board assembly | 9660 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570187778321891357/5570187778321891357_signature_28.png | Ryan Baldelobar | https://www.jotform.com/uploads/macarewecare/222927978372067/5570187778321891357/5570187778321891357_signature_31.png
| 2023-04-11 | 2023-04-11 15:49 | East avenue medical center | East avenue q.c | NCR | NCR 2nd - Quezon City | Bloodbank | YES | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 24172 | MACARE-HBDC | eamc.biomed@gmail.com | 27 |  | 2023-03-02 | MCS + | 15H044 | No display defective cpu board | Performed troubleshooting, delivered and installed for stand by cpu board ,checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | Troubleshooting | Unit operational | Returned defective cpu board for safe keeping | CPU Board assembly | 9660 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570181648324337665/5570181648324337665_signature_28.png | Ryan Baldelobar | https://www.jotform.com/uploads/macarewecare/222927978372067/5570181648324337665/5570181648324337665_signature_31.png
| 2023-04-11 | 2023-04-05 15:51 | Baliuag District Hospital | Baliwag, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 24171 | PBSP | reyeshannahxx@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 80 ASSAY: Xpert MTB/RIF Ultra, COUNT: 371 ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 | 2023-04-05 | GXIV | 812770 | Module B3 (686201) for replacement | #NAME? | Replaced B3 | Unit is operational. |  | GeneXpert Module SN: 210071169 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570100991212802333/5570100991212802333_signature_28.png | Hannah Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5570100991212802333/5570100991212802333_signature_31.png
| 2023-04-11 | 2023-04-11 15:25 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24170 | DOH | laureanjo@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 898 ASSAY: Xpert MTB/RIF Ultra, COUNT: 708 ASSAY: Xpert SARS-CoV-2, COUNT: 18 |  | GX IV | 838527 | Defective module 4 sn769730 | New module 4 installed sn714503 | Replace new module, clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570075262319036319/5570075262319036319_signature_28.png | Laurean jo cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5570075262319036319/5570075262319036319_signature_31.png
| 2023-04-11 | 2023-04-11 14:03 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 24169 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 195 ASSAY: Xpert MTB/RIF, COUNT: 1061 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1295 |  | GX4 | 838285 | Module A3 - h/w fail | Replace module A3 temporarily with service module SN: 708880 | Replace module A3/676993 temporarily with service module SN: 708880 | All modules are operational. | Service module SN. 708880 installed and defective module pulled out. For Xpertcheck verification upon calibration. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570072369526896893/5570072369526896893_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5570072369526896893/5570072369526896893_signature_31.png
| 2023-04-11 | 2023-04-11 13:57 | Vicente Sotto Memorial Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology | NO | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 24168 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 203 ASSAY: Xpert MTB/RIF, COUNT: 926 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1364 | 2023-03-24 | GX4 | 820206 | Module A3/721922 - defective | Replace defective module A3/721922 with new module SN: 711106 | Module Replacement | All modules are operational. |  | 900-0491RS GeneXpert Module 6-Color SN. 711106 | 9782 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570072369527408904/5570072369527408904_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5570072369527408904/5570072369527408904_signature_31.png
| 2023-04-11 | 2023-04-11 14:29 | Mindanao Central Sanitarium | Pasabolong, Zamboanga City, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24167 | DOH | ybioamie@gmail.com | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1895 | 2023-04-11 | GX4 | 821214 | Module 1 and 2 error 4012 and 2014 | Preventive Maintenence, Checked for new errors and room temperature | Preventive Maintenence | Module 3 and 4 available. Module 1 and 2 must undergone Xpertcheck | Module 1 and 2 have already been reported multiple times. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570066793845860434/5570066793845860434_signature_28.png | Amie B. Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5570066793845860434/5570066793845860434_signature_31.png
| 2023-04-11 | 2023-04-04 14:33 | Calatagan Rural Health Unit | Ayala St., Calatagan | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24166 | DOH | 1001melani@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 387 | 2023-01-16 | GXIV-D | 820210 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | All Modules are Operational | All Modules Passed | XpertCheck Kit | 9468 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570057721217485997/5570057721217485997_signature_28.png | Melani Buenaventura | https://www.jotform.com/uploads/macarewecare/222927978372067/5570057721217485997/5570057721217485997_signature_31.png
| 2023-04-11 | 2023-04-11 14:29 | Bislig District Hospital | Poblacion, Bislig City | 13 | Surigao del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 24165 | PBSP | gvnovero@yahoo.com | 31.65 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1399 |  | GX-4 | 805280 | None | None | Replacing Module B1 with SN: 750170 to SN: 662236 Replacing Module B4 with SN: 902726 to SN: 674605 | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570044507328345671/5570044507328345671_signature_28.png | Georgina Novero | https://www.jotform.com/uploads/macarewecare/222927978372067/5570044507328345671/5570044507328345671_signature_31.png
| 2023-04-11 | 2023-04-11 13:43 | HINIGARAN RURAL HEALTH UNIT | Aguilar St., Hinigaran, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 24164 | PBSP | maryjoycarton1031@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 232 ASSAY: Xpert MTB/RIF Ultra, COUNT: 771 |  | GX-IV | 815283 | Defective modules 633207/A2 and 686121/A3 | #NAME? | - Replaced module 633207/A2 with new module SN 220004692. - Replaced module 686121/A3 with new module SN 609119 | All four modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5570019002619105207/5570019002619105207_signature_28.png | mary joy carton | https://www.jotform.com/uploads/macarewecare/222927978372067/5570019002619105207/5570019002619105207_signature_31.png
| 2023-04-11 | 2023-04-11 12:46 | Nazareth Health Center | Brgy. Nazareth, Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24163 | DOH | nanettebtolentino@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 979 |  | GeneXpert | 838269 | H/W Fail on module B3/676363 | Replaced module | Module replacement New module SN/688652 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569980230162115176/5569980230162115176_signature_28.png | Nanette B. Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/5569980230162115176/5569980230162115176_signature_31.png
| 2023-04-11 | 2023-04-11 10:44 | Maria Aurora Community Hospital | Saturno st. Brgy 1, Maria Aurora, Aurora | 3 | Aurora | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24162 | PBSP | mariabelenth@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 578 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2170 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 | 2023-03-14 | Genexpert IV | 807565 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement = 601597 - 628831 | Genexpert Module = 628831 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569967475211800761/5569967475211800761_signature_28.png | Mary Grace Pajarit | https://www.jotform.com/uploads/macarewecare/222927978372067/5569967475211800761/5569967475211800761_signature_31.png
| 2023-04-11 | 2023-04-11 11:51 | Minglanilla RHU | Ward 1, Pider Minglanilla Cebu | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24161 | DOH | brendagonzagaparedes@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 688 |  | Gx4 | 838512 | 4009 | Reset machine | Get Data, IQR and Sys log, Gxx files Clean filter, and ICore, clean bay | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569950684219928261/5569950684219928261_signature_28.png | Brenda Gonzaga Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/5569950684219928261/5569950684219928261_signature_31.png
| 2023-04-11 | 2023-04-11 11:51 | Minglanilla RHU | Ward 1, Pider Minglanilla Cebu | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24160 | DOH | brendagonzagaparedes@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 688 |  | Gx4 | 838512 | 4009 | Reset machine | Get Data, IQR and Sys log, Gxx files Clean filter, and ICore, clean bay | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569950214216311850/5569950214216311850_signature_28.png | Brenda Gonzaga Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/5569950214216311850/5569950214216311850_signature_31.png
| 2023-04-11 | 2023-04-11 10:09 | Tuburan Rural Health Unit | Poblacion,Tuburan,Cebu | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 24159 | PBSP | marielgracia20@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 56 ASSAY: Xpert MTB/RIF, COUNT: 106 | 2023-04-04 | Gx4 | 815485 | Module B2 With Serial Number 699653 Hardware Fail | Replace Module B2 with new serial number 791790 | Replace Module | All modules are functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569891457226802624/5569891457226802624_signature_28.png | Mariel R. Gracia | https://www.jotform.com/uploads/macarewecare/222927978372067/5569891457226802624/5569891457226802624_signature_31.png
| 2023-04-11 | 2023-04-04 11:27 | San Dionisio Health Center | Pelaez St. San Dionisio | NCR | NCR 4th - Parañaque | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24157 | DOH | marubi1019@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 41 ASSAY: Xpert MTB/RIF, COUNT: 290 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1209 | 2023-01-18 | GX-IV-D | 831898 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. All modules passed. |  | XpertCheck-CE-5 | DR# 9496 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569817491214200521/5569817491214200521_signature_28.png | Fatima Joyce Tresvalles | https://www.jotform.com/uploads/macarewecare/222927978372067/5569817491214200521/5569817491214200521_signature_31.png
| 2023-04-11 | 2023-04-04 14:07 | San Isidro Health Center | Dr. A Santos Ave | NCR | NCR 4th - Parañaque | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24155 | PBSP | herminadoa@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1564 ASSAY: Xpert HIV-1 Viral Load, COUNT: 84 | 2023-04-04 | GX-IV-D | 838192 | Modules B1 and B2 for replacement. Module B3 and B4 for XpertCheck | Replaced defective modules, run XpertCheck. | Performed maintenance Replaced defective modules Run XpertCheck Generated reports | Unit is operational. | Module B3/806897, B4/806898 failed in XpertCheck. | Old: B1/806894, B2/806895 New: B1/784193, B2/788537 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569817491218600406/5569817491218600406_signature_28.png | Amelia G.Herminado | https://www.jotform.com/uploads/macarewecare/222927978372067/5569817491218600406/5569817491218600406_signature_31.png
| 2023-04-11 | 2023-04-05 09:04 | Batis Health Center | F. Manalo St., Brgy. Batis | NCR | NCR 2nd - San Juan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24158 | DOH | almacenaio16@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 565 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1451 ASSAY: Xpert HIV-1 Viral Load, COUNT: 58 | 2023-03-27 | GX-IV-D | 831598 | Module A1 and A3 for replacement. | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. |  | Old: A1/754112, A3/677100 New: A1/630159, A2/613865 | DR# 9807 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569817491219653225/5569817491219653225_signature_28.png | Rosario T. Almacen | https://www.jotform.com/uploads/macarewecare/222927978372067/5569817491219653225/5569817491219653225_signature_31.png
| 2023-04-11 | 2023-04-05 11:29 | Cardinal Santos Medical Center | Wilson St. | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24156 | COMMERCIAL | egatchalian@csmc.ph | 22 | ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert HPV, COUNT: 16 ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert MTB/RIF Ultra, COUNT: 805 ASSAY: Others, COUNT: 9 |  | GX-IV-D-10C | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569817491219816868/5569817491219816868_signature_28.png | Velasco-Gatchalian, Ellaine Marie | https://www.jotform.com/uploads/macarewecare/222927978372067/5569817491219816868/5569817491219816868_signature_31.png
| 2023-04-10 | 2023-04-04 16:28 | Pili rhu | Pili Camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24154 | DOH | lfsanao724@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 70 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1113 | 2023-03-28 | GX4 | 838408 | For module replacement and maintenance B2 682797 - 220005368 B4 861254 - 662296 | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational b1,b2,b4 |  | 900 0491 rs | 9815 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569437608322497057/5569437608322497057_signature_28.png | Liezl F. Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/5569437608322497057/5569437608322497057_signature_31.png
| 2023-04-10 | 2023-04-04 12:36 | San jose RHU | San jose Camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24153 | DOH | jhuneric2008@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 52 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1893 | 2023-03-28 | GX4 | 83838928 | For module replacement B2 714964 - 607968 | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Module hw failed unable to performed re xpertcheck |  | 900 0491rs | 9816 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569436258321833513/5569436258321833513_signature_28.png | Rosmar V. Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5569436258321833513/5569436258321833513_signature_31.png
| 2023-04-10 | 2023-03-31 15:17 | Tropical disease foundation inc | Amorsolo st makati city | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24152 | MACARE-HBDC | rrjflores@tdf.org.ph | 26 | ASSAY: Xpert MTB/RIF, COUNT: 470 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1130 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX4 | 838134 | For maintenance | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569434398329100529/5569434398329100529_signature_28.png | Rupert Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5569434398329100529/5569434398329100529_signature_31.png
| 2023-04-10 | 2023-03-31 12:54 | Doña marta health center | Revilla st.,Pasay city | NCR | NCR 4th - Pasay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24151 | DOH | cerissalynsabado@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 754 | 2023-03-27 | GX4 | 831700 | For module replacement A2 699271 - 694702 | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | For Module replacement and maintenance | Operational |  | 900 - 0491 rs | 9795 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569428368325368787/5569428368325368787_signature_28.png | Cerissa Lyn Sabado | https://www.jotform.com/uploads/macarewecare/222927978372067/5569428368325368787/5569428368325368787_signature_31.png
| 2023-04-10 | 2023-03-31 14:37 | Tropical disease foundation inc | Amorsolo st.,Makati city | NCR | NCR 4th - Makati | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24150 | CONSORTIUM | rrjflores@tdf.org.ph | 23 | ASSAY: Xpert MTB/RIF, COUNT: 470 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1130 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-03-28 | GX4 | 812560 | For maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed xpertcheck | Maintenance and xpertcheck | All module passed |  | Xpertcheck-ce-5 | 9818 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569429128329580237/5569429128329580237_signature_28.png | Rupert Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5569429128329580237/5569429128329580237_signature_31.png
| 2023-04-10 | 2023-03-31 12:54 | Doña marta health center | Revilla st.,Pasay city | NCR | NCR 4th - Pasay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24149 | DOH | cerissalynsabado@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 754 | 2023-03-27 | GX4 | 831700 | For module replacement A2 699271 - 694702 | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | For Module replacement and maintenance | Operational |  | 900 - 0491 rs | 9795 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569428368325368787/5569428368325368787_signature_28.png | Cerissa Lyn Sabado | https://www.jotform.com/uploads/macarewecare/222927978372067/5569428368325368787/5569428368325368787_signature_31.png
| 2023-04-10 | 2023-03-31 10:04 | Best diagnostic corporation | Masikap st, Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24148 | COMMERCIAL | purchasing.bestdiagnostic@gmail.com | 30 | ASSAY: Xpert HBV VL, COUNT: 53 ASSAY: Xpert HCV VL, COUNT: 3 | 2023-03-16 | GX4 | 110008529 | For module replacement B1 220001787 - 210072548 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit, performed module replacement | Module replacement and maintenance | Operational |  | 900-0508 | 9697 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569426618329550964/5569426618329550964_signature_28.png | Bonafe,clariza | https://www.jotform.com/uploads/macarewecare/222927978372067/5569426618329550964/5569426618329550964_signature_31.png
| 2023-04-10 | 2023-03-30 15:30 | Domingo santiago health center | D.santiago st sampaloc manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24147 | DOH | fei_paas@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 989 | 2023-03-20 | GX4 | 831527 | Module replacement and maintenance ,ups unable to back defective battery | Module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | A1 disabled reported,A2, A3, A4 operational | For replacement of ups battery | 900 0491rs | 97309705 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569425008321484083/5569425008321484083_signature_28.png | Flory Fe A. Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5569425008321484083/5569425008321484083_signature_31.png
| 2023-04-10 | 2023-03-30 13:36 | Tatalon health center | Victory ave. Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24146 | DOH | jel472020@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 160 ASSAY: Xpert MTB/RIF Ultra, COUNT: 498 | 2023-03-27 | GX4 | 831610 | For module replacement A4 754334 - 210062953 | Module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | A1 and A4 operational, A2 and A3 reported |  | 900 0491rs | 9801 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569423828324564765/5569423828324564765_signature_28.png | Jasmin A. magtalas | https://www.jotform.com/uploads/macarewecare/222927978372067/5569423828324564765/5569423828324564765_signature_31.png
| 2023-04-10 | 2023-03-30 11:56 | Bagomg pag asa health center | 35 rd. 9 cor. Road 10 brgy bagong pag asa, quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 24145 | DOH | sisonborjalnora@gmail.com | 25 |  | 2023-03-27 | Gx4 | 831507 | For module replacement A2 754335 - 914754 | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900-0491 rs | 9811 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569421698328679196/5569421698328679196_signature_28.png | Nora B. Sison | https://www.jotform.com/uploads/macarewecare/222927978372067/5569421698328679196/5569421698328679196_signature_31.png
| 2023-04-10 | 2023-03-29 13:41 | FEU NRMF | Dahlia ave, Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24144 | CONSORTIUM | kahlilsilvestre92@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 28 ASSAY: Xpert MTB/RIF Ultra, COUNT: 334 | 2023-03-28 | GX4 | 840975 | For maintenance and xpertcheck | Performed xpertcheck and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and xpertcheck | All module passed |  | Xpertcheck -ce-5 | 9819 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569420198325194198/5569420198325194198_signature_28.png | Rienalyn Taripe | https://www.jotform.com/uploads/macarewecare/222927978372067/5569420198325194198/5569420198325194198_signature_31.png
| 2023-04-10 | 2023-04-05 08:29 | Josefina belmonte duran memorial district hospital | Purok 2 tuburan ligao city albay | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24143 | PBSP | johnrayds.22@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 252 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1877 | 2023-04-05 | GX4 | 814371 | For module replacement B3 210068728 - 210058809 | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | B2,b4 disable b1,b3 operational |  | 9000491 rs | O |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569418918327702070/5569418918327702070_signature_28.png | JOHNRAY DE LOS SANTOS | https://www.jotform.com/uploads/macarewecare/222927978372067/5569418918327702070/5569418918327702070_signature_31.png
| 2023-04-10 | 2023-04-05 10:37 | BSLSNRL | Rizal st,legaspi city | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24142 | PBSP | rbalzaga.pbsp@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 716 ASSAY: Xpert MTB/RIF, COUNT: 179 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2042 ASSAY: Xpert HIV-1 Viral Load, COUNT: 163 | 2023-04-05 | GX4 | 807579 | For module replacement b3 686316 - 629180 | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 9000491 rd | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5569417748325420849/5569417748325420849_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5569417748325420849/5569417748325420849_signature_31.png
| 2023-04-07 | 2023-04-04 14:54 | Baao RHU | Baao Camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24141 | PBSP | baaomho@gmail.com | 25 |  | 2023-04-05 | GX4 | 808639 | For module replacement and maintenance B2 757187 - 210005380 B3 667607 - 210008802 B4 646599 - 210009609 | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Unit operational |  |  | 900 0491rs |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5566631648012304989/5566631648012304989_signature_28.png | RHU BAAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5566631648012304989/5566631648012304989_signature_31.png
| 2023-04-07 | 2023-04-04 14:54 | Baao RHU | Baao Camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24140 | PBSP | baaomho@gmail.com | 25 |  | 2023-04-05 | GX4 | 808639 | For module replacement and maintenance B2 757187 - 210005380 B3 667607 - 210008802 B4 646599 - 210009609 | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Unit operational |  |  | 900 0491rs |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5566631088015391711/5566631088015391711_signature_28.png | RHU BAAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5566631088015391711/5566631088015391711_signature_31.png
| 2023-04-07 | 2023-04-07 11:03 | The Medical City South Luzon | United Blvd., Sta Rosa | 4A | Laguna | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24139 | COMMERCIAL | tmcsl-molecularlab@themedicalcity.com | 25 | ASSAY: Xpert HBV VL, COUNT: 7 ASSAY: Xpert HPV, COUNT: 6 ASSAY: Xpert MTB/RIF, COUNT: 39 ASSAY: Xpert MTB/RIF Ultra, COUNT: 149 ASSAY: Xpert HIV-1 Viral Load, COUNT: 9 |  | GXIV | 839233 | Unit for scheduled Preventive Maintenance. | Performed preventive maintenance. Generated reports. | Preventive Maintenance. | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5566462953014978367/5566462953014978367_signature_28.png | Ella Asa | https://www.jotform.com/uploads/macarewecare/222927978372067/5566462953014978367/5566462953014978367_signature_31.png
| 2023-04-06 | 2023-04-05 11:05 | Iligan Society of Internists | Iligan City | 10 | Lanao del Norte | TB DOTS | YES | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24138 | PBSP | dayolaermajessica@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2693 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 |  | GeneXpert | 805090 | H/W Fail on Module C2/769626 | Replaced module New module SN/644791 | Module Replacement | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5565677650818681111/5565677650818681111_signature_28.png | Erma Jessica M. Dayola | https://www.jotform.com/uploads/macarewecare/222927978372067/5565677650818681111/5565677650818681111_signature_31.png
| 2023-04-05 | 2023-04-05 16:02 | Escalanate City Health Office | Escalante, Negros Occidental | 6 | Negros Occidental | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24137 | DOH | sheila.negrido@yahoo.co | 30 | ASSAY: Xpert MTB/RIF, COUNT: 446 |  | Gx4 | 838523 | Module SN 707593 - Error 4009 Module Sn 862542 - Error 4012 | Module replacement | Module replacement with the following new modules: Module 1 SN # - 624898 Module 3 SN #- 710947 | All Modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564915847222400757/5564915847222400757_signature_28.png | Sheila Mae Negrido | https://www.jotform.com/uploads/macarewecare/222927978372067/5564915847222400757/5564915847222400757_signature_31.png
| 2023-04-05 | 2023-04-05 13:02 | Chinese General Hospital | Blumentritt, Sta. Cruz | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24136 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 10 ASSAY: Xpert CT/NG, COUNT: 27 ASSAY: Xpert HBV VL, COUNT: 50 ASSAY: Xpert HPV, COUNT: 27 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1271 ASSAY: Xpert HCV VL, COUNT: 11 ASSAY: Xpert HIV-1 Viral Load, COUNT: 17 |  | GXIV | 835604 | Unit for PM. For activation of calibration. | Performed PM. Performed activation of calibration. Generated reports. | PM and activation. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564810960748928504/5564810960748928504_signature_28.png | Maria Odessa Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/5564810960748928504/5564810960748928504_signature_31.png
| 2023-04-05 | 2023-04-05 12:13 | Carigara District Hospital | Carigara Leyte | 8 | Leyte | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24135 | PBSP | chezcairo@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 633 ASSAY: Xpert MTB/RIF Ultra, COUNT: 186 | 2023-04-05 | GXIV | 807619 | Module A1 for replacement | - Replace defective module A1/775309 , new module SN: 765577 - Performed functionality test - Generate Reports | #NAME? | Unit is Operational |  | GeneXpert Module 6 Color | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564784016822166171/5564784016822166171_signature_28.png | charisse ragrag | https://www.jotform.com/uploads/macarewecare/222927978372067/5564784016822166171/5564784016822166171_signature_31.png
| 2023-04-05 | 2023-04-05 11:06 | A.H. Lacson Health Center | Sta. Ana Hospital | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  |  |  |  | 24134 | PBSP | acg_medtech@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: NA |  | GX XVI | 839292 | Unit for installation (from Sta. Ana Hospital) | #NAME? | Installed unit | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564771804469892790/5564771804469892790_signature_28.png | Agnes Ranches | https://www.jotform.com/uploads/macarewecare/222927978372067/5564771804469892790/5564771804469892790_signature_31.png
| 2023-04-05 | 2023-04-05 11:02 | A.H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION | CEPHEID | Ronnel Belizar |  |  |  |  | 24133 | PBSP | acg_medtech@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 241 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1704 | 2023-04-05 | GXIV | 805067 | Unit for calibration. | #NAME? | Performed XpertCheck | Unit is operational. A1: passed A2: failed A3: passed A4: passed |  | XpertCheck Kit | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564770794467074514/5564770794467074514_signature_28.png | Agnes Ranches | https://www.jotform.com/uploads/macarewecare/222927978372067/5564770794467074514/5564770794467074514_signature_31.png
| 2023-04-05 | 2023-04-05 08:48 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | COVID Laboratory | NO | Unit transfer | CEPHEID | Ronnel Belizar |  |  |  |  | 24132 | PBSP | atumbaga1079@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 466 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1800 ASSAY: Xpert SARS-CoV-2, COUNT: 4968 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GXIV | 839292 | Unit for transfer to A.H. Lacson Health Center | - disassembled unit * GeneXpert GX XVI with LAN cable and power cord * HP Monitor w/ power cord & VGA cable (CNK94302FS) * Lenovo Desktop w/ power cord, wireless keyboard and mouse (CPH0520F08995) * barcode scanner w/ USB to LAN cable and stand (20052010551944) - 1 KVA UPS (AT1000190402ST0019) * HP printer w/ power cord and USB cable (VNC7703219) * ethernet adapter - unit picked up by PBSP | Disassembled system ready for PBSP pick up | Unit is operational. Ready for transfer. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564729224467383523/5564729224467383523_signature_28.png | Allan Tumbaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5564729224467383523/5564729224467383523_signature_31.png
| 2023-04-05 | 2023-04-05 10:35 | Bansalan Rural Health Unit | Poblacion Dos, Bansalan Davao Del Sur | 11 | Davao del Sur | TB DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24131 | PBSP | ellen68nebrada@gmail.com | 28.6 |  |  | GX-IV 4D | 838181 | Module B1 for Replacement | Replaced the Defective Module B1 SN: 807129 with new Module SN: 768092 | Perform Module Replacement | All Modules are Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564725853417059500/5564725853417059500_signature_28.png | Ellen Nebrada Mayor RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5564725853417059500/5564725853417059500_signature_31.png
| 2023-04-05 | 2023-04-05 08:48 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | COVID Laboratory | NO | Unit transfer | CEPHEID | Ronnel Belizar |  |  |  |  | 24130 | PBSP | atumbaga1079@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 466 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1800 ASSAY: Xpert SARS-CoV-2, COUNT: 4968 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GXIV | 839292 | Unit for transfer to A.H. Lacson Health Center | - disassembled unit * GeneXpert GX XVI with LAN cable and power cord * HP Monitor w/ power cord & VGA cable (CNK94302FS) * Lenovo Desktop w/ power cord, wireless keyboard and mouse (CPH0520F08995) * barcode scanner w/ USB to LAN cable and stand (20052010551944) - 1 KVA UPS (AT1000190402ST0019) * HP printer w/ power cord and USB cable (VNC7703219) * ethernet adapter - unit picked up by PBSP | Disassembled system ready for PBSP pick up | Unit is operational. Ready for transfer. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564708024462725828/5564708024462725828_signature_28.png | Allan Tumbaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5564708024462725828/5564708024462725828_signature_31.png
| 2023-04-05 | 2023-04-05 09:45 | Palompon RHU | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 24129 | DOH | luzvimindanavidad2@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 274 ASSAY: Xpert MTB/RIF Ultra, COUNT: 580 |  | GX4 | 838297 | None | None | #NAME? | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564687228112637052/5564687228112637052_signature_28.png | Luzviminda Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5564687228112637052/5564687228112637052_signature_31.png
| 2023-04-05 | 2023-03-31 14:12 | Caint City Health Office | Cainta | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 24126 | PBSP | rowenatec@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 22 ASSAY: Xpert MTB/RIF Ultra, COUNT: 705 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 |  | GX-IV-D | 838182 | Module B2 for troubleshooting | Performed Cepheid troubleshooting for module B2 | Performed maintenance Troubleshoot module A2 Generated reports | Unit is operational. Module B2/714059 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451102159342/5564679451102159342_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451102159342/5564679451102159342_signature_31.png
| 2023-04-05 | 2023-03-31 16:05 | Angono Rural Health Unit | Angono | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24128 | DOH | siennablanco@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 825 | 2023-03-20 | GX-IV-D | 819864 | Module A2 for module replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | Unit is operational. Module A4/719514 not operational, for reporting. |  | Old: A2/686363 New: A2/784784 | DR# 9728 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451102961503/5564679451102961503_signature_28.png | Rose Sienna Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451102961503/5564679451102961503_signature_31.png
| 2023-04-05 | 2023-03-30 14:31 | Department of Health Treatment and Rehabilitation Center - Bicutan | Camp Bagong Diwa | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24124 | PBSP | dohtrcbicutan.lab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 250 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1667 | 2023-03-27 | GX-IV-D | 807452 | Module C3/801912 error 4008 | Replaced gateway board | Performed maintenance Replaced gateway board Generated reports | Unit is operational. Module C3/801912 still not operational. |  | Gateway board | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451107568160/5564679451107568160_signature_28.png | April Chakee Javier | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451107568160/5564679451107568160_signature_31.png
| 2023-04-05 | 2023-03-31 14:17 | Cainta City Health Office | Caintq | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24127 | DOH | rowenatec@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 23 ASSAY: Xpert MTB/RIF Ultra, COUNT: 711 ASSAY: Xpert HIV-1 Viral Load, COUNT: 52 | 2023-03-27 | GX-IV-D | 831623 | Module A3 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | Unit is operational. |  | Old: A3/754520 New:A3/679622 | DR# 9803 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451108991116/5564679451108991116_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451108991116/5564679451108991116_signature_31.png
| 2023-04-05 | 2023-03-31 09:49 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | Relocation | CEPHEID | John Monsour Doctor |  |  |  |  | 24125 | COMMERCIAL | moleculardx@themedicalcity.com | 26 |  |  | GX-XVI-D-10C | 110015767 | For relocation | Transferred to other room | Transferred the GX XVI to other room Performed maintenance Generated reports | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451109313449/5564679451109313449_signature_28.png | Julie Rose Boyles | https://www.jotform.com/uploads/macarewecare/222927978372067/5564679451109313449/5564679451109313449_signature_31.png
| 2023-04-05 | 2023-03-29 13:36 | Santolan Super Health Center | Evangelista Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24123 | PBSP | samandlet17@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 276 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3551 | 2023-03-27 | GX-IV-D | 838898 | Module B4 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | Unit is operational. Module B3/882702 not operational, reported already. | UPS shutting down. Recommend to replace. | Old: B4/738573 New: B4/619237 | N/A |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5564678471107276205/5564678471107276205_signature_28.png | LOLITA FORMALEJO | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5564678471107276205/5564678471107276205_signature_31.png
| 2023-04-05 | 2023-03-29 10:29 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Molecular | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 24122 | COMMERCIAL | moleculardx@themedicalcity.com | 29 | ASSAY: Xpert HBV VL, COUNT: 12 ASSAY: Xpert HPV, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 91 ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 |  | GX-XVI-D-10C | 110015767 | Module B3/210082269 has errors 4017 and 4009 | Checked module B3 | Performed maintenance Checked module B3 Generated reports Disabled Module B3 | Unit is operational. Moduke B3/210082269 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564678471108918637/5564678471108918637_signature_28.png | PAOLO CRISPIN | https://www.jotform.com/uploads/macarewecare/222927978372067/5564678471108918637/5564678471108918637_signature_31.png
| 2023-04-05 | 2023-03-30 11:52 | North Daang Hari Super Health Center | North Daang Hari | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24121 | PBSP | onlychumz@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1708 | 2023-03-27 | GX-IV-D | 804702 | Module A3 for replacement | Replaced module A3 | Performed maintenance Replaced defective module Generated reports | All modules operational. |  | Old: A3/731867 New: A3/784192 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564678471101701576/5564678471101701576_signature_28.png | Charmaine Trinidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5564678471101701576/5564678471101701576_signature_31.png
| 2023-04-05 | 2023-03-29 15:31 | Sumilang Super Health Center | Dr. Garcia St, Sumilang | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24119 | PBSP | maricor.obnamia@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1341 | 2023-03-27 | GX-IV-D | 814163 | Module B1 for replacement. | Replaced defective module | Performed maintenance Replaced defective module Generated reports | Unit is operational. |  | Old: B1/601852 New: B1/637354 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564678471106866294/5564678471106866294_signature_28.png | Maria Corazon D. Obnamia | https://www.jotform.com/uploads/macarewecare/222927978372067/5564678471106866294/5564678471106866294_signature_31.png
| 2023-04-05 | 2023-03-29 10:35 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Molecular | NO | Pull out | CEPHEID | John Monsour Doctor |  |  |  |  | 24120 | COMMERCIAL | moleculardx@themedicalcity.com | 29 |  |  | GX-IV-L | 802117 | For pull out | Pulled out the GeneXpert system | Pulled out the GeneXpert unit, laptop and barcode scanner | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564678471108288157/5564678471108288157_signature_28.png | Paolo Crispin | https://www.jotform.com/uploads/macarewecare/222927978372067/5564678471108288157/5564678471108288157_signature_31.png
| 2023-04-05 | 2023-04-05 07:17 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB Dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24118 | DOH | marteetingzon20@gmail.com | 26 |  | 2023-04-05 | GXIV | 838421 | #NAME? | - Replace defective modules B2/675662, B3/861170 - new Module SN: B2/617811, B3/210005379 - Performed preventive maintenance - Performed functionality test - Generate Reports | #NAME? | All modules are in good condition |  | 2 - GeneXpert Module 6 Color | DOH |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564658130814924518/5564658130814924518_signature_28.png | Martee H. Tingzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5564658130814924518/5564658130814924518_signature_31.png
| 2023-04-05 | 2023-03-30 13:59 | Quezon City Jail | K-3rd St. Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 24117 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: NA | 2023-03-30 | GXIV | 803089 | Modules A3(658318), A4(675322) for replacement. | #NAME? | Replaced A3, A4 | Unit is operational. |  | GeneXpert Module SN: 662287, 661626 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564648904463934267/5564648904463934267_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5564648904463934267/5564648904463934267_signature_31.png
| 2023-04-04 | 2023-04-04 12:29 | Jose B. Lingad Memorial General Hospital | San Fernando, Pampanga | 3 | Pampanga | Pathology | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 24116 | COMMERCIAL | jblmgh.is@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1180 ASSAY: Xpert HCV VL, COUNT: 26 ASSAY: Xpert HBV VL, COUNT: 78 | 2023-04-03 | GXIV | 812571 | Module A2 (684507) for replacement | #NAME? | Replaced A2 | Unit is operational. |  | GeneXpert Module SN: 889844 | 9869 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564011882227822220/5564011882227822220_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5564011882227822220/5564011882227822220_signature_31.png
| 2023-04-04 | 2023-04-04 14:43 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24114 | DOH | elma.bragat@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1000 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2200 ASSAY: Xpert HIV-1 Viral Load, COUNT: 135 | 2023-03-24 | GX-IV | 838508 | FOR MODULE B3 REPLACEMENT | #NAME? | -PERFORM MODULE B3 DEFECTIVE SN: 862394 REPLACEMENT WITH NEW SN: 695932 -PERFORM PM AND CLEANING | EQUIPMENT IS OPERATIONAL | #NAME? | GX MODULE 6-COLOR | 9790 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564011739223054433/5564011739223054433_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5564011739223054433/5564011739223054433_signature_31.png
| 2023-04-04 | 2023-04-04 14:38 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | INSTALLATION | CEPHEID | Joseph Acebedo |  |  |  |  | 24113 | PBSP | elma.bragat@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1000 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2200 ASSAY: Xpert HIV-1 Viral Load, COUNT: 135 |  | GX-IV | 815275 | MONITOR REPLACEMENT | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564011739221917776/5564011739221917776_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5564011739221917776/5564011739221917776_signature_31.png
| 2023-04-04 | 2023-04-04 14:50 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24115 | DOH | elma.bragat@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1000 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2200 ASSAY: Xpert HIV-1 Viral Load, COUNT: 135 | 2023-03-28 | GX-IV | 838510 | MODULE C2 AND C3 REPLACEMENT | #NAME? | -PERFORM MODULE C2 DEF SN: 862256 REPLACMENT WITH NEW SN: 685729 -PERFORM MODULE C3 DEF SN: 862392 REPLACMENT WITH NEW SN: 210015791 -PERFORM PM AND CLEANING | EQUIPMENT IS OPERATIONAL | MODULE C4 DISABLED | GX MODULE 6-COLOR | 9836, 9837 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564011739229697826/5564011739229697826_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5564011739229697826/5564011739229697826_signature_31.png
| 2023-04-04 | 2023-04-04 14:37 | Masinloc Rural Health Unit | Sta. Lucia St. South Poblacion Masinloc, Zambales | 3 | Zambales | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24112 | DOH | rodenne@yahoo.com | 26.5 | ASSAY: Xpert MTB/RIF, COUNT: 255 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1738 | 2023-02-01 | GXIV-D | 831608 | A3:754353 for replacement Due for calibration | Preventive maintenance Calibration Module replacement | Performed surface cleaning Performed module replacement Archived test results Performed preventive maintenance Performed calibration Gathered reports Performed functionality test Installtion of HIV assay | Unit is operational | All modules passed calibration | 1 genexpert module, 3 xpertcheck kits | 9764 and 9561 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5564009039414144675/5564009039414144675_signature_28.png | Rodenn Eduvala | https://www.jotform.com/uploads/macarewecare/222927978372067/5564009039414144675/5564009039414144675_signature_31.png
| 2023-04-04 | 2023-04-04 14:18 | Lanuza Health Center | Lanuza, Sta Cruz | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24111 | PBSP | joneevamarie@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1341 | 2023-04-04 | GXIV | 805111 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports | PM and Xpertcheck. | Unit is operational. Module B1 (630369) , service required. |  | Xpertcheck kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563986320742286775/5563986320742286775_signature_28.png | Eva Marie C. Jone | https://www.jotform.com/uploads/macarewecare/222927978372067/5563986320742286775/5563986320742286775_signature_31.png
| 2023-04-04 | 2023-04-04 08:55 | Dimasalang Health Center | Dimasalang, Sta. Cruz | 4A | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24110 | PBSP | kikokendz_060830@yahoo.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 ASSAY: Xpert MTB/RIF, COUNT: 177 ASSAY: Xpert MTB/RIF Ultra, COUNT: 814 | 2023-04-04 | GXIV | 707072 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generated reports. | PM and Xpertcheck. | Unit is operational. Module B2 (693886) and B4 (634018) , service required. |  | Xpertcheck kit | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563986320749559679/5563986320749559679_signature_28.png | Maria joy cleofas . Abenales | https://www.jotform.com/uploads/macarewecare/222927978372067/5563986320749559679/5563986320749559679_signature_31.png
| 2023-04-04 | 2023-04-04 11:00 | Philippine Red Cross - Ormoc Chapter | Ormoc City, Leyte | 8 | Leyte | Bloodbank | NO | REPAIR | HAEMONETICS | Wynnona Reoma |  |  |  |  | 24109 | COMMERCIAL | cochingvenus@gmail.com | 20 |  |  | MCS+ | 15H043G | Rubber rollers for replacement | Replace AC, blood, and transfer pumps with new rubber rollers (6 pcs) | Replace AC, blood, and transfer pumps with new rubber rollers (6 pcs) | Unit is operational. | Pull-out old AC, blood, and transfer pumps |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563981095311465882/5563981095311465882_signature_28.png | Venus D.Coching | https://www.jotform.com/uploads/macarewecare/222927978372067/5563981095311465882/5563981095311465882_signature_31.png
| 2023-04-04 | 2023-04-04 11:00 | Philippine Red Cross - Ormoc Chapter | Ormoc City, Leyte | 8 | Leyte | Bloodbank | NO | REPAIR | HAEMONETICS | Wynnona Reoma |  |  |  |  | 24108 | COMMERCIAL | cochingvenus@gmail.com | 20 |  |  | MCS+ | 15H043G | Rubber rollers for replacement | Replace AC, blood, and transfer pumps with new rubber rollers (6 pcs) | Replace AC, blood, and transfer pumps with new rubber rollers (6 pcs) | Unit is operational. | Pull-out old AC, blood, and transfer pumps |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563980925313403662/5563980925313403662_signature_28.png | Venus D.Coching | https://www.jotform.com/uploads/macarewecare/222927978372067/5563980925313403662/5563980925313403662_signature_31.png
| 2023-04-04 | 2023-04-04 10:34 | Presentacion RHU | Presentacion ,Camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24106 | PBSP | tabo.juliet@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 509 ASSAY: Xpert MTB/RIF, COUNT: 71 |  | GX 2 | 834808 | For module replacement B1 693691 - 643064 B2 643219 - 699326 | Performed module replacement Cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974297122498035/5563974297122498035_signature_28.png | Julieta Tabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974297122498035/5563974297122498035_signature_31.png
| 2023-04-04 | 2023-03-30 09:35 | Bernadette Bautista medical clinic | 7 Bf road,Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24107 | CONSORTIUM | richeclevia@gmail.com | 28 |  | 2023-03-28 | GX4 | 835603 | For maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit and performed xpertcheck | Maintenance and xpertcheck | All module passed xpertcheck |  | Xpertcheck-ce-5 | 9817 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974307125742951/5563974307125742951_signature_28.png | Richard eclevia | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974307125742951/5563974307125742951_signature_31.png
| 2023-04-04 | 2023-04-04 10:34 | Presentacion RHU | Presentacion ,Camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24105 | PBSP | tabo.juliet@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 509 ASSAY: Xpert MTB/RIF, COUNT: 71 |  | GX 2 | 834808 | For module replacement B1 693691 - 643064 B2 643219 - 699326 | Performed module replacement Cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974237125278187/5563974237125278187_signature_28.png | Julieta Tabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974237125278187/5563974237125278187_signature_31.png
| 2023-04-04 | 2023-03-30 09:35 | Bernadette Bautista medical clinic | 7 Bf road,Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24104 | CONSORTIUM | richeclevia@gmail.com | 28 |  | 2023-03-28 | GX4 | 835603 | For maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit and performed xpertcheck | Maintenance and xpertcheck | All module passed xpertcheck |  | Xpertcheck-ce-5 | 9817 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974237129360035/5563974237129360035_signature_28.png | Richard eclevia | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974237129360035/5563974237129360035_signature_31.png
| 2023-04-04 | 2023-03-30 09:35 | Bernadette Bautista medical clinic | 7 Bf road,Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24102 | CONSORTIUM | richeclevia@gmail.com | 28 |  | 2023-03-28 | GX4 | 835603 | For maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit and performed xpertcheck | Maintenance and xpertcheck | All module passed xpertcheck |  | Xpertcheck-ce-5 | 9817 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974177121781581/5563974177121781581_signature_28.png | Richard eclevia | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974177121781581/5563974177121781581_signature_31.png
| 2023-04-04 | 2023-04-04 10:34 | Presentacion RHU | Presentacion ,Camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 24103 | PBSP | tabo.juliet@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 509 ASSAY: Xpert MTB/RIF, COUNT: 71 |  | GX 2 | 834808 | For module replacement B1 693691 - 643064 B2 643219 - 699326 | Performed module replacement Cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974177128162152/5563974177128162152_signature_28.png | Julieta Tabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5563974177128162152/5563974177128162152_signature_31.png
| 2023-04-04 | 2023-04-04 13:33 | Kidapawan City Health Office | Quezon Blvd, Kidapawan City | 12 | Cotabato | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24101 | DOH | manz.franz@yahoo.com | 25.9 | ASSAY: Xpert MTB/RIF, COUNT: 319 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2043 |  | GX-IV 4d | 838545 | Stuck Cartridge and For Preventive Maintenance | Aligned the valve position of the Module B1 to prevent further stuck cartridge issue, Clean fan filters, Cartridge Bay, Plunger Rod, Icore Lenses, Generate IQR, Assay Stat and System Log | Check the valve position of the module and perform preventive maintenance | Module B3 and B4 disabled due to H/W |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563962625222783697/5563962625222783697_signature_28.png | Ma. Luisa Francisco | https://www.jotform.com/uploads/macarewecare/222927978372067/5563962625222783697/5563962625222783697_signature_31.png
| 2023-04-04 | 2023-04-04 11:23 | Lugait Rural Health Unit | Lugait | 10 | Misamis Oriental | TB DOTS | NO | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24100 | DOH | rodelynlimocon@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 739 | 2023-03-28 | GeneXpert | 838562 | Error 4012 on module B2/691477 | Replaced module new module SN/690358 | Module Replacement Preventive Maintenance Generate report data files Database back up | Operational |  | GeneXpert Module 6 color SN/690358 | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563885567211201241/5563885567211201241_signature_28.png | Rodelyn Limocon | https://www.jotform.com/uploads/macarewecare/222927978372067/5563885567211201241/5563885567211201241_signature_31.png
| 2023-04-04 | 2023-04-04 09:55 | San Leonardo Rural Health Unit | Diversion, San Leonardo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24099 | PBSP | ligayadomingo20@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 190 ASSAY: Xpert MTB/RIF Ultra, COUNT: 728 | 2023-04-04 | Genexpert IV | 846042 | UPS no backup | Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test Archive Test Result 2021-2022 | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 8pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563869865911542109/5563869865911542109_signature_28.png | Ligaya Domimgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5563869865911542109/5563869865911542109_signature_31.png
| 2023-04-04 | 2023-04-04 10:28 | San Antonio Rural Health Unit | Brgy. Rizal San Antonio, Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24098 | PBSP | markanthonyyambao@yahoo.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 448 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1165 | 2023-03-10 | GXIV-D | 814377 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration | 4 Xpertcheck Kits | 39477 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563858931224076300/5563858931224076300_signature_28.png | Mark Anthony S. Yambao | https://www.jotform.com/uploads/macarewecare/222927978372067/5563858931224076300/5563858931224076300_signature_31.png
| 2023-04-04 | 2023-04-04 10:01 | Parang Rural Health Unit | Parang, Maguindanao | BARMM | Maguindanao | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24097 | PBSP | jiina26untal@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 63 ASSAY: Xpert MTB/RIF Ultra, COUNT: 479 | 2023-04-04 | Gx-IV | 838135 | Module 2 for service required | Perform Genexpert machine Preventive maintenance Replace module 2 SN: 866870 with new module SN: 714948 | Perform Genexpert machine Preventive maintenance Replace module 2 SN: 866870 with new module SN: 714948 | 4 operational modules |  | Genexpert module 6 colors | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563834533773085484/5563834533773085484_signature_28.png | Johara Untal | https://www.jotform.com/uploads/macarewecare/222927978372067/5563834533773085484/5563834533773085484_signature_31.png
| 2023-04-04 | 2023-04-04 09:23 | Cebu TB reference Lab | Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 24096 | DOH | asmantiquilla.pbsp@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 1280 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3216 ASSAY: Xpert SARS-CoV-2, COUNT: 99 ASSAY: Xpert HIV-1 Viral Load, COUNT: 126 |  | GXIV | 838263 | Defective module 3 and 4 | New module installed in module 3 sn 616725and 4 sn220004971 | New module installed in module 3 sn 616725and 4 sn220004971 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563811132311397452/5563811132311397452_signature_28.png | ALLAN S. MANTIQUILLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5563811132311397452/5563811132311397452_signature_31.png
| 2023-04-04 | 2023-04-04 08:42 | Hilongos Rural Health Unit | Hilongos Leyte | 8 | Leyte | TB dots | NO | REPAIR CALIBRATION | CEPHEID | John Paul Labutap |  |  |  |  | 24095 | PBSP | leovillaflor@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 262 ASSAY: Xpert MTB/RIF Ultra, COUNT: 624 | 2023-04-04 | GXIV | 838143 | #NAME? | - Replace module B2 (806815) - new module SN: 210073300 - Run Xpertcheck on B1 -Performed preventive maintenance - Generate Reports | #NAME? | Unit is Operational | #NAME? | - GeneXpert Module 6 Color - 1 xpc cartridge | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563791383413838771/5563791383413838771_signature_28.png | Leonides Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5563791383413838771/5563791383413838771_signature_31.png
| 2023-04-04 | 2023-04-04 08:12 | Olongapo City Health Office | Rizal St. West Bajac-bajac Olongapo City, Zambales | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24094 | PBSP | roselynmaganomilan@gmail.com | 29 |  |  | GXIV-D | 804700 | D1 for re-calibration Error 5006 Not genarating calibration info and failed module report D2 unavailable for calibration | Troubleshooting Calibration Preventive maintenance | Performed preventive maintenance Archived test results Performed calibration Gathered reports Performed functionality test | Unit is operational except D1 | For reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563772181225920475/5563772181225920475_signature_28.png | Roselyn M. Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5563772181225920475/5563772181225920475_signature_31.png
| 2023-04-03 | 2023-04-03 14:40 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Molecular laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24093 | PBSP | tindocvelche@gmail.com | 22.8 | ASSAY: Xpert SARS-CoV-2, COUNT: 5788 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 | 2023-04-03 | Gx-IV | 834442 | Module 1 and 2 H/W failed | Perform Genexpert machine Preventive maintenance Replace module 1 SN:66088 with new module SN: 733420 | Perform Genexpert machine Preventive maintenance Replace module 1 SN:66088 with new module SN: 733420 | 3 operational modules |  | Genexpert module 6 colors | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563140147712312005/5563140147712312005_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5563140147712312005/5563140147712312005_signature_31.png
| 2023-04-03 | 2023-03-31 14:48 | Amparo Health Center | Amparo, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24092 | PBSP | jjpsjasperjoyce@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1246 | 2023-03-31 | GXIV | 838873 | Module B1 (769519) for replacement | #NAME? | Replaced B1 | Unit is operational. |  | GeneXpert Module SN: 769519 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563136231215837821/5563136231215837821_signature_28.png | Jasper Joyce Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/5563136231215837821/5563136231215837821_signature_31.png
| 2023-04-03 | 2023-03-31 11:29 | San Lorenzo Ruiz Women's Hospital | Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24089 | PBSP | regielyncamu@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 688 | 2023-03-31 | GXIV | 805009 | Module B1 (607475) for replacement | #NAME? | Replaced B1 | Unit is operational. |  | GeneXpert Module SN: 662289 | NQ |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911215781489/5563132911215781489_signature_28.png | Regielyn Camu | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911215781489/5563132911215781489_signature_31.png
| 2023-04-03 | 2023-03-29 11:53 | Mandaluyong City Health Office | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 24090 | PBSP | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2644 ASSAY: Xpert HIV-1 Viral Load, COUNT: 124 | 2023-03-29 | GXIV | 805062 | Module A1 (601438) for replacement | #NAME? | Replaced A1 | Unit is operational. |  | GeneXpert Module SN: 661307 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911217228679/5563132911217228679_signature_28.png | Teresita Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911217228679/5563132911217228679_signature_31.png
| 2023-04-03 | 2023-03-31 10:04 | Bagong Barrio Health Center | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 24091 | DOH | cruzernitamay@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 38 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1312 | 2023-03-20 | GXIV | 831906 | Module A3 (653357) for replacement. | #NAME? | Replaced A3 | Unit is operational. |  | GeneXpert Module SN: 738260 | 9706 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911219932187/5563132911219932187_signature_28.png | Ernita May Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911219932187/5563132911219932187_signature_31.png
| 2023-04-03 | 2023-03-29 14:22 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24087 | PBSP | itee@loveyourself.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 154 ASSAY: Xpert MTB/RIF Ultra, COUNT: 198 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1254 | 2023-03-29 | GXIV | 808649 | Module B1 (667732) for replacement. | #NAME? | Replaced module B1 | Unit is operational. |  | GeneXpert Module SN: 635671 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911215800058/5563132911215800058_signature_28.png | Ivan Tee | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911215800058/5563132911215800058_signature_31.png
| 2023-04-03 | 2023-03-29 10:42 | A.H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24088 | PBSP | dinajaen1990@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 281 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1628 | 2023-03-29 | GXIV | 805067 | Modules A3 (643151), A4 (681297) for replacement | #NAME? | Replaced modules A3, A4 | Unit is operational. |  | GeneXpert Module | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911219088362/5563132911219088362_signature_28.png | Dina Jaen | https://www.jotform.com/uploads/macarewecare/222927978372067/5563132911219088362/5563132911219088362_signature_31.png
| 2023-04-03 | 2023-04-03 13:53 | Abucay Rural Health Unit | Decena Ave. Salian Abucay, Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24086 | DOH | parreraledelina22176@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 169 ASSAY: Xpert MTB/RIF Ultra, COUNT: 295 | 2023-03-22 | GXIV-D | 831878 | A4 for moduke replacement | Preventive maintenance Module replacement | Performed preventive maintenance Performed surface cleaning Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 9754 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563108658228534200/5563108658228534200_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5563108658228534200/5563108658228534200_signature_31.png
| 2023-04-03 | 2023-04-03 12:21 | Samal Rural Health Unit | A. Bonifacio St. East Daan Bago Samal, Bataan | 3 | Bataan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24085 | DOH | vickytholits2002@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 204 ASSAY: Xpert MTB/RIF Ultra, COUNT: 438 | 2023-03-22 | GXIV-D | 831729 | A1:757531 for module replacement | Module replacement Preventive maintenance | Performed surface cleaning Performed oreventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 9755 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563053608222708441/5563053608222708441_signature_28.png | Maria Victoria P. Banzuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5563053608222708441/5563053608222708441_signature_31.png
| 2023-04-03 | 2023-04-03 11:27 | Cuyapo Rural Health Unit | District 1, Cuyapo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24083 | DOH | shane.verlaine@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 390 ASSAY: Xpert MTB/RIF Ultra, COUNT: 824 | 2023-02-07 | Genexpert IV | 831672 | Module 2 and 4 Temperature error | Performed Xpertcheck | Performed Maintenance Gathered Reports Performed Calibratiom Performed Functionality Test | Unit is Operational | Xpertcheck Result - All modules Passed | Xpertcheck 4pc | 9594 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026785913244711/5563026785913244711_signature_28.png | VERLAINE ASCENCION NERVI A. VERUECO | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026785913244711/5563026785913244711_signature_31.png
| 2023-04-03 | 2023-03-31 13:12 | Mexico Rural Health Unit I | Poblacion Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 24084 | DOH | jasperdizon20@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-03-22 | Genexpert IV | 831877 | Module 3 Error 1002 | Performed Xpertcheck on module 3 and 4 | Performed Maintenance Gathered Reports Performed Module Replacement (802416 - 661387) Performed Functionality Test | Unit is Operational | Xpertcheck result - A3 Failed & A4 - Passed | Xpertcheck 4pc & Genexpert Module (661387) | 9766 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026795914431349/5563026795914431349_signature_28.png | Jasper Dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026795914431349/5563026795914431349_signature_31.png
| 2023-04-03 | 2023-03-30 13:44 | Rizal Rural Health Unit | Poblacion Norte, Rizal, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24080 | PBSP | abigailestoque@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 60 ASSAY: Xpert MTB/RIF Ultra, COUNT: 412 | 2023-03-30 | Genexpert IV | 834691 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026775913790344/5563026775913790344_signature_28.png | Abigail Estoque | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026775913790344/5563026775913790344_signature_31.png
| 2023-04-03 | 2023-03-30 15:46 | Bongabon District Hospital | Brgy. Curva Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24082 | DOH | tsantuajr0128@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 129 ASSAY: Xpert MTB/RIF Ultra, COUNT: 250 | 2023-03-21 | Genexpert IV | 831562 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement (721744 - 758096) & (754108 - 692115) Performed Functionality Test | Unit is Operational | N/a | Genexpert Modules (768096 & 692115) | 9748 & 9749 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026785915624147/5563026785915624147_signature_28.png | Teodorico Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026785915624147/5563026785915624147_signature_31.png
| 2023-04-03 | 2023-03-28 13:40 | Cabanatuan City Health Office | Kapt. Pepe, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24081 | DOH | cholab2021@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1961 ASSAY: Xpert MTB/RIF, COUNT: 765 | 2023-03-22 | Genexpert IV | 831592 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement (784673 - 636666) Disabled Module A1 Performed Functionality Test | Unit is Operational | N/a | Genexpert Module - 636666 | 9767 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026785916668608/5563026785916668608_signature_28.png | Lea Veronica Samsom | https://www.jotform.com/uploads/macarewecare/222927978372067/5563026785916668608/5563026785916668608_signature_31.png
| 2023-04-03 | 2023-04-03 11:36 | Biliran Provincial Hospital | Naval, Biliran | 8 | Biliran | Tb Dots | YES | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 24079 | PBSP | ghelmission13@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 737 ASSAY: Xpert MTB/RIF Ultra, COUNT: 710 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 | 2023-04-03 | GX4 | 815273 | Module B1 defective | Replace defective module B1/736722 with new module SN: 220003268 | #NAME? | All modules are functional. | Waiting for monitor allocation. | 900 GX - 6 color module | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5563025992762358504/5563025992762358504_signature_28.png | Gheland Corto | https://www.jotform.com/uploads/macarewecare/222927978372067/5563025992762358504/5563025992762358504_signature_31.png
| 2023-04-03 | 2023-04-03 09:27 | Talakag Rural Health Unit | Talakag | 10 | Misamis Oriental | TB DOTS | NO | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24078 | DOH | mhotalakaglaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 194 | 2023-03-28 | GeneXpert | 838566 | Module B1 and B2 Failed during calibration | Replaced Module | Preventive Maintenance Module Replacement Generate report data files Database back up | Operational |  | GeneXpert Module 6-color SN/632850 GeneXpert Module 6-color SN/673860 | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5562960283618493545/5562960283618493545_signature_28.png | Khiza Tinoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5562960283618493545/5562960283618493545_signature_31.png
| 2023-04-03 | 2023-03-30 12:45 | Nasugbu Rural Health Unit | Nasugbu-Ternate Highway, Nasugbu | 4A | Batangas | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24076 | PBSP | lucenavicente15@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 861 |  | GXIV-D | 807581 | Module 1 with Lost Communication Error | Cepheid Troubleshooting | Performed Maintenance Troubleshoot Problem Gather Additional Data for Follow-up Reporting Generate Reports | 3 Modules are Operational, 1 Module Reported |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919301212122236/5562919301212122236_signature_28.png | Lucena I. Vicente | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919301212122236/5562919301212122236_signature_31.png
| 2023-04-03 | 2023-03-31 10:18 | Dasmariñas City Health Office I | D. Mangubat, Dasmariñas | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24077 | PBSP | cho1lab@yahoo.com | 23 | ASSAY: Xpert HPV, COUNT: 29 ASSAY: Xpert HIV-1 Viral Load, COUNT: 83 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2153 |  | GXIV-D | 838893 | Module 4 with Temperature Error | Module Replacement | Replaced Module 4 Performed Maintenance Performed Self-Test Generate Reports | All Modules are Operational | B4: 699489 -> 210036657 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919301219142491/5562919301219142491_signature_28.png | Donie P.Canteras | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919301219142491/5562919301219142491_signature_31.png
| 2023-04-03 | 2023-03-29 09:31 | Alabang Health Center | Alabang-Zapote Rd., Alabang | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24073 | DOH | cathycarasig@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 67 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1313 | 2023-03-27 | GXIV-D | 831913 | Module 4 with Temperature Error | Module Replacement | Replaced Module 4 Performed Maintenance Generate Reports | All Modules are Operational | A4: 893809-> 210015908 | GeneXpert Module 6-Colors | 9812 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201212635987/5562919201212635987_signature_28.png | Catherine R. Carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201212635987/5562919201212635987_signature_31.png
| 2023-04-03 | 2023-03-28 08:26 | Rosario Reyes Health Center | San Andres St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24075 | DOH | moninabatil16@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2235 | 2023-03-27 | GXIV-D | 831843 | Module 3 with Plunger Problem Unit is due for XpertCheck | Module Replacement XpertCheck | Performed Maintenance Replaced Module 3 Performed XpertCheck Generate Reports | 3 Modules Failed in XpertCheck, 1 Newly Replaced | B3: 757668->694064 | GX Module 6-Colors | 9804 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201212785580/5562919201212785580_signature_28.png | Monina B. Batil | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201212785580/5562919201212785580_signature_31.png
| 2023-04-03 | 2023-03-28 10:04 | M. Icasiano Health Center | Pedro Gil St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24071 | DOH | lorrainefabros@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1306 | 2023-03-27 | GXIV-D | 831684 | Module 2 has temperature errors Unit is due for XpertCheck | Module Replacement XpertCheck | Replaced Module 2 Performed Maintenance Performed XpertCheck Generate Reports | 3 Modules Failed in XpertCheck, 1 Newly Replaced | A2: 757053-> 210045104 | GeneXpert Module 6-Colors | 9797 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201215670939/5562919201215670939_signature_28.png | Fabros,Lorraine Lucila L. | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201215670939/5562919201215670939_signature_31.png
| 2023-04-03 | 2023-03-30 10:46 | Lemery Rural Health Unit | Ilustre Ave., Lemery | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24072 | PBSP | maritonie1912@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1052 |  | GXIV-D | 838179 | Module 2 with Temperature Errors Unit is due for XpertCheck | Module Replacement XpertCheck | Replaced Module 2 Performed Maintenance Performed XpertCheck Generate Reports | 2 Modules are operational, 2 Modules Reported | A2: 675429 -> 677307 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201216522816/5562919201216522816_signature_28.png | Maritonie Gadiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201216522816/5562919201216522816_signature_31.png
| 2023-04-03 | 2023-03-29 08:05 | Bagong Barangay Health Center | W. Zamora St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24074 | DOH | thea.peneyra@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 850 | 2023-03-27 | GXIV-D | 831694 | Module Failed in XpertCheck | Module Replacement | Module Replacement in Module 4 Performed Maintenance Generate Reports | 2 Modules are Operational, 2 Modules Reported | A4: 757241 -> 637824 | GeneXpert Module 6-Colors | 9793 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201218963422/5562919201218963422_signature_28.png | Maria Anthea S. Peneyra | https://www.jotform.com/uploads/macarewecare/222927978372067/5562919201218963422/5562919201218963422_signature_31.png
| 2023-04-02 | 2023-03-30 09:29 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24070 | DOH | burauenlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 97 ASSAY: Xpert MTB/RIF Ultra, COUNT: 209 | 2023-03-30 | GXIV | 838551 | N/a | #NAME? | Preventive Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5562511980825294611/5562511980825294611_signature_28.png | Mary Ann Sudaro | https://www.jotform.com/uploads/macarewecare/222927978372067/5562511980825294611/5562511980825294611_signature_31.png
| 2023-04-02 | 2023-03-31 10:04 | Bagong Barrio Health Center | Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 24068 | DOH | cruzernitamay@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 38 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1312 | 2023-03-20 | GXIV | 831906 | Module A3 (653357) for replacement. | #NAME? | Replaced A3 | Unit is operational. |  | GeneXpert Module SN: 738260 | 9706 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462223700423/5561874462223700423_signature_28.png | Ernita May Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462223700423/5561874462223700423_signature_31.png
| 2023-04-02 | 2023-03-29 10:42 | A.H. Lacson Health Center | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24067 | PBSP | dinajaen1990@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 281 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1628 | 2023-03-29 | GXIV | 805067 | Modules A3 (643151), A4 (681297) for replacement | #NAME? | Replaced modules A3, A4 | Unit is operational. |  | GeneXpert Module | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462226910609/5561874462226910609_signature_28.png | Dina Jaen | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462226910609/5561874462226910609_signature_31.png
| 2023-04-02 | 2023-03-31 11:29 | San Lorenzo Ruiz Women's Hospital | Malabon | NCR | NCR 3rd - Malabon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24069 | PBSP | regielyncamu@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 688 | 2023-03-31 | GXIV | 805009 | Module B1 (607475) for replacement | #NAME? | Replaced B1 | Unit is operational. |  | GeneXpert Module SN: 662289 | NQ |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462227067552/5561874462227067552_signature_28.png | Regielyn Camu | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462227067552/5561874462227067552_signature_31.png
| 2023-04-02 | 2023-03-29 14:22 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24067 | PBSP | itee@loveyourself.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 154 ASSAY: Xpert MTB/RIF Ultra, COUNT: 198 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1254 | 2023-03-29 | GXIV | 808649 | Module B1 (667732) for replacement. | #NAME? | Replaced module B1 | Unit is operational. |  | GeneXpert Module SN: 635671 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462229724222/5561874462229724222_signature_28.png | Ivan Tee | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462229724222/5561874462229724222_signature_31.png
| 2023-04-02 | 2023-03-29 11:53 | Mandaluyong City Health Office | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 24067 | PBSP | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2644 ASSAY: Xpert HIV-1 Viral Load, COUNT: 124 | 2023-03-29 | GXIV | 805062 | Module A1 (601438) for replacement | #NAME? | Replaced A1 | Unit is operational. |  | GeneXpert Module SN: 661307 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462229789458/5561874462229789458_signature_28.png | Teresita Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5561874462229789458/5561874462229789458_signature_31.png
| 2023-04-01 | 2023-03-31 13:53 | Malungon Rural Health Unit | Poblacion Malungon | 12 | Sarangani | TB-DOTS | NO | REPAIR TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24066 | DOH | madameerrmie@gmail.com | 25.8 |  |  | GX-4D | 838546 | System Unit is defective | Replace the old system unit SN: CPH0320F05932 with new System Unit SN: CPH0320F05906 Adjust the Avr Power output | Replaced the System Unit | Genexpert System unit is operational, Module B2 encounter multiple temperature error, data will be sent to cepheid for further assessment | Unstable voltage |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5561263117018331018/5561263117018331018_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5561263117018331018/5561263117018331018_signature_31.png
| 2023-04-01 | 2023-03-31 13:53 | Malungon Rural Health Unit | Poblacion Malungon | 12 | Sarangani | TB-DOTS | NO | REPAIR TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24065 | DOH | madameerrmie@gmail.com | 25.8 |  |  | GX-4D | 838546 | System Unit is defective | Replace the old system unit SN: CPH0320F05932 with new System Unit SN: CPH0320F05906 Adjust the Avr Power output | Replaced the System Unit | Genexpert System unit is operational, Module B2 encounter multiple temperature error, data will be sent to cepheid for further assessment | Unstable voltage |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5561262757012037661/5561262757012037661_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5561262757012037661/5561262757012037661_signature_31.png
| 2023-03-31 | 2023-03-31 15:11 | Ramos Rural Health Unit | Poblacion Ramos, Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24064 | DOH | jadeeh8143@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 121 ASSAY: Xpert MTB/RIF Ultra, COUNT: 250 |  | GXIV-D | 831721 | UPS for troubleshooting For Re-Xpertcheck | Preventive maintenace Calibration Troubleshooting | Performed UPS Troubleshooting Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | Used 4 xpertcheck kits, UPS for observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5560577864614477178/5560577864614477178_signature_28.png | Emily Jade Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/5560577864614477178/5560577864614477178_signature_31.png
| 2023-03-31 | 2023-03-31 11:23 | Benguet Provincial Health Office | La Trinidad, Benguet | CAR | Benguet | Laboratory | NO | CPU installation | CEPHEID | Harold Toco |  |  |  |  | 24063 | PBSP | benguetphortdl@gmail.com | 30 |  | 2023-03-31 | Gx4 | 807621 | For CPU replacement | Setup CPU | Performed CPU installation | All 4 modules operational |  | CPU S/N: CPH0322F32901 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5560430524816111139/5560430524816111139_signature_28.png | Frannie G. Calixto | https://www.jotform.com/uploads/macarewecare/222927978372067/5560430524816111139/5560430524816111139_signature_31.png
| 2023-03-31 | 2023-03-31 11:08 | Cabadbaran City Health Office | S. Curato St. Brgy. 5, Cabadbaran City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION | CEPHEID | Faizar Avenido |  |  |  |  | 24062 | DOH | samstercem@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF, COUNT: 441 ASSAY: Xpert MTB/RIF Ultra, COUNT: 381 |  | GX-4 | 838550 | 1 module found service required from XPC | None | Calibration and Preventive Maintenance | 3 modules working, 1 module for replacement (service required) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5560417654028483159/5560417654028483159_signature_28.png | Sammy A. Cempron | https://www.jotform.com/uploads/macarewecare/222927978372067/5560417654028483159/5560417654028483159_signature_31.png
| 2023-03-31 | 2023-03-31 10:35 | Soccsksargen General Hospital | Surallah | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24061 | PBSP | johnellamarisbautista@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 169 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1271 | 2023-03-31 | Gx-IV | 815432 | Module 3 for service required | Perform Genexpert machine Preventive maintenance Replace module 3 SN: 696322 with new module SN: 614553 | Perform Genexpert machine Preventive maintenance Replace module 3 SN: 696322 with new module SN: 614553 | 4 operational modules |  | Genexpert module 6-color | 614553 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5560402352323468165/5560402352323468165_signature_28.png | Johnella Maris Bautista, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5560402352323468165/5560402352323468165_signature_31.png
| 2023-03-31 | 2023-03-31 09:36 | San Pedro Rural Health Unit 2 (Langgam Health Center) | Brgy. Langgam, San Pedro | NCR | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24060 | DOH | leniemahalin@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 1028 ASSAY: Xpert MTB/RIF Ultra, COUNT: 992 ASSAY: Xpert HIV-1 Viral Load, COUNT: 127 | 2023-03-27 | GXIV | 821189 | Unit for PM. For module replacement of modules A3 and A4. | Performed PM. Performed module replacement on modules A3 and A4. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module A3 (601706) and A4 (883941) in replacement to defective A3 (783485) and A4 (610367). | DR 9806 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5560361978623098478/5560361978623098478_signature_28.png | Ma. Celenita Mahalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5560361978623098478/5560361978623098478_signature_31.png
| 2023-03-31 | 2023-03-31 08:34 | Baguio General Hospital | Baguo City | CAR | Benguet | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 24059 | PBSP | keithtagudar@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 353 ASSAY: Xpert MTB/RIF, COUNT: 192 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1330 | 2023-03-31 | Gx4 | 805212 | Module B2 for replacement. For maintenance | Replaced module B2: 602852 | Performed maintenance and module replacement. | All 4 modules operational |  | New mod 657183 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5560332261527084105/5560332261527084105_signature_28.png | Keith Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5560332261527084105/5560332261527084105_signature_31.png
| 2023-03-31 | 2023-03-30 13:22 | Tondo Health Center | Gagalangin | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24057 | DOH | ma.lorinamanera@yahoo.com | 25 |  | 2023-03-27 | GXIV | 831620 | Unit for PM. For module replacement of modules A2 and A3. | Performed PM. Performed module replacement on modules A2 and A3. | PM and module replacement. | Unit is operational. |  | Gx module A2 (210071179) and A3 (616329) in replacement to defective A2 (678005) and A3 (702275). | DR 9810 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5560313868625629902/5560313868625629902_signature_28.png | L0ntoc ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5560313868625629902/5560313868625629902_signature_31.png
| 2023-03-31 | 2023-03-30 13:53 | Atang dela Rama Health Center | Gagalangin | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24058 | DOH | monskie_22@yahoo.com | 28 |  | 2023-03-27 | GXIV | 831644 | Unit for PM. For module replacement of module B3. | Performed PM. Performed module replacement of B3. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B3 (614658) in replacement to defective B3 (756760). | DR 9798 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5560313868627415470/5560313868627415470_signature_28.png | ABDOn, EvELyn | https://www.jotform.com/uploads/macarewecare/222927978372067/5560313868627415470/5560313868627415470_signature_31.png
| 2023-03-30 | 2023-03-30 12:33 | Dinas Rural Health Unit | Purok Avocado, Don Jose, Dinas Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24055 | PBSP | bernantereynacres@gmail.com | 25 | ASSAY: Others, COUNT: 0 | 2023-03-30 | GX 4 | 838533 | System Unit unable to start up | Replaced System unit | Replaced System unit, Undergone Preventive Maintenence | Available and working | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559724309121059145/5559724309121059145_signature_28.png | REYNA CRES B. LAZAGA | https://www.jotform.com/uploads/macarewecare/222927978372067/5559724309121059145/5559724309121059145_signature_31.png
| 2023-03-30 | 2023-03-30 15:20 | Zamboanga del sur Medical Center | Dao, Pagadian City, Zamboanga drl Sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24056 | DOH | limboemillie@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1115 ASSAY: Xpert HIV-1 Viral Load, COUNT: 49 | 2023-03-30 | GX 4 | 838204 | Module 3 error 1002 | Underground Preventive Maintenence | Preventive Maintenence, checked if any heat source is near. Checked room temperature and module temperatures. | Module 1, 2 and 4 available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559724309126697874/5559724309126697874_signature_28.png | Emillie Limbo | https://www.jotform.com/uploads/macarewecare/222927978372067/5559724309126697874/5559724309126697874_signature_31.png
| 2023-03-30 | 2023-03-30 14:08 | Macabebe Rural Health Unit | Macabebe Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24054 | DOH | salvamari13@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 | 2023-03-22 | GXIV-D | 831893 | A2 and A3 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 2 genexpert module | 9744 and 9760 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559673022518413508/5559673022518413508_signature_28.png | Salva Sese | https://www.jotform.com/uploads/macarewecare/222927978372067/5559673022518413508/5559673022518413508_signature_31.png
| 2023-03-30 | 2023-03-30 14:08 | Macabebe Rural Health Unit | Macabebe Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24053 | DOH | salvamari13@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 | 2023-03-22 | GXIV-D | 831893 | A2 and A3 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 2 genexpert module | 9744 and 9760 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559672932514213201/5559672932514213201_signature_28.png | Salva Sese | https://www.jotform.com/uploads/macarewecare/222927978372067/5559672932514213201/5559672932514213201_signature_31.png
| 2023-03-30 | 2023-03-30 14:07 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 24052 | DOH | joydarlie72@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 958 ASSAY: Xpert MTB/RIF Ultra, COUNT: 670 ASSAY: Xpert SARS-CoV-2, COUNT: 19 |  | GX IV | 838527 | Due date for Preventive Maintenance | Copy data of IQR, system log, assay statistics | Clean Fan Filter | Only 2 modules functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559664014222018391/5559664014222018391_signature_28.png | Joy Ma. Darlie Labrado | https://www.jotform.com/uploads/macarewecare/222927978372067/5559664014222018391/5559664014222018391_signature_31.png
| 2023-03-30 | 2023-03-30 14:06 | Democrito O. Plaza Memorial Hospital | National Highway, Brgy. Patin-ay, Prosperidad | 13 | Agusan del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 24051 | PBSP | lermabpabillan@yahoo.com | 21.65 | ASSAY: Xpert MTB/RIF, COUNT: 487 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2806 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-4 | 805335 | Error 4004 found at Module B1 CPU doesn't booting up. Stuck Cartridge at Module B1 | None | Removing Stuck Cartridge Installation of Laptop (Acer E5-471) as temporary. Reporting of defective module and CPU | 1 module disabled and reported 3 modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559662624421380451/5559662624421380451_signature_28.png | Lerma B. Pabillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5559662624421380451/5559662624421380451_signature_31.png
| 2023-03-30 | 2023-03-30 14:04 | Escalante City Heath Office | Escalante, Poblacion Negros Occidental | 7 | Negros Occidental | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24050 | DOH | sheila.negrido@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 287 |  | Gx4 | 838523 | Module 1 and 3 temp error | Conduct Preventive Maintenance | Conduct preventive maintenance, Clean Icore, clean Plunger and clean filter | All modules are operational but the area need to be added for additional aircon unit or upgrade for Higher Power aircon |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559658329318185057/5559658329318185057_signature_28.png | Sheila Mae Negrido | https://www.jotform.com/uploads/macarewecare/222927978372067/5559658329318185057/5559658329318185057_signature_31.png
| 2023-03-30 | 2023-03-30 12:34 | Apalit Rural Health Unit | Apalit Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24049 | DOH | duenasblesilda@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 138 ASSAY: Xpert MTB/RIF Ultra, COUNT: 590 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 | 2023-03-22 | GXIV-D | 831900 | A3 and A4 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 2 genexpert module | 9772 and 9758 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559655442514918914/5559655442514918914_signature_28.png | Blesilda Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5559655442514918914/5559655442514918914_signature_31.png
| 2023-03-30 | 2023-03-30 13:48 | LAPULAPU CITY HEALTH OFFICE | LAPULAPU CITY | 7 | Cebu | TB DOTS | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24048 | DOH | tanya_pax@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1549 ASSAY: Xpert MTB/RIF, COUNT: 884 |  | GX-IV | 837661 | MODULE A2 AND A3 TEMP ERRORS | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | #NAME? |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559648956513308428/5559648956513308428_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5559648956513308428/5559648956513308428_signature_31.png
| 2023-03-30 | 2023-03-30 09:22 | Dona Remedios Rural Health Unit | Brgy. Pulong Sampaloc, Dona Remedios Trinidad, Bulacan | 3 | Bulacan | Laboratory | NO | Back up laptop installation | CEPHEID | Harold Toco |  |  |  |  | 24047 | PBSP | maricel.rmt24@gmail.com | 28 |  | 2023-03-29 | Gx4 | 838872 | Back up laptop ibstallation due to monitor display issue | Laptop instsllation and setup | Performed back up laptop installation | All 4 modules operational |  | Laptop | TOS number: G 5421 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559648621212960112/5559648621212960112_signature_28.png | CECILIA ANGELES | https://www.jotform.com/uploads/macarewecare/222927978372067/5559648621212960112/5559648621212960112_signature_31.png
| 2023-03-30 | 2023-03-30 13:45 | Sultan kudarat Provincial Hospital | Isulan, Sultan kudarat | 12 | Sultan Kudarat | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24046 | PBSP | charmfrias1270@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 100 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2849 | 2023-03-30 | Gx-IV | 807405 | Module 4 for Service Required | Perform Genexpert machine Preventive maintenance Replace module 4 SN: 731302 with new module SN: 210019503 | Perform Genexpert machine Preventive maintenance Replace module 4 SN: 731302 with new module SN: 210019503 | 4 operational modules |  | Genexpert module 6 color | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559648242714269664/5559648242714269664_signature_28.png | Charma Grace E. Frias | https://www.jotform.com/uploads/macarewecare/222927978372067/5559648242714269664/5559648242714269664_signature_31.png
| 2023-03-30 | 2023-03-30 13:04 | Compostela Rural Health Unit | Poblacion Compostela, Davao de Oro | 11 | Davao de Oro | TB - DOTS | NO | REPAIR CALIBRATION TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24045 | PBSP | nielmaquiling20@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF, COUNT: 110 ASSAY: Xpert MTB/RIF Ultra, COUNT: 990 |  | GX-IV 4D | 807360 | Module B3 for Replacement Due for Calibration Printer always encounter paperjamming issues | Replaced defective Module B3 with SN: 693105 with New Module SN: 714940 Generate IQR, Assay Stat and System Log Run XpertCheck on modules B1, B2 and B4 only Reset Printer Cartridge Connections to see any obstacle or debris that can cause paperjamming issue | Module Replacement Perform XpertCheck Calibration Troubleshoot Printer | Modules B1, B2 and B4 passed the calibration Printer still having paper jamming issue even no malfunctions seen in the cartridges | Module B1 for under observation, encountered temp related error, data will be sent to Cepheid for further assestment |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559628917793597738/5559628917793597738_signature_28.png | Niel C. Maquiling | https://www.jotform.com/uploads/macarewecare/222927978372067/5559628917793597738/5559628917793597738_signature_31.png
| 2023-03-30 | 2023-03-30 11:30 | Navotas Medical and Wellness Center | M. Naval St. | NCR | NCR 3rd - Navotas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24044 | PBSP | angelicagionson@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 22 ASSAY: Xpert MTB/RIF Ultra, COUNT: 915 | 2023-03-27 | GXIV | 807592 | Unit for PM. For module replacement of module B4. | Performed PM. Performed module replacement on B4. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module B4 (634856) in replacement to defective B4 (661957). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559613640017735143/5559613640017735143_signature_28.png | Angelica gionson | https://www.jotform.com/uploads/macarewecare/222927978372067/5559613640017735143/5559613640017735143_signature_31.png
| 2023-03-30 | 2023-03-30 12:33 | Dinas Rural Health Unit | Purok Avocado, Don Jose, Dinas Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24043 | PBSP | bernantereynacres@gmail.com | 25 | ASSAY: Others, COUNT: 0 | 2023-03-30 | GX 4 | 838533 | System Unit unable to start up | Replaced System unit | Replaced System unit, Undergone Preventive Maintenence | Available and working | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559604074417966040/5559604074417966040_signature_28.png | REYNA CRES B. LAZAGA | https://www.jotform.com/uploads/macarewecare/222927978372067/5559604074417966040/5559604074417966040_signature_31.png
| 2023-03-30 | 2023-03-29 14:02 | Guagua Rural Health Unit | Gil Puyat St. Sto. Niño Guagua, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24042 | DOH | abigailgsula@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 154 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1135 | 2023-03-22 | GXIV-D | 831812 | A2 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational except A1 |  | 1 Genexpert module | 9765 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559600872514703169/5559600872514703169_signature_28.png | Abigail Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5559600872514703169/5559600872514703169_signature_31.png
| 2023-03-30 | 2023-03-30 12:11 | Dinas Rural Health Unit | Purok Avocado, Don Jose, Dinas Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24041 | PBSP | bernantereynacres@gmail.com | 25 | ASSAY: Others, COUNT: 0 | 2023-03-30 | SOC400 | CPHO320F05909 | System unit unable to start | System Unit replacement | System Unit replacement | Available and working | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559597164417956844/5559597164417956844_signature_28.png | Reyna Cres B. Lazaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5559597164417956844/5559597164417956844_signature_31.png
| 2023-03-30 | 2023-03-30 11:09 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24040 | DOH | cauayanmho@gmail.com | 4 | ASSAY: Xpert MTB/RIF, COUNT: 467 ASSAY: Xpert MTB/RIF Ultra, COUNT: 857 |  | GX 4 | 838477 | Module B1 with frequently error code 1002 | Conduct PM and run Xpert check | Run Xpert check for module B1 | Module B1 passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559569065474135410/5559569065474135410_signature_28.png | Michelle Gregorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5559569065474135410/5559569065474135410_signature_31.png
| 2023-03-30 | 2023-03-30 09:22 | Dona Remedios Rural Health Unit | Brgy. Pulong Sampaloc, Dona Remedios Trinidad, Bulacan | 3 | Bulacan | Laboratory | NO | Back up laptop installation | CEPHEID | Harold Toco |  |  |  |  | 24039 | PBSP | maricel.rmt24@gmail.com | 28 |  | 2023-03-29 | Gx4 | 838872 | Back up laptop ibstallation due to monitor display issue | Laptop instsllation and setup | Performed back up laptop installation | All 4 modules operational |  | Laptop | TOS number: G 5421 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559491368397533061/5559491368397533061_signature_28.png | CECILIA ANGELES | https://www.jotform.com/uploads/macarewecare/222927978372067/5559491368397533061/5559491368397533061_signature_31.png
| 2023-03-30 | 2023-03-30 08:29 | Cadiz City Health Office | Cadiz, Poblacion Negros Occidental | 7 | Negros Occidental | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 24038 | DOH | dazakellyjean@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 187 | 2023-03-30 | Gx4 | 807614 | Error 4017 Thermistor Verification problem | Replaced new module in B3 serial number 618705 | Module replacement in Module 3 Serial Number 683828 Get data results of 3 months, Generate IQR, System log, Assay statistics Clean Icore, clean valve, clean plunger | Machines operational |  | Module | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5559458967018862424/5559458967018862424_signature_28.png | Jelly Jean Daza | https://www.jotform.com/uploads/macarewecare/222927978372067/5559458967018862424/5559458967018862424_signature_31.png
| 2023-03-29 | 2023-03-29 17:36 | Toledo City Health Office | Toledo City | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  |  |  |  | 24037 | PBSP | maryjune1977@yahoo.com | 23 |  | 2023-03-29 | Gx4 | 808640 | Due for Xpertcheck | Perform Xpertcheck | Perform Xpertcheck | All modules functional | Module B2 fail for Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558921198422489935/5558921198422489935_signature_28.png | Mary June Elemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5558921198422489935/5558921198422489935_signature_31.png
| 2023-03-29 | 2023-03-29 14:59 | Perpetual Soccour Hospital | Cebu City | 7 | Cebu | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Raymond Laurito |  |  |  |  | 24036 | COMMERCIAL | pshclinlab@gmail.com | 22 |  |  | MCS + Haemonitics | 2D167 | Due date for Calibration | Calibration/Maintenance | Calibration/Maintenance | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558830508292036615/5558830508292036615_signature_28.png | Roda Fatima Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5558830508292036615/5558830508292036615_signature_31.png
| 2023-03-29 | 2023-03-29 15:00 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Joseph Acebedo |  |  |  |  | 24035 | COMMERCIAL | pshclinlab@gmail.com | 22 |  |  | MCS+ | 11H307 | FOR MAINTENANCE AND CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558830011799804592/5558830011799804592_signature_28.png | Roda Fatima Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5558830011799804592/5558830011799804592_signature_31.png
| 2023-03-29 | 2023-03-29 14:26 | Juan Posadas Health Center | Rodriguez St. | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24034 | PBSP | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1123 | 2023-03-27 | GXIV | 805060 | Unit for PM. For module replacement of module B1. For re-Xpertcheck of module B4. | Performed PM. Performed module replacement on B1. Performed Xpertcheck on B4. Generated reports. | PM, module replacement and Xpertcheck. | Unit is operational. Module B4 (61932), service required. |  | Gx module B1 (761520) in replacement to defective B1 (649046). Xpertcheck Kit. | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558825109514382091/5558825109514382091_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/5558825109514382091/5558825109514382091_signature_31.png
| 2023-03-29 | 2023-03-29 12:17 | Porac Rural Health Unit | Poblacion Porac Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24033 | DOH | rhulabporac@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 206 ASSAY: Xpert MTB/RIF Ultra, COUNT: 577 | 2023-03-22 | GXIV-D | 831911 | C3 and C4 for module replacement | Preventive maintenance Module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 2 Genexpert modules | 9745 and 9753 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558729878912055967/5558729878912055967_signature_28.png | Marife Lansangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5558729878912055967/5558729878912055967_signature_31.png
| 2023-03-29 | 2023-03-29 10:28 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | TB-DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24032 | PBSP | cyrilcordova0515@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4142 | 2023-04-28 | GX4 | 812020 | None | Save file GXC File, IQR, System Log and Assay Stat | Run Xpert check calibration for 4 modules | All modules passed the calibration |  |  | Lot 1000410214 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558664539215372983/5558664539215372983_signature_28.png | Mr. Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5558664539215372983/5558664539215372983_signature_31.png
| 2023-03-29 | 2023-03-29 10:21 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24031 | PBSP | cyrilcordova0515@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4142 |  | GX4 | 804267 | Module E4 required service | Save file for GXC file, IQR, SYSTEM LOG & Assay Stats for reporting. | Run Xpert check calibration | Module E1, E2 and E3 passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558660959213661649/5558660959213661649_signature_28.png | Mr. Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5558660959213661649/5558660959213661649_signature_31.png
| 2023-03-29 | 2023-03-29 10:21 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 24030 | PBSP | cyrilcordova0515@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4142 |  | GX4 | 804267 | Module E4 required service | Save file for GXC file, IQR, SYSTEM LOG & Assay Stats for reporting. | Run Xpert check calibration | Module E1, E2 and E3 passed the calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558660279212718618/5558660279212718618_signature_28.png | Mr. Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5558660279212718618/5558660279212718618_signature_31.png
| 2023-03-29 | 2023-03-29 09:40 | Buluan District Hospital | Buluan, Maguindanao | BARMM | Maguindanao | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 24029 | PBSP | bdhclinicallaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 421 ASSAY: Xpert MTB/RIF Ultra, COUNT: 711 | 2023-03-29 | Gx-IV | 815490 | Module 3 and 4 for service required | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | 3 operational modules |  | Xpertcheck -CE-5 | 1000378669 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558653025721868695/5558653025721868695_signature_28.png | Jose Francis Recinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5558653025721868695/5558653025721868695_signature_31.png
| 2023-03-29 | 2023-03-29 09:18 | Tondo Medical Center | H. Lopez Blvd. | NCR | NCR 1st - Manila | Covid Lab | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24028 | PBSP | microscan@tmc.doh.gov.ph | 28 | ASSAY: Xpert HPV, COUNT: 472 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1192 ASSAY: Xpert SARS-CoV-2, COUNT: 11044 | 2023-03-27 | GXIV | 844121 | Unit for PM. For module replacement of B4. | Performed PM. Performed module replacement on B4. Generated reports. | PM and module replacement. | Unit is operational. For reporting of module A1. |  | Gx module B4 (210071372) in replacement to defective B4 (642801). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558626306214291819/5558626306214291819_signature_28.png | cynthia fuertes,rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5558626306214291819/5558626306214291819_signature_31.png
| 2023-03-29 | 2023-03-29 08:45 | Tondo Medical Center | H. Lopez Blvd. | NCR | NCR 1st - Manila | Covid Lab | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24027 | DOH | microscan@tmc.doh.gov.ph | 28 | ASSAY: Xpert HPV, COUNT: 150 ASSAY: Xpert MTB/RIF Ultra, COUNT: 250 ASSAY: Xpert SARS-CoV-2, COUNT: 2022 | 2023-03-27 | GXIV | 831872 | Unit for PM. Foe module replacement of module A4. | Performed PM. Performed module replacement on A4. Generated reports. | PM and module replacement. | Unit is operational. |  | Gx module A4 (688318) in replacement to defective A4 (634910). | DR 9800 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558625726213585464/5558625726213585464_signature_28.png | Cynthia Fuertes ,Rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5558625726213585464/5558625726213585464_signature_31.png
| 2023-03-29 | 2023-03-24 22:48 | Davao del Sur Provincial Hospital - Digos | Poblacion Digos, Davao del Sur | 11 | Davao del Sur | TB DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24026 | PBSP | shammy0308@gmail.com | 25.9 |  | 2023-03-23 | GX-IV 4D | 807464 | Module A2, A3 and A4 for Replacement | Replaced Defective Module A2 SN: 687556 with new module SN:672190 Replaced the Defective Module A3 SN: 689053 with new Module SN: 672526 Replaced the Defective Module A4 SN: 766131 with new Module SN: 676407 | Perform Module Replacement | Unit is fully operational |  | 900-0491rs | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558591047912130785/5558591047912130785_signature_28.png | Sharmella Joy Doguiles Pinote | https://www.jotform.com/uploads/macarewecare/222927978372067/5558591047912130785/5558591047912130785_signature_31.png
| 2023-03-29 | 2023-03-24 22:48 | Davao del Sur Provincial Hospital - Digos | Poblacion Digos, Davao del Sur | 11 | Davao del Sur | TB DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 24025 | PBSP | shammy0308@gmail.com | 25.9 |  | 2023-03-23 | GX-IV 4D | 807464 | Module A2, A3 and A4 for Replacement | Replaced Defective Module A2 SN: 687556 with new module SN:672190 Replaced the Defective Module A3 SN: 689053 with new Module SN: 672526 Replaced the Defective Module A4 SN: 766131 with new Module SN: 676407 | Perform Module Replacement | Unit is fully operational |  | 900-0491rs | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558590877911236303/5558590877911236303_signature_28.png | Sharmella Joy Doguiles Pinote | https://www.jotform.com/uploads/macarewecare/222927978372067/5558590877911236303/5558590877911236303_signature_31.png
| 2023-03-29 | 2023-03-28 13:00 | Felipe Abrigo Memorial Hospital | Guiuan, Eastern Samar | 8 | Eastern Samar | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 24024 | PBSP | emileramirez161821@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 285 ASSAY: Xpert MTB/RIF Ultra, COUNT: 614 ASSAY: Xpert HIV-1 Viral Load, COUNT: 7 | 2023-03-28 | GXIV | 815271 | N/a | #NAME? | Preventive Maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558411096722154739/5558411096722154739_signature_28.png | EMILIANA B. RAMIREZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5558411096722154739/5558411096722154739_signature_31.png
| 2023-03-28 | 2023-03-28 13:27 | Ospital ng Lungsod ng San Jose Del Monte | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 24023 | PBSP | krizetteamorio@gmail.com | 29.1 | ASSAY: Xpert MTB/RIF, COUNT: 252 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1415 |  | GXIV | 807624 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. B1: failed B2: failed B3: passed B4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558016432227168909/5558016432227168909_signature_28.png | Krizette Marie Amorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5558016432227168909/5558016432227168909_signature_31.png
| 2023-03-28 | 2023-03-28 11:09 | San Jose Del Monte City Health Center | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 24022 | PBSP | theadelrosario@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 858 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 |  | GXIV | 814161 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. B1: passed B2: failed B3: insufficient kit B4: insufficient kit |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558016432225130220/5558016432225130220_signature_28.png | Thea Del Rosario | https://www.jotform.com/uploads/macarewecare/222927978372067/5558016432225130220/5558016432225130220_signature_31.png
| 2023-03-28 | 2023-03-28 11:02 | San Jose Del Monte City Health Center VI | San Jose Del Monte, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION | CEPHEID | Ronnel Belizar |  |  |  |  | 24021 | PBSP | theadelrosario@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1025 ASSAY: Xpert MTB/RIF Ultra, COUNT: 904 ASSAY: Xpert HIV-1 Viral Load, COUNT: 15 |  | GXIV | 807620 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. C1: passed C2: insufficient cartridge C3: passed C4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5558016432227418348/5558016432227418348_signature_28.png | Thea Del Rosario | https://www.jotform.com/uploads/macarewecare/222927978372067/5558016432227418348/5558016432227418348_signature_31.png
| 2023-03-28 | 2023-03-28 14:28 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 24020 | PBSP | elma.bragat@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1069 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2080 ASSAY: Xpert HIV-1 Viral Load, COUNT: 134 |  | GX-IV | 815275 | FOR MODULE D1 REPLACEMENT | #NAME? | - PERFORM MODULE D1 DEFECTIVE SN:683160 WITH NEW MODULE SN:717195 -PERFORM CLEANING OF GX EQPT | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557947873612974688/5557947873612974688_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5557947873612974688/5557947873612974688_signature_31.png
| 2023-03-28 | 2023-03-28 13:27 | Julian Felipe Health Center | Julian Felipe St. Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24019 | PBSP | maoi.lloyd@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1255 ASSAY: Xpert MTB/RIF, COUNT: 43 | 2023-03-28 | GXIV | 838884 | Unit for PM. Unit for Xpertcheck. | Performed PM. Performed Xpertcheck. Generared reports. | PM and Xpertcheck. | Unit is operational. Module B4 (863606) is service required. |  | Xpertcheck | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557917978463447006/5557917978463447006_signature_28.png | Lloyd Maoi M. De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5557917978463447006/5557917978463447006_signature_31.png
| 2023-03-28 | 2023-03-28 13:29 | Carcar cho | Valladollid Carcar | 7 | Cebu | Tb dots | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 24018 | DOH | chaska.lp14@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 492 | 2023-03-28 | Gx4 | 838283 | Error 1002 | Get data, system log, assay statistics, IQR Archive results in 3 months | Clean Up filter, clean icore | All module operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557908210023011752/5557908210023011752_signature_28.png | Lucille Mae Potoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5557908210023011752/5557908210023011752_signature_31.png
| 2023-03-28 | 2023-03-28 13:21 | Las Nieves Rural Health Unit | Purok 5, Poblacion, Las Nieves | 13 | Agusan del Norte | RTDL | NO | REPAIR TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 24017 | WHO | daisermt@yahoo.com | 25.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 142 |  | GX-IV | 847338 | Stuck Cartridge at Module B4 | None | Removing Cartridge using manual method. | All modules are fully functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557906142818521092/5557906142818521092_signature_28.png | Daisy G. Siega | https://www.jotform.com/uploads/macarewecare/222927978372067/5557906142818521092/5557906142818521092_signature_31.png
| 2023-03-28 | 2023-03-28 12:49 | Caloocan City Health Office | C4 road, Sangandaan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24016 | PBSP | crissyname@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 831 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1407 ASSAY: Xpert HIV-1 Viral Load, COUNT: 171 | 2023-03-27 | GXIV | 805271 | Unit for PM. For module replacement of module A2. | Performed PM. Performed module replacement of A2 | PM and module replacement. | Unit is operational. |  | Gx module A2 (757103) in replacement to defective A2 (802378). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557885168468203881/5557885168468203881_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5557885168468203881/5557885168468203881_signature_31.png
| 2023-03-28 | 2023-03-28 12:08 | Carmen Health Center | Brgy. Carmen, Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 24015 | DOH | murielmelchor11@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 459 |  | GeneXpert | 838442 | None | None | Preventive maintenance Generate report data files Performed database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557869531221813299/5557869531221813299_signature_28.png | Muriel melchor | https://www.jotform.com/uploads/macarewecare/222927978372067/5557869531221813299/5557869531221813299_signature_31.png
| 2023-03-28 | 2023-03-28 11:55 | Datu Odin Sinsuat District Hospital | Datu Odin Sinsuat, Maguindanao | 12 | Maguindanao | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 24014 | DOH | rowenasalazar1234@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 305 |  | Gx-IV | 838205 | Module 2 multiple temperature error Module 4 for service required | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 2 and 4 for Troubleshooting | Perform Genexpert machine Preventive maintenance Perform Xpertcheck at module 2 and 4 for Troubleshooting | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557867740022986606/5557867740022986606_signature_28.png | Rowena Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5557867740022986606/5557867740022986606_signature_31.png
| 2023-03-28 | 2023-03-28 11:17 | Velasquez Health Center | Varona St. Tondo | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 24013 | DOH | wcpatron11x4@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 167 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1873 | 2023-03-27 | GXIV | 831534 | Unit for PM. For module replacement of B2. | Performed PM. Performed module replacement on B2. | PM and Module Replacement | Unit is operational. |  | Gx module B2 (711157) in replacement to defective module B2 (753694). | DR 9796 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557829938462291445/5557829938462291445_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5557829938462291445/5557829938462291445_signature_31.png
| 2023-03-28 | 2023-03-27 13:06 | San Luis Rural Health Unit | Poblacion, San Luis, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24012 | DOH | mohan_santos@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 54 ASSAY: Xpert MTB/RIF Ultra, COUNT: 205 | 2023-02-07 | Genexpert IV | 831714 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = A1, A2, A4 - Passed & A3 - Serviced Required | Xpertcheck 5pc | 9585 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557773252022523453/5557773252022523453_signature_28.png | Janine Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/5557773252022523453/5557773252022523453_signature_31.png
| 2023-03-28 | 2023-03-27 11:15 | Candaba Rural Health Unit | Poblacion, Candaba, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 24011 | DOH | rochelannlontoc@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 524 ASSAY: Xpert MTB/RIF, COUNT: 82 | 2023-03-22 | Genexpert IV | 831873 | Module 3, Hardware Failure | Performed Xpertcheck in A3 | Performed Maintenance Performed Module Replacement ( 713126 & 757746 = 708506 & 626301) Performed Calibration Gathered Reports Installed HIV assay Performed Functionality Test | Unit is Operational | Xpertcheck Result - Passed | Genexpert Module - 708506 & 626301 | 9761 & 9756 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557773252025756106/5557773252025756106_signature_28.png | Rochel ann lontoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5557773252025756106/5557773252025756106_signature_31.png
| 2023-03-27 | 2023-03-24 13:34 | Philippine Red Cross - Sta. Cruz | JP Rizal Memorial Hospital, Manila S Rd., Calamba | 4A | Laguna | Blood bank | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  |  |  |  | 24010 | MACARE-HBDC | xernan.dawinan@redcross.org.ph | 26 |  | 2023-03-24 | MCS+ | 02J046 | Defective centrifuge | Replaced defective centrifuge | Replaced defective centrifuge Performed maintenance Simulation run Calibrated the unit | Unit is operational. | Unit is currently installed and repaired in PRC Calamba | Centrifuge assembly SN: 1242 | DR# 9693 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085921218452951/5557085921218452951_signature_28.png | Xernan Dawinan | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085921218452951/5557085921218452951_signature_31.png
| 2023-03-27 | 2023-03-22 09:40 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24008 | DOH | amymitra22@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 286 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1288 | 2023-03-20 | GX-IV-D | 831588 | Modules A2 and A4 for replacement. | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | Unit is operational. Module A3/657408 not operational. |  | Old: A2/792154, A4/754217 New: A2/630566, A4/609675 | DR# 9727, 9726 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861216523047/5557085861216523047_signature_28.png | Amelia Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861216523047/5557085861216523047_signature_31.png
| 2023-03-27 | 2023-03-23 15:10 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24009 | DOH | gemmponce@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 22 ASSAY: Xpert MTB/RIF Ultra, COUNT: 609 | 2023-03-20 | GX-IV-D | 821292 | Module A1 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | Unit is operational. Module A2/783209 not operational but already reported. |  | Old: A1/792273 New: A1/674291 | DR# 9708 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861219244628/5557085861219244628_signature_28.png | Gemma M. Ponce | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861219244628/5557085861219244628_signature_31.png
| 2023-03-27 | 2023-03-22 13:33 | Amang Rodriguez Memorial Medical Center | Sumulong Highway | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24006 | PBSP | bacteriology.armmc@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 467 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2258 | 2023-03-20 | GX-IV-D | 807982 | Module B3 for replacement | Replaced defective module. | Performed maintenence Replaced defective module Generated reports | All modules operational. |  | Old: B3/608964 New: B3/706298 | G 5293 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861214247272/5557085861214247272_signature_28.png | Maricel Calderon | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861214247272/5557085861214247272_signature_31.png
| 2023-03-27 | 2023-03-23 13:30 | Lumban Rural Health Unit | Gen. Luna St., Lumban | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24007 | PBSP | sheilamarylegisma@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 160 ASSAY: Xpert MTB/RIF Ultra, COUNT: 914 | 2023-03-20 | GX-IV-D | 838148 | Module B3 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | All modules operational. |  | Old: B3/806807 New: B3/630625 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861217058644/5557085861217058644_signature_28.png | Rizalyne sacdalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861217058644/5557085861217058644_signature_31.png
| 2023-03-27 | 2023-03-23 10:29 | Sta. Cruz Rural Health Unit | Pio Valenzuela St., Sta. Cruz | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 24005 | DOH | jezzaleal@gmail.co | 24 | ASSAY: Xpert MTB/RIF, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1013 | 2023-03-20 | GX-IV-D | 820229 | Module D4 for replacement | Replaced module D4 | Performed maintenance Replaced defective module Performed maintenance | Unit is operational. Module D1/722064 not operational but reported already. |  | Old: D4/722098 New: D4/713370 | DR# 9733 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861218657608/5557085861218657608_signature_28.png | Jezza Nequinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5557085861218657608/5557085861218657608_signature_31.png
| 2023-03-27 | 2023-03-27 14:02 | Tarlac City Health Center X | Tibag Tarlac City | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24004 | DOH | jayteecornelius@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF, COUNT: 261 ASSAY: Xpert MTB/RIF Ultra, COUNT: 454 | 2023-03-22 | GXIV-D | 831587 | A1 and A4 for module replacement | Repair Module replacement Preventive maintenance | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 2 genexpert module | 9771 and 9759 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557074676686558085/5557074676686558085_signature_28.png | Jaytee Cornelius Bachiller | https://www.jotform.com/uploads/macarewecare/222927978372067/5557074676686558085/5557074676686558085_signature_31.png
| 2023-03-27 | 2023-03-27 12:03 | Concepcion District Hospital | L. Cort3z St. Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24003 | PBSP | rowenapunsalan04@gmail.com | 26.6 | ASSAY: Xpert MTB/RIF, COUNT: 195 ASSAY: Xpert MTB/RIF Ultra, COUNT: 903 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GXIV-D | 807562 | Genexpert module troubleshooting | Troubkeshooting Preventive maintenance | Performed surfqce cleaning Performed Cepheid troubleshooting procedures Gathered reports Performed functionality test | Unit is operational except B1 and B2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557061566682768366/5557061566682768366_signature_28.png | Maria Rowena Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5557061566682768366/5557061566682768366_signature_31.png
| 2023-03-27 | 2023-03-27 12:03 | Concepcion District Hospital | L. Cort3z St. Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 24002 | PBSP | rowenapunsalan04@gmail.com | 26.6 | ASSAY: Xpert MTB/RIF, COUNT: 195 ASSAY: Xpert MTB/RIF Ultra, COUNT: 903 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GXIV-D | 807562 | Genexpert module troubleshooting | Troubkeshooting Preventive maintenance | Performed surfqce cleaning Performed Cepheid troubleshooting procedures Gathered reports Performed functionality test | Unit is operational except B1 and B2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557009556688652781/5557009556688652781_signature_28.png | Maria Rowena Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5557009556688652781/5557009556688652781_signature_31.png
| 2023-03-27 | 2023-03-27 12:09 | West Metro Medical Centre | Veterans Avenue Extension, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 24001 | COMMERCIAL | labresult@westmetro.com.ph | 21 | ASSAY: Xpert SARS-CoV-2, COUNT: 2450 | 2023-03-27 | GX 4 | 834112 | Module B1 and B2 error 4009, 2014 | Preventive maintenence | Module Replacement | Machine not in use | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5557001969016940422/5557001969016940422_signature_28.png | Ma. Razel Canumay | https://www.jotform.com/uploads/macarewecare/222927978372067/5557001969016940422/5557001969016940422_signature_31.png
| 2023-03-27 | 2023-03-23 12:59 | Dasmariñas City Health Office II | Brgy. Sta. Cruz, Dasmariñas | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE ReXpertCheck | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23998 | DOH | hazelmilan0777@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2556 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 | 2023-03-20 | GXIV-D | 838535 | Modules 2, 3 and 4 with Temperature Errors | Module Replacement | Replaced Modules 2 and 4 ReXpertCheck Module 3 Generate Reports | 3 Modules are Operational, 1 Defective | 862309, 654888 -> 638084, 724603 , XC: B3- Failed | GeneXpert Module 6-Colors | 9732, 9735 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980301211749944/5556980301211749944_signature_28.png | Hazel B. Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980301211749944/5556980301211749944_signature_31.png
| 2023-03-27 | 2023-03-23 11:06 | Imus City Health Office II | Plaridel Subd., Imus | 4A | Cavite | Laboratory | NO | MAINTENANCE ReXpertCheck | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23997 | DOH | amielynnmangulabnan1217@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1103 |  | GXIV-D | 821284 | Module 1 and 3 with Temperature Error | ReXpertCheck | Performed Maintenance Performed XpertCheck on Modules 1 and 3 Generate Reports | 3 Modules are Operational, 1 Defective | Module 1 is Unavailable in XpertCheck, Module 3 Passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980301215419886/5556980301215419886_signature_28.png | Amielynn Mangulabnan | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980301215419886/5556980301215419886_signature_31.png
| 2023-03-27 | 2023-03-24 09:38 | Alfonso Rural Health Unit | Burgos St, Alfonso | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23999 | DOH | aim_1225@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 573 | 2023-03-20 | GXIV-D | 821279 | Modules 2,3,4 with Temperature Errors | Module Replacement | Replaced Modules 2, 3, and 4 Performed Maintenance Generate Reports | All Modules are Operational | 730963, 673017, 657115 -> 653212, 867876, 614310 | GeneXpert Module 6-Colors | 9722, 9723, 9724 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980301217946898/5556980301217946898_signature_28.png | Aim Arroyo | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980301217946898/5556980301217946898_signature_31.png
| 2023-03-27 | 2023-03-24 11:20 | Silang Rural Health Unit | Del Pilar St, Silang | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 24000 | DOH | bencq78@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1342 | 2023-03-20 | GXIV-D | 820708 | Modules 1 and 4 Failed in XpertCheck | Module Replacement | Replaced Modules 1 and 4 Performed Maintenance Generate Reports | All Modules are Operational | 723022, 723243 -> 634972, 629212 | GeneXpert Module 6-Colors | 9712, 9713 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980301219450334/5556980301219450334_signature_28.png | Benessie Madlangsakay | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980301219450334/5556980301219450334_signature_31.png
| 2023-03-27 | 2023-03-21 09:52 | General Trias City Health Office | Hospital Road, General Trias | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE ReXpertCheck | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23993 | DOH | chogentri.lab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1540 ASSAY: Xpert HIV-1 Viral Load, COUNT: 140 | 2023-03-20 | GXIV-D | 821340 | 3 Modules with Temperature Errors | Module Replacement | Performed Maintenance Replaced Modules 3 and 4 ReXpertCheck Module 4 | 3 Modules Operational, 1 Defective | 720511, 744626 -> 626065, 601653 | GeneXpert Module Kit 6-Colors | DR 9710, 9711 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071212723230/5556980071212723230_signature_28.png | Mack Alfred L. Chico | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071212723230/5556980071212723230_signature_31.png
| 2023-03-27 | 2023-03-21 11:43 | Tanza Rural Health Unit | Municipal Compound, Tanza | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23994 | PBSP | gutierrezrhian@gmail.com | 25 |  |  | GXIV-D | 838162 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 2 Modules Passed, 2 Modules Failed | XC: 3 & 4 Failed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071213539756/5556980071213539756_signature_28.png | Mary Ann C. Gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071213539756/5556980071213539756_signature_31.png
| 2023-03-27 | 2023-03-21 11:27 | Tanza Rural Health Unit | Municipal Compound, Tanza | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23996 | PBSP | gutierrezrhian@gmail.com | 25 |  | 2023-03-20 | GXIV-D | 812835 | CPU is Defective | Back-up Laptop XpertCheck | Install Back-up Laptop Performed Maintenance Performed XpertCheck Generate Reports | All Modules are Operational, All Modules Passed the XpertCheck | 2UA70329CD -> CPH1019F03275 | Laptop L480 | G 5266 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071214379883/5556980071214379883_signature_28.png | Mary Ann C. Gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071214379883/5556980071214379883_signature_31.png
| 2023-03-27 | 2023-03-22 09:31 | Balayan Rural Health Unit | Brgy. Gumamela, Balayan | 4A | Batangas | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23995 | PBSP | veronica.dalida@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 617 |  | GXIV-D | 838451 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 3 Modules are Operational, 1 for Observation with Temperature Errors | XC: All Modules Passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071216268768/5556980071216268768_signature_28.png | Veronics Dalida | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071216268768/5556980071216268768_signature_31.png
| 2023-03-27 | 2023-03-22 08:33 | Calaca Rural Health Unit | Dela Salud St., Poblacion 3, Calaca | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23995 | DOH | loidamarasigan123@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1162 | 2023-03-20 | GXIV-D | 821300 | Module 2 has temperature error Unit is due for XpertCheck | Module Replacememt XpertCheck | Performed Maintenance Performed XpertCheck Replaced Module 2 Generate Reports | All Modules Passed and Operational | A2: 731300 -> 730076 | GeneXpert Module 6-Colors | 9731 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071218699351/5556980071218699351_signature_28.png | Loida C. Marasigan | https://www.jotform.com/uploads/macarewecare/222927978372067/5556980071218699351/5556980071218699351_signature_31.png
| 2023-03-27 | 2023-03-27 10:06 | Ospital ning Capas | Sto. Domingo 2 Capas, Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23992 | DOH | onclaboratory@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 221 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1349 | 2023-03-21 | GXIV-D | 831591 | A1 and A2 for module replacemeny | Preventive maintenance Module replacement | Performed surface and internal cleaning Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 2 genexpert module | 9746 and 9747 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556925546688342754/5556925546688342754_signature_28.png | Afrilyn Roque | https://www.jotform.com/uploads/macarewecare/222927978372067/5556925546688342754/5556925546688342754_signature_31.png
| 2023-03-26 | 2023-03-24 10:33 | Binan City Health Office 2 | Romana Subd. | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23991 | DOH | eloicyparra@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 84 ASSAY: Xpert MTB/RIF, COUNT: 195 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1163 ASSAY: Xpert HIV-1 Viral Load, COUNT: 235 | 2023-03-20 | GXIV | 821278 | Unit for preventive maintenance. For module replacement of modules B1 and B2. | Performed preventive maintenance. Performed module replacement on modules B1 and B2. Generated reports | Preventive maintenance and module replacement. | Unit is operational. Module B3 for monitoring. Awaiting module for module B4. |  | Gx module B1 (783692) and B2 (750375) in replacement to defective B1 (783692) and B2 (750375). | DR |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058963322261427/5556058963322261427_signature_28.png | Eloicy Parra | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058963322261427/5556058963322261427_signature_31.png
| 2023-03-26 | 2023-03-23 12:21 | Calamba City Health Office | Old Municipal Site, Calamba | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23990 | DOH | 2052456@aup.edu.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1554 | 2023-03-20 | GXIV | 838225 | Unit for preventive maintenance. For module replacement of module B3. | Performed preventive maintenance. Performed module replacement on B3. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. |  | Gx module B3 (687052) in replacement to defective B3 (807734). | DR 9709 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058863323577063/5556058863323577063_signature_28.png | Mary Jane Lourdes Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058863323577063/5556058863323577063_signature_31.png
| 2023-03-26 | 2023-03-23 12:14 | Calamba City Health Office | Old Municipal Site, Calamba | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23989 | PBSP | 2052456@aup.edu.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1936 | 2023-03-20 | GXIV | 807570 | Unit for preventive maintenance. For module replacement of modules B2 and B3. | Performed preventive maintenance. Performed module replacement of module B2 and B3. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. |  | Gx module B2 (684352) and B3 (684562) in replacement to defective B2 (654606) and B3 (684919). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058863328856349/5556058863328856349_signature_28.png | Mary Jane Lourdes Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058863328856349/5556058863328856349_signature_31.png
| 2023-03-26 | 2023-03-23 16:19 | Los Banos Rural Health Unit | Los Banos Road | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23988 | PBSP | rmbtaguic@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 796 | 2023-03-20 | GXIV | 804716 | Unit for preventive maintenance. For module replacement of module B3. | Performed preventive maintenance. Performed module replacement. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. |  | Gx module B3 (210046115) in replacement to defective B3 (706162). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823325407321/5556058823325407321_signature_28.png | Rhea Marie Dorado | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823325407321/5556058823325407321_signature_31.png
| 2023-03-26 | 2023-03-22 12:52 | Mabini Rural Health Unit | Poblacion Mabini | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23987 | PBSP | glee.maravilla@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 995 | 2023-03-20 | GXIV | 838867 | Unit for preventive maintenance. For module replacement of module B1. | Performed preventive maintenance. Performed module replacement on module B1. Generated repirts. | Preventive maintenance and module replacement. | Unit is operational |  | Gx module B1 (627970) in replacement to detective B1 (863439). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823325578925/5556058823325578925_signature_28.png | Gliceria C. Maravilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823325578925/5556058823325578925_signature_31.png
| 2023-03-26 | 2023-03-23 11:29 | Bay Rural Health Unit | J. Tanalega Bay | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23987 | PBSP | rhubaylaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 22 ASSAY: Xpert MTB/RIF Ultra, COUNT: 560 | 2023-03-20 | GXIV | 838132 | Unit for preventive maintenance. For module replacement of module B4. | Performed preventive maintenance. Performed module replacement. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. |  | Gx module B4 (694696) in replacement to defective module B4 (806709). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823328668769/5556058823328668769_signature_28.png | Sheena Mae A. Baquiran | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823328668769/5556058823328668769_signature_31.png
| 2023-03-26 | 2023-03-22 08:30 | Juan Posadas Health Center | Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23987 | PBSP | elizajunio05@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1077 | 2023-03-20 | GXIV | 805060 | Unit for preventive maintenance. For module replacement of module B3. | Performed preventive maintenance. Performed module replacement on module B3. Generated reports, . | Preventive maintenance and module replacement. | Unit is operational. |  | Gx Module B3 (629063) in replacement to defective B3 (654188). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823328978940/5556058823328978940_signature_28.png | Eliza Junio | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823328978940/5556058823328978940_signature_31.png
| 2023-03-26 | 2023-03-24 09:04 | Binan City Health Office 1 | Zone 1, Good Morning Subd. | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23987 | PBSP | annefriam.almazora@yahoo.com | 28 | ASSAY: Xpert HPV, COUNT: 149 ASSAY: Xpert MTB/RIF Ultra, COUNT: 259 | 2023-03-20 | GXIV | 812780 | Unit for preventive maintenance. For module replacement of module B3. | Performed preventive maintenance. Performed module replacement on module B3. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. |  | Gx module B3 (881212) in replacement to defective module B3 (624415). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823329632915/5556058823329632915_signature_28.png | Annefriam Almazora | https://www.jotform.com/uploads/macarewecare/222927978372067/5556058823329632915/5556058823329632915_signature_31.png
| 2023-03-25 | 2023-03-24 08:16 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23986 | DOH | ammegzedd@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 422 ASSAY: Xpert MTB/RIF Ultra, COUNT: 696 | 2023-03-24 | GXIV | 838421 | N/A | #NAME? | Preventive Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5555551469841110548/5555551469841110548_signature_28.png | Gemma Malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5555551469841110548/5555551469841110548_signature_31.png
| 2023-03-25 | 2023-03-25 10:21 | Chinese General Hospital | 286 Blumentritt St., Sta. Cruz, Manila | NCR | NCR 1st - Manila | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 23985 | CONSORTIUM | cgh_pathology@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 10 ASSAY: Xpert CT/NG, COUNT: 26 ASSAY: Xpert HBV VL, COUNT: 44 ASSAY: Xpert HPV, COUNT: 25 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1253 | 2023-03-20 | Gx4 | 835604 | For module replacement and Xpertcheck calibration | Replaced modules B1 and B2 Performed xpertcheck calibration. | Performed maintenance Performed module replacement. Performed xpertcheck calibration. Performed functionality test. | All 4 modules operational. | Defective modules SN:790874 and 790897 | Good mods sn 210025281 and 690208 | DR number: 9719 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5555273543518757904/5555273543518757904_signature_28.png | Mary Joyce S. Gaudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5555273543518757904/5555273543518757904_signature_31.png
| 2023-03-25 | 2023-03-25 09:30 | Hi Precision Diagnostics | E. Long Bldg 440-442 Del Monte Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 23984 | CONSORTIUM | jose@hi-precision.com.ph | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert HPV, COUNT: 43 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4146 | 2023-03-24 | GxIV | 835602 | Module A2 for replacement | Replaced module A2 | Performed maintenance Performed module replacement. | All 4 modules operational. | Defective mod sn: 791652 | Good module SN 692016 | DR number 9781 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5555181343515209635/5555181343515209635_signature_28.png | Joselito Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5555181343515209635/5555181343515209635_signature_31.png
| 2023-03-25 | 2023-03-25 09:24 | Hi Precision Diagnostics | W. Long Bldg. 440-442 Del Monte Ave., Quezon City | NCR | NCR 2nd - Quezon City | Molecular Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 23983 | CONSORTIUM | jose@hi-precision.com.ph | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert HPV, COUNT: 77 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4211 | 2023-03-24 | GxIV | 835601 | Modules B2 and B3 for module replacement | Replaced modules B2 and B3 | Performed PM Perpormed module replacement. Performed functionality test. | All 4 modules operatioanal. | Defective mods SN 787789 and 891796 | Good Modules SN 787789 and 891796 | DR number 9781 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5555180043516933245/5555180043516933245_signature_28.png | Joselito Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5555180043516933245/5555180043516933245_signature_31.png
| 2023-03-24 | 2023-03-24 14:42 | San Rafael Rural Health Unit | San Rafael, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23982 | DOH | madssvalerio@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 381 | 2023-01-19 | GXIV | 831692 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. A1: A2: A3: A4: |  | Xpertcheck kit | 831692 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554639015519420096/5554639015519420096_signature_28.png | Marilyn Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5554639015519420096/5554639015519420096_signature_31.png
| 2023-03-24 | 2023-03-24 14:58 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23981 | MACARE-HBDC | spmcrhivda@gmail.com | 25.7 |  |  | GX-IV 4D | 839152 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554510610713989417/5554510610713989417_signature_28.png | Flordeliza M. Go RMT MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/5554510610713989417/5554510610713989417_signature_31.png
| 2023-03-24 | 2023-03-24 13:39 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23980 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 195 ASSAY: Xpert MTB/RIF, COUNT: 1136 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1172 |  | GX4 | 838285 | None | None | #NAME? | All four modules are functional | Observed frequent temp error in module A3. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554462838813187650/5554462838813187650_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5554462838813187650/5554462838813187650_signature_31.png
| 2023-03-24 | 2023-03-24 13:42 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23979 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 203 ASSAY: Xpert MTB/RIF, COUNT: 976 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1207 |  | GX4 | 820206 | None | None | Preventive Maintenance | All four modules are functional | Waiting for module A3 replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554462838813833245/5554462838813833245_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5554462838813833245/5554462838813833245_signature_31.png
| 2023-03-24 | 2023-03-24 11:41 | Manila Medical Center | 850 UN Ave., Paco, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | Cienne Mae Valerio |  |  |  |  | 23978 | CONSORTIUM | bacezar@manilamed.com.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 842 | 2023-03-20 | GXIV | 835540 | Due for Module Replacement Due for Xpertcheck | Replaced Defective Modules Cleaned Cartridge Bay, Icore and Plunger Run Xpertcheck 1.5A Gathered Reports | Performed Module Replacement on B2 and B3 Performed Xpertcheck on B1 and B4 | Unit is operational B2 Passed B4 Xpertcheck Retry - Service Required for Reporting | OLD MODULE 790340 NEW 210036762 OLD 636508 NEW 616066 | 900-GX | 9720 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554412035936697875/5554412035936697875_signature_28.png | Boylyn Cezar | https://www.jotform.com/uploads/macarewecare/222927978372067/5554412035936697875/5554412035936697875_signature_31.png
| 2023-03-24 | 2023-03-24 11:41 | Manila Medical Center | 850 UN Ave., Paco, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | Cienne Mae Valerio |  |  |  |  | 23977 | CONSORTIUM | bacezar@manilamed.com.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 842 | 2023-03-20 | GXIV | 835540 | Due for Module Replacement Due for Xpertcheck | Replaced Defective Modules Cleaned Cartridge Bay, Icore and Plunger Run Xpertcheck 1.5A Gathered Reports | Performed Module Replacement on B2 and B3 Performed Xpertcheck on B1 and B4 | Unit is operational B2 Passed B4 Xpertcheck Retry - Service Required | OLD MODULE 790340 NEW 210036762 OLD 636508 NEW 616066 | 900-GX | 9720 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554411975937752855/5554411975937752855_signature_28.png | Boylyn Cezar | https://www.jotform.com/uploads/macarewecare/222927978372067/5554411975937752855/5554411975937752855_signature_31.png
| 2023-03-24 | 2023-03-24 11:43 | Dr. Paulino J. Garcia Memorial Research and Medical Center | Mabini Street Extension,Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23976 | PBSP | bryanmichaelbalunes@gmail.com | 27 |  | 2023-03-24 | Genexpert XVI | 844111 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration (B3 = Failed) & (All modules Passed except B3) Disabled B3 Performed Functionality Test | Unit is Operational | Xpertcheck B3 = Service Required | Xpertcheck 17pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554411379938195666/5554411379938195666_signature_28.png | Bryan Michael Balunes | https://www.jotform.com/uploads/macarewecare/222927978372067/5554411379938195666/5554411379938195666_signature_31.png
| 2023-03-24 | 2023-03-23 12:06 | Guimba Community Hospital | Saranay District (Pob.),Guimba, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Rexpertcheck module 3 | CEPHEID | Matthew Roxas |  |  |  |  | 23975 | PBSP | marquezandrelyn18@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 129 ASSAY: Xpert MTB/RIF Ultra, COUNT: 834 | 2023-03-23 | Genexpert IV | 804560 | UPS not working in power shortage Printer automatic shutdown | Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Xpertcheck in Module B3 (633463) = Passed Performed Functionality Test | Unit is Operational | UPS for Replacement | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554411379939217268/5554411379939217268_signature_28.png | Andrelyn Marquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5554411379939217268/5554411379939217268_signature_31.png
| 2023-03-24 | 2023-03-24 11:03 | Norala District Hospital | Norala, South Cotabato | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23974 | DOH | NoralaDHlaboratory@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF, COUNT: 287 ASSAY: Xpert MTB/RIF Ultra, COUNT: 788 |  | Gx-lV | 838543 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554384491615911311/5554384491615911311_signature_28.png | Reziel Deano Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/5554384491615911311/5554384491615911311_signature_31.png
| 2023-03-24 | 2023-03-23 14:10 | Our Lady of Lourdes Hospital | Sta. Mesa, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23973 | CONSORTIUM | louie.sgebiomedtech@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 219 ASSAY: Xpert MTB/RIF Ultra, COUNT: 177 | 2023-03-20 | GXIV | 840974 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. B1: passed B2: passed B3: passed B4: passed |  | Xpertcheck kit | 9717 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329779213869334/5554329779213869334_signature_28.png | Mea Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329779213869334/5554329779213869334_signature_31.png
| 2023-03-24 | 2023-03-22 14:18 | Pandi Rural Health Unit | Pandi, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23971 | WHO | pandirhulaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 202 ASSAY: Xpert MTB/RIF Ultra, COUNT: 792 |  | GXIV | 838883 | Unit for calibration. | #NAME? | Performed Xpertcheck | Unit is operational. B1: passed B2: insufficient kit B3: insufficient kit B4: passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419211211303/5554329419211211303_signature_28.png | Jhoana Marie Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419211211303/5554329419211211303_signature_31.png
| 2023-03-24 | 2023-03-23 10:17 | Victor R. Potenciano Medical Center | Edsa, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Pathology | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23969 | CONSORTIUM | oro_janet@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 221 ASSAY: Xpert MTB/RIF Ultra, COUNT: 201 | 2023-03-20 | GXIV | 835600 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. B1: passed B2: failed B3: insufficient kit B4: passed B2 for replacement. B3 is under observation. |  | Xpertcheck kit | 9715 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419216563381/5554329419216563381_signature_28.png | Janet Oro | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419216563381/5554329419216563381_signature_31.png
| 2023-03-24 | 2023-03-22 12:51 | Bustos Rural Health Unit | Bustos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23972 | DOH | ronald_galvez2002@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 144 ASSAY: Xpert MTB/RIF Ultra, COUNT: 472 | 2023-03-20 | GXIV | 831874 | Module B4 (793120) for replacement. | #NAME? | Replaced module B4 | Unit is operational. |  | GeneXpert Module SN: 909803 | 9734 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419216801022/5554329419216801022_signature_28.png | Ronald Galvez | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419216801022/5554329419216801022_signature_31.png
| 2023-03-24 | 2023-03-23 12:55 | University of The East Ramon Magsaysay General Hospital | Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23969 | CONSORTIUM | patholab@uerm.edu.ph | 25 | ASSAY: Xpert C. difficile BT, COUNT: 47 ASSAY: Xpert CT/NG, COUNT: 60 ASSAY: Xpert MTB/RIF, COUNT: 421 ASSAY: Xpert MTB/RIF Ultra, COUNT: 128 ASSAY: Others, COUNT: 2 | 2023-03-20 | GXIV | 847578 | Unit for calibration. | #NAME? | Performed Xpertcheck. | Unit is operational. B1: passed B2: passed B3: passed B4: passed |  | Xpertcheck kit | 9716 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419218636158/5554329419218636158_signature_28.png | Gilbert Claveria | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419218636158/5554329419218636158_signature_31.png
| 2023-03-24 | 2023-03-22 10:29 | Meycauayan City Health Center VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23970 | DOH | ronamarieiris_25@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 181 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1081 ASSAY: Xpert HIV-1 Viral Load, COUNT: 44 | 2023-03-20 | GXIV | 831904 | Module A3 (758218) for replacement. | - replaced module A3 by 793458 - checked and cleaned cartridge bay, plunger rod, filter - Generated reports | Replaced A3 | Unit is operational. |  | GeneXpert Module SN: 793458 | 9714 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419218833414/5554329419218833414_signature_28.png | Rona Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5554329419218833414/5554329419218833414_signature_31.png
| 2023-03-24 | 2023-03-23 16:23 | Sta lucia health center | Gomes st.Novaliches,Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23968 | DOH | josie_dingle2004@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 128 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1037 | 2023-03-22 | GX4 | 831901 | A3 hw failed for rexpertcheck module replacement 758241 - 695005 | Performed module replacement and re xpertcheck,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and re xpertcheck | A1 module reported,A3 failed xpertcheck |  | 900-0491rs | 9752 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554270322773760539/5554270322773760539_signature_28.png | Josephine L.Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5554270322773760539/5554270322773760539_signature_31.png
| 2023-03-24 | 2023-03-23 16:23 | Sta lucia health center | Gomes st.Novaliches,Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23967 | DOH | josie_dingle2004@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 128 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1037 | 2023-03-22 | GX4 | 831901 | A3 hw failed for rexpertcheck module replacement 758241 - 695005 | Performed module replacement and re xpertcheck,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and re xpertcheck | A1 module reported,A3 failed xpertcheck |  | 900-0491rs | 9752 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554260362772628775/5554260362772628775_signature_28.png | Josephine L.Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5554260362772628775/5554260362772628775_signature_31.png
| 2023-03-24 | 2023-03-23 12:55 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23966 | DOH | vanessadelmonte401@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 158 ASSAY: Xpert MTB/RIF Ultra, COUNT: 383 | 2023-03-23 | GXIV | 838409 | #NAME? | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554256749158045051/5554256749158045051_signature_28.png | Vanessa mabag | https://www.jotform.com/uploads/macarewecare/222927978372067/5554256749158045051/5554256749158045051_signature_31.png
| 2023-03-24 | 2023-03-23 16:23 | Sta lucia health center | Gomes st.Novaliches,Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23965 | DOH | josie_dingle2004@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 128 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1037 | 2023-03-22 | GX4 | 831901 | A3 hw failed for rexpertcheck module replacement 758241 - 695005 | Performed module replacement and re xpertcheck,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and re xpertcheck | A1 module reported,A3 failed xpertcheck |  | 900-0491rs | 9752 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554191832772896954/5554191832772896954_signature_28.png | Josephine L.Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5554191832772896954/5554191832772896954_signature_31.png
| 2023-03-24 | 2023-03-23 14:07 | Commonwealth health center | Katuparan st,Commonwealth ,Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23964 | DOH | commonwealthclinical@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 579 ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1526 | 2023-03-20 | GX4 | 831615 | For module replacement A4 encounter temp error for rexpertcheck 793225 - 625759 | Performed module replacement and performed xpertcheck,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and re xpertcheck | A4 passed xpertcheck all module operational |  | 900-0491rs | 9729 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554190062773104138/5554190062773104138_signature_28.png | BLUE ROB PACURE | https://www.jotform.com/uploads/macarewecare/222927978372067/5554190062773104138/5554190062773104138_signature_31.png
| 2023-03-24 | 2023-03-23 14:07 | Commonwealth health center | Katuparan st,Commonwealth ,Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23963 | DOH | commonwealthclinical@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 579 ASSAY: Xpert MTB/RIF, COUNT: 82 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1526 | 2023-03-20 | GX4 | 831615 | For module replacement A4 encounter temp error for rexpertcheck 793225 - 625759 | Performed module replacement and performed xpertcheck,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and re xpertcheck | A4 passed xpertcheck all module operational |  | 900-0491rs | 9729 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5554189692779988616/5554189692779988616_signature_28.png | BLUE ROB PACURE | https://www.jotform.com/uploads/macarewecare/222927978372067/5554189692779988616/5554189692779988616_signature_31.png
| 2023-03-23 | 2023-03-23 15:13 | Sultan Naga Dimaporo Rural Health Unit | Sultan Naga Dimaporo | 10 | Lanao del Norte | TB DOTS | NO | INSTALLATION CALIBRATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23962 | PBSP | rhu.snd@gmail.com | 30 |  |  | GeneXpert | 838193 | None | None | GeneXpert Installation XpertCheck Calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5553650829794665670/5553650829794665670_signature_28.png | JOHAIMEN M MACA-ALANG | https://www.jotform.com/uploads/macarewecare/222927978372067/5553650829794665670/5553650829794665670_signature_31.png
| 2023-03-23 | 2023-03-23 11:43 | Philippine Red Cross (Port Area) | Port area,Manila | NCR | NCR 1st - Manila | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 23961 | MACARE-HBDC | dolly.banares@redcross.org.ph | 28 |  |  | MCS + | 02A087 | As per ens user they encounter during priming of transfer pump | Checked pumps, checked calibration parameters,check line sensors,display panel,dpm,spm,cuff,performed calibration, checked cable connection,performed functionality test | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5553617016321311123/5553617016321311123_signature_28.png | Dolly bañares | https://www.jotform.com/uploads/macarewecare/222927978372067/5553617016321311123/5553617016321311123_signature_31.png
| 2023-03-23 | 2023-03-23 13:48 | Negros Oriental Provincial Hospital RTDL | Dumaguete City | 7 | Cebu | Laboratory | YES | REPAIR | CEPHEID | Jovy Noel |  |  |  |  | 23960 | PBSP | rezelmarielagos@gmail.com | 24 |  | 2023-03-22 | Gx4 | 805275 | Plunger Rod stockup cannot move up. | Refresh software | Refresh software | All modules functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5553613557227524275/5553613557227524275_signature_28.png | Rezelmarie Lagos | https://www.jotform.com/uploads/macarewecare/222927978372067/5553613557227524275/5553613557227524275_signature_31.png
| 2023-03-23 | 2023-03-23 13:22 | Western Visayas Medical Inc. | Ilo Ilo City | 6 | Iloilo | Laboratory | YES | REPAIR | CEPHEID | Jovy Noel |  |  |  |  | 23959 | PBSP | mitorilla.pbsp@gmail.com | 23 |  | 2023-03-16 | Gx4 | 848241 | Module B3 210002438 | Replace module | Replace module B3 with new serial number 210007529 | All modules functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5553595957227360353/5553595957227360353_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5553595957227360353/5553595957227360353_signature_31.png
| 2023-03-23 | 2023-03-23 13:14 | Pontevedra Rural Health Unit | Pontevedra Capiz | 6 | Iloilo | Laboratory | YES | REPAIR | CEPHEID | Jovy Noel |  |  |  |  | 23958 | DOH | rtabangjr@gmail.com | 24 |  | 2023-03-15 | Gx4 | 838296 | Module B1 and B2 | Replace module | Replace Module B1 And B2 with new serial number 694490,694269 | All modules are function ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5553581197221564960/5553581197221564960_signature_28.png | Reynaldo Tabang | https://www.jotform.com/uploads/macarewecare/222927978372067/5553581197221564960/5553581197221564960_signature_31.png
| 2023-03-23 | 2023-03-23 11:43 | Philippine Red Cross (Port Area) | Port area,Manila | NCR | NCR 1st - Manila | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 23957 | MACARE-HBDC | dolly.banares@redcross.org.ph | 28 |  |  | MCS + | 02A087 | As per ens user they encounter during priming of transfer pump | Checked pumps, checked calibration parameters,check line sensors,display panel,dpm,spm,cuff,performed calibration, checked cable connection,performed functionality test | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5553541946329104376/5553541946329104376_signature_28.png | Dolly bañares | https://www.jotform.com/uploads/macarewecare/222927978372067/5553541946329104376/5553541946329104376_signature_31.png
| 2023-03-23 | 2023-03-23 11:32 | Kapatagan Rural Health Unit | Poblacion, Kapatagan | 10 | Lanao del Norte | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23956 | DOH | jhayrajaculbe@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 558 |  | GeneXpert | 838224 | System Unit not turning on, power indicator blinking | Reset system units power Reset RAM insertion | System Unit Repair Preventive Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5553525043124074916/5553525043124074916_signature_28.png | Jayralyn J. Maquiling | https://www.jotform.com/uploads/macarewecare/222927978372067/5553525043124074916/5553525043124074916_signature_31.png
| 2023-03-23 | 2023-03-23 10:55 | Pontevedra Rural Health Unit | Pontevedra Negros Occidental | 6 | Negros Occidental | Laboratory Dept. | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 23955 | DOH | maeannjbonilla@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 633 |  | GX-IV | 838491 | Error 4012 | Calibrate the module B1 and conduct PM | Run Xpert check calibration for module B1 with SN 862261. | Module B1 failed in Xpert check and need service. Module B2, B3 and B4 are in good condition. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5553500010323919994/5553500010323919994_signature_28.png | Mae Ann Pabalinas | https://www.jotform.com/uploads/macarewecare/222927978372067/5553500010323919994/5553500010323919994_signature_31.png
| 2023-03-23 | 2023-03-22 14:15 | Daram Rural Health Unit | Daram, Samar | 8 | Samar | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23954 | PBSP | irene85aragon@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 240 ASSAY: Xpert MTB/RIF Ultra, COUNT: 291 | 2023-03-22 | GXIV | 838868 | #NAME? | #NAME? | #NAME? | Unit is Operational | #NAME? | Xpertcheck | XPERTCHECK - CE - 5 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5553399004583875407/5553399004583875407_signature_28.png | Irene Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/5553399004583875407/5553399004583875407_signature_31.png
| 2023-03-22 | 2023-03-17 10:19 | Doña Marta Health Center | Revilla St.,Pasay city | NCR | NCR 4th - Pasay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23953 | DOH | cerissalynsabado@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 | 2023-01-18 | GX4 | 831700 | For maintenance and xpertcheck | Xpert check,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Xpertcheck and maintenance | A2 module reported already, | A1,A3,A4 passed xpertcheck | Xpertcheck kits | 9492 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552949782771659241/5552949782771659241_signature_28.png | Cerissa Lyn Sabado | https://www.jotform.com/uploads/macarewecare/222927978372067/5552949782771659241/5552949782771659241_signature_31.png
| 2023-03-22 | 2023-03-17 10:19 | Doña Marta Health Center | Revilla St.,Pasay city | NCR | NCR 4th - Pasay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23952 | DOH | cerissalynsabado@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 | 2023-01-18 | GX4 | 831700 | For maintenance and xpertcheck | Xpert check,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Xpertcheck and maintenance | A2 module reported already, | A1,A3,A4 passed xpertcheck | Xpertcheck kits | 9492 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552949522779522091/5552949522779522091_signature_28.png | Cerissa Lyn Sabado | https://www.jotform.com/uploads/macarewecare/222927978372067/5552949522779522091/5552949522779522091_signature_31.png
| 2023-03-22 | 2023-03-22 15:20 | East avenue medical center | East avenue quezon city | NCR | NCR 2nd - Quezon City | Bloodbank bank | YES | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 23951 | MACARE-HBDC | eamc.biomed@gmail.com | 28 |  | 2023-03-02 | MCS + | 15H044G | No display,defective cpu board | Performed troubleshooting, replaced cpu board for testing after replacing the cpu board the machine have display, peformed isolation running completes process , checked calibration parameters. | Troubleshooting | Unit unoperational for quotation cpu board |  | Cpu board for testing purpose, | 9660 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552948472777362548/5552948472777362548_signature_28.png | Ryan Baldelobar | https://www.jotform.com/uploads/macarewecare/222927978372067/5552948472777362548/5552948472777362548_signature_31.png
| 2023-03-22 | 2023-03-22 14:38 | Victoria Rural Health Unit | Brgy. Bulo, Victoria Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23950 | DOH | diazrmt@yahoo.com | 23.8 | ASSAY: Xpert MTB/RIF, COUNT: 289 ASSAY: Xpert MTB/RIF Ultra, COUNT: 488 | 2023-02-01 | GXIV_D | 831669 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except A3:756996 failed xpertcheck | For reporting | 4 xpertcheck kits | 9566 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552859410285574458/5552859410285574458_signature_28.png | Alma Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5552859410285574458/5552859410285574458_signature_31.png
| 2023-03-22 | 2023-03-22 15:12 | Cabanatuan City Health Office | Kapt. Pepe, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23949 | DOH | cholab2021@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1927 ASSAY: Xpert MTB/RIF, COUNT: 782 | 2023-03-22 | Genexpert IV | 831592 | Module A1 error 1002 | Gathered Data | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552858242024170529/5552858242024170529_signature_28.png | Maricris So | https://www.jotform.com/uploads/macarewecare/222927978372067/5552858242024170529/5552858242024170529_signature_31.png
| 2023-03-22 | 2023-03-22 14:38 | Victoria Rural Health Unit | Brgy. Bulo, Victoria Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23948 | DOH | diazrmt@yahoo.com | 23.8 | ASSAY: Xpert MTB/RIF, COUNT: 289 ASSAY: Xpert MTB/RIF Ultra, COUNT: 488 | 2023-02-01 | GXIV_D | 831669 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except A3:756996 failed xpertcheck | For reporting | 4 xpertcheck kits | 9566 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552844340284609237/5552844340284609237_signature_28.png | Alma Estayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5552844340284609237/5552844340284609237_signature_31.png
| 2023-03-22 | 2023-03-22 16:30 | Ciudad Medical Zamboanga | Mayor vitaliano agan Ave., Zamboanga City | 9 | Zamboanga del Sur | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23947 | COMMERCIAL | abdurahmanrajiv@gmail.com | 26 | ASSAY: Xpert CT/NG, COUNT: 9 ASSAY: Xpert HBV VL, COUNT: 63 ASSAY: Xpert HPV, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 359 ASSAY: Xpert HCV VL, COUNT: 38 | 2023-03-22 | GX 2 | 834128 | Module 1 h/w failed | Preventive maintenance, Troubleshooting | Preventive maintenance, Troubleshooting | Module 1 error 4017, module 2 available | Module 1 recommended for replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552843951854631896/5552843951854631896_signature_28.png | Rajiv B. Abdurahman | https://www.jotform.com/uploads/macarewecare/222927978372067/5552843951854631896/5552843951854631896_signature_31.png
| 2023-03-22 | 2023-03-21 15:54 | Calbayog City Health Office | Talahib, Calbayog Samar | 8 | Samar | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23946 | PBSP | maganteelena@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 627 ASSAY: Xpert MTB/RIF Ultra, COUNT: 603 | 2023-03-21 | GXIV | 804265 | #NAME? | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552734746016074051/5552734746016074051_signature_28.png | Elena M. Villa | https://www.jotform.com/uploads/macarewecare/222927978372067/5552734746016074051/5552734746016074051_signature_31.png
| 2023-03-22 | 2023-03-22 12:13 | Bagumbayan Rural Health Unit | Bagumbayan, Sultan kudarat | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23944 | DOH | marizalenoid@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 112 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1424 | 2023-03-22 | Gx-IV | 838261 | Module 4 for service required Module 3 stuck cartridge | Replace module 4 SN: 808021 with new module SN: 722484 Fixed module 3 stuck cartridge Perform Genexpert machine Preventive maintenance | Replace module 4 SN: 808021 with new module SN: 722484 Fixed module 3 stuck cartridge Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 - color | DR no. 9681 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552693618348824604/5552693618348824604_signature_28.png | Ma. RIZA R. DIONELA | https://www.jotform.com/uploads/macarewecare/222927978372067/5552693618348824604/5552693618348824604_signature_31.png
| 2023-03-22 | 2023-03-21 13:53 | Calumpang Rural Health Unit | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23945 | DOH | jan19.cam@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 976 ASSAY: Xpert MTB/RIF, COUNT: 566 |  | Gx-IV | 838504 | Module 4 H/W failed due to multiple temperature error | Perform Genexpert machine Calibration/Xpertcheck at module 4 for troubleshooting Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Calibration/Xpertcheck at module 4 for troubleshooting Perform Genexpert machine Preventive maintenance | Gx-IV | Module 4 passed the Xpertcheck/calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552693628345810620/5552693628345810620_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5552693628345810620/5552693628345810620_signature_31.png
| 2023-03-22 | 2023-03-22 11:46 | Bunawan District Health Center | Bunawan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23943 | PBSP | amutcee@gmail.com | 28.9 |  | 2023-03-21 | GX-IV 4D | 815392 | Module B4 for Module Replacement | Replace the defective Module B4 S/N: 695992 with new Module SN: 750033 | Perform Module Replacement | Unit is fully operational | Module B3 Encounter Error, data will be sent to Cepheid for further evaluation | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552665322528593928/5552665322528593928_signature_28.png | AMRAIDA S. CAMID, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5552665322528593928/5552665322528593928_signature_31.png
| 2023-03-22 | 2023-03-22 11:46 | Bunawan District Health Center | Bunawan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23942 | PBSP | amutcee@gmail.com | 28.9 |  | 2023-03-21 | GX-IV 4D | 815392 | Module B4 for Module Replacement | Replace the defective Module B4 S/N: 695992 with new Module SN: 750033 | Perform Module Replacement | Unit is fully operational | Module B3 Encounter Error, data will be sent to Cepheid for further evaluation | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552665272524532533/5552665272524532533_signature_28.png | AMRAIDA S. CAMID, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5552665272524532533/5552665272524532533_signature_31.png
| 2023-03-22 | 2023-03-17 12:28 | Oriental Mindoro Provincial Hospital | Calapan | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23941 | PBSP | mariatheresayaco@gmail.com | 20 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1574 ASSAY: Xpert SARS-CoV-2, COUNT: 3820 | 2023-03-13 | GX-XVI-D | 842229 | Module C1/883434 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | All modules operational. |  | Old: C1/883434 New: C1/ 705875 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552580672609596129/5552580672609596129_signature_28.png | Maria Theresa M. Yaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5552580672609596129/5552580672609596129_signature_31.png
| 2023-03-22 | 2023-03-21 09:47 | The Medical City | Ortigas Ave, | NCR | NCR 2nd - Pasig | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  |  |  |  | 23940 | COMMERCIAL | moleculardx@themedicalcity.com | 25 |  |  | GX-XVI-10C | 110015767 | For installation | Installed the GeneXpert system | Installed the GeneXpert unit, workstation, printer, barcode scanner and UPS. Functionality test Import assays | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552580602601604064/5552580602601604064_signature_28.png | Xchosa Kay Repatacodo | https://www.jotform.com/uploads/macarewecare/222927978372067/5552580602601604064/5552580602601604064_signature_31.png
| 2023-03-22 | 2023-03-14 11:20 | Occidental Mindoro Provincial Hospital | Mamburao | 4B | Occidental Mindoro | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23939 | PBSP | jaysonlopez12@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 271 ASSAY: Xpert SARS-CoV-2, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 17 |  | GX-IV-D | 807987 | For maintenance Module B3/662114 continuous booting | Check module B3 connections Performed maintenance | Performed maintenance Checked internal connections Checked moudle B3 internal cables Generated reports | Modules B1, B2, and B4 operational. Module B3/662114 not operational, reported already. | Work with Shed Daroya |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579592608167524/5552579592608167524_signature_28.png | Jayson Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579592608167524/5552579592608167524_signature_31.png
| 2023-03-22 | 2023-03-14 15:53 | San Jose Public Health and Diagnostic Center | Poblacion 7, San Jose | 4B | Occidental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23935 | PBSP | naty_garcia1962@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 150 ASSAY: Xpert MTB/RIF Ultra, COUNT: 750 |  | GX-IV-D | 838197 | For XpertCheck | Run XpertCheck | Performed maintenance Generated reports Run XpertCheck | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579582601295299/5552579582601295299_signature_28.png | Natividad Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579582601295299/5552579582601295299_signature_31.png
| 2023-03-22 | 2023-03-14 11:46 | Occidental Mindoro Provincial Hospital | Mamburao | 4B | Occidental Mindoro | Covid laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23936 | MACARE-HBDC | oksicovid19lab@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 215 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1668 ASSAY: Xpert SARS-CoV-2, COUNT: 1377 |  | GX-XVI-D | 845005 | For maintenance | Performed maintenance | Performed maintenance Checked module Generated reports | Unit is operational. Modules A1/903695, C4/902997, D1/902984, D3/903602, D4/902994 not operational. | With Shed Daroya |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579582605115639/5552579582605115639_signature_28.png | LARRY AMITA | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579582605115639/5552579582605115639_signature_31.png
| 2023-03-22 | 2023-03-17 09:11 | Puerto Galera Molecular Laboratory | Poblacion, Puerto Galera | 4B | Oriental Mindoro | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23937 | PBSP | tessur669@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 53 ASSAY: Xpert MTB/RIF Ultra, COUNT: 238 ASSAY: Xpert SARS-CoV-2, COUNT: 228 |  | GX-IV-D-2 | 834810 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579582608721382/5552579582608721382_signature_28.png | Maria Teresa A. Sur | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579582608721382/5552579582608721382_signature_31.png
| 2023-03-22 | 2023-03-14 13:51 | San Sebastian District Hospital | Balud-Busuran Rd. Sablayan | 4B | Occidental Mindoro | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23938 | PBSP | ssdhlaboratory@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 365 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1639 |  | GX-IV-D | 812791 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579592607884965/5552579592607884965_signature_28.png | Anthony Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5552579592607884965/5552579592607884965_signature_31.png
| 2023-03-22 | 2023-03-16 10:40 | San Mariano Rural Health Unit | Sta Filomena, San Mariano, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23931 | PBSP | jonalyndumlao78@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1279 | 2023-03-16 | Genexpert IV | 812789 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement (B4) Performed Functionality Test | Unit is Operational | Module Replacement = 677512 - 210039596 | Genexpert Module = 210039596 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062022842413/5552573062022842413_signature_28.png | Jonalyn F. DUMLAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062022842413/5552573062022842413_signature_31.png
| 2023-03-22 | 2023-03-17 16:16 | Panopdopan District Hospital | Lamut, Ifugao | CAR | Ifugao | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23933 | PBSP | ilamawa.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1242 | 2023-03-17 | Genexpert IV | 807617 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Calibration (A3, A4 = Passed) Performed Functionality Test | Unit is Operational | Module Replacement = 662356 - 672908 | Genexpert Module = 672908 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062023976881/5552573062023976881_signature_28.png | Iolanthe Dawn L. Amawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062023976881/5552573062023976881_signature_31.png
| 2023-03-22 | 2023-03-21 13:14 | Santa Rosa Rural Health | Rizal st. Santa Rosa, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23934 | DOH | addc1995@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 85 | 2023-03-21 | Genexpert IV | 831599 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062024144716/5552573062024144716_signature_28.png | Moy Espino | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062024144716/5552573062024144716_signature_31.png
| 2023-03-22 | 2023-03-20 14:18 | San Isidro Rural Health Unit | Poblacion, San Isidro, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23932 | PBSP | rmumali44@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 523 | 2023-03-20 | Genexpert IV | 805107 | UPS not working | Troubleshoot the UPS No power | Performed Maintenance Gathered Reports Performed Module Replacement (644491 - 714934) Performed Functionality Test | Unit is Operational | UPS for Replacement | Genexpert Module - 714934 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062024240368/5552573062024240368_signature_28.png | ROWENA M. UMALI | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062024240368/5552573062024240368_signature_31.png
| 2023-03-22 | 2023-03-17 10:34 | Maddela District Hospital | Diduyon, Maddela, Quirino | 2 | Quirino | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23931 | DOH | mdhlaboratory00@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 182 ASSAY: Xpert MTB/RIF, COUNT: 142 | 2023-03-17 | Genexpert IV | 838469 | Error 2008 3 modules | Performed Xpertcheck on all modules | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062026949330/5552573062026949330_signature_28.png | Kathlen Joy P. Budomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5552573062026949330/5552573062026949330_signature_31.png
| 2023-03-22 | 2023-03-20 14:18 | San Isidro Rural Health Unit | Poblacion, San Isidro, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23929 | PBSP | rmumali44@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 523 | 2023-03-20 | Genexpert IV | 805107 | UPS not working | Troubleshoot the UPS No power | Performed Maintenance Gathered Reports Performed Module Replacement (644491 - 714934) Performed Functionality Test | Unit is Operational | UPS for Replacement | Genexpert Module - 714934 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932021894826/5552572932021894826_signature_28.png | ROWENA M. UMALI | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932021894826/5552572932021894826_signature_31.png
| 2023-03-22 | 2023-03-17 16:16 | Panopdopan District Hospital | Lamut, Ifugao | CAR | Ifugao | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23928 | PBSP | ilamawa.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1242 | 2023-03-17 | Genexpert IV | 807617 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Calibration (A3, A4 = Passed) Performed Functionality Test | Unit is Operational | Module Replacement = 662356 - 672908 | Genexpert Module = 672908 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932023319642/5552572932023319642_signature_28.png | Iolanthe Dawn L. Amawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932023319642/5552572932023319642_signature_31.png
| 2023-03-22 | 2023-03-17 10:34 | Maddela District Hospital | Diduyon, Maddela, Quirino | 2 | Quirino | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23927 | DOH | mdhlaboratory00@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 182 ASSAY: Xpert MTB/RIF, COUNT: 142 | 2023-03-17 | Genexpert IV | 838469 | Error 2008 3 modules | Performed Xpertcheck on all modules | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932025590853/5552572932025590853_signature_28.png | Kathlen Joy P. Budomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932025590853/5552572932025590853_signature_31.png
| 2023-03-22 | 2023-03-16 10:40 | San Mariano Rural Health Unit | Sta Filomena, San Mariano, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23927 | PBSP | jonalyndumlao78@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1279 | 2023-03-16 | Genexpert IV | 812789 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement (B4) Performed Functionality Test | Unit is Operational | Module Replacement = 677512 - 210039596 | Genexpert Module = 210039596 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932026691998/5552572932026691998_signature_28.png | Jonalyn F. DUMLAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932026691998/5552572932026691998_signature_31.png
| 2023-03-22 | 2023-03-21 13:14 | Santa Rosa Rural Health | Rizal st. Santa Rosa, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23930 | DOH | addc1995@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 85 | 2023-03-21 | Genexpert IV | 831599 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932027499431/5552572932027499431_signature_28.png | Moy Espino | https://www.jotform.com/uploads/macarewecare/222927978372067/5552572932027499431/5552572932027499431_signature_31.png
| 2023-03-21 | 2023-03-21 15:18 | Zamboanga City Medical Centre | Dr Evangelista Street, Zamboanga City, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | Installation of module | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23926 | WHO | jbcua.pbsp@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 348 ASSAY: Xpert SARS-CoV-2, COUNT: 2462 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 ASSAY: Others, COUNT: 297 | 2023-03-07 | GX 4 | 848250 | None | Installed module 210075576 | Installed module 210075576 and preventive Maintenence | All modules are available | None | 900-0807 | G 5235 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551964602025891063/5551964602025891063_signature_28.png | James peniel b. Cua | https://www.jotform.com/uploads/macarewecare/222927978372067/5551964602025891063/5551964602025891063_signature_31.png
| 2023-03-21 | 2023-03-21 15:25 | Caraga Regional Hospital | Rizal St. Brgy. Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 23925 | PBSP | abpabellan.pbsp@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF, COUNT: 286 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3764 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 |  | GX-4 | 815489 | Module C1 found error 2014 | Troubleshooting and reporting of C1 | Replacing D3 with SN: 643989 to SN: 688658 | 7 modules are working, 1 module for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551931610192817321/5551931610192817321_signature_28.png | Andrew Pabellan | https://www.jotform.com/uploads/macarewecare/222927978372067/5551931610192817321/5551931610192817321_signature_31.png
| 2023-03-21 | 2023-03-21 14:19 | Davao TB Reference Laboratory | DOH Compound, Bajada, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23924 | WHO | dtrlmark.pbsp@gmail.com | 29.7 |  | 2023-03-21 | GX-IV 4D 10 Colours | 848248 | Module B3 for Replacement | Replaced the Defective Module B3 SN: 210002512 with new SN: 2100011400 Generate IQR, Assay Stat, and System Log Archived Data test results | Perform Module Replacement | Unit is Fully operational |  | 900-0650RS | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551904355211460633/5551904355211460633_signature_28.png | Irene Abejuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5551904355211460633/5551904355211460633_signature_31.png
| 2023-03-21 | 2023-03-21 12:32 | Bulakan Rural Health Unit | Bulakan, Bulacan | NCR | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23923 | DOH | l_gemino@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 145 ASSAY: Xpert MTB/RIF Ultra, COUNT: 598 | 2023-03-20 | GXIV | 831641 | Modules A1 (754477) and A3 (754490) for replacement. | - replaced modules A1, A3 by 745864 and 635985 - checked and cleaned cartridge bay, plunger rod, filter - Generated reports | Replaced modules | Unit is operational. |  | GeneXpert Module | 9725, 9707 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551902882225269466/5551902882225269466_signature_28.png | Lilibeth Gemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5551902882225269466/5551902882225269466_signature_31.png
| 2023-03-21 | 2023-03-21 14:28 | Ramos Rural Health Unit | Poblacion South Ramos Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23922 | DOH | jadeeh8143@yahoo.com | 28 |  | 2023-02-01 | GXIV-D | 831721 | Due for Calibration | Calibration Preventive maintenance Troubleshooting | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test Troubleshoot sudden shutdown during calibration | Unit is operational except, UPS for replacement | Sudden shutdown during calibration. UPS for replacement, for recalibration | 4 Genexpert module | 9562 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551897597127007741/5551897597127007741_signature_28.png | Emily Jade Gabriel | https://www.jotform.com/uploads/macarewecare/222927978372067/5551897597127007741/5551897597127007741_signature_31.png
| 2023-03-21 | 2023-03-21 11:12 | Caloocan City Health Office | Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23921 | PBSP | caloocancentrallaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 90 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1344 ASSAY: Xpert HIV-1 Viral Load, COUNT: 124 | 2023-03-20 | GXIV | 814154 | Unit for preventive maintenance. For module replacement of modules B2 and B4. | Performed preventive maintenance. Performed module replacement on modules B2 and B4. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. |  | Gx Module B2 (632364) and B4 (635260) in replacement to defective B2 (690080) and B4 (696343). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551778108424893036/5551778108424893036_signature_28.png | Albert Igama | https://www.jotform.com/uploads/macarewecare/222927978372067/5551778108424893036/5551778108424893036_signature_31.png
| 2023-03-21 | 2023-03-21 10:40 | Moncada Rural Health Unit | Poblacion 1 Moncada Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23920 | PBSP | taming_17@yahoo.com | 30.5 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 56 | 2023-03-10 | GXIV-D | 808292 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational |  | 5 xpertcheck kits | 39468 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551773067123895415/5551773067123895415_signature_28.png | Jessa Pabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5551773067123895415/5551773067123895415_signature_31.png
| 2023-03-21 | 2023-03-21 10:40 | Moncada Rural Health Unit | Poblacion 1 Moncada Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23919 | PBSP | taming_17@yahoo.com | 30.5 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 56 | 2023-03-10 | GXIV-D | 808292 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational |  | 5 xpertcheck kits | 39468 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551772607125871219/5551772607125871219_signature_28.png | Jessa Pabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5551772607125871219/5551772607125871219_signature_31.png
| 2023-03-21 | 2023-03-21 10:49 | Malapatan Rural Health Unit | Malapatan, Sarangani | 12 | Sarangani | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23918 | DOH | rhumalapatan12@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 124 ASSAY: Xpert MTB/RIF, COUNT: 7 | 2023-03-21 | Gx-IV | 838419 | Module 2 and 3 H/W failed | Replace module 2 SN:861212 with new module SN: 871118 Perform Genexpert machine Preventive maintenance | Replace module 2 SN:861212 with new module SN: 871118 Perform Genexpert machine Preventive maintenance | 3 operational modules | Module 3 waiting for module to available | Genexpert 6 colors module | DR no. 9682 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551765583424900582/5551765583424900582_signature_28.png | Ceres Mae Jimenez,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5551765583424900582/5551765583424900582_signature_31.png
| 2023-03-21 | 2023-03-21 10:24 | Minglanilla Rhu | Ward 1, Pider, Poblacion Minglanilla | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 23917 | DOH | brendagonzagaparedes@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 584 |  | Gx4 | 838512 | Error 2008 in Module 2 and Module 3 | Clean cartridge bay, clean Icore, clean filter, perform self test in all modules Get data sys log, IQR, 3 months gxx files Back up Database | Clean cartridge bay, clean Icore, clean filter, perform self test in all modules | All modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551751384211676633/5551751384211676633_signature_28.png | Brenda Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/5551751384211676633/5551751384211676633_signature_31.png
| 2023-03-21 | 2023-03-21 09:46 | Pulilan Rural Health Unit | Pulilan, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23916 | DOH | celsptcruz19@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 107 ASSAY: Xpert MTB/RIF Ultra, COUNT: 387 | 2023-03-21 | GXIV | 831915 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. B2 failed calibration. B1, B3, B4 passed calibration. |  | XpertCheck Kit | 9517 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551739737655870106/5551739737655870106_signature_28.png | Cecile Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5551739737655870106/5551739737655870106_signature_31.png
| 2023-03-21 | 2023-03-21 09:46 | Pulilan Rural Health Unit | Pulilan, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23915 | DOH | celsptcruz19@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 107 ASSAY: Xpert MTB/RIF Ultra, COUNT: 387 | 2023-03-21 | GXIV | 831915 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. B2 failed calibration. B1, B3, B4 passed calibration. |  | XpertCheck Kit | 9517 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551739427652107233/5551739427652107233_signature_28.png | Cecile Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5551739427652107233/5551739427652107233_signature_31.png
| 2023-03-21 | 2023-03-21 09:06 | Catarman Municipal Health Office | Catarman | 10 | Camiguin | TB DOTS | YES | INSTALLATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23914 | PBSP | irishcamacho1225@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1159 |  | GeneXpert | 835063 | H/W Fail on Module B1/787847 | Replaced module New module SN/702140 | Module Replacement Defective Module SN/787847 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551730147611371954/5551730147611371954_signature_28.png | Irish cmacho | https://www.jotform.com/uploads/macarewecare/222927978372067/5551730147611371954/5551730147611371954_signature_31.png
| 2023-03-21 | 2023-03-21 09:38 | BIOPATH CAP | CEBU CITY | 7 | Cebu | MOLECULAR LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23913 | COMMERCIAL | jkee@biopath.ph | 28 | ASSAY: Xpert SARS-CoV-2, COUNT: 12906 |  | GX-XVI | 838657 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551725513427396859/5551725513427396859_signature_28.png | Mary Joy Kee | https://www.jotform.com/uploads/macarewecare/222927978372067/5551725513427396859/5551725513427396859_signature_31.png
| 2023-03-21 | 2023-03-15 14:02 | Bongabong Municipal Health Office | Aplaya 2, Bongabong | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23912 | PBSP | cortezmariel30@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 246 | 2023-03-13 | GX-IV-D | 834687 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Modules A2, A3 and A4 operational. Modules A1/786829 and A4/786837 service required. |  | XpertCheck | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716219931839/5546734716219931839_signature_28.png | Mariel A. Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716219931839/5546734716219931839_signature_31.png
| 2023-03-21 | 2023-03-15 14:02 | Bongabong Municipal Health Office | Aplaya 2, Bongabong | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23911 | PBSP | cortezmariel30@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 246 | 2023-03-13 | GX-IV-D | 834687 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Modules A2, A3 and A4 operational. Modules A1/786829 and A4/786837 service required. |  | XpertCheck | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716219931839/5546734716219931839_signature_28.png | Mariel A. Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716219931839/5546734716219931839_signature_31.png
| 2023-03-21 | 2023-03-21 09:04 | Gat Andres Bonifacio Memorial Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23910 | PBSP | sjavila04@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 956 | 2023-03-20 | GXIV | 805103 | Unit for preventive maintenance. For module replacement of module B2. | Performed preventive maintenance. Performed module replacement on module B2. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. |  | Gx module B2 (744744) in replacement to defective module B2 (652075 | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551701108426079651/5551701108426079651_signature_28.png | Shana Avila | https://www.jotform.com/uploads/macarewecare/222927978372067/5551701108426079651/5551701108426079651_signature_31.png
| 2023-03-20 | 2023-03-20 10:32 | Bicol South Luzon Sub National Reference Laboratory | BRTTH Compound, Daraga | 5 | Albay | Laboratory | NO | INSTALLATION | CEPHEID | Patrick Los Baños |  |  |  |  | 23909 | PBSP | rbalzaga.pbsp@gmail.com | 25 |  |  | GXIV | 110016954 | Unit for installation | Performed unit installation | Performed unit installation | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551314682235103195/5551314682235103195_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5551314682235103195/5551314682235103195_signature_31.png
| 2023-03-20 | 2023-03-20 10:44 | Tuburan Rural Health Unit | Poblacion,Tuburan,Cebu | 7 | Cebu | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Jovy Noel |  |  |  |  | 23907 | PBSP | marielgracia20@gmail.com | 23 |  | 2023-03-06 | Gx4 | 815485 | Module B2 Hardware Fail Error Code 2034,1006,4011 | Not Ok to use | Perform Preventive Maintenace, Clean all modules,Clean all cartrige bay and plunger rod maintenance and resit all wire from the modules. Save IQR and Print IQR and Archive test report. | Not ok to use module B2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551073367223950579/5551073367223950579_signature_28.png | Mariel R. Gracia | https://www.jotform.com/uploads/macarewecare/222927978372067/5551073367223950579/5551073367223950579_signature_31.png
| 2023-03-20 | 2023-03-17 16:16 | Panopdopan District Hospital | Lamut, Ifugao | CAR | Ifugao | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23905 | PBSP | ilamawa.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1242 | 2023-03-17 | Genexpert IV | 807617 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Calibration (A3, A4 = Passed) Performed Functionality Test | Unit is Operational | Module Replacement = 662356 - 672908 | Genexpert Module = 672908 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551029881832317882/5551029881832317882_signature_28.png | Iolanthe Dawn L. Amawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5551029881832317882/5551029881832317882_signature_31.png
| 2023-03-20 | 2023-03-20 14:18 | San Isidro Rural Health Unit | Poblacion, San Isidro, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23906 | PBSP | rmumali44@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 523 | 2023-03-20 | Genexpert IV | 805107 | UPS not working | Troubleshoot the UPS No power | Performed Maintenance Gathered Reports Performed Module Replacement (644491 - 714934) Performed Functionality Test | Unit is Operational | UPS for Replacement | Genexpert Module - 714934 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551029881838345187/5551029881838345187_signature_28.png | ROWENA M. UMALI | https://www.jotform.com/uploads/macarewecare/222927978372067/5551029881838345187/5551029881838345187_signature_31.png
| 2023-03-20 | 2023-03-17 10:34 | Maddela District Hospital | Diduyon, Maddela, Quirino | 2 | Quirino | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23903 | DOH | mdhlaboratory00@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 182 ASSAY: Xpert MTB/RIF, COUNT: 142 | 2023-03-17 | Genexpert IV | 838469 | Error 2008 3 modules | Performed Xpertcheck on all modules | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck Result = All Modules Passed | Xpertcheck 4pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551029881839152795/5551029881839152795_signature_28.png | Kathlen Joy P. Budomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5551029881839152795/5551029881839152795_signature_31.png
| 2023-03-20 | 2023-03-16 10:40 | San Mariano Rural Health Unit | Sta Filomena, San Mariano, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23904 | PBSP | jonalyndumlao78@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1279 | 2023-03-16 | Genexpert IV | 812789 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement (B4) Performed Functionality Test | Unit is Operational | Module Replacement = 677512 - 210039596 | Genexpert Module = 210039596 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551029881839377951/5551029881839377951_signature_28.png | Jonalyn F. DUMLAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5551029881839377951/5551029881839377951_signature_31.png
| 2023-03-20 | 2023-03-20 13:29 | Malungon Rural Health Unit | Malungon, sarangani | 12 | Sarangani | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23902 | DOH | ulbenarioe@yahoo.com | 24 |  |  | Gx-IV | 838546 | Genexpert CPU problem | Perform Genexpert CPU Troubleshooting | Perform Genexpert CPU Troubleshooting | Genexpert CPU subject for replacement due to power supply/ mother board problem |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551006561428433164/5551006561428433164_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5551006561428433164/5551006561428433164_signature_31.png
| 2023-03-20 | 2023-03-17 13:39 | Maysan 3S Health Center | F. Alarcon St., Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23899 | PBSP | cesguardiano@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 86 ASSAY: Xpert HIV-1 Viral Load, COUNT: 5 |  | GXIV | 814151 | Module A2 for replacement | - replaced A2 by 701103 - Generated reports - checked and cleaned cartridge bay, plunger rod, filter - performed functionality test | Replaced module A2 | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551001391213019031/5551001391213019031_signature_28.png | Cecilia Guardiano | https://www.jotform.com/uploads/macarewecare/222927978372067/5551001391213019031/5551001391213019031_signature_31.png
| 2023-03-20 | 2023-03-17 11:29 | Dr. Jose N. Rodriguez Memorial Hospital Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23901 | PBSP | diazjimvianne@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 15 ASSAY: Xpert HPV, COUNT: 13 ASSAY: Xpert MTB/RIF, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2107 ASSAY: Xpert SARS-CoV-2, COUNT: 14266 ASSAY: Xpert HIV-1 Viral Load, COUNT: 119 ASSAY: Others, COUNT: 9 |  | GXIV | 707523 | Module A1 for replacement | - replaced module A1 by 684743 - checked and cleaned cartridge bay, plunger rod - generated reports | Replaced module A1 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551001391214772310/5551001391214772310_signature_28.png | Jim Vianne Diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5551001391214772310/5551001391214772310_signature_31.png
| 2023-03-20 | 2023-03-17 11:25 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23898 | PBSP | diazjimvianne@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 15 ASSAY: Xpert HPV, COUNT: 13 ASSAY: Xpert MTB/RIF, COUNT: 12 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2107 ASSAY: Xpert SARS-CoV-2, COUNT: 14266 |  | GXIV | 814159 | Module B4 for replacement | - replaced B4 by 775634 - checked and cleaned cartridge bay, plunger rod, filter - Generated reports | Replaced module B4 | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551001391216056826/5551001391216056826_signature_28.png | Jim Vianne Diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5551001391216056826/5551001391216056826_signature_31.png
| 2023-03-20 | 2023-03-17 11:17 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Molecular Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 23900 | PBSP | diazjimvianne@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF, COUNT: 164 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1183 ASSAY: Xpert SARS-CoV-2, COUNT: 5422 |  | GXIV | 834431 | Module D3 for replacement | - replaced module D3 by 775961 - Generated reports - checked and cleaned cartridge bay, plunger rod, filter | Replaced module D3 | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5551001391217300683/5551001391217300683_signature_28.png | Jim Vianne Diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5551001391217300683/5551001391217300683_signature_31.png
| 2023-03-20 | 2023-03-17 14:12 | Imus City Health Office III | Green Gate Subd., Imus | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23897 | PBSP | tess19_chua@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1062 |  | GXIV-D | 838166 | Module 1 for Replacement | Module Replacement | Replaced Module 1 Performed Maintenance Generate Reports | All Modules are Operational | B1: 807095 -> 695564 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995821219328494/5550995821219328494_signature_28.png | Marites M. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995821219328494/5550995821219328494_signature_31.png
| 2023-03-20 | 2023-03-17 12:01 | De La Salle Health and Science Institute | Congressional Ave., Dasmariñas City | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23896 | CONSORTIUM | acpabruada@dlshsi.edu.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 932 | 2023-03-16 | GX16-D | 842166 | B1 with Temperature Error Unit is due for XpertCheck | Module Replacement XpertCheck | Replaced Module B1 Performed Maintenance Performed XpertCheck Generate Reports | All Modules are Operational. | B1: 880285 -> 210019244 , XC: All Passed | GX Module Kit 6-C | DR 9700 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995821211137674/5550995821211137674_signature_28.png | Angelita C. Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995821211137674/5550995821211137674_signature_31.png
| 2023-03-20 | 2023-03-15 14:17 | Naic Rural Health Unit | Capt. C Nazareno, Naic | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23891 | PBSP | jengcasamar2007@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1357 |  | GXIV-D | 812006 | 2 Modules Failed in XpertCheck | Module Replacement | Replaced 2 Modules Performed Maintenance Generate Reports | All Modules are Operational | 688783, 682790 -> 687250, 684672 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201212774963/5550995201212774963_signature_28.png | CASAMAR,JENNIFER | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201212774963/5550995201212774963_signature_31.png
| 2023-03-20 | 2023-03-15 14:14 | Naic Rural Health Unit | Capt. C Nazareno St., Naic | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23893 | PBSP | jengcasamar2007@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1357 |  | GXIV-D | 807567 | 2 Modules Failed im XpertCheck | Module Replacement | Replaced Modules 1 and 2 Performed Maintenance Generate Reports | All Modules are Operational | 738685, 731238 -> 697534, 691190 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201212942128/5550995201212942128_signature_28.png | Casamar,Jennifer | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201212942128/5550995201212942128_signature_31.png
| 2023-03-20 | 2023-03-16 11:36 | Lingap Kalusugan CHU I - Alima | Evangelista St., Alima | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23892 | PBSP | eng112803@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1258 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GXIV-D | 838185 | Module 2 with Temperature Error | Module Replacement | Module Replacement in B2 Performed Maintenance Generate Reports | All Modules are Operational | 807163->745780 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201216703777/5550995201216703777_signature_28.png | Aide E. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201216703777/5550995201216703777_signature_31.png
| 2023-03-20 | 2023-03-16 10:41 | Pamplona Uno Health Center | Alabang-Zapote Road., Pamplona Uno | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23894 | PBSP | carolinevecino5821@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 864ca |  | GXIV-D | 838541 | 3 Modules Failed in XpertCheck | Module Replacement | Replaced Modules 1,3,4 Performed Maintenance Generate Reports | All Modules are Operational | 862138, 862319, 862136 -> 657495, 658503, 661004 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201216759705/5550995201216759705_signature_28.png | Caroline R. Vecino | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201216759705/5550995201216759705_signature_31.png
| 2023-03-20 | 2023-03-17 11:56 | De La Salle Health and Science Institute | Congressional Ave., Dasmariñas City | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23895 | CONSORTIUM | acpabruada@dlshsi.edu.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 396 | 2023-03-16 | GXIV-D | 812380 | Module 3 H/W Failed Unit is due for XpertCheck | Module Replacement XpertCheck | Replaced Module 3 Performed XpertCheck Performed Maintenance Generate Reports | 3 Modules Passed and 1 Newly Replaced | B3: 802293-> 210005363 | GX Module Kit 6-C | DR 9700 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201217542236/5550995201217542236_signature_28.png | Angelita C. Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5550995201217542236/5550995201217542236_signature_31.png
| 2023-03-20 | 2023-03-20 12:24 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23890 | PBSP | spmcrhivda@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1602 ASSAY: Xpert HIV-1 Viral Load, COUNT: 350 | 2023-03-20 | GX-IV 4D | 808298 | Module B2 and B3 for Replacement | Replace the defective Module B2 S/N: 602506 with new Module S/N: 603064 Replace the defective Module B3 S/N: 662625 with new Module S/N: 644429 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Module Replacement and Preventive Maintenance | Module B1 S/N: 671900 encounter temperature Error, Data will be sent to Cepheid for further assessment |  | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550965306216900358/5550965306216900358_signature_28.png | Flordeliza B. Go RMT, MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/5550965306216900358/5550965306216900358_signature_31.png
| 2023-03-20 | 2023-03-20 12:24 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23889 | PBSP | spmcrhivda@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1602 ASSAY: Xpert HIV-1 Viral Load, COUNT: 350 | 2023-03-20 | GX-IV 4D | 808298 | Module B2 and B3 for Replacement | Replace the defective Module B2 S/N: 602506 with new Module S/N: 603064 Replace the defective Module B3 S/N: 662625 with new Module S/N: 644429 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log | Perform Module Replacement and Preventive Maintenance | Module B1 S/N: 671900 encounter temperature Error, Data will be sent to Cepheid for further assessment |  | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550965106213222744/5550965106213222744_signature_28.png | Flordeliza B. Go RMT, MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/5550965106213222744/5550965106213222744_signature_31.png
| 2023-03-20 | 2023-03-20 11:43 | Tarlac Provincial Hospital | Hospital drive San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23888 | PBSP | molecularlab.tph@gmail.com | 25.5 |  | 2023-02-24 | GXIV-D | 805108 | C3 for module replacement | Preventive maintenance Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | Genexpert module | 39293 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550934931021352969/5550934931021352969_signature_28.png | Julie B. Guillermo | https://www.jotform.com/uploads/macarewecare/222927978372067/5550934931021352969/5550934931021352969_signature_31.png
| 2023-03-20 | 2023-03-20 11:43 | Tarlac Provincial Hospital | Hospital drive San Vicente, Tarlac City | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23887 | PBSP | molecularlab.tph@gmail.com | 25.5 |  | 2023-02-24 | GXIV-D | 805108 | C3 for module replacement | Preventive maintenance Module replacement | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | Genexpert module | 39293 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550934871029191746/5550934871029191746_signature_28.png | Julie B. Guillermo | https://www.jotform.com/uploads/macarewecare/222927978372067/5550934871029191746/5550934871029191746_signature_31.png
| 2023-03-20 | 2023-03-20 10:32 | Bicol South Luzon Sub National Reference Laboratory | BRTTH Compound, Daraga | 5 | Albay | Laboratory | NO | INSTALLATION | CEPHEID | Patrick Los Baños |  |  |  |  | 23886 | PBSP | rbalzaga.pbsp@gmail.com | 25 |  |  | GXIV | 110016954 | Unit for installation | Performed unit installation | Performed unit installation | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550904900509451465/5550904900509451465_signature_28.png | Richard Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5550904900509451465/5550904900509451465_signature_31.png
| 2023-03-20 | 2023-03-20 09:49 | Santa Ignacia Rural Health Unit | Ignacio St. Poblacion West Santa Ignacia Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23885 | DOH | jdi0973@yahoo.com | 29.8 | ASSAY: Xpert MTB/RIF, COUNT: 118 ASSAY: Xpert MTB/RIF Ultra, COUNT: 205 | 2023-02-02 | GXIV-D | 831909 | A2 error 2035 A4 error 1001,4012 Due for calibration | Calibration Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All passed | 4 Xpertcheck kits | 9571 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5550876104812329291/5550876104812329291_signature_28.png | Jesus D. Ibarra Jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5550876104812329291/5550876104812329291_signature_31.png
| 2023-03-18 | 2023-03-17 00:57 | Naujan Municipak Health Office | Brgy Santiago, Naujan | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23884 | DOH | baldoza.shane25@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 410 ASSAY: Xpert MTB/RIF Ultra, COUNT: 258 | 2023-03-07 | GXIV | 838226 | Unit for preventive maintenance. Unit for Xpertcheck. | Preventive maintenance and Xpertcheck. | Performed preventive maintenance . Performed Xpertcheck. Generated reports | Unit is operational. Modules B2 and B3, service required. | Used 1 spare Xpertcheck | Xpertcheck | DR 9668 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201215917229473/5549201215917229473_signature_28.png | Shiena Roanne Baldoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201215917229473/5549201215917229473_signature_31.png
| 2023-03-18 | 2023-03-15 11:36 | Oriental Mindoro Southern District Hospital | Western Nautical Highway, Roxas | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23880 | PBSP | markedisonmartinez@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 254 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1759 |  | GX-IV-D | 807590 | Modules B1, B2 and B4 for replacement. | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025913220294/5549201025913220294_signature_28.png | Mark Edison G. MARTINEZ, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025913220294/5549201025913220294_signature_31.png
| 2023-03-18 | 2023-03-15 14:02 | Bulalacao Municipal Health Office | Aplaya 2, Bongabong | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23883 | PBSP | cortezmariel30@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 246 | 2023-03-13 | GX-IV-D | 834687 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Modules A2, A3 and A4 operational. Modules A1/786829 and A4/786837 service required. |  | XpertCheck | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025913833698/5549201025913833698_signature_28.png | Mariel A. Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025913833698/5549201025913833698_signature_31.png
| 2023-03-18 | 2023-03-16 23:46 | Victoria Rural Health Unit | Victoria | 4B | Oriental Mindoro | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23881 | PBSP | lractrebor24@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 185 ASSAY: Xpert MTB/RIF Ultra, COUNT: 604 |  | GX-IV-D-2 | 835178 | For maintenance | Performed maintenance | Performed maintenance Generated reports Checked module B2 | Module B1 operational. Module B2 not operational, for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025914862090/5549201025914862090_signature_28.png | Carl Robert B. Racellis | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025914862090/5549201025914862090_signature_31.png
| 2023-03-18 | 2023-03-16 09:13 | Pinamalayan Rural Health Unit | Municipal Compound, Pinamalayan | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23882 | DOH | mholaboratory.pinamalayan@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 216 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1057 | 2023-03-07 | GX-IV-D | 838230 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Module A1 not operational. All modules passed. |  | XpertCheck Kit | DR# 9669 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025915778622/5549201025915778622_signature_28.png | Dave Erwin M. Festin, RN, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025915778622/5549201025915778622_signature_31.png
| 2023-03-18 | 2023-03-15 09:18 | Bulalacao Municipal Health Center | Poblacion, Bulalacao | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23880 | PBSP | queence.macaraig@gmail.com | 25 |  | 2023-03-15 | GX-IV-D-2 | 835185 | For XpertCheck | Performed maintenance | Performed maintenance. Run XpertCheck Generated reports Installed new UPS | All modules operational. Module B1 passed, Module B2 service required. |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025918411513/5549201025918411513_signature_28.png | Queencelyn Festin | https://www.jotform.com/uploads/macarewecare/222927978372067/5549201025918411513/5549201025918411513_signature_31.png
| 2023-03-18 | 2023-03-17 15:37 | Batangas Medical Center | Kumintang Ibaba Batangas City | 4A | Batangas | TB Culture Lab | YES | CALIBRATION | CEPHEID | Harold Toco |  |  |  |  | 23879 | PBSP | batmc.tbculturelab@gmail.com | 30 |  | 2023-03-17 | Gx4 | 834627 | For Xpertcheck | Performed xpertcheck and maintenanca | Performed Xpertcheck and mauntenance | 3 modules ooerational 1 mod with HW failed | 3 modules passed the xpertcheck | 1Kit of Xpertcheck | PBPP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548988855362078800/5548988855362078800_signature_28.png | Patricia Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5548988855362078800/5548988855362078800_signature_31.png
| 2023-03-17 | 2023-03-17 14:20 | Lung center of the Philippines | Quezon ave, Quezon city | NCR | NCR 1st - Manila | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 23878 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | MCS + | 08H188 | As per end user they encounter liquid detected in the chamber of centrifuge | Check centrifuge, check sensor, cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed functionality test | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548618042772690351/5548618042772690351_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/5548618042772690351/5548618042772690351_signature_31.png
| 2023-03-17 | 2023-03-17 14:25 | Lung center of the Philippines | Quezon ave, Quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 23878 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | MCS + | 02D166 | As per end user they encounter autoshut down of the machine when all tubings are set up and start pressing the draw button. | Check UPS Voltage output,check power cord,cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,centrifuge,performed functionality test | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548618042776374052/5548618042776374052_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/5548618042776374052/5548618042776374052_signature_31.png
| 2023-03-17 | 2023-03-17 14:54 | Batangas Medical Center | Kumintang Ibaba Batangas City | 4A | Batangas | TB Culture Laboratory | NO | INSTALLATION | CEPHEID | Harold Toco |  |  |  |  | 23877 | PBSP | batmc.tbculturelab@gmail.com | 29 |  | 2023-03-17 | Gx4 10C | 110016955 | For Installation | Performed unit installation | Performed unit installation | All 4 modules operational |  | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548481278023878465/5548481278023878465_signature_28.png | Patricia Louisse A.Baltazat | https://www.jotform.com/uploads/macarewecare/222927978372067/5548481278023878465/5548481278023878465_signature_31.png
| 2023-03-17 | 2023-03-17 14:58 | Batangas Medical Center | Kumintang Ibaba Batangas City | 4A | Batangas | TB Culture Lab | YES | CALIBRATION | CEPHEID | Harold Toco |  |  |  |  | 23876 | PBSP | batmc.tbculturelab@gmail.com | 28 |  | 2023-03-17 | Gx16 | 844120 | For Xpertcheck Calibration and Maintenance | Performed maintenance and calibration. | Performed maintenance and calibration. | 15 modules operational but 2 of them service required..1 with HW failed. |  | 4 kits Xpertcheck | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548481268024687118/5548481268024687118_signature_28.png | Patricia Louisse Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5548481268024687118/5548481268024687118_signature_31.png
| 2023-03-17 | 2023-03-17 14:45 | Camalig MHO | R. Nasol St., Camalig | 5 | Albay | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 23875 | PBSP | lgrageda@gmail.com | 24 |  | 2023-03-08 | GXIV | 835187 | Module B2 for replacement | Replaced module B2 (717205 > 210046961) | Performed module replacement, maintenance cleaning | Unit is operational |  | Genexpert module 6c | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548459106012210005/5548459106012210005_signature_28.png | Leilani M. Grageda | https://www.jotform.com/uploads/macarewecare/222927978372067/5548459106012210005/5548459106012210005_signature_31.png
| 2023-03-17 | 2023-03-17 14:22 | Mindanao Central Sanitarium | Pasabolong, Zamboanga city, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23874 | DOH | ybioamie@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1851 | 2023-03-17 | GX 4 | 821214 | Module 1 and 2 error 4012 | Preventive Maintenence | Preventive Maintenence | Module 1 and 2 are not in use, Module 3 and 4 are in use. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548455610212781860/5548455610212781860_signature_28.png | Amie Ybio | https://www.jotform.com/uploads/macarewecare/222927978372067/5548455610212781860/5548455610212781860_signature_31.png
| 2023-03-17 | 2023-03-17 14:25 | Iligan Society of Internists | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | YES | REPAIR | CEPHEID | Wynnona Reoma |  |  |  |  | 23873 | PBSP | dayolaermajessica@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2635 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 | 2023-03-17 | GXIV-4D | 707842 | Module D3 - Xpert Check Service Required | -Replace defective module D3/675780 with new module SN: 693982 | #NAME? | All four modules are functional. |  | 900-GX 6 Colors | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548438842888648213/5548438842888648213_signature_28.png | Erma Jessica Dayola | https://www.jotform.com/uploads/macarewecare/222927978372067/5548438842888648213/5548438842888648213_signature_31.png
| 2023-03-17 | 2023-03-17 13:38 | Bayugan City Health Office | Brgy. Poblacion, Bayugan City | 13 | Agusan del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 23872 | PBSP | diolaugay@gmail.com | 21.67 | ASSAY: Xpert MTB/RIF, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1790 |  | GX-4 | 815391 | None | None | Replacing Module B3 with SN: 672206 to SN: 694385 | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548414957919767450/5548414957919767450_signature_28.png | Lourdes Diola | https://www.jotform.com/uploads/macarewecare/222927978372067/5548414957919767450/5548414957919767450_signature_31.png
| 2023-03-17 | 2023-03-16 13:23 | Duay health center | Hidago st.,Pasay city | NCR | NCR 4th - Pasay | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23871 | PBSP | rhemejetbotaquia@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1376 | 2023-03-16 | GX4 | 812833 | For module replacement 685785 - 635266, 685796 - 638162 | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | Operational, b1,b3,b4 | B2 as per end user they invalid error | 900 - 0491rs | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548402354424335894/5548402354424335894_signature_28.png | Rheme Tañeza | https://www.jotform.com/uploads/macarewecare/222927978372067/5548402354424335894/5548402354424335894_signature_31.png
| 2023-03-17 | 2023-03-17 12:51 | BACOLOD CITY HEALTH OFFICE | BBB St., Bacolod City | 6 | Negros Occidental | TB Laboratory | YES | ON-SITE ASSISTANCE Inspection/Evaluation of New Site for Genexpert machines Units | CEPHEID | Gerald Fritz Selga |  |  |  |  | 23870 | PBSP | cyrilcordova0515@gmail.com | 27 | ASSAY: , COUNT: N/A |  | GX-IV | 804838 | For evaluation/site inspection of new location for Genexpert machines (for transfer from old laboratory) | Site Inspection | - Check new room for Genexpert units; - Suggest location and quantity of electrical wall outlets; - Suggest location for ACU and exhaust fan; - Check functionality of Solar Generator; | - Solar generator battery defective with worn out plates, tried recharge of battery but evaluation of battery plates showed damaged plates and rendered battery defective; - Solar generator inverter and charge controller for testing upon availability of new battery. - Solar panels for testing. | Replacement battery needed with following specifications: 12 V, 100 AH, model SMN100L (Motolite) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548387172368686727/5548387172368686727_signature_28.png | CYRIL CORDOVA, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5548387172368686727/5548387172368686727_signature_31.png
| 2023-03-17 | 2023-03-17 11:27 | Poro Rural Health Unit | Poro Camotes | 7 | Cebu | Tb Laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 23869 | DOH | avilacharmaine6@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 259 | 2023-03-17 | Gx4 | 838279 | Error 2014 | Replaced New Module | Replaced New modules 608877 and 749916 Pulled out defective mdoles 808038 and 808039 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548333044419470630/5548333044419470630_signature_28.png | Charmaine Apat | https://www.jotform.com/uploads/macarewecare/222927978372067/5548333044419470630/5548333044419470630_signature_31.png
| 2023-03-17 | 2023-03-17 10:48 | Malilipot RHU | Rizal St, Malilipot | 5 | Albay | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | Patrick Los Baños |  |  |  |  | 23868 | PBSP | faith.buban2124@gmail.com | 26 |  | 2023-03-08 | GXIV | 808280 | Modules B1 and B3 for replacement, for XpertCheck | Replaced modules B1 and B3 (655916 > 870447, 621226 > 870950) | Performed module replacement, maintenance cleaning, XpertCheck | Unit is operational, Module B4 failed XpertCheck |  | Genexpert module 6c | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5548312378714197355/5548312378714197355_signature_28.png | Alyssa Nicole Marina B. Jamora | https://www.jotform.com/uploads/macarewecare/222927978372067/5548312378714197355/5548312378714197355_signature_31.png
| 2023-03-16 | 2023-03-16 11:47 | Marilao Rural Health Unit | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23867 | DOH | lab.rhu1marilao@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 443 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1610 ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 | 2023-01-20 | GXIV | 831897 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. A1, A2 failed Xpertcheck. For reporting. |  | Xpertcheck Kit | 9523 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547893372223399738/5547893372223399738_signature_28.png | JONATHAN JENILO | https://www.jotform.com/uploads/macarewecare/222927978372067/5547893372223399738/5547893372223399738_signature_31.png
| 2023-03-16 | 2023-03-16 14:36 | Sta. Maria Rural Health Unit | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23866 | PBSP | mhosantamaria@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 220 ASSAY: Xpert MTB/RIF Ultra, COUNT: 909 |  | GXIV | 838866 | Module B3 for replacement. | #NAME? | Replaced module B3 (861660) | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547893372226160914/5547893372226160914_signature_28.png | Athena Daquioag | https://www.jotform.com/uploads/macarewecare/222927978372067/5547893372226160914/5547893372226160914_signature_31.png
| 2023-03-16 | 2023-03-16 11:47 | Marilao Rural Health Unit | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23865 | DOH | lab.rhu1marilao@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 443 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1610 ASSAY: Xpert HIV-1 Viral Load, COUNT: 73 | 2023-01-20 | GXIV | 831897 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. A1, A2 failed Xpertcheck. For reporting. |  | Xpertcheck Kit | 9523 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547884362223891139/5547884362223891139_signature_28.png | JONATHAN JENILO | https://www.jotform.com/uploads/macarewecare/222927978372067/5547884362223891139/5547884362223891139_signature_31.png
| 2023-03-16 | 2023-03-16 14:36 | Sta. Maria Rural Health Unit | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23864 | PBSP | mhosantamaria@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 220 ASSAY: Xpert MTB/RIF Ultra, COUNT: 909 |  | GXIV | 838866 | Module B3 for replacement. | #NAME? | Replaced module B3 (861660) | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547884362222214184/5547884362222214184_signature_28.png | Athena Daquioag | https://www.jotform.com/uploads/macarewecare/222927978372067/5547884362222214184/5547884362222214184_signature_31.png
| 2023-03-16 | 2023-03-16 13:23 | Duay health center | Hidago st.,Pasay city | NCR | NCR 4th - Pasay | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23863 | PBSP | rhemejetbotaquia@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1376 | 2023-03-16 | GX4 | 812833 | For module replacement 685785 - 635266, 685796 - 638162 | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | Operational, b1,b3,b4 | B2 as per end user they invalid error | 900 - 0491rs | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547847582779955042/5547847582779955042_signature_28.png | Rheme Tañeza | https://www.jotform.com/uploads/macarewecare/222927978372067/5547847582779955042/5547847582779955042_signature_31.png
| 2023-03-16 | 2023-03-16 17:01 | St lukes medical center | E.Rodriguez, Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23862 | COMMERCIAL | megapo@stlukes.com.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 237 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2933 | 2023-03-16 | GX4 | 807155 | For module replacement 603082 - 210072550, 613835- 210072753 | Performed module replacement,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | All module operational |  | 900-0508 | 9698 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547845672772109982/5547845672772109982_signature_28.png | Eunice Expectacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5547845672772109982/5547845672772109982_signature_31.png
| 2023-03-16 | 2023-03-16 17:01 | St lukes medical center | E.Rodriguez, Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23861 | COMMERCIAL | megapo@stlukes.com.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 237 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2933 | 2023-03-16 | GX4 | 807155 | For module replacement 603082 - 210072550, 613835- 210072753 | Performed module replacement,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | All module operational |  | 900-0508 | 9698 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547845012779077142/5547845012779077142_signature_28.png | Eunice Expectacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5547845012779077142/5547845012779077142_signature_31.png
| 2023-03-16 | 2023-03-16 17:01 | St lukes medical center | E.Rodriguez, Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23860 | COMMERCIAL | biomed.qc@stlukes.com.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 237 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2933 | 2023-03-16 | GX4 | 807155 | For module replacement 603082 - 210072550, 613835- 210072753 | Performed module replacement,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | All module operational |  | 900-0508 | 9698 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547844762779241089/5547844762779241089_signature_28.png | Eunice Expectacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5547844762779241089/5547844762779241089_signature_31.png
| 2023-03-16 | 2023-03-16 17:01 | St lukes medical center | E.Rodriguez, Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23859 | COMMERCIAL | biomed.qc@stlukes.com.ph | 30 | ASSAY: Xpert CT/NG, COUNT: 237 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2933 | 2023-03-16 | GX4 | 807155 | For module replacement 603082 - 210072550, 613835- 210072753 | Performed module replacement,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | All module operational |  | 900-0508 | 9698 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547843832777471351/5547843832777471351_signature_28.png | Eunice Expectacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5547843832777471351/5547843832777471351_signature_31.png
| 2023-03-16 | 2023-03-16 10:20 | Faiview Health Center | Dahlia st., Fairview Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23858 | PBSP | fairviewclinicallab@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 15 ASSAY: Xpert HPV, COUNT: 461 ASSAY: Xpert MTB/RIF, COUNT: 343 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1108 | 2023-03-16 | GX4 | 838876 | For module replacement 768659 - 684563, 862432 - 684983 | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900-0491 rs | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842822771640555/5547842822771640555_signature_28.png | Hiromi O. Leandado | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842822771640555/5547842822771640555_signature_31.png
| 2023-03-16 | 2023-03-15 16:38 | Rodriguez RHU | J.P Rizal Rodriguez rizal | 4A | Rizal | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23856 | PBSP | karlbanal25@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2436 ASSAY: Xpert HIV-1 Viral Load, COUNT: 89 |  | GX4 | 838894 | Workstation no display booting problem | Check workstation,reset ram, reset workstation,check cable connection | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842292773550971/5547842292773550971_signature_28.png | Karl Michael Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842292773550971/5547842292773550971_signature_31.png
| 2023-03-16 | 2023-03-15 14:24 | E.Rodriguez health center | Pajo st., Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23857 | DOH | clairebaptista32786@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 164 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1154 | 2023-03-13 | GX4 | 831848 | UPS no power | Replaced ups, test run unit,test back up supply of ups | UPS replacement | UPS operational,modules 1,2,3 disabled failed xpertcheck |  | Powercom ups | 9688 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842292777182433/5547842292777182433_signature_28.png | Mary claire p baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842292777182433/5547842292777182433_signature_31.png
| 2023-03-16 | 2023-03-14 14:37 | Lung center of the Philippines | Quezon ave, Quezon city | NCR | NCR 2nd - Quezon City | Koica laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23855 | PBSP | jennelynlouis@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: , COUNT: 1460 | 2023-03-14 | GX4 | 704584 | For module replacement 652610 - 650425,laptop defective unable to boot, ups defective unable to backup | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational unit network in another gx system |  | 9000491 rs | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842192777019162/5547842192777019162_signature_28.png | Jennelyn louis | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842192777019162/5547842192777019162_signature_31.png
| 2023-03-16 | 2023-03-15 12:18 | Tondo foreshore health center | Pacheco st.,Tondo ,Manila | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23854 | MACARE-HBDC | carina_moncada@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2113 ASSAY: Others, COUNT: 190 | 2023-03-14 | GX4 (10colors) | 110001650 | For module replacement A4 2100 10365-210048255 | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | Operational |  | 900- 0650rs | 9691 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842192773751744/5547842192773751744_signature_28.png | CARINA S. MONCADA | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842192773751744/5547842192773751744_signature_31.png
| 2023-03-16 | 2023-03-14 14:18 | Lung center of the Philippines | Quezon ave.Quezon CIty | NCR | NCR 2nd - Quezon City | Koica laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23854 | PBSP | jennelynlouis@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1460 | 2023-03-14 | GX4 | 707450 | For module replacement 688291 - 644613, 674546 - 645554 def ups unable to backup | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 9000491rs | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842192779207460/5547842192779207460_signature_28.png | Jennelyn louis | https://www.jotform.com/uploads/macarewecare/222927978372067/5547842192779207460/5547842192779207460_signature_31.png
| 2023-03-16 | 2023-03-16 18:02 | Davao Regional Medical Center | Apokon Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23853 | PBSP | vincentmilestone20@gmail.com | 23.6 |  | 2023-03-16 | GX-IV 4D | 834626 | Module B2 and B4 are for replacement | Replace the Service Module B2 / SN: 682931 with new Module SN: 890841 Replace the defective Module B4 / SN: 655108 with new Module SN: 892119 | Perform Module Replacement | 3 modules are operational Module B1 Encouter Multiple Error, Datas will be sent to Cepheid for further Assessment |  | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547796076419273648/5547796076419273648_signature_28.png | DONGIE REY A. ARAPOC | https://www.jotform.com/uploads/macarewecare/222927978372067/5547796076419273648/5547796076419273648_signature_31.png
| 2023-03-16 | 2023-03-16 18:02 | Davao Regional Medical Center | Apokon Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23852 | PBSP | vincentmilestone20@gmail.com | 23.6 |  | 2023-03-16 | GX-IV 4D | 834626 | Module B2 and B4 are for replacement | Replace the Service Module B2 / SN: 682931 with new Module SN: 890841 Replace the defective Module B4 / SN: 655108 with new Module SN: 892119 | Perform Module Replacement | 3 modules are operational Module B1 Encouter Multiple Error, Datas will be sent to Cepheid for further Assessment |  | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547720596413174928/5547720596413174928_signature_28.png | DONGIE REY A. ARAPOC | https://www.jotform.com/uploads/macarewecare/222927978372067/5547720596413174928/5547720596413174928_signature_31.png
| 2023-03-16 | 2023-03-16 18:09 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23851 | PBSP | vincentmilestone20@gmail.com | 23.5 |  | 2023-03-16 | GX-IV 4d | 834440 | Module D3 for replacement | Replace the Service Module D3/SN 869918 with new Module SN: 895742 | Perform Module Replacement | Unit is fully operational | Module D2 encounter error, error will be sent to Cepheid for further Assessment | 900-0491rs | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547720106414784311/5547720106414784311_signature_28.png | DONGIE REY A. ARAPOC | https://www.jotform.com/uploads/macarewecare/222927978372067/5547720106414784311/5547720106414784311_signature_31.png
| 2023-03-16 | 2023-03-16 12:09 | Sison Rural Health Unit | Pobalcion Central Sison Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23850 | PBSP | sisonrhu@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF, COUNT: 103 ASSAY: Xpert MTB/RIF Ultra, COUNT: 263 | 2023-03-09 | GXIV-D | 834684 | Due for Calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except B2:784440 failed xpertcheck, for reporting | B2 for replacement | 5 Xpertcheck kits | 39406 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547641585313205687/5547641585313205687_signature_28.png | Jeannette O. Alcaide | https://www.jotform.com/uploads/macarewecare/222927978372067/5547641585313205687/5547641585313205687_signature_31.png
| 2023-03-16 | 2023-03-16 15:35 | Sorsogon Provincial Hospital | Sorsogon Diversion Road, Sorsogon | 5 | Sorsogon | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  |  |  |  | 23849 | PBSP | wizard682000@yahoo.com | 24 |  | 2023-03-08 | GXIV | 835214 | Module B2 for replacement | Replaced module B2 (888907 > 633491) | Replaced module B2, performed maintenance cleaning | Unit is operational |  | Genexpert module 6c | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547617616014917619/5547617616014917619_signature_28.png | Reynaldo B. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5547617616014917619/5547617616014917619_signature_31.png
| 2023-03-16 | 2023-03-16 14:43 | GOVERNOR CELESTINO GALLARES MEMORIAL HOSPITAL | TAGBILARAN BOHOL | 7 | Bohol | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23848 | DOH | majieocmen@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 985 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1341 ASSAY: Xpert HIV-1 Viral Load, COUNT: 42 |  | GX-IV | 820228 | PREVENTIVE MAINTENANCE | COLLECT IQR, SYSTEM LOG, ARCHIVE | PERFORM CLEANING OF GX EQUIPMENT AND PREVENTIVE MAINTENANCE | EQUIPMENT IS OPERATIONAL | MODULE A4 DISABLED DUE TO ERRORS |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547593243779647527/5547593243779647527_signature_28.png | MA. JERNAI SOSOBAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5547593243779647527/5547593243779647527_signature_31.png
| 2023-03-16 | 2023-03-16 14:47 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | OPD OB | YES | MAINTENANCE | MEDGYN | Roi Kenneth Fagutao |  |  |  |  | 23846 | COMMERCIAL | benedict.alconera@drmc.doh.gov.ph | 23 |  |  | MGC - 200 Cryotherapy | 11263 | For Preventive Maintenance | Clean Surface Area, visual inspection of the machine | Perform Preventive Maintenance | Unit is not operational due to missing parts - Nitrogen Tank |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547589541717183967/5547589541717183967_signature_28.png | Benedict john m. Alconera | https://www.jotform.com/uploads/macarewecare/222927978372067/5547589541717183967/5547589541717183967_signature_31.png
| 2023-03-16 | 2023-03-16 14:04 | GOVERNOR CELESTINO GALLARES MEMORIAL HOSPITAL | TAGBILARAN BOHOL | 7 | Bohol | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23845 | PBSP | majieocmen@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 506 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1393 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 |  | GX-IV | 838878 | FOR MODULE B4 REPLACEMENT | #NAME? | -PERFORM MODULE B4 DEF. S/N:863568 REPLACEMENT WITH NEW S/N:618997 -PERFORM PREV. MAINTENANCE | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547562582775931898/5547562582775931898_signature_28.png | Ma. Jernai Sosoban | https://www.jotform.com/uploads/macarewecare/222927978372067/5547562582775931898/5547562582775931898_signature_31.png
| 2023-03-16 | 2023-03-16 13:06 | Bulusan RHU | Central Bulusan | 5 | Sorsogon | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  |  |  |  | 23844 | PBSP | arleneragasa@gmail.com | 26 |  | 2023-03-08 | GXIV | 835197 | Module B1 for replacement | Replaced module B1 (788957 > 671724) | Replaced module B1, performed maintenance cleaning | Unit is operational |  | Genexpert module 6C | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547529349683665462/5547529349683665462_signature_28.png | Maria Arlene G. Ragasa | https://www.jotform.com/uploads/macarewecare/222927978372067/5547529349683665462/5547529349683665462_signature_31.png
| 2023-03-16 | 2023-03-16 11:57 | Samar Provincial Hospital | Catbalogan Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Paul Labutap |  |  |  |  | 23843 | PBSP | leslieleones36@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1171 ASSAY: Xpert MTB/RIF Ultra, COUNT: 870 | 2023-03-16 | GXIV | 808291 | Defective Monitor | #NAME? | #NAME? | #NAME? | Defective Monitor is for replacement (SN : 509NTKFGR189) |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547506831222962721/5547506831222962721_signature_28.png | melanie joy lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5547506831222962721/5547506831222962721_signature_31.png
| 2023-03-16 | 2023-03-16 12:09 | Sison Rural Health Unit | Pobalcion Central Sison Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23842 | PBSP | sisonrhu@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF, COUNT: 103 ASSAY: Xpert MTB/RIF Ultra, COUNT: 263 | 2023-03-09 | GXIV-D | 834684 | Due for Calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except B2:784440 failed xpertcheck, for reporting | B2 for replacement | 5 Xpertcheck kits | 39406 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547502907812725690/5547502907812725690_signature_28.png | Jeannette O. Alcaide | https://www.jotform.com/uploads/macarewecare/222927978372067/5547502907812725690/5547502907812725690_signature_31.png
| 2023-03-16 | 2023-03-16 11:24 | Agusan del Norte Provincial Health Office | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 23841 | PBSP | ronniebalansag74@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF, COUNT: 936 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1083 |  | GX-4 | 807461 | None | None | Replacing Module A2 with SN: 775673 to SN: 663320 | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547469540302166092/5547469540302166092_signature_28.png | Ronnie M. Balansag | https://www.jotform.com/uploads/macarewecare/222927978372067/5547469540302166092/5547469540302166092_signature_31.png
| 2023-03-16 | 2023-03-16 10:23 | Mlang Rural Health Unit | M'lang, North Cotabato | 12 | Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23840 | PBSP | kathlynlacabe7@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 674 | 2023-03-16 | Gx-IV | 815434 | Module 3 H/W failed | Perform Genexpert machine Preventive maintenance Replace module 3 SN: 686389 with new module SN: 640319 | Perform Genexpert machine Preventive maintenance Replace module 3 SN: 686389 with new module SN: 640319 | 4 operational modules |  | Genexpert module 6-color | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547459134525937144/5547459134525937144_signature_28.png | Kathlyn mae y. Lacabe | https://www.jotform.com/uploads/macarewecare/222927978372067/5547459134525937144/5547459134525937144_signature_31.png
| 2023-03-16 | 2023-03-16 11:04 | Bulan Rural Health Unit | C. Gotladera St., Bulan | 5 | Sorsogon | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  |  |  |  | 23839 | DOH | bing.gonzales33@gmail.com | 28 |  | 2023-03-08 | GXIV | 807983 | Module B4 for replacement | Replaced module B4 (664775 > 690369) | Replaced module B4, performed maintenance cleaning | Unit is operational |  | Genexpert module 6C | PBSP |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547453501617586790/5547453501617586790_signature_28.png | Claribel G. Nacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5547453501617586790/5547453501617586790_signature_31.png
| 2023-03-16 | 2023-03-16 10:27 | Guiguinto Rural Health Unit II | Guiguinto, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23838 | PBSP | joie_calica@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 160 ASSAY: Xpert MTB/RIF Ultra, COUNT: 678 ASSAY: Xpert HIV-1 Viral Load, COUNT: 188 |  | GXIV | 814165 | Temperaturr errors on modules B1, B2, B3 | #NAME? | #NAME? | Unit is operational. B2 hardware failed. B1, B3 temp. errors. For follow up/reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547432880227190515/5547432880227190515_signature_28.png | Razhen Matias | https://www.jotform.com/uploads/macarewecare/222927978372067/5547432880227190515/5547432880227190515_signature_31.png
| 2023-03-16 | 2023-03-16 10:11 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23837 | PBSP | cyrilcordova0515@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4110 |  | GX IV | 804267 | Defective module 606617 | Clean fan filter | New module installed 702067 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547422775322343239/5547422775322343239_signature_28.png | Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5547422775322343239/5547422775322343239_signature_31.png
| 2023-03-16 | 2023-03-16 09:01 | LA CARLOTA CITY HEALTH OFFICE | La Carlota City, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | REPAIR | CEPHEID | Gerald Fritz Selga |  |  |  |  | 23836 | PBSP | tinbellotinbello@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 402 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1388 |  | GX-IV | 807594 | Defective module 639820/A3 | Replaced defective module | - Replaced defective module 639820/A3 with new module SN 677745 - Performed Preventive Maintenance (module by cleaning, plunger rod maintenace, filter cleaning, optic slots cleaning) - Generate reports for defective module 690194/A1. | 3 modules operational, module A1 for report to Jotform. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547398642398985111/5547398642398985111_signature_28.png | Valerie Christine G. Bello | https://www.jotform.com/uploads/macarewecare/222927978372067/5547398642398985111/5547398642398985111_signature_31.png
| 2023-03-16 | 2023-03-16 08:53 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB-Dots | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23835 | DOH | materesaaliman78@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 278 ASSAY: Xpert MTB/RIF Ultra, COUNT: 430 | 2023-03-16 | GXIV | 838406 | n/a | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547379682391521865/5547379682391521865_signature_28.png | Ma.Teresa Aliman | https://www.jotform.com/uploads/macarewecare/222927978372067/5547379682391521865/5547379682391521865_signature_31.png
| 2023-03-16 | 2023-03-15 09:02 | Tinambacan Main Health Center | Tinambacan Sur, Calbayog Samar | 8 | Samar | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23833 | PBSP | porlaressteph@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 184 ASSAY: Xpert MTB/RIF Ultra, COUNT: 224 | 2023-03-15 | GXIV-2D | 834824 | Unit for Calibration and Preventive Maintenance | #NAME? | #NAME? | Unit is Operational | Module B1 (Retry and Passed), Module B2 Passed | Xpertcheck Cartridge | XPERTCHECK - CE - 5 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547055663028110466/5547055663028110466_signature_28.png | stephanie jis rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/5547055663028110466/5547055663028110466_signature_31.png
| 2023-03-16 | 2023-03-15 15:17 | Calbayog City Health Office | Talahib, Calbayog Samar | 8 | Samar | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23834 | PBSP | mardelen_l@yahoo.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 658 ASSAY: Xpert MTB/RIF Ultra, COUNT: 582 | 2023-03-15 | GXIV | 804265 | for Preventive maintenance | #NAME? | #NAME? | Module B4 disabled (multiple temperature error) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5547055663028211405/5547055663028211405_signature_28.png | Mardelen C. Langomes | https://www.jotform.com/uploads/macarewecare/222927978372067/5547055663028211405/5547055663028211405_signature_31.png
| 2023-03-15 | 2023-03-15 15:53 | Bayambang District Hospital | Bayambang | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23832 | PBSP | bevzneri@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 739 |  | GXIV | 812830 | All modules for replacement | - replaced all modules (New SN: 791850, 667602, 861757, 667606 - Generated reports - checked & cleaned cartridge bay, plunger rod, I core slot, filter - Performed functionality test | Replaced modules B1: 802139 B2: 685779 B3: 690916 B4: 879481 | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546979842227081229/5546979842227081229_signature_28.png | Beverly neri | https://www.jotform.com/uploads/macarewecare/222927978372067/5546979842227081229/5546979842227081229_signature_31.png
| 2023-03-15 | 2023-03-14 14:18 | Lung center of the Philippines | Quezon ave.Quezon CIty | NCR | NCR 2nd - Quezon City | Koica laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23831 | PBSP | jennelynlouis@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1460 | 2023-03-14 | GX4 | 707450 | For module replacement 688291 - 644613, 674546 - 645554 def ups unable to backup | Performed module replacement and cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 9000491rs | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546963682774195918/5546963682774195918_signature_28.png | Jennelyn louis | https://www.jotform.com/uploads/macarewecare/222927978372067/5546963682774195918/5546963682774195918_signature_31.png
| 2023-03-15 | 2023-03-14 14:37 | Lung center of the Philippines | Quezon ave, Quezon city | NCR | NCR 2nd - Quezon City | Koica laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23830 | PBSP | jennelynlouis@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: , COUNT: 1460 | 2023-03-14 | GX4 | 704584 | For module replacement 652610 - 650425,laptop defective unable to boot, ups defective unable to backup | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational unit network in another gx system |  | 9000491 rs | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546962682773595619/5546962682773595619_signature_28.png | Jennelyn louis | https://www.jotform.com/uploads/macarewecare/222927978372067/5546962682773595619/5546962682773595619_signature_31.png
| 2023-03-15 | 2023-03-15 12:18 | Tondo foreshore health center | Pacheco st.,Tondo ,Manila | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23829 | MACARE-HBDC | carina_moncada@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2113 ASSAY: Others, COUNT: 190 | 2023-03-14 | GX4 (10colors) | 110001650 | For module replacement A4 2100 10365-210048255 | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | Operational |  | 900- 0650rs | 9691 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546960482777599512/5546960482777599512_signature_28.png | CARINA S. MONCADA | https://www.jotform.com/uploads/macarewecare/222927978372067/5546960482777599512/5546960482777599512_signature_31.png
| 2023-03-15 | 2023-03-15 14:24 | E.Rodriguez health center | Pajo st., Quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | Noel Floresca |  |  |  |  | 23828 | DOH | clairebaptista32786@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 164 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1154 | 2023-03-13 | GX4 | 831848 | UPS no power | Replaced ups, test run unit,test back up supply of ups | UPS replacement | UPS operational,modules 1,2,3 disabled failed xpertcheck |  | Powercom ups | 9688 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546959192779373515/5546959192779373515_signature_28.png | Mary claire p baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/5546959192779373515/5546959192779373515_signature_31.png
| 2023-03-15 | 2023-03-15 16:38 | Rodriguez RHU | J.P Rizal Rodriguez rizal | 4A | Rizal | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23827 | PBSP | karlbanal25@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2436 ASSAY: Xpert HIV-1 Viral Load, COUNT: 89 |  | GX4 | 838894 | Workstation no display booting problem | Check workstation,reset ram, reset workstation,check cable connection | Troubleshooting | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546957742776471369/5546957742776471369_signature_28.png | Karl Michael Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5546957742776471369/5546957742776471369_signature_31.png
| 2023-03-15 | 2023-03-15 20:53 | Sagay City Health Office | Sagay, Poblacion | 7 | Negros Occidental | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 23826 | PBSP | maravelesdebra@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 321 | 2023-03-15 | Gx4 | 838144 | Hardware failure due to Error 1002/1001/2014 | Install new Module Serial Number 868334 | Replaced defective Module 806310 | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546944341611771549/5546944341611771549_signature_28.png | Debra M. Abellar | https://www.jotform.com/uploads/macarewecare/222927978372067/5546944341611771549/5546944341611771549_signature_31.png
| 2023-03-15 | 2023-03-15 15:46 | Bacolod City Health Office | Bacolod City | 6 | Negros Occidental | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 23825 | COMMERCIAL | cyrilcordova0515@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4110 |  | GXIV | 804838 | Defective module C4 645428 | Clean fan filter | New module installed 699776 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546760711625691714/5546760711625691714_signature_28.png | Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5546760711625691714/5546760711625691714_signature_31.png
| 2023-03-15 | 2023-03-15 13:27 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 23824 | DOH | zionclydematote@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 899 ASSAY: Xpert MTB/RIF, COUNT: 855 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2785 ASSAY: Xpert HIV-1 Viral Load, COUNT: 235 | 2023-03-15 | Genexpert IV | 831718 | Error 2035, 2014, 2097 | Performed Xpertcheck on module A1, A2 and A3 | Performed Maintenance Gathered Reports Performed Xpertcheck (A1, A2, A3) Enabled A1 and A2 Performed Functionality Test | Unit is Operational | Xpertcheck Result A1, A2, A3 = Passed | Xpertcheck 3pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546752253225535675/5546752253225535675_signature_28.png | Sunshine Matote | https://www.jotform.com/uploads/macarewecare/222927978372067/5546752253225535675/5546752253225535675_signature_31.png
| 2023-03-15 | 2023-03-15 11:51 | Alcala Municipal Health Office | Centro Norte, Alcala, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 23823 | DOH | rhualcala@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 188 ASSAY: Xpert MTB/RIF Ultra, COUNT: 489 | 2023-03-15 | Genexpert IV | 838531 | Module 2 Hardware Failure | Performed Xpertcheck on B2 | Performed Maintenance Gathered Reports Performed Xpertcheck Performed Functionality Test | Unit is Operational | Xpertcheck - Service Required | Xpertchech 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546752253225709376/5546752253225709376_signature_28.png | Clarisa Pascual | https://www.jotform.com/uploads/macarewecare/222927978372067/5546752253225709376/5546752253225709376_signature_31.png
| 2023-03-15 | 2023-03-15 09:18 | Bulalacao Municipal Health Center | Poblacion, Bulalacao | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23822 | PBSP | queence.macaraig@gmail.com | 25 |  | 2023-03-15 | GX-IV-D-2 | 835185 | For XpertCheck | Performed maintenance | Performed maintenance. Run XpertCheck Generated reports Installed new UPS | All modules operational. Module B1 passed, Module B2 service required. |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716211798921/5546734716211798921_signature_28.png | Queencelyn Festin | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716211798921/5546734716211798921_signature_31.png
| 2023-03-15 | 2023-03-15 14:02 | Bulalacao Municipal Health Office | Aplaya 2, Bongabong | 4B | Oriental Mindoro | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23821 | PBSP | cortezmariel30@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 30 ASSAY: Xpert MTB/RIF Ultra, COUNT: 246 | 2023-03-13 | GX-IV-D | 834687 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Modules A2, A3 and A4 operational. Modules A1/786829 and A4/786837 service required. |  | XpertCheck | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716219931839/5546734716219931839_signature_28.png | Mariel A. Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716219931839/5546734716219931839_signature_31.png
| 2023-03-15 | 2023-03-15 11:36 | Oriental Mindoro Southern District Hospital | Western Nautical Highway, Roxas | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23820 | PBSP | markedisonmartinez@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 254 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1759 |  | GX-IV-D | 807590 | Modules B1, B2 and B4 for replacement. | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716218305031/5546734716218305031_signature_28.png | Mark Edison G. MARTINEZ, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5546734716218305031/5546734716218305031_signature_31.png
| 2023-03-15 | 2023-03-15 13:50 | Estevez Memorial Hospital Inc | Don Juan Estevez Sr. St. | 5 | Albay | Laboratory | YES | INSTALLATION | CEPHEID | Patrick Los Baños |  |  |  |  | 23819 | COMMERCIAL | estevezmemorialhospitalinc@gmail.com | 28 |  | 2023-03-07 | GXIV | 110017942 | None, for unit installation | Installed unit and accessories, performed end-user training, generated reports | Performed unit installation | Unit is operational |  | Cepheid GeneXpert GXIV 10C S/N: 110017942 | DR#9670 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546708348477274249/5546708348477274249_signature_28.png | Estevez Memorial Hospital Inc. | https://www.jotform.com/uploads/macarewecare/222927978372067/5546708348477274249/5546708348477274249_signature_31.png
| 2023-03-15 | 2023-03-15 13:45 | Malalag Rural Health Unit | Poblacion Malalag, Davao del Sur | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23818 | PBSP | krstadimacuta@gmail.com | 28.9 |  |  | GX-IV 4D | 835220 | For Preventive Maintenance and Machine always encounter module disconnection | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Reset GeneXpert connections | Perform Preventive Maintenance | Unit is operational but the UPS of the GeneXpert system is faulty. | When UPS is at battery mode, less than 3 minute mark it will automatically drained and the GeneXpert with its Computer shutoff abruptly |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546707331875119621/5546707331875119621_signature_28.png | Krista Mae Dimacuta | https://www.jotform.com/uploads/macarewecare/222927978372067/5546707331875119621/5546707331875119621_signature_31.png
| 2023-03-15 | 2023-03-15 13:36 | Bongabong Municipal Health Office | Aplaya Bongabong | 4B | Oriental Mindoro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23817 | PBSP | cortezmariel30@gmail.com | 28 |  | 2023-03-13 | GxIV | 834811 | Unit for preventive maintenance. For installation of gateway board. | Preventive maintenance and installation of gateway board. | Performed preventive maintenance. Performed installation of gateway board. | Unit is operational. |  | Gatewayboard | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546699515992799427/5546699515992799427_signature_28.png | Mariel A. Cortez | https://www.jotform.com/uploads/macarewecare/222927978372067/5546699515992799427/5546699515992799427_signature_31.png
| 2023-03-15 | 2023-03-15 13:46 | LORENZO D. ZAYCO DISTRICT HOSPITAL | National Highway, Kabankalan, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 23816 | PBSP | beuahravante@mail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 526 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1404 |  | GX-IV | 807455 | HW Fail on modules 630292/A1 and 609383/A2 | Replaced defective modules, Preventive Maintenance (module bay cleaning, filter cleaning, plunger rod maintenace, optic slots cleaning) | - Replaced module 630292/A1 with new module SN 682751. - Replaced module 609383/A2 with new module SN 671781 | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546692475227334376/5546692475227334376_signature_28.png | BEULAH AGRAVANTE, R.M.T. | https://www.jotform.com/uploads/macarewecare/222927978372067/5546692475227334376/5546692475227334376_signature_31.png
| 2023-03-15 | 2023-03-15 13:47 | Region 1 Medical Center | Arellano St. Dagupan City | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23815 | PBSP | romarmtamayo@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 418 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1080 ASSAY: Xpert SARS-CoV-2, COUNT: 4158 | 2023-03-08 | GXXVI | 842230 | C2,C4, B3 and B4 For module replacement | Preventive maintenance Module Replacement | Performed preventive maintenance Performed module Replacement Gathered reports Performed functionality test | Unit is operational |  | 4 Genexpert module | 39370 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546690685931091302/5546690685931091302_signature_28.png | Romar Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5546690685931091302/5546690685931091302_signature_31.png
| 2023-03-15 | 2023-03-15 12:49 | Mayor Hilarion A. Ramiro Sr. | Ozamiz City | 10 | Misamis Occidental | Molecular Laboratory | YES | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23814 | PBSP | mharsmc.covidlabgenex@gmail.com | 29 | ASSAY: Xpert SARS-CoV-2, COUNT: 5627 |  | GeneXpert | 834437 | None | None | Preventive Maintenance Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546683793208397060/5546683793208397060_signature_28.png | Gretchen Lanas | https://www.jotform.com/uploads/macarewecare/222927978372067/5546683793208397060/5546683793208397060_signature_31.png
| 2023-03-15 | 2023-03-15 12:49 | Mayor Hilarion A. Ramiro Sr. | Ozamiz City | 10 | Misamis Occidental | Molecular Laboratory | YES | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23813 | PBSP | mharsmc.covidlabgenex@gmail.com | 29 | ASSAY: Xpert SARS-CoV-2, COUNT: 5627 |  | GeneXpert | 834437 | None | None | Preventive Maintenance Database back up | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546683623207719669/5546683623207719669_signature_28.png | Gretchen Lanas | https://www.jotform.com/uploads/macarewecare/222927978372067/5546683623207719669/5546683623207719669_signature_31.png
| 2023-03-15 | 2023-03-15 13:13 | TUBIGON RURAL HEALTH UNIT | TUBIGON BOHOL | 7 | Bohol | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  |  |  |  | 23812 | PBSP | josephraytan@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 1055 ASSAY: Xpert MTB/RIF Ultra, COUNT: 564 |  | GX-IV | 815431 | MODULE B4 ERROR 2005 | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL, MODULE B4 IS DISABLED DUE TO MULTIPLE ERRORS |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546667919712192454/5546667919712192454_signature_28.png | Joseph Ray F. Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5546667919712192454/5546667919712192454_signature_31.png
| 2023-03-15 | 2023-03-15 12:13 | Democrito O. Plaza Memorial Hospital | National Highway, Brgy. Patin-ay, Prosperidad | 13 | Agusan del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 23811 | PBSP | lermabpabillan@yahoo.com | 20.1 | ASSAY: Xpert MTB/RIF, COUNT: 487 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2806 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-4 | 805335 | None | None | Replacing module C1 with SN: 210068711 to SN: 885698 | All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546654132541587378/5546654132541587378_signature_28.png | Lerma B. Pabillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5546654132541587378/5546654132541587378_signature_31.png
| 2023-03-15 | 2023-03-15 12:18 | Tondo foreshore health center | Pacheco st.,Tondo ,Manila | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23810 | MACARE-HBDC | carina_moncada@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2113 ASSAY: Others, COUNT: 190 | 2023-03-14 | GX4 (10colors) | 110001650 | For module replacement A4 2100 10365-210048255 | Performed module replacement cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement | Operational |  | 900- 0650rs | 9691 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546635401997785262/5546635401997785262_signature_28.png | CARINA S. MONCADA | https://www.jotform.com/uploads/macarewecare/222927978372067/5546635401997785262/5546635401997785262_signature_31.png
| 2023-03-15 | 2023-03-14 09:43 | San Lazaro Hospital | Quiricada St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE ReXpertCheck | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23809 | PBSP | acong.pbsp@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 119 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GXIV-D | 804766 | Unit is for ReXpertCheck | XpertCheck | ReXpertCheck Modules 2 and 4 Performed Maintenance Generate Reports | 1 module failed and 1 passed in ReXpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602011640245/5546630602011640245_signature_28.png | Ashley Cynna Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602011640245/5546630602011640245_signature_31.png
| 2023-03-15 | 2023-03-14 09:36 | San Lazaro Hospital | Quiricada St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23806 | PBSP | acong.pbsp@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 119 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GXIV-D | 707843 | Modules Failed in XpertCheck | Module Replacement | Replaced Modules 1 and 4 Performed Maintenance Performed Self Test Generate reports | 3 modules are operational, 1 for reporting. CPU reported. | 700528-> 662707, 602779->699744 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602012925381/5546630602012925381_signature_28.png | Ashley Cynna Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602012925381/5546630602012925381_signature_31.png
| 2023-03-15 | 2023-03-15 11:49 | Tanza Rural Health Unit | Municipal Compound, Tanza | 4A | Cavite | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23808 | PBSP | gutierrezrhian@gmail.com | 25 |  |  | GXIV-D | 812385 | CPU has Booting Problem | Troubleshoot | Performed Cepheid Troubleshooting Checked HDD Connections Swapped RAM place | CPU is defective and for Replacement | CPU S/N: 2UA70329CD |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602012966912/5546630602012966912_signature_28.png | Mary Ann C. Gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602012966912/5546630602012966912_signature_31.png
| 2023-03-15 | 2023-03-15 10:31 | Noveleta Rural Health Unit | Municipal Compound, Noveleta | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23807 | PBSP | jacque_estinoso23@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 380 |  | GXIV-D | 835186 | Module 1 with Temperature Error | Module Replacement | Module Replacement in B1 Performed Maintenance Generate Reports | All Modules are Operational | 788976->761468 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602017381797/5546630602017381797_signature_28.png | Jacqueline Estinoso | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602017381797/5546630602017381797_signature_31.png
| 2023-03-15 | 2023-03-14 14:34 | Tunasan Health Center | E.Rodriguez St., Tunasan | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23805 | PBSP | gigietan0911@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1294 ASSAY: Xpert HIV-1 Viral Load, COUNT: 130 |  | GXIV-D | 807613 | Module 3 Failed in XpertCheck | Module Replacememt | Module Replacement in Module 3 Performed Maintenance Generate Reports | All module are Operational | 662483->678800 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602019223935/5546630602019223935_signature_28.png | Geraldine Trazo-Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5546630602019223935/5546630602019223935_signature_31.png
| 2023-03-15 | 2023-03-15 11:39 | Mayor Hilarion A. Ramiro Sr. Medical Center | Ozamiz City | 10 | Misamis Occidental | Microbiology Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23804 | PBSP | felixvicentcampos@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 892 ASSAY: Xpert MTB/RIF, COUNT: 34 |  | GeneXpert | 804765 | None | None | Preventive Maintenance Xpertcheck calibration Database back up | Operational | Service required module B4/618492 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546629743205335548/5546629743205335548_signature_28.png | Felix Vicent C. Campos | https://www.jotform.com/uploads/macarewecare/222927978372067/5546629743205335548/5546629743205335548_signature_31.png
| 2023-03-15 | 2023-03-15 11:39 | Mayor Hilarion A. Ramiro Sr. Medical Center | Ozamiz City | 10 | Misamis Occidental | Microbiology Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23803 | PBSP | felixvicentcampos@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 892 ASSAY: Xpert MTB/RIF, COUNT: 34 |  | GeneXpert | 804765 | None | None | Preventive Maintenance Xpertcheck calibration Database back up | Operational | Service required module B4/618492 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546629653208959848/5546629653208959848_signature_28.png | Felix Vicent C. Campos | https://www.jotform.com/uploads/macarewecare/222927978372067/5546629653208959848/5546629653208959848_signature_31.png
| 2023-03-15 | 2023-03-14 11:05 | Sindangan Rural Health Unit | Sindangan | 9 | Zamboanga del Norte | TB DOTS | YES | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23802 | PBSP | lryano0618@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1466 |  | GeneXpert | 807596 | None | Module replacement new modules B2/750227 B3/757019 | Module replacement of defective modules B2/671327 B3/642704 Preventive Maintenance Generate reports data files | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546607293203128095/5546607293203128095_signature_28.png | Leni Riza O. Yano | https://www.jotform.com/uploads/macarewecare/222927978372067/5546607293203128095/5546607293203128095_signature_31.png
| 2023-03-15 | 2023-03-15 10:14 | Cagayan Valley Center for Health Development, Department of Health Regional Office 2 | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23801 | PBSP | kate.christinejk@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 420 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1688 ASSAY: Xpert SARS-CoV-2, COUNT: 4875 | 2023-03-15 | Genexpert XVI | 842894 | Module D1 and D4 Hardware Failure | Module D1 and D4 to be reported | Performed Maintenance Gathered Reports Performed Module Replacement (B2) Performed Functionality Test | Unit is Operational | Module Replacement = 889068 - 626812 | Genexpert Module - 626812 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546570875958412910/5546570875958412910_signature_28.png | Kate Christine D. Calimag | https://www.jotform.com/uploads/macarewecare/222927978372067/5546570875958412910/5546570875958412910_signature_31.png
| 2023-03-15 | 2023-03-15 09:56 | Cagayan Valley Center for Health Development, Department of Health Regional Office 2 | Carig Norte, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23800 | PBSP | kate.christinejk@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 153 ASSAY: Xpert MTB/RIF Ultra, COUNT: 667 ASSAY: Xpert SARS-CoV-2, COUNT: 1891 | 2023-03-15 | Genexpert IV | 834434 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement (B2) Performed Functionality Test | Unit is Operational | Module Replacement = 776897 - 630557 | Genexpert Module - 630557 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546569285956855619/5546569285956855619_signature_28.png | Kate Christine D. Calimag | https://www.jotform.com/uploads/macarewecare/222927978372067/5546569285956855619/5546569285956855619_signature_31.png
| 2023-03-15 | 2023-03-15 09:54 | Parang Rural Health Unit | Parang, Maguindanao | BARMM | Maguindanao | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23799 | PBSP | Nadjimbagua20@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 83 ASSAY: Xpert MTB/RIF Ultra, COUNT: 442 | 2023-03-15 | Gx-IV | 838135 | Module 2 for service required | Perform Genexpert machine Calibration/Xpertcheck Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Calibration/Xpertcheck Perform Genexpert machine Preventive maintenance | 3 operational modules |  | XpertCheck -CE-5 | Lot 1000378669 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546561205377736637/5546561205377736637_signature_28.png | Nadjim Bagua | https://www.jotform.com/uploads/macarewecare/222927978372067/5546561205377736637/5546561205377736637_signature_31.png
| 2023-03-15 | 2023-03-15 09:19 | GARCIA MEMORIAL PROVINCIAL HOSPITAL | TALIBON BOHOL | 7 | Bohol | TB DOTS | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Joseph Acebedo |  |  |  |  | 23798 | MACARE-HBDC | ybanezrosalina19@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 288 ASSAY: Xpert MTB/RIF Ultra, COUNT: 495 |  | GX-IV | 838203 | #NAME? | #NAME? | #NAME? | MODULE B2 INTERMITTENT ERROR, MODULE B3 DISABLED / H/W FAILED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546528567762203658/5546528567762203658_signature_28.png | Rosalina Ybanez | https://www.jotform.com/uploads/macarewecare/222927978372067/5546528567762203658/5546528567762203658_signature_31.png
| 2023-03-15 | 2023-03-14 12:15 | Doña Remedios Trinidad Rural Health Unit | DRT, Bulacan | NCR | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23797 | PBSP | maricel.rmt24@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 90 ASSAY: Xpert MTB/RIF Ultra, COUNT: 178 |  | GXIV | 838872 | Unit is not operational. | #NAME? | Checked monitor | Unit is operational. B1, B3, B4 passed, B2 failed calibration. Temporarily installed back up monitor. Monitor is displaying white screen. Original monitor is for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546491982223548062/5546491982223548062_signature_28.png | Cecilia Angeles | https://www.jotform.com/uploads/macarewecare/222927978372067/5546491982223548062/5546491982223548062_signature_31.png
| 2023-03-15 | 2023-03-14 15:39 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23797 | PBSP | bmclab_micro@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 586 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1289 ASSAY: Xpert HCV VL, COUNT: 11 ASSAY: Xpert HIV-1 Viral Load, COUNT: 219 |  | GXIV | 805112 | Desktop is not booting up | #NAME? | Checked desktop | Unit is operational. Desktop is under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546491982225884245/5546491982225884245_signature_28.png | Angelica Guintu | https://www.jotform.com/uploads/macarewecare/222927978372067/5546491982225884245/5546491982225884245_signature_31.png
| 2023-03-15 | 2023-03-14 10:35 | St. Lukes medical center | E.rodriguez quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23796 | COMMERCIAL | biomed.qc@stlukes.com.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 236 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2928 |  | GX4 | 807155 | C4 module encounter error 5007 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Troubleshooting and maintenance | C2 h/w failed | Unit originally frm c.i transfer to microbiology for temporary use |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546448521999800544/5546448521999800544_signature_28.png | Phillip Jordan M. Salvador, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5546448521999800544/5546448521999800544_signature_31.png
| 2023-03-15 | 2023-03-14 10:35 | St. Lukes medical center | E.rodriguez quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23795 | COMMERCIAL | biomed.qc@stlukes.com.ph | 28 | ASSAY: Xpert CT/NG, COUNT: 236 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2928 |  | GX4 | 807155 | C4 module encounter error 5007 | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Troubleshooting and maintenance | C2 h/w failed | Unit originally frm c.i transfer to microbiology for temporary use |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546447821999113844/5546447821999113844_signature_28.png | Phillip Jordan M. Salvador, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5546447821999113844/5546447821999113844_signature_31.png
| 2023-03-15 | 2023-03-14 09:59 | St.lukes medical center | E.rodriguez ,Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | YES | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23794 | COMMERCIAL | kfuy@macare-medicals.com | 26 | ASSAY: Xpert CT/NG, COUNT: 236 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2982 | 2023-03-14 | GX4 | 839234 | B1 encounter temp error 4012,B2 encounter 5007 error ,B4 encounter 5007 error | Performed xpertcheck in b1,b2,b4,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Troubleshooting | All module failed in xpertcheck | For module replacement | Xpertcheck kits | 0 |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5546446671996089603/5546446671996089603_signature_28.png | Phillip Jordan M.Salvador RMT | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5546446671996089603/5546446671996089603_signature_31.png
| 2023-03-15 | 2023-03-14 09:59 | St.lukes medical center | E.rodriguez ,Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | YES | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23793 | COMMERCIAL | biomed.qc@stlukes.com.ph | 26 | ASSAY: Xpert CT/NG, COUNT: 236 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2982 | 2023-03-14 | GX4 | 839234 | B1 encounter temp error 4012,B2 encounter 5007 error ,B4 encounter 5007 error | Performed xpertcheck in b1,b2,b4,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Troubleshooting | All module failed in xpertcheck | For module replacement | Xpertcheck kits | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5546446331991265168/5546446331991265168_signature_28.png | Phillip Jordan M.Salvador RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5546446331991265168/5546446331991265168_signature_31.png
| 2023-03-14 | 2023-03-14 15:01 | Cotabato City Health Office | Cotabato City | BARMM | Maguindanao | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23792 | DOH | onehobe834@gmail.com | 24.2 | ASSAY: Xpert MTB/RIF, COUNT: 505 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1157 |  | Gx-IV | 838528 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545923369742096863/5545923369742096863_signature_28.png | Emelyn Cadiz | https://www.jotform.com/uploads/macarewecare/222927978372067/5545923369742096863/5545923369742096863_signature_31.png
| 2023-03-14 | 2023-03-14 15:05 | Olongapo City Health Office | Arthur st. West Bajac-Bajac Olongapo City | 3 | Zambales | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23791 | PBSP | roselynmaganomilan@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 773 ASSAY: Xpert MTB/RIF Ultra, COUNT: 939 | 2023-03-10 | GXIV-D | 804700 | D1:613358 error | Calibration Preventive maintenance | Performed surface cleaning Cleaning of icore, cartridge bay, plunger rod and filter Xpertcheck application update to 1.5a Archived test results Performed calibration Gathered reports Performed functionality test | Unit is operational except D1 for reporting | Calibration info and failed module reports did not generte | 2 xpertcheck kits | 39470 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545874231118978559/5545874231118978559_signature_28.png | Roselyn Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5545874231118978559/5545874231118978559_signature_31.png
| 2023-03-14 | 2023-03-14 14:59 | Olongapo City Health Office | Arthur st. West Bajac-Bajac Olongapo City | 3 | Zambales | Labortory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23790 | PBSP | roselynmaganomilan@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF, COUNT: 702 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1112 | 2023-02-08 | GXIV-D | 834807 | C2:777166 for module replacement | Module replacement Preventive maintenance | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational except C1: for reporting | Temperature error | Genexpert module | 39366 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545874221115152553/5545874221115152553_signature_28.png | Roselyn Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5545874221115152553/5545874221115152553_signature_31.png
| 2023-03-14 | 2023-03-14 14:52 | Olongapo City Health Office | Ruzal ave. Olongapo City Zambales | 3 | Zambales | Labortory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23789 | PBSP | roselynmaganomilan@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 773 ASSAY: Xpert MTB/RIF Ultra, COUNT: 939 | 2023-02-08 | GXIV-D | 804700 | D4:685616 for moduoe replacement | Module replacement Preventive maintenance | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational except B1 |  | Genexpert module | 39366 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545862651118106300/5545862651118106300_signature_28.png | Rosely Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5545862651118106300/5545862651118106300_signature_31.png
| 2023-03-14 | 2023-03-14 14:08 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Molecular laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23788 | PBSP | tindocvelche@gmail.com | 21.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 6304 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | Gx-IV | 834441 | Module 4 H/W failed due | Perform Genexpert machine Preventive maintenance Report module problem | Perform Genexpert machine Preventive maintenance Report module Problem | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545843683104593991/5545843683104593991_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5545843683104593991/5545843683104593991_signature_31.png
| 2023-03-14 | 2023-03-14 14:11 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Molecular laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23787 | PBSP | tindocvelche@gmail.com | 22.9 | ASSAY: Xpert SARS-CoV-2, COUNT: 6304 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | Gx-IV | 834442 | Module 1 and 2 H/W failed | Perform Genexpert machine Preventive maintenance Report module problem | Perform Genexpert machine Preventive maintenance Report module problem | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545843103102767183/5545843103102767183_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5545843103102767183/5545843103102767183_signature_31.png
| 2023-03-14 | 2023-03-14 14:03 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Molecular lab | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23786 | DOH | tindocvelche@gmail.com | 19.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 6304 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | Gx-IV | 837662 | Module 1 and 3 for service required | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration for module 1 and 3 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration for module 1 and 3 for troubleshooting | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545842213106149246/5545842213106149246_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5545842213106149246/5545842213106149246_signature_31.png
| 2023-03-14 | 2023-03-14 13:36 | Valentina G. Plaza Memorial Hospital | Purok 2, San Teodoro, Bunawan | 13 | Agusan del Sur | RTDL | NO | REPAIR TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 23785 | PBSP | bdhclinlab@gmail.com | 23.28 | ASSAY: Xpert MTB/RIF, COUNT: 975 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3715 |  | GX-4 | 815436 | Module A2 was found E2032 | Done reporting module A2 | Troubleshooting and documentation of Module A2 Replacing module A3 with SN: 804866 to SN: 870956 Replacing module A4 with SN: 210016522 to SN: 870963 | 3 modules working and 1 module for replacement | None |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545822093697754474/5545822093697754474_signature_28.png | Rosephine Toyong | https://www.jotform.com/uploads/macarewecare/222927978372067/5545822093697754474/5545822093697754474_signature_31.png
| 2023-03-14 | 2023-03-14 13:09 | Zamboanga City Health Office | Rizal Street, Zamboanga City, Zamboanga del sur | 9 | Zamboanga del Sur | Labiratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23784 | PBSP | teresitaflores99@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 484 | 2023-03-14 | GX 4 | 838128 | None | Module Replacement, Preventive maintenence | Module Replacement, Preventive maintenence | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545801416271058918/5545801416271058918_signature_28.png | Teresita J. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5545801416271058918/5545801416271058918_signature_31.png
| 2023-03-14 | 2023-03-14 11:57 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23783 | DOH | dyrosemarie71@gmail.com | 27.3 |  |  | GX-IV 4D | 838401 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses | Perform Preventive Maintenance | Unit is fully operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545759298423346244/5545759298423346244_signature_28.png | Rosemarie Dy RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5545759298423346244/5545759298423346244_signature_31.png
| 2023-03-14 | 2023-03-14 11:29 | Morong Rural Health Unit | Juan Luna Morong Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23782 | DOH | lielyn.lim@yahoo.com | 27.7 | ASSAY: Xpert MTB/RIF, COUNT: 129 ASSAY: Xpert MTB/RIF Ultra, COUNT: 245 | 2023-02-07 | GXIV-D | 831607 | Due for Calibrtion | Calibration Preventive maintenance | -Performed surface cleaning -Performed cleaning of cartridge bay,icore,plunger rod and filter -Xpercheck application update to 1.5a -Performed calibration -Generate reports -Archived test results -Performed functionality test | Unit is operational except A1 for replacement | A1:754382 failed xpertcheck, for reporting | 4 Xpertcheck kit | 9596 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545746996583151321/5545746996583151321_signature_28.png | Lielyn Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5545746996583151321/5545746996583151321_signature_31.png
| 2023-03-14 | 2023-03-14 11:05 | Sindangan Rural Health Unit | Sindangan | 9 | Zamboanga del Norte | TB DOTS | YES | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23781 | PBSP | lryano0618@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1466 |  | GeneXpert | 807596 | None | Module replacement new modules B2/750227 B3/757019 | Module replacement of defective modules B2/671327 B3/642704 Preventive Maintenance Generate reports data files | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545738314466472651/5545738314466472651_signature_28.png | Leni Riza O. Yano | https://www.jotform.com/uploads/macarewecare/222927978372067/5545738314466472651/5545738314466472651_signature_31.png
| 2023-03-14 | 2023-03-14 10:57 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | TB-Lab | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23780 | PBSP | tindocvelche@gmail.com | 19.3 | ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 67 |  | Gx-IV | 807463 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545726603101656267/5545726603101656267_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5545726603101656267/5545726603101656267_signature_31.png
| 2023-03-14 | 2023-03-14 10:57 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | TB-Lab | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23779 | PBSP | tindocvelche@gmail.com | 19.3 | ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 67 |  | Gx-IV | 807463 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545726533102042266/5545726533102042266_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5545726533102042266/5545726533102042266_signature_31.png
| 2023-03-14 | 2023-03-14 10:51 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | TB-Lab | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23778 | PBSP | tindocvelche@gmail.com | 19.8 | ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert MTB/RIF, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 16 ASSAY: Xpert HIV-1 Viral Load, COUNT: 67 | 2023-03-14 | Gx-IV | 807607 | N/A | Perform Genexpert machine Preventive maintenance Replace Genexpert CPU SN: *CPH0520F08599* with new CPU SN:*CPH0322F32916* | Perform Genexpert machine Preventive maintenance Replace Genexpert CPU SN: *CPH0520F08599* with new CPU SN:*CPH0322F32916* | 4 operational modules |  | Lenovo ThinkCentre | SN: *CPH0322F32916* |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545725493104506628/5545725493104506628_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5545725493104506628/5545725493104506628_signature_31.png
| 2023-03-14 | 2023-03-14 10:38 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | TB-Lab | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23777 | PBSP | tindocvelche@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1803 ASSAY: Xpert MTB/RIF, COUNT: 745 ASSAY: Xpert HPV, COUNT: 31 ASSAY: Xpert HBV VL, COUNT: 30 ASSAY: Xpert HIV-1 Viral Load, COUNT: 129 | 2023-03-14 | Gx-XVI | 839307 | N/A | Perform Genexpert machine Calibration/Xpertcheck Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Calibration/Xpertcheck Perform Genexpert machine Preventive maintenance | 16 operational modules | All modules passed the annual calibration | Xpertcheck -CE-5 | Lot 1000378669 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545724593109344517/5545724593109344517_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5545724593109344517/5545724593109344517_signature_31.png
| 2023-03-14 | 2023-03-14 10:43 | Cebu TB Reference Lab | Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Raymond Laurito |  |  |  |  | 23776 | PBSP | cebutbreferencelab@gmail.com | 25 |  | 2023-03-14 | GXIV | 838882 | Defective LCD monitor | New LCD monitor installation | New LCD monitor installation | Systems operational | LCD monitor: ThinkVision | ThinkVision S22e-20 Flat panel monitor | Sn V5FRK606 MTM (Model) 62C6-KAT1US |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545718524611135755/5545718524611135755_signature_28.png | JOE FER ARES | https://www.jotform.com/uploads/macarewecare/222927978372067/5545718524611135755/5545718524611135755_signature_31.png
| 2023-03-14 | 2023-03-14 08:50 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | CRMC-TB lab | YES | INSTALLATION | CEPHEID | Jesson Tupas |  |  |  |  | 23775 | PBSP | efmd@crmc.doh.gov.ph | 26 |  |  | Gx-IV (10C) | 110016958 | N/A | Install Genexpert machine 10C Install Genexpert CPU Install Genexpert Monitor install Genexpert UPS Install Genexpert wireless mouse W/ mousepad and keyboard Install MTB-XDR assay | Genexpert machine Installation | 4 operational modules | Printer not installed due to enduser choice |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545653723101687343/5545653723101687343_signature_28.png | Velche B. Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5545653723101687343/5545653723101687343_signature_31.png
| 2023-03-13 | 2023-03-13 14:49 | Concepcion District Hospital | L. Cortez St. Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23774 | PBSP | cdh.laboratorydept@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 224 ASSAY: Xpert MTB/RIF Ultra, COUNT: 862 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 | 2023-03-08 | GXIV-D | 807562 | B4 module replacement B2:4004, Valve motion B1: temperature error | Preventive maintenance Troubleshooting Module replacement | Performed preventive maintenance Check-up moduoe B1 and B2 Performed module replacement Gathered reports Performed functionality test | Unit is operational except B1 and B2, bustedR220 and C142 for module B2 | B1 and B2 for reporting | Genexpert module | 39359 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545180185313502931/5545180185313502931_signature_28.png | Maria Rowena Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5545180185313502931/5545180185313502931_signature_31.png
| 2023-03-13 | 2023-03-13 13:11 | Ospital ning Capas | Sto. Domingo 2nd Municipal Compound Capas, Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23773 | DOH | onclaboratory@gmail.com | 31 | ASSAY: Xpert MTB/RIF, COUNT: 259 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1278 | 2023-02-01 | GXIV-D | 831591 | Due for calibration UPS not charging | Calibration Preventive maintenance Troubleshooting | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test Troubleshoot UPS | Unit is operational except A1,A2,A4, Defective UPS not charging | A1,A2,A4 failed xpertcheck for reporting. UPS battery for replacement | 7 xpertcheck kits | 9564 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545180185317270574/5545180185317270574_signature_28.png | Afrilyn M. Roque | https://www.jotform.com/uploads/macarewecare/222927978372067/5545180185317270574/5545180185317270574_signature_31.png
| 2023-03-13 | 2023-03-13 17:14 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Tb Lab | YES | INSTALLATION | CEPHEID | Jesson Tupas |  |  |  |  | 23772 | PBSP | tindocvelche@gmail.com | 26 |  |  | Gx-IV (10C) | 110016958 | N/A | Genexpert machine installation | Genexpert machine installation | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545119513101367413/5545119513101367413_signature_28.png | Velche Tindoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5545119513101367413/5545119513101367413_signature_31.png
| 2023-03-13 | 2023-03-10 12:56 | Nasugbu Rural Health Unit | Nasugbu | 4A | Batangas | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23771 | PBSP | lucenavicente15@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 803 |  | GXIV-D | 807581 | B1 Lost Communication Error | Reliability Test Reset Gateway Board | Performed Maintenance Run Reliability Test in Module 1 Checked Gateway-Module Cable Reset Gateway Board Checked Ultrasonic Cable if intact Generate Reports | All modules are operational. | Monitor display is violet because of CPU VGA output |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545075691214060636/5545075691214060636_signature_28.png | Lucena I. Vicente | https://www.jotform.com/uploads/macarewecare/222927978372067/5545075691214060636/5545075691214060636_signature_31.png
| 2023-03-13 | 2023-03-13 15:49 | Cabanatuan City Health Office | Kapt. Pepe subd. Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE Xpertcheck | CEPHEID | Matthew Roxas |  |  |  |  | 23770 | DOH | cholab2021@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 860 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1855 | 2023-03-13 | Genexpert IV | 831592 | Module 1 Temperature Error | Performed Xpertcheck on Module 1 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck result on A1 - Passed | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545038737413626016/5545038737413626016_signature_28.png | Lea samson | https://www.jotform.com/uploads/macarewecare/222927978372067/5545038737413626016/5545038737413626016_signature_31.png
| 2023-03-13 | 2023-03-13 13:44 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | Replace service module | CEPHEID | Wynnona Reoma |  |  |  |  | 23769 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 22.9 | ASSAY: Xpert HPV, COUNT: 181 ASSAY: Xpert MTB/RIF, COUNT: 1037 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1094 |  | GXIV-4D | 820206 | None | None | -Replace service module installed in slot A3 SN: 651877 with another service module SN: 708880 | All four modules are functional. | Waiting for module A3 allocation. Pull-out module with SN 651877 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5545003923418302906/5545003923418302906_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5545003923418302906/5545003923418302906_signature_31.png
| 2023-03-13 | 2023-03-10 13:30 | Angono Rural Health Unit | Angono | 4A | Rizal | Laboratory | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23768 | DOH | angonomholab@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 63 ASSAY: Xpert MTB/RIF Ultra, COUNT: 762 |  | GX-IV-D | 819864 | Module A4/719514 has error 2014, 1001, 1002 | Run XpertCheck on module A4 | Performed maintenance Run XpertCheck on module A4 Generated reports Run reliability test | Modules A1 and A3 operational. Modules A2 and A4 not operational. Module A4 passed the XpertCheck but under observation. | Used 1 cartridge of XpertCheck. Module A2 reported already. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996541212171731/5544996541212171731_signature_28.png | Rose Sienna T.Blanco | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996541212171731/5544996541212171731_signature_31.png
| 2023-03-13 | 2023-03-09 14:13 | Pagsanjan Rural Health Unit | Pagsanjan | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23767 | PBSP | iamleiarnie@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 357 ASSAY: Xpert HIV-1 Viral Load, COUNT: 172 |  | GX-IV-D | 807568 | CPU not operational. | Reseat internal connections | Performed maintenance Reseat internal connections Generated reports | Unit is not operational. | UPS has no backup. CPU not operational. Recommend to replace CPU and UPS. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996471217852438/5544996471217852438_signature_28.png | Lei Arnie Abrigo | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996471217852438/5544996471217852438_signature_31.png
| 2023-03-13 | 2023-03-07 10:39 | Moonwalk Health Center | Brgy. Moonwalk | NCR | NCR 4th - Parañaque | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23763 | PBSP | almaaurea@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 210 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1018 |  | GX-IV-D | 805113 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996461218038208/5544996461218038208_signature_28.png | Alma Aurea Limson | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996461218038208/5544996461218038208_signature_31.png
| 2023-03-13 | 2023-03-08 11:14 | Jose Reyes Memorial Medical Center | Santa Cruz | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23764 | MACARE-HBDC | covidlab@jrrmmc.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 979 ASSAY: Xpert SARS-CoV-2, COUNT: 3558 | 2023-03-07 | GX-XVI-D-10C | 110013856 | Modules D1 to D4 for replacement. | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. |  | Old: D1/210069230, D2/210069253,D3/210069288, D4/210069289 New: D1/21002843, D2/210010680, D3/210017264, D4/210037681 | DR# 9672 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996461219547637/5544996461219547637_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996461219547637/5544996461219547637_signature_31.png
| 2023-03-13 | 2023-03-07 13:13 | San Dionisio Health Center | San Dionisio | NCR | NCR 4th - Parañaque | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23765 | DOH | marubi1019@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 388 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1069 |  | GX-IV-D | 831898 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore, Checked filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996471216296138/5544996471216296138_signature_28.png | Fatima Joyce Tresvalles | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996471216296138/5544996471216296138_signature_31.png
| 2023-03-13 | 2023-03-09 13:16 | Sta. Cruz Rural Health Unit | Sta. Cruz | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23766 | DOH | jezzaleal@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 939 |  | GX-IV-D | 820229 | Module D1 and D4 for re-xpertcheck | Run XpertCheck on module D1 and D4 | Performed maintenance Run XpertCheck on module D1 and D4 Generated reports | Modules D2 and D3 operational. Module D1 and D4 not operational. | Used 2 XpertCheck cartridges. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996471217493653/5544996471217493653_signature_28.png | Jezza Nequinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5544996471217493653/5544996471217493653_signature_31.png
| 2023-03-13 | 2023-03-13 13:32 | Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23762 | MACARE-HBDC | rhwclab@gmail.com | 24.7 |  |  | GX-IV 4D | 839236 | For Preventive Maintenance and Modules Encounter Errors | Clean Surface Area, Cartridge Bay, Plunger Rod and fan filters Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Modules Encounter Error 2097, reported to JotForm for further assessment of the Product Application Specialist |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544988458429236417/5544988458429236417_signature_28.png | Reproductive Health and Wellness Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5544988458429236417/5544988458429236417_signature_31.png
| 2023-03-13 | 2023-03-13 14:30 | Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23761 | DOH | rhwclab@gmail.com | 24.9 |  |  | GX-IV 4d | 838262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod and fan filters Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Modules Encounter Error 2097, files needed sent to Jotform and Product Application Specialist Team for further assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544987188423840687/5544987188423840687_signature_28.png | Reproductive Health and Wellness Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5544987188423840687/5544987188423840687_signature_31.png
| 2023-03-13 | 2023-03-13 14:30 | Reproductive Health and Wellness Center | Jacinto St, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23760 | DOH | rhwclab@gmail.com | 24.9 |  |  | GX-IV 4d | 838262 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod and fan filters Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Modules Encounter Error 2097, files needed sent to Jotform and Product Application Specialist Team for further assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544987068427652686/5544987068427652686_signature_28.png | Reproductive Health and Wellness Center | https://www.jotform.com/uploads/macarewecare/222927978372067/5544987068427652686/5544987068427652686_signature_31.png
| 2023-03-13 | 2023-03-07 15:13 | National kidney and transplant institute | East ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23759 | MACARE-HBDC | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 184 ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Others, COUNT: 26 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert MTB/RIF, COUNT: 401 ASSAY: Xpert HCV VL, COUNT: 33 |  | Gxxvi | 804476 | For maintenance, disabled a1,a2,a3,b1,b3,b4,c1,c2,c3,c4 | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and checkup | A4 and b2 in use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544805291211845564/5544805291211845564_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5544805291211845564/5544805291211845564_signature_31.png
| 2023-03-13 | 2023-03-07 15:13 | National kidney and transplant institute | East ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23758 | MACARE-HBDC | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 184 ASSAY: Xpert CT/NG, COUNT: 5 ASSAY: Others, COUNT: 26 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert MTB/RIF, COUNT: 401 ASSAY: Xpert HCV VL, COUNT: 33 |  | Gxxvi | 804476 | For maintenance, disabled a1,a2,a3,b1,b3,b4,c1,c2,c3,c4 | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and checkup | A4 and b2 in use |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544804971215813910/5544804971215813910_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5544804971215813910/5544804971215813910_signature_31.png
| 2023-03-13 | 2023-03-07 15:58 | National kidney and transplant institute | East ave qc | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23757 | MACARE-HBDC | molecularlabnkti@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 219 ASSAY: Others, COUNT: 14 ASSAY: Xpert HBV VL, COUNT: 29 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1186 ASSAY: Xpert SARS-CoV-2, COUNT: 18206 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert SARS-CoV-2, COUNT: 9 |  | GXXVI | 110005587 | A2 D1,D4 encounter temperature error | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | A2,D1,D4 disable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544804501215977832/5544804501215977832_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5544804501215977832/5544804501215977832_signature_31.png
| 2023-03-13 | 2023-03-07 15:58 | National kidney and transplant institute | East ave qc | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23756 | MACARE-HBDC | molecularlabnkti@gmail.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 219 ASSAY: Others, COUNT: 14 ASSAY: Xpert HBV VL, COUNT: 29 ASSAY: Xpert HPV, COUNT: 111 ASSAY: Xpert MTB/RIF, COUNT: 1186 ASSAY: Xpert SARS-CoV-2, COUNT: 18206 ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert SARS-CoV-2, COUNT: 9 |  | GXXVI | 110005587 | A2 D1,D4 encounter temperature error | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | A2,D1,D4 disable |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544804341218005663/5544804341218005663_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5544804341218005663/5544804341218005663_signature_31.png
| 2023-03-13 | 2023-03-08 09:03 | Casimiro ynares sr.memorial hospital | Becco rd. Burgos Rodriguez rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23755 | DOH | mariannegervacio1@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 64 ASSAY: Xpert MTB/RIF Ultra, COUNT: 349 | 2023-02-20 | Gx4 | 831642 | For maintenance and xpertcheck | Performed xpertcheck, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and xpertcheck | Operational | All module passed | Xpertcheck kits | 9639 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544803261218756222/5544803261218756222_signature_28.png | Marianne Gervacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5544803261218756222/5544803261218756222_signature_31.png
| 2023-03-13 | 2023-03-08 09:03 | Casimiro ynares sr.memorial hospital | Becco rd. Burgos Rodriguez rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23754 | DOH | mariannegervacio1@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 64 ASSAY: Xpert MTB/RIF Ultra, COUNT: 349 | 2023-02-20 | Gx4 | 831642 | For maintenance and xpertcheck | Performed xpertcheck, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and xpertcheck | Operational | All module passed | Xpertcheck kits | 9639 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544802591217356133/5544802591217356133_signature_28.png | Marianne Gervacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5544802591217356133/5544802591217356133_signature_31.png
| 2023-03-13 | 2023-03-08 13:37 | Esperanza health center | Teresa st sta mesa manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23753 | DOH | jong.agdamag@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF, COUNT: 276 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1025 | 2023-01-19 | Gx4 | 831891 | For maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and xpertcheck | A1,A2 failed xpertcheck, A3,A4 passed xpertcheck |  | Xpertcheck kits | 9512 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544801431216641295/5544801431216641295_signature_28.png | Victor Agdamag jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5544801431216641295/5544801431216641295_signature_31.png
| 2023-03-13 | 2023-03-08 13:37 | Esperanza health center | Teresa st sta mesa manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23752 | DOH | jong.agdamag@gmail.com | 30 | ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF, COUNT: 276 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1025 | 2023-01-19 | Gx4 | 831891 | For maintenance and xpertcheck | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and xpertcheck | A1,A2 failed xpertcheck, A3,A4 passed xpertcheck |  | Xpertcheck kits | 9512 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544800891219368039/5544800891219368039_signature_28.png | Victor Agdamag jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5544800891219368039/5544800891219368039_signature_31.png
| 2023-03-13 | 2023-03-08 14:47 | Domingo santiago health center | D.santiago st sampaloc,manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23751 | DOH | fei_paas@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 249 ASSAY: Xpert MTB/RIF Ultra, COUNT: 941 | 2023-01-19 | Gx4 | 831527 | For maintenance and xpertcheck A3 and A4 reported already | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and xpertcheck | A3 and A4 reported already A4 passed | A1 failed xpertcheck | Xpertcheck kits | 9511 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544800281215559617/5544800281215559617_signature_28.png | Flory fe Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5544800281215559617/5544800281215559617_signature_31.png
| 2023-03-13 | 2023-03-08 14:47 | Domingo santiago health center | D.santiago st sampaloc,manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23750 | DOH | fei_paas@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 249 ASSAY: Xpert MTB/RIF Ultra, COUNT: 941 | 2023-01-19 | Gx4 | 831527 | For maintenance and xpertcheck A3 and A4 reported already | cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and xpertcheck | A3 and A4 reported already A4 passed | A1 failed xpertcheck | Xpertcheck kits | 9511 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544800001214841532/5544800001214841532_signature_28.png | Flory fe Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5544800001214841532/5544800001214841532_signature_31.png
| 2023-03-13 | 2023-03-07 14:29 | National kidney and transplant institute | East ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23749 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 880 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Others, COUNT: 51 ASSAY: Xpert HBV VL, COUNT: 442 ASSAY: Xpert HPV, COUNT: 19 ASSAY: Xpert MTB/RIF, COUNT: 540 ASSAY: Xpert SARS-CoV-2, COUNT: 450 ASSAY: Xpert HCV VL, COUNT: 74 |  | Gx 4 | 838524 | B3 continous 1002,,for maintenance,dirty filter | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and check up | B3 disable, unable b1, | B1,b2,b4 ready to use |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544785341212083210/5544785341212083210_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5544785341212083210/5544785341212083210_signature_31.png
| 2023-03-13 | 2023-03-07 14:29 | National kidney and transplant institute | East ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23748 | MACARE-HBDC | molecularlabnkti@gmail.com | 26 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 880 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Others, COUNT: 51 ASSAY: Xpert HBV VL, COUNT: 442 ASSAY: Xpert HPV, COUNT: 19 ASSAY: Xpert MTB/RIF, COUNT: 540 ASSAY: Xpert SARS-CoV-2, COUNT: 450 ASSAY: Xpert HCV VL, COUNT: 74 |  | Gx 4 | 838524 | B3 continous 1002,,for maintenance,dirty filter | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and check up | B3 disable, unable b1, | B1,b2,b4 ready to use |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544785001213209933/5544785001213209933_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5544785001213209933/5544785001213209933_signature_31.png
| 2023-03-13 | 2023-03-09 14:47 | National children's hospital | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23747 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2197 ASSAY: Xpert SARS-CoV-2, COUNT: 6462 ASSAY: Xpert SARS-CoV-2, COUNT: 18 | 2023-03-07 | GX XVI | 110014171 | For module replacement b1 210069274 - 210002752 | Performed module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900-0650rs 10 colors modules | 9671 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544783951211793525/5544783951211793525_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5544783951211793525/5544783951211793525_signature_31.png
| 2023-03-13 | 2023-03-09 14:47 | National children's hospital | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23746 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2197 ASSAY: Xpert SARS-CoV-2, COUNT: 6462 ASSAY: Xpert SARS-CoV-2, COUNT: 18 | 2023-03-07 | GX XVI | 110014171 | For module replacement b1 210069274 - 210002752 | Performed module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900-0650rs 10 colors modules | 9671 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544783561217908360/5544783561217908360_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5544783561217908360/5544783561217908360_signature_31.png
| 2023-03-13 | 2023-03-09 14:47 | National children's hospital | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23745 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2197 ASSAY: Xpert SARS-CoV-2, COUNT: 6462 ASSAY: Xpert SARS-CoV-2, COUNT: 18 | 2023-03-07 | GX XVI | 110014171 | For module replacement b1 210069274 - 210002752 | Performed module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900-0650rs 10 colors modules | 9671 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544783011213405285/5544783011213405285_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5544783011213405285/5544783011213405285_signature_31.png
| 2023-03-13 | 2023-03-10 09:47 | Batasan hills super health center | Batasan qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23744 | PBSP | ashwinkevin@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 710 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1837 | 2023-03-08 | Gx4 | 805254 | For module replacement and maintenance D1 672771 - 622105 | Performed module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900 0491 rs | G5243 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544782561214087456/5544782561214087456_signature_28.png | Kevin Orlando Cabuso | https://www.jotform.com/uploads/macarewecare/222927978372067/5544782561214087456/5544782561214087456_signature_31.png
| 2023-03-13 | 2023-03-10 09:47 | Batasan hills super health center | Batasan qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23743 | PBSP | ashwinkevin@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 710 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1837 | 2023-03-08 | Gx4 | 805254 | For module replacement and maintenance D1 672771 - 622105 | Performed module replacement, cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900 0491 rs | G5243 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544782251219897645/5544782251219897645_signature_28.png | Kevin Orlando Cabuso | https://www.jotform.com/uploads/macarewecare/222927978372067/5544782251219897645/5544782251219897645_signature_31.png
| 2023-03-13 | 2023-03-10 11:35 | Wenceslao health center | Culiat qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23742 | DOH | ameliafclemente2001@yahoo.com | 27 | ASSAY: Xpert HPV, COUNT: 199 ASSAY: Xpert MTB/RIF, COUNT: 199 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1782 | 2023-01-20 | Gx4 | 831876 | For maintenance and xpertcheck | Perf9rmed xpertcheck,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and xpertcheck | All module passed |  | Xpertcheck | 9532 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544765191211174386/5544765191211174386_signature_28.png | Amelia F. Clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5544765191211174386/5544765191211174386_signature_31.png
| 2023-03-13 | 2023-03-10 11:35 | Wenceslao health center | Culiat qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23741 | DOH | ameliafclemente2001@yahoo.com | 27 | ASSAY: Xpert HPV, COUNT: 199 ASSAY: Xpert MTB/RIF, COUNT: 199 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1782 | 2023-01-20 | Gx4 | 831876 | For maintenance and xpertcheck | Perf9rmed xpertcheck,cleaned unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance and xpertcheck | All module passed |  | Xpertcheck | 9532 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544764901218063911/5544764901218063911_signature_28.png | Amelia F. Clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5544764901218063911/5544764901218063911_signature_31.png
| 2023-03-13 | 2023-03-07 11:01 | East avenue medical center | East ave qc | NCR | NCR 2nd - Quezon City | Tb dots | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23740 | PBSP | eamcdots@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 154 ASSAY: Xpert MTB/RIF, COUNT: 940 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3071 | 2023-03-07 | Gx4 | 803087 | For module replacement c1 611387 - 688470 c4 661735 - 661016 | Performed module replacement cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Opearational |  | 900 gx | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544762941219368202/5544762941219368202_signature_28.png | rose ann parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5544762941219368202/5544762941219368202_signature_31.png
| 2023-03-13 | 2023-03-07 11:01 | East avenue medical center | East ave qc | NCR | NCR 2nd - Quezon City | Tb dots | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23739 | PBSP | eamcdots@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 154 ASSAY: Xpert MTB/RIF, COUNT: 940 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3071 | 2023-03-07 | Gx4 | 803087 | For module replacement c1 611387 - 688470 c4 661735 - 661016 | Performed module replacement cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Opearational |  | 900 gx | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544762541213560727/5544762541213560727_signature_28.png | rose ann parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5544762541213560727/5544762541213560727_signature_31.png
| 2023-03-13 | 2023-03-10 15:16 | Tondo medical center | Tondo manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23738 | DOH | microscan@tmc.doh.gov.ph | 23 | ASSAY: Xpert HPV, COUNT: 315 ASSAY: Xpert MTB/RIF, COUNT: 194 ASSAY: Xpert MTB/RIF Ultra, COUNT: 479 ASSAY: Xpert SARS-CoV-2, COUNT: 4176 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 |  | Gx4 | 831872 | Cpu unable to boot | Check ram,reset ram cleaned terminal and transfer another slot location, check hdd cable and reset | Troubleshooting | For observation |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5544751261215020770/5544751261215020770_signature_28.png | Analyn N. Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5544751261215020770/5544751261215020770_signature_31.png
| 2023-03-11 | 2023-03-11 10:13 | TLY Anglo | 2F/#715 3F Anglo Bldg Shaw Blvd Wak wack Greenhills, Mandaluyong City | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | Unit relocation | CEPHEID | Harold Toco |  |  |  |  | 23737 | MACARE-HBDC | gestoya@loveyourself.ph | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 223 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1174 | 2023-03-11 | Gx-4 | 831895 | None, Unit relocation. | Unit relocation and setup. | Performed unit relocation. | Unit is operational. |  | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5543111341315101124/5543111341315101124_signature_28.png | Gem anthony Estoya | https://www.jotform.com/uploads/macarewecare/222927978372067/5543111341315101124/5543111341315101124_signature_31.png
| 2023-03-10 | 2023-03-10 17:21 | St Lukes Extension Clinic | J Bocobo St, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Patrick Los Baños |  |  |  |  | 23736 | MACARE-HBDC | laboratory@slec.ph | 22 |  | 2023-03-03 | Infinity 48s | 809643 | Module A4 and H4 for replacement | Replaced modules A4 and H4 (690624 > 210050828, 684147 > 672198) | Replaced defective modules , performed preventive maintenance cleaning | Unit is operational |  | Genexpert Module 6-color | DR# 9661 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542595723203176810/5542595723203176810_signature_28.png | Maria Teresita Baclig | https://www.jotform.com/uploads/macarewecare/222927978372067/5542595723203176810/5542595723203176810_signature_31.png
| 2023-03-10 | 2023-03-10 15:13 | Porac Rural Health Unit | Porac Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23735 | DOH | rhulabporac@gmail.com | 29.65 | ASSAY: Xpert MTB/RIF, COUNT: 263 ASSAY: Xpert MTB/RIF Ultra, COUNT: 535 | 2023-03-03 | GXIV-D | 831911 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except C3 and C4 |  | 2 xpertcheck kits | 9584 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542564565316396781/5542564565316396781_signature_28.png | Marife Lansangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5542564565316396781/5542564565316396781_signature_31.png
| 2023-03-10 | 2023-03-10 15:58 | Compostela Rural Health Unit | Poblacion, Compostela Davao De Oro | 11 | Davao de Oro | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23734 | PBSP | comp.clin.lab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF, COUNT: 130 ASSAY: Xpert MTB/RIF Ultra, COUNT: 935 |  | GX-IV 4d | 807360 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Reseat, printer cartridge | Perform Preventive Maintenance | Module 3 encounter temp error, error will be sent to Cepheid for further Assessment Printer always having paper jam, even after reseating cartridge and checking for blockages |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542468866915460607/5542468866915460607_signature_28.png | MARY ANN P MONTEDERAMOS, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5542468866915460607/5542468866915460607_signature_31.png
| 2023-03-10 | 2023-03-09 14:51 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23731 | PBSP | abigailgsula@gmail.com | 26.3 | ASSAY: Xpert MTB/RIF, COUNT: 207 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1060 | 2023-02-03 | GXIV-D | 831912 | Due for Calibration | Calibration Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except A1 and A2 | 2 modules passed calibration | 2 xpertcheck kits | 9582 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542320914216009264/5542320914216009264_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5542320914216009264/5542320914216009264_signature_31.png
| 2023-03-10 | 2023-03-09 12:03 | Minalin Rural Health Unit | Minalin Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23730 | PBSP | bernadetteann0318@gmail.com | 30.5 | ASSAY: Xpert MTB/RIF, COUNT: 94 ASSAY: Xpert MTB/RIF Ultra, COUNT: 547 |  | GXIV-D | 834688 | Defective UPS( Sudden shutdown) | Troubleshooting | Test UPS Battery voltage Performed functionality test | Unit is operational, Deffective UPS for reporting | UPS for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542320914219270222/5542320914219270222_signature_28.png | Bernadette Ann Lazaro | https://www.jotform.com/uploads/macarewecare/222927978372067/5542320914219270222/5542320914219270222_signature_31.png
| 2023-03-10 | 2023-03-10 12:22 | Mabalacat Rural Health Unit | Mabalacat Pampanga | 3 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23732 | PBSP | perezriyadh@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 204 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1024 ASSAY: Xpert HIV-1 Viral Load, COUNT: 26 | 2023-03-08 | GXIV-D | 807564 | A1 and A4 moduoe replacement | Module replacement Preventive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 2 genexpert modules | 39363 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542320924211003676/5542320924211003676_signature_28.png | Riyadh Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5542320924211003676/5542320924211003676_signature_31.png
| 2023-03-10 | 2023-03-10 12:29 | Mabalacat Rural Health Unit | Mabalacat Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23733 | PBSP | perezriyadh@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF, COUNT: 81 ASSAY: Xpert MTB/RIF Ultra, COUNT: 870 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 | 2023-03-08 | GXIV-2 | 835104 | B2 For replacement | Module replacement Prevenyive maintenance | Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | 1 genexpert module | 39363 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542320924214613733/5542320924214613733_signature_28.png | Riyadh Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5542320924214613733/5542320924214613733_signature_31.png
| 2023-03-10 | 2023-03-10 11:39 | Bongabon District Hospital | Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23729 | DOH | teody_santua@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 136 ASSAY: Xpert MTB/RIF Ultra, COUNT: 227 | 2023-02-03 | Genexpert IV | 831562 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck - A1, A3 - Passed | Xpertcheck 3pc | 9588 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542309857412390373/5542309857412390373_signature_28.png | Teodorico S. Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5542309857412390373/5542309857412390373_signature_31.png
| 2023-03-10 | 2023-03-10 11:39 | Bongabon District Hospital | Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23728 | DOH | teody_santua@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 136 ASSAY: Xpert MTB/RIF Ultra, COUNT: 227 | 2023-02-03 | Genexpert IV | 831561 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck - A1, A3 - Passed | Xpertcheck 3pc | 9588 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542309447413092533/5542309447413092533_signature_28.png | Teodorico S. Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5542309447413092533/5542309447413092533_signature_31.png
| 2023-03-10 | 2023-03-10 11:54 | Dinas Rural Health Unit | Don jose, Dinas, Zamboanga Del sur | 9 | Zamboanga del Sur | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23727 | PBSP | genosograce92@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2023-03-10 | Gx 4 | 838533 | System unit malfunction | Pulled out the said System unit for troubleshooting. | Used different ram slots, replaced CMOS battery reconnect all wires inside the System unit. | Under repair/ Troubleshooting | System unit pulled out | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542301908615390568/5542301908615390568_signature_28.png | Grace Geñoso | https://www.jotform.com/uploads/macarewecare/222927978372067/5542301908615390568/5542301908615390568_signature_31.png
| 2023-03-10 | 2023-03-10 11:37 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 23726 | DOH | ccmclaboratory01@gmail.com | 25.1 | ASSAY: Xpert MTB/RIF, COUNT: 595 ASSAY: Xpert MTB/RIF Ultra, COUNT: 974 |  | GXIV-4D | 838561 | Module B1 - H/W failure | #NAME? | #NAME? | Module B1 are within calibrated parameters. Module B1 and B3 are operational. | Used one XPC cartridge from stock in office. Waiting for module B2 and B4 module allocation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542300705019928708/5542300705019928708_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5542300705019928708/5542300705019928708_signature_31.png
| 2023-03-10 | 2023-03-09 16:02 | Jose B. Lingad Memorial Regional Hospital | MacArthur Hwy, San Fernando Pampanga | 3 | Pampanga | Microbiology | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23725 | PBSP | peachmobi@gmail.com | 27 | ASSAY: Others, COUNT: 0 | 2023-03-09 | Genexpert IV | 110016957 | N/a | N/a | Installed Genexpert Workstation Installed Genexpert Installed Barcode Scanner and Printer Installed Drivers Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Package | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542290057414671620/5542290057414671620_signature_28.png | Phoebe Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5542290057414671620/5542290057414671620_signature_31.png
| 2023-03-10 | 2023-03-09 15:24 | Jose B. Lingad Memorial Regional Hospital | San Fernando City, Pampanga | 3 | Pampanga | HIV Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23724 | COMMERCIAL | concepcionalfonso2@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 62 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 29 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1198 | 2023-03-07 | Genexpert IV | 812571 | Laptop Keyboard not working, Barcode scanner sometimes not working | Calibrate the bar code scanner | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement - 684514 - 210003967 | Genexpert Module - 210003967 | 9673 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542290057416115325/5542290057416115325_signature_28.png | CONCEPCION G. ALFONSO | https://www.jotform.com/uploads/macarewecare/222927978372067/5542290057416115325/5542290057416115325_signature_31.png
| 2023-03-10 | 2023-03-10 11:19 | Bogo City Health Office | Bogo City, Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23723 | PBSP | ajerahmee@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 906 ASSAY: Xpert MTB/RIF Ultra, COUNT: 443 |  | GX IV | 808630 | Defective module 3 sn210003928 | Clean fan filter, replace module 3 | New module installed module 3 sn210049371 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542287076375901030/5542287076375901030_signature_28.png | Jerahmee Anticamara | https://www.jotform.com/uploads/macarewecare/222927978372067/5542287076375901030/5542287076375901030_signature_31.png
| 2023-03-10 | 2023-03-09 14:51 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23722 | PBSP | abigailgsula@gmail.com | 26.3 | ASSAY: Xpert MTB/RIF, COUNT: 207 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1060 | 2023-02-03 | GXIV-D | 831912 | Due for Calibration | Calibration Preventive maintenance | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except A1 and A2 | 2 modules passed calibration | 2 xpertcheck kits | 9582 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542180405315244746/5542180405315244746_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5542180405315244746/5542180405315244746_signature_31.png
| 2023-03-10 | 2023-03-09 12:03 | Minalin Rural Health Unit | Minalin Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23721 | PBSP | bernadetteann0318@gmail.com | 30.5 | ASSAY: Xpert MTB/RIF, COUNT: 94 ASSAY: Xpert MTB/RIF Ultra, COUNT: 547 |  | GXIV-D | 834688 | Defective UPS( Sudden shutdown) | Troubleshooting | Test UPS Battery voltage Performed functionality test | Unit is operational, Deffective UPS for reporting | UPS for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5542180405317211530/5542180405317211530_signature_28.png | Bernadette Ann Lazaro | https://www.jotform.com/uploads/macarewecare/222927978372067/5542180405317211530/5542180405317211530_signature_31.png
| 2023-03-09 | 2023-03-09 11:25 | BILIRAN PROVINCIAL HOSPITAL | Naval Biliran | 8 | Biliran | TBdots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23720 | PBSP | ghelmission13@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 684 ASSAY: Xpert MTB/RIF Ultra, COUNT: 598 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 | 2023-03-09 | GXIV | 815273 | Module B3 for Replacement | -Module replacement B3 SN 890236 - 803273(new) -Performed preventive maintenance -Performed functionality test -Generate reports | #NAME? | Unit is Operational. Module B1 Disabled (Temperature Error) |  | GeneXpert Module 6 Color SN 803273 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541820211527220742/5541820211527220742_signature_28.png | GHELAND M. CORTO | https://www.jotform.com/uploads/macarewecare/222927978372067/5541820211527220742/5541820211527220742_signature_31.png
| 2023-03-09 | 2023-03-09 15:41 | Caraga Regional Hospital | Rizal Street, Brgy. Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | INSTALLATION | CEPHEID | Faizar Avenido |  |  |  |  | 23719 | PBSP | apabellan.pbsp@gmail.com | 18.9 | ASSAY: Others, COUNT: 0 |  | GX-4 (10 colors) | 110016959 | None | None | Installation of GeneXpert XDR Machine and its peripherals | Unit is operational. Desktop CPU Information CPU Name: ThinkCenter M70s SFF Cepheid S/N: CPH0322F32974 Cepheid P/N: 800-0560GXDX Model No.:11DBS07100 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541602044895128967/5541602044895128967_signature_28.png | Andrew Pabellan | https://www.jotform.com/uploads/macarewecare/222927978372067/5541602044895128967/5541602044895128967_signature_31.png
| 2023-03-09 | 2023-03-09 15:50 | Danao City Health Office | Danao City, Cebu | 7 | Cebu | Laboratory | NO | INSTALLATION REPAIR Upgrage OS to 5.1 | CEPHEID | Raymond Laurito |  |  |  |  | 23718 | PBSP | chockenriquez@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1260 ASSAY: , COUNT: 1056 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GX IV | 808286 | Defective module 2 sn 675213 | Update OS version to 5.1 | New module 2 sn700309 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541586389174750871/5541586389174750871_signature_28.png | Nicanor Enriquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5541586389174750871/5541586389174750871_signature_31.png
| 2023-03-09 | 2023-03-09 10:40 | OFW Hospital | PEO Compound, Sindalan, City of San Fernando, Pampanga | 3 | Pampanga | Microbiology | NO | INSTALLATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23716 | MACARE-HBDC | jtmromero19@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HPV, COUNT: 0 | 2023-03-08 | Genexpert IV | 110018556 | N/a | N/a | Installed CPU Installed Printer Installed Monitor Installed Genexpert Installed Ups, barcode scanner Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | Genexpert Package with AVR and Printer | 94877 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541558309814272430/5541558309814272430_signature_28.png | John Terrence Romero | https://www.jotform.com/uploads/macarewecare/222927978372067/5541558309814272430/5541558309814272430_signature_31.png
| 2023-03-09 | 2023-03-09 13:57 | Jose B. Lingad Memorial Regional Hospital | MacArthur Hwy, San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23717 | PBSP | peachmobi@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 399 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1762 | 2023-03-09 | Genexpert IV | 804763 | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Performed Maintenance | Unit is Operational | Module Replacement - 677185 - 721058 | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541558319813497017/5541558319813497017_signature_28.png | Phoebe L. Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5541558319813497017/5541558319813497017_signature_31.png
| 2023-03-09 | 2023-03-09 13:03 | Nabunturan Rural Health Unit | Poblacion, Nabunturan Davao De Oro | 11 | Davao de Oro | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23715 | WHO | danneth_pcanastra@yahoo.com | 26.6 | ASSAY: Xpert MTB/RIF, COUNT: 325 ASSAY: Xpert MTB/RIF Ultra, COUNT: 691 |  | GX-IV | 847253 | For Preventive Maintenance Multiple Error at Module B3 | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Peform Functionality Test on Module B3 | Perform Preventive Maintenance | Module B3 S/N: 910202 Encounter Multiple Temperature Related Error, Datas will be sent to Cepheid for Further Assessment 3 Modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541519556949861113/5541519556949861113_signature_28.png | Danneth P. Canastra | https://www.jotform.com/uploads/macarewecare/222927978372067/5541519556949861113/5541519556949861113_signature_31.png
| 2023-03-09 | 2023-03-09 12:29 | GMA Rural Health Unit II | General Emilio Aguinaldo | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23714 | PBSP | ldlconti@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1150 |  | GXIV-D | 814152 | Module 3 with temperature error | Module Replacement | Performed Maintenance Replaced Module 3 Generate Reports | All modules are operational | 696058->653731 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506378313603695/5541506378313603695_signature_28.png | Lorena L. Conti | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506378313603695/5541506378313603695_signature_31.png
| 2023-03-09 | 2023-03-08 12:00 | Paliparan CHO IV | Paliparan, Dasmariñas | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23713 | PBSP | krlandaya25@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1571 |  | GXIV-D | 838139 | Modules failed in XpertCheck | Module Replacement | Replaced 3 Modules Performed Maintenance Generate reports | All modules are operational | 210070791, 806695, 910289 -> 731304, 731947, 750228 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038313304059/5541506038313304059_signature_28.png | Khristofer Ralph L. Andaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038313304059/5541506038313304059_signature_31.png
| 2023-03-09 | 2023-03-07 12:55 | Almanza Uno Health Center | Alabang-Zapote Rd., Almanza Uno | NCR | NCR 4th - Las Piñas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23712 | DOH | kmbutingan@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1302 ASSAY: Xpert HIV-1 Viral Load, COUNT: 188 | 2023-01-16 | GXIV-D | 831918 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 2 modules for replacement, 2 modules are operational | XC: 2 modules passed | XpertCheck kit | DR 9464 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038314698746/5541506038314698746_signature_28.png | Alyssa Kiana Ysabel B. Rovillos | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038314698746/5541506038314698746_signature_31.png
| 2023-03-09 | 2023-03-07 11:13 | Alabang Health Center | Alabang | NCR | NCR 4th - Muntinlupa | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23710 | DOH | cathycarasig@yahoo.com | 26 | ASSAY: Xpert HPV, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1194 | 2023-01-16 | GXIV-D | 831913 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 1 module reported, 3 passed | XC: 3 modules passed | XpertCheck Kit | DR 9465 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038316039361/5541506038316039361_signature_28.png | Catherine Carasig | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038316039361/5541506038316039361_signature_31.png
| 2023-03-09 | 2023-03-09 10:38 | Bulihan Rural Health Unit | Bulihan, Silang | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23711 | PBSP | alei22775@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 686 |  | GXIV-D | 838446 | Module 2 has temperature error | Module Replacement | Replaced module 2 Performed Maintenance Generate Reports | 3 modules are operational, 1 module for reporting | 861471->626198 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038316944196/5541506038316944196_signature_28.png | Alejandra Alegre | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038316944196/5541506038316944196_signature_31.png
| 2023-03-09 | 2023-03-06 15:06 | Bagong Barangay Health Center | West Zamora, Pandacan | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23709 | DOH | thea.peneyra@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 696 | 2023-01-16 | GXIV-D | 831694 | Unit is due for XpertCheck | Maintenance XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 2 modules for replacement, 2 modules are operational | XpertCheck: 2 modules failed, 2 modules passed | XpertCheck Kit | DR 9474 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038311697346/5541506038311697346_signature_28.png | Maria Anthea Peneyra | https://www.jotform.com/uploads/macarewecare/222927978372067/5541506038311697346/5541506038311697346_signature_31.png
| 2023-03-09 | 2023-03-01 10:56 | San Diego Health Center | San Diego St., Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23708 | MACARE-HBDC | johndarrellturato@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 14 ASSAY: Xpert HIV-1 Viral Load, COUNT: 839 | 2023-02-20 | GXIV-D-10-C | 848262 | Module 3 with Temperature Error | Moduke Replacement | Replaced Module 3 Performed Maintenance Generate Reports | All modules are operational | B3:210002569->210078083 | GeneXpert Module Kit 10-Colors | DR 9640 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868314025652/5541505868314025652_signature_28.png | John Darrell Turato | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868314025652/5541505868314025652_signature_31.png
| 2023-03-09 | 2023-02-28 12:26 | University of Perpetual Help Dalta Medical Center | Alabang-Zapote Rd. | NCR | NCR 4th - Las Piñas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23705 | COMMERCIAL | uphdmc_lab@yahoo.com | 28 | ASSAY: Xpert CT/NG, COUNT: 9 ASSAY: Xpert HBV VL, COUNT: 12 ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 145 | 2023-02-22 | GXIV-D | 831316 | Unit is due for XpertCheck | XpertCheck Maintenance | Cleaned Cartridge Bay, Plunger Rod, Icore Lens and Filter Performed XpertCheck Generate Reports | All Modules are Operational | All Modules Passed the XpertCheck | XpertCheck Kit | SI# 94797 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868316591246/5541505868316591246_signature_28.png | Jimenez, Inger Lia | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868316591246/5541505868316591246_signature_31.png
| 2023-03-09 | 2023-03-01 16:34 | San Nicolas Health Center | San Nicolas, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23706 | DOH | guadex7@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 990 | 2023-01-16 | GXIV-D | 831618 | Unit is due for XpertCheck | XpertCheck | Performed XpertCheck Performed Maintenance Generate Reports | 2 modules are operational, 2 modules are defective | XC: 2 modules passed, 1 failed | XpertCheck Kit | DR 9463 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868316703401/5541505868316703401_signature_28.png | Guadalupe Cube | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868316703401/5541505868316703401_signature_31.png
| 2023-03-09 | 2023-02-28 13:43 | Pamplona Uno Health Center | Pamplona Uno | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23704 | PBSP | carolinevecino5821@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 811 |  | GXIV-D | 838541 | Defective Monitor | Monitor Replacement | Replaced Monitor Performed Maintemance Generate Repors | 3 modules for replacement, 1 module operational | Pulled Out backed-up crolinrmonitor of PBSP |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868317189917/5541505868317189917_signature_28.png | Caroline R. Vecino | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868317189917/5541505868317189917_signature_31.png
| 2023-03-09 | 2023-03-02 10:45 | De La Salle Health and Science Institue | D.Mangubat,Dasmariñas City | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23707 | PBSP | acpabruada@dlshsi.edu.ph | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 418 |  | GXIV-L | 707514 | B1 has temperature errors | Module Replacement | Replaced Module 1 Performed Maintenance Generate Reports | All modules are operational | B1: 802387 -> 643051 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868318795985/5541505868318795985_signature_28.png | Angelita Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5541505868318795985/5541505868318795985_signature_31.png
| 2023-03-09 | 2023-03-09 12:57 | Lapulapu. City Health office | Lapu. Lapu City, Pajo | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 23703 | PBSP | tanyapax@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1158 | 2023-03-09 | GXIV | 808637 | Error 4004 valve motion error | Replaced new Module 698837 | Replaced module 688617 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541474324783737621/5541474324783737621_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5541474324783737621/5541474324783737621_signature_31.png
| 2023-03-09 | 2023-03-09 11:14 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Molecular Laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  |  |  |  | 23702 | COMMERCIAL | shenarossmujer@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 9 ASSAY: Xpert MTB/RIF, COUNT: 42 ASSAY: Xpert MTB/RIF Ultra, COUNT: 456 |  | GXII | 812916 | Unit for relocation. | Performed relocation from Main Laboratory to Molecular Laboratory. Generated reports. | Relocation | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541413464936598475/5541413464936598475_signature_28.png | SHENAROSS E. MUJER | https://www.jotform.com/uploads/macarewecare/222927978372067/5541413464936598475/5541413464936598475_signature_31.png
| 2023-03-09 | 2023-03-07 12:01 | East avenue medical center | East ave qc | NCR | NCR 2nd - Quezon City | Bloodbank | YES | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 23701 | MACARE-HBDC | eamc.biomed@gmail.com | 26 |  | 2023-03-02 | Mcs + | 15H044G | No display,suspect cpu board defective | Performed troubleshooting, restart cpu board error no display still occur, test spare cpu board for testing,after replace cpu board machine appear display and system initialization completed | Trouble shooting | Unit unoperational | For replacement of cpu board | Cpu board | 9660 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541412853734179339/5541412853734179339_signature_28.png | Jocelyn Rote | https://www.jotform.com/uploads/macarewecare/222927978372067/5541412853734179339/5541412853734179339_signature_31.png
| 2023-03-09 | 2023-03-07 12:01 | East avenue medical center | East ave qc | NCR | NCR 2nd - Quezon City | Bloodbank | YES | TROUBLESHOOTING | HAEMONETICS | Noel Floresca |  |  |  |  | 23700 | MACARE-HBDC | eamc.biomed@gmail.com | 26 |  | 2023-03-02 | Mcs + | 15H044G | No display,suspect cpu board defective | Performed troubleshooting, restart cpu board error no display still occur, test spare cpu board for testing,after replace cpu board machine appear display and system initialization completed | Trouble shooting | Unit unoperational | For replacement of cpu board | Cpu board | 9660 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541412503737512421/5541412503737512421_signature_28.png | Jocelyn Rote | https://www.jotform.com/uploads/macarewecare/222927978372067/5541412503737512421/5541412503737512421_signature_31.png
| 2023-03-09 | 2023-03-09 10:37 | Balud Rural Health Unit | Balud, Masbate | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23699 | PBSP | rhu_balud1@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 304 | 2023-03-08 | GXIV | 835199 | Barcode scanner for replacement | #NAME? | Replaced barcode | Unit is operational |  | Barcodr Scanner SN: 19255010556824 | G 5250 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541392575426818964/5541392575426818964_signature_28.png | Sharon Oliva | https://www.jotform.com/uploads/macarewecare/222927978372067/5541392575426818964/5541392575426818964_signature_31.png
| 2023-03-09 | 2023-03-09 10:37 | Balud Rural Health Unit | Balud, Masbate | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23698 | PBSP | rhu_balud1@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 304 | 2023-03-08 | GXIV | 835199 | Barcode scanner for replacement | #NAME? | Replaced barcode | Unit is operational |  | Barcodr Scanner SN: 19255010556824 | G 5250 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541392455428914652/5541392455428914652_signature_28.png | Sharon Oliva | https://www.jotform.com/uploads/macarewecare/222927978372067/5541392455428914652/5541392455428914652_signature_31.png
| 2023-03-09 | 2023-03-09 10:37 | Balud Rural Health Unit | Balud, Masbate | 5 | Masbate | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23697 | PBSP | rhu_balud1@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 47 ASSAY: Xpert MTB/RIF Ultra, COUNT: 304 | 2023-03-08 | GXIV | 835199 | Barcode scanner for replacement | #NAME? | Replaced barcode | Unit is operational |  | Barcodr Scanner SN: 19255010556824 | G 5250 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5541389425428648914/5541389425428648914_signature_28.png | Sharon Oliva | https://www.jotform.com/uploads/macarewecare/222927978372067/5541389425428648914/5541389425428648914_signature_31.png
| 2023-03-08 | 2023-03-08 15:16 | Philippine General Hospital | Taft avenue Manila | NCR | NCR 1st - Manila | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  | sdaroya.macare@gmail.com |  |  | 23696 | MACARE-HBDC | pghmolecularlab@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 20245 |  | GX16 | 844975 | Unit for preventive maintenance. For troubleshooting of problematic modules. | Performed preventive maintenance and troubleshooting. Generated reports. | Preventive maintenance and troubleshooting. | Unit is operational. Modules C1 (783584) has temp error and C2 (672269) has valve error. C1 and C2 are confirmed replacement. Re-enabled modules D2 and D4 after preventive maintenance. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540691739124965476/5540691739124965476_signature_28.png | Michelle Ann C. Sanchez | https://www.jotform.com/uploads/macarewecare/222927978372067/5540691739124965476/5540691739124965476_signature_31.png
| 2023-03-08 | 2023-03-08 14:50 | Philippine Redcross Zamboanga | Pettite barracks, Zamboanga City | 9 | Zamboanga del Sur | Blood Collection Section | NO | MAINTENANCE | HAEMONETICS | Gian Carlo Aizon |  |  |  |  | 23695 | COMMERCIAL | zamboangacity@redcross.org.ph | 25 |  | 2023-03-08 | Macare 9000 | 01M256 | None | Preventive maintenance | Preventive maintenance | Machine is available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540690894991607708/5540690894991607708_signature_28.png | Concesa M Pangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5540690894991607708/5540690894991607708_signature_31.png
| 2023-03-08 | 2023-03-08 14:50 | Adela Serra Ty Memorial Hospital | Capitol Hills, Telaje, Tandag City | 13 | Surigao del Sur | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 23694 | PBSP | astmmclaboratory17@gmail.com | 24.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4033 |  | GX-4 | 804255 | XPC Result Monitor does not power on | Troubleshooting | Preventive Maintenance Calibration | 1 module passed and 3 modules service required 2 modules working as of now Temporary Monitor has been used | None |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540681895402230537/5540681895402230537_signature_28.png | Editha Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5540681895402230537/5540681895402230537_signature_31.png
| 2023-03-08 | 2023-03-08 14:06 | Philippine General Hospital | Taft Ave., Manila | NCR | NCR 1st - Manila | UP Medical Research Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Shiedrick Daroya |  |  |  |  | 23693 | PBSP | mmcajucom@up.edu.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 179 ASSAY: Xpert MTB/RIF Ultra, COUNT: 55 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 | 2023-03-08 | GXIV | 802511 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. |  | Xpertcheck | PL-8207 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540680399122895639/5540680399122895639_signature_28.png | MARC AGNEW M. CAJUCOM | https://www.jotform.com/uploads/macarewecare/222927978372067/5540680399122895639/5540680399122895639_signature_31.png
| 2023-03-08 | 2023-03-08 14:00 | Philippine General Hospital | Taft Ave. | NCR | NCR 1st - Manila | UP Medircal Research Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23692 | PBSP | mmcajucom@up.edu.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 179 ASSAY: Xpert MTB/RIF Ultra, COUNT: 55 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 | 2023-03-08 | GXIV | 838145 | Unit for preventive maintenance. Unit for Xpertcheck | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B4 (806707) is service required. |  | Xpertcheck | PL#9949 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540661889121673275/5540661889121673275_signature_28.png | MARC AGNEW M. CAJUCOM | https://www.jotform.com/uploads/macarewecare/222927978372067/5540661889121673275/5540661889121673275_signature_31.png
| 2023-03-08 | 2023-03-08 14:00 | Philippine General Hospital | Taft Ave. | NCR | NCR 1st - Manila | UP Medircal Research Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23691 | PBSP | mmcajucom@up.edu.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 179 ASSAY: Xpert MTB/RIF Ultra, COUNT: 55 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 | 2023-03-08 | GXIV | 838145 | Unit for preventive maintenance. Unit for Xpertcheck | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B4 (806707) is service required. |  | Xpertcheck | PL#9949 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540661869128238210/5540661869128238210_signature_28.png | MARC AGNEW M. CAJUCOM | https://www.jotform.com/uploads/macarewecare/222927978372067/5540661869128238210/5540661869128238210_signature_31.png
| 2023-03-08 | 2023-03-08 13:30 | Macare In-House | Q.C. | NCR | NCR 2nd - Quezon City | Engineering | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23690 | COMMERCIAL | rcb.macare@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 0 |  | GXIV | 801889 | Unit for troubleshooting | #NAME? | Checked up unit | Unit is operational. A1 hardware failed. A2 & A4 service required. XC performed on 2-15-22 by Noel Floresca. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540629721213334547/5540629721213334547_signature_28.png | Macare In-House | https://www.jotform.com/uploads/macarewecare/222927978372067/5540629721213334547/5540629721213334547_signature_31.png
| 2023-03-08 | 2023-03-08 12:36 | Senator Ninoy Aquino Rural Health Unit | Kulaman, Senator Ninoy Aquino | 12 | Sultan Kudarat | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23689 | PBSP | benjieramos1975@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 130 ASSAY: Xpert MTB/RIF Ultra, COUNT: 319 | 2023-03-08 | Gx-IV | 838129 | Module B4 for Service Required | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/ExpertCheck | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/ExpertCheck | 4 operational modules |  | Xpertcheck -CE-5 | 1000378669 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540607122916710761/5540607122916710761_signature_28.png | Biengie E. Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5540607122916710761/5540607122916710761_signature_31.png
| 2023-03-08 | 2023-03-08 11:32 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | JARO ILOILO CITY | 6 | Iloilo | MOLECULAR LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23688 | DOH | mc-molecular@wvsu.edu.ph | 27 | ASSAY: Xpert HPV, COUNT: HPV 16_18-45 : 15 ASSAY: Xpert HPV, COUNT: HPV HR_16_18-45 : 247 |  | GX-IV | 838290 | PREVENTIVE MAINTENANCE | COLLECT IQR, SYSTEM LOG, AND ASSAY STAT | PERFORM CLEANING AND PM OF GX EQUIPMENT | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540559579367083632/5540559579367083632_signature_28.png | Karla Danielle T. Peralta | https://www.jotform.com/uploads/macarewecare/222927978372067/5540559579367083632/5540559579367083632_signature_31.png
| 2023-03-08 | 2023-03-08 11:25 | Macare In-House | Q.C. | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Ronnel Belizar |  |  |  | UPS | 23687 | COMMERCIAL | rcb.macare@gmail.com | 25 |  | 2023-03-07 | AT1000VA | AT1000181128ST0073 | Unit for repair. | #NAME? | Replaced battery | Unit is operational. |  | Battery SN: 459865, 459872 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540554331213922340/5540554331213922340_signature_28.png | Macare In-House | https://www.jotform.com/uploads/macarewecare/222927978372067/5540554331213922340/5540554331213922340_signature_31.png
| 2023-03-08 | 2023-03-08 11:18 | Macare In-House | Quezon City | NCR | NCR 2nd - Quezon City | Engineering | NO | REPAIR | OTHERS | Ronnel Belizar |  |  |  | UPS | 23686 | COMMERCIAL | rcb.macare@gmail.com | 25 |  | 2023-03-07 | RPT-1025AP | 40024441610 | Unit for repair. | #NAME? | Replaced battery | Unit is operational. |  | Battery SN: 459843, 459870 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540549991217192709/5540549991217192709_signature_28.png | Macare In-house | https://www.jotform.com/uploads/macarewecare/222927978372067/5540549991217192709/5540549991217192709_signature_31.png
| 2023-03-08 | 2023-03-07 12:44 | Banlat Health Center | Banlat, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23685 | PBSP | lorinemmendiola@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 185 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1371 |  | GXIV | 804688 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. A1, A4 insufficient kit. A2 passed. A3 failed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540547041216675375/5540547041216675375_signature_28.png | Lorine Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5540547041216675375/5540547041216675375_signature_31.png
| 2023-03-08 | 2023-03-08 11:07 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | E Lopez St Jaro Iloilo | 6 | Iloilo | TB DOTS | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23684 | PBSP | tvanrlujan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 76 |  | GX-IV | 707531 | PREVENTIVE MAINTENANCE | COLLECT IQR,SYSTEM LOG, AND ASSAY | PERFORM CLEANING AND PREVENTIVE MAINTENANCE OF GX EQUIPMENT | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540543689364633247/5540543689364633247_signature_28.png | Van Ryan R. Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5540543689364633247/5540543689364633247_signature_31.png
| 2023-03-08 | 2023-03-08 11:07 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | E Lopez St Jaro Iloilo | 6 | Iloilo | TB DOTS | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23683 | PBSP | tvanrlujan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 76 |  | GX-IV | 707531 | PREVENTIVE MAINTENANCE | COLLECT IQR,SYSTEM LOG, AND ASSAY | PERFORM CLEANING AND PREVENTIVE MAINTENANCE OF GX EQUIPMENT | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540542989367624164/5540542989367624164_signature_28.png | Van Ryan R. Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5540542989367624164/5540542989367624164_signature_31.png
| 2023-03-08 | 2023-03-08 10:42 | WEST VISAYAS STATE UNIVERSITY MEDICAL CENTER | E Lopez St Jaro Iloilo | 6 | Iloilo | TB DOTS | NO | ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23682 | PBSP | tvanrlujan@gmail.com | 25 | ASSAY: Others, COUNT: NA |  | GX-IV | 804699 | SYSTEM UNIT POWER/BOOT ISSUE | #NAME? | -PERFORM TROUBLESHOOTING, CONFIRM SYSTEM UNIT SN:2UA3520VX3 | SYSTEM UNIT IS DEFECTIVE |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540531219168135132/5540531219168135132_signature_28.png | Van Ryan R. Lujan | https://www.jotform.com/uploads/macarewecare/222927978372067/5540531219168135132/5540531219168135132_signature_31.png
| 2023-03-08 | 2023-03-08 10:11 | Nazareth Health Center | Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23681 | DOH | mymy.cruz0824@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 878 |  | GeneXpert | 838269 | None | None | Preventive Maintenance Generate data files | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540509949277823326/5540509949277823326_signature_28.png | Cruz,Michelle Theresa | https://www.jotform.com/uploads/macarewecare/222927978372067/5540509949277823326/5540509949277823326_signature_31.png
| 2023-03-07 | 2023-03-07 13:31 | Paquibato District Health Center | Malabog, Paquibato Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23680 | PBSP | abelongelfie@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF, COUNT: 226 ASSAY: Xpert MTB/RIF Ultra, COUNT: 57 | 2023-03-07 | GX-IV | 838538 | Module B2 for Replacement | Replaced defective Module B2 with S/N: 862295 with new Module S/N : 628622 | Perform Module Replacement | Unit is fully operational |  | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5540125217113714009/5540125217113714009_signature_28.png | Elfie Amor Abelong | https://www.jotform.com/uploads/macarewecare/222927978372067/5540125217113714009/5540125217113714009_signature_31.png
| 2023-03-07 | 2023-03-07 11:21 | Maria Reyna - Xavier University Hospital | Cagayan de Oro City | 10 | Misamis Oriental | Blood Bank | YES | REPAIR | HAEMONETICS | Kim Ryan Arrojado |  |  |  |  | 23679 | COMMERCIAL | bb.mrxuh@gmail.com | 29 |  |  | MCS+ | 05M131 | Computer BIOS CRC test failed | None | Swap CPU board from other machine to isolate the issue Replacement of CPU battery Cleaning of machine | Non-operational | For replacement all pump rotor roller and for calibration. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539939823324744198/5539939823324744198_signature_28.png | MARIA REYNA XAVIER UNIVERSITY HOSPITAL INC. | https://www.jotform.com/uploads/macarewecare/222927978372067/5539939823324744198/5539939823324744198_signature_31.png
| 2023-03-07 | 2023-03-07 11:21 | Maria Reyna - Xavier University Hospital | Cagayan de Oro City | 10 | Misamis Oriental | Blood Bank | YES | REPAIR | HAEMONETICS | Kim Ryan Arrojado |  |  |  |  | 23678 | COMMERCIAL | bb.mrxuh@gmail.com | 29 |  |  | MCS+ | 05M131 | Computer BIOS CRC test failed | None | Swap CPU board from other machine to isolate the issue Replacement of CPU battery Cleaning of machine | Non-operational | For replacement all pump rotor roller and for calibration. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539939603323696429/5539939603323696429_signature_28.png | MARIA REYNA XAVIER UNIVERSITY HOSPITAL INC. | https://www.jotform.com/uploads/macarewecare/222927978372067/5539939603323696429/5539939603323696429_signature_31.png
| 2023-03-07 | 2023-03-07 15:31 | Bataan General Hospital | Balanga Bataan | 3 | Bataan | PCR Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23677 | PBSP | honeyeichel@gmail.com | 25.2 | ASSAY: Xpert SARS-CoV-2, COUNT: 16487 | 2023-02-24 | GXIV-D | 834433 | Due for D2 module replacement | Preventive maintenance Repair (Module replacement) | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Unit is operational |  | Genexpert module | 39289 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539859485712125013/5539859485712125013_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5539859485712125013/5539859485712125013_signature_31.png
| 2023-03-07 | 2023-03-07 14:09 | Pilar Rural Health Unit | Pilar Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23676 | DOH | aimna_24@yahoo.com | 27.7 | ASSAY: Xpert MTB/RIF, COUNT: 135 ASSAY: Xpert MTB/RIF Ultra, COUNT: 123 | 2023-02-07 | GXIV-D | 831613 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational except B3 | 3 modules passed Calibration | 3 xpertcheck kits | 9597 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539859485717661508/5539859485717661508_signature_28.png | Marciana Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5539859485717661508/5539859485717661508_signature_31.png
| 2023-03-07 | 2023-03-07 15:43 | Bataan General Hospital | Balanga Bataan | 3 | Bataan | PCR Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23675 | PBSP | honeyeichel@gmail.com | 25.2 | ASSAY: Xpert SARS-CoV-2, COUNT: 16487 |  | GXIV-D | 834433 | Due for calibration | Preventive maintenance CALIBRATION | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed Calibration, 3 xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539859455712126092/5539859455712126092_signature_28.png | Honey Eichel Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5539859455712126092/5539859455712126092_signature_31.png
| 2023-03-07 | 2023-03-07 13:58 | Sto. Domingo District Hospital | Puring Bulo, Sto. Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23674 | DOH | lhaydeesel@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 214 ASSAY: Xpert MTB/RIF Ultra, COUNT: 730 | 2023-02-07 | Genexpert IV | 831582 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck - All Modules Passed | Xpertcheck kit | 9595 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539806081911612728/5539806081911612728_signature_28.png | Cyndrell B. Samatra | https://www.jotform.com/uploads/macarewecare/222927978372067/5539806081911612728/5539806081911612728_signature_31.png
| 2023-03-07 | 2023-03-07 13:49 | WESTERN VISAYAS MEDICAL CENTER | Q ABETO MANDURRIAO ILOILO CITY | 6 | Iloilo | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23673 | WHO | mitorilla.pbsp@gmail.com | 20 | ASSAY: Others, COUNT: MTB XDR - 1000 |  | GX-IV 10C | 848241 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | -MODULE 3 IS DISABLED DUE TO H/W FAIL |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539776283314738523/5539776283314738523_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5539776283314738523/5539776283314738523_signature_31.png
| 2023-03-07 | 2023-03-07 13:43 | WESTERN VISAYAS MEDICAL CENTER | Q ABETO MANDURRIAO ILOILO CITY | 6 | Iloilo | LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23672 | PBSP | mitorilla.pbsp@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 5580 ASSAY: Xpert MTB/RIF, COUNT: 921 |  | GX-XVI | 839316 | FOR PREVENTIVE MAINTENANCE | COLLECT IQR ,SYSTEM,LOG AND ASSAY STAT | PERFORM CLEANING AND PM OF GX EQUIPMEMT | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539772833311742864/5539772833311742864_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5539772833311742864/5539772833311742864_signature_31.png
| 2023-03-07 | 2023-03-07 13:33 | Chinese General Hospital | Blumentritt, Sta. Cruz | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | Shiedrick Daroya |  |  |  |  | 23671 | COMMERCIAL | jmescolano@cghmc.com.ph | 25 |  |  | MCS+ | 08M463 | Unit for preventive maintenance. | Performed preventive maintenance. Verified calibration parameters. Performed simulation running. | Preventive maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539769385811931434/5539769385811931434_signature_28.png | Joan Escolano | https://www.jotform.com/uploads/macarewecare/222927978372067/5539769385811931434/5539769385811931434_signature_31.png
| 2023-03-07 | 2023-03-07 13:32 | WESTERN VISAYAS MEDICAL CENTER | Q ABETO MANDURRIAO ILOILO CITY | 6 | Iloilo | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 23670 | PBSP | mitorilla.pbsp@gmail.com | 20 | ASSAY: Xpert HPV, COUNT: 1424 ASSAY: Xpert MTB/RIF, COUNT: 193 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1004 ASSAY: Xpert HIV-1 Viral Load, COUNT: 256 |  | GX-4 | 706910 | MODULE A2 REPLACEMENT | #NAME? | -PERFORM MODULE A2 DEF SN:635736 REPLACEMENT WITH NEW SN:914749 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539768883318323341/5539768883318323341_signature_28.png | Michel Joy Torilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5539768883318323341/5539768883318323341_signature_31.png
| 2023-03-07 | 2023-03-07 12:42 | Abucay Rural Health Unit | Abucay Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23669 | DOH | parreraledelina22176@gmail.com | 27.7 | ASSAY: Xpert MTB/RIF, COUNT: 179 ASSAY: Xpert MTB/RIF Ultra, COUNT: 216 | 2023-02-07 | GXIV-D | 831878 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational excel B2 and B4 | All modules passed calibration | 2 xpertcheck kits | 9604 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539748767716571920/5539748767716571920_signature_28.png | Ledelina M. Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5539748767716571920/5539748767716571920_signature_31.png
| 2023-03-07 | 2023-03-07 12:46 | Chinese General Hospital | Blumentritt, Sta. Cruz | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23668 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 8 ASSAY: Xpert CT/NG, COUNT: 24 ASSAY: Xpert HBV VL, COUNT: 38 ASSAY: Xpert HPV, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1205 ASSAY: Xpert HCV VL, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | GXIV | 835604 | Unit for preventive maintenance. For troubleshooting of modules with error. | Performed preventive maintenance. Performed troubleshooting. Generated reports. | Preventive maintenance and teoubleshooting. | Unit is operational. Module B1 already reported. Module B3 and B4 has sample related errors. Module. B2 for reporting due to temperature related errors. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539739725812610250/5539739725812610250_signature_28.png | Maria Odessa R. Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/5539739725812610250/5539739725812610250_signature_31.png
| 2023-03-07 | 2023-03-07 11:57 | Soccsksargen General Hospital | Surallah | 12 | South Cotabato | Laboratory | YES | CALIBRATION | CEPHEID | Jesson Tupas |  |  |  |  | 23666 | PBSP | Johnellamarisbautista@gmail.com | 24.9 | ASSAY: Xpert MTB/RIF, COUNT: 178 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1061 | 2023-03-07 | Gx-IV | 815432 | Module B3 for Service Required | Perform Genexpert machine Calibration/Xpertcheck | Perform Genexpert machine Calibration/Xpertcheck | 4 operational modules |  | Xpertcheck -CE-5 | 1000330696 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539714748464528938/5539714748464528938_signature_28.png | Johnella Maris R. Bautista, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5539714748464528938/5539714748464528938_signature_31.png
| 2023-03-07 | 2023-03-06 14:52 | Soccsksargen General Hospital | Surallah | 12 | South Cotabato | Laboratory | YES | REPAIR | CEPHEID | Jesson Tupas |  |  |  |  | 23667 | PBSP | johnellamarisbautista@gmail.com | 29.4 | ASSAY: Xpert MTB/RIF, COUNT: 181 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1040 |  | Gx-IV | 815432 | Genexpert CPU Problem | Troubleshoot and repair Genexpert CPU Replace CMOS battery | Troubleshoot and repair Genexpert CPU Replace CMOS battery | 4 operational modules and CPU problem solved |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539714748467329250/5539714748467329250_signature_28.png | Johnella Maris Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5539714748467329250/5539714748467329250_signature_31.png
| 2023-03-07 | 2023-03-07 11:15 | Samal Rural Health Unit | Samal Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23665 | DOH | vickytholits2022@gmail.com | 27.7 | ASSAY: Xpert MTB/RIF, COUNT: 262 ASSAY: Xpert MTB/RIF Ultra, COUNT: 367 | 2023-02-07 | GXIV-D | 831729 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed calibration | 3 xpertcheck kits | 9603 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539693277717404799/5539693277717404799_signature_28.png | Maria Victoria Banzuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5539693277717404799/5539693277717404799_signature_31.png
| 2023-03-07 | 2023-03-07 11:21 | Maria Reyna - Xavier University Hospital | Cagayan de Oro City | 10 | Misamis Oriental | Blood Bank | YES | REPAIR | HAEMONETICS | Kim Ryan Arrojado |  |  |  |  | 23664 | COMMERCIAL | bb.mrxuh@gmail.com | 29 |  |  | MCS+ | 05M131 | Computer BIOS CRC test failed | None | Swap CPU board from other machine to isolate the issue Replacement of CPU battery Cleaning of machine | Non-operational | For replacement all pump rotor roller and for calibration. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539691034012233577/5539691034012233577_signature_28.png | MARIA REYNA XAVIER UNIVERSITY HOSPITAL INC. | https://www.jotform.com/uploads/macarewecare/222927978372067/5539691034012233577/5539691034012233577_signature_31.png
| 2023-03-07 | 2023-03-07 10:05 | Orani Rural Health Unit | Orani Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23663 | DOH | oranirhu@yahoo.com | 28.2 | ASSAY: Xpert MTB/RIF, COUNT: 168 ASSAY: Xpert MTB/RIF Ultra, COUNT: 260 | 2023-02-07 | GXIV-D | 831668 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed | 4 Xpertcheck kits | 9602 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539667507716747405/5539667507716747405_signature_28.png | Rosalie Tayag | https://www.jotform.com/uploads/macarewecare/222927978372067/5539667507716747405/5539667507716747405_signature_31.png
| 2023-03-07 | 2023-03-07 10:05 | Orani Rural Health Unit | Orani Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23662 | DOH | oranirhu@yahoo.com | 28.2 | ASSAY: Xpert MTB/RIF, COUNT: 168 ASSAY: Xpert MTB/RIF Ultra, COUNT: 260 | 2023-02-07 | GXIV-D | 831668 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of plunger rod,icore,filter and cartridge bay Archived test results Xpertcheck application update to 1.5a Performed calibration Gathered reports Performed functionality test | Unit is operational | All modules passed | 4 Xpertcheck kits | 9602 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539665587713832260/5539665587713832260_signature_28.png | Rosalie Tayag | https://www.jotform.com/uploads/macarewecare/222927978372067/5539665587713832260/5539665587713832260_signature_31.png
| 2023-03-07 | 2023-03-07 08:44 | Hermosa Rural Health Unit | Hermosa Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23661 | DOH | datudianemichelle@yahoo.com | 24.2 | ASSAY: Xpert MTB/RIF, COUNT: 179 ASSAY: Xpert MTB/RIF Ultra, COUNT: 289 | 2023-02-07 | GXIV-D | 831645 | Due for calibration | Preventive maintenance Calibration | Performed surface cleaning Cleaning of filter,plunger rod,icore and cartridge bay Xpertcheck application update from 1.4 to 1.5a Archived test results Performed calibration Gathered reports Performed functionality test | Unit is operational except A2:756767 | Failed Xpertcheck A2 for reporting | 5 pcs Xpertcheck kits | 9598 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539609521612131711/5539609521612131711_signature_28.png | Diane Ersando | https://www.jotform.com/uploads/macarewecare/222927978372067/5539609521612131711/5539609521612131711_signature_31.png
| 2023-03-07 | 2023-03-07 08:34 | Cadiz City Health Office | Cadiz City | 7 | Negros Occidental | Tb Laboratory | YES | INSTALLATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 23660 | COMMERCIAL | zhevebacomo@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 194 | 2023-03-03 | Gx4 | 838499 | UPs problem | Install Ups for genexpert machine | Install Ups for the 2 GX units | Operational |  | Ups | Ups |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5539589537224783970/5539589537224783970_signature_28.png | Zherlyn Eve Bacomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5539589537224783970/5539589537224783970_signature_31.png
| 2023-03-06 | 2023-03-06 15:37 | St. Luke's Medical center | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Microbiology | YES | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23659 | COMMERCIAL | kfuy@macare-medicals.com | 25 | ASSAY: Xpert CT/NG, COUNT: 237 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2902 | 2023-03-06 | Gx4 | 839234 | For troubleshooting B3 encounter intermittent h/w failed b4 module encounter 5007 | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Performed xpertcheck and maintenance | Retry b1,b2,b4 failed b3 xpert check | Transfer gx 807155 from c.i to micro for temporary use of the machine, | Xpertcheck | 0 |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5539162312777833202/5539162312777833202_signature_28.png | LIZA Apo | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5539162312777833202/5539162312777833202_signature_31.png
| 2023-03-06 | 2023-03-06 13:59 | Maria Reyna Xavier University Hospital, Inc. | Cagayan de Oro City | 10 | Misamis Oriental | Blood Bank | YES | TROUBLESHOOTING | HAEMONETICS | Kim Ryan Arrojado |  |  |  |  | 23658 | COMMERCIAL | bb.mrxuh@gmail.com | 29 |  |  | MCS+ | 05M131 | Computer Bios crc test failed | N/A | Troubleshooting of the cpu board | Non-operational | Need replacement of cpu board battery, UPS and replacement of all assay rotors |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538940521458156882/5538940521458156882_signature_28.png | Maria Reyna Xavier University Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/5538940521458156882/5538940521458156882_signature_31.png
| 2023-03-06 | 2023-03-06 13:59 | Maria Reyna Xavier University Hospital, Inc. | Cagayan de Oro City | 10 | Misamis Oriental | Blood Bank | YES | TROUBLESHOOTING | HAEMONETICS | Kim Ryan Arrojado |  |  |  |  | 23657 | COMMERCIAL | bb.mrxuh@gmail.com | 29 |  |  | MCS+ | 05M131 | Computer Bios crc test failed | N/A | Troubleshooting of the cpu board | Non-operational | Need replacement of cpu board battery, UPS and replacement of all assay rotors |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538940431458476397/5538940431458476397_signature_28.png | Maria Reyna Xavier University Hospital | https://www.jotform.com/uploads/macarewecare/222927978372067/5538940431458476397/5538940431458476397_signature_31.png
| 2023-03-06 | 2023-03-06 12:41 | Panabo City Health Office | Poblacion Panabo, Davao Del Norte | 11 | Davao del Norte | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23656 | DOH | nadineeust67@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1590 | 2023-03-06 | GX-IV 4D | 838295 | For Preventive Maintenance Module B1 reported to have Multiple Error | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Run XpertCheck on Module B1 and Dummy Test | Perform Preventive Maintenance Troubleshoot Module B1 | Module B1 Passed the XpertCheck but still Error 4019 exist, Data will be sent to Cepheid for further Assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538912478125025904/5538912478125025904_signature_28.png | Nadine Gemma Eustaquio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5538912478125025904/5538912478125025904_signature_31.png
| 2023-03-06 | 2023-03-06 12:41 | Panabo City Health Office | Poblacion Panabo, Davao Del Norte | 11 | Davao del Norte | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23655 | DOH | nadineeust67@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1590 | 2023-03-06 | GX-IV 4D | 838295 | For Preventive Maintenance Module B1 reported to have Multiple Error | Clean Surface Area, Cartridge Bay, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Run XpertCheck on Module B1 and Dummy Test | Perform Preventive Maintenance Troubleshoot Module B1 | Module B1 Passed the XpertCheck but still Error 4019 exist, Data will be sent to Cepheid for further Assessment |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538912438127873176/5538912438127873176_signature_28.png | Nadine Gemma Eustaquio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5538912438127873176/5538912438127873176_signature_31.png
| 2023-03-06 | 2023-03-06 13:13 | Cristino M. Paragas Memorial Hospital | Quiniput, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23654 | PBSP | fatimaabegailabdurahman@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2023-03-06 | GX 4 | 815269 | System Unit unable to Start | Troubleshoot, double check and reconnected all cables and memory inside the system unit | Preventive Maintemance, Troubleshoot | Unusable | System unable to start | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538906376646409178/5538906376646409178_signature_28.png | Fatima Abegail Abdurahman | https://www.jotform.com/uploads/macarewecare/222927978372067/5538906376646409178/5538906376646409178_signature_31.png
| 2023-03-06 | 2023-03-06 13:00 | Democrito O. Plaza Memorial Hospital | National Highway, Brgy. Patin-ay, Prosperidad | 13 | Agusan del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 23653 | PBSP | lermabpabillan@yahoo.com | 21.4 | ASSAY: Xpert MTB/RIF, COUNT: 530 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2692 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 |  | GX-4 | 805335 | C1 was found error 1002 | None | Replacing Module C2 with SN: 210068708 to SN: 695878 Replacing Module C3 with SN: 681208 to SN: 610085 Replacing Module C4 with SN: 738198 to SN: 686369 | 1 module for replacement. 3 modules working | There is upcoming module for Module C1. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538887333125484094/5538887333125484094_signature_28.png | Lerma B. Pabillan | https://www.jotform.com/uploads/macarewecare/222927978372067/5538887333125484094/5538887333125484094_signature_31.png
| 2023-03-06 | 2023-03-06 11:04 | South Cotabato Provincial Hospital | Koronadal City | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23652 | PBSP | anamariermt22@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF, COUNT: 175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2063 ASSAY: Xpert HIV-1 Viral Load, COUNT: 311 | 2023-03-06 | Gx-IV | 815380 | Module B2, B3, and B4 for Service Required | Replace module B2 SN:696206 with new module SN: 658121 Replace module B3 SN:696210 with new module SN: 676181 Replace module B4 SN:684397 with new module SN: 655764 Genexpert machine Preventive maintenance | Replace module B2 SN:696206 with new module SN: 658121 Replace module B3 SN:696210 with new module SN: 676181 Replace module B4 SN:684397 with new module SN: 655764 Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 colors | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538859015955192212/5538859015955192212_signature_28.png | ANA MARIE GRACE C. EGASAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5538859015955192212/5538859015955192212_signature_31.png
| 2023-03-06 | 2023-03-06 11:04 | South Cotabato Provincial Hospital | Koronadal City | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23651 | PBSP | anamariermt22@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF, COUNT: 175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2063 ASSAY: Xpert HIV-1 Viral Load, COUNT: 311 | 2023-03-06 | Gx-IV | 815380 | Module B2, B3, and B4 for Service Required | Replace module B2 SN:696206 with new module SN: 658121 Replace module B3 SN:696210 with new module SN: 676181 Replace module B4 SN:684397 with new module SN: 655764 Genexpert machine Preventive maintenance | Replace module B2 SN:696206 with new module SN: 658121 Replace module B3 SN:696210 with new module SN: 676181 Replace module B4 SN:684397 with new module SN: 655764 Genexpert machine Preventive maintenance | 4 operational modules |  | Genexpert module 6 colors | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538858855956359366/5538858855956359366_signature_28.png | ANA MARIE GRACE C. EGASAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5538858855956359366/5538858855956359366_signature_31.png
| 2023-03-06 | 2023-03-02 15:08 | Cabuyao City Health Office 1 | Cabuyao City Hall | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23650 | PBSP | shirleypalcantara18@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 154 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1345 | 2023-02-27 | GXIV | 812794 | Unit for preventive maintenance. For module replacement of module B3. | Performed preventive maintenance. Performed module replacement on module B3. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. For replacement of monitor. |  | GX module B3 (730957) in replacement to defective B3 (709361). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538842321216972974/5538842321216972974_signature_28.png | Shirley Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5538842321216972974/5538842321216972974_signature_31.png
| 2023-03-06 | 2023-03-02 11:23 | Sta. Rosa Ciry Health Office 2 | Sta. Rosa | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23649 | DOH | annie.soliman@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 137 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1666 | 2023-01-17 | GXIV | 821156 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck | DR 9481 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538842321218275738/5538842321218275738_signature_28.png | Annie Soliman | https://www.jotform.com/uploads/macarewecare/222927978372067/5538842321218275738/5538842321218275738_signature_31.png
| 2023-03-06 | 2023-03-03 13:43 | Sto. Tomas City Health Office | Sampalocan rd. Sto. Tomas | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23648 | PBSP | bathanevelyn@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1085 | 2023-02-27 | GXIV | 812786 | Unit for preventive maintenance. For module replacement of module B4. | Performed preventive maintenance. Performed module replacement on module B4. Generated reports. | Preventive maintenance and module replacement. | Unit is opetational. |  | Gx Module B4 (805766) in replacement to defective module B4 (671461). | G |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5538842321219111236/5538842321219111236_signature_28.png | Evelyn A. Bathan | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5538842321219111236/5538842321219111236_signature_31.png
| 2023-03-06 | 2023-03-03 09:43 | Mabini Rural Health Unit | Mabini | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23647 | PBSP | glee.maravilla@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 28 ASSAY: Xpert MTB/RIF Ultra, COUNT: 962 | 2023-02-27 | GXIV | 838867 | Unit for preventive maintenance. For module replacement of module B2. | Performed preventive maintenance. Performed module replacement on module B2. Generated reports. | Preventive maintenance and module replacement. | Unit is operational |  | GX module B2 (799769) in replacement to defective B2 (863440). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538842321219437934/5538842321219437934_signature_28.png | Gliceria C. Maravilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5538842321219437934/5538842321219437934_signature_31.png
| 2023-03-06 | 2023-03-03 12:19 | Dagupan City Health Office | Dagupan Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23646 | DOH | thisisformyfamily@yahoo.com | 25.07 | ASSAY: Xpert MTB/RIF, COUNT: 337 ASSAY: Xpert MTB/RIF Ultra, COUNT: 939 |  | GXIV-D | 838286 | Due for Calibration | Performed surface cleaning Cleaning of filter,cartridge bay,icore and plunger rod Xpertcheck application update to 1.5A Performed calibration Gathered reports Performed functionality test | Preventive maintenance Calibration | Unit is operational | 6 pcs Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538786625315149690/5538786625315149690_signature_28.png | Rudolf Joseph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5538786625315149690/5538786625315149690_signature_31.png
| 2023-03-06 | 2023-03-03 14:20 | Rosario Super Health Center | Rosario | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23645 | WHO | eden.cawayan@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 250 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1324 |  | GX-IV-D | 843924 | Module B1/638232 has valve error | Performed Cepheid troubleshooting | Performed maintenance Performed Cepheid troubleshooting Generated reports | Unit is operational. Module B1/638232 not operational but with existing report. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698331213479221/5538698331213479221_signature_28.png | Eden R. Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698331213479221/5538698331213479221_signature_31.png
| 2023-03-06 | 2023-03-01 13:20 | Morong Rural Health Unit | Morong | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23640 | DOH | morongrhu@yahoo.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 133 ASSAY: Xpert MTB/RIF Ultra, COUNT: 330 |  | GX-IV-D | 821213 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698131215768318/5538698131215768318_signature_28.png | Necitas V Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698131215768318/5538698131215768318_signature_31.png
| 2023-03-06 | 2023-03-01 10:26 | Cardona Rural Health Unit | Cardona | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23642 | DOH | rhucardona@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 102 ASSAY: Xpert MTB/RIF Ultra, COUNT: 227 |  | GX-IV-D | 821215 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698131216121502/5538698131216121502_signature_28.png | Renalyn Villaran | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698131216121502/5538698131216121502_signature_31.png
| 2023-03-06 | 2023-03-02 13:55 | Teresa Rural Health Unit | Teresa | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23643 | DOH | teresa_rhu@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 42 ASSAY: Xpert MTB/RIF Ultra, COUNT: 647 |  | GX-IV-D | 820710 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698131217466188/5538698131217466188_signature_28.png | Rosbenilda A. Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698131217466188/5538698131217466188_signature_31.png
| 2023-03-06 | 2023-03-02 10:48 | Camp Capinpin Army Station Hospital | Plaza Aldea, Tanay | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23641 | DOH | kaye.alseco@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 63 ASSAY: Xpert MTB/RIF Ultra, COUNT: 138 |  | GX-IV-D | 831896 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Modules A2 and A3 operational. Modules A1/757276 and A4/ 758194 not operational reported already. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698131219104791/5538698131219104791_signature_28.png | Ma. Corazon S. Alseco, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698131219104791/5538698131219104791_signature_31.png
| 2023-03-06 | 2023-03-03 09:44 | Taytay Rural Health Unit | Taytay | 4A | Rizal | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23644 | PBSP | dannelli11@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1085 |  | GX-IV-D | 831836 | For CPU replacement. | Replaced CPU | Performed maintenance Replacement CPU Run XpertCheck Generated reports | All modules operational. Module B2/639463 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698141219191316/5538698141219191316_signature_28.png | Dino Dannelli Delos Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5538698141219191316/5538698141219191316_signature_31.png
| 2023-03-05 | 2023-03-04 15:53 | St.lukes medical center q.c | E.rodriguez | NCR | NCR 2nd - Quezon City | Microbiology | YES | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23639 | COMMERCIAL | biomed.qc@stlukes.com.ph | 25 | ASSAY: Xpert C. difficile BT, COUNT: 0 ASSAY: Xpert CT/NG, COUNT: 238 ASSAY: Others, COUNT: 4 ASSAY: Xpert HPV, COUNT: 5 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2894 |  | Gx4 | 839234 | All module encountered continuous error b1 4012 b2 5007 error b3 hw failed and 4012,b4 5007 error | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,performed reliability test, | Troubleshooting | Un operational | Room temp 25 degrees |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225562777771339/5538225562777771339_signature_28.png | Phillip Jordan M. Salvador, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225562777771339/5538225562777771339_signature_31.png
| 2023-03-05 | 2023-03-03 09:49 | Rodriguez rhu | Jp rizal Rodriguez rizal | 4A | Rizal | Tb dots | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23637 | PBSP | cchlaboratory2018@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2238 | 2023-03-03 | Gx 4 | 807580 | For module replacement b1 688668 - 783498 b4 666171- 803899 | Performed module replacement, cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900gx | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482773085132/5538225482773085132_signature_28.png | Karl Michael M. Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482773085132/5538225482773085132_signature_31.png
| 2023-03-05 | 2023-03-01 11:21 | Marikina cho | Marikina city | NCR | NCR 2nd - Marikina | Laboratory | NO | INSTALLATION | CEPHEID | Noel Floresca |  |  |  |  | 23636 | PBSP | amymitra22@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | Gx4 | 805211 | Unit from kalinga health to be transfer @ marikina cho without barcode stand | Transfer unit,cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Transfer and re installtion | Operational | Units from kalinga health to be transfer @marikina cho,pullout defective laptop dell 9yk2sm1 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482775222150/5538225482775222150_signature_28.png | Amelia Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482775222150/5538225482775222150_signature_31.png
| 2023-03-05 | 2023-03-03 12:06 | Vetetrans memorial medical center | North ave. Qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23638 | MACARE-HBDC | microlabdop@gmail.com | 31 | ASSAY: Xpert MTB/RIF, COUNT: 460 ASSAY: Xpert MTB/RIF Ultra, COUNT: 514 |  | Gx4 | 839150 | As per end user encounter error 2008, for maintenance | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | Operational | Last xpertcheck 7/8/2021 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482777116320/5538225482777116320_signature_28.png | Maireen Balleras vmmc | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482777116320/5538225482777116320_signature_31.png
| 2023-03-05 | 2023-03-03 09:37 | Rodriguez rhu | J.p rizal Rodriguez rizal | 4A | Rizal | Tb dots | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23634 | PBSP | cchlaboratory2018@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2497 ASSAY: Xpert HIV-1 Viral Load, COUNT: 89 | 2023-03-03 | Gx4 | 838894 | For module replacement b4 862178 - 804586 | Performed module replacement cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Module replacement and maintenance | Operational |  | 900gx | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482777601498/5538225482777601498_signature_28.png | Karl Michael M. Banal | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482777601498/5538225482777601498_signature_31.png
| 2023-03-05 | 2023-02-28 15:07 | Medilinx | Amorsolo st makati city | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23635 | COMMERCIAL | jeselleanneegipto.medilinx@outlook.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 24 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 117 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 101 ASSAY: Xpert MTB/RIF Ultra, COUNT: 404 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 349 |  | Gx4 (10c) | 110001647 | For maintenance | cleared unit,filter,plunger rod,cartridge bay,i core,gathered reports,performed visual check,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482779685549/5538225482779685549_signature_28.png | Jeselle Anne Egipto | https://www.jotform.com/uploads/macarewecare/222927978372067/5538225482779685549/5538225482779685549_signature_31.png
| 2023-03-04 | 2023-03-04 11:42 | Chong Hua Hospital | Fuente, Capitol Site, Cebu | 7 | Cebu | Bloodbank | YES | REPAIR | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 23633 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 23 |  | 2023-03-04 | Mcs+ | 11J219 | Power Switch Defective | Pulled out and replaced Power switch | Replaced New Power Switch | Operational |  | Power Switch button | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5537108710517145683/5537108710517145683_signature_28.png | Arriane Guererro | https://www.jotform.com/uploads/macarewecare/222927978372067/5537108710517145683/5537108710517145683_signature_31.png
| 2023-03-03 | 2023-03-03 12:19 | Dagupan City Health Office | Dagupan Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23632 | DOH | thisisformyfamily@yahoo.com | 25.07 | ASSAY: Xpert MTB/RIF, COUNT: 337 ASSAY: Xpert MTB/RIF Ultra, COUNT: 939 |  | GXIV-D | 838286 | Due for Calibration | Performed surface cleaning Cleaning of filter,cartridge bay,icore and plunger rod Xpertcheck application update to 1.5A Performed calibration Gathered reports Performed functionality test | Preventive maintenance Calibration | Unit is operational | 6 pcs Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536530765313364503/5536530765313364503_signature_28.png | Rudolf Joseph Guillen | https://www.jotform.com/uploads/macarewecare/222927978372067/5536530765313364503/5536530765313364503_signature_31.png
| 2023-03-03 | 2023-03-03 15:05 | Mexico RHU IV | Pandacaqui, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23631 | DOH | annbathan79@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-03-03 | Genexpert IV | 831902 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck - A1, A3 Passed & A2, A4 Service Required | Xpertcheck kit | 9583 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536402642028737221/5536402642028737221_signature_28.png | Mary Ann D. Carreon | https://www.jotform.com/uploads/macarewecare/222927978372067/5536402642028737221/5536402642028737221_signature_31.png
| 2023-03-03 | 2023-03-03 15:05 | Mexico RHU IV | Pandacaqui, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23630 | DOH | annbathan79@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-03-03 | Genexpert IV | 831902 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck - A1, A3 Passed & A2, A4 Service Required | Xpertcheck kit | 9583 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536402252025500695/5536402252025500695_signature_28.png | Mary Ann D. Carreon | https://www.jotform.com/uploads/macarewecare/222927978372067/5536402252025500695/5536402252025500695_signature_31.png
| 2023-03-03 | 2023-03-03 14:43 | Klinika Bernardo Health Center | Ermin Garcia, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23629 | PBSP | klinikabernardolaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 297 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1294 |  | GXIV | 808618 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536400512315218648/5536400512315218648_signature_28.png | Maileen Reyna | https://www.jotform.com/uploads/macarewecare/222927978372067/5536400512315218648/5536400512315218648_signature_31.png
| 2023-03-03 | 2023-03-03 13:29 | Hinigaran Rural Health Unit | Hinigaran, Negros Occidental | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 23628 | PBSP | rhuhinigaran@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 300 ASSAY: Xpert MTB/RIF Ultra, COUNT: 679 | 2079-08-02 | GX IV | 815283 | Module A2 & A3 | Get soft copy of calibration info for reporting, IQR, System Log and Assay Stat | Run Xpert check calibration for 4 modules | Module A2 & A3 failed in calibration | Conduct PM | Xpert Check | Lot No. 1000378669 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536317552016352811/5536317552016352811_signature_28.png | Christine Goza | https://www.jotform.com/uploads/macarewecare/222927978372067/5536317552016352811/5536317552016352811_signature_31.png
| 2023-03-03 | 2023-03-03 10:47 | Quezon City Jail | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23627 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV | 803089 | CPU for troubleshooting | #NAME? | Cleaned RAM | Unit is operational. CPU under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536216933418105627/5536216933418105627_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5536216933418105627/5536216933418105627_signature_31.png
| 2023-03-03 | 2023-03-03 10:47 | Quezon City Jail | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23626 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV | 803089 | CPU for troubleshooting | #NAME? | Cleaned RAM | Unit is operational. CPU under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536211363413779276/5536211363413779276_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5536211363413779276/5536211363413779276_signature_31.png
| 2023-03-03 | 2023-03-03 10:47 | Quezon City Jail | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23625 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GXIV | 803089 | CPU for troubleshooting | #NAME? | Cleaned RAM | Unit is operational. CPU under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536211253414170596/5536211253414170596_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5536211253414170596/5536211253414170596_signature_31.png
| 2023-03-03 | 2023-03-03 10:47 | Quezon City Jail | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23624 | PBSP | carleecruz0@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 803089 | CPU for troubleshooting | #NAME? | Cleaned RAM | Unit is operational. CPU under observation. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536210673417471945/5536210673417471945_signature_28.png | Carlee Mae Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5536210673417471945/5536210673417471945_signature_31.png
| 2023-03-03 | 2023-03-01 13:31 | Palimbang Rural Health Unit | Palimbang, Sultan kudarat | 12 | Sultan Kudarat | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23623 | WHO | tina_villegas2319@yahoo.com | 21.3 | ASSAY: Xpert MTB/RIF, COUNT: 238 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 |  | Gx -lV | 847257 | Module B2 H/W failed due to multiple temperature error 1002 | Perform Genexpert machine Preventive maintenance Module B2 troubleshooting Install assay MTB-RIF Ultra | Perform Genexpert machine Preventive maintenance Module B2 troubleshooting Install assay MTB-RIF Ultra | 3 operational modules | Module B2 Waiting for module to available |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536137390415309592/5536137390415309592_signature_28.png | Cristina Nilda E. Villegas | https://www.jotform.com/uploads/macarewecare/222927978372067/5536137390415309592/5536137390415309592_signature_31.png
| 2023-03-03 | 2023-03-02 14:06 | Rosario District Hospital | Rosario, La Union | 1 | La Union | Laboratory . | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23622 | PBSP | rosariodhlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2023-03-02 | GXIV | 815396 | CPU for replacement | - Replaced defective CPU HP Z240 SN: 2UA7021LZP - performed xpertcheck - Generated reports - Performed functionality test | #NAME? | Unit is operational. All modules passed. |  | CPU Lenovo Desktop SN: CPH0322F32993 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5536101172225514518/5536101172225514518_signature_28.png | Aime Lou Supsupin | https://www.jotform.com/uploads/macarewecare/222927978372067/5536101172225514518/5536101172225514518_signature_31.png
| 2023-03-02 | 2023-03-02 17:29 | Region 1 Medical Center | Dagupan Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23621 | PBSP | romarmtamayo@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 302 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1216 ASSAY: Xpert SARS-CoV-2, COUNT: 72 ASSAY: Xpert HIV-1 Viral Load, COUNT: 165 | 2023-02-24 | GXIV-D | 804726 | Module replacement for modules C1, C3 and C4 | Performed surface cleaning Generate pre-iqr Performed module replacement Generate post iqr Performed preventive maintenance Gathered reports Performed functionality test | Module replacement Preventive maintenance | Unit is operational |  | 3 genexpert Modules | 39310 |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5535590452021191704/5535590452021191704_signature_28.png | Romar Tamayo | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5535590452021191704/5535590452021191704_signature_31.png
| 2023-03-02 | 2023-03-02 14:42 | Region 1 Medical Center | Dagupan Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23620 | PBSP | romarmtamayo@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF, COUNT: 467 ASSAY: , COUNT: 994 ASSAY: Xpert SARS-CoV-2, COUNT: 4194 | 2023-02-24 | GXXVI-D | 738529 | Module replacemet for defective modules | Performed surface cleaning Generate Pre-IQR Performed module replacement Performed preventive maintenance Generate post-iqr Gathered reports Performed functionalitybtest | Preventive maintenace Module replacement | Unit is operational |  | 10 Genexpert modules | 39294 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535590102029952728/5535590102029952728_signature_28.png | Romar Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5535590102029952728/5535590102029952728_signature_31.png
| 2023-03-02 | 2023-03-02 14:42 | Region 1 Medical Center | Dagupan Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23619 | PBSP | romarmtamayo@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF, COUNT: 467 ASSAY: , COUNT: 994 ASSAY: Xpert SARS-CoV-2, COUNT: 4194 | 2023-02-24 | GXXVI-D | 842230 | Module replacemet for defective modules | Performed surface cleaning Generate Pre-IQR Performed module replacement Performed preventive maintenance Generate post-iqr Gathered reports Performed functionalitybtest | Preventive maintenace Module replacement | Unit is operational |  | 9 Genexpert modules | 39294 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535585482026117566/5535585482026117566_signature_28.png | Romar Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5535585482026117566/5535585482026117566_signature_31.png
| 2023-03-02 | 2023-03-02 12:28 | Ormoc City Health Office | Ormoc City, Leyte | 8 | Leyte | TB Dots | NO | ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23618 | PBSP | rainazerene@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 297 |  | GXIV-4D | 804253 | Module B2 - H/W fail | None | #NAME? | All modules are available and operational. | Module B2 under observation |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535488969613367245/5535488969613367245_signature_28.png | Raina Zerene P. Ablen | https://www.jotform.com/uploads/macarewecare/222927978372067/5535488969613367245/5535488969613367245_signature_31.png
| 2023-03-02 | 2023-03-02 13:48 | Mabalacat RHU | Mabalacat City, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23617 | PBSP | perezriyadh@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 250 ASSAY: Xpert MTB/RIF Ultra, COUNT: 942 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 | 2023-03-02 | Genexpert IV | 807564 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement - 683625 - 623586 & 685562 - 626193 | Genexpert Module - 623586 & 626193 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535485940522002577/5535485940522002577_signature_28.png | Riyadh L Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5535485940522002577/5535485940522002577_signature_31.png
| 2023-03-02 | 2023-03-02 13:48 | Mabalacat RHU | Mabalacat City, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23616 | PBSP | perezriyadh@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 250 ASSAY: Xpert MTB/RIF Ultra, COUNT: 942 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 | 2023-03-02 | Genexpert IV | 807564 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement - 683625 - 623586 & 685562 - 626193 | Genexpert Module - 623586 & 626193 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535485450522229540/5535485450522229540_signature_28.png | Riyadh L Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5535485450522229540/5535485450522229540_signature_31.png
| 2023-03-02 | 2023-03-02 13:48 | Mabalacat RHU | Mabalacat City, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23615 | PBSP | perezriyadh@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 250 ASSAY: Xpert MTB/RIF Ultra, COUNT: 942 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 | 2023-03-02 | Genexpert IV | 807564 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement - 683625 - 623586 & 685562 - 626193 | Genexpert Module - 623586 & 626193 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535485390526795935/5535485390526795935_signature_28.png | Riyadh L Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5535485390526795935/5535485390526795935_signature_31.png
| 2023-03-02 | 2023-03-02 14:30 | MABINAY MEDICARE COMMUNITY HOSPITAL | Mabinay | 7 | Negros Oriental | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 23614 | DOH | mmchlaboratory2017@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 269 ASSAY: Xpert MTB/RIF Ultra, COUNT: 253 |  | GX-IV | 838481 | E5011, E2008 | Xpertcheck run | Preventive Maintenance - cleaned module bays, plunger rods, filter, optic slots. | 4 modules operational, module 2 and 3 for observation run using new lot numbers of mtbrif ultra. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535484803116817223/5535484803116817223_signature_28.png | DIANNE C. MI-OT | https://www.jotform.com/uploads/macarewecare/222927978372067/5535484803116817223/5535484803116817223_signature_31.png
| 2023-03-02 | 2023-03-02 14:25 | Victorias City Health Office | Victorias City Negros Occidental | 6 | Negros Occidental | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 23613 | PBSP | vincent_llamoleonor@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 407 |  | Gx4 | 838141 | Defective Module 681375 | Module replaced new 692294 Clean up Icore, Filter, and plunger | Module replaced new 692294 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535478372613103848/5535478372613103848_signature_28.png | Vincent L. Leonor | https://www.jotform.com/uploads/macarewecare/222927978372067/5535478372613103848/5535478372613103848_signature_31.png
| 2023-03-02 | 2023-03-02 13:13 | Sasa District Health Center | Sasa, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23612 | PBSP | mariarosario_mapayo@yahoo.com.sg | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 327 | 2023-03-02 | GX-IV | 838889 | Module B1 and B3 for Module Replacement | Replace the defective Module B1 S/N: 863324 with new Module SN: 669006 Replaced the defective Module B3 SN: 863784 with new Module SN: 656621 Generate IQR, Assay Stat and System Log Archived Data test results | Perform Module Replacement | Unit is fully operational | N/A | 900-0491rs | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535462946811057642/5535462946811057642_signature_28.png | Maria Rosario M. Pino RMT RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5535462946811057642/5535462946811057642_signature_31.png
| 2023-03-02 | 2023-03-02 13:13 | Sasa District Health Center | Sasa, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23611 | PBSP | mariarosario_mapayo@yahoo.com.sg | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 327 | 2023-03-02 | GX-IV | 838889 | Module B1 and B3 for Module Replacement | Replace the defective Module B1 S/N: 863324 with new Module SN: 669006 Replaced the defective Module B3 SN: 863784 with new Module SN: 656621 Generate IQR, Assay Stat and System Log Archived Data test results | Perform Module Replacement | Unit is fully operational | N/A | 900-0491rs | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535462596812738235/5535462596812738235_signature_28.png | Maria Rosario M. Pino RMT RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5535462596812738235/5535462596812738235_signature_31.png
| 2023-03-02 | 2023-03-02 13:13 | Sasa District Health Center | Sasa, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23610 | PBSP | mariarosario_mapayo@yahoo.com.sg | 26.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 327 | 2023-03-02 | GX-IV | 838889 | Module B1 and B3 for Module Replacement | Replace the defective Module B1 S/N: 863324 with new Module SN: 669006 Replaced the defective Module B3 SN: 863784 with new Module SN: 656621 Generate IQR, Assay Stat and System Log Archived Data test results | Perform Module Replacement | Unit is fully operational | N/A | 900-0491rs | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535462526814690232/5535462526814690232_signature_28.png | Maria Rosario M. Pino RMT RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5535462526814690232/5535462526814690232_signature_31.png
| 2023-03-02 | 2023-03-02 12:26 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23609 | DOH | sajecopooh20@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF, COUNT: 668 ASSAY: Xpert MTB/RIF Ultra, COUNT: 303 |  | GXIV-4D | 838555 | None | None | #NAME? | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535434248324942458/5535434248324942458_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5535434248324942458/5535434248324942458_signature_31.png
| 2023-03-02 | 2023-03-02 12:26 | Baybay City Health Office | Baybay City, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23608 | DOH | sajecopooh20@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF, COUNT: 668 ASSAY: Xpert MTB/RIF Ultra, COUNT: 303 |  | GXIV-4D | 838555 | None | None | #NAME? | All modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535434098324339110/5535434098324339110_signature_28.png | Sarah Jean Gabieta | https://www.jotform.com/uploads/macarewecare/222927978372067/5535434098324339110/5535434098324339110_signature_31.png
| 2023-03-02 | 2023-03-02 10:27 | Hilongos Rural Health Unit | Hilongos Leyte | 8 | Leyte | TB dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23607 | PBSP | leovillaflor@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 568 ASSAY: Xpert MTB/RIF Ultra, COUNT: 568 | 2023-03-02 | GXIV | 838143 | B4 for module replacement | -Module replacement B4 SN 806456 into 686518 -Performed Preventive maintenance -Performed functionality test -Generate reports | #NAME? | Module B2 H/W Failed ( Temperature Error) |  | GeneXpert Module 6 Color SN 686518 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535345607216119369/5535345607216119369_signature_28.png | Leonides V. Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5535345607216119369/5535345607216119369_signature_31.png
| 2023-03-02 | 2023-03-02 10:19 | Maitum Rural Health Unit | Maitum, Sarangani | 12 | Sarangani | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23606 | PBSP | mikexxxballesteros@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 316 ASSAY: Xpert MTB/RIF Ultra, COUNT: 563 |  | Gx -IV (2 module) | 835075 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535342232813884389/5535342232813884389_signature_28.png | Michael F. Ballesteros, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5535342232813884389/5535342232813884389_signature_31.png
| 2023-03-02 | 2023-03-02 09:39 | Tarlac City Health Center 2 | Matatalaib Tarlac City | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23605 | DOH | audreymarcelo11@gmail.com | 28.125 | ASSAY: Xpert MTB/RIF, COUNT: 181 ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 | 2023-02-02 | GXIV-D | 831917 | Due for calibration | Performed surface cleaning Cleaning of plunger rod,icore,cattridge bay and filter Archived test results Xpertcheck application update to 1.5A Performed calibration Gathered reports Performed functionality test | Preventive maintenance Calibration | Unit is operationap | All modules passed calibration | 5 Xpertcheck kits used | 9572 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535322882029855606/5535322882029855606_signature_28.png | Audrey Capiendo | https://www.jotform.com/uploads/macarewecare/222927978372067/5535322882029855606/5535322882029855606_signature_31.png
| 2023-03-02 | 2023-03-01 14:06 | Tarlac Provincial Hospital | Hospital Drive San Vicente Tarlac City | 3 | Tarlac | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23604 | PBSP | molecularlab.tph@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 663 | 2023-03-01 | GXIV-D | 805108 | C2 for module replacement | Performed surface cleaning Generate pre-iqr Performed module replacement on module C2 Generate pre-iqr Archived test result Gathered reports Performed functionality test | Module replacement Preventive maintenanca | Unit is operational except C3 |  | 1 Genexpert Module | 39293 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535307072022620300/5535307072022620300_signature_28.png | Camille Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5535307072022620300/5535307072022620300_signature_31.png
| 2023-03-02 | 2023-02-28 10:52 | Tondo Health Center | Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23602 | DOH | ma.lorinamanera@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 0 ASSAY: Xpert MTB/RIF, COUNT: 359 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1271 | 2023-01-17 | GXIV | 831620 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck on module 4. | Unit is operational. 1 working module. |  |  | DR |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535298155422528389/5535298155422528389_signature_28.png | Loren M. Lontoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5535298155422528389/5535298155422528389_signature_31.png
| 2023-03-02 | 2023-03-02 09:03 | The Medical City South Luzon | Sta. Rosa | 4A | Laguna | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23603 | COMMERCIAL | tmcsl-molecularlab@themedicalcity.com | 25 | ASSAY: Xpert HBV VL, COUNT: 6 ASSAY: Xpert HPV, COUNT: 6 ASSAY: Xpert MTB/RIF, COUNT: 40 ASSAY: Xpert MTB/RIF Ultra, COUNT: 131 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2023-02-27 | GXIV | 839233 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports | Preventive maintenance and Xpertcheck. | Unit is operational. All modulee passed Xpertcheck. |  | Xpertcheck | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535298155424100941/5535298155424100941_signature_28.png | Ela Maria Victoria Asa | https://www.jotform.com/uploads/macarewecare/222927978372067/5535298155424100941/5535298155424100941_signature_31.png
| 2023-03-02 | 2023-02-28 12:55 | Navotas Wellness and Medical Center | Navotas City | NCR | NCR 3rd - Navotas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23601 | PBSP | alejandrinoknowell@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 725 | 2023-02-28 | GXIV | 807592 | Unit for preventive maintenance. For module replacement of module B3. For generating of reports for module B4. | Performed preventive maintenance. Performed module replacement on module B3. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. For reporting of module B4. |  | GX module B3 (210012275) in replacement to defective module (665858). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535298155425897488/5535298155425897488_signature_28.png | Noel Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5535298155425897488/5535298155425897488_signature_31.png
| 2023-03-02 | 2023-02-28 13:00 | Navotas Medical and Wellness Center | Navotas City | NCR | NCR 3rd - Navotas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23600 | DOH | alejandrinoknowell@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 725 |  | GXIV | 831626 | Unit for preventive maintenance. For Xpertcheck of module A1 due to errors. | Performed preventive maintenance. Performed Xpertcheck on module B1. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B1 passed Xpertcheck. | Used 1 spare Xpertcheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535287265423233471/5535287265423233471_signature_28.png | Noel Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5535287265423233471/5535287265423233471_signature_31.png
| 2023-03-02 | 2023-03-01 11:04 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23599 | DOH | luzvimindanavidad2@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF, COUNT: 326 ASSAY: Xpert MTB/RIF Ultra, COUNT: 477 |  | GXIV-4D | 838297 | None | None | #NAME? | All four modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5535043468618196747/5535043468618196747_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5535043468618196747/5535043468618196747_signature_31.png
| 2023-03-01 | 2023-03-01 10:56 | Detoxicare | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | INSTALLATION | CEPHEID | Ronnel Belizar |  |  |  |  | 23598 | COMMERCIAL | xtpcr.mdl@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: NA |  | GXIV | 110018312 | Unit for installation. | #NAME? | Installed Unit | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534935062224673273/5534935062224673273_signature_28.png | Mary Joy Bagsic | https://www.jotform.com/uploads/macarewecare/222927978372067/5534935062224673273/5534935062224673273_signature_31.png
| 2023-03-01 | 2023-03-01 15:22 | LIBACAO RURAL HEALTH UNIT | Libacao, Aklan | 6 | Aklan | TB Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 23597 | DOH | haladeeb18@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 20 ASSAY: Xpert MTB/RIF Ultra, COUNT: 41 |  | GX-IV | 838540 | HW fail on modules 1 and 4 | Xpertcheck run | Module bay cleaning, plunger cleaning, filter cleaning, Xpertcheck run on all modules | 2 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534658655912621865/5534658655912621865_signature_28.png | MARIA HALA DEEB ZAULDA | https://www.jotform.com/uploads/macarewecare/222927978372067/5534658655912621865/5534658655912621865_signature_31.png
| 2023-03-01 | 2023-03-01 15:12 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | LABORATORY | NO | PULL OUT | CEPHEID | Joseph Acebedo |  |  |  |  | 23596 | COMMERCIAL | pshclinlab@gmail.com | 26 | ASSAY: Xpert HCV VL, COUNT: 6 ASSAY: Xpert BCR-ABL Ultra, COUNT: 12 ASSAY: Xpert CT/NG, COUNT: 16 ASSAY: Xpert HBV VL, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 20 ASSAY: Xpert MTB/RIF, COUNT: 8 |  | GX-II | 802314 | PULL OUT (END OF CONTRACT) | #NAME? | -COLLECT IQR, SYSTEM LOG, AND ASSAY STAT -PULL OUT OF GX 2 EQUIPMENT, BARCODE SCANNER SN: 1214400502325 ; ACER LAPTOP SN: NXGCUSP001624150A07600, AND HP DESKJET PRINTER SN:CNKK319JM2 | ALL EQUIPMENTS ARE OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534649257915820005/5534649257915820005_signature_28.png | Madeline Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5534649257915820005/5534649257915820005_signature_31.png
| 2023-03-01 | 2023-03-01 15:00 | Carigara District Hospital | Carigara Leyte | 8 | Leyte | TB dots | NO | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23595 | PBSP | chezcairo@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 692 ASSAY: Xpert MTB/RIF Ultra, COUNT: 64 | 2023-03-01 | GXIV | 807619 | Module (A3, A4) for module replacement | -Module replacement A3 - 600865 into 680829 A4 - 745385 into 661864 -Performed preventive maintenance -Performed functionality test -Generate reports | #NAME? | Module A1 - H/W Failed (Temperature Error) Module (A2, A3, A4) - Operational |  |  | GeneXpert Module 6 Color SN( 680829, 661864) |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534649241196495750/5534649241196495750_signature_28.png | charisse c.ragrag | https://www.jotform.com/uploads/macarewecare/222927978372067/5534649241196495750/5534649241196495750_signature_31.png
| 2023-03-01 | 2023-03-01 14:35 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23594 | PBSP | laboratory@loveyourself.ph | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 |  | GXIV | 808649 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. B1 failed, B2, B3, B4 insufficient cartridge. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534648942259677349/5534648942259677349_signature_28.png | Happy Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5534648942259677349/5534648942259677349_signature_31.png
| 2023-03-01 | 2023-03-01 14:49 | Toril Health Center | Poblacion, Toril Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23593 | PBSP | leabuelis@yahoo.com | 24.7 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert HIV-1 Viral Load, COUNT: 978 | 2023-03-01 | GX-IV | 838865 | Module B2 and B4 for Replacement | Replaced the Defective Module B2 with SN: 861528 with new Module SN 607055 Replaced the Defect Module B3 with SN: 861633 with new Module SN: 602619 Generate PRE and POST IQR, Assay Stat and System Log | Replacement of Modules | Unit is fully operational | N/A | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534640286815882459/5534640286815882459_signature_28.png | Lea L. Buelis RMT RN, ASCPI | https://www.jotform.com/uploads/macarewecare/222927978372067/5534640286815882459/5534640286815882459_signature_31.png
| 2023-03-01 | 2023-03-01 14:49 | Toril Health Center | Poblacion, Toril Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23592 | PBSP | leabuelis@yahoo.com | 24.7 | ASSAY: Xpert MTB/RIF, COUNT: 50 ASSAY: Xpert HIV-1 Viral Load, COUNT: 978 | 2023-03-01 | GX-IV | 838865 | Module B2 and B4 for Replacement | Replaced the Defective Module B2 with SN: 861528 with new Module SN 607055 Replaced the Defect Module B3 with SN: 861633 with new Module SN: 602619 Generate PRE and POST IQR, Assay Stat and System Log | Replacement of Modules | Unit is fully operational | N/A | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534640126817943318/5534640126817943318_signature_28.png | Lea L. Buelis RMT RN, ASCPI | https://www.jotform.com/uploads/macarewecare/222927978372067/5534640126817943318/5534640126817943318_signature_31.png
| 2023-03-01 | 2023-03-01 14:06 | Tarlac Provincial Hospital | Hospital Drive San Vicente Tarlac City | 3 | Tarlac | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23591 | PBSP | molecularlab.tph@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 663 | 2023-03-01 | GXIV-D | 805108 | C2 for module replacement | Performed surface cleaning Generate pre-iqr Performed module replacement on module C2 Generate pre-iqr Archived test result Gathered reports Performed functionality test | Module replacement Preventive maintenanca | Unit is operational except C3 |  | 1 Genexpert Module | 39293 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534634072026461649/5534634072026461649_signature_28.png | Camille Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5534634072026461649/5534634072026461649_signature_31.png
| 2023-03-01 | 2023-03-01 14:35 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23590 | PBSP | laboratory@loveyourself.ph | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 |  | GXIV | 808649 | Unit for calibration | #NAME? | Performed XpertCheck | Unit is operational. B1 failed, B2, B3, B4 insufficient cartridge. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534620122255901548/5534620122255901548_signature_28.png | Happy Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5534620122255901548/5534620122255901548_signature_31.png
| 2023-03-01 | 2023-03-01 14:03 | Bantayan Rural Health Unit | Brgy. Poblacion, bantayan | 7 | Cebu | Laboratory | YES | CALIBRATION MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23589 | DOH | lailane.gdiono@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 492 ASSAY: Xpert BCR-ABL Ultra, COUNT: 503 |  | GXIV | 838509 | Defective Modules 1 and 4 | Run xpertcheck module 1 & 4 | Run xpertcheck module 1 & 4 | Module 1 and 4 Service Require during xpertcheck calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534602978816585877/5534602978816585877_signature_28.png | Lailane Dionio | https://www.jotform.com/uploads/macarewecare/222927978372067/5534602978816585877/5534602978816585877_signature_31.png
| 2023-03-01 | 2023-03-01 10:45 | URDANETA DISTRICT HOSPITAL | Urdaneta Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23588 | PBSP | warlyacosta@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF, COUNT: 849 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1497 | 2023-02-24 | GXIV-D | 805109 | B3 for module replacement | Performed surface cleaning Generate Pre-IQR Performed module replacement Generate Post-IQR Performe preventive maintenance Gathered report files Performed functionality test | Mosule replacement Preventive maintenance | Unit is operational except B2 |  | 1 Genexpert Module | 39295 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534600412022056248/5534600412022056248_signature_28.png | Warlito Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5534600412022056248/5534600412022056248_signature_31.png
| 2023-03-01 | 2023-03-01 11:36 | Dr. Emigdio C. Cruz Sr. Mermorial Hospital | Arayat, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23587 | PBSP | labeccmh@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 0 | 2023-03-01 | Genexpert IV | 812790 | N/a | N/a | Performed Maintenance Gathered Reports Installed New PC Pull out back up PC Installed MTB ULTRA assay Installed Driver for Printer Performed Functionality Test | Unit is Operational | N/a | Genexpert PC - CPH0322F32851 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534522050811341538/5534522050811341538_signature_28.png | Leilani Marie Apostol | https://www.jotform.com/uploads/macarewecare/222927978372067/5534522050811341538/5534522050811341538_signature_31.png
| 2023-03-01 | 2023-03-01 11:04 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23586 | DOH | luzvimindanavidad2@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF, COUNT: 326 ASSAY: Xpert MTB/RIF Ultra, COUNT: 477 |  | GXIV-4D | 838297 | None | None | #NAME? | All four modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534495525529961161/5534495525529961161_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5534495525529961161/5534495525529961161_signature_31.png
| 2023-03-01 | 2023-03-01 11:04 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23585 | DOH | luzvimindanavidad2@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF, COUNT: 326 ASSAY: Xpert MTB/RIF Ultra, COUNT: 477 |  | GXIV-4D | 838297 | None | None | #NAME? | All four modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534495345526880753/5534495345526880753_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5534495345526880753/5534495345526880753_signature_31.png
| 2023-03-01 | 2023-03-01 11:04 | Palompon Rural Health Unit | Palompon, Leyte | 8 | Leyte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23584 | DOH | luzvimindanavidad2@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF, COUNT: 326 ASSAY: Xpert MTB/RIF Ultra, COUNT: 477 |  | GXIV-4D | 838297 | None | None | #NAME? | All four modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5534495185521903906/5534495185521903906_signature_28.png | Luzviminda A. Navidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5534495185521903906/5534495185521903906_signature_31.png
| 2023-02-28 | 2023-02-28 15:26 | Balagtas Rural Health Unit I | Balagtas, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23583 | WHO | april.joyce04@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 114 ASSAY: Xpert MTB/RIF Ultra, COUNT: 351 | 2023-02-28 | GXIV | 846037 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. UPS is beeping non stop. UPS is recommended for replacement/repair. All modules passed xpertcheck. |  | Xpertcheck Kit | 9653 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533979092224056401/5533979092224056401_signature_28.png | Joyce Arela | https://www.jotform.com/uploads/macarewecare/222927978372067/5533979092224056401/5533979092224056401_signature_31.png
| 2023-02-28 | 2023-02-28 12:53 | Norzagaray Rural Health Unit | Norzagaray, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23583 | PBSP | chelseajempolicarpio@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 197 ASSAY: Xpert MTB/RIF Ultra, COUNT: 667 | 2023-02-28 | GXIV | 838167 | Module B3 for replacement | - checked & cleaned cartridge bay, plunger rod, filter - replaced module B3 by 708483 - Generated reports - Performed reliability test | Replaced module B3 | Unit is operational. |  | Genexpert module | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533979092224142720/5533979092224142720_signature_28.png | Rowena Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5533979092224142720/5533979092224142720_signature_31.png
| 2023-02-28 | 2023-02-28 15:26 | Calumpang Rural Health Unit | Brgy. Calumpang | 12 | South Cotabato | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23582 | DOH | jan19.cam@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF, COUNT: 566 ASSAY: Xpert MTB/RIF Ultra, COUNT: 976 |  | Gx-IV | 838504 | Module B4 H/W failed due to temperature error | Perform Genexpert machine Preventive maintenance Observe module B4 during actual sample running for module Troubleshooting | Perform Genexpert machine Preventive maintenance Observe module B4 during actual sample running for module Troubleshooting | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533831159415118259/5533831159415118259_signature_28.png | Janice R. Camiña | https://www.jotform.com/uploads/macarewecare/222927978372067/5533831159415118259/5533831159415118259_signature_31.png
| 2023-02-28 | 2023-02-28 15:08 | Guimba Community Hospital | Guimba, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23581 | PBSP | laboratorygch@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 186 ASSAY: Xpert MTB/RIF Ultra, COUNT: 768 | 2023-02-28 | Genexpert IV | 807560 | Printer sometimes not working Barcode scanner sometimes not working | Troubleshoot the printer and bar code scanner | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational Printer and barcode scanner for replacement | Module Replacement - 629882 - 868986 | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533785100816259032/5533785100816259032_signature_28.png | Andrelyn Marquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5533785100816259032/5533785100816259032_signature_31.png
| 2023-02-28 | 2023-02-28 15:08 | Guimba Community Hospital | Guimba, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23580 | PBSP | laboratorygch@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 186 ASSAY: Xpert MTB/RIF Ultra, COUNT: 768 | 2023-02-28 | Genexpert IV | 807560 | Printer sometimes not working Barcode scanner sometimes not working | Troubleshoot the printer and bar code scanner | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational Printer and barcode scanner for replacement | Module Replacement - 629882 - 868986 | Genexpert Module | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533784930813440691/5533784930813440691_signature_28.png | Andrelyn Marquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5533784930813440691/5533784930813440691_signature_31.png
| 2023-02-28 | 2023-02-28 14:01 | Himamaylan City Health Office | Himamaylan City, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 23579 | PBSP | zhu_rin04@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 576 ASSAY: Xpert MTB/RIF Ultra, COUNT: 844 |  | GX-IV | 838138 | None | Xpertcheck run | Module bay cleaning, Plunger rod cleaning, Filter cleaning, Xpertcheck run | 4 modulea operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533738499918280565/5533738499918280565_signature_28.png | Mary Cecil C. Jandumon | https://www.jotform.com/uploads/macarewecare/222927978372067/5533738499918280565/5533738499918280565_signature_31.png
| 2023-02-28 | 2023-02-28 13:11 | Central Luzon Doctor's Hospital | Hospital Drive San Vicente Tarlac City | 3 | Tarlac | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23578 | COMMERCIAL | rohmhardameg@gmail.com | 22.9 | ASSAY: Xpert MTB/RIF, COUNT: 22 ASSAY: Xpert SARS-CoV-2, COUNT: 361 | 2023-02-27 | GXIV-D | 830841 | A1 For moduoe replacement | Performed surface cleaning Generate Pre-IQR Performed module replacement Generate Post-IQR Gathered reports Archived test results Performed functionality test | Module replacement Preventive Mwintenance | Unit is operational |  | Genexpert module | 9654 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533716254211838982/5533716254211838982_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/5533716254211838982/5533716254211838982_signature_31.png
| 2023-02-28 | 2023-02-28 13:07 | Biliran Provincial Hospital | Naval, Biliran | 8 | Biliran | TB Dots | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 23577 | PBSP | ghelmission13@yahoo.com | 24.3 | ASSAY: Xpert MTB/RIF, COUNT: 701 ASSAY: Xpert MTB/RIF Ultra, COUNT: 565 | 2023-02-28 | GXIV-4D | 815273 | #NAME? | -Replace defective module B2/805711 with new module SN: 715981 | #NAME? | Module B1 - for reporting. Module B3 - waiting for module allocation. Module B2 & B4 are functional. |  | 900-0491RS | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533705442758115369/5533705442758115369_signature_28.png | Gheland Corto | https://www.jotform.com/uploads/macarewecare/222927978372067/5533705442758115369/5533705442758115369_signature_31.png
| 2023-02-28 | 2023-02-28 11:34 | Tarlac Provincial Hospital | Hospital Drive San Vicente Tarlac City | 3 | Tarlac | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23576 | PBSP | molecularlab.tph@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 648 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 |  | GXIV-D | 805108 | Additional Troubleshooting procedure on module C3 error 4008 | Performed cepheid troubleshooting procedures Checked F1 fuse if opened and 0V detected Performed module swap, ( C3 to C4 vice versa) Performed functionality test | Troubleshooting | Unit is operatial except module C2 | Module C4:772914 HW Failed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533703544213250355/5533703544213250355_signature_28.png | Camille Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5533703544213250355/5533703544213250355_signature_31.png
| 2023-02-28 | 2023-02-28 11:57 | Tagbina Rural Health Unit | Brgy. Poblacion, Tagbina | 13 | Surigao del Sur | RTDL | NO | INSTALLATION | CEPHEID | Faizar Avenido |  |  |  |  | 23575 | PBSP | ellaquimalbarracin@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 65 ASSAY: Xpert MTB/RIF Ultra, COUNT: 430 |  | GX-4 | 838877 | None | None | Replacing of Temporary Monitor to Old Cepheid Monitor | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533687112915571217/5533687112915571217_signature_28.png | Ellaquim L. Albarracin | https://www.jotform.com/uploads/macarewecare/222927978372067/5533687112915571217/5533687112915571217_signature_31.png
| 2023-02-28 | 2023-02-28 11:20 | Padada Rural Health | Poblacion, Padada Davao Del Sur | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23574 | DOH | zhakie.marvz@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 238 | 2023-02-28 | GX-IV 4D | 838511 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Unit is fully operational | N/A | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533645097311863694/5533645097311863694_signature_28.png | Marvielyn Lusica | https://www.jotform.com/uploads/macarewecare/222927978372067/5533645097311863694/5533645097311863694_signature_31.png
| 2023-02-28 | 2023-02-28 11:20 | Padada Rural Health | Poblacion, Padada Davao Del Sur | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23573 | DOH | zhakie.marvz@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 238 | 2023-02-28 | GX-IV 4D | 838511 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Perform Preventive Maintenance | Unit is fully operational | N/A | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533645057314294262/5533645057314294262_signature_28.png | Marvielyn Lusica | https://www.jotform.com/uploads/macarewecare/222927978372067/5533645057314294262/5533645057314294262_signature_31.png
| 2023-02-28 | 2023-02-28 11:04 | Bogo City Health Office | Bogo City, Mountain Breeze | 7 | Cebu | Tb Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 23572 | PBSP | bogocholab@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 947 | 2023-02-28 | Gx4 | 808630 | Error 4009, 1001 and 1002 | Module replacement of new Module 614200 | Module replacement of Module 634956 | All modules are operational |  | Module serial number 614200 | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533636399317554389/5533636399317554389_signature_28.png | Reggie Dahili | https://www.jotform.com/uploads/macarewecare/222927978372067/5533636399317554389/5533636399317554389_signature_31.png
| 2023-02-27 | 2023-02-27 20:34 | Chong Hua Hospital | Fuente Osmena, Capitol Site | 7 | Cebu | Bloodbank | YES | REPAIR | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 23571 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 22 |  | 2023-02-27 | Mcs+ | 02F035 | Bowl Leaked | Repair and Cleaned bowl area | Repair and Cleaned bowl area | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5533107495952456299/5533107495952456299_signature_28.png | Mayleen Jakosalem | https://www.jotform.com/uploads/macarewecare/222927978372067/5533107495952456299/5533107495952456299_signature_31.png
| 2023-02-27 | 2023-02-27 14:24 | Sultan kudarat Provincial Hospital | Isulan | 12 | Sultan Kudarat | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23570 | PBSP | charmfrias1270@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF, COUNT: 271 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2521 | 2023-02-27 | Gx-IV | 807405 | Module B4 Service Required | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | 4 operational modules |  | Xpertcheck -CE-5 | Lot 1000378669 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532933110515158994/5532933110515158994_signature_28.png | Charma Grace Frias | https://www.jotform.com/uploads/macarewecare/222927978372067/5532933110515158994/5532933110515158994_signature_31.png
| 2023-02-27 | 2023-02-27 13:30 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 23569 | DOH | cauayanmho@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 519 ASSAY: Xpert MTB/RIF Ultra, COUNT: 717 |  | GX-IV | 838477 | None | Preventive Maintenance | Maintenance bay cleaning, Plunger rod maintenace, Filter cleaning | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532863358027901515/5532863358027901515_signature_28.png | Jose Ali C. Treyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5532863358027901515/5532863358027901515_signature_31.png
| 2023-02-27 | 2023-02-27 13:30 | Cauayan Municipal Health Office | Cauayan, Negros Occidental | 6 | Negros Occidental | TB Laboratory | YES | MAINTENANCE | CEPHEID | Gerald Fritz Selga |  |  |  |  | 23568 | DOH | cauayanmho@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 519 ASSAY: Xpert MTB/RIF Ultra, COUNT: 717 |  | GX-IV | 838477 | None | Preventive Maintenance | Maintenance bay cleaning, Plunger rod maintenace, Filter cleaning | 4 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532863268027510472/5532863268027510472_signature_28.png | Jose Ali C. Treyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5532863268027510472/5532863268027510472_signature_31.png
| 2023-02-27 | 2023-02-21 13:42 | Jose Reyes Memorial Medical Center | R. Ave.,Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23564 | DOH | covidlabjrrmmc@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 629 ASSAY: Xpert MTB/RIF Ultra, COUNT: 35 ASSAY: Xpert SARS-CoV-2, COUNT: 1365 | 2023-01-16 | GXIV- D | 831530 | Unit is due for Xpertcheck | Xpert Check | Performed Maintenance Performed xpertcheck Generate Reports | All modules passed. |  | Xpertcheck kit | DR# 9466 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532766811213979942/5532766811213979942_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5532766811213979942/5532766811213979942_signature_31.png
| 2023-02-27 | 2023-02-23 11:11 | Bulihan Rural Health Unit | Bulihan, Silang | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23566 | PBSP | alei22775@yahoo.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 657 |  | GXIV-D | 838446 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 1 module for replacement, 3 modules are operational | XC: All modules passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532766811216688996/5532766811216688996_signature_28.png | Alejandra Alegre | https://www.jotform.com/uploads/macarewecare/222927978372067/5532766811216688996/5532766811216688996_signature_31.png
| 2023-02-27 | 2023-02-21 13:35 | Jose Reyes Memorial Medical Center | Rizal Ave.,Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23567 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 923 ASSAY: Xpert SARS-CoV-2, COUNT: 3373 | 2023-02-20 | GX16 - D - 10-C | 1100130856 | Temp problem in module A4 | Module Replacement | Replaced module 4 Performed Maintenance Generate Reports | 12 modules are operational, 4 modules for replacement | 210069235→ 210078094 | GX Module kit 10-c | DR# 9641 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532766811217912030/5532766811217912030_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5532766811217912030/5532766811217912030_signature_31.png
| 2023-02-27 | 2023-02-22 13:20 | Rosario Rural Health Unit | Rosario | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23565 | DOH | arlenepenaarmijo@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 958 |  | GXIV-D | 838537 | Unit has temp. errors on module 1 | Xpertcheck Maintenance | Performed Maintenance Performed Xpertcheck on reported module Generate Reports | 1 module operational, 3 modules for replacement. Module 1 passed but H/W failed in GX software. | Used Spare Cartridge |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532766811219815375/5532766811219815375_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/5532766811219815375/5532766811219815375_signature_31.png
| 2023-02-27 | 2023-02-24 09:11 | Ilocos Training and Regional Medical Center | San Fernando City | 1 | La Union | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23563 | PBSP | itrmc.tblaboratory@gmail.com | 24.5 | ASSAY: Xpert HPV, COUNT: 672 ASSAY: Xpert MTB/RIF, COUNT: 725 ASSAY: Xpert MTB/RIF Ultra, COUNT: 712 ASSAY: Others, COUNT: 572 ASSAY: Xpert SARS-CoV-2, COUNT: 282 | 2023-02-09 | GXIV-D | 848249 | Module replacement for modules B1:210002714 B3:210002718 | Performed surface cleaning Generate Pre IQR Performed module replacement Generate Post IQR Performed preventive maintenance Generate reports Archived test results Performed functionality test | Module replacement Preventive maintenance | Unit is operational except B4 | New modules B1:210078192 and B3:210078088 | 2 Genexpert modules | 38939 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532753881219710602/5532753881219710602_signature_28.png | Marie Jennifer Culo | https://www.jotform.com/uploads/macarewecare/222927978372067/5532753881219710602/5532753881219710602_signature_31.png
| 2023-02-27 | 2023-02-23 13:58 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23561 | PBSP | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 213 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1556 ASSAY: Xpert HIV-1 Viral Load, COUNT: 158 |  | GX-IV-D | 804254 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore Generated reports | Modules D1 and D4 operational. | Modules D2/690385 and D3/709295 not operational. Reported already. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801211807241/5532676801211807241_signature_28.png | Amelia C. Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801211807241/5532676801211807241_signature_31.png
| 2023-02-27 | 2023-02-23 14:04 | Marikina City Health Office | Shoe Ave | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23562 | DOH | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 286 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1349 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 | 2023-02-07 | GX-IV-D | 831888 | For UPS replacement | Replaced UPS | Performed maintenance Replaced UPS Generated reports | Module A1 and A4 operational. Module A2 operational but with existing report. Module A3 not operational. |  | Accupower 1kva UPS | DR# 9600 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801214391868/5532676801214391868_signature_28.png | Amelia C. Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801214391868/5532676801214391868_signature_31.png
| 2023-02-27 | 2023-02-23 10:22 | Cainta City Health Office | Cainta | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23558 | DOH | rowenatec@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 153 ASSAY: Xpert MTB/RIF Ultra, COUNT: 618 ASSAY: Xpert HIV-1 Viral Load, COUNT: 52 |  | GX-IV-D | 831623 | Monitor shutting down | Performed maintenance | Performed maintenance Checked monitor Generated reports | Modules A1, A2 and A4 operational. Recommend to replace monitor. | Module A3/754520 not operational. Reported already. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801214415734/5532676801214415734_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801214415734/5532676801214415734_signature_31.png
| 2023-02-27 | 2023-02-23 10:37 | Cainta City Health Office | Cainta | 4A | Rizal | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23560 | PBSP | rowenatec@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 163 ASSAY: Xpert MTB/RIF Ultra, COUNT: 580 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 |  | GX-IV-D | 838182 | Communication erron on module B2 | Check internal connections | Performed maintenance Checked modules connection Checked gateway board Generated reports | Modules B1, B3 and B4 operational. | Module B2 not operational. For reporting. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801219588516/5532676801219588516_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801219588516/5532676801219588516_signature_31.png
| 2023-02-27 | 2023-02-23 10:32 | Cainta City Health Office | Cainta | 4A | Rizal | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23559 | PBSP | rowenatec@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 192 ASSAY: Xpert MTB/RIF Ultra, COUNT: 900 |  | GX-IV-D | 804687 | Stuck cartridge | Removed stuck cartridge | Performed maintenance Removed sticker cartridge Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801219970679/5532676801219970679_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676801219970679/5532676801219970679_signature_31.png
| 2023-02-27 | 2023-02-21 15:19 | The Medical City | Ortigas Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23556 | COMMERCIAL | moleculardx@themedicalcity.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 8 ASSAY: Xpert HBV VL, COUNT: 14 ASSAY: Xpert HPV, COUNT: 52 ASSAY: Xpert MTB/RIF, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 9 ASSAY: Xpert HIV-1 Viral Load, COUNT: 53 |  | GX-IV-L | 802117 | For maintenance | Performed maintenance | Performed maintenance Connect to network switch Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721211901293/5532676721211901293_signature_28.png | PAOLO | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721211901293/5532676721211901293_signature_31.png
| 2023-02-27 | 2023-02-22 11:20 | East Avenue Medical Center | East Avenue | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23557 | MACARE-HBDC | eamcmolecularlab@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 304 ASSAY: Xpert SARS-CoV-2, COUNT: 1087 ASSAY: Others, COUNT: 2 | 2023-02-21 | GX-IV-D | 110014085 | Module B4/210069598 for replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | All modules operational. |  | Old B4/210069598 New B4/210077551 | DR# 9645 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721213751071/5532676721213751071_signature_28.png | Ann Gylene Sisante | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721213751071/5532676721213751071_signature_31.png
| 2023-02-27 | 2023-02-21 15:26 | The Medical City | Ortigas Avenue | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23553 | COMMERCIAL | moleculardx@themedicalcity.com | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 5 ASSAY: Others, COUNT: 29 ASSAY: Others, COUNT: 16 ASSAY: Xpert CT/NG, COUNT: 8 ASSAY: Others, COUNT: 4 ASSAY: Xpert HBV VL, COUNT: 201 ASSAY: Xpert HPV, COUNT: 233 ASSAY: Xpert MTB/RIF, COUNT: 144 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2049 ASSAY: Xpert SARS-CoV-2, COUNT: 7 ASSAY: Xpert HCV VL, COUNT: 53 |  | GX-IV-L | 802630 | Laptop charger not working | Checked charger | Performed maintenance Checked charger Checked battery Connect to network switch Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721214022001/5532676721214022001_signature_28.png | PAOLO | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721214022001/5532676721214022001_signature_31.png
| 2023-02-27 | 2023-02-22 14:23 | National Children's Hospital | E. Rodriguez | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23555 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2064 ASSAY: Xpert SARS-CoV-2, COUNT: 6224 ASSAY: Others, COUNT: 18 | 2023-02-21 | GX-XVI-D | 110014171 | Modules A3 and D2 for replacement | Replaced modules | Performed maintenance Replaced defective modules Generated reports | Unit is operational. Module B1/210069274 not operational but already reported. |  | Old A3/210069033, D2/210069280 New A3/210077159, D2/210078182 | DR# 9646 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721216529275/5532676721216529275_signature_28.png | Ice Catalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721216529275/5532676721216529275_signature_31.png
| 2023-02-27 | 2023-02-22 16:39 | National Children's Hospital | E. Rodriguez | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23554 | PBSP | nchmolecularlab2021@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 497 ASSAY: Xpert MTB/RIF Ultra, COUNT: 166 ASSAY: Xpert SARS-CoV-2, COUNT: 2327 |  | GX-IV-D | 803109 | For maintenance | Performed maintenance | Performed maintenance Cleaned plunger rod, icore, cartridge bay Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721218031649/5532676721218031649_signature_28.png | Ice Catalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676721218031649/5532676721218031649_signature_31.png
| 2023-02-27 | 2023-02-25 08:32 | Tondo Medical Center | Tondo, Manila | NCR | NCR 1st - Manila | COVID Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23551 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 43 ASSAY: Others, COUNT: 423 ASSAY: Xpert MTB/RIF, COUNT: 461 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1041 ASSAY: Xpert SARS-CoV-2, COUNT: 11066 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2023-02-25 | GX XVI | 844121 | For module replacement | #NAME? | Replaced modules A1, A2, A4, B2, B3, C4 | Unit is operational. Module B4 for reporting |  | Genexpert module SN: 210057331, 210052269, 220003328, 220004106, 608884, 618126 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676381211728262/5532676381211728262_signature_28.png | Kyle Emmanuel Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676381211728262/5532676381211728262_signature_31.png
| 2023-02-27 | 2023-02-25 11:51 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 23552 | PBSP | juliecaladormt@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 960 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1276 ASSAY: Xpert SARS-CoV-2, COUNT: 5226 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 | 2023-02-25 | GX XVI | 839292 | For module replacement | #NAME? | Replaced modules | Unit is operational. |  | GeneXpert Module SN: 210019248, 210051875, 210051368, 210045352, 210033753, 210029448, 210020537, 210020134 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676381218563448/5532676381218563448_signature_28.png | Julie Anne Calado | https://www.jotform.com/uploads/macarewecare/222927978372067/5532676381218563448/5532676381218563448_signature_31.png
| 2023-02-27 | 2023-02-23 08:52 | Talisay Rural Health Unit | Talisay Batangas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23550 | DOH | medcertrhutalisay@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 371 | 2023-01-17 | GXIV | 820201 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules A1 A3 and A4 passed Xpertcheck. |  | Xpertcheck | DR 9484 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532675741214080814/5532675741214080814_signature_28.png | Adettlou A. Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5532675741214080814/5532675741214080814_signature_31.png
| 2023-02-27 | 2023-02-23 13:08 | Tanauan Rural Health Unit | Tanauan Batangas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23550 | DOH | ramilomher27@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 164 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1425 | 2023-01-17 | GXIV | 821299 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. |  | Xpertcheck | DR 9488 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5532675741218702192/5532675741218702192_signature_28.png | Edita Ramilo | https://www.jotform.com/uploads/macarewecare/222927978372067/5532675741218702192/5532675741218702192_signature_31.png
| 2023-02-26 | 2023-01-25 09:52 | Mataas na Kahoy Rural Health Unit | Brgy II, Mataas na Kahoy | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23549 | PBSP | dmlescano1219@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 282 | 2023-01-23 | GXIV | 838325 | Unit for preventive maintenance. Unit for Xpertcheck. For module replacement of module B3. | Performed preventive maintenance. Performed module replacement on B3. Performed Xpertcheck. Generated reports. | Preventive maintenance, Xpertcheck and Module replacement. | Unit is operational. |  | GX module for B3(210068186) in replacement to defective module B3(860135). | G5112 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504283053963764095/5504283053963764095_signature_28.png | Debbie Lescano | https://www.jotform.com/uploads/macarewecare/222927978372067/5504283053963764095/5504283053963764095_signature_31.png
| 2023-02-26 | 2023-01-25 09:52 | Mataas na Kahoy Rural Health Unit | Brgy II, Mataas na Kahoy | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23548 | PBSP | dmlescano1219@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 282 | 2023-01-23 | GXIV | 838325 | Unit for preventive maintenance. Unit for Xpertcheck. For module replacement of module B3. | Performed preventive maintenance. Performed module replacement on B3. Performed Xpertcheck. Generated reports. | Preventive maintenance, Xpertcheck and Module replacement. | Unit is operational. |  | GX module for B3(210068186) in replacement to defective module B3(860135). | G5112 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504283053963764095/5504283053963764095_signature_28.png | Debbie Lescano | https://www.jotform.com/uploads/macarewecare/222927978372067/5504283053963764095/5504283053963764095_signature_31.png
| 2023-02-25 | 2023-02-25 16:51 | Chong Hua Hospital | Fuente Capitol Site, Cebu City | 7 | Cebu | Bloodbank | YES | REPAIR | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 23547 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  | 2023-02-25 | Mcs+ | 02F035 | Wheel of the stand damage | Repair wheel and align | Repair wheel and set up alignment | Operational |  | Bolts and nut with washer | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5531246637885023691/5531246637885023691_signature_28.png | Arianne Mae Guerrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5531246637885023691/5531246637885023691_signature_31.png
| 2023-02-24 | 2022-11-10 19:47 | Lung center of the Philippines | Quezon ave.diliman qc | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 23545 | MACARE-HBDC | lcppathology@gmail.com | 26 |  |  | MCS + | 02D166 | For maintenance | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,,performed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5530506992774706530/5530506992774706530_signature_28.png | Rhona Babilonai | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5530506992774706530/5530506992774706530_signature_31.png
| 2023-02-24 | 2023-02-22 13:54 | Bicol medical center | Bmc road,naga city | 5 | Camarines Sur | Blood donation | NO | INSTALLATION | HAEMONETICS | Noel Floresca |  |  |  |  | 23546 | MACARE-HBDC | bmc.bloodbank@gmail.com | 26 |  |  | MCS + | 11J221 | For installation | Performed installation,performed functionality test | Installation, | Operational ready to use | Unit calibrated 2/20/2023 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5530506992778109803/5530506992778109803_signature_28.png | Ma. Charity P de los Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5530506992778109803/5530506992778109803_signature_31.png
| 2023-02-23 | 2023-02-23 15:39 | Mexico RHU I | Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23544 | DOH | jasperdizon02@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 157 ASSAY: Xpert MTB/RIF Ultra, COUNT: 268 | 2023-02-03 | Genexpert IV | 831877 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck - All Modules Passed | Xpertcheck 3pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529728912024215963/5529728912024215963_signature_28.png | JASPER P. DIZON | https://www.jotform.com/uploads/macarewecare/222927978372067/5529728912024215963/5529728912024215963_signature_31.png
| 2023-02-23 | 2023-02-23 16:36 | Antipas Rural Health Unit | Poblacion Antipas | 12 | Cotabato | TB-DOTS | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23543 | DOH | rhuantipas@gmail.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 442 | 2023-02-23 | GX-IV | 838259 | Module B1 encounter multiple Temperature Error For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration at Module B1 only | Perform Preventive Maintenance and XpertCheck Calibration for module B1 only | Unit has 3 operational Modules Module B1 for service required | N/A | Xpertcheck CE cartridge | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529528279021212151/5529528279021212151_signature_28.png | Alvin A. Codilla, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5529528279021212151/5529528279021212151_signature_31.png
| 2023-02-23 | 2023-02-23 15:35 | Baguio Health Services Office | T.Alonzo Baguio City | CAR | Benguet | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23542 | DOH | sherryannemarrero@gmail.com | 18.9 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF, COUNT: 76 ASSAY: Xpert MTB/RIF Ultra, COUNT: 876 |  | GXIV-D | 838281 | Temperature error modules B1 and B2 | Performed surface cleaning Cleaning of plunger rod,icore,cartridge bay and filter Archived test results Calibrate B1 and B2 Gathered reports Performed functionality test | Performed preventive maintenance Performed calibration | Unit is operational except | 2 xpertcheck kits used |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529488294013627017/5529488294013627017_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5529488294013627017/5529488294013627017_signature_31.png
| 2023-02-23 | 2023-02-23 14:56 | Macare In- House | Aurora Grand, Quezon City | NCR | NCR 2nd - Quezon City | Accounting | NO | REPAIR | OTHERS | Ronnel Belizar |  |  |  | UPS | 23541 | COMMERCIAL | adminassistant@macare-medicals.com | 25 |  | 2023-02-23 | APC BX625CI - MS | 3B1717X11292 | Battery for replacement | #NAME? | Replaced battery | Unit is operational. |  | Battery SN: 459864 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529454471212204291/5529454471212204291_signature_28.png | Renelyn Laurel | https://www.jotform.com/uploads/macarewecare/222927978372067/5529454471212204291/5529454471212204291_signature_31.png
| 2023-02-23 | 2023-02-23 11:15 | Baguio General Hospital and Medical Center | Gov. Pack Rd. Baguio City | CAR | Benguet | Molecular | NO | INSTALLATION | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23540 | MACARE-HBDC | molecularlab@bghmc.doh.gov.ph | 22.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 0 | 2023-02-21 | GXIV-D | 110016112 | Installation | Installation | Installation and setup of Genexpert Setup Genexpert users Installation of printers Gathered reports Performed functionality test | Unit is operational |  | Genexpert Machine set | 9643 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529444114017219802/5529444114017219802_signature_28.png | April Joy Gundran | https://www.jotform.com/uploads/macarewecare/222927978372067/5529444114017219802/5529444114017219802_signature_31.png
| 2023-02-23 | 2023-02-23 13:57 | Labuan General Hospital | Labuan, Zamboanga City, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23539 | DOH | irenecasepe@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1345 | 2023-02-23 | GX 4 | 838568 | Module 1 H/W failed error 4012 | Conducted Self test error, checked the room and modules temperatures, inspected if an heat source is available. Relocated the machine to a different building. Preventive Maintenance | Conducted Self test error, checked the room and modules temperatures, inspected if an heat source is available. Relocated the machine to a different building. Preventive Maintenance | Module 1 keeps on H/W failed. Module 2, 3 and 4 are available. | Module 1 wich is currently h/w failed, After conducting self test error becomes available again | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529418568222943562/5529418568222943562_signature_28.png | Irene A. Casepe | https://www.jotform.com/uploads/macarewecare/222927978372067/5529418568222943562/5529418568222943562_signature_31.png
| 2023-02-23 | 2023-02-23 13:28 | Legazpi City Health Office | Legazpi City, Albay | 5 | Albay | TB Dots | YES | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23538 | WHO | janzchel_neneh10@yahoo.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 303 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1178 |  | GXIV-4D | 846354 | None | #NAME? | NA | Module B2 is experiencing H/W failure, waiting for module allocation. Three other modules are operational. | Currently GX machine is not functional due to non-functional UPS. Waiting for UPS replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529396943427364980/5529396943427364980_signature_28.png | Janzhel Marie Chan | https://www.jotform.com/uploads/macarewecare/222927978372067/5529396943427364980/5529396943427364980_signature_31.png
| 2023-02-23 | 2023-02-23 11:25 | Bagumbayan Rural Health Unit | Bagumbayan | 12 | Sultan Kudarat | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23537 | DOH | marizalenoid@yahoo.com | 26.8 | ASSAY: Xpert MTB/RIF, COUNT: 175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1166 |  | Gx-IV | 838261 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529324557429992203/5529324557429992203_signature_28.png | Ma. Riza R. Dionela | https://www.jotform.com/uploads/macarewecare/222927978372067/5529324557429992203/5529324557429992203_signature_31.png
| 2023-02-23 | 2023-02-21 15:13 | Bustos Rural Health Unit | Bustos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23535 | DOH | ronald_galvez2002@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 165 ASSAY: Xpert MTB/RIF Ultra, COUNT: 419 | 2023-01-19 | GXIV | 831874 | Unit for xpertcheck | #NAME? | Performed Xpertcheck except on B4 since it is already reported for replacement. | Unit is operational. B1 & B3 passed calibration. B2 failed calibration. B2 for reporting. |  | XpertCheck kit | 9515 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529306331211850038/5529306331211850038_signature_28.png | Ronald Galvez | https://www.jotform.com/uploads/macarewecare/222927978372067/5529306331211850038/5529306331211850038_signature_31.png
| 2023-02-23 | 2023-02-21 13:43 | Baliuag District Hospital | Baliuag, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23534 | PBSP | reyeshannahxx@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 116 ASSAY: Xpert MTB/RIF Ultra, COUNT: 282 ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 |  | GXIV | 812770 | Unit for calibration | #NAME? | Performed Xpertcheck | Unit is operational. B1, B2, B4 passed calibration. B3 failed calibration. B3 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529306331212494026/5529306331212494026_signature_28.png | Hannah Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5529306331212494026/5529306331212494026_signature_31.png
| 2023-02-23 | 2023-02-22 13:04 | Mandaluyong City Health Office | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23536 | DOH | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 190 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1916 ASSAY: Xpert HIV-1 Viral Load, COUNT: 100 | 2023-01-20 | GXIV | 831596 | Unit for calibration | #NAME? | performed Xpertcheck except C2 since it is already reported for replacement. | Unit is operational. C3 passed calibration. C1, C4 failed calibration. |  | Xpert Check kit | 9536 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529306331218569529/5529306331218569529_signature_28.png | Teresita Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5529306331218569529/5529306331218569529_signature_31.png
| 2023-02-23 | 2023-02-23 09:54 | Talisay District Hospital | Talisay City Poblacion | 7 | Cebu | Tb. Laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 23533 | DOH | josephlorenzomercado@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 600 | 2023-02-23 | GX 4 | 838559 | Error 2008 in Module 2 | Clean Plunger and Icore and monitor the module | Clean Module 2 | All modules are operational |  | Na | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529273347115536482/5529273347115536482_signature_28.png | Joseph Lorenzo Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/5529273347115536482/5529273347115536482_signature_31.png
| 2023-02-23 | 2023-02-23 09:35 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23532 | DOH | materesaaliman78@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 302 ASSAY: Xpert MTB/RIF Ultra, COUNT: 353 |  | GXIV | 838406 | For Preventive maintenance | #NAME? | Preventive maintenance | Module (B1,B2,B4) Operational, Module B3 disabled (Waiting for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529257320791224811/5529257320791224811_signature_28.png | Ma.Teresa E.Aliman | https://www.jotform.com/uploads/macarewecare/222927978372067/5529257320791224811/5529257320791224811_signature_31.png
| 2023-02-23 | 2023-02-23 07:58 | Riverside Medical Center, Inc. | B.S Aquino Drive, Bacolod City | 6 | Negros Occidental | Laboratory Department | NO | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 23531 | COMMERCIAL | laboratory@rivermedcenter.net | 21 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 41 ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 33 ASSAY: Xpert HCV VL, COUNT: 49 | 2023-02-23 | GX IV | 110005268 | They requested for module validation for newly replaced module for B2 as requirement for their Audit. | Module B2 passed the calibration. Save copy of GXX file and IQR for reporting. | Perform Xpert check calibration for module B2 with S/N 210070907. | All 4 modules are functional |  | Lot No. 05502 | 1 Xpert check |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5529215160809250926/5529215160809250926_signature_28.png | Nicholle Maravilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5529215160809250926/5529215160809250926_signature_31.png
| 2023-02-22 | 2023-02-22 16:00 | Apalit Rural Health Unit | Apalit Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23530 | DOH | duenasblesilda@gmail.com | 30.8 | ASSAY: Xpert MTB/RIF, COUNT: 204 ASSAY: Xpert MTB/RIF Ultra, COUNT: 471 ASSAY: Xpert HIV-1 Viral Load, COUNT: 10 | 2023-02-03 | GXIV-D | 831900 | Due for calibration | Performed surface cleaning Cleaning of plunger rod,cartridge bay,icore and filter Archived test results Performed calibration Gathered reports Performed functionality test | Preventive maintenance Calibration | Unit is operational except modules A1,A3 and A4 | A1 failed xpertcheck for reporting | 3 pcs Genexpert Xpertcheck kits | 9581 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528639169128567679/5528639169128567679_signature_28.png | Blesilda G. Dueñas | https://www.jotform.com/uploads/macarewecare/222927978372067/5528639169128567679/5528639169128567679_signature_31.png
| 2023-02-22 | 2023-02-22 14:22 | Macabebe Rural Health Unit | Macabebe Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23529 | DOH | salvamari13@yahoo.com | 25.68 | ASSAY: Xpert MTB/RIF, COUNT: 161 ASSAY: Xpert MTB/RIF Ultra, COUNT: 679 | 2023-01-16 | GXIV-D | 831893 | Error 4012 on module A1 Due for calibration | Performed surface cleaning Cleaning of icore,cartridge bay, plunger rod and filter Archived test results Performed calibration Gathered reports Performed functionality test | Preventive maintenance Calibration | Unit is operational except modules A2 and A3 | All modules passed calibration | 2 Xpertcheck kits | 9473 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528577919127048079/5528577919127048079_signature_28.png | Salva Sese | https://www.jotform.com/uploads/macarewecare/222927978372067/5528577919127048079/5528577919127048079_signature_31.png
| 2023-02-22 | 2023-02-22 14:30 | Jose B. Lingad Memorial Regional Hospital | San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23528 | COMMERCIAL | concepcionalfonso2@gmail.com | 29 | ASSAY: Xpert HBV VL, COUNT: 56 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 28 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1196 | 2023-02-22 | Genexpert IV | 812571 | Error 2096, 2097 | Performed Xpertcheck in module 1, 2, 3 | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck - 2, 3, 4 - Service Required | Xpertcheck 7pc | N/a |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5528572247263166244/5528572247263166244_signature_28.png | Concepcion Alfonso | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5528572247263166244/5528572247263166244_signature_31.png
| 2023-02-22 | 2023-02-22 13:23 | Tomas Claudio Health Center | Corner E. Quirino - Camus Street, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23527 | DOH | gaelmarellita@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 260 ASSAY: Xpert MTB/RIF Ultra, COUNT: 287 | 2023-02-22 | GX-IV | 838492 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses | Perfrom Preventive Maintenance | Unit is fully operational | Module B1 encounter temperature error, for under observation | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528531447811965220/5528531447811965220_signature_28.png | Gaelmare M. Rellita | https://www.jotform.com/uploads/macarewecare/222927978372067/5528531447811965220/5528531447811965220_signature_31.png
| 2023-02-22 | 2023-02-22 13:23 | Tomas Claudio Health Center | Corner E. Quirino - Camus Street, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23526 | DOH | gaelmarellita@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 260 ASSAY: Xpert MTB/RIF Ultra, COUNT: 287 | 2023-02-22 | GX-IV | 838492 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses | Perfrom Preventive Maintenance | Unit is fully operational | Module B1 encounter temperature error, for under observation | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528531427812956195/5528531427812956195_signature_28.png | Gaelmare M. Rellita | https://www.jotform.com/uploads/macarewecare/222927978372067/5528531427812956195/5528531427812956195_signature_31.png
| 2023-02-22 | 2023-02-22 11:48 | Bulusan Rural Health Unit | Bulusan, Sorsogon | 5 | Sorsogon | TB Dots | YES | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23525 | PBSP | arleneragasa@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF, COUNT: 74 ASSAY: Xpert MTB/RIF Ultra, COUNT: 727 |  | GXIV-2D | 835197 | None | #NAME? | NA | All two modules are operational. | Module B1 failed XpertCheck last January 30, 2023. Waiting for module allocation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528472384809646560/5528472384809646560_signature_28.png | Kristoffer Madrona | https://www.jotform.com/uploads/macarewecare/222927978372067/5528472384809646560/5528472384809646560_signature_31.png
| 2023-02-22 | 2023-02-22 11:01 | Binan City Health Office 2 | Binan Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23524 | DOH | eloicyparra@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 287 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1023 ASSAY: Xpert HIV-1 Viral Load, COUNT: 235 | 2023-02-20 | GXIV | 821278 | For installation of back-up monitor. For gathering of report on module A4. | Performed installation of back-up monitor. Performed self test and troubleshooting on module A4. Generated reports. | Installation and troubleshooting. | Unit is operational. For reporting of module A4 due to H/W Failure. |  | Back-up monitor HP P19A (CNK629119V) in replacement to defective monitor HP P19A (CNK831071J). | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528445826424502689/5528445826424502689_signature_28.png | Eloicy Parra | https://www.jotform.com/uploads/macarewecare/222927978372067/5528445826424502689/5528445826424502689_signature_31.png
| 2023-02-22 | 2023-02-22 10:34 | Kumalarang Rural Health Unit | Kumalarang, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23523 | PBSP | roseimelda@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 998 | 2023-02-22 | GX 4 | 815278 | System unit unable to start up. | Change CMOS battery and change boot start up option. | Change CMOS battery and change boot start up option. | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528431011859195935/5528431011859195935_signature_28.png | Rose Imelda P.Uy | https://www.jotform.com/uploads/macarewecare/222927978372067/5528431011859195935/5528431011859195935_signature_31.png
| 2023-02-22 | 2023-02-22 10:34 | Kumalarang Rural Health Unit | Kumalarang, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23522 | PBSP | roseimelda@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 998 | 2023-02-22 | GX 4 | 815278 | System unit unable to start up. | Change CMOS battery and change boot start up option. | Change CMOS battery and change boot start up option. | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528431011859195935/5528431011859195935_signature_28.png | Rose Imelda P.Uy | https://www.jotform.com/uploads/macarewecare/222927978372067/5528431011859195935/5528431011859195935_signature_31.png
| 2023-02-22 | 2023-02-22 10:34 | Kumalarang Rural Health Unit | Kumalarang, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23521 | PBSP | roseimelda@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 998 | 2023-02-22 | GX 4 | 815278 | System unit unable to start up. | Change CMOS battery and change boot start up option. | Change CMOS battery and change boot start up option. | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528431011859195935/5528431011859195935_signature_28.png | Rose Imelda P.Uy | https://www.jotform.com/uploads/macarewecare/222927978372067/5528431011859195935/5528431011859195935_signature_31.png
| 2023-02-22 | 2023-02-22 09:17 | St. Elizabeth Hospital Inc. | General Santos City | 12 | South Cotabato | Molecular laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23520 | COMMERCIAL | laboratory@sehi.ph | 19.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 40 ASSAY: Xpert SARS-CoV-2, COUNT: 6434 |  | Gx -IV | 831313 | Module B3 H/W failed | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528387709612809694/5528387709612809694_signature_28.png | Roselle Jave Dignadice | https://www.jotform.com/uploads/macarewecare/222927978372067/5528387709612809694/5528387709612809694_signature_31.png
| 2023-02-22 | 2023-02-22 08:54 | St. Elizabeth Hospital Inc. | General Santos City | 12 | South Cotabato | Molecular laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23519 | COMMERCIAL | laboratory@sehi.ph | 20.64 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 109 ASSAY: Xpert SARS-CoV-2, COUNT: 6169 |  | Gx - IV (10C) | 110003713 | Module B1 and B4 H/W failed | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528377289618410240/5528377289618410240_signature_28.png | Roselle Jave Dignadice | https://www.jotform.com/uploads/macarewecare/222927978372067/5528377289618410240/5528377289618410240_signature_31.png
| 2023-02-22 | 2023-02-22 08:10 | Sorsogon Medical Mission Group and Health Services Cooperative | Sorsogon City, Sorsogon | 5 | Sorsogon | TB Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23518 | PBSP | lotino.ana@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 696 | 2023-02-22 | GXIV-4D R1 | 707452 | Module B4/644792 defective | #NAME? | -Replace defective module SN: 644792 with new module SN: 210068733 | Module B3 is disabled. Three other modules are operational. |  | 900-0491 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528375458419466708/5528375458419466708_signature_28.png | Ana Liza Lotino | https://www.jotform.com/uploads/macarewecare/222927978372067/5528375458419466708/5528375458419466708_signature_31.png
| 2023-02-22 | 2023-02-22 08:27 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23517 | PBSP | microscan@tmc.doh.gov.ph | 22 | ASSAY: Xpert HPV, COUNT: 43 ASSAY: Xpert HPV, COUNT: 423 ASSAY: Xpert MTB/RIF, COUNT: 461 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1008 ASSAY: Xpert SARS-CoV-2, COUNT: 11072 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GX16 | 844121 | Unit for visit. For checking of status of GX16. | Performed on-site checking of GX16. Generated reports. | On-site checking. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528351898466346311/5528351898466346311_signature_28.png | Julie Ann D. Maniego | https://www.jotform.com/uploads/macarewecare/222927978372067/5528351898466346311/5528351898466346311_signature_31.png
| 2023-02-22 | 2023-02-22 08:27 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23516 | PBSP | microscan@tmc.doh.gov.ph | 22 | ASSAY: Xpert HPV, COUNT: 43 ASSAY: Xpert HPV, COUNT: 423 ASSAY: Xpert MTB/RIF, COUNT: 461 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1008 ASSAY: Xpert SARS-CoV-2, COUNT: 11072 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GX16 | 844121 | Unit for visit. For checking of status of GX16. | Performed on-site checking of GX16. Generated reports. | On-site checking. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528351868467361586/5528351868467361586_signature_28.png | Julie Ann D. Maniego | https://www.jotform.com/uploads/macarewecare/222927978372067/5528351868467361586/5528351868467361586_signature_31.png
| 2023-02-22 | 2023-02-22 08:27 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23515 | PBSP | microscan@tmc.doh.gov.ph | 22 | ASSAY: Xpert HPV, COUNT: 43 ASSAY: Xpert HPV, COUNT: 423 ASSAY: Xpert MTB/RIF, COUNT: 461 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1008 ASSAY: Xpert SARS-CoV-2, COUNT: 11072 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GX16 | 844121 | Unit for visit. For checking of status of GX16. | Performed on-site checking of GX16. Generated reports. | On-site checking. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5528351608466095513/5528351608466095513_signature_28.png | Julie Ann D. Maniego | https://www.jotform.com/uploads/macarewecare/222927978372067/5528351608466095513/5528351608466095513_signature_31.png
| 2023-02-21 | 2023-02-21 09:28 | E.rodriguez health center | Pajo st quzon city | NCR | NCR 2nd - Quezon City | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23514 | DOH | clairebaptista32786@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert HPV, COUNT: 162 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1148 | 2023-01-20 | Gx 4 | 831848 | Defective ups no power,for maintenance and xpertcheck | Performed maintenance,cleaned unit,cleaned filter,plunger rod,icore,catridge bay, gathered reports, replaced ups fuse,,checked power supply | Pms and xpertpert check | A1,A2,A3 failed xpert check A4 passed | For replacement of ups | Xpertcheck | 9525 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527850552772770732/5527850552772770732_signature_28.png | Mary Claire P Baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/5527850552772770732/5527850552772770732_signature_31.png
| 2023-02-21 | 2023-02-21 09:28 | E.rodriguez health center | Pajo st quzon city | NCR | NCR 2nd - Quezon City | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23513 | DOH | clairebaptista32786@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert HPV, COUNT: 162 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1148 | 2023-01-20 | Gx 4 | 831848 | Defective ups no power,for maintenance and xpertcheck | Performed maintenance,cleaned unit,cleaned filter,plunger rod,icore,catridge bay, gathered reports, replaced ups fuse,,checked power supply | Pms and xpertpert check | A1,A2,A3 failed xpert check A4 passed | For replacement of ups | Xpertcheck | 9525 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527849759111256717/5527849759111256717_signature_28.png | Mary Claire P Baptista | https://www.jotform.com/uploads/macarewecare/222927978372067/5527849759111256717/5527849759111256717_signature_31.png
| 2023-02-21 | 2022-11-10 14:21 | Lung center of the Philippines | Quezon ave.qc | NCR | NCR 2nd - Quezon City | Bloodbank | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 23512 | MACARE-HBDC | lcppathology@gmail.com | 26 |  |  | MCS + | 08H188 | For maintenance | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527836862774004547/5527836862774004547_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/5527836862774004547/5527836862774004547_signature_31.png
| 2023-02-21 | 2023-02-21 15:07 | General Tinio RHU | General Tinio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23511 | DOH | manuelharrisf@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 433 | 2023-02-07 | Genexpert IV | 831695 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck - All Modules Passed | Xpertcheck 5pc | 9590 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527816872021282278/5527816872021282278_signature_28.png | Harris F. Manuel | https://www.jotform.com/uploads/macarewecare/222927978372067/5527816872021282278/5527816872021282278_signature_31.png
| 2023-02-21 | 2023-02-21 13:57 | Tarlac Provincial Hospital | Hospital Drive San Vicemte Tarlac City | 3 | Tarlac | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23510 | PBSP | molecularlab.tph@gmail.com | 28.15 | ASSAY: Xpert MTB/RIF, COUNT: 78 ASSAY: Xpert MTB/RIF Ultra, COUNT: 613 |  | GXIV-D | 805108 | Error 2032 on module C3 | Performed surface cleaning Performed self-test and system restart Reset the gateway board Troubleshoot ultrasonic horn by pushin it downwards Valve position check-up Performed functionality test | Cepheid troubleshooting | Unit is operational except C2 and C3 | C3 HW FAILED for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527739843155099954/5527739843155099954_signature_28.png | Camille Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5527739843155099954/5527739843155099954_signature_31.png
| 2023-02-21 | 2023-02-21 15:26 | Visayas Med Hospital | Osmena Blvd. Cebu City | 7 | Cebu | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 23509 | COMMERCIAL | vcmclab123@gmail.com | 25 |  | 2023-02-21 | Mcs+ | 5M204 | Due for Calibration | Perform calibration and Preventive maintenance | Calibrate and conduct Preventive Maintenance | Units are within the standard for calibration. PASSED during calibration |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527738647214089304/5527738647214089304_signature_28.png | Maria Mirasol U. Trinidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5527738647214089304/5527738647214089304_signature_31.png
| 2023-02-21 | 2023-02-21 14:43 | Matnog Medicare Community Hospital | Matnog, Sorsogon | 5 | Sorsogon | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 23508 | PBSP | mernadithg@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF, COUNT: 97 ASSAY: Xpert MTB/RIF Ultra, COUNT: 687 |  | GXIV-2D | 835200 | #NAME? | #NAME? | #NAME? | Module B2 is operational. Module B1 to be reported for relplacement. CPU is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527714792106841241/5527714792106841241_signature_28.png | Mernadith Garcera | https://www.jotform.com/uploads/macarewecare/222927978372067/5527714792106841241/5527714792106841241_signature_31.png
| 2023-02-21 | 2023-02-21 14:06 | Malungon Rural Health Unit | Malungon | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23507 | DOH | ulbenarioe@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 300 ASSAY: Xpert MTB/RIF Ultra, COUNT: 280 | 2023-02-21 | Gx-IV | 838546 | Cpu Problem | Replace Genexpert CPU SN: CPHO520F08267 with service CPU SN: CPH0320F05932 Perform Genexpert machine Preventive maintenance | Replace Genexpert CPU SN: CPHO520F08267 with service CPU SN: CPH0320F05932 Perform Genexpert machine Preventive maintenance | 4 operational modules |  | Lenovo ThinkCentre M730s | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527700998965394104/5527700998965394104_signature_28.png | Ermie D. Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5527700998965394104/5527700998965394104_signature_31.png
| 2023-02-21 | 2023-02-21 13:43 | Siay Rural Health Unit | Poblacion Siay, Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | INSTALLATION REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23506 | PBSP | juvyouano@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 887 | 2023-02-21 | GX 4 | 815304 | Monitor malfunction, stacked Cartridge | Replaced With a new Monitor. Removed stacked Cartridge and Preventive Maintenance | Replaced With a new Monitor. Removed stacked Cartridge and Preventive Maintenance | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527687774279133665/5527687774279133665_signature_28.png | Juvy E. Ouano | https://www.jotform.com/uploads/macarewecare/222927978372067/5527687774279133665/5527687774279133665_signature_31.png
| 2023-02-21 | 2023-02-21 13:19 | Buhangin District Health Center | NHA Village Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23505 | DOH | buhanginlaboratory@gmail.com | 26.8 | ASSAY: Xpert MTB/RIF, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 662 | 2023-02-21 | GX-IV | 838232 | For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat, and System Log. Archived Data test results | Perform Preventive Maintenance | Unit is Fully Operational | N/A | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527664527819773358/5527664527819773358_signature_28.png | Katherine Deocampo | https://www.jotform.com/uploads/macarewecare/222927978372067/5527664527819773358/5527664527819773358_signature_31.png
| 2023-02-21 | 2023-02-21 12:59 | Bulan Rural Health Unit | Bulan, Sorsogon | 5 | Sorsogon | TB Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23504 | PBSP | bing.gonzales33@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 344 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2704 | 2023-02-21 | GXIV-4D | 807983 | -Module B2/698310 defective -Stuck cartridge at module B3 | #NAME? | -Replace defective module B2/698310 with new module SN: 210070787 -Remove stuck cartridge in module B3/210070808 | -Module B1, B2, & B3 are functional -Module B4 is disabled (reported defective) |  | 900-0508 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527651396811673138/5527651396811673138_signature_28.png | Claribel G. Nacion | https://www.jotform.com/uploads/macarewecare/222927978372067/5527651396811673138/5527651396811673138_signature_31.png
| 2023-02-21 | 2023-02-21 11:14 | Gerona Rural Health Unit | Brgy. Pinasling Gerona, Tarlac | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23503 | DOH | huwakie15@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 251 | 2023-02-02 | GXIV-D | 831693 | Due for Calibration | Performed surface cleaning Cleaning of filter,plunger rod,cartridge bay and icore Archived test results Performed calibration Performed functionality test | Preventive maintenance and Calibration | Unit is operational | All modules passed calibration | 3 pcs Xpertcheck kits | 9569 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527617379016243399/5527617379016243399_signature_28.png | Artemio I. Nool III | https://www.jotform.com/uploads/macarewecare/222927978372067/5527617379016243399/5527617379016243399_signature_31.png
| 2023-02-21 | 2023-02-21 11:31 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23502 | PBSP | wcpatron11x4@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 189 ASSAY: Xpert MTB/RIF Ultra, COUNT: 750 | 2023-02-20 | GXIV | 812561 | Unit for preventive maintenance. For installation of PSU and gateway board. | Performed installation of new PSU and gateway board. Performed preventive maintenance. Generated reports. | Repair and precentive maintenance. | Unit is now operational. |  | Power supply unit Gateway board kit | DR 9642 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527599573619220203/5527599573619220203_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5527599573619220203/5527599573619220203_signature_31.png
| 2023-02-21 | 2023-02-21 09:19 | CULION SANITARIUM AND GENERAL HOSPITAL | CULION, PALAWAN | 4B | Palawan | LABORATORY | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 23501 | PBSP | culsanlab@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 62 ASSAY: Xpert MTB/RIF Ultra, COUNT: 51 ASSAY: Xpert SARS-CoV-2, COUNT: 1617 |  | GX-II | 835101 | FOR MODULE A1 AND MODULE A2 REPLACEMENT | #NAME? | -PERFORM MODULE A1 DEF SN:607206 REPLACEMENT WITH NEW SN: 210068599 -PERFORM MODULE A2 DEF SN:624779 REPLACEMENT WITH NEW SN: 210068606 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527524404715452968/5527524404715452968_signature_28.png | Victoria p. derpo | https://www.jotform.com/uploads/macarewecare/222927978372067/5527524404715452968/5527524404715452968_signature_31.png
| 2023-02-21 | 2023-02-20 14:46 | Munoz CHO | Munoz City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23500 | DOH | scomrtdl@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 354 ASSAY: Xpert MTB/RIF Ultra, COUNT: 496 | 2023-02-07 | Genexpert IV | 831673 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | All modules passed | Xpertcheck 2pc | 9593 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5527492092024036830/5527492092024036830_signature_28.png | Catherine I. Celestino | https://www.jotform.com/uploads/macarewecare/222927978372067/5527492092024036830/5527492092024036830_signature_31.png
| 2023-02-20 | 2023-02-20 16:38 | CORON RURAL HEALTH UNIT | Coron , Palawan | 4B | Palawan | TB-DOTS | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23499 | PBSP | lycking@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 228 ASSAY: Xpert MTB/RIF Ultra, COUNT: 431 |  | GX-IV | 807612 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | -EQUIPMENT IS OPERATIONAL, MODULE B4 IS DISABLED DUE TO ERROR 4017 AND H/W FAIL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526920639416143649/5526920639416143649_signature_28.png | Lyca Torrible | https://www.jotform.com/uploads/macarewecare/222927978372067/5526920639416143649/5526920639416143649_signature_31.png
| 2023-02-20 | 2023-02-20 14:00 | Danao City Health Office | Danao City, Cebu | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23498 | PBSP | chockenriquez@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1379 ASSAY: Xpert BCR-ABL Ultra, COUNT: 977 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GX IV | 808286 | Module B2 and Module B4 | Clean Fan Filter | Clean Fan Filter, assist end-user on using of jotform apps | Module B2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526858500347875863/5526858500347875863_signature_28.png | Nicanor Enriquez Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5526858500347875863/5526858500347875863_signature_31.png
| 2023-02-20 | 2023-02-20 14:00 | Danao City Health Office | Danao City, Cebu | 7 | Cebu | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23497 | PBSP | chockenriquez@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1379 ASSAY: Xpert BCR-ABL Ultra, COUNT: 977 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GX IV | 808286 | Module B2 and Module B4 | Clean Fan Filter | Clean Fan Filter, assist end-user on using of jotform apps | Module B2 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526856968724034758/5526856968724034758_signature_28.png | Nicanor Enriquez Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5526856968724034758/5526856968724034758_signature_31.png
| 2023-02-20 | 2023-02-20 14:34 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23496 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 158 ASSAY: Xpert MTB/RIF, COUNT: 1579 ASSAY: Xpert MTB/RIF Ultra, COUNT: 886 |  | GXIV-4D | 838285 | NA | #NAME? | NA | All four modules are operational | Module A3 for observation - temp errors occured recently |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526849038129133013/5526849038129133013_signature_28.png | Christine Ivy Paula S. Agtuca | https://www.jotform.com/uploads/macarewecare/222927978372067/5526849038129133013/5526849038129133013_signature_31.png
| 2023-02-20 | 2023-02-20 14:39 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23495 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 162 ASSAY: Xpert MTB/RIF, COUNT: 1113 ASSAY: Xpert MTB/RIF Ultra, COUNT: 846 |  | GXIV-4D | 820206 | NA | #NAME? | NA | All four modules are functional | Still waiting for module 2 allocation, temp used service module |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526848318127023518/5526848318127023518_signature_28.png | Christine Ivy Paula Agtuca | https://www.jotform.com/uploads/macarewecare/222927978372067/5526848318127023518/5526848318127023518_signature_31.png
| 2023-02-20 | 2023-02-20 13:51 | Rosales Rural Health Unit | Rosales Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23494 | DOH | allan.vinluan2013@gmail.com | 31.5 | ASSAY: Xpert MTB/RIF, COUNT: 424 ASSAY: Xpert MTB/RIF Ultra, COUNT: 115 | 2023-02-02 | GXIV-D | 838472 | Due for calibration | Preventive maintenance and calibration | Performed surface cleaning Cleaning of plunger rod,cartridge bay,icore and filter Archived test results Performed calibration Gathered reports Performed funcyionality test | Unit is operational except module B3 Failed Xpertcheck | B3 for reporting | 6 Pcs Xpertcheck kits | 9574 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526829899017434065/5526829899017434065_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5526829899017434065/5526829899017434065_signature_31.png
| 2023-02-20 | 2023-02-20 13:51 | Rosales Rural Health Unit | Rosales Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23493 | DOH | allan.vinluan2013@gmail.com | 31.5 | ASSAY: Xpert MTB/RIF, COUNT: 424 ASSAY: Xpert MTB/RIF Ultra, COUNT: 115 | 2023-02-02 | GXIV-D | 838472 | Due for calibration | Preventive maintenance and calibration | Performed surface cleaning Cleaning of plunger rod,cartridge bay,icore and filter Archived test results Performed calibration Gathered reports Performed funcyionality test | Unit is operational except module B3 Failed Xpertcheck | B3 for reporting | 6 Pcs Xpertcheck kits | 9574 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526829719016636147/5526829719016636147_signature_28.png | Allan Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5526829719016636147/5526829719016636147_signature_31.png
| 2023-02-20 | 2023-02-20 12:22 | Urdaneta District Hospital | Urdaneta Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23492 | PBSP | warlyacosta@gmail.com | 29.25 | ASSAY: Xpert MTB/RIF, COUNT: 903 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1469 |  | GXIV-D | 805109 | Filter replacement Cpu bios discharged | Troubleshooting | Replaced filter Performed preventive maintenance Performed functionality test | Unit is operational except B2 and B3 | B2 for reporting, replaced filter 2pcs |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526804889019951488/5526804889019951488_signature_28.png | Warlito Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5526804889019951488/5526804889019951488_signature_31.png
| 2023-02-20 | 2023-02-16 16:20 | Quezon Medical Center | Quezon Avje. St. | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23489 | PBSP | joannpalermo0613@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 227 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2164 |  | GXIV - D | 804692 | Unit is due for xpertcheck | xpertcheck | Performed xpertcheck Performed Maintenance Generate Reports | 1 module failed, 2 passed, 1 newly replaced |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526743881211837241/5526743881211837241_signature_28.png | Jo-Ann P. Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5526743881211837241/5526743881211837241_signature_31.png
| 2023-02-20 | 2023-02-16 11:48 | Bondoc Peninsula District Hospital | Catanauan | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23491 | PBSP | trinity_carisse136@yahoo.com.ph | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 58 | 2023-02-13 | GXIV-D - L | 805339 | 4 modules with temperature errors | Replacement | Module Replacement of 4 modules Performed Maintenance Generate Reports | All Modules are newly replaced | 603919 → 210068514,687770 → 210068591,647644 → 210068501,647675 →210068502 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526743881216705396/5526743881216705396_signature_28.png | Trinity carisse puyal | https://www.jotform.com/uploads/macarewecare/222927978372067/5526743881216705396/5526743881216705396_signature_31.png
| 2023-02-20 | 2023-02-14 10:31 | De La Salle Health And Sciences Institute | D. Mangubat, Dasmariñas City | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23490 | PBSP | acpabruada@dlshi.edu.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 367 | 2023-02-13 | GXIV-D | 707514 | Module 2 has temperature errors | Module Replacement | Replace Module 2 Performed Maintenance Generate Reports | 3 operational, 1 defective | B2:672006 → 724670 | GX MODULE 6-C | G 5193 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526743881217905400/5526743881217905400_signature_28.png | Angie C. Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5526743881217905400/5526743881217905400_signature_31.png
| 2023-02-20 | 2023-02-20 10:36 | Dr Jorge P. Royeca Hospital | General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23488 | PBSP | fatchoi12001@yahoo.com | 32.8 | ASSAY: Xpert MTB/RIF, COUNT: 352 ASSAY: Xpert MTB/RIF Ultra, COUNT: 491 ASSAY: Xpert HIV-1 Viral Load, COUNT: 37 |  | Gx-IV | 838130 | Genexpert room no Air-condition installed Genexpert module temperature not within the acceptable range 15-30°C | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules | Genexpert machine need to install air-conditioning unit for the required temperature |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526701715968064217/5526701715968064217_signature_28.png | Maria Ivy A. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5526701715968064217/5526701715968064217_signature_31.png
| 2023-02-20 | 2023-02-17 12:15 | Panopdopan District Hospital | Bannit, Lamut, Ifugao | CAR | Ifugao | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23487 | PBSP | meannebahiwag@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 104 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1194 | 2023-02-17 | Genexpert IV | 807167 | Module 4 Hardware Failure | Gathered Reports to report | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement - 662351 - 684468 | Genexpert Module - 684468 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526690287213920951/5526690287213920951_signature_28.png | Mary Ann Bahiwag | https://www.jotform.com/uploads/macarewecare/222927978372067/5526690287213920951/5526690287213920951_signature_31.png
| 2023-02-20 | 2023-02-16 15:52 | Jones RHU | Jones, Isabela | 2 | Isabela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23486 | DOH | oroprecy19@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 294 ASSAY: Xpert MTB/RIF Ultra, COUNT: 762 | 2023-02-16 | Genexpert IV | 838484 | Error 2005 | Performed Xpertcheck | Performed Maintenance Performed Calibration Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck Result - All Modules Passed | Xpertcheck kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526690287217450970/5526690287217450970_signature_28.png | Precy Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5526690287217450970/5526690287217450970_signature_31.png
| 2023-02-20 | 2023-02-17 14:29 | Sta. Maria Rural Health Unit | Sta. Maria | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23485 | PBSP | maidauridorado@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 85 ASSAY: Xpert MTB/RIF Ultra, COUNT: 186 |  | GX-IV-D | 838133 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. | All modules passed. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526630391212160900/5526630391212160900_signature_28.png | Maida U. Dorado | https://www.jotform.com/uploads/macarewecare/222927978372067/5526630391212160900/5526630391212160900_signature_31.png
| 2023-02-20 | 2023-02-17 12:09 | Kamuning Health Center | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23484 | DOH | vicvic.chua25@yahoo.com | 25 | ASSAY: Others, COUNT: 419 ASSAY: Xpert MTB/RIF, COUNT: 92 ASSAY: Xpert MTB/RIF Ultra, COUNT: 418 | 2023-01-18 | GXIV | 831612 | Unit for calibration (A2 & A4) | #NAME? | Performed Xpertcheck | Unit is operational. A1 & A3 already reported for replacement. A2 & A4 passed calibration. |  | Xpert Check kit | 9499 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526625851215361421/5526625851215361421_signature_28.png | Erica Fergel G. Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/5526625851215361421/5526625851215361421_signature_31.png
| 2023-02-20 | 2023-02-17 15:01 | East Avenue Medical Center | East Ave., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | REPAIR TROUBLESHOOTING | HAEMONETICS | Cienne Mae Valerio |  |  |  |  | 23483 | COMMERCIAL | bloodbankeamc@gmaiI.com | 25 |  | 2023-02-17 | MCS+ | 15H044G | No Display | Checked Connections Checked CPU Board Measured voltage of CPU Board | Performed Troubleshooting Reseat CPU Board | Unit is not operational. Suspected CPU Board Related Issue. | Suspected CPU Related Issue | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526619031213125908/5526619031213125908_signature_28.png | Noralen Magalong | https://www.jotform.com/uploads/macarewecare/222927978372067/5526619031213125908/5526619031213125908_signature_31.png
| 2023-02-20 | 2023-02-17 09:12 | San Pedro Rural Health Unit 2 | Brgy. Langgam San Pedro | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23480 | DOH | leniemahalin@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 861 ASSAY: Xpert MTB/RIF, COUNT: 188 ASSAY: Xpert MTB/RIF Ultra, COUNT: 801 ASSAY: Xpert HIV-1 Viral Load, COUNT: 127 | 2023-02-13 | GXIV | 821189 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and xpertcheck. | Unit is operational. Module A2 (784156) service required. |  | Xpertcheck. | DR 9614 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526617381211555589/5526617381211555589_signature_28.png | Ma. Celenita Mahalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5526617381211555589/5526617381211555589_signature_31.png
| 2023-02-20 | 2023-02-17 14:42 | San Pedro City Health Office 1 | San Pedro Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23481 | DOH | venersantos0425@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 435 ASSAY: Xpert MTB/RIF, COUNT: 184 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1120 | 2023-02-13 | GXIV | 821224 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules A1, A2 and A4, passed Xpertcheck. |  | Xpertcheck | DR 9613 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526617381217317444/5526617381217317444_signature_28.png | Vener R. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5526617381217317444/5526617381217317444_signature_31.png
| 2023-02-20 | 2023-02-16 10:49 | Binan City Health Office 2 | Binan Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23482 | DOH | eloicyparra@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 318 ASSAY: Xpert MTB/RIF Ultra, COUNT: 984 ASSAY: Xpert HIV-1 Viral Load, COUNT: 235 | 2023-01-17 | GXIV | 821278 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules A1(730690), A2(626255) and A3(730710) are service required. |  | Xpertcheck | DR 9480 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526617381218860154/5526617381218860154_signature_28.png | Eloicy Parra | https://www.jotform.com/uploads/macarewecare/222927978372067/5526617381218860154/5526617381218860154_signature_31.png
| 2023-02-20 | 2023-02-15 09:58 | National Tuberculosis Reference Laboratory | Filinvest | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23477 | PBSP | mspogosa.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1361 | 2023-02-13 | GXIV-D | 707512 | 2 modules with temp errors | Replacement | Module replacement in modules 1 and 4 Performed Maintenance Generate Reports | Unit is operational | A1: 695979->702147 & A4: 644321-> 706443 | Gx Module | G 5189 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526613831212907121/5526613831212907121_signature_28.png | Maria Guileane Pogosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5526613831212907121/5526613831212907121_signature_31.png
| 2023-02-20 | 2023-02-15 09:51 | National Tuberculosis Reference Laboratory | Filinvest | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23479 | PBSP | mspogosa.pbsp@gmail.com | 26 | ASSAY: Others, COUNT: 3743 | 2023-02-13 | GXIV-D-10C | 848244 | Defective monitor Module 4 errors | Replacement | Module and monitor replacement Performed maintenance Generate reports | Unit is operational | B4: 210038638-->210078090 | Monitor | G 5189 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526613831214211228/5526613831214211228_signature_28.png | Maria Guileane Pogosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5526613831214211228/5526613831214211228_signature_31.png
| 2023-02-20 | 2023-02-15 10:03 | National Tuberculosis Reference Laboratory | Filinvest | NCR | NCR 4th - Muntinlupa | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23478 | PBSP | mspogosa.pbsp@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1361 | 2023-02-13 | GXIV-D | 807606 | Module 4 with temperature problems | Replacement | Module replacement in B4 Performed Maintenance Generate reports | Unit is operational | B4: 619472-> 757462 | Gx Module | G 5189 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5526613831216564062/5526613831216564062_signature_28.png | Maria Guileane Pogosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5526613831216564062/5526613831216564062_signature_31.png
| 2023-02-17 | 2023-02-17 15:36 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23476 | PBSP | vincentmilestone20@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3032 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 ASSAY: Xpert SARS-CoV-2, COUNT: 9953 | 2023-02-17 | GX-IV | 834236 | For Preventive Maintenance, and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration Replaced the Defective Module B2: SN:783534 with service Module SN:682931 | Perform Preventive Maintenance and XpertCheck Calibration | Module B2 and B4 for Service Required Module B1 and B3 passed the Calibration | Service module is from Gx16 | Xpertcheck Ce5 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524295249781865011/5524295249781865011_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5524295249781865011/5524295249781865011_signature_31.png
| 2023-02-17 | 2023-02-17 15:36 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23475 | PBSP | vincentmilestone20@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3032 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 ASSAY: Xpert SARS-CoV-2, COUNT: 9953 | 2023-02-17 | GX-IV | 834236 | For Preventive Maintenance, and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration Replaced the Defective Module B2: SN:783534 with service Module SN:682931 | Perform Preventive Maintenance and XpertCheck Calibration | Module B2 and B4 for Service Required Module B1 and B3 passed the Calibration | Service module is from Gx16 | Xpertcheck Ce5 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524295199786183719/5524295199786183719_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5524295199786183719/5524295199786183719_signature_31.png
| 2023-02-17 | 2023-02-17 15:40 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23474 | PBSP | vincentmilestone20@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3032 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 ASSAY: Xpert SARS-CoV-2, COUNT: 9953 | 2023-02-17 | GX-IV | 805063 | For Preventive Maintenance, and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat, and System Log Archived Data test results Perform XPERTCHECK Calibration | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration | Na | Xpertcheck CE 5 | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294739786585952/5524294739786585952_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294739786585952/5524294739786585952_signature_31.png
| 2023-02-17 | 2023-02-17 15:40 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23473 | PBSP | vincentmilestone20@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3032 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 ASSAY: Xpert SARS-CoV-2, COUNT: 9953 | 2023-02-17 | GX-IV | 805063 | For Preventive Maintenance, and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat, and System Log Archived Data test results Perform XPERTCHECK Calibration | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration | Na | Xpertcheck CE 5 | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294719786066676/5524294719786066676_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294719786066676/5524294719786066676_signature_31.png
| 2023-02-17 | 2023-02-17 13:55 | Davao Regional and Medical Center | Brgy. Apokon tagum city | 11 | Davao del Norte | Molecular laboratory | YES | REPAIR TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23472 | PBSP | vincentmilestone20@gmail.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 947 ASSAY: Xpert SARS-CoV-2, COUNT: 3540 | 2023-02-16 | Gx-VXI | 838756 | Power supply Problem | Power supply troubleshooting Replace 1 power supply of GX-XVI Replace module SN: 866583 with service module SN: 870474 | Power supply troubleshooting Replace 1 power supply of GX-XVI Replace module SN: 866583 with service module SN: 870474 | 8 operational modules | Subject for power supply replacement | Power supply LE240F | 100-4795 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294639483558782/5524294639483558782_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294639483558782/5524294639483558782_signature_31.png
| 2023-02-17 | 2023-02-17 13:55 | Davao Regional and Medical Center | Brgy. Apokon tagum city | 11 | Davao del Norte | Molecular laboratory | YES | REPAIR TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23471 | PBSP | vincentmilestone20@gmail.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 947 ASSAY: Xpert SARS-CoV-2, COUNT: 3540 | 2023-02-16 | Gx-VXI | 838756 | Power supply Problem | Power supply troubleshooting Replace 1 power supply of GX-XVI Replace module SN: 866583 with service module SN: 870474 | Power supply troubleshooting Replace 1 power supply of GX-XVI Replace module SN: 866583 with service module SN: 870474 | 8 operational modules | Subject for power supply replacement | Power supply LE240F | 100-4795 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294619487284971/5524294619487284971_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294619487284971/5524294619487284971_signature_31.png
| 2023-02-17 | 2023-02-17 15:43 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23470 | PBSP | vincentmilestone20@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3032 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 ASSAY: Xpert SARS-CoV-2, COUNT: 9953 | 2023-02-17 | Gx IV | 834440 | For Preventive Maintenance, and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration Replaced Defective Module D3 SN:864397 with service Module SN:869918 | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration but Module 3 still error exist | Module D1 for under observation and Service Module is from Gx16 | Xpertcheck CE 5 | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294299782259471/5524294299782259471_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294299782259471/5524294299782259471_signature_31.png
| 2023-02-17 | 2023-02-17 15:43 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23469 | PBSP | vincentmilestone20@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3032 ASSAY: Xpert HIV-1 Viral Load, COUNT: 263 ASSAY: Xpert SARS-CoV-2, COUNT: 9953 | 2023-02-17 | Gx IV | 834440 | For Preventive Maintenance, and XpertCheck Calibration | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Perform XPERTCHECK Calibration Replaced Defective Module D3 SN:864397 with service Module SN:869918 | Perform Preventive Maintenance and XpertCheck Calibration | All Modules Passed the Calibration but Module 3 still error exist | Module D1 for under observation and Service Module is from Gx16 | Xpertcheck CE 5 | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294159788582194/5524294159788582194_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5524294159788582194/5524294159788582194_signature_31.png
| 2023-02-17 | 2023-02-17 13:45 | Minalin Rural Health Unit | San Nicolas Minalin | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23468 | PBSP | bernadetteann0318@gmail.com | 30.6 | ASSAY: Xpert MTB/RIF, COUNT: 108 ASSAY: Xpert MTB/RIF Ultra, COUNT: 513 | 2022-08-12 | GXIV-D | 834688 | Machine shutdown Due for calibration | Performed UPS troubleshooting Checked Cable Connection Performed surfqce cleaning Cleaning of Icore,cartridge bay,plunger rod and filter Installation of Xpertcheck application V1.5a Performed calibration Performed functionality test | Preventive maintenance Genexpert Calibration Gathered reports | Unit is operational,UPS under observation | All modules passade calibration | 4pcs Genexpert Xpertcheck kits | 3312 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524237399515360873/5524237399515360873_signature_28.png | Bernadette Lazaro | https://www.jotform.com/uploads/macarewecare/222927978372067/5524237399515360873/5524237399515360873_signature_31.png
| 2023-02-17 | 2023-02-17 13:33 | Dr. Joerge P. Royeca Hospital | General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23467 | PBSP | fatchoi12001@yahoo.com | 34.4 | ASSAY: Xpert MTB/RIF, COUNT: 257 ASSAY: Xpert MTB/RIF Ultra, COUNT: 384 ASSAY: Xpert HIV-1 Viral Load, COUNT: 47 | 2023-02-10 | Gx- IV | 805007 | Module B1, B2 and B3 for Service Required Genexpert machine modules temperature not within the acceptable range of 15-30°C | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | 4 operational modules | Genexpert machine need to install air-conditioning unit for required temperature |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524219409489170755/5524219409489170755_signature_28.png | Maria Ivy A. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5524219409489170755/5524219409489170755_signature_31.png
| 2023-02-17 | 2023-02-17 13:33 | Dr. Joerge P. Royeca Hospital | General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23466 | PBSP | fatchoi12001@yahoo.com | 34.4 | ASSAY: Xpert MTB/RIF, COUNT: 257 ASSAY: Xpert MTB/RIF Ultra, COUNT: 384 ASSAY: Xpert HIV-1 Viral Load, COUNT: 47 | 2023-02-10 | Gx- IV | 805007 | Module B1, B2 and B3 for Service Required Genexpert machine modules temperature not within the acceptable range of 15-30°C | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | 4 operational modules | Genexpert machine need to install air-conditioning unit for required temperature |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524219379484997316/5524219379484997316_signature_28.png | Maria Ivy A. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5524219379484997316/5524219379484997316_signature_31.png
| 2023-02-17 | 2023-02-17 13:22 | Tagoloan Rural Health Unit | Poblacion, Tagoloan | 10 | Misamis Oriental | TB DOTS | NO | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23465 | DOH | venzbacoy@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 26 ASSAY: Xpert MTB/RIF Ultra, COUNT: 788 |  | GeneXpert | 838574 | N/A | Cleaned/sanitized cartridge bay, plunger and equipment case Washed fan filter Archive tests Database backup | Preventive Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524209583328207979/5524209583328207979_signature_28.png | Venz Bacoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5524209583328207979/5524209583328207979_signature_31.png
| 2023-02-17 | 2023-02-17 12:18 | Siniloan Rural Health Unit | Siniloan | 4A | Laguna | Laboratory | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23463 | PBSP | godfreyrnvillanueva@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 205 ASSAY: Xpert MTB/RIF Ultra, COUNT: 433 |  | GX-IV-D | 838165 | UPS shutting down | Checked UPS | Performed maintenance Run XpertCheck Checked UPS Generated reports | Unit is operational. UPS shutting down and has low output voltage. | Recommend to replace UPS. All modules passed. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524183593957582756/5524183593957582756_signature_28.png | Godfrey Villanueva | https://www.jotform.com/uploads/macarewecare/222927978372067/5524183593957582756/5524183593957582756_signature_31.png
| 2023-02-17 | 2023-02-15 13:33 | Lumban Rural Health Unit | Lumban | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23463 | PBSP | merlitadeguzman1216@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 213 ASSAY: Xpert MTB/RIF Ultra, COUNT: 908 |  | GX-IV-D | 838148 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Module B1, B2 and B4 passed and operational. Module B3 service required. | Module B2/806806 under observation. Module B3/806807 for reporting. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524183593957583142/5524183593957583142_signature_28.png | Merlita C. DE Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5524183593957583142/5524183593957583142_signature_31.png
| 2023-02-17 | 2023-02-16 12:04 | Philippine Orthopedic Center | Maria Clara St. | NCR | NCR 2nd - Quezon City | TB Dots | NO | REPAIR | CEPHEID | John Monsour Doctor |  |  |  |  | 23464 | PBSP | cmblanceta@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 201 | 2023-02-13 | GX-IV-D | 803088 | Installation of new CPU | Installed new CPU | Performed maintenance Installed new CPU Import assay Installed printer driver Generated reports | Unit is operational. | Module B2/618609 not operational but reported already. | Lenovo CPU | G 5192 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524183593958002115/5524183593958002115_signature_28.png | STELLA MAREE C.BERNARDO,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5524183593958002115/5524183593958002115_signature_31.png
| 2023-02-17 | 2023-02-17 12:09 | Kamuning Health Center | Kamuning, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23462 | DOH | vicvic.chua25@yahoo.com | 25 | ASSAY: Others, COUNT: 419 ASSAY: Xpert MTB/RIF, COUNT: 92 ASSAY: Xpert MTB/RIF Ultra, COUNT: 418 | 2023-01-18 | GXIV | 831612 | Unit for calibration (A2 & A4) | #NAME? | Performed Xpertcheck | Unit is operational. A1 & A3 already reported for replacement. A2 & A4 passed calibration. |  | Xpert Check kit | 9499 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524182157852159273/5524182157852159273_signature_28.png | Erica Fergel G. Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/5524182157852159273/5524182157852159273_signature_31.png
| 2023-02-17 | 2023-02-13 15:45 | To love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | Relocation | CEPHEID | Ronnel Belizar |  |  |  |  | 23459 | PBSP | laboratory@loveyourself.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: o |  | GXIV | 808619 | Unit for relocation | Relocated unit | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597851629402/5524161597851629402_signature_28.png | Wico Consul | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597851629402/5524161597851629402_signature_31.png
| 2023-02-17 | 2023-02-13 15:30 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | relocation. | CEPHEID | Ronnel Belizar |  |  |  |  | 23460 | PBSP | laboratory@loveyourself.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: .' |  | GXIV | 807956 | Unit for relocation | Relocated unit from 2nd flr. to 3rd flr. | #NAME? | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597852702549/5524161597852702549_signature_28.png | Wico Consul | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597852702549/5524161597852702549_signature_31.png
| 2023-02-17 | 2023-02-14 10:31 | Bocaue Rural Health Unit I | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23458 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 296 ASSAY: Xpert MTB/RIF Ultra, COUNT: 661 |  | GXIV | 831691 | Module B4 for re xpertcheck | Performed xpertcheck on B4 | #NAME? | Unit is operational. B4 failed xpertCheck. B4 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597856080453/5524161597856080453_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597856080453/5524161597856080453_signature_31.png
| 2023-02-17 | 2023-02-13 15:41 | To Love Yourself | Shaw Blvd., Mandalayong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | Relocation | CEPHEID | Ronnel Belizar |  |  |  |  | 23457 | PBSP | laboratory@loveyourself.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: o |  | GXIV | 807112 | Unit for relocation | Relocated unit from 2nd flr. to 3rd flr. | #NAME? | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597857480337/5524161597857480337_signature_28.png | wico Consul | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597857480337/5524161597857480337_signature_31.png
| 2023-02-17 | 2023-02-13 15:50 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | Relocation | CEPHEID | Ronnel Belizar |  |  |  |  | 23461 | DOH | laboratory@loveyourself.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: O |  | GXIV | 831895 | Unit for relocation | #NAME? | #NAME? | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597857637031/5524161597857637031_signature_28.png | wico Consul | https://www.jotform.com/uploads/macarewecare/222927978372067/5524161597857637031/5524161597857637031_signature_31.png
| 2023-02-17 | 2023-02-17 10:22 | RIVERSIDE MEDICAL CENTER INC | B.S AQUINO DRIVE, BACOLOD CITY | 6 | Negros Occidental | Laboratory | YES | INSTALLATION REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 23456 | COMMERCIAL | laboratory@rivermedcenter.net | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 39 ASSAY: Xpert HPV, COUNT: 16 ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 29 |  | GX IV | 110005268 | Defective Module B2 sn 210029540 | New module installed 210070907 | New module installed 210070907 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524106700301057110/5524106700301057110_signature_28.png | Jelyn L. Fernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5524106700301057110/5524106700301057110_signature_31.png
| 2023-02-17 | 2023-02-16 15:42 | Bulacan Medical center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | Ronnel Belizar |  |  |  |  | 23455 | PBSP | bmclab_micro@yahoo.com | 25 | ASSAY: Others, COUNT: NA | 2023-02-13 | GXIV | 805112 | Module B2 & gateway board for replacement | #NAME? | Replaced B2 & gateway board | Unit is operational. |  | GeneXpert Module SN: 654090 Gateway Board SN: 6140S22420080 | G 5190 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962414223713/5524084962414223713_signature_28.png | Marichu Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962414223713/5524084962414223713_signature_31.png
| 2023-02-17 | 2023-02-15 14:38 | Bulakan Rural Health Unit II | Bulakan, Bulacan | NCR | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23454 | DOH | l_gemino@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 184 ASSAY: Xpert MTB/RIF Ultra, COUNT: 545 | 2023-01-20 | GXIV | 831641 | Unit for xpertcheck | #NAME? | Performed Xpertcheck | Unit is operational. A1, A2, A4 passed calibration. A3 failed xpertcheck. A3 for reporting / follow-up. |  | XpertCheck kit | 9526 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962415162380/5524084962415162380_signature_28.png | Lilibeth Gemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962415162380/5524084962415162380_signature_31.png
| 2023-02-17 | 2023-02-15 11:30 | Maysan 3S Health Station | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23451 | PBSP | neldimatera@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 616 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2524 ASSAY: Xpert HIV-1 Viral Load, COUNT: 170 |  | GXIV | 814151 | Unit for xpertcheck | #NAME? | Performed Xpertcheck | Unit is operational. A3 insufficient xpertcheck kit. A2 failed xpertcheck. A4 & A4 passed xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962411946764/5524084962411946764_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962411946764/5524084962411946764_signature_31.png
| 2023-02-17 | 2023-02-15 11:37 | Maysan 3S Health Station | Maysan, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23452 | PBSP | neldimatera@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 616 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2524 ASSAY: Xpert HIV-1 Viral Load, COUNT: 170 |  | GXIV | 815381 | Unit for xpertcheck | #NAME? | Performed Xpertcheck | Unit is operational. All modules passed. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962416128967/5524084962416128967_signature_28.png | Nelliza Dimatera | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962416128967/5524084962416128967_signature_31.png
| 2023-02-17 | 2023-02-16 13:46 | Sta. Maria Rural Health Unit | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23453 | PBSP | mhosantamaria@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 318 ASSAY: Xpert MTB/RIF Ultra, COUNT: 799 |  | GXIV | 838866 | Unit for xpertcheck | #NAME? | Performed xpertCheck | Unit is operational. All modules passed calibration except B3. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962419336096/5524084962419336096_signature_28.png | Mary Grace Chi | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084962419336096/5524084962419336096_signature_31.png
| 2023-02-17 | 2023-02-14 10:31 | Bocaue Rural Health Unit I | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23449 | DOH | zaychi20@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 296 ASSAY: Xpert MTB/RIF Ultra, COUNT: 661 |  | GXIV | 831691 | Module B4 for re xpertcheck | Performed xpertcheck on B4 | #NAME? | Unit is operational. B4 failed xpertCheck. B4 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084852411414311/5524084852411414311_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084852411414311/5524084852411414311_signature_31.png
| 2023-02-17 | 2023-02-13 15:50 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | Relocation | CEPHEID | Ronnel Belizar |  |  |  |  | 23448 | DOH | laboratory@loveyourself.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: O |  | GXIV | 831895 | Unit for relocation | #NAME? | #NAME? | unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084852412026816/5524084852412026816_signature_28.png | wico Consul | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084852412026816/5524084852412026816_signature_31.png
| 2023-02-17 | 2023-02-13 15:45 | To love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | Relocation | CEPHEID | Ronnel Belizar |  |  |  |  | 23446 | PBSP | laboratory@loveyourself.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: o |  | GXIV | 808619 | Unit for relocation | Relocated unit | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084852418432041/5524084852418432041_signature_28.png | Wico Consul | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084852418432041/5524084852418432041_signature_31.png
| 2023-02-17 | 2023-02-13 15:30 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | relocation. | CEPHEID | Ronnel Belizar |  |  |  |  | 23447 | PBSP | laboratory@loveyourself.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: .' |  | GXIV | 807956 | Unit for relocation | Relocated unit from 2nd flr. to 3rd flr. | #NAME? | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084852419285413/5524084852419285413_signature_28.png | Wico Consul | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084852419285413/5524084852419285413_signature_31.png
| 2023-02-17 | 2023-02-13 15:41 | To Love Yourself | Shaw Blvd., Mandalayong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | Relocation | CEPHEID | Ronnel Belizar |  |  |  |  | 23450 | PBSP | laboratory@loveyourself.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: o |  | GXIV | 807112 | Unit for relocation | Relocated unit from 2nd flr. to 3rd flr. | #NAME? | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084862415884050/5524084862415884050_signature_28.png | wico Consul | https://www.jotform.com/uploads/macarewecare/222927978372067/5524084862415884050/5524084862415884050_signature_31.png
| 2023-02-17 | 2023-02-16 17:53 | Datu Halun Sakilan Memorial Hospital | Lamiang Bongao Tawi-Tawi | BARMM | Tawi-Tawi | Laboratory | NO | INSTALLATION CALIBRATION | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23445 | IOM | joel_kamlian2003@yahoo.com | 23 | ASSAY: , COUNT: 0 | 2023-02-16 | GX4 | 110008652 | None | Installatio and calibration | Installatio and calibration | All modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524073105119412702/5524073105119412702_signature_28.png | Ĵoel kamlian | https://www.jotform.com/uploads/macarewecare/222927978372067/5524073105119412702/5524073105119412702_signature_31.png
| 2023-02-17 | 2023-02-16 18:09 | Datu Halun Sakilan Memorial Hospital | Lamion, bongao Tawi-Tawi | BARMM | Tawi-Tawi | Laboratory | NO | INSTALLATION | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23443 | DOH | joel_kamlian2003@yahoo.com | 26 | ASSAY: , COUNT: 0 | 2023-02-16 | GX 4 | 838834 | Module 1 hardware failed | Gathered data and reported | Gathered data and reported | Module 1 hardware failed | Reccomended for module replacement | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524073075115530333/5524073075115530333_signature_28.png | Joel kamlian | https://www.jotform.com/uploads/macarewecare/222927978372067/5524073075115530333/5524073075115530333_signature_31.png
| 2023-02-17 | 2023-02-16 15:22 | Datu Halun Memorial Hospital | Bongao Tawi-Tawi | BARMM | Tawi-Tawi | Laboratory | NO | Xpertcheck Calibration, Preventive Maintenance | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23444 | DOH | rappaiballho@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 446 ASSAY: Xpert MTB/RIF Ultra, COUNT: 469 | 2023-02-16 | GX 4 | 838834 | Module 1 h/w failed | Module 2, 3 and 4 Undergone Xpertcheck | Module 2, 3 and 4 Undergone Xpertcheck, Reported module 1. | Module 1 currently hardware failed. Module 2, 3 and 4 passed the calibration | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524073075117419245/5524073075117419245_signature_28.png | Rappai O.Ballaho | https://www.jotform.com/uploads/macarewecare/222927978372067/5524073075117419245/5524073075117419245_signature_31.png
| 2023-02-17 | 2023-02-16 17:40 | Datu Halun Memorial Hospital | Ilmoh ,Lamion Bongao Tawi-Tawi | BARMM | Tawi-Tawi | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23442 | PBSP | rappaiballaho@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 115 ASSAY: Xpert MTB/RIF Ultra, COUNT: 376 | 2023-02-16 | GX - 4 | 808643 | Module 4 h/w failed | Xpertcheck calibration for module 1,2 and 3 | Xpertcheck calibration for module 1,2 and 3 | Module 4 h/w failed | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524073055114003642/5524073055114003642_signature_28.png | Rappai O Ballaho | https://www.jotform.com/uploads/macarewecare/222927978372067/5524073055114003642/5524073055114003642_signature_31.png
| 2023-02-17 | 2023-02-17 08:02 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23441 | DOH | ammegzedd@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 501 ASSAY: Xpert MTB/RIF Ultra, COUNT: 527 |  | GXIV | 838421 | For Preventive maintenance | #NAME? | #NAME? | B1 & B4 Operational. B2 & B3 (disabled) waiting for module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524017975523687488/5524017975523687488_signature_28.png | Gemma Malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5524017975523687488/5524017975523687488_signature_31.png
| 2023-02-17 | 2023-02-17 08:02 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23440 | DOH | ammegzedd@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 501 ASSAY: Xpert MTB/RIF Ultra, COUNT: 527 |  | GXIV | 838421 | For Preventive maintenance | #NAME? | #NAME? | B1 & B4 Operational. B2 & B3 (disabled) waiting for module replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5524017605529779260/5524017605529779260_signature_28.png | Gemma Malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5524017605529779260/5524017605529779260_signature_31.png
| 2023-02-16 | 2023-02-16 17:08 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | CALIBRATION | CEPHEID | Ronnel Belizar |  |  |  |  | 23439 | PBSP | bmcgenexpertlab@yahoo.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 4936 |  | GXIV | 834623 | Unit for calibration | - checked and cleaned cartridge bay, plunger rod, i core slot, filter - performed xpertcheck - generated reports - archived test reports from the beginning until 2022 - performed self tests | Performed XpertCheck | Unit is operational. All modules passed calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523520513912527693/5523520513912527693_signature_28.png | Michaella Sayo / Alenvie Penilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5523520513912527693/5523520513912527693_signature_31.png
| 2023-02-16 | 2023-02-16 17:08 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | CALIBRATION | CEPHEID | Ronnel Belizar |  |  |  |  | 23438 | PBSP | bmcgenexpertlab@yahoo.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 4936 |  | GXIV | 834623 | Unit for calibration | - checked and cleaned cartridge bay, plunger rod, i core slot, filter - performed xpertcheck - generated reports - archived test reports from the beginning until 2022 - performed self tests | Performed XpertCheck | Unit is operational. All modules passed calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523501943918980755/5523501943918980755_signature_28.png | Michaella Sayo / Alenvie Penilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5523501943918980755/5523501943918980755_signature_31.png
| 2023-02-16 | 2023-02-16 16:06 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | MAINTENANCE | HAEMONETICS | Joseph Acebedo |  |  |  |  | 23437 | COMMERCIAL | cduhlabresults2@yahoo.com | 20 |  |  | MCS+ | 14F014 | PREVENTIVE MAINTENANCE | #NAME? | CHECK AND INSPECT ROTOR ROLLER RUBBER -INSPECT CENTRIFUGE ASSEMBLY AND SPILL DETECTOR, PASSED | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523442150145956694/5523442150145956694_signature_28.png | Lovely Orongan | https://www.jotform.com/uploads/macarewecare/222927978372067/5523442150145956694/5523442150145956694_signature_31.png
| 2023-02-16 | 2023-02-16 15:50 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | MAINTENANCE | HAEMONETICS | Joseph Acebedo |  |  |  |  | 23436 | COMMERCIAL | cduhlabresults2@yahoo.com | 20 |  |  | MCS+ | 17H639SPG | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523439500141652651/5523439500141652651_signature_28.png | Lovely Orongan | https://www.jotform.com/uploads/macarewecare/222927978372067/5523439500141652651/5523439500141652651_signature_31.png
| 2023-02-16 | 2023-02-16 14:49 | Caraga Regional Hospital | Rizal Street, Brgy. Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 23435 | PBSP | abpabellan.pbsp@gmail.com | 26.1 | ASSAY: Xpert MTB/RIF, COUNT: 586 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3237 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 |  | GX-4 | 807602 | GX Serial No. 815489 - Module D3 found temperature related error | D3 already reported in jotform | Replacing Module C1 with SN: 652389 to SN: 642566 Replacing Module C2 with SN: 625609 to SN: 210018392 Replacing Module C3 with SN: 675554 to SN: 731550 Replacing Module C4 with SN: 686046 to SN: 890309 | 7 modules working and 1 module for reporting | Done getting additional data in Module D3 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523413573815839078/5523413573815839078_signature_28.png | Andrew Pabellan | https://www.jotform.com/uploads/macarewecare/222927978372067/5523413573815839078/5523413573815839078_signature_31.png
| 2023-02-16 | 2023-02-16 14:30 | Talakag Rural Health Unit | Talakag | 10 | Bukidnon | TB DOTS | NO | MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23434 | DOH | jhudielkathleen@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 167 |  | GeneXpert | 838566 | None | Cleaned plunger, cartridge bay and equipment body case Washed fan filter Generate data files Archive tests Data back up | Preventive maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523388653326871622/5523388653326871622_signature_28.png | Jhudiel Kathleen Agol | https://www.jotform.com/uploads/macarewecare/222927978372067/5523388653326871622/5523388653326871622_signature_31.png
| 2023-02-16 | 2023-02-16 12:57 | Maolboal Rural Health Unit | Moalboal Cebu | 7 | Cebu | Laboratory | YES | CALIBRATION | CEPHEID | Jovy Noel |  |  |  |  | 23433 | PBSP | roavaul@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 311 | 2023-02-16 | GX4 | 838892 | Due for Xpertcheck | Perform Xpertcheck | Perform Xpertcheck | All Modules pass for Xpertcheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523370582124660799/5523370582124660799_signature_28.png | Ava A. Gako | https://www.jotform.com/uploads/macarewecare/222927978372067/5523370582124660799/5523370582124660799_signature_31.png
| 2023-02-16 | 2023-02-16 13:49 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23432 | DOH | lornsabaigar@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 173 ASSAY: Xpert MTB/RIF Ultra, COUNT: 293 |  | GXIV | 838409 | for Preventive maintenance | #NAME? | Preventive maintenance | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523363767122921129/5523363767122921129_signature_28.png | Lorna Cabrigas Abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/5523363767122921129/5523363767122921129_signature_31.png
| 2023-02-16 | 2023-02-16 11:30 | San Manuel RHU | District 3, San Manuel, Isabela | 2 | Isabela | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23431 | PBSP | tonyantoniojr100@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 199 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1280 | 2023-02-16 | Genexpert IV | 834686 | Stuck Cartridge | Resolved stuck cartridge | Performed Maintenance Gathered Reports Enabled module 2 Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523279707122376305/5523279707122376305_signature_28.png | TONY Q. ANTONIO JR | https://www.jotform.com/uploads/macarewecare/222927978372067/5523279707122376305/5523279707122376305_signature_31.png
| 2023-02-16 | 2023-02-15 14:37 | Ilagan CHO | Ilagan City, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23430 | PBSP | cho1ilagan.laboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 652 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1016 | 2023-02-15 | Genexpert IV | 808635 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement - 684832 - 686640 | Genexpert Module - 686640 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523279707123577417/5523279707123577417_signature_28.png | April Jenneth B. Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/5523279707123577417/5523279707123577417_signature_31.png
| 2023-02-16 | 2023-02-15 14:37 | Ilagan CHO | Ilagan City, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23429 | PBSP | cho1ilagan.laboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 652 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1016 | 2023-02-15 | Genexpert IV | 808635 | N/a | N/a | Performed Maintenance Gathered Reports Performed Module Replacement Performed Functionality Test | Unit is Operational | Module Replacement - 684832 - 686640 | Genexpert Module - 686640 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523277128682728589/5523277128682728589_signature_28.png | April Jenneth B. Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/5523277128682728589/5523277128682728589_signature_31.png
| 2023-02-16 | 2023-02-16 10:39 | Victorias CHO | Montinola St., Victorias City Negros Occ. | 6 | Negros Occidental | TB DOTS | YES | CALIBRATION | CEPHEID | Emerson Bejemino |  |  |  |  | 23428 | PBSP | vincent_llamoleonor@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1000 ASSAY: Xpert MTB/RIF, COUNT: 452 |  | GX4 | 838141 | None | Save GXX file, IQR, System Log & Assay Statistics | Perform Xpert check calibration for module 1 & 2. | Module 1, 3 & 4 are functional. Module B2 has already allocation for replacement. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523253618615537603/5523253618615537603_signature_28.png | Mr. Vincent Leonor | https://www.jotform.com/uploads/macarewecare/222927978372067/5523253618615537603/5523253618615537603_signature_31.png
| 2023-02-16 | 2023-02-16 09:43 | Binan City Health Office 1 | Binan Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23427 | MACARE-HBDC | annefriam.almazora@yahoo.com | 28 | ASSAY: Others, COUNT: 0 | 2023-02-13 | GXIV | 812780 | Unit for PM. Unit for Xpertcheck. For installation of UPS and CPU. | Performed installation of CPU and UPS. Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance, installation and Xpertcheck. | Unit is operational. Module B3 (624415) is service required. |  | CPU (CPH0322F32990 | G5191 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523213621916532494/5523213621916532494_signature_28.png | Annefriam Almazora | https://www.jotform.com/uploads/macarewecare/222927978372067/5523213621916532494/5523213621916532494_signature_31.png
| 2023-02-16 | 2023-02-16 09:43 | Binan City Health Office 1 | Binan Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23426 | MACARE-HBDC | annefriam.almazora@yahoo.com | 28 | ASSAY: Others, COUNT: 0 | 2023-02-13 | GXIV | 812780 | Unit for PM. Unit for Xpertcheck. For installation of UPS and CPU. | Performed installation of CPU and UPS. Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance, installation and Xpertcheck. | Unit is operational. Module B3 (624415) is service required. |  | CPU (CPH0322F32990 | G5191 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523212651919651774/5523212651919651774_signature_28.png | Annefriam Almazora | https://www.jotform.com/uploads/macarewecare/222927978372067/5523212651919651774/5523212651919651774_signature_31.png
| 2023-02-16 | 2023-02-15 14:14 | Tondo Health Center | Tondo Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23423 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 444 ASSAY: Xpert MTB/RIF, COUNT: 472 ASSAY: Xpert MTB/RIF Ultra, COUNT: 960 ASSAY: Xpert SARS-CoV-2, COUNT: 11112 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GX16 | 844121 | CPU not powering on. For on-site troubleshooting and repair. | Performed troubleshooting. Checked and reset all internal connections. Reset RAM and insert it to other slot. | Troubleshooting and repair. | Unit is now operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523132179213007443/5523132179213007443_signature_28.png | Julie Ann D. Maniego | https://www.jotform.com/uploads/macarewecare/222927978372067/5523132179213007443/5523132179213007443_signature_31.png
| 2023-02-16 | 2023-02-15 21:59 | Tondo Foreshore Super Health Center | Pacheco St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23425 | MACARE-HBDC | olartecharlie@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1919 ASSAY: Others, COUNT: 156 |  | GXIV-10c | 110001650 | Unit for preventive maintenance. For generating of data on module A4. | Performed preventive maintenance. Performed Xpertcheck on module A4. Generated report. | Preventive maintenance. | Unit is operational. 3 operational modules. Module A4 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523132179214606833/5523132179214606833_signature_28.png | Charlie Olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5523132179214606833/5523132179214606833_signature_31.png
| 2023-02-16 | 2023-02-14 13:43 | Bureau of Quarantine | Intramuros Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23424 | MACARE-HBDC | dklo.lab@quarantine.doh.gov.ph | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 603 |  | GX16 | 844418 | Unit for Preventive Maintenance. | Performed Preventive Maintenance. Generated reports. | Preventive Maintenance. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523132179215726916/5523132179215726916_signature_28.png | Dulce Karmela Lo | https://www.jotform.com/uploads/macarewecare/222927978372067/5523132179215726916/5523132179215726916_signature_31.png
| 2023-02-16 | 2023-02-14 09:59 | Tondo Health Center | Gagalangin Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23423 | DOH | ma.lorinamanera@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 388 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1221 |  | GXIV | 831620 | Unit for preventive maintenance. | Performed preventive maintenance. | Preventive maintenance. | Unit is operational. 1 module operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5523132179218445189/5523132179218445189_signature_28.png | Lontoc, Ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5523132179218445189/5523132179218445189_signature_31.png
| 2023-02-15 | 2023-02-15 20:23 | Cagayan de Oro City Health Office | Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23422 | DOH | chocdodiagnosticlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2564 |  | GeneXpert | 838441 | H/W Fail on module B4/806415 Frequent communication lost | Cleaned/sanitized equipment case, cartridge bay and plunger Replaced fan filter Resetting of the gateway board | Preventive Maitenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522736093327485993/5522736093327485993_signature_28.png | Charice Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5522736093327485993/5522736093327485993_signature_31.png
| 2023-02-15 | 2023-02-15 14:16 | Bunawan District Health Center | Bunawan Proper, Davao City | 11 | Davao del Sur | TB DOTS | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23421 | PBSP | amutcee@gmail.com | 28.9 | ASSAY: Xpert MTB/RIF, COUNT: 72 ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 | 2023-02-15 | GX-IV | 815392 | For Calibration and Maintenance Module B3 reported Hardware Failure | Run Xpertcheck Calibration in four modules Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat, System Log | Perform XpertCheck Calibration and Maintenance | Module B1, B2 and B3 passed the Calibration Module B4 for service required | Module B3 for under observation, GeneXpert Machine ambient temperature is high, | Xpertcheck CE-5 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522547637715300837/5522547637715300837_signature_28.png | Amraida S. Camid | https://www.jotform.com/uploads/macarewecare/222927978372067/5522547637715300837/5522547637715300837_signature_31.png
| 2023-02-15 | 2023-02-15 14:16 | Bunawan District Health Center | Bunawan Proper, Davao City | 11 | Davao del Sur | TB DOTS | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23420 | PBSP | amutcee@gmail.com | 28.9 | ASSAY: Xpert MTB/RIF, COUNT: 72 ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 | 2023-02-15 | GX-IV | 815392 | For Calibration and Maintenance Module B3 reported Hardware Failure | Run Xpertcheck Calibration in four modules Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat, System Log | Perform XpertCheck Calibration and Maintenance | Module B1, B2 and B3 passed the Calibration Module B4 for service required | Module B3 for under observation, GeneXpert Machine ambient temperature is high, | Xpertcheck CE-5 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522547607715823643/5522547607715823643_signature_28.png | Amraida S. Camid | https://www.jotform.com/uploads/macarewecare/222927978372067/5522547607715823643/5522547607715823643_signature_31.png
| 2023-02-15 | 2023-02-15 14:34 | Ammw Jadsac District Hospital | Pudtol Apayao | CAR | Apayao | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23419 | PBSP | dlgblacar.pbsp@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF, COUNT: 639 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1775 | 2023-01-27 | GXIV-D | 807609 | A4 for moduoe replacement A3 HW failed, temperature error | Repair and preventive maintenance | Performed surface cleaning Performed module replacement Performed preventive maintenance Gathered reports Perfromed functionality test | Unit is operational exceot module A3 | Module A3 for reporting | Genexpert module | 38603 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522525886123350845/5522525886123350845_signature_28.png | Daisy Lacar | https://www.jotform.com/uploads/macarewecare/222927978372067/5522525886123350845/5522525886123350845_signature_31.png
| 2023-02-15 | 2023-02-15 11:09 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Xpertcheck 1 module | CEPHEID | Matthew Roxas |  |  |  |  | 23416 | DOH | zionclydematote@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 708 ASSAY: Xpert MTB/RIF, COUNT: 1043 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2607 ASSAY: Xpert HIV-1 Viral Load, COUNT: 222 | 2023-02-15 | Genexpert IV | 831718 | Error 2126, Reset when running MTB- Ultra | Performed Troubleshooting Reset the Gateway Board Performed Xpertcheck on Module 4 | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck result - Service Required | Xpertcheck 3pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522517517801193996/5522517517801193996_signature_28.png | Sunshine Matote | https://www.jotform.com/uploads/macarewecare/222927978372067/5522517517801193996/5522517517801193996_signature_31.png
| 2023-02-15 | 2023-02-15 10:02 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23417 | MACARE-HBDC | ruellamorena@gmail.com | 27 | ASSAY: Xpert SARS-CoV-2, COUNT: 522 | 2023-02-15 | Genexpert IV | 110008457 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522517517805810632/5522517517805810632_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5522517517805810632/5522517517805810632_signature_31.png
| 2023-02-15 | 2023-02-14 16:53 | Solano RHU | Brgy Roxas, Solano, Nueva Vizcaya | 2 | Nueva Vizcaya | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23418 | DOH | rhusolanolab@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 203 ASSAY: Xpert MTB/RIF Ultra, COUNT: 450 | 2023-02-14 | Genexpert IV | 838201 | Error 2037 | (Cartridge problem) | Performed Maintenance Gathered Reports Performed Calibration on all modules Performed Functionality Test | Unit is Operational | All modules passed | Xpertcheck kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522517517808642941/5522517517808642941_signature_28.png | May Vanessa A. Dela Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5522517517808642941/5522517517808642941_signature_31.png
| 2023-02-15 | 2023-02-15 14:03 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | MAINTENANCE | HAEMONETICS | Joseph Acebedo |  |  |  |  | 23415 | COMMERCIAL | pshclinlab@gmail.com | 23 |  |  | MCS* | 11H307 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522504588227932529/5522504588227932529_signature_28.png | Roda Fatima Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5522504588227932529/5522504588227932529_signature_31.png
| 2023-02-15 | 2023-02-15 13:54 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | MAINTENANCE | HAEMONETICS | Joseph Acebedo |  |  |  |  | 23414 | COMMERCIAL | pshclinlab@gmail.com | 23 |  |  | MCS+ | 02D167 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522500718222643333/5522500718222643333_signature_28.png | Roda Fatima Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5522500718222643333/5522500718222643333_signature_31.png
| 2023-02-15 | 2023-02-15 12:40 | Valentina G. Plaza Memorial Hospital | Brgy. Poblacion, Bunawan | 13 | Agusan del Sur | RTDL | NO | REPAIR CALIBRATION | CEPHEID | Faizar Avenido |  |  |  |  | 23413 | PBSP | vgpmh.official@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF, COUNT: 1276 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3490 |  | GX-4 | 815436 | B3 and B4 found Service Required | Initial Troubleshooting | Replacing Module A1 with SN: 645126 to SN:731391 XpertCheck Calibration | 2 modules working, 2 modules service required (for replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522480812519230815/5522480812519230815_signature_28.png | Honey Mae B. Panes | https://www.jotform.com/uploads/macarewecare/222927978372067/5522480812519230815/5522480812519230815_signature_31.png
| 2023-02-15 | 2023-02-15 10:52 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23411 | DOH | rhukalayaanlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 111 ASSAY: Xpert MTB/RIF Ultra, COUNT: 546 | 2023-02-07 | GX-IV-D | 821292 | Defective barcode scanner | Replaced defective barcode scanner | Performed maintenance Replaced defective barcode scanner Generated reports | Modules A1/792273 and A2/783209 not operational. Modules A3 and A4 operational. | Module A1 and A2 reported. | GeneXpert barcode scanner | DR# 9601 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522396840133179868/5522396840133179868_signature_28.png | Leslee Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5522396840133179868/5522396840133179868_signature_31.png
| 2023-02-15 | 2023-02-14 14:16 | East Avenue Medical Center | East Ave Diliman | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | HAEMONETICS | John Monsour Doctor |  |  |  |  | 23412 | COMMERCIAL | paulinobernaldo_1419@yahoo.com | 27 |  |  | MCS+ | 15H044G | Sometimes not proceeding to self test. | Reset CPU board | Checked internal connections Checked CPU board Checked calibration parameters Simulation run | Unit is operational but for further troubleshooting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5522396840136902368/5522396840136902368_signature_28.png | Christopher Paul Bernaldo | https://www.jotform.com/uploads/macarewecare/222927978372067/5522396840136902368/5522396840136902368_signature_31.png
| 2023-02-14 | 2023-02-14 15:43 | Panabo City Health Office | Poblacion District, Panabo City | 11 | Davao del Norte | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23410 | DOH | nadineeust67@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1450 | 2023-02-14 | GX-IV 4D | 838295 | Module B2 SN 808297 encounter multiple Temperature Error | Perform XPC at Module B2 SN 808297 ; | Perform XPC at the Certain Module | Module B2 SN 808297 service required | Data will be sent to headoffice for further Assessment | XPERTCHECK Cartridge | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5521704257714460130/5521704257714460130_signature_28.png | Nadine Gemma Eustaquio | https://www.jotform.com/uploads/macarewecare/222927978372067/5521704257714460130/5521704257714460130_signature_31.png
| 2023-02-14 | 2023-02-14 15:43 | Panabo City Health Office | Poblacion District, Panabo City | 11 | Davao del Norte | TB DOTS | NO | TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23409 | DOH | nadineeust67@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1450 | 2023-02-14 | GX-IV 4D | 838295 | Module B2 SN 808297 encounter multiple Temperature Error | Perform XPC at Module B2 SN 808297 ; | Perform XPC at the Certain Module | Module B2 SN 808297 service required | Data will be sent to headoffice for further Assessment | XPERTCHECK Cartridge | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5521704217711505611/5521704217711505611_signature_28.png | Nadine Gemma Eustaquio | https://www.jotform.com/uploads/macarewecare/222927978372067/5521704217711505611/5521704217711505611_signature_31.png
| 2023-02-14 | 2023-02-14 14:12 | BIOPATH CAP | CEBU CITY | 7 | Cebu | MOLECULAR LABORATORY | NO | MAINTENANCE | RBC BIOSCIENCE | Joseph Acebedo |  |  |  |  | 23408 | COMMERCIAL | jkee@biopath.ph | 22 |  |  | MAGCORE PLUS II AUTOMATED NUCLEIC ACID EXTRACTOR | P26R021820135 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5521645748302614674/5521645748302614674_signature_28.png | Kee, Mary Joy C. | https://www.jotform.com/uploads/macarewecare/222927978372067/5521645748302614674/5521645748302614674_signature_31.png
| 2023-02-14 | 2023-02-14 13:17 | BIOPATH CAP | CEBU CITY | 7 | Cebu | MOLECULAR LABORATORY | NO | MAINTENANCE | RBC BIOSCIENCE | Joseph Acebedo |  |  |  |  | 23407 | COMMERCIAL | jkee@biopath.ph | 22 |  |  | MAGCORE PLUS II AUTOMATED NUCLEIC ACID EXTRACTOR | P26R021920135 | FOR PREVENTIVE MAINTENANCE | #NAME? | #NAME? | #NAME? |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5521643238304185657/5521643238304185657_signature_28.png | Mary Joy C. kee | https://www.jotform.com/uploads/macarewecare/222927978372067/5521643238304185657/5521643238304185657_signature_31.png
| 2023-02-14 | 2023-02-14 13:35 | Northern Mindanao Medical Center | Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | YES | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23406 | PBSP | nmmcgenexpress@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 364 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7032 ASSAY: Xpert SARS-CoV-2, COUNT: 13266 ASSAY: Xpert HIV-1 Viral Load, COUNT: 426 | 2023-01-11 | GeneXpert | 844118 | Monitor display issue | Replacement of LCD Monitor New LCD monitor: ThinkVision S22e-20 Flat Panel Monitor | Replaced LCD Monitor | Operational |  | ThinkVision S22e-20 Flat Panel Monitor | #01-2023-2588 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5521625713715743455/5521625713715743455_signature_28.png | Rione Frances Lomongo | https://www.jotform.com/uploads/macarewecare/222927978372067/5521625713715743455/5521625713715743455_signature_31.png
| 2023-02-14 | 2023-02-14 12:46 | Moncada Rural Health Unit | Moncada Tarlac | 3 | Tarlac | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23405 | PBSP | mananganjeslen30@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 | 2023-02-09 | GXIV-D | 808292 | For CPU replacement | Installation and setup of new CPU Setup of user admin Installation of assays Performed preventive maintenance Gathered Reports Performed functionality test | Installation and Preventive maintenance | Unit is operational |  | 800-0560gxdx(config Desktop,Lenovo) | 38940 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5521598455415660610/5521598455415660610_signature_28.png | Jessa Pabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5521598455415660610/5521598455415660610_signature_31.png
| 2023-02-14 | 2023-02-14 09:36 | Minglanilla Rural Health Unit | Ward II, Minglanilla Poblacion, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 23404 | DOH | brendagonzagaparedes@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 480 | 2023-02-14 | Gx4 | 838512 | 5007 | Clean Icore and Plunger | Clean Icore and Plunger | All module are operational. The error in Module Serial number 862343 are sample related error. |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5521483467013678503/5521483467013678503_signature_28.png | Brenda Paredes, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5521483467013678503/5521483467013678503_signature_31.png
| 2023-02-14 | 2023-02-10 10:56 | Dr. Jorge P. Royeca Hospital | General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23402 | PBSP | fatchoi12001@yahoo.com | 32.8 | ASSAY: Xpert MTB/RIF, COUNT: 352 ASSAY: Xpert MTB/RIF Ultra, COUNT: 491 ASSAY: Xpert HIV-1 Viral Load, COUNT: 37 |  | Gx - IV | 838130 | Genexpert room no Air-condition installed Genexpert module temperature not within the acceptable range of 15-30°C | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | 4 operational modules | Genexpert machine need to install air-conditioning unit for required temperature |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5521374271756437696/5521374271756437696_signature_28.png | Maria Ivy A. Martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5521374271756437696/5521374271756437696_signature_31.png
| 2023-02-13 | 2023-02-13 14:46 | Bayugan City Health Office | Brgy. Poblacion, Bayugan City | 13 | Agusan del Sur | RTDL | NO | REPAIR | CEPHEID | Faizar Avenido |  |  |  |  | 23401 | PBSP | diolaugay@gmail.com | 26.46 | ASSAY: Xpert MTB/RIF, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1418 |  | GX-4 6 colors | 815391 | B3 was found E1002 and E4009 | Initial Troubleshooting | Replacing Module B1 with SN: 668567 to SN: 634834 | 3 modules working and 1 defective module |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520816136475028817/5520816136475028817_signature_28.png | Lourdes Diola | https://www.jotform.com/uploads/macarewecare/222927978372067/5520816136475028817/5520816136475028817_signature_31.png
| 2023-02-13 | 2023-02-13 14:56 | CEBU TB RERENCE LAB | Cebu city | 7 | Cebu | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23400 | PBSP | asmantiquilla.pbsp@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 1695 |  | GX IV | 838882 | Due date for Xpertcheck | Perform xpertcheck | Perform xpertcheck, clean fan filter | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520810818142197494/5520810818142197494_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5520810818142197494/5520810818142197494_signature_31.png
| 2023-02-13 | 2023-02-13 14:51 | Toledo city health office | Toledo city | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Jovy Noel |  |  |  |  | 23399 | PBSP | maryjune1977@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1186 | 2023-02-13 | HP Z230 | SGH530PDFQ | Installation | Install CPU and Barcode Scanner | Installation | CPU and Barcode functional ok to use. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520807702717989476/5520807702717989476_signature_28.png | Mary June Elemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5520807702717989476/5520807702717989476_signature_31.png
| 2023-02-13 | 2023-02-10 16:21 | Toledo cho | Toledo city | 7 | Cebu | Laboratory | YES | Pullout | CEPHEID | Jovy Noel |  |  |  |  | 23398 | PBSP | maryjune1977@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 11186 | 2023-02-10 | HP Z230 | SGH530PDFQ | Barcode Scanner Not Detected | Pull out | Pull out CPU and Barcode | Not ok |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520802522713184989/5520802522713184989_signature_28.png | Mary June Elemino | https://www.jotform.com/uploads/macarewecare/222927978372067/5520802522713184989/5520802522713184989_signature_31.png
| 2023-02-13 | 2023-02-13 13:58 | Mini Forest Health Center - District C | Quezon Boulevard, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23397 | DOH | miniforestlaboratory@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF, COUNT: 228 ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 | 2023-02-13 | GX-IV 4D | 838507 | Module B4 Encounter an Error For Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters Generate IQR, Assay Stat and System Log | Perform Preventive Maintenance | Module Before for under observation Other Modules are Operational | N/A | N/A | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520773564211445643/5520773564211445643_signature_28.png | Joy Nio | https://www.jotform.com/uploads/macarewecare/222927978372067/5520773564211445643/5520773564211445643_signature_31.png
| 2023-02-13 | 2023-02-10 15:23 | Plaridel Rural Health Unit IV | Plaridel, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23396 | DOH | Yelle0908@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 198 ASSAY: Xpert MTB/RIF Ultra, COUNT: 579 | 2023-01-19 | GX IV | 831711 | Unit for xpertcheck | Performed xpertcheck | #NAME? | Unit is operational. All modules passed calibration. |  | xpertcheck kit | 9516 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610271212636397/5520610271212636397_signature_28.png | Ahlelei G. Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610271212636397/5520610271212636397_signature_31.png
| 2023-02-13 | 2023-02-09 14:16 | Marulas 3S Health Center | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23392 | DOH | melissamartin0828@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 314 ASSAY: Xpert MTB/RIF Ultra, COUNT: 815 | 2023-01-18 | GXIV | 831696 | Unit for xpertcheck | Performed xpertcheck | #NAME? | Unit is operational. All modules passed calibration. |  | Xpertcheck Kit | 9495 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231212252309/5520610231212252309_signature_28.png | Melissa Martin | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231212252309/5520610231212252309_signature_31.png
| 2023-02-13 | 2023-02-08 15:22 | Bagong Barrio Health Center | Bagong Barrio, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23393 | PBSP | cruzernitamay@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 129 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1057 |  | GXIV | 831906 | Unit for relocation | Relocated unit to suitable area | #NAME? | Unit is operational. UPS for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231213450983/5520610231213450983_signature_28.png | Ernita May Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231213450983/5520610231213450983_signature_31.png
| 2023-02-13 | 2023-02-09 11:43 | Punturin 3S Health Station | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23395 | DOH | lukegonzaga09@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 184 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1018 | 2023-01-18 | GXIV | 831914 | Unit for xpertcheck | Performed xpertcheck | #NAME? | Unit is operational. B1 failed. B2, B4 passed calibration. |  | XpertCheck Kit | 9494 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231214506471/5520610231214506471_signature_28.png | Judelyn Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231214506471/5520610231214506471_signature_31.png
| 2023-02-13 | 2023-02-08 12:47 | Camarin Health Center | Camarin, Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23394 | PBSP | jasminfalala@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 24 ASSAY: Xpert MTB/RIF Ultra, COUNT: 790 | 2023-02-06 | GXIV | 815395 | Module B2 (745915) for replacement | Replaced module B2 | #NAME? | Unit is operational. |  | Genexpert Module SN: 670428 | G 5163 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231215107510/5520610231215107510_signature_28.png | Arianne Jasmin Resurreccion | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231215107510/5520610231215107510_signature_31.png
| 2023-02-13 | 2023-02-10 13:01 | Jose B. Lingad Memorial General Hospital | San Fernando, Pampanga | 3 | Pampanga | Histopath | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23393 | COMMERCIAL | jblmgh.is@gmail.com | 25 | ASSAY: Xpert HBV VL, COUNT: 49 ASSAY: Others, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 28 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1183 | 2023-02-10 | GXIV | 812571 | Unit for xpertcheck | Performed xpertcheck | - installed XPC 1.5a - checked & cleaned cartridge bay, plunger rod, I core slot, filter - performed xpertcheck - Generated reports - performed functionality test | Unit is operational. A2, A3, A4 insufficient kit. A1 failed xpertcheck. |  | XpertCheck Kit | 9578 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231219966990/5520610231219966990_signature_28.png | Michelle Sitchon | https://www.jotform.com/uploads/macarewecare/222927978372067/5520610231219966990/5520610231219966990_signature_31.png
| 2023-02-13 | 2023-02-10 10:45 | Cardinal Santos Medical Center | Wilson St. | NCR | NCR 2nd - San Juan | Bloodbank | NO | MAINTENANCE | HAEMONETICS | John Monsour Doctor |  |  |  |  | 23390 | COMMERCIAL | ftejada@csmc.ph | 24 |  |  | MCS+ | 08M461 | For maintenance | Performed maintenance | Performed maintenance Checked pump rollers Verified calibration parameters Simulation run. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569471212372604/5520569471212372604_signature_28.png | Fidel Tejada | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569471212372604/5520569471212372604_signature_31.png
| 2023-02-13 | 2023-02-10 13:28 | Rosario Super Health Center | Dr. Sixto Antonio Ave | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23391 | WHO | eden.cawayan@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 278 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1119 ASSAY: Xpert SARS-CoV-2, COUNT: 113 |  | GX-IV-D | 843924 | Module B1/638232 has error 1018 | Run reliability test | Performed maintenance Run reliability test on module B1 Generated reports | Module B1/638232 not operational. Modules B2, B3 and B4 operational. | Module B1 disabled. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569471214077093/5520569471214077093_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569471214077093/5520569471214077093_signature_31.png
| 2023-02-13 | 2023-02-07 11:42 | Rizal Medical Center | Shaw Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23385 | PBSP | molecularlab@rmc.doh.gov.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 159 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2726 ASSAY: Xpert SARS-CoV-2, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 29 | 2023-02-06 | GX-IV-D | 814374 | Modules B2 and B3 for replacement. | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. |  | Old B2/601366, B3/673989 New B2/210068691, B3/210068672 | G 5159 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431212737048/5520569431212737048_signature_28.png | JUSTINE LORENZ FALZADO, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431212737048/5520569431212737048_signature_31.png
| 2023-02-13 | 2023-02-10 10:17 | Cardinal Santos Medical Center | Wilson St. | NCR | NCR 2nd - San Juan | Microbiology | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23388 | COMMERCIAL | egatchalian@csmc.ph | 24 | ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert HPV, COUNT: 13 ASSAY: Xpert MTB/RIF, COUNT: 67 ASSAY: Xpert MTB/RIF Ultra, COUNT: 758 ASSAY: Others, COUNT: 9 |  | GX-IV-D | 110003714 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431214965147/5520569431214965147_signature_28.png | Ellaine Marie Velasco-Gatchalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431214965147/5520569431214965147_signature_31.png
| 2023-02-13 | 2023-02-07 12:45 | Rizal Medical Center | Shaw Blvd | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23386 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2 ASSAY: Xpert SARS-CoV-2, COUNT: 15864 ASSAY: Others, COUNT: 9 |  | GX-XVI-D | 838627 | Modules A1 and B4 error 4012 | Run XpertCheck on modules A1 and B4 | Performed maintenance Run XpertCheck on modules A1 and B4 Generated reports | Unit is operational. Modules A1/860914 and B4/860924 not operational. | Modules A1 and B4 failed in XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431215242255/5520569431215242255_signature_28.png | JUSTINE LORENZ V. FALZADO | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431215242255/5520569431215242255_signature_31.png
| 2023-02-13 | 2023-02-09 13:38 | Claro M. Recto Memorial District Hospital | Infanta | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23389 | PBSP | katelynaveryumerez@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 245 ASSAY: Xpert MTB/RIF Ultra, COUNT: 775 |  | GX-IV-D | 807608 | Modules C1 and C4 for replacement. | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. | Old C1/623870, C4/808388 New C1/743403, C4/210068710 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431218316411/5520569431218316411_signature_28.png | Lorelyn U. Miras | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431218316411/5520569431218316411_signature_31.png
| 2023-02-13 | 2023-02-07 15:35 | Batis Health Center | Batis | NCR | NCR 2nd - San Juan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23387 | DOH | almacenaio16@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 673 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1056 ASSAY: Xpert HIV-1 Viral Load, COUNT: 56 | 2023-01-20 | GX-IV-D | 831598 | For XpertCheck | Performed XpertCheck | Performed maintenance Run XpertCheck Generated reports | Modules A1 and A3 not operational. Module A2/604349 service required. Module A4 passed and operational. | Modules A1/754112 and A3/677100 reported already. Rosario | XpertCheck-CE-5 | DR# 9533 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431219898488/5520569431219898488_signature_28.png | Rosario T. Almacen | https://www.jotform.com/uploads/macarewecare/222927978372067/5520569431219898488/5520569431219898488_signature_31.png
| 2023-02-12 | 2023-02-12 14:15 | Nasud MHO | Basud Camarines Norte | 5 | Camarines Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 23384 | PBSP | gbechano@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 314 | 2023-02-12 | GX4 | 807600 | For Module replacement. For CPU Replacement | Replaced Module A4 | Performed module repkacement. Perforned CPU replacement. Performed maintenance. | All 4 modules operational. |  | New module 210070836 CPU SN: CHP1121F28280. | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5519950492424941659/5519950492424941659_signature_28.png | Grace B. Echano | https://www.jotform.com/uploads/macarewecare/222927978372067/5519950492424941659/5519950492424941659_signature_31.png
| 2023-02-11 | 2023-02-10 11:06 | Bay Rural Health Unit | Bay Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23382 | PBSP | princessanncatungos@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 490 | 2023-02-10 | GXIV | 838132 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B4 (806709) service required. |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518945642114685047/5518945642114685047_signature_28.png | Princess Ann Catungos | https://www.jotform.com/uploads/macarewecare/222927978372067/5518945642114685047/5518945642114685047_signature_31.png
| 2023-02-11 | 2023-02-10 14:20 | Los Banos Rural Health Unit | Los Banos Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23383 | PBSP | rmbtaguic@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 586 |  | GXIV | 804716 | Unit for preventive maintenance. For checking of errors via jotform. | Performed preventive maintenance. Monitored running for checking if error will persist. Generated reports. | Prevntive maintenance | Unit is operational. No errors after running. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518945642119683112/5518945642119683112_signature_28.png | Rhea Marie Dorado | https://www.jotform.com/uploads/macarewecare/222927978372067/5518945642119683112/5518945642119683112_signature_31.png
| 2023-02-11 | 2023-02-10 14:20 | Los Banos Rural Health Unit | Los Banos Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23380 | PBSP | rmbtaguic@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 49 ASSAY: Xpert MTB/RIF Ultra, COUNT: 586 |  | GXIV | 804716 | Unit for preventive maintenance. For checking of errors via jotform. | Performed preventive maintenance. Monitored running for checking if error will persist. Generated reports. | Prevntive maintenance | Unit is operational. No errors after running. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518945212111683886/5518945212111683886_signature_28.png | Rhea Marie Dorado | https://www.jotform.com/uploads/macarewecare/222927978372067/5518945212111683886/5518945212111683886_signature_31.png
| 2023-02-11 | 2023-02-10 11:06 | Bay Rural Health Unit | Bay Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23381 | PBSP | princessanncatungos@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 490 | 2023-02-10 | GXIV | 838132 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B4 (806709) service required. |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518945212111943171/5518945212111943171_signature_28.png | Princess Ann Catungos | https://www.jotform.com/uploads/macarewecare/222927978372067/5518945212111943171/5518945212111943171_signature_31.png
| 2023-02-10 | 2023-02-10 14:24 | Talavera General Hospital | Poblacion, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23379 | MACARE-HBDC | talaverashope@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 144 ASSAY: Xpert HIV-1 Viral Load, COUNT: 51 | 2023-02-10 | Genexpert IV | 110013773 | N/a | N/a | Performed Maintenance (Cleaned Icore, Cleaned Plunger Rod, Cleaned Cartridge Bay, Cleaned Filters) Gathered Reports Installed HCV Assay Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518257263912348461/5518257263912348461_signature_28.png | Fitz Gerald Manuzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5518257263912348461/5518257263912348461_signature_31.png
| 2023-02-10 | 2023-02-10 14:24 | Talavera General Hospital | Poblacion, Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23378 | MACARE-HBDC | talaverashope@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 144 ASSAY: Xpert HIV-1 Viral Load, COUNT: 51 | 2023-02-10 | Genexpert IV | 110013773 | N/a | N/a | Performed Maintenance (Cleaned Icore, Cleaned Plunger Rod, Cleaned Cartridge Bay, Cleaned Filters) Gathered Reports Installed HCV Assay Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518257183912134896/5518257183912134896_signature_28.png | Fitz Gerald Manuzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5518257183912134896/5518257183912134896_signature_31.png
| 2023-02-10 | 2023-02-10 16:20 | Philippine heart center | East ave qc | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23377 | PBSP | phcmolecularlab@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 263 ASSAY: Others, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 2074 | 2023-02-06 | Gx4 | 834443 | For maintenance and module replacement, b2 210067784 - 210068732 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance and module replacement | B1 unable to detect b2 b3 b4 operational |  | 900gx | G5157 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518251145128736647/5518251145128736647_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5518251145128736647/5518251145128736647_signature_31.png
| 2023-02-10 | 2023-02-10 16:20 | Philippine heart center | East ave qc | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23376 | PBSP | phcmolecularlab@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 263 ASSAY: Others, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 2074 | 2023-02-06 | Gx4 | 834443 | For maintenance and module replacement, b2 210067784 - 210068732 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance and module replacement | B1 unable to detect b2 b3 b4 operational |  | 900gx | G5157 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518251065127249489/5518251065127249489_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5518251065127249489/5518251065127249489_signature_31.png
| 2023-02-10 | 2023-02-10 16:30 | Philippine heart center | East ave qc | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23375 | MACARE-HBDC | phcmolecularlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 263 ASSAY: Others, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 2074 |  | Gx4 | 804260 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518249615125899773/5518249615125899773_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5518249615125899773/5518249615125899773_signature_31.png
| 2023-02-10 | 2023-02-10 16:30 | Philippine heart center | East ave qc | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23374 | MACARE-HBDC | phcmolecularlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 11 ASSAY: Xpert MTB/RIF Ultra, COUNT: 263 ASSAY: Others, COUNT: 4 ASSAY: Xpert SARS-CoV-2, COUNT: 2074 |  | Gx4 | 804260 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518249495122110931/5518249495122110931_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5518249495122110931/5518249495122110931_signature_31.png
| 2023-02-10 | 2023-02-10 15:13 | Bicol Regional Blood Center | BRTTH Compound Legaspi City Albay | 5 | Albay | Blood Bank | NO | CALIBRATION TROUBLESHOOTING | HAEMONETICS | Harold Toco |  |  |  |  | 23373 | COMMERCIAL | bslsnrl.doh@gmail.com | 25 |  | 2023-02-10 | MCS+ | 15H032G | Centrifuge bowl not detected.. | Performed and verify calibation parameters. Checked centrifuge.. | Performed calibration.. Checked centrifuge performance. Run simulation test.. Performed troubleshooting. | Unit is operational. |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518231050501842327/5518231050501842327_signature_28.png | Verson Olicia | https://www.jotform.com/uploads/macarewecare/222927978372067/5518231050501842327/5518231050501842327_signature_31.png
| 2023-02-10 | 2023-02-10 14:44 | Bicol South Luzon Sub National Reference Lab | BRTTH Compound Legaspi City Albay | 5 | Albay | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 23372 | PBSP | rbalzaga.pbsp@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 97 ASSAY: Xpert MTB/RIF Ultra, COUNT: 512 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 | 2023-02-10 | Gx4 | 808636 | Modules B1 and B3 for replacement | Replaced modules B1 and B3 | Replaced modules B1 And B3 Performed maintwnance. | All 4 modules operational |  | 862526 210068731 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518212030506574678/5518212030506574678_signature_28.png | Richard B. Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5518212030506574678/5518212030506574678_signature_31.png
| 2023-02-10 | 2023-02-10 13:23 | MALAY MUNICIPAL HEALTH OFFICE | Malay, Aklan | 6 | Aklan | LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23371 | DOH | rhedal30@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 360 ASSAY: Xpert MTB/RIF Ultra, COUNT: 338 |  | GX-IV | 838457 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518161097275933086/5518161097275933086_signature_28.png | Redal D Taytay | https://www.jotform.com/uploads/macarewecare/222927978372067/5518161097275933086/5518161097275933086_signature_31.png
| 2023-02-10 | 2023-02-09 18:56 | Sta lucia health center | Gomez st novaliches qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23370 | DOH | josie_dingle2004@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 232 ASSAY: Xpert MTB/RIF Ultra, COUNT: 825 | 2023-01-20 | Gx4 | 831901 | For xpertcheck and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpertcheck | Xpertcheck and maintenance | A1 failed xpertcheck,A2 reported already,A3,A4 passed |  | Xpertcheck | 9535 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518134466824127963/5518134466824127963_signature_28.png | Josephine Dingle | https://www.jotform.com/uploads/macarewecare/222927978372067/5518134466824127963/5518134466824127963_signature_31.png
| 2023-02-10 | 2023-02-09 19:02 | FEU -NRMF | Dahlia ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23369 | CONSORTIUM | kahlilsilvestre92@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 22 ASSAY: Xpert MTB/RIF Ultra, COUNT: 316 |  | Gx4 | 840975 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518134446825833671/5518134446825833671_signature_28.png | Rienalyn Taripe | https://www.jotform.com/uploads/macarewecare/222927978372067/5518134446825833671/5518134446825833671_signature_31.png
| 2023-02-10 | 2023-02-09 03:37 | Don Modesto Formilleza Sr. Memorial Hospital | Looc | 4B | Romblon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23368 | PBSP | dmfsmh92@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 567 |  | GXIV-D-2 | 835089 | Module 1 Failed in XpertCheck | Module Replacement | Performed Maintenance Replaced module 1 Generate Reports | All modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130712221924860/5518130712221924860_signature_28.png | Denzel Kent M. Pasion | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130712221924860/5518130712221924860_signature_31.png
| 2023-02-10 | 2023-02-07 12:21 | San Lazaro Hospital | Alvarez, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23365 | PBSP | eisel21@yahoo.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1608 ASSAY: Xpert HIV-1 Viral Load, COUNT: 921 |  | GXIV-D | 707843 | Unit is due for XpertCheck | XpertCheck Maintenance | Performed Maintenance Performed XpertCheck Generate reports | 2 modules Failed, 1 module passed,1 module unavailable in XpertCheck |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130702221237620/5518130702221237620_signature_28.png | Leslie Grace Bundalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130702221237620/5518130702221237620_signature_31.png
| 2023-02-10 | 2023-02-07 12:03 | San Lazaro Hospital | Alvarez, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23366 | PBSP | eisel21@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1608 ASSAY: Xpert HIV-1 Viral Load, COUNT: 921 | 2023-02-06 | GXIV-D | 804766 | Unit is due for module replacement | Module replacement | Performed Module Replacement Performed self test Generate reports | 1 module newly replaced | B3:685262--> 662573 | GX Module 6-C | G 5160 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130702221630050/5518130702221630050_signature_28.png | Leslie Grace Bundalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130702221630050/5518130702221630050_signature_31.png
| 2023-02-10 | 2023-02-07 15:17 | Lingap Kalusugan City Health Unit I- Alima | Alima,Bacoor | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23367 | PBSP | eng112803@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1105 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 |  | GXIV-D | 838185 | Unut is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130702228030596/5518130702228030596_signature_28.png | Aide E. Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130702228030596/5518130702228030596_signature_31.png
| 2023-02-10 | 2023-02-03 11:06 | Trece Martires City Health Office | Trece Martires | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23363 | PBSP | loricavivian@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1244 |  | GXIV-D | 838160 | Unit is due for Maintenance | Maintenance | Performed Maintenance Generate Reports | 3 modules operational, 1 defective | Last XpertCheck: October 14, 2022 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652221624792/5518130652221624792_signature_28.png | Vivian Lorica | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652221624792/5518130652221624792_signature_31.png
| 2023-02-10 | 2023-01-31 08:58 | UST Hospital | Dapitan | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23364 | CONSORTIUM | ustimp@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 674 |  | GXIV-D | 110000711 | Unit is due for Maintenance | Maintenance | Cleaned Plunger Rod, Cartridge Bay, and Filter Performed Self-Test Generate Report File | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652226239661/5518130652226239661_signature_28.png | Lourdes fajardo | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652226239661/5518130652226239661_signature_31.png
| 2023-02-10 | 2023-02-02 10:23 | Imus City Health Office III | Green gate Subd, Imus | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23361 | PBSP | marylainemartinez@yahoo.com.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 904 |  | GXIV-D | 838166 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports | 3 modules passed, 1 module failed | XpertCheck: B1/807095 - Failed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652226365342/5518130652226365342_signature_28.png | Mary laine martinez | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652226365342/5518130652226365342_signature_31.png
| 2023-02-10 | 2023-01-31 11:26 | Tunasan Health Center | Tunasan | NCR | NCR 4th - Muntinlupa | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23360 | PBSP | gigietan0911@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1078 ASSAY: Xpert HIV-1 Viral Load, COUNT: 145 |  | GXIV-D | 807613 | Bios Battery Problem Unit is due for XpertCheck | Battery Replacement XpertCheck | Performed Maintenance Performed XpertCheck Replaced Bios Battery Generate Reports | 3 Modules Passed, 1 Failed | XpertCheck: FAILED- 662483/B3 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652226837131/5518130652226837131_signature_28.png | Geraldine Trazo-Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652226837131/5518130652226837131_signature_31.png
| 2023-02-10 | 2023-02-01 16:25 | Jose Reyes Memorial Medical Center | Rizal Ave. | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23362 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 881 ASSAY: Xpert SARS-CoV-2, COUNT: 3160 |  | GX16-D-10 Colors | 110013856 | 4008 Problem | Troubleshoot | Performed Cepheid Troubleshooting Generate Reports Capture ohms value and gateway board sides | 5 modules not operational. 11 modules operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652229401796/5518130652229401796_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5518130652229401796/5518130652229401796_signature_31.png
| 2023-02-10 | 2023-02-10 10:49 | International organization for migration | HV dela costa makati city | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23359 | COMMERCIAL | jsabalboro@iom.int | 25 | ASSAY: Xpert MTB/RIF, COUNT: 54 ASSAY: Xpert MTB/RIF Ultra, COUNT: 471 ASSAY: Xpert SARS-CoV-2, COUNT: 54 |  | Gx4 | 805288 | For network in 1 lapto and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gatheredreports,test run unit performed network | Maintenance and network | B2,b3 failed xpertcheck,b1,b4 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518098943026530397/5518098943026530397_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5518098943026530397/5518098943026530397_signature_31.png
| 2023-02-10 | 2023-02-10 10:49 | International organization for migration | HV dela costa makati city | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23358 | COMMERCIAL | jsabalboro@iom.int | 25 | ASSAY: Xpert MTB/RIF, COUNT: 54 ASSAY: Xpert MTB/RIF Ultra, COUNT: 471 ASSAY: Xpert SARS-CoV-2, COUNT: 54 |  | Gx4 | 805288 | For network in 1 lapto and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gatheredreports,test run unit performed network | Maintenance and network | B2,b3 failed xpertcheck,b1,b4 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518098853026267792/5518098853026267792_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5518098853026267792/5518098853026267792_signature_31.png
| 2023-02-10 | 2023-02-10 10:56 | International organization for migration | H.V dela costa makati city | NCR | NCR 4th - Makati | Labiratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23357 | COMMERCIAL | jsabalboro@iom.int | 28 | ASSAY: Xpert MTB/RIF, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 691 ASSAY: Xpert SARS-CoV-2, COUNT: 157 |  | Gx4 | 110002868 | For maintenance and network in 1 laptop | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed network | Maintenance and network | Operationak |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518097703026199716/5518097703026199716_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5518097703026199716/5518097703026199716_signature_31.png
| 2023-02-10 | 2023-02-10 10:56 | International organization for migration | H.V dela costa makati city | NCR | NCR 4th - Makati | Labiratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23356 | COMMERCIAL | jsabalboro@iom.int | 28 | ASSAY: Xpert MTB/RIF, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 691 ASSAY: Xpert SARS-CoV-2, COUNT: 157 |  | Gx4 | 110002868 | For maintenance and network in 1 laptop | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed network | Maintenance and network | Operationak |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518097633025679991/5518097633025679991_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5518097633025679991/5518097633025679991_signature_31.png
| 2023-02-10 | 2023-02-10 11:32 | International organization for migration | H.V dela costa makati city | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23355 | COMMERCIAL | jsabalboro@iom.int | 28 | ASSAY: Xpert MTB/RIF, COUNT: ASSAY: Xpert MTB/RIF Ultra, COUNT: ASSAY: Xpert SARS-CoV-2, COUNT: |  | Gx4 | 110004816 | For maintenance and network | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,yest run unit,performed network | Maintenance and network | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518096703028779559/5518096703028779559_signature_28.png | Joy Sabalboro | https://www.jotform.com/uploads/macarewecare/222927978372067/5518096703028779559/5518096703028779559_signature_31.png
| 2023-02-10 | 2023-02-09 11:06 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23354 | PBSP | 2052456@aup.edu.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 166 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1654 | 2023-01-09 | GXIV | 807570 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules B1(624926) B2(654606) B3(684919) are service required. | Used 2 spare xpertcheck | Xpertcheck | 37930 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050850524088137/5518050850524088137_signature_28.png | Mary Jane Lourdes Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050850524088137/5518050850524088137_signature_31.png
| 2023-02-10 | 2023-02-09 11:06 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23353 | PBSP | 2052456@aup.edu.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 166 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1654 | 2023-01-09 | GXIV | 807570 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules B1(624926) B2(654606) B3(684919) are service required. | Used 2 spare xpertcheck | Xpertcheck | 37930 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050850524088137/5518050850524088137_signature_28.png | Mary Jane Lourdes Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050850524088137/5518050850524088137_signature_31.png
| 2023-02-10 | 2023-02-09 11:01 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23352 | DOH | 2052456@aup.edu.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 144 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1405 | 2023-02-09 | GXIV | 838225 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules B3(807734) and B4(807732) are service required. |  | Xpertcheck kot | DR |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050900527749581/5518050900527749581_signature_28.png | Mary Jane Lourdes Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050900527749581/5518050900527749581_signature_31.png
| 2023-02-10 | 2023-02-09 11:06 | Calamba Ciry Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23351 | PBSP | 2052456@aup.edu.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 166 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1654 | 2023-01-09 | GXIV | 807570 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules B1(624926) B2(654606) B3(684919) are service required. | Used 2 spare xpertcheck | Xpertcheck | 37930 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050850524088137/5518050850524088137_signature_28.png | Mary Jane Lourdes Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050850524088137/5518050850524088137_signature_31.png
| 2023-02-10 | 2023-02-09 13:39 | Cabuyao City Health Office 1 | Cabuyao Laguna | 4A | Laguna | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23350 | PBSP | shirleypalcantara18@gmail.com | 29 | ASSAY: , COUNT: 154 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1249 | 2023-02-06 | GXIV | 812794 | Unit for preventive maintenance. For module replacement of modules B1 & B4. For checking of monitor. | Performed preventive maintenance. Performed module replacement on module B1 & B4. Performed checking of monitor. Generated reports. | Preventive maintenance, module replacement and checking. | Unit is operational. 3 operational modules. Monitor HP ProDisplay P19A (CNK6291216) is lagging on bottom part and is recommended for replacement |  | GX Module for B1 (721198) and B4 (889236) in replacement for defective B1 (684186) and B4 (640819). | G |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050340523182770/5518050340523182770_signature_28.png | Shirley Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5518050340523182770/5518050340523182770_signature_31.png
| 2023-02-10 | 2023-02-10 10:03 | Hinaplanon Health Center | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23349 | DOH | villacortaamy50@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 785 |  | GXIV-4D | 838227 | Due for quarter PM | None | #NAME? | All four modules are operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5518040827766942557/5518040827766942557_signature_28.png | Amy S. Villacorta | https://www.jotform.com/uploads/macarewecare/222927978372067/5518040827766942557/5518040827766942557_signature_31.png
| 2023-02-09 | 2023-02-09 15:53 | Tomas Cabili Health Center | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23348 | DOH | mariacielovenezuela@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 819 |  | GXIV-4D | 838294 | Due for quarterly PM | NA | #NAME? | All four modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517386475517551864/5517386475517551864_signature_28.png | Maria Cielo Venezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5517386475517551864/5517386475517551864_signature_31.png
| 2023-02-09 | 2023-02-09 15:02 | Lung center of the philippines | East ave qc | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 23346 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | Mcs+ | 08H188 | For calibration and maintenance | Cleaned unit,check valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | Calibration and maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517383941914454898/5517383941914454898_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/5517383941914454898/5517383941914454898_signature_31.png
| 2023-02-09 | 2023-02-09 15:31 | Lung center of the philippines | East ave qc | NCR | NCR 2nd - Quezon City | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 23347 | MACARE-HBDC | rhonababilonia_19@yahoo.com | 26 |  |  | Mcs + | 02D166 | For calibration and maintenance | Cleaned unit,checked valve,rollers,line sensors,display panel,dpm,spm,cuff,performed calibration ,performed functionality test | Calibration and maintenance | Operational | Without protocol card |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517383941915113567/5517383941915113567_signature_28.png | Rhona Babilonia | https://www.jotform.com/uploads/macarewecare/222927978372067/5517383941915113567/5517383941915113567_signature_31.png
| 2023-02-09 | 2023-02-09 14:54 | Bureau of Plant Industry Manila | Quirino Avenue, Manila | NCR | NCR 1st - Manila | NPQSD | NO | INSTALLATION | CARL ZEISS | Patrick Los Baños |  |  |  |  | 23345 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 26 |  | 2023-02-09 | Primo Star | 3116003729 | Accessories for installation | Installed accessories | - installed binocular with phototube - installed Axiocam 208C - performed functionality test | Unit is operational |  | -binocular with phototube -Axiocam 208C (S/N: 5321003334> | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517356220027390917/5517356220027390917_signature_28.png | Jesriel escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/5517356220027390917/5517356220027390917_signature_31.png
| 2023-02-09 | 2023-02-09 14:54 | Bureau of Plant Industry Manila | Quirino Avenue, Manila | NCR | NCR 1st - Manila | NPQSD | NO | INSTALLATION | CARL ZEISS | Patrick Los Baños |  |  |  |  | 23344 | COMMERCIAL | npqsd.laboratoryprocurement@gmail.com | 26 |  | 2023-02-09 | Primo Star | 3116003729 | Accessories for installation | Installed accessories | - installed binocular with phototube - installed Axiocam 208C - performed functionality test | Unit is operational |  | -binocular with phototube -Axiocam 208C (S/N: 5321003334> | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517356200025828656/5517356200025828656_signature_28.png | Jesriel escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/5517356200025828656/5517356200025828656_signature_31.png
| 2023-02-09 | 2023-02-09 14:32 | Mindanao Central Sanitarium | Pasabolong, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23343 | DOH | laboratory@mcsgh.com.ph | 19 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1760 | 2023-02-09 | GX 4 | 821214 | Module 1 and 2 temperature errors | Preventive maintenance | Preventive maintenance | Module 1 and 2 disabled due to excessive temperature errors | Module 1 and 2 already reported way before 2022 | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517347399661336449/5517347399661336449_signature_28.png | Amie B. YBIO | https://www.jotform.com/uploads/macarewecare/222927978372067/5517347399661336449/5517347399661336449_signature_31.png
| 2023-02-09 | 2023-02-09 12:57 | Manolo Fortich Rural Health Unit | Manolo Fortich | 10 | Bukidnon | RTDL | YES | REPAIR | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23342 | PBSP | mfrhutbdots@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 753 |  | GeneXpert | 815286 | Error 2126 | Reset power supply and gateway board | Preventive maintenance Power supply reset Gateway board reset | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517344989763575612/5517344989763575612_signature_28.png | Elaine Chung | https://www.jotform.com/uploads/macarewecare/222927978372067/5517344989763575612/5517344989763575612_signature_31.png
| 2023-02-09 | 2023-02-09 13:29 | Sto Domingo District Hospital | Sto Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23341 | DOH | acharizaann@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 249 ASSAY: Xpert MTB/RIF Ultra, COUNT: 670 | 2023-02-09 | Genexpert IV | 831582 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Unit is for Calibration this March | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517301977116766879/5517301977116766879_signature_28.png | Chariza Ann T.Aluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5517301977116766879/5517301977116766879_signature_31.png
| 2023-02-09 | 2023-02-09 13:40 | Fairview health center | Dahlia ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 23340 | PBSP | hiromiomallio@yahoo.com.sg | 25 |  |  | Gx4 | 838876 | For maintenance and xpertcheck | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit and performed xpertcheck | Maintenance and xpertcheck | B3 and b4 failed xpertcheck,b2 retry,b1 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517295572847135366/5517295572847135366_signature_28.png | Hiromi O. Leandado | https://www.jotform.com/uploads/macarewecare/222927978372067/5517295572847135366/5517295572847135366_signature_31.png
| 2023-02-09 | 2023-02-09 13:09 | Obrero Main Health Center | Brgy. Obrero, Butuan City | 13 | Agusan del Norte | RTDL | NO | TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 23339 | DOH | dorothygundaya@gmail.com | 23.775 | ASSAY: Xpert MTB/RIF, COUNT: 708 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1515 ASSAY: Xpert HIV-1 Viral Load, COUNT: 47 |  | GX-4 | 838403 | B1 was found Error 4009 B3 was found Error 1002 B4 was found Error 1002 | Troubleshooting | XpertCheck Calibration | 1 module failed from XpertCheck (B4) 1 module failed from self-test (B1) but passed xpc 1 module for observation | 2 modules working, 2 modules for replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517293840229279753/5517293840229279753_signature_28.png | Dorothy Gundaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5517293840229279753/5517293840229279753_signature_31.png
| 2023-02-09 | 2023-02-09 12:42 | Malungon Rural Health Unit | Malungon, Sarangani province | 12 | Sarangani | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23338 | DOH | ulbenarioe@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 230 |  | Gx -IV | 838546 | Genexpert machine CPU Problem | Perform Genexpert machine CPU Troubleshooting | Check wirings- CPU Power Cord functional Replace CMOS battery Unplug and change ram to other Pin Check CPU Fun and lighting - none functional | Genexpert machine CPU is for replacement due to power supply or mother board problem |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517283538917348667/5517283538917348667_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5517283538917348667/5517283538917348667_signature_31.png
| 2023-02-09 | 2023-02-09 12:42 | Malungon Rural Health Unit | Malungon, Sarangani province | 12 | Sarangani | Laboratory | YES | TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23337 | DOH | ulbenarioe@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 230 |  | Gx -IV | 838546 | Genexpert machine CPU Problem | Perform Genexpert machine CPU Troubleshooting | Check wirings- CPU Power Cord functional Replace CMOS battery Unplug and change ram to other Pin Check CPU Fun and lighting - none functional | Genexpert machine CPU is for replacement due to power supply or mother board problem |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517283518919863338/5517283518919863338_signature_28.png | Ermie Ulbenario | https://www.jotform.com/uploads/macarewecare/222927978372067/5517283518919863338/5517283518919863338_signature_31.png
| 2023-02-09 | 2023-02-09 11:06 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23336 | PBSP | 2052456@aup.edu.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 166 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1654 | 2023-01-09 | GXIV | 807570 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules B1(624926) B2(654606) B3(684919) are service required. | Used 2 spare xpertcheck | Xpertcheck | 37930 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517230172911735834/5517230172911735834_signature_28.png | Mary Jane Lourdes Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5517230172911735834/5517230172911735834_signature_31.png
| 2023-02-09 | 2023-02-09 11:01 | Calamba City Health Office | Calamba Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23335 | DOH | 2052456@aup.edu.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 144 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1405 | 2023-02-09 | GXIV | 838225 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules B3(807734) and B4(807732) are service required. |  | Xpertcheck kot | DR |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517229652919390411/5517229652919390411_signature_28.png | Mary Jane Lourdes Gayadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5517229652919390411/5517229652919390411_signature_31.png
| 2023-02-09 | 2023-02-09 11:05 | NABAS RURAL HEALTH UNIT | Nabas, Aklan | 6 | Aklan | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23334 | PBSP | lambito.godelyn61@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: N/A |  | GX-IV | 815387 | #NAME? | #NAME? | #NAME? | SYSTEM UNIT IS DEFECTIVE |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517216381628808973/5517216381628808973_signature_28.png | Godelyn Lambito | https://www.jotform.com/uploads/macarewecare/222927978372067/5517216381628808973/5517216381628808973_signature_31.png
| 2023-02-09 | 2023-02-09 10:38 | Davao del Norte Provincial Hospital-Igacos Samal Zone | Island Garden City of Samal | 11 | Davao del Norte | Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23333 | PBSP | jftacastacas@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1096 | 2023-02-09 | GX-IV 4D | 815383 | Module B3 and B4 for Replacement | Perform Module Replacement | Replaced the Defective Module B3/ SN 696056 with new Module SN:2100708378 Replaced the defective Module B4/SN:696057 with new Module SN: 210070837 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is fully operational | N/A | 900-0508 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517201290147011555/5517201290147011555_signature_28.png | JANICE FE PASTOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5517201290147011555/5517201290147011555_signature_31.png
| 2023-02-09 | 2023-02-09 10:38 | Davao del Norte Provincial Hospital-Igacos Samal Zone | Island Garden City of Samal | 11 | Davao del Norte | Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23332 | PBSP | jftacastacas@gmail.com | 26.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1096 | 2023-02-09 | GX-IV 4D | 815383 | Module B3 and B4 for Replacement | Perform Module Replacement | Replaced the Defective Module B3/ SN 696056 with new Module SN:2100708378 Replaced the defective Module B4/SN:696057 with new Module SN: 210070837 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is fully operational | N/A | 900-0508 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517201250143720897/5517201250143720897_signature_28.png | JANICE FE PASTOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5517201250143720897/5517201250143720897_signature_31.png
| 2023-02-09 | 2023-02-08 14:25 | Talaingod Rural Health Unit | JBL, Brgy. Sto. Niño, Talaingod, Davao del Norte | 11 | Davao del Norte | TB-DOTS | NO | REPAIR CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23331 | PBSP | denniseolayan@gmail.com | 23.7 | ASSAY: Xpert MTB/RIF, COUNT: 298 ASSAY: Xpert MTB/RIF Ultra, COUNT: 593 | 2023-02-08 | GX-IV 4D | 838859 | Module B3 for Replacement For XpertCheck Calibration and Maintenance | Perform Preventive Maintenance, XpertCheck Calibration and Module Replacement | Replaced the Defective Module B3/SN: 863141 with new Module SN: 660296 Calibrate the Machine using XpertCheck Clean Surface Area, Cartridge Bay, Fan Filters, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is operational, modules passed the Calibration | Module B4 for under observation | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517188682662689753/5517188682662689753_signature_28.png | Christine Dennise Olayan RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5517188682662689753/5517188682662689753_signature_31.png
| 2023-02-09 | 2023-02-09 10:07 | Iligan Society of Internists | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | YES | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23330 | PBSP | dayolaermajessica@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1474 | 2023-02-09 | GXIV-4D R1 | 707842 | Due for calibration | #NAME? | #NAME? | All four modules are operational | Module D1 - retry then passed, module D2 & D4 - passed, module D3 - retry then skipped | Xpertcheck CE-5 Lot No. 05502 Exp. 2024 02 11 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517180456215331466/5517180456215331466_signature_28.png | Erma Jessica Dayola | https://www.jotform.com/uploads/macarewecare/222927978372067/5517180456215331466/5517180456215331466_signature_31.png
| 2023-02-09 | 2023-02-09 09:27 | Iligan Society of Internists | Iligan City, Lanao del Norte | 10 | Lanao del Norte | TB Dots | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23329 | PBSP | dayolaermajessica@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1106 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 | 2023-02-09 | GXIV-4D | 805090 | Module C4 defective | Replace defective module C4/685784 with new module SN 802400 | #NAME? | All four modules are operational. |  | 900-0491RS | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5517155476104781942/5517155476104781942_signature_28.png | Erma Jessica Dayola | https://www.jotform.com/uploads/macarewecare/222927978372067/5517155476104781942/5517155476104781942_signature_31.png
| 2023-02-08 | 2023-02-08 15:18 | Bagong pag asa health center | Bagong pag asa qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23328 | DOH | erlindalorilla026@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 69 ASSAY: Xpert MTB/RIF, COUNT: 238 ASSAY: Xpert MTB/RIF Ultra, COUNT: 949 |  | Gx 4 | 831597 | For maintenance and encounter loss communication error A2 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit.relocate position module 2 to position 1 no error of loss communication during swapping of module position.. | Maintenance and troubleshooting | For observation module A2,A1 reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516603895327701101/5516603895327701101_signature_28.png | Erlinda Lorilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5516603895327701101/5516603895327701101_signature_31.png
| 2023-02-08 | 2023-02-08 15:18 | Bagong pag asa health center | Bagong pag asa qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23327 | DOH | erlindalorilla026@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 69 ASSAY: Xpert MTB/RIF, COUNT: 238 ASSAY: Xpert MTB/RIF Ultra, COUNT: 949 |  | Gx 4 | 831597 | For maintenance and encounter loss communication error A2 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit.relocate position module 2 to position 1 no error of loss communication during swapping of module position.. | Maintenance and troubleshooting | For observation module A2,A1 reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516603845328496946/5516603845328496946_signature_28.png | Erlinda Lorilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5516603845328496946/5516603845328496946_signature_31.png
| 2023-02-08 | 2023-02-08 09:30 | Biopath quezon city | E.rodriguez,quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23326 | COMMERCIAL | ainocentes@biopath.ph | 33 | ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 2100 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Others, COUNT: 48 |  | GXXVI | 839306 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516602065323765965/5516602065323765965_signature_28.png | Amanda inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/5516602065323765965/5516602065323765965_signature_31.png
| 2023-02-08 | 2023-02-08 09:30 | Biopath quezon city | E.rodriguez,quezon city | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23325 | COMMERCIAL | ainocentes@biopath.ph | 33 | ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 2100 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Others, COUNT: 48 |  | GXXVI | 839306 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516601885322549071/5516601885322549071_signature_28.png | Amanda inocentes | https://www.jotform.com/uploads/macarewecare/222927978372067/5516601885322549071/5516601885322549071_signature_31.png
| 2023-02-08 | 2023-02-07 15:58 | St lukes medical center | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23324 | COMMERCIAL | megapo@stlukes.com.ph | 27 | ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 226 ASSAY: Others, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2789 |  | Gx 4 | 839234 | B1 and b2 failed xpertcheck,b3 encounter intermittent temp error 4012,for repair broken door hinges b4 module | Performed maintenance,Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit, performed realibility test | Repair,maintenance,troubleshooting | B4 fixed door hinges, | All module passed reliability test |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516601805326834314/5516601805326834314_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5516601805326834314/5516601805326834314_signature_31.png
| 2023-02-08 | 2023-02-08 14:55 | Jose B. Lingad Memorial Regional Hospital | Dolores, San Fernando City Pampanga | 3 | Pampanga | Immuno Laboratory (HIV) | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23323 | MACARE-HBDC | concepcionalfonso2@gmail.com | 29 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1174 ASSAY: Xpert HCV VL, COUNT: 28 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 46 | 2023-02-08 | Genexpert IV | 813571 | Barcode Scanner defective Laptop keyboard some keys not working | Troubleshoot the scanner Reconfigure the scanner | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | Barcode scanner for replacement, Laptop keyboard not working | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516492789511436881/5516492789511436881_signature_28.png | Concepcion G. Alfonso | https://www.jotform.com/uploads/macarewecare/222927978372067/5516492789511436881/5516492789511436881_signature_31.png
| 2023-02-08 | 2023-02-08 14:55 | Jose B. Lingad Memorial Regional Hospital | Dolores, San Fernando City Pampanga | 3 | Pampanga | Immuno Laboratory (HIV) | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23322 | MACARE-HBDC | concepcionalfonso2@gmail.com | 29 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1174 ASSAY: Xpert HCV VL, COUNT: 28 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 46 | 2023-02-08 | Genexpert IV | 813571 | Barcode Scanner defective Laptop keyboard some keys not working | Troubleshoot the scanner Reconfigure the scanner | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | Barcode scanner for replacement, Laptop keyboard not working | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516492749515415372/5516492749515415372_signature_28.png | Concepcion G. Alfonso | https://www.jotform.com/uploads/macarewecare/222927978372067/5516492749515415372/5516492749515415372_signature_31.png
| 2023-02-08 | 2023-02-08 14:55 | Jose B. Lingad Memorial Regional Hospital | Dolores, San Fernando City Pampanga | 3 | Pampanga | Immuno Laboratory (HIV) | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23321 | MACARE-HBDC | concepcionalfonso2@gmail.com | 29 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1174 ASSAY: Xpert HCV VL, COUNT: 28 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 46 | 2023-02-08 | Genexpert IV | 813571 | Barcode Scanner defective Laptop keyboard some keys not working | Troubleshoot the scanner Reconfigure the scanner | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | Barcode scanner for replacement, Laptop keyboard not working | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516492079518771958/5516492079518771958_signature_28.png | Concepcion G. Alfonso | https://www.jotform.com/uploads/macarewecare/222927978372067/5516492079518771958/5516492079518771958_signature_31.png
| 2023-02-08 | 2023-02-08 14:40 | Zamboanga City Medical Center | Evangelista Street, Snta Catalina Rd, Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23320 | PBSP | zcmctblab@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 441 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4110 ASSAY: Xpert SARS-CoV-2, COUNT: 22894 ASSAY: Xpert HIV-1 Viral Load, COUNT: 240 | 2023-02-08 | GX 4 | 839309 | None | Module Replacement | Replaced module 638590, 21007 68701, 210068700, 210068678 with Modules 661531, 665298, 210068723, 210068675 | All 16 modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516487855427589083/5516487855427589083_signature_28.png | Aira Faith S. Runas | https://www.jotform.com/uploads/macarewecare/222927978372067/5516487855427589083/5516487855427589083_signature_31.png
| 2023-02-08 | 2023-02-08 12:43 | DR. RAFAEL S. TUMBOKON MEMORIAL HOSPITAL | KALIBO AKLAN | 6 | Aklan | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23319 | PBSP | gonzagaevarmt@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3132 ASSAY: Xpert HIV-1 Viral Load, COUNT: 66 ASSAY: Xpert MTB/RIF, COUNT: 1096 |  | GX-IV | 804680 | FOR MODULE D2 REPLACEMENT | -PERFORM MODULE D2 REPLACEMENT DEFECTIVE SN:655239 WITH NEW SN:210068729 | #NAME? | #NAME? | #NAME? |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516411587368366202/5516411587368366202_signature_28.png | Eva Raquel D. Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5516411587368366202/5516411587368366202_signature_31.png
| 2023-02-08 | 2023-02-08 09:33 | Asian Hospital and Medical Center | Alabang Muntinlupa | NCR | NCR 4th - Muntinlupa | Molecular Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23317 | COMMERCIAL | labteamlead@asianhospital.com | 25 | ASSAY: Others, COUNT: 32 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert MTB/RIF, COUNT: 41 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1022 ASSAY: Xpert HIV-1 Viral Load, COUNT: 20 | 2023-02-02 | GXIV | 803869 | Unit for preventive maintenance. For relocation of GeneXpert system from micro to molecular lab. For installation of UPS. | Performed relocation of GeneXpert. Performed installation of UPS. Performed preventive maintenance. Generated reports. | Installation, preventive maintenance and relocation | Unit is operational. |  | Accupower UPS (AT1000170815ST0163) in replacement to defective Emerson UPS (1435400116PS27M). | DR 9567 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516401132911806554/5516401132911806554_signature_28.png | Christine Fallaria | https://www.jotform.com/uploads/macarewecare/222927978372067/5516401132911806554/5516401132911806554_signature_31.png
| 2023-02-08 | 2023-02-07 09:15 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23318 | PBSP | microscan@tmc.doh.gov.ph | 27 | ASSAY: Xpert HPV, COUNT: 445 ASSAY: Xpert MTB/RIF, COUNT: 501 ASSAY: Xpert MTB/RIF Ultra, COUNT: 892 ASSAY: Xpert SARS-CoV-2, COUNT: 11012 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2023-02-06 | GX16 | 844121 | Unit for preventive maintenance. For module replacement of C1, C3, D2, D3, D4. | Performed preventive maintenance. Performed module replacement. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. |  | Gx Module 210070820, 210070826, 210070827, 210073563, 210073914 | G 5161 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516401132916470792/5516401132916470792_signature_28.png | Cynthia fuertes rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5516401132916470792/5516401132916470792_signature_31.png
| 2023-02-08 | 2023-02-07 13:59 | Velasquez Health Center | Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23316 | DOH | wcpatron11x4@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 453 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1532 | 2023-01-18 | GXIV | 831534 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module A1 (761593) service required. |  | Xpertcheck | DR 9500 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516401132916770136/5516401132916770136_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5516401132916770136/5516401132916770136_signature_31.png
| 2023-02-08 | 2023-02-08 11:45 | Tagbina Rural Health Unit RTDL | Purok 3, Sto. Niño, Poblacion, Tagbina | 13 | Surigao del Sur | RTDL | NO | CALIBRATION | CEPHEID | Faizar Avenido |  |  |  |  | 23315 | PBSP | ellalaguros@yahoo.com | 26.1 | ASSAY: Xpert MTB/RIF, COUNT: 65 ASSAY: Xpert MTB/RIF Ultra, COUNT: 376 |  | GX-4 | 838877 | None | None | XpertCheck Calibration | All Modules Working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516383223183404243/5516383223183404243_signature_28.png | Ellaquim L. Albarracin | https://www.jotform.com/uploads/macarewecare/222927978372067/5516383223183404243/5516383223183404243_signature_31.png
| 2023-02-08 | 2023-02-08 09:49 | Midsayap Rural Health Unit | Midsayap North Cotabato | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23314 | PBSP | oliverbeltran18rmt@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 196 ASSAY: Xpert MTB/RIF Ultra, COUNT: 439 | 2023-02-08 | Gx -IV (2 module) | 835183 | Module B1 for Service Required Module B2 for Service Required | Perform Genexpert machine Preventive maintenance Replace module B1 SN: 787983 with new module SN: 210068707 Replace module B2 SN:788463 with new module SN: 210068709 | Perform Genexpert machine Preventive maintenance Replace module B1 SN: 787983 with new module SN: 210068707 Replace module B2 SN:788463 with new module SN: 210068709 | Genexpert machine has 2 operational modules |  | Genexpert 6-colors module | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516369452427936519/5516369452427936519_signature_28.png | Oliver A. Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5516369452427936519/5516369452427936519_signature_31.png
| 2023-02-08 | 2023-02-08 11:09 | Tarlac City Health Center X | Tibag Tarlac City | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23313 | DOH | jayteecornelius@gmail.com | 28.125 | ASSAY: Xpert MTB/RIF, COUNT: 337 ASSAY: Xpert MTB/RIF Ultra, COUNT: 368 | 2023-02-01 | GXIV-D | 831587 | Due for calibration | Performed preventive maintenance Archived test results Performed calibration Gathered reports Oerformed functionality test | Preventive maintenance and calibration | Unit is operational excel modules A1 and A4 ( for replacement, reported) | A2 and A3 passed calibration | 2 Xpertcheck kits | 9563 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516367585212786385/5516367585212786385_signature_28.png | Jaytee Bachiller | https://www.jotform.com/uploads/macarewecare/222927978372067/5516367585212786385/5516367585212786385_signature_31.png
| 2023-02-08 | 2023-02-08 11:09 | Tarlac City Health Center X | Tibag Tarlac City | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23312 | DOH | jayteecornelius@gmail.com | 28.125 | ASSAY: Xpert MTB/RIF, COUNT: 337 ASSAY: Xpert MTB/RIF Ultra, COUNT: 368 | 2023-02-01 | GXIV-D | 831587 | Due for calibration | Performed preventive maintenance Archived test results Performed calibration Gathered reports Oerformed functionality test | Preventive maintenance and calibration | Unit is operational excel modules A1 and A4 ( for replacement, reported) | A2 and A3 passed calibration | 2 Xpertcheck kits | 9563 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516367425211821936/5516367425211821936_signature_28.png | Jaytee Bachiller | https://www.jotform.com/uploads/macarewecare/222927978372067/5516367425211821936/5516367425211821936_signature_31.png
| 2023-02-08 | 2023-02-07 15:43 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23311 | PBSP | rcb.macare@gmail.com | 27 | ASSAY: Xpert SARS-CoV-2, COUNT: 4947 |  | GXIV | 834623 | Temperature error on C2 (785115) | checked system log & test results to analyze concern | - checked system log - checked test results of 2023 - Generated reports | Unit is operational. Module C2 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516331442411173692/5516331442411173692_signature_28.png | Michaella B. Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5516331442411173692/5516331442411173692_signature_31.png
| 2023-02-08 | 2023-02-07 16:48 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23310 | PBSP | bmclab_micro@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA | 2023-02-06 | GXIV | 805112 | Modules D1 & D2 for replacement | Replaced D1 (687392), D4 (648734) | #NAME? | Unit is operational |  | Genexpert Module SN: 210068596, 210068588 | G 5162 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516330952414270250/5516330952414270250_signature_28.png | Marichu Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5516330952414270250/5516330952414270250_signature_31.png
| 2023-02-08 | 2023-02-07 16:43 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23309 | PBSP | bmclab_micro@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 543 ASSAY: Xpert MTB/RIF Ultra, COUNT: 860 ASSAY: Xpert HCV VL, COUNT: 3 ASSAY: Xpert HIV-1 Viral Load, COUNT: 146 | 2023-02-06 | GXIV | 807466 | Module D2 (686244) for replacement | Replaced module D2 | #NAME? | Unit is operational |  | GeneXpert Module SN: 210068621 | G 5162 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516330952415255785/5516330952415255785_signature_28.png | Marichu Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5516330952415255785/5516330952415255785_signature_31.png
| 2023-02-08 | 2023-02-07 12:27 | Meycauayan city Health Center VII | Pandayan, Meycauayan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23308 | DOH | ronamarieiris_25@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 287 ASSAY: Xpert MTB/RIF Ultra, COUNT: 876 ASSAY: Xpert HIV-1 Viral Load, COUNT: 33 | 2023-01-20 | GXIV | 831904 | Unit for calibration | Performed xpertcheck on A1, A2, A4 | #NAME? | Unit is operational. A3 is already confirmed for module replacement. A1, A4 failed xpertcheck. A2 passed. |  | Xpertcheck kit | 9524 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516330952419517286/5516330952419517286_signature_28.png | Iris De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5516330952419517286/5516330952419517286_signature_31.png
| 2023-02-08 | 2023-02-07 14:55 | Bongabon District Hospital | Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23307 | DOH | teody_santua@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 169 | 2023-02-07 | Genexpert IV | 831562 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | Unit is for Calibration this March | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5516244816027152509/5516244816027152509_signature_28.png | Teodorico S. Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5516244816027152509/5516244816027152509_signature_31.png
| 2023-02-07 | 2023-02-07 14:55 | Bongabon District Hospital | Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23306 | DOH | teody_santua@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 169 | 2023-02-07 | Genexpert IV | 831562 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | Unit is for Calibration this March | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515918726029086652/5515918726029086652_signature_28.png | Teodorico S. Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5515918726029086652/5515918726029086652_signature_31.png
| 2023-02-07 | 2023-02-07 14:55 | Bongabon District Hospital | Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23305 | DOH | teody_santua@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 169 | 2023-02-07 | Genexpert IV | 831562 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | Unit is for Calibration this March | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515918626023252247/5515918626023252247_signature_28.png | Teodorico S. Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5515918626023252247/5515918626023252247_signature_31.png
| 2023-02-07 | 2023-02-06 13:42 | Palayan CHO | Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23304 | DOH | riazarizzamae@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 304 ASSAY: Xpert MTB/RIF Ultra, COUNT: 630 | 2023-02-06 | Genexpert iV | 831637 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibration Performed Functionality Test | Unit is Operational | N/a | Xpertcheck kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515918246024192569/5515918246024192569_signature_28.png | Rizza Mae A. Riaza,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5515918246024192569/5515918246024192569_signature_31.png
| 2023-02-07 | 2023-02-07 15:20 | Pikit Rural Health Unit | Pikit, North Cotabato | 12 | Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23302 | PBSP | litosepelagior@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 157 | 2023-02-07 | Gx-IV | 835181 | Module C2 for Service Required | Perform Genexpert machine Preventive maintenance Replace module SN:788671 with new module SN: 210068730 | Perform Genexpert machine Preventive maintenance Replace module SN:788671 with new module SN: 210068730 | Genexpert machine has 2 operational modules |  | Genexpert 6-color module | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515647522956489423/5515647522956489423_signature_28.png | Lito R. Sepelagio | https://www.jotform.com/uploads/macarewecare/222927978372067/5515647522956489423/5515647522956489423_signature_31.png
| 2023-02-07 | 2023-02-07 13:51 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23301 | DOH | ccmclaboratory01@gmail.com | 24.5 | ASSAY: Xpert MTB/RIF, COUNT: 720 ASSAY: Xpert MTB/RIF Ultra, COUNT: 815 |  | GXIV-4D | 838561 | None | None | #NAME? | Module B1 and B3 are operational. Waiting for module allocation for module B2 and B4 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515584715211159677/5515584715211159677_signature_28.png | Fatimah Mariveles | https://www.jotform.com/uploads/macarewecare/222927978372067/5515584715211159677/5515584715211159677_signature_31.png
| 2023-02-07 | 2023-02-07 13:47 | CEBU DOCTOR’S UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | MOLECULAR LABORATORY | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23300 | COMMERCIAL | cduhcovidpcrlab@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 45 |  | GX-IV | 110008524 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515583702226343283/5515583702226343283_signature_28.png | Alyssa Sayson | https://www.jotform.com/uploads/macarewecare/222927978372067/5515583702226343283/5515583702226343283_signature_31.png
| 2023-02-07 | 2023-02-07 09:59 | Tarlac City Health Center 3 | San Miguel Tarlac City | 3 | Tarlac | Laborqtiry | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23299 | DOH | daganlea572@gmail.com | 28.9 | ASSAY: Xpert MTB/RIF, COUNT: 189 ASSAY: Xpert MTB/RIF Ultra, COUNT: 177 | 2023-02-02 | GXIV-D | 831697 | Due for calibration | Performed preventive maintenance Archived test results Update Xpertcheck application from 1.5 to 1.5a Performed calibration Performed functionality test | Preventive maintenance and calibration | Unit is operational | All modules passed calibration | 6 pcs Xpertcheck kit | 9570 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515579694114359489/5515579694114359489_signature_28.png | Lea A. Dagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5515579694114359489/5515579694114359489_signature_31.png
| 2023-02-07 | 2023-02-07 13:23 | Bansalan Rural Health Unit | Poblacion Dos, Bansalan Davao del Sur | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23298 | PBSP | ellen68nebrada@gmail.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 190 ASSAY: Xpert MTB/RIF Ultra, COUNT: 448 | 2023-02-07 | GX-IV 4D | 838181 | Module B3 and B4 for Replacement | Module Replacement | Replaced the Defective Module B3/ SN: 807134 with new Module SN: 210073548 Replaced the defective Module B4/ SN: 807078 with new Module SN: 210073544 Generate IQR, Assay Stat and System Log Replaced Fan Filters Activate XpertCheck Code | Unit is fully operational | N/A | 900-0508 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515573086811458200/5515573086811458200_signature_28.png | Ellen Nebrada Mayor, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5515573086811458200/5515573086811458200_signature_31.png
| 2023-02-07 | 2023-02-07 13:23 | Bansalan Rural Health Unit | Poblacion Dos, Bansalan Davao del Sur | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23297 | PBSP | ellen68nebrada@gmail.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 190 ASSAY: Xpert MTB/RIF Ultra, COUNT: 448 | 2023-02-07 | GX-IV 4D | 838181 | Module B3 and B4 for Replacement | Module Replacement | Replaced the Defective Module B3/ SN: 807134 with new Module SN: 210073548 Replaced the defective Module B4/ SN: 807078 with new Module SN: 210073544 Generate IQR, Assay Stat and System Log Replaced Fan Filters Activate XpertCheck Code | Unit is fully operational | N/A | 900-0508 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515573056817562978/5515573056817562978_signature_28.png | Ellen Nebrada Mayor, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5515573056817562978/5515573056817562978_signature_31.png
| 2023-02-07 | 2023-02-07 11:43 | Laguindingan Rural Health Unit | Laguindingan | 10 | Misamis Oriental | TB DOTS | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23296 | DOH | chellesambaan@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 38 |  | GeneXpert | 838575 | Stuck Cartridge Error 1001 on Module B3/863072 | Realigment of the valve drive Recalibration of Module B3 | Preventive maintenance Valve drive realigment Xpertcheck calibration | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515523241716883166/5515523241716883166_signature_28.png | Rochelle M. Sambaan | https://www.jotform.com/uploads/macarewecare/222927978372067/5515523241716883166/5515523241716883166_signature_31.png
| 2023-02-07 | 2023-02-07 12:03 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23295 | DOH | meggysantiago@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 451 ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 |  | GXIV | 838548 | For Preventive maintenance | #NAME? | Preventive Maintenance | Unit it Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515522447967811976/5515522447967811976_signature_28.png | Mary Grace Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5515522447967811976/5515522447967811976_signature_31.png
| 2023-02-07 | 2023-02-07 12:03 | Anahawan District Hospital | Anahawan Southern Leyte | 8 | Southern Leyte | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23294 | DOH | meggysantiago@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 451 ASSAY: Xpert MTB/RIF Ultra, COUNT: 203 |  | GXIV | 838548 | For Preventive maintenance | #NAME? | Preventive Maintenance | Unit it Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515522327968203950/5515522327968203950_signature_28.png | Mary Grace Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5515522327968203950/5515522327968203950_signature_31.png
| 2023-02-07 | 2023-02-07 11:16 | Loreto Rural Health Unit RTDL | Brgy. Poblacion, Loreto | 13 | Agusan del Sur | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 23293 | PBSP | ruthcasmagz@gmail.com | 24.4 | ASSAY: Xpert MTB/RIF, COUNT: 470 ASSAY: Xpert MTB/RIF Ultra, COUNT: 617 |  | GX-2 | 835077 | None | None | Replacing Module B1 w/SN: 655081 to SN: 675191 | All modules working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515500208515119704/5515500208515119704_signature_28.png | Ruth Magadan | https://www.jotform.com/uploads/macarewecare/222927978372067/5515500208515119704/5515500208515119704_signature_31.png
| 2023-02-07 | 2023-02-07 09:21 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23292 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 310 ASSAY: Xpert MTB/RIF, COUNT: 277 ASSAY: Xpert MTB/RIF Ultra, COUNT: 422 ASSAY: Xpert SARS-CoV-2, COUNT: 4223 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 | 2023-02-06 | GXIV | 834430 | Unit for preventive maintenance. For module replacement of B2 and B3. | Performed preventive maintenance. Performed module replacement on B2 and B3. Generated reports. | Preventive maintenance and module replacement. | Unit is operational. |  | Gx Module B2 (210070789) and B3 (210070807). | G 5161 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5515470165424780142/5515470165424780142_signature_28.png | Cynthia Fuertes,rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5515470165424780142/5515470165424780142_signature_31.png
| 2023-02-06 | 2023-02-06 14:17 | Mabalacat Rural Health Unit | Mabalacat Pampanga | 3 | Tarlac | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23291 | PBSP | perezriyadh@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF, COUNT: 340 ASSAY: Xpert MTB/RIF Ultra, COUNT: 925 ASSAY: Xpert HIV-1 Viral Load, COUNT: 21 | 2023-01-09 | GXIV-D | 807564 | A1 temp erature error (1001 and 2014) | Performed preventive maintenance Update xpertcheck application to 1.5a Performed calibration for modulen A4 Gathered reports Performed functionality test | Prevenyive maintenance and calibration | Unit is not operational | A1 and A4 for reporting | 2 Xpertcheck kit | 37969 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514965572983005159/5514965572983005159_signature_28.png | Riyadh L. Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5514965572983005159/5514965572983005159_signature_31.png
| 2023-02-06 | 2023-02-06 14:52 | Victorias CHO | Victorias City, Negros Occ. | 6 | Negros Occidental | TB- DOTS | YES | REPAIR | CEPHEID | Emerson Bejemino |  |  |  |  | 23290 | PBSP | vincent_llamoleonor@yahoo.com | 26.9 | ASSAY: Xpert MTB/RIF, COUNT: 524 ASSAY: Xpert MTB/RIF Ultra, COUNT: 927 |  | GX IV | 838141 | Defective module | Replace module SN 210070828 with new SN 210070812 | Replace module SN 210070828 with new SN 210070812 | 3 modules operational and 1 already reported |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514760686897726960/5514760686897726960_signature_28.png | Vincent Leonor | https://www.jotform.com/uploads/macarewecare/222927978372067/5514760686897726960/5514760686897726960_signature_31.png
| 2023-02-06 | 2023-02-06 12:41 | Davao TB Reference Laboratory | DOH Compound, JP Laurel, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23289 | PBSP | dtrlmark.pbsp@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 314 | 2023-02-06 | GX IV- 4D R1 | 707513 | Module B1 encounter multiple Temperature Error | Module Replacement | Replaced the Defective Module B1/ SN:636452 with new Module SN: 2100073546 Generate IQR, Assay Stat and System Log | Unit is operational | N/A | 900-0508 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514693508999515092/5514693508999515092_signature_28.png | Joseph Mark B. Anque | https://www.jotform.com/uploads/macarewecare/222927978372067/5514693508999515092/5514693508999515092_signature_31.png
| 2023-02-06 | 2023-02-06 12:41 | Davao TB Reference Laboratory | DOH Compound, JP Laurel, Davao City | 11 | Davao del Sur | TB-DOTS | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23288 | PBSP | dtrlmark.pbsp@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 314 | 2023-02-06 | GX IV- 4D R1 | 707513 | Module B1 encounter multiple Temperature Error | Module Replacement | Replaced the Defective Module B1/ SN:636452 with new Module SN: 2100073546 Generate IQR, Assay Stat and System Log | Unit is operational | N/A | 900-0508 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514693448993884652/5514693448993884652_signature_28.png | Joseph Mark B. Anque | https://www.jotform.com/uploads/macarewecare/222927978372067/5514693448993884652/5514693448993884652_signature_31.png
| 2023-02-06 | 2023-02-02 13:40 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23285 | PBSP | diazjimvianne@gmail.com | 25 | ASSAY: Xpert CT/NG, COUNT: 10 ASSAY: Others, COUNT: 11 ASSAY: Xpert MTB/RIF, COUNT: 27 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1749 ASSAY: Xpert SARS-CoV-2, COUNT: 13977 ASSAY: Xpert HIV-1 Viral Load, COUNT: 88 ASSAY: Others, COUNT: 9 |  | GXIV | 814159 | Unit for xpertcheck | Performed xpertcheck | #NAME? | Unit is operational. B4 failed xpertcheck. B2 & B3 insufficient. xpertcheck kit. B1 passed xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642211216745245/5514642211216745245_signature_28.png | Jim Vianne Diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642211216745245/5514642211216745245_signature_31.png
| 2023-02-06 | 2023-02-03 13:51 | To Love Yourself | Shaw Blvd., Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23287 | DOH | laboratory@loveyourself.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 57 | 2023-01-20 | GXIV | 831895 | Unit for xpertcheck | Performed xpertcheck | #NAME? | Unit is operational. A2 & A3 insufficient cartridge. A1 & A4 passed calibration. UPS for reporting. |  | XpertCheck Kit | 9530 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642211217109987/5514642211217109987_signature_28.png | Zarah Jane Aquilino | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642211217109987/5514642211217109987_signature_31.png
| 2023-02-06 | 2023-02-02 15:58 | Bagumbong Health Center | Bagumbong, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23286 | DOH | jenilyn_duenas@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 224 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1802 | 2023-01-18 | GXIV | 831907 | Unit for xpertcheck | Performed xpertcheck | #NAME? | Unit is operational. |  | xpertcheck kit | 9497 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642211218430432/5514642211218430432_signature_28.png | Jenilyn Duenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642211218430432/5514642211218430432_signature_31.png
| 2023-02-06 | 2023-02-01 13:02 | Bagong Silang ph. 1 Health Center | Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23284 | PBSP | alcaydeirene07@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 142 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1304 |  | GXIV | 838856 | Unit for maintenance | Performed maintenance | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171211718833/5514642171211718833_signature_28.png | Irene Alcayde | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171211718833/5514642171211718833_signature_31.png
| 2023-02-06 | 2023-01-31 13:45 | Bagong Barrio Health Center | Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23280 | DOH | cruzernitamay@gmail.com | 38 | ASSAY: Xpert MTB/RIF, COUNT: 139 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1103 | 2023-01-18 | GXIV | 831906 | Unit for calibration | Performed Xpert Check | #NAME? | Unit is operational. The environment is not suitable for GX operation (unable to install aircon. temp. too high 35°C). Module A3, A4 failed calibration. |  | xpertcheck kit | 9493 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171212998750/5514642171212998750_signature_28.png | Ernita Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171212998750/5514642171212998750_signature_31.png
| 2023-02-06 | 2023-02-01 13:31 | Bagong Silang ph. 2 Health Center | Bagong Silang, Caloocan city | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23281 | DOH | laguitannataniel@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 146 ASSAY: Xpert MTB/RIF Ultra, COUNT: 623 | 2023-01-19 | GXIV | 831677 | Unit for Xpert Check | Performed Xpert Check | #NAME? | Unit is operational. A4 failed calibration. A4 for reporting. |  | XpertCheck kit | 9506 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171215530421/5514642171215530421_signature_28.png | Nathaniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171215530421/5514642171215530421_signature_31.png
| 2023-02-06 | 2023-02-02 13:11 | Dr. Jose N. Rodriguez Memorial Hospital | Tala, Caloocan | NCR | NCR 3rd - Caloocan | Molecular Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23282 | PBSP | diazjimvianne@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: NA |  | GXIV | 707523 | #NAME? | - Installed GX unit to switch together w/ GX units 834432 & 814159 - performed xpertcheck | - re installed GX 4.8 - performed manual SQL installation - manually enabled SQL services - installed GX to switch - checked & cleaned cartridge bay, plunger rod, I core slot, filter - performed xpertcheck - performed functionality test | Unit is operational. Unable to restore GX software. Original laptop is recommended for upgrade. C2 & C3 insufficient xpertcheck kit. C1 failed, C4 passed xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171215788588/5514642171215788588_signature_28.png | Jim Vianne Diaz | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171215788588/5514642171215788588_signature_31.png
| 2023-02-06 | 2023-02-01 13:38 | Bagong Silang ph. 2 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23283 | DOH | laguitannataniel@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 331 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1050 | 2023-01-18 | GXIV | 831845 | Unit for calibration | Performed Xpertcheck | #NAME? | Unit is operational. All modules passed calibration. |  | XpertCheck kit | 9498 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171218639307/5514642171218639307_signature_28.png | Nathaniel Laguitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5514642171218639307/5514642171218639307_signature_31.png
| 2023-02-06 | 2023-02-03 11:04 | Jose B. Lingad Memorial General Hospital | Dolores City of San Fernando | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23279 | MACARE-HBDC | jblmgh.is@gmail.com | 34.05 | ASSAY: Xpert HBV VL, COUNT: 46 ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HCV VL, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1169 |  | GXIV-L | 812571 | Scanner error, invalid scan. | Performed surface cleaning Performed preventive maintenance Reconfigure Scanner Test scan of cartridge Gathered reports Performed functionality test | Preventive maintenance and repair | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514592122985102627/5514592122985102627_signature_28.png | Charlene Laxamana | https://www.jotform.com/uploads/macarewecare/222927978372067/5514592122985102627/5514592122985102627_signature_31.png
| 2023-02-06 | 2023-02-03 11:32 | Jose B. Lingad Memorial Genaral Hospital | Dolores City of San Fernando | 3 | Pampanga | Laboratory Bacte | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23278 | PBSP | jblmrhlab@gnail.com | 33 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 547 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1497 |  | GXIV-D | 804763 | Module A2 HW failed High ambient temperature | Performed surface cleaning Performed preventive maintenance Gathered reports Performed functionality test | Preventive maintenance | Unit is operational except for module A2 | Module A2 for reporting HW failed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514592122981706495/5514592122981706495_signature_28.png | Janice S. Condes | https://www.jotform.com/uploads/macarewecare/222927978372067/5514592122981706495/5514592122981706495_signature_31.png
| 2023-02-06 | 2023-02-02 10:53 | Laoang Rural Health Unit | Laoang Northern Samar | 8 | Northern Samar | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23277 | DOH | roseandarayan26@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 450 ASSAY: Xpert MTB/RIF Ultra, COUNT: 406 |  | GXIV | 838202 | For PM | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514567742312704080/5514567742312704080_signature_28.png | Rosario Andarayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5514567742312704080/5514567742312704080_signature_31.png
| 2023-02-06 | 2023-02-03 14:03 | Marikina City Health Office | Shoe Ave. | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23276 | DOH | amymitra22@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 358 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1229 ASSAY: Xpert HIV-1 Viral Load, COUNT: 105 | 2023-01-16 | GX-IV-D | 831588 | Module A1/661455 and A3/657408 error 4012 For XpertCheck | Performed XpertCheck | Performed maintenance Run XpertCheck Generated reports | Module A1 passed but under observation. Module A4 passed and operational. Module A3/657408 not operational. Module A2/792154 operational with existing report. |  | XpertCheck-CE-5 | DR# 9475 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514548741212983559/5514548741212983559_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5514548741212983559/5514548741212983559_signature_31.png
| 2023-02-06 | 2023-02-03 14:03 | Marikina City Health Office | Shoe Ave. | NCR | NCR 2nd - Marikina | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23275 | DOH | amymitra22@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 358 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1229 ASSAY: Xpert HIV-1 Viral Load, COUNT: 105 | 2023-01-16 | GX-IV-D | 831588 | Module A1/661455 and A3/657408 error 4012 For XpertCheck | Performed XpertCheck | Performed maintenance Run XpertCheck Generated reports | Module A1 passed but under observation. Module A4 passed and operational. Module A3/657408 not operational. Module A2/792154 operational with existing report. |  | XpertCheck-CE-5 | DR# 9475 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514546981219437176/5514546981219437176_signature_28.png | Amy Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5514546981219437176/5514546981219437176_signature_31.png
| 2023-02-06 | 2023-02-03 10:03 | The Medical City South Luzon | Sta. Rosa Laguna | 4A | Laguna | Molecular laboratory | NO | INSTALLATION | CEPHEID | Shiedrick Daroya |  |  |  |  | 23273 | COMMERCIAL | tmcsl-molecularlab@themedicalcity.com | 25 | ASSAY: Xpert HBV VL, COUNT: 3 ASSAY: Xpert HPV, COUNT: 5 ASSAY: Xpert MTB/RIF, COUNT: 43 ASSAY: Xpert MTB/RIF Ultra, COUNT: 116 ASSAY: Xpert HIV-1 Viral Load, COUNT: 7 |  | GXIV | 839233 | GeneXpert System for relocation from Pathology to Molecular Laboratory. | Performed installation of GeneXpert System to Molecular Laboratory. Generated reports. | Installation. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514544771212067970/5514544771212067970_signature_28.png | Ella Asa | https://www.jotform.com/uploads/macarewecare/222927978372067/5514544771212067970/5514544771212067970_signature_31.png
| 2023-02-06 | 2023-02-03 13:01 | Caloocan City Health Office | Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR | CEPHEID | Shiedrick Daroya |  |  |  |  | 23274 | PBSP | crissyname@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 406 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2279 ASSAY: Xpert HIV-1 Viral Load, COUNT: 169 |  | GXIV | 805271 | For troubleshooting of stucked cartridge on Module A1. | Performed repair on module A1 due to stucked cartridge. Performed functionality test. Generated reports. | Repair. | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514544771215246306/5514544771215246306_signature_28.png | Maricris Name | https://www.jotform.com/uploads/macarewecare/222927978372067/5514544771215246306/5514544771215246306_signature_31.png
| 2023-02-06 | 2023-02-03 14:07 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23272 | COMMERCIAL | spmcrhivda@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 57 ASSAY: Xpert MTB/RIF Ultra, COUNT: 497 ASSAY: Xpert HIV-1 Viral Load, COUNT: 523 | 2023-02-03 | GX-IV | 839152 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, and Icore Lenses | Unit is Operational but lost Filter | Improvised a temporary filter while waiting for the replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5514241617993721747/5514241617993721747_signature_28.png | Flordeliza Go | https://www.jotform.com/uploads/macarewecare/222927978372067/5514241617993721747/5514241617993721747_signature_31.png
| 2023-02-05 | 2023-02-05 13:28 | Luis Hora Memorial Regional Hospital | Abatan, Bauko, Mountain Province | CAR | Mountain Province | Laboratory | NO | CALIBRATION | CEPHEID | Harold Toco |  |  |  |  | 23271 | DOH | kazzjyll@gmail.com | 37 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1419 | 2023-02-01 | Gx4 | 831671 | For Xpertcheck Calubration/Maintenance.. UPS unable to back up power during power outage. | Performed Xpertcheck | Performed Naintenance/Cleaning/ Performed xpertcheck calibration. Check and troubleshoot UPS | All 4 modules PASS and operational. Defective UPS. | Pulled out defective UPS.. | Xpertcheck kit | DR number 9560 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5513883008516532597/5513883008516532597_signature_28.png | CHARLENE P. PASNGADAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5513883008516532597/5513883008516532597_signature_31.png
| 2023-02-05 | 2023-02-03 09:55 | St lukes medical center | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Micobiology | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23270 | COMMERCIAL | biomed.qc@stlukes.com.ph | 24 | ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 227 ASSAY: Others, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2779 |  | Gx4 | 839234 | For checkedup and troubleshootimg,broken door module hinge,b3 encounter intermittent h/w failed error. | Perforemed maintenance Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance and troubleshooting | B1 and b2 failed xpert last 7/10/22 B4 for pullout to be repair @ office |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5513849421313304824/5513849421313304824_signature_28.png | Arnel Isidro | https://www.jotform.com/uploads/macarewecare/222927978372067/5513849421313304824/5513849421313304824_signature_31.png
| 2023-02-05 | 2023-02-03 09:55 | St lukes medical center | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Micobiology | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23269 | COMMERCIAL | biomed.qc@stlukes.com.ph | 24 | ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 227 ASSAY: Others, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2779 |  | Gx4 | 839234 | For checkedup and troubleshootimg,broken door module hinge,b3 encounter intermittent h/w failed error. | Perforemed maintenance Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance and troubleshooting | B1 and b2 failed xpert last 7/10/22 B4 for pullout to be repair @ office |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5513849181314569972/5513849181314569972_signature_28.png | Arnel Isidro | https://www.jotform.com/uploads/macarewecare/222927978372067/5513849181314569972/5513849181314569972_signature_31.png
| 2023-02-05 | 2023-02-03 15:42 | Medilinx | Amorsolo st makati city | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23268 | COMMERCIAL | mellicettbaronia.medilinx@outlook.com | 28 | ASSAY: Xpert HPV, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 222 ASSAY: Xpert BCR-ABL Ultra, COUNT: 16 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 71 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Xpert HIV-1 Viral Load, COUNT: 254 |  | Gx4 10c | 110001647 | As per end user they encounter 5007 error | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5513848791313619786/5513848791313619786_signature_28.png | Mellicett Baronia | https://www.jotform.com/uploads/macarewecare/222927978372067/5513848791313619786/5513848791313619786_signature_31.png
| 2023-02-05 | 2023-02-05 10:36 | Bontoc General Hospital | Bontoc Mountain Province | CAR | Mountain Province | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Harold Toco |  |  |  |  | 23267 | PBSP | regaspimerrylkait@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2023-02-05 | Gx4 | 807582 | CPU troubleshooting as requested by Cepheid Tech Support | Performed CPU troubleshooting | Performed CPU troubleshooting | CPU issue did not resolved.. | Pulled out the CPU | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5513743068515005814/5513743068515005814_signature_28.png | Merryl kait Regaspi | https://www.jotform.com/uploads/macarewecare/222927978372067/5513743068515005814/5513743068515005814_signature_31.png
| 2023-02-05 | 2023-02-05 09:42 | Gabriela Silang General Hospital | Vigan | 1 | Ilocos Sur | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23266 | PBSP | mary.alexis.tahil@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 743 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GXIV-D | 807578 | Stuck Cartridge in Module 4 | Troubleshoot | Performed Maintenance Performed Self Test Removes Stucked Cartridge Generate Reports | Module 4 failed in XpertCheck, 3 modules are operational | UPS is Defective |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5513710911214516018/5513710911214516018_signature_28.png | Mary alexis Filarca-Tahil | https://www.jotform.com/uploads/macarewecare/222927978372067/5513710911214516018/5513710911214516018_signature_31.png
| 2023-02-04 | 2023-02-03 22:35 | Kolambugan Rural Health Unit | Kolambugan, Lanao del Norte | 10 | Lanao del Norte | TB DOTS | YES | Module Replacement | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23265 | PBSP | yojecinaj@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 570 |  | GeneXpert | 815272 | Error 1002 on Module A3/695631 | Replaced module | Module Replacement New Module SN: 610218 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512496413329814078/5512496413329814078_signature_28.png | Janice Joy R. Siso | https://www.jotform.com/uploads/macarewecare/222927978372067/5512496413329814078/5512496413329814078_signature_31.png
| 2023-02-03 | 2023-01-25 11:47 | Lugait Rural Health Unit | Lugait | 10 | Misamis Oriental | TBDOTS | NO | CALIBRATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23264 | DOH | dianekate22@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 341 |  | GeneXpert | 838562 | H/W Fail on Module B2 and B4 | Recalibration of failed modules 691477/B2 862890/B4 | Recalibration | Operational | B2 Service Required, B4 Passed |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512440843329338183/5512440843329338183_signature_28.png | Kate Diane Liquit | https://www.jotform.com/uploads/macarewecare/222927978372067/5512440843329338183/5512440843329338183_signature_31.png
| 2023-02-03 | 2023-02-03 14:41 | Santa Rosa RHU | Santa Rosa, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23263 | DOH | santarosatbdots1@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 157 ASSAY: Xpert MTB/RIF Ultra, COUNT: 84 | 2023-02-03 | Genexpert IV | 831599 | N/a | N/a | Performed Maintenance (Cleaned Icore, Cleaned Filters, Cleaned Cartridge Bay, Cleaned Plunger Rod) Performed Calibration Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck - All Modules Passed | Xpertcheck kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512223383919309073/5512223383919309073_signature_28.png | Ana S. Cabalong | https://www.jotform.com/uploads/macarewecare/222927978372067/5512223383919309073/5512223383919309073_signature_31.png
| 2023-02-03 | 2023-02-03 14:35 | CEBU TB REFERENCE LAB | Cebu City | 7 | Cebu | Laboratoy | YES | REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 23262 | PBSP | asmantiquilla.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1695 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4441 ASSAY: Xpert SARS-CoV-2, COUNT: 137 ASSAY: Xpert HIV-1 Viral Load, COUNT: 318 |  | GX IV | 838882 | Defective module 4 sn863607 | New installed module 4 sn210068078 | New installed module 4 sn210068078 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512158704416891143/5512158704416891143_signature_28.png | Allan s. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5512158704416891143/5512158704416891143_signature_31.png
| 2023-02-03 | 2023-02-03 14:07 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23261 | COMMERCIAL | spmcrhivda@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 57 ASSAY: Xpert MTB/RIF Ultra, COUNT: 497 ASSAY: Xpert HIV-1 Viral Load, COUNT: 523 | 2023-02-03 | GX-IV | 839152 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, and Icore Lenses | Unit is Operational but lost Filter | Improvised a temporary filter while waiting for the replacement | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512153423917077399/5512153423917077399_signature_28.png | Flordeliza Go | https://www.jotform.com/uploads/macarewecare/222927978372067/5512153423917077399/5512153423917077399_signature_31.png
| 2023-02-03 | 2023-02-03 14:06 | North Daang Hari Super Health Center | North Daang Hari | NCR | NCR 4th - Taguig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23260 | PBSP | onlychumz@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1180 |  | GX4D | 804702 | Unit is Due for Calibration | Calibration Preventive Maintenance | Performed Xpertcheck Performed Maintenance A1 661743 passed A2 713660 passed A3 731867 failed A4 685536 passed | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512140010114528075/5512140010114528075_signature_28.png | Charmaine D. Trinidad | https://www.jotform.com/uploads/macarewecare/222927978372067/5512140010114528075/5512140010114528075_signature_31.png
| 2023-02-03 | 2023-02-03 10:51 | Cainta City Health Office | Cainta | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23259 | PBSP | rowenatec@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 204 ASSAY: Xpert MTB/RIF Ultra, COUNT: 525 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 |  | GX-IV-D | 838182 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Modules B1, B3 and B4 passed and operational. | Module B2/714059 cannot run XpertCheck and continuous booting. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119685026317430/5512119685026317430_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119685026317430/5512119685026317430_signature_31.png
| 2023-02-03 | 2023-02-02 11:56 | Metro Antipolo Hospital and Medical Center | 1348 Marikina-Infanta Hwy, Brgy. Mayamot | 4A | Rizal | Laboratory | NO | INSTALLATION | CEPHEID | John Monsour Doctor |  |  |  |  | 23257 | COMMERCIAL | mahmci_labcmt@yahoo.com | 23 | ASSAY: Others, COUNT: N/A |  | GX-IV-D | 110005265 | For Installation | Installed the GeneXpert | Installed the GeneXpert system Installed CPU, monitor, barcode scanner, UPS and printer. Functionality test | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645022733835/5512119645022733835_signature_28.png | Graziela Madrid | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645022733835/5512119645022733835_signature_31.png
| 2023-02-03 | 2023-02-01 13:12 | Bureau of Plant Industry | Malate | NCR | NCR 1st - Manila | Laboratory | NO | Installation of accessories | CARL ZEISS | John Monsour Doctor |  |  |  |  | 23255 | COMMERCIAL | npqsd.laboratoryoperations@gmail.com | 26 |  | 2023-01-27 | Axiolab. A1 | 3135010036 | For Installation of accessories | Installed camera accessories | Installed camera accessory Checked computer connections Setup microscope camera Functionality test | Unit is operational. |  | Axiocam ERc | SI# 94685 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645025397268/5512119645025397268_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645025397268/5512119645025397268_signature_31.png
| 2023-02-03 | 2023-02-01 09:59 | Bureau of Plant Industry | Malate | NCR | NCR 1st - Manila | Laboratory | NO | Installation of accessories | CARL ZEISS | John Monsour Doctor |  |  |  |  | 23256 | COMMERCIAL | npqsd.laboratoryoperations@gmail.com | 26 |  | 2023-01-27 | Stemi 305 | 3943010058 | For Installation of accessories | Installed camera accessories | Installed camera accessory Checked computer connections Setup microscope camera Functionality test | Unit is operational. |  | Axiocam ERc | SI# 94685 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645026514683/5512119645026514683_signature_28.png | Jesriel Escalada | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645026514683/5512119645026514683_signature_31.png
| 2023-02-03 | 2023-02-03 10:28 | Cainta City Health Office | Cainta | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23258 | PBSP | rowenatec@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 236 ASSAY: Xpert MTB/RIF Ultra, COUNT: 824 |  | GX-IV-D | 804687 | For Maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645029591072/5512119645029591072_signature_28.png | Richard Ballester+ | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645029591072/5512119645029591072_signature_31.png
| 2023-02-03 | 2023-02-03 10:43 | Cainta City Health Office | Cainta | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23254 | DOH | rowenatec@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 196 ASSAY: Xpert MTB/RIF Ultra, COUNT: 564 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 | 2023-01-20 | GX-IV-D | 831623 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Modules A1, A2 and A4 passed and operational. | Module A3/754520 not operational. Reported already. | XpertCheck-CE-5 | DR# 9538 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645023243211/5512119645023243211_signature_28.png | Richard Ballester | https://www.jotform.com/uploads/macarewecare/222927978372067/5512119645023243211/5512119645023243211_signature_31.png
| 2023-02-03 | 2023-02-03 13:07 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23253 | PBSP | spmcrhivda@gmail.com | 23.2 | ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1558 ASSAY: Xpert HIV-1 Viral Load, COUNT: 349 | 2023-02-03 | GX-IV | 808298 | For Calibration Module B2 and B3 encounter Error | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Module B2 and B3 encountered multiple Temperature Related Error | Errors Report to JotForm for further Assessment | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512104113914460991/5512104113914460991_signature_28.png | Flordeliza Go | https://www.jotform.com/uploads/macarewecare/222927978372067/5512104113914460991/5512104113914460991_signature_31.png
| 2023-02-03 | 2023-02-03 10:43 | Benguet PHO | Bnguet Gen Hospital Conpound La Trinidad, Benguet | CAR | Benguet | Laboratory | NO | REPAIR | CEPHEID | Harold Toco |  |  |  |  | 23252 | PBSP | benguetphortdl@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2023-02-02 | Gx4 | 807621 | Workstation back up laptop for installation. | Install back up laptop | Install back up laptop..Performed maintenance | All 4 modules operational. |  | Laptop | G 5135 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5512020314813597769/5512020314813597769_signature_28.png | Frannie G. Calixto | https://www.jotform.com/uploads/macarewecare/222927978372067/5512020314813597769/5512020314813597769_signature_31.png
| 2023-02-03 | 2023-02-03 07:42 | Baguio General Hospital | Baguio City | CAR | Benguet | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Harold Toco |  |  |  |  | 23251 | PBSP | dmgonzales.pbsp@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 994 ASSAY: Xpert MTB/RIF, COUNT: 423 ASSAY: Xpert HPV, COUNT: 224 | 2023-01-31 | Gx4 | 805212 | Module B1 for replacement | Replaced module B1 | Replaced module B1..Performed maibtwnance. Performed functionality test. | 3 modules operational 1 module disabled status due to error and for reforting. | None | New mod 801381 | G 5125 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511922659011472342/5511922659011472342_signature_28.png | Donnel Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5511922659011472342/5511922659011472342_signature_31.png
| 2023-02-02 | 2023-02-02 11:56 | Columbio Rural Health Unit | Columbio | 12 | Sultan Kudarat | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23250 | DOH | covivido211@gmail.com | 27.3 | ASSAY: Xpert MTB/RIF, COUNT: 200 ASSAY: Xpert MTB/RIF Ultra, COUNT: 334 |  | Gx -IV | 838556 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | Genexpert machine has 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511426283916490974/5511426283916490974_signature_28.png | Rolan E. Coloma | https://www.jotform.com/uploads/macarewecare/222927978372067/5511426283916490974/5511426283916490974_signature_31.png
| 2023-02-02 | 2023-02-02 15:52 | San Francisco Rural Health Unit | San Francisco, Camotes, Cebu | 7 | Cebu | TB Dots | YES | CALIBRATION MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23249 | PBSP | ejoey22@gmail.com | 23.6 | ASSAY: Xpert MTB/RIF, COUNT: 418 ASSAY: Xpert MTB/RIF Ultra, COUNT: 257 | 2023-02-02 | GXIV-2D | 835058 | Due for calibration | #NAME? | #NAME? | All modules within calibrated parameters. All two modules are operational. | Collected three extra XPC cartridges. | XpertCheck-CE-5, Lot no. 1000378669, Exp. 2024 02 11 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511351017412005421/5511351017412005421_signature_28.png | Josiecora Albarracin | https://www.jotform.com/uploads/macarewecare/222927978372067/5511351017412005421/5511351017412005421_signature_31.png
| 2023-02-02 | 2023-02-02 16:06 | West Rembo Health Center | 21st street, west rembo | NCR | NCR 4th - Makati | Laboratory | NO | INSTALLATION | CEPHEID | Eron Dumagieng |  |  |  |  | 23248 | PBSP | mcabalosjr.pbsp@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF, COUNT: 0 |  | GX4D | 838875 | Unit is due for installation | Installation | Unit installation of GX machine 838875 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511346547412884780/5511346547412884780_signature_28.png | Jun Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/5511346547412884780/5511346547412884780_signature_31.png
| 2023-02-02 | 2023-02-02 15:02 | Poro Rural Health Unit | Poro, Camotes, Cebu | 7 | Cebu | TBDots | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 23247 | DOH | avilacharmaine6@gmail.com | 24.1 | ASSAY: Xpert MTB/RIF, COUNT: 299 ASSAY: Xpert MTB/RIF Ultra, COUNT: 204 |  | GXIV-4D | 838279 | -UPS doesn't turn on | -Check fuse if still intact (fuse is okay) -Check power supply voltage, 220V (within range) | #NAME? | All four modules are operational. Module B3 & B4 are for module replacement (Xpertcheck service required on cal 6/15/22) | UPS was not recovered. Still no power after troubleshooting. Recommend for replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511309521338395243/5511309521338395243_signature_28.png | Charmaine Apat | https://www.jotform.com/uploads/macarewecare/222927978372067/5511309521338395243/5511309521338395243_signature_31.png
| 2023-02-02 | 2023-02-02 14:22 | San Isidro RHU | Poblacion, San Isidro, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23246 | PBSP | urian_marlyn@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 46 ASSAY: Xpert MTB/RIF Ultra, COUNT: 426 | 2023-02-02 | Genexpert IV | 805107 | Module 3 error 2008 | Gathered Reports | Performed Maintenance (Cleaned Icore, Cleaned Cartridge Bay, Cleaned Filters, Cleaned Plunger Rod) Gathered Reports Performed Module Replacement 662521 - 210045948 Performed Calibration Performed Functionality Test | Unit is Operational | Xpertcheck - B3 Passed - B4 Service Required | Xpertcheck Kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511308213913990970/5511308213913990970_signature_28.png | Marlyn M.Urian | https://www.jotform.com/uploads/macarewecare/222927978372067/5511308213913990970/5511308213913990970_signature_31.png
| 2023-02-02 | 2023-02-02 14:21 | Bauan Rural Health Unit II | Lapaz Ext. Road, Bauan | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23245 | DOH | rhubauan2@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 108 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1077 | 2023-01-17 | GXIV | 820204 | Unit for preventive maintenanc. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules A1(721909) and A4(721920), service required. | Used 1 spare xpertcheck | Xpertcheck | DR 9486 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511303153386914879/5511303153386914879_signature_28.png | Maribel Aguirre Abrigonda | https://www.jotform.com/uploads/macarewecare/222927978372067/5511303153386914879/5511303153386914879_signature_31.png
| 2023-02-02 | 2023-02-02 14:13 | Dalaguete Rural Health Unit | Poblacion, Dalaguete Cebu | 7 | Cebu | Tb Lab | YES | REPAIR MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 23244 | PBSP | ajreef5517@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 1238 | 2023-02-02 | Gx4 | 815464 | Defective Module 2 : 696460 | Replaced New Module serial 884841 | Replaced New Module Get IQR, Sys Log, Gxx files | All modules operation |  | Na | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511281592817749792/5511281592817749792_signature_28.png | Allysa Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5511281592817749792/5511281592817749792_signature_31.png
| 2023-02-02 | 2023-02-02 12:51 | Pozorrubio Rural Health Unit | Sison St. Pozorrubio Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23243 | DOH | rhupozorrubio@yahoo.com | 29.8 | ASSAY: Xpert MTB/RIF, COUNT: 147 ASSAY: Xpert MTB/RIF Ultra, COUNT: 99 |  | GXIV-D | 838467 | Due for calibration | Performed surface cleaning Performed preventive maintenance Xpertcheck application update from V.1.5 to V.1.5a Performed calibration Gathered reports Performed functionality test | Preventive maintenance and calibration | Unit is operational | All modules passed calibration, 1 retry. Used 5 spare xpertcheck kits |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511258273741366341/5511258273741366341_signature_28.png | Jinky Grace Lumague | https://www.jotform.com/uploads/macarewecare/222927978372067/5511258273741366341/5511258273741366341_signature_31.png
| 2023-02-02 | 2023-02-02 13:26 | Panabo City Health Office | Poblacion, Panabo City | 11 | Davao del Norte | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23242 | DOH | nadineeust67@gmail.com | 22.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1368 | 2023-02-02 | GX-IV | 838295 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Machine is Operational; Module B2 Encounter Multiple Temperature Related Error | Data reported to Office for further assessment | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511257196016262605/5511257196016262605_signature_28.png | Nadine Gemma S. Estaquio | https://www.jotform.com/uploads/macarewecare/222927978372067/5511257196016262605/5511257196016262605_signature_31.png
| 2023-02-02 | 2023-02-02 12:18 | Batangas City Health Office | Batangas City | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23240 | DOH | gjanlorenzo@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 172 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1046 | 2023-01-17 | GXIV | 821317 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. All modules passed xpertcheck. |  | Xpertcheck | DR 9487 |  | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5511253515827496259/5511253515827496259_signature_28.png | Jan Lorenzo Godoy | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5511253515827496259/5511253515827496259_signature_31.png
| 2023-02-02 | 2023-02-02 11:10 | San Isidro RTDL | Dr. A. Santos, San Isidro, Parañaque City | NCR | NCR 4th - Taguig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23239 | PBSP | herminadoa@yahoo.com | 24.2 | ASSAY: Xpert MTB/RIF, COUNT: 127 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1378 ASSAY: Xpert HIV-1 Viral Load, COUNT: 72 |  | GX4D | 838192 | Unit is due for Calibration | Calibration | Performed Maintenance Performed XpertCheck B1 806894 failed- service requied B2 806895 retry then failed - service required B3 806897 retry B4 806898 retry | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511174131616174012/5511174131616174012_signature_28.png | Amelia G. Herminado | https://www.jotform.com/uploads/macarewecare/222927978372067/5511174131616174012/5511174131616174012_signature_31.png
| 2023-02-02 | 2023-02-02 10:01 | INTEGRATED PROVINCIAL HEALTH OFFICE-SIQUIJOR | SIQUIJOR | 7 | Siquijor | GENEXPERT LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23238 | PBSP | jtcabanlit.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 627 ASSAY: Xpert MTB/RIF Ultra, COUNT: 35 |  | GX-II | 835056 | PREVENTIVE MAINTENANCE | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511130150145534930/5511130150145534930_signature_28.png | James Cabanlit | https://www.jotform.com/uploads/macarewecare/222927978372067/5511130150145534930/5511130150145534930_signature_31.png
| 2023-02-02 | 2023-02-02 09:41 | INTEGRATED PROVINCIAL HEALTH OFFICE-SIQUIJOR | SIQUIJOR | 7 | Siquijor | GENEXPERT LAB | NO | REPAIR | CEPHEID | Joseph Acebedo |  |  |  |  | 23237 | PBSP | jtcabanlit.pbsp@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1198 ASSAY: Xpert MTB/RIF Ultra, COUNT: 33 |  | GX-IV | 808638 | FOR MODULE B2,B3, and B4 REPLACEMENT | -REPLACE MODULE B2 DEF SN:667655 WITH NEW SN:882343 MODULE B3 DEF SN:667656 WITH NEW : 210068080 MODULE B4 DEF SN:893060 WITH SN:639819 | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511117950144298485/5511117950144298485_signature_28.png | James Cabanlit | https://www.jotform.com/uploads/macarewecare/222927978372067/5511117950144298485/5511117950144298485_signature_31.png
| 2023-02-02 | 2023-02-01 13:30 | Chinese General Hospital | Blumentritt, Manila | NCR | NCR 1st - Manila | Bacteriology | NO | INSTALLATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23235 | CONSORTIUM | cgh_clinicalpathology@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 5 ASSAY: Xpert CT/NG, COUNT: 24 ASSAY: Xpert HBV VL, COUNT: 37 ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1148 ASSAY: Xpert HCV VL, COUNT: 11 ASSAY: Xpert HIV-1 Viral Load, COUNT: 17 |  | GXIV | 835604 | Unit for preventive maintenance. GeneXpert system for relocation within laboratory. | Performed relocation of Genexpert system. Installed Genexpert system to new location. Performed preventive maintenance. Generated reports. | Preventive maintenance and relocation of Genexpert system. | Unit is operational. For reporting of module B1 due to temperature. error. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105062024680453/5511105062024680453_signature_28.png | Maria ODessa R. Mallillin | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105062024680453/5511105062024680453_signature_31.png
| 2023-02-02 | 2023-02-01 15:20 | Atang dela Rama Health Center | Gagalangin Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23236 | DOH | monskie_22@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 383 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1028 | 2023-01-19 | GXIV | 831644 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed xpertcheck. Generated reports | Preventive maintenance and Xpertcheck. | Unit is operational. Modules B2(756759) B3(756760) and B4(756764) are service required. 1 operational module. |  | Xpertcheck | DR 9507 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105062029393047/5511105062029393047_signature_28.png | Monaliza taruc | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105062029393047/5511105062029393047_signature_31.png
| 2023-02-02 | 2023-01-31 08:25 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23233 | DOH | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 155 ASSAY: Xpert MTB/RIF, COUNT: 147 ASSAY: Xpert MTB/RIF Ultra, COUNT: 201 ASSAY: Xpert SARS-CoV-2, COUNT: 2100 | 2023-01-14 | GXIV | 831872 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is not operational. 4 defective modules. |  | Xpertcheck | DR 9514 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022021410103/5511105022021410103_signature_28.png | Analyn N.Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022021410103/5511105022021410103_signature_31.png
| 2023-02-02 | 2023-02-01 08:50 | Gat. Andres Bonifacio Memorial Medical Center | Delpan, Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23234 | PBSP | laboratory.gabmmc@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 107 ASSAY: Xpert MTB/RIF Ultra, COUNT: 791 | 2023-02-01 | GXIV | 805103 | Unit for preventive maintenance. Unit for Xpertcheck. For troubleshooting of module B2. | Performed preventive maintenance. Performed Xpertcheck. Replaced filter. Troubleshoot module B2. Generated reports. | Preventive maintenance, Xpertcheck and trobleshooting. | Unit is operational. B2 is defective due to broken ultrasonic horn. Modules B1, B3 and B4 passed Xpertcheck. |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022022488015/5511105022022488015_signature_28.png | Alma Lacson | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022022488015/5511105022022488015_signature_31.png
| 2023-02-02 | 2023-01-25 14:48 | Sto. Tomas Rural Health Unit | Poblacion Dos, Sto. Tomas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23230 | PBSP | vickybana_412@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 193 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1007 | 2023-01-25 | GXIV | 812786 | Unit for preventive maintenance. Umit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B4(671461), service required. 3 operational modules. |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022024052356/5511105022024052356_signature_28.png | Victoria B. Menor | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022024052356/5511105022024052356_signature_31.png
| 2023-02-02 | 2023-01-31 14:29 | Parola Health Center | Parola Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23232 | DOH | friaslynette@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 154 | 2023-01-30 | GXIV | 831871 | Unit for preventive maintenance. Unit for Xpertcheck. For installation of monitor. | Performed preventive maintenance. Performed Xpertcheck. Installed new monitor. Generated reports. | Preventive maintenance, xpertcheck and installation of monitor. | Unit is operational. |  | AOC 22B1HS Monitor (QSDN8HA087214) in replacment to defective monitor HP ProDisplay (CNK83224NB). Xpertcheck Kit. | DR 9553, DR9509 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022024093777/5511105022024093777_signature_28.png | Lynette Olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022024093777/5511105022024093777_signature_31.png
| 2023-02-02 | 2023-01-31 08:32 | Tondo Medical Center | Tondo Manila | NCR | NCR 1st - Manila | Covid Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23231 | PBSP | microscan@tmc.doh.gov.ph | 25 | ASSAY: Xpert HPV, COUNT: 445 ASSAY: Xpert MTB/RIF, COUNT: 501 ASSAY: Xpert MTB/RIF Ultra, COUNT: 829 ASSAY: Xpert SARS-CoV-2, COUNT: 10903 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2023-01-31 | GX16 | 844121 | Unit for preventive maintenance. Unit for Xpertcheck. Generated reports. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules A1, A2, A4, B2, B3, C4 is service required. |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022027835375/5511105022027835375_signature_28.png | Analyn N.Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5511105022027835375/5511105022027835375_signature_31.png
| 2023-02-02 | 2023-02-02 09:13 | Pilar Municipal Health Office and Lying-in Clinic | Pilar, Camotes, Cebu | 7 | Cebu | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23229 | PBSP | steveyong3384@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 85 | 2023-02-02 | GXIV-2D | 834813 | Module B2/787177 - XpertCheck Service Required | Replace defective module B2/787177 with module SN 210068686 | Preventive Maintenance | All two modules are functional. | Collected two XPC cartridges. | GeneXpert Module 6 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511100396907508657/5511100396907508657_signature_28.png | Steve Cuering | https://www.jotform.com/uploads/macarewecare/222927978372067/5511100396907508657/5511100396907508657_signature_31.png
| 2023-02-02 | 2023-02-02 09:13 | Pilar Municipal Health Office and Lying-in Clinic | Pilar, Camotes, Cebu | 7 | Cebu | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23228 | PBSP | steveyong3384@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 85 | 2023-02-02 | GXIV-2D | 834813 | Module B2/787177 - XpertCheck Service Required | Replace defective module B2/787177 with module SN 210068686 | Preventive Maintenance | All two modules are functional. | Collected two XPC cartridges. | GeneXpert Module 6 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511100146904988471/5511100146904988471_signature_28.png | Steve Cuering | https://www.jotform.com/uploads/macarewecare/222927978372067/5511100146904988471/5511100146904988471_signature_31.png
| 2023-02-02 | 2023-02-02 08:55 | San Roque Rural Health Unit | San Roque Northern Samar | 8 | Northern Samar | Laboratory | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23227 | DOH | lynnjarito@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 265 ASSAY: Xpert MTB/RIF Ultra, COUNT: 351 |  | GXIV | 838498 | For Preventive Maintenance | #NAME? | #NAME? | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511089112761311225/5511089112761311225_signature_28.png | Lynn N. Jarito | https://www.jotform.com/uploads/macarewecare/222927978372067/5511089112761311225/5511089112761311225_signature_31.png
| 2023-02-02 | 2023-02-01 08:55 | Lung cemter of the Philippines | East ave.qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23226 | PBSP | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1824 ASSAY: Xpert SARS-CoV-2, COUNT: 2014 | 2023-01-31 | GX XVI | 801080 | For maintenance and module replacement b3 796657 - 210070829 C2 686590 - 210068734 | Performed module replacement,Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance, module replacement | Operational |  | 900 gx | G5123 , G5124 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511073901213166671/5511073901213166671_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5511073901213166671/5511073901213166671_signature_31.png
| 2023-02-02 | 2023-02-01 08:55 | Lung cemter of the Philippines | East ave.qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23225 | PBSP | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1824 ASSAY: Xpert SARS-CoV-2, COUNT: 2014 | 2023-01-31 | GX XVI | 801080 | For maintenance and module replacement b3 796657 - 210070829 C2 686590 - 210068734 | Performed module replacement,Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance, module replacement | Operational |  | 900 gx | G5123 , G5124 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511073861212692065/5511073861212692065_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5511073861212692065/5511073861212692065_signature_31.png
| 2023-02-02 | 2023-02-01 09:01 | Lung center of the Philippines | East ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23224 | PBSP | lcpmicrolab@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 359 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2906 ASSAY: Xpert SARS-CoV-2, COUNT: 3164 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | Gx4 | 804205 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511073361219349920/5511073361219349920_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5511073361219349920/5511073361219349920_signature_31.png
| 2023-02-02 | 2023-02-01 09:01 | Lung center of the Philippines | East ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23223 | PBSP | lcpmicrolab@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 359 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2906 ASSAY: Xpert SARS-CoV-2, COUNT: 3164 ASSAY: Xpert HIV-1 Viral Load, COUNT: 16 |  | Gx4 | 804205 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511073321216317538/5511073321216317538_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5511073321216317538/5511073321216317538_signature_31.png
| 2023-02-02 | 2023-02-01 09:12 | Lung center of the Philippines | East ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23222 | MACARE-HBDC | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 ASSAY: Xpert SARS-CoV-2, COUNT: 895 ASSAY: Xpert HIV-1 Viral Load, COUNT: 28 ASSAY: Others, COUNT: 1 ASSAY: Others, COUNT: 1 |  | Gx 4 | 838837 | For maintenance, h/w failed module b4 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | 3 module operational b4 reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511072571217886517/5511072571217886517_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5511072571217886517/5511072571217886517_signature_31.png
| 2023-02-02 | 2023-02-01 09:12 | Lung center of the Philippines | East ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23221 | MACARE-HBDC | lcpmicrolab@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 44 ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 ASSAY: Xpert SARS-CoV-2, COUNT: 895 ASSAY: Xpert HIV-1 Viral Load, COUNT: 28 ASSAY: Others, COUNT: 1 ASSAY: Others, COUNT: 1 |  | Gx 4 | 838837 | For maintenance, h/w failed module b4 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | 3 module operational b4 reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511072531219729602/5511072531219729602_signature_28.png | Laurence Laquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5511072531219729602/5511072531219729602_signature_31.png
| 2023-02-02 | 2023-02-01 11:18 | Lung center of the Philippines | East ave qc | NCR | NCR 2nd - Quezon City | Koica | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23220 | PBSP | nilabis.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 0 | 2023-02-01 | Gx4 r1 | 704584 | For maintenance,xpertcheck,defectivelaptop corrupted o.s and defective lan cable | Performed xpertcheck Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance,xpertcheck | B2 failed xpert check,B1,B3,B4 passed xpertcheck | Unit unable to use bec.defective laptop | Xpertcheck | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511072111214880519/5511072111214880519_signature_28.png | Melissa Dawn Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5511072111214880519/5511072111214880519_signature_31.png
| 2023-02-02 | 2023-02-01 11:18 | Lung center of the Philippines | East ave qc | NCR | NCR 2nd - Quezon City | Koica | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23219 | PBSP | nilabis.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 0 | 2023-02-01 | Gx4 r1 | 704584 | For maintenance,xpertcheck,defectivelaptop corrupted o.s and defective lan cable | Performed xpertcheck Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance,xpertcheck | B2 failed xpert check,B1,B3,B4 passed xpertcheck | Unit unable to use bec.defective laptop | Xpertcheck | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511072061211498789/5511072061211498789_signature_28.png | Melissa Dawn Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5511072061211498789/5511072061211498789_signature_31.png
| 2023-02-02 | 2023-02-01 11:48 | Lung center of the Philippines | East ave qc | NCR | NCR 2nd - Quezon City | Koica | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23218 | PBSP | nilabis.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1273 | 2023-02-01 | Gx4 r1 | 707450 | For maintenance and xpertcheck | Performed xpertcheck,Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance and xpertcheck | C2,C3 failed xpert check C1 and C4 passed |  | Xpertcheck | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511071731219494774/5511071731219494774_signature_28.png | Melissa Dawn Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5511071731219494774/5511071731219494774_signature_31.png
| 2023-02-02 | 2023-02-01 11:48 | Lung center of the Philippines | East ave qc | NCR | NCR 2nd - Quezon City | Koica | NO | CALIBRATION TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23217 | PBSP | nilabis.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 229 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1273 | 2023-02-01 | Gx4 r1 | 707450 | For maintenance and xpertcheck | Performed xpertcheck,Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance and xpertcheck | C2,C3 failed xpert check C1 and C4 passed |  | Xpertcheck | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511071621219868881/5511071621219868881_signature_28.png | Melissa Dawn Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5511071621219868881/5511071621219868881_signature_31.png
| 2023-02-02 | 2023-02-01 14:26 | St lukes medical center | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23216 | COMMERCIAL | biomed.qc@stlukes.com.ph | 28 | ASSAY: Xpert HPV, COUNT: 389 |  | Gx4 | 807155 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | C2 module h/w failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511071181213365979/5511071181213365979_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5511071181213365979/5511071181213365979_signature_31.png
| 2023-02-02 | 2023-02-01 14:26 | St lukes medical center | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23215 | COMMERCIAL | biomed.qc@stlukes.com.ph | 28 | ASSAY: Xpert HPV, COUNT: 389 |  | Gx4 | 807155 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | C2 module h/w failed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511071101216703855/5511071101216703855_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5511071101216703855/5511071101216703855_signature_31.png
| 2023-02-02 | 2023-02-01 14:05 | St.lukes medical center | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23214 | COMMERCIAL | biomed.qc@stlukes.com.ph | 26 | ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 225 ASSAY: Others, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2769 |  | Gx4 | 839234 | For maintenance,struck cartridge | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | B4 module broken b1 and b2 reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511070231213075091/5511070231213075091_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5511070231213075091/5511070231213075091_signature_31.png
| 2023-02-02 | 2023-02-01 14:05 | St.lukes medical center | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Microbiology | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23213 | COMMERCIAL | biomed.qc@stlukes.com.ph | 26 | ASSAY: Xpert C. difficile BT, COUNT: 1 ASSAY: Xpert CT/NG, COUNT: 225 ASSAY: Others, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2769 |  | Gx4 | 839234 | For maintenance,struck cartridge | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | B4 module broken b1 and b2 reported already |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5511070171216157870/5511070171216157870_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5511070171216157870/5511070171216157870_signature_31.png
| 2023-02-02 | 2023-02-01 15:26 | National children's hospital | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Covid lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23212 | PBSP | nchmolecularlab2021@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 589 ASSAY: Xpert MTB/RIF Ultra, COUNT: 145 ASSAY: Xpert SARS-CoV-2, COUNT: 2491 | 2023-02-01 | Gx4 | 803109 | For maintenence and xpertcheck | Performed xpertcheck,Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance,xpertcheck | B3 retry, 3 modules passed |  | Xpertcheck kits | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510975351313448851/5510975351313448851_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5510975351313448851/5510975351313448851_signature_31.png
| 2023-02-02 | 2023-02-01 15:21 | National childrens hospital | E.rodriguez qc | NCR | NCR 2nd - Quezon City | Covid lab | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23212 | MACARE-HBDC | nchmolecularlab2021@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 16 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1861 ASSAY: Xpert SARS-CoV-2, COUNT: 5827 ASSAY: Others, COUNT: 18 |  | Gxxvi (10c) | 110014171 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | A3 and D2 reported already 14 module operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510975351315408171/5510975351315408171_signature_28.png | Mark Magsino | https://www.jotform.com/uploads/macarewecare/222927978372067/5510975351315408171/5510975351315408171_signature_31.png
| 2023-02-01 | 2023-02-01 16:28 | St. Lukes Medical Center | 5th ave. BGC | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23211 | MACARE-HBDC | jlalina@stlukes.com.ph | 24.3 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 42 ASSAY: Xpert BCR-ABL Ultra, COUNT: 98 ASSAY: Xpert C. difficile BT, COUNT: 0 ASSAY: Xpert CT/NG, COUNT: 1005 ASSAY: Others, COUNT: 11 ASSAY: Xpert HBV VL, COUNT: 51 ASSAY: Xpert HPV, COUNT: 4 ASSAY: Others, COUNT: 963 ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1675 |  | GX4D | 808411 | Unit is due for module Replacement | Module Replacement Preventive Maintenance | Performed Module Replacement C3 692466 to C3 702387 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510506967329189834/5510506967329189834_signature_28.png | Jethro Alina | https://www.jotform.com/uploads/macarewecare/222927978372067/5510506967329189834/5510506967329189834_signature_31.png
| 2023-02-01 | 2023-02-01 16:28 | St. Lukes Medical Center | 5th ave. BGC | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23210 | MACARE-HBDC | jlalina@stlukes.com.ph | 24.3 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 42 ASSAY: Xpert BCR-ABL Ultra, COUNT: 98 ASSAY: Xpert C. difficile BT, COUNT: 0 ASSAY: Xpert CT/NG, COUNT: 1005 ASSAY: Others, COUNT: 11 ASSAY: Xpert HBV VL, COUNT: 51 ASSAY: Xpert HPV, COUNT: 4 ASSAY: Others, COUNT: 963 ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1675 |  | GX4D | 808411 | Unit is due for module Replacement | Module Replacement Preventive Maintenance | Performed Module Replacement C3 692466 to C3 702387 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510506767325683615/5510506767325683615_signature_28.png | Jethro Alina | https://www.jotform.com/uploads/macarewecare/222927978372067/5510506767325683615/5510506767325683615_signature_31.png
| 2023-02-01 | 2023-02-01 16:39 | St. Lukes Medical Center BGC | 5th Avenue, BGC | NCR | NCR 4th - Taguig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23209 | COMMERCIAL | jlalina@stlukes.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1675 ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert HPV, COUNT: 963 ASSAY: Xpert HBV VL, COUNT: 51 ASSAY: Xpert CT/NG, COUNT: 1005 ASSAY: Xpert BCR-ABL Ultra, COUNT: 98 ASSAY: Xpert BCR-ABL Ultra, COUNT: 42 |  | GX4D | 839231 | Unit is due for Calibration | Preventive Maintenance Calibration | Performed Calibration C1 866691 passed C2 866696 passed C3 866793 passed C4 866803 failed | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510505487326479272/5510505487326479272_signature_28.png | Jethro Alina | https://www.jotform.com/uploads/macarewecare/222927978372067/5510505487326479272/5510505487326479272_signature_31.png
| 2023-02-01 | 2023-02-01 14:06 | South Cotabato Provincial Hospital | Koronadal City | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23208 | PBSP | anamariermt22@gmail.com | 26.6 | ASSAY: Xpert MTB/RIF, COUNT: 275 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1827 ASSAY: Xpert HIV-1 Viral Load, COUNT: 389 | 2023-02-01 | Gx - IV | 815380 | Module B1 H/W failed | Replace module B1 SN:671816 with new module SN:210070834 | Replace module B1 SN:671816 with new module SN:210070834 | Genexpert machine has 4 operational modules |  | Genexpert module 6-Color | 900-0508 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510420669655646529/5510420669655646529_signature_28.png | ANA MARIE GRACE EGASAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5510420669655646529/5510420669655646529_signature_31.png
| 2023-02-01 | 2023-02-01 13:53 | San Carlos City Health Office | San Carlos Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23207 | DOH | gbmmbg@yahoo.com | 28.3 | ASSAY: Xpert MTB/RIF, COUNT: 202 ASSAY: Xpert MTB/RIF Ultra, COUNT: 266 | 2023-02-01 | GXIV-D | 838437 | Temperature error on Modules B1 and B2 | Performed aurfqce cleaning Analyze reported errors Performed preventive maintenance Updated Xpertcheck application to 1.5A Performed calibration on modules B1 and B2 Performed functionality test | Preventive maintenance and calibration | Unit is operational except B2 failed calibration, B1 passed the calibration | B2 for reporting | 2 Xpertcheck kits | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510412858987409246/5510412858987409246_signature_28.png | Maricar Galleon | https://www.jotform.com/uploads/macarewecare/222927978372067/5510412858987409246/5510412858987409246_signature_31.png
| 2023-02-01 | 2023-02-01 13:45 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23206 | PBSP | laureanjo@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1678 ASSAY: Xpert MTB/RIF Ultra, COUNT: 95 ASSAY: Xpert SARS-CoV-2, COUNT: 40 |  | GX IV | 804840 | No power on CPU | Restart CPU | Restart CPU | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510406983949808194/5510406983949808194_signature_28.png | Laurean Jo F. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5510406983949808194/5510406983949808194_signature_31.png
| 2023-02-01 | 2023-02-01 12:40 | Reproductive Health and Wellness Center | Jacinto St, Poblacion District Davao City | 11 | Davao del Sur | HIV | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23205 | MACARE-HBDC | rhwclab@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 26 ASSAY: Xpert HIV-1 Viral Load, COUNT: 413 | 2023-02-01 | GX-IV 4D | 839236 | For Calibration | Perfrom XpertCheck Calibration | Calibrate the 4 Modules | All Modules Passed the Calibration | Module B1 for observation | XPERTCHECK CE 5 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510404641715186131/5510404641715186131_signature_28.png | Elaine Malano, Rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5510404641715186131/5510404641715186131_signature_31.png
| 2023-02-01 | 2023-02-01 12:40 | Reproductive Health and Wellness Center | Jacinto St, Poblacion District Davao City | 11 | Davao del Sur | HIV | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23204 | MACARE-HBDC | rhwclab@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 26 ASSAY: Xpert HIV-1 Viral Load, COUNT: 413 | 2023-02-01 | GX-IV 4D | 839236 | For Calibration | Perfrom XpertCheck Calibration | Calibrate the 4 Modules | All Modules Passed the Calibration | Module B1 for observation | XPERTCHECK CE 5 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510404661719292484/5510404661719292484_signature_28.png | Elaine Malano, Rmt | https://www.jotform.com/uploads/macarewecare/222927978372067/5510404661719292484/5510404661719292484_signature_31.png
| 2023-02-01 | 2023-02-01 11:16 | Koronadal City Health Office | Koronadal City | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23203 | PBSP | calvokenneth12@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF, COUNT: 734 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1232 | 2023-02-01 | Gx - IV | 707529 | Module A2 h/w failed | Perform Genexpert machine Preventive maintenance Replace module A2 SN:685790 with new module SN:668136 | Perform Genexpert machine Preventive maintenance Replace module A2 SN:685790 with new module SN:668136 | Genexpert machine has 4 operational modules |  | Genexpert module 6- color | 900-0491RS |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510311969655425266/5510311969655425266_signature_28.png | Kenneth Paolo T. Calvo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5510311969655425266/5510311969655425266_signature_31.png
| 2023-02-01 | 2023-02-01 09:53 | Siargao Island Medical Center | Dapa, Surigao del Norte | 13 | Surigao del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 23202 | DOH | sheena_rmt@yahoo.com | 30.14 | ASSAY: Xpert MTB/RIF, COUNT: 120 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1161 ASSAY: Xpert SARS-CoV-2, COUNT: 2565 |  | GX-4 | 838444 | B1 and B2 was found service required last November | None | Preventive Maintenance Data Collection | 2 modules working and 2 modules for replacement | 2 modules was reported. Waiting parts from Cepheid |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510278252794545798/5510278252794545798_signature_28.png | Sheena Rose T. Reazon | https://www.jotform.com/uploads/macarewecare/222927978372067/5510278252794545798/5510278252794545798_signature_31.png
| 2023-02-01 | 2023-02-01 08:40 | MANDAUE CITY HEALTH OFFICE | 838508 | 7 | Cebu | LABORATORY | NO | MERGE GX UNITS | CEPHEID | Joseph Acebedo |  |  |  |  | 23201 | DOH | elma.bragat@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1346 ASSAY: Xpert MTB/RIF Ultra, COUNT: 900 ASSAY: Xpert HIV-1 Viral Load, COUNT: 104 |  | GX-IV | 838508 | Connect GX units 838508,838510, and 815275 due to space limitation | Connect GX equipment using switch provided by end user | Collect iqr,system log, and assay stat | Equipment is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5510238140218780127/5510238140218780127_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5510238140218780127/5510238140218780127_signature_31.png
| 2023-01-31 | 2023-01-31 12:04 | Taguig Pateros District Hospital | East Service Rd. Western Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23200 | DOH | tpdhmicrobiology@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF, COUNT: 632 |  | GX4D | 831638 | Unit is due for Calibration | Preventive Maintenance Calibration | Performed Xpertcheck A1 710693 passed A2 670055 passed A3 754637 retry then failed A4 754638 retry | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5509616951212022368/5509616951212022368_signature_28.png | Pearl Gamboa | https://www.jotform.com/uploads/macarewecare/222927978372067/5509616951212022368/5509616951212022368_signature_31.png
| 2023-01-31 | 2023-01-31 12:04 | Taguig Pateros District Hospital | East Service Rd. Western Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23199 | DOH | tpdhmicrobiology@gmail.com | 24.6 | ASSAY: Xpert MTB/RIF, COUNT: 632 |  | GX4D | 831638 | Unit is due for Calibration | Preventive Maintenance Calibration | Performed Xpertcheck A1 710693 passed A2 670055 passed A3 754637 retry then failed A4 754638 retry | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5509616751216854930/5509616751216854930_signature_28.png | Pearl Gamboa | https://www.jotform.com/uploads/macarewecare/222927978372067/5509616751216854930/5509616751216854930_signature_31.png
| 2023-01-31 | 2023-01-31 14:01 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23198 | PBSP | cdh.laboratorydept@gmail.com | 28.4 | ASSAY: Xpert MTB/RIF, COUNT: 310 ASSAY: Xpert MTB/RIF Ultra, COUNT: 751 ASSAY: Xpert HIV-1 Viral Load, COUNT: 15 | 2023-01-13 | GXIV-D | 807562 | Module B3 for repkacement Module B4 for reporting error 2037 | Performed surface cleaning Replaced defective module Performed preventive maintenance Gathered reports Performed functionality test | Preventive maintenance and repair | Unit is operational except for modukea B1 and B4 | B4 for reporting | Genexpert module new SN:210068087 Defective module SN:688510 | 38197 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5509545200025043637/5509545200025043637_signature_28.png | Maria Rowena Punsalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5509545200025043637/5509545200025043637_signature_31.png
| 2023-01-31 | 2023-01-31 13:29 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23197 | DOH | dyrosemarie71@gmail.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 893 | 2023-01-31 | GX-IV | 838401 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data Test results | Unit is fully operational | N/A | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5509526841715317350/5509526841715317350_signature_28.png | Rosemarie B. Dy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5509526841715317350/5509526841715317350_signature_31.png
| 2023-01-31 | 2023-01-31 13:29 | Talomo South Health Center | Puan, Davao City | 11 | Davao del Sur | TB-DOTS | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23196 | DOH | dy_rosemarie@gmail.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 93 ASSAY: Xpert MTB/RIF Ultra, COUNT: 893 | 2023-01-31 | GX-IV | 838401 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data Test results | Unit is fully operational | N/A | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5509526811713847267/5509526811713847267_signature_28.png | Rosemarie B. Dy, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5509526811713847267/5509526811713847267_signature_31.png
| 2023-01-31 | 2023-01-31 11:53 | Polomolok Rural Health Unit | Poblacion Polomolok | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23195 | DOH | jmmirasolrmt@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF, COUNT: 359 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1181 |  | Gx -IV | 838265 | N/A | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | Genexpert machine has 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5509467740144491257/5509467740144491257_signature_28.png | June Michele J. Mirasol | https://www.jotform.com/uploads/macarewecare/222927978372067/5509467740144491257/5509467740144491257_signature_31.png
| 2023-01-31 | 2023-01-31 11:19 | Pilar Rural Health Unit | Brgy. Pilaring, Pilar, Surigao del Norte | 13 | Surigao del Norte | RTDL | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 23194 | PBSP | DexterJohnTDotillos@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | GX-2 | 834833 | None | None | Installation of new CPU and Monitor Preventive Maintenance XpertCheck Calibration | All modules are working | All modules passed in XpertCheck |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5509451038011544339/5509451038011544339_signature_28.png | Dexter John T. Dotillos | https://www.jotform.com/uploads/macarewecare/222927978372067/5509451038011544339/5509451038011544339_signature_31.png
| 2023-01-31 | 2023-01-31 11:00 | IPIL RURAL HEALTH UNIT | Poblacion ipil zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | MAINTENANCE Site inspection | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23193 | PBSP | rheasolana@gmail.com | 34 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1382 | 2023-01-31 | GX 4 | 808615 | Aircondition not working | Preventive maintenance | Preventive maintenance, Site inspection | All 4 modules are available but under high temperature | Air-condition of the room Must be replaced emmedietly. Room temperature above required temperature. | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5509441950753234252/5509441950753234252_signature_28.png | Rheamay E. Solana | https://www.jotform.com/uploads/macarewecare/222927978372067/5509441950753234252/5509441950753234252_signature_31.png
| 2023-01-31 | 2023-01-31 08:39 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY | 7 | Cebu | LABORATORY | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23192 | PBSP | elma.bragat@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1000 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1024 ASSAY: Xpert HIV-1 Viral Load, COUNT: 138 |  | GX-IV | 815275 | DISTORTED MONITOR DISPLAY | #NAME? | #NAME? | -CONFIRM MONITOR WITH SN 611NTVS6B383 IS DEFECTIVE | #NAME? |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5509405268811453724/5509405268811453724_signature_28.png | Elma Gonzales | https://www.jotform.com/uploads/macarewecare/222927978372067/5509405268811453724/5509405268811453724_signature_31.png
| 2023-01-30 | 2023-01-30 14:32 | South Cotabato Provincial Hospital | Koronadal City | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23191 | PBSP | anamariermt22@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF, COUNT: 285 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1810 ASSAY: Xpert HIV-1 Viral Load, COUNT: 389 | 2023-01-30 | Gx-4 | 815380 | Module B1 waiting for module to available Module B2,B3, and B4 for Service Required | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck | Genexpert machine has 3 operational modules |  | XpertCheck -CE-5 | Lot #: 1000330696 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508703701711118231/5508703701711118231_signature_28.png | ANA MARIE GRACE C. EGASAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5508703701711118231/5508703701711118231_signature_31.png
| 2023-01-30 | 2023-01-26 09:30 | General Emilio Aguinaldo Memorial Hospital | Trece Martires | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23188 | PBSP | djaemiel@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1625 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 | 2023-01-23 | GXIV-L | 802860 | Booting problem in module 3 | Module replacement | Replaced module 3 Performed Maintenance Generate reports Performed Self test | All modules are operational | B3: 632179 --> 702383 | Gx Module 6-Colors | G 5105 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685331213367805/5508685331213367805_signature_28.png | Jaemiel Juvirens A. Dimaranan | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685331213367805/5508685331213367805_signature_31.png
| 2023-01-30 | 2023-01-27 10:47 | Dasmariñas City Health Office II | DBB-E, Dasmariñas City | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23189 | DOH | hazelmilan@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2338 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 |  | GXIV-D | 838535 | Due for Maintenance | Preventive Maintenance | Performed Maintenance Generate Reports Performed Self Test | 3 Modules are Defective, 1 module operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685331213562899/5508685331213562899_signature_28.png | Roberto c, borila | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685331213562899/5508685331213562899_signature_31.png
| 2023-01-30 | 2023-01-27 12:10 | Paliparan City Health Office IV | Paliparan | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23190 | PBSP | krlandaya25@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1391 ASSAY: Xpert HIV-1 Viral Load, COUNT: 27 |  | GXIV-D | 838139 | Due for XpertCheck | XPertCheck Maintenance | Performed Maintenance Performed XpertCheck Generate Reports Performed Self-Test | 3 Modules Failed, 1 Module Passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685341211129957/5508685341211129957_signature_28.png | Khristofer Ralph L. Andaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685341211129957/5508685341211129957_signature_31.png
| 2023-01-30 | 2023-01-24 11:02 | Pamplona Uno Health Center | Pamplona Uno | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23183 | PBSP | carolinevecino5821@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 700 | 2023-01-23 | GXIV-D | 838541 | H/W failure | Module Replacement | Replaced Module 2 Performed Maintenance Performed XpertCheck Generate Reports | 1 newly replaced module | B2: 862469 --> 210068184 | Gx module 6-C | G 5111 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291211474189/5508685291211474189_signature_28.png | Caroline Vecino | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291211474189/5508685291211474189_signature_31.png
| 2023-01-30 | 2023-01-24 08:57 | Pulang Lupa Dos Health Center | Pulang Lupa Dos | NCR | NCR 4th - Las Piñas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23184 | PBSP | pulanglupados@yahoo.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1031 |  | GXIV-D | 838183 | Unit due for XpertCheck | XpertCheck Maintenance | Performed Maintenance Performed XpertCheck Generate Reports Performed Self Test | Unit is operational All modules passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291211719064/5508685291211719064_signature_28.png | Elvira Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291211719064/5508685291211719064_signature_31.png
| 2023-01-30 | 2023-01-25 13:22 | Indang Rural Health Unit | J. Mabiling St. , Indang | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23185 | PBSP | loloambo_2006@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 608 |  | GXIV-D | 838137 | Unit is due for XpertCheck | XpertCheck Maintenance | Performed Maintenance Performed XpertCheck Performed Self-Test Generate Reports | All modules passed and operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291215236642/5508685291215236642_signature_28.png | Marvin Sierra | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291215236642/5508685291215236642_signature_31.png
| 2023-01-30 | 2023-01-25 11:18 | Naic Rural Health Unit | Naic | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23186 | PBSP | jengcasamar2007@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 | 2023-01-16 | GXIV-D | 807567 | Temperature Error | Module Replacement | Replaced Module 4 Performed Maintenance Performed XpertCheck Generate Reports Test functionality of UPS | 2 modules Failed, 2 modules newly replaced. Recommendation for UPS replacement (2 ports have no output) | 4: 678360 --> 874939 | Gx module 6-C | G 5039 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291216319799/5508685291216319799_signature_28.png | Jennifer S.Casamar | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291216319799/5508685291216319799_signature_31.png
| 2023-01-30 | 2023-01-25 11:27 | Naic Rural Health Unit | Naic | 4A | Cavite | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 23187 | PBSP | jengcasamar2007@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 978 | 2023-01-25 | GXIV-D | 812006 | Booting errors in mod 1 | Gateway board replacement | Performed Maintenance Replaced and configure gateway board Performed xpertcheck Generate reports | 2 modules failed, 2 modules newly replaced |  | Gateway board kit | G 5115 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291216536665/5508685291216536665_signature_28.png | Jennifer Casamar | https://www.jotform.com/uploads/macarewecare/222927978372067/5508685291216536665/5508685291216536665_signature_31.png
| 2023-01-30 | 2023-01-30 13:40 | Zamboanga Sibugay Provincial Hospital | Sanito ipil, Zamboanga sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23182 | PBSP | exelrose.remotigue@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 547 | 2023-01-30 | GX 4 | 834831 | Monitor broken | Reported for replacement | Preventive maintenance | All 2 modules are available | Monitor malfunction | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508668647415268822/5508668647415268822_signature_28.png | Exel Rose Remotigue | https://www.jotform.com/uploads/macarewecare/222927978372067/5508668647415268822/5508668647415268822_signature_31.png
| 2023-01-30 | 2023-01-30 12:03 | Zaragoza RHU I | Del Pilar East Zaragoza Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23181 | PBSP | palapalzarah2425@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 68 ASSAY: Xpert MTB/RIF Ultra, COUNT: 319 | 2023-01-30 | Genexpert IV | 838888 | N/a | N/a | Performed Maintenance Gathered Reports Performed Calibratiom Performed Functionality Test | Unit is Operational | Xpertcheck - All Modules Passed | Xpertcheck kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508632990124938434/5508632990124938434_signature_28.png | Zarah A Palapal | https://www.jotform.com/uploads/macarewecare/222927978372067/5508632990124938434/5508632990124938434_signature_31.png
| 2023-01-30 | 2023-01-27 11:32 | Rafael Lazatin Memorial Medical Hospital | Angeles Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23180 | PBSP | wilper143@yahoo.com | 29.4 | ASSAY: Xpert MTB/RIF, COUNT: 264 ASSAY: Xpert MTB/RIF Ultra, COUNT: 702 | 2023-01-09 | GXIV-D | 838874 | Due for Calibrtion | Performed preventive maintenance Install xpertcheck application Performed machine calibration Gathered reports Performed functionality test | Preventive maintenance and Calibrtion | Unit is operational | All modules pssed calibration | Xpertcheck kits | 38013 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508499391214516650/5508499391214516650_signature_28.png | Perla Garcia | https://www.jotform.com/uploads/macarewecare/222927978372067/5508499391214516650/5508499391214516650_signature_31.png
| 2023-01-30 | 2023-01-27 11:27 | San Lorenzo Ruiz Women's Hospital | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23178 | PBSP | slrgh_lab1999@yoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 344 ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 |  | GX IV | 805009 | Unit for xpertcheck | Performed calibration | #NAME? | Unit is operational. All modules except B1 failed calibration. B1 for reporting. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463311212840237/5508463311212840237_signature_28.png | Myralene Piamonte | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463311212840237/5508463311212840237_signature_31.png
| 2023-01-30 | 2023-01-26 12:51 | University of The East - Ramon Magsaysay Memorial Hospital | Aurora BIvd., Quezon City | NCR | NCR 2nd - Quezon City | Blood Bank | NO | MAINTENANCE | HAEMONETICS | Ronnel Belizar |  |  |  |  | 23177 | COMMERCIAL | rcb.macare@gmail.com | 25 |  |  | MCS+ | 13C270 | Unit for maintenance | Performed maintenance | #NAME? | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463311213128227/5508463311213128227_signature_28.png | Ruel Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463311213128227/5508463311213128227_signature_31.png
| 2023-01-30 | 2023-01-27 12:37 | Catmon Health Center | Catmon, Malabon | NCR | NCR 3rd - Malabon | laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23179 | DOH | edrick.dianzon@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1442 ASSAY: Xpert HIV-1 Viral Load, COUNT: 70 |  | GXIV | 838880 | Unit for calibration | Performed XpertCheck | #NAME? | Unit is operational. All modules passed calibration |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463311213647541/5508463311213647541_signature_28.png | Erick Dianzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463311213647541/5508463311213647541_signature_31.png
| 2023-01-30 | 2023-01-25 15:16 | Norzagaray Rural Health Unit I | Norzagaray, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23176 | PBSP | chelseajempolicarpio@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 253 ASSAY: Xpert MTB/RIF Ultra, COUNT: 555 |  | GXIV | 838167 | unit for calibration | Performed Xpertcheck | #NAME? | Unit is operational. B1, B2, B4 passed the calibration. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271211107081/5508463271211107081_signature_28.png | Chelsea Policarpio | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271211107081/5508463271211107081_signature_31.png
| 2023-01-30 | 2023-01-24 13:53 | Mandaluyong City Health Office | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Ronnel Belizar |  |  |  |  | 23175 | DOH | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 331 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1723 ASSAY: Xpert HIV-1 Viral Load, COUNT: 92 |  | Gx IV | 831596 | Module C2 for xpertcheck | Performed xpertcheck on C2 | #NAME? | unit is operational. C2 for reporting /follow-up. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271213470015/5508463271213470015_signature_28.png | Revena Sibayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271213470015/5508463271213470015_signature_31.png
| 2023-01-30 | 2023-01-25 12:11 | Malolos Rural Health Unit VI | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23172 | WHO | meimbanj@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 495 ASSAY: Xpert MTB/RIF Ultra, COUNT: 724 |  | GXIV | 838862 | Unit for calibration | performed xpertcheck | #NAME? | Unit is operational. All modules passed xpertcheck. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271215832555/5508463271215832555_signature_28.png | Felix Jericho Meimban | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271215832555/5508463271215832555_signature_31.png
| 2023-01-30 | 2023-01-24 13:47 | Mandaluyong City Health office | Boni, Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23173 | PBSP | rsjalova1974@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 488 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2053 ASSAY: Xpert HIV-1 Viral Load, COUNT: 89 | 2023-01-23 | GX IV | 805062 | #NAME? | -Replaced module A2 (667784) by 682608 | #NAME? | Unit is operational. All modules passed except A1. A1 for reporting. |  | GeneXpert Module SN: 682608 | G 5103 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271218285196/5508463271218285196_signature_28.png | Revena Sibayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271218285196/5508463271218285196_signature_31.png
| 2023-01-30 | 2023-01-26 12:31 | University of The East - Ramon Maysaysay Memorial Hospital | Aurora Blvd., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Ronnel Belizar |  |  |  |  | 23174 | CONSORTIUM | osangcomendador19@gmail.com | 23 | ASSAY: Others, COUNT: 53 ASSAY: Xpert CT/NG, COUNT: 60 ASSAY: Xpert MTB/RIF, COUNT: 490 ASSAY: Xpert MTB/RIF Ultra, COUNT: 44 ASSAY: Others, COUNT: 2 |  | GXIV | 847578 | Unit for maintenance | Performed maintenance | #NAME? | Unit is operational. No issues found. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271219293422/5508463271219293422_signature_28.png | Osang Comendador | https://www.jotform.com/uploads/macarewecare/222927978372067/5508463271219293422/5508463271219293422_signature_31.png
| 2023-01-30 | 2023-01-28 07:56 | Bicol Medical Center | Naga | 5 | Camarines Sur | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  |  |  |  | 23171 | COMMERCIAL | bmc.bloodbank@gmail.com | 27 |  |  | MCS+ | 15H027G | For maintenance | Performed maintenance | Performed maintenance Checked rollers Verified calibration parameters Simulation run | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447751212801674/5508447751212801674_signature_28.png | Leo Collao | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447751212801674/5508447751212801674_signature_31.png
| 2023-01-30 | 2023-01-28 07:18 | Bicol Medical Center | Naga | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23168 | PBSP | bmc.mtip@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 1362 ASSAY: Xpert MTB/RIF, COUNT: 1368 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5140 ASSAY: Xpert SARS-CoV-2, COUNT: 2385 ASSAY: Xpert HIV-1 Viral Load, COUNT: 261 |  | GX-XVI-D | 842225 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod and filter. Generated reports | Unit is operational. | Module D3/886347 not operational. For reporting. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711211554604/5508447711211554604_signature_28.png | Monaleen Lavadia | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711211554604/5508447711211554604_signature_31.png
| 2023-01-30 | 2023-01-28 07:25 | Bicol Medical Center | Naga | 5 | Camarines Sur | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23167 | MACARE-HBDC | bmc.mtip@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 1772 ASSAY: Xpert MTB/RIF Ultra, COUNT: 34 ASSAY: Xpert SARS-CoV-2, COUNT: 192 |  | GX-IV-D | 110008458 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711211801434/5508447711211801434_signature_28.png | Monaleen lavadia | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711211801434/5508447711211801434_signature_31.png
| 2023-01-30 | 2023-01-28 07:53 | Bicol Medical Center | Naga | 5 | Camarines Sur | Blood bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | John Monsour Doctor |  |  |  |  | 23170 | COMMERCIAL | bmc.bloodbank@gmail.com | 27 |  |  | MCS+ | 15H031G | For calibration | Performed maintenance | Performed maintenance Checked rollers Verified calibration parameters Simulation run | Unit is operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711217200081/5508447711217200081_signature_28.png | LEO COLLAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711217200081/5508447711217200081_signature_31.png
| 2023-01-30 | 2023-01-27 14:33 | Pili Rural Health Unit | Pili | 5 | Albay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23167 | DOH | lfsanao724@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 175 ASSAY: Xpert MTB/RIF Ultra, COUNT: 920 |  | GX-IV-D | 838408 | Modules B2, B3 and B4 for xpertcheck | Performed XpertCheck | Performed maintenance Run XpertCheck Generated reports | Module B1/688686 operational. Module B2/682797 passed the XpertCheck but under observation. Modules B3/861110 and B4/861254 service required. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711217534491/5508447711217534491_signature_28.png | Liezl Sanao | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711217534491/5508447711217534491_signature_31.png
| 2023-01-30 | 2023-01-27 12:51 | Libon Rural Health Unit | Santiago St. Libon | 5 | Albay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23169 | DOH | marilyn@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 259 ASSAY: Xpert MTB/RIF Ultra, COUNT: 798 |  | GX-IV-D | 838397 | Module B1, B2 and B4 for re-xpertcheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Modules B3/860996 Operational. Module B1/640264 and B2/860937 passed the XpertCheck but under observation. Module B4/861019 | Some part of the keyboard not functioning. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711218361139/5508447711218361139_signature_28.png | Marilyn S. Razal | https://www.jotform.com/uploads/macarewecare/222927978372067/5508447711218361139/5508447711218361139_signature_31.png
| 2023-01-28 | 2023-01-28 15:21 | Governor Celestino Gallares Hospital | Tagbilaran BOHOL | 7 | Bohol | TB dots | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 23166 | PBSP | majieocmen@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1221 | 2023-01-28 | Gx4 | 820228 | Error 4009 | Calibrate module 4 serial number 722094 | Calibrate module 722094 | Failed during calibration for replacement |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5507000783618851445/5507000783618851445_signature_28.png | Jernai Sosoban | https://www.jotform.com/uploads/macarewecare/222927978372067/5507000783618851445/5507000783618851445_signature_31.png
| 2023-01-28 | 2023-01-28 11:03 | Cong. Simeon Toribio Memorial Hospital | Carmen, Poblacion Bohol | 7 | Bohol | Tb dots | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 23165 | DOH | Himaya.vale@yahoo.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 909 | 2023-01-28 | Gx4 | 838394 | Error 1002 | Performed calibration | Calibration on Module 1 Serial 838394 | 4 modules available |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506857995417736177/5506857995417736177_signature_28.png | Himaya Vale | https://www.jotform.com/uploads/macarewecare/222927978372067/5506857995417736177/5506857995417736177_signature_31.png
| 2023-01-27 | 2023-01-25 14:48 | Sto. Tomas Rural Health Unit | Poblacion Dos, Sto. Tomas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23164 | PBSP | vickybana_412@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 193 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1007 | 2023-01-25 | GXIV | 812786 | Unit for preventive maintenance. Umit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B4(671461), service required. 3 operational modules. |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506404935318047594/5506404935318047594_signature_28.png | Victoria B. Menor | https://www.jotform.com/uploads/macarewecare/222927978372067/5506404935318047594/5506404935318047594_signature_31.png
| 2023-01-27 | 2023-01-25 14:48 | Sto. Tomas Rural Health Unit | Poblacion Dos, Sto. Tomas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23163 | PBSP | vickybana_412@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 193 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1007 | 2023-01-25 | GXIV | 812786 | Unit for preventive maintenance. Umit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B4(671461), service required. 3 operational modules. |  | Xpertcheck | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506404075319328967/5506404075319328967_signature_28.png | Victoria B. Menor | https://www.jotform.com/uploads/macarewecare/222927978372067/5506404075319328967/5506404075319328967_signature_31.png
| 2023-01-27 | 2023-01-26 11:32 | San Sebastian District Hospital | Sablayan Occ Mindoro | 4B | Occidental Mindoro | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23162 | PBSP | ssdhlaboratory@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 622 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1290 | 2023-01-26 | GXIV | 812791 | Unit for preventive maintenance. Unit for Xpertcheck. For module replacement of B4. | Performed preventive maintenance. Performed module replacement on B4. Performed Xpertcheck. Generated reports. | Preventive maintenance, Xpertcheck and Module replacement. | Unit is operational. All modules passed Xpertcheck. |  | GX module for B4(878695) in replacement for defective B4(793158). | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506299065311696857/5506299065311696857_signature_28.png | Anthony D. Castro | https://www.jotform.com/uploads/macarewecare/222927978372067/5506299065311696857/5506299065311696857_signature_31.png
| 2023-01-27 | 2023-01-27 18:03 | Garcia Memorial Provincial Hospital | Talibon Bohol | 7 | Bohol | Tb Laboratory | YES | CALIBRATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 23161 | DOH | ybanezrosalina19@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 349 | 2023-01-27 | Gx4 | 838203 | Error 2014 and UPs problem | Calibrate genexpert | Calibrate Machine | All modules are operational |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506232913765383412/5506232913765383412_signature_28.png | Rosalina Ybanez | https://www.jotform.com/uploads/macarewecare/222927978372067/5506232913765383412/5506232913765383412_signature_31.png
| 2023-01-27 | 2023-01-27 15:48 | Batasan hills super health center | Batasan q..c | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23160 | PBSP | ashwinkevin@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 130 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1585 ASSAY: Xpert HPV, COUNT: 643 | 2023-01-27 | Gx4 | 805254 | For xpert check and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Xpert check and maintenance | D3 passed | D2/D4 retry xpertcheck D1 failed | Xpertcheck | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506225641311466779/5506225641311466779_signature_28.png | Kevin Orlando D. Cabuso | https://www.jotform.com/uploads/macarewecare/222927978372067/5506225641311466779/5506225641311466779_signature_31.png
| 2023-01-27 | 2023-01-27 15:48 | Batasan hills super health center | Batasan q..c | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23159 | PBSP | ashwinkevin@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 130 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1585 ASSAY: Xpert HPV, COUNT: 643 | 2023-01-27 | Gx4 | 805254 | For xpert check and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Xpert check and maintenance | D3 passed | D2/D4 retry xpertcheck D1 failed | Xpertcheck | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506225591311969324/5506225591311969324_signature_28.png | Kevin Orlando D. Cabuso | https://www.jotform.com/uploads/macarewecare/222927978372067/5506225591311969324/5506225591311969324_signature_31.png
| 2023-01-27 | 2023-01-27 13:04 | Bernadette bautista medical clinic | Commonwealth qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23158 | CONSORTIUM | richeclevia@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 0 |  | Gx4 | 835603 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506224911318663798/5506224911318663798_signature_28.png | Richard eclevia | https://www.jotform.com/uploads/macarewecare/222927978372067/5506224911318663798/5506224911318663798_signature_31.png
| 2023-01-27 | 2023-01-27 13:04 | Bernadette bautista medical clinic | Commonwealth qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23157 | CONSORTIUM | richeclevia@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 0 |  | Gx4 | 835603 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506224861312400439/5506224861312400439_signature_28.png | Richard eclevia | https://www.jotform.com/uploads/macarewecare/222927978372067/5506224861312400439/5506224861312400439_signature_31.png
| 2023-01-27 | 2023-01-27 09:06 | Holy spirit health center | Catalina st q.c | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23156 | DOH | triciamarie435@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 441 ASSAY: Xpert MTB/RIF Ultra, COUNT: 369 | 2023-01-20 | Gx4 | 831586 | For xpert check and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Xpert check and maintenance | Operational all module passed |  | Xpertcheck kits | 9531 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506223551317642929/5506223551317642929_signature_28.png | Patricia Marie C. Tabug | https://www.jotform.com/uploads/macarewecare/222927978372067/5506223551317642929/5506223551317642929_signature_31.png
| 2023-01-27 | 2023-01-27 09:06 | Holy spirit health center | Catalina st q.c | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23155 | DOH | triciamarie435@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 441 ASSAY: Xpert MTB/RIF Ultra, COUNT: 369 | 2023-01-20 | Gx4 | 831586 | For xpert check and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Xpert check and maintenance | Operational all module passed |  | Xpertcheck kits | 9531 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506223481318807760/5506223481318807760_signature_28.png | Patricia Marie C. Tabug | https://www.jotform.com/uploads/macarewecare/222927978372067/5506223481318807760/5506223481318807760_signature_31.png
| 2023-01-27 | 2023-01-26 15:34 | Payatas b health center | Bulacan st payatas qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23154 | DOH | racquelmrocero@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 110 ASSAY: Xpert MTB/RIF Ultra, COUNT: 959 | 2023-01-20 | GX4 | 831899 | For xpertcheck and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Xpert check and maintenance | A2 failed xpertcheck,A3 retry xpertcheck |  | Xpertcheck kits | 9527 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506222781317441283/5506222781317441283_signature_28.png | Racquel M.Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5506222781317441283/5506222781317441283_signature_31.png
| 2023-01-27 | 2023-01-26 15:34 | Payatas b health center | Bulacan st payatas qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23153 | DOH | racquelmrocero@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 110 ASSAY: Xpert MTB/RIF Ultra, COUNT: 959 | 2023-01-20 | GX4 | 831899 | For xpertcheck and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Xpert check and maintenance | A2 failed xpertcheck,A3 retry xpertcheck |  | Xpertcheck kits | 9527 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506222691317941829/5506222691317941829_signature_28.png | Racquel M.Rocero | https://www.jotform.com/uploads/macarewecare/222927978372067/5506222691317941829/5506222691317941829_signature_31.png
| 2023-01-27 | 2023-01-26 12:20 | Commonwealth health center | Katuparan st, Commonwealth q.c | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23152 | DOH | commonwealthclinical@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 212 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1152 ASSAY: Xpert HPV, COUNT: 487 | 2023-01-20 | GX4 | 831615 | For xpert check and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Xpert check and maintenance | A3 reported already, A1,A2,A4 passed xpertcheck |  | Xpertcheck kits | 9528 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506221691318099588/5506221691318099588_signature_28.png | Rizza B. Clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5506221691318099588/5506221691318099588_signature_31.png