| 2023-01-27 | 2023-01-26 12:20 | Commonwealth health center | Katuparan st, Commonwealth q.c | NCR | NCR 2nd - Quezon City | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23151 | DOH | commonwealthclinical@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 212 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1152 ASSAY: Xpert HPV, COUNT: 487 | 2023-01-20 | GX4 | 831615 | For xpert check and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Xpert check and maintenance | A3 reported already, A1,A2,A4 passed xpertcheck |  | Xpertcheck kits | 9528 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506221521317244624/5506221521317244624_signature_28.png | Rizza B. Clemente | https://www.jotform.com/uploads/macarewecare/222927978372067/5506221521317244624/5506221521317244624_signature_31.png
| 2023-01-27 | 2023-01-26 08:52 | San Mateo rhu V | Modesta village san mateo | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23150 | PBSP | mmacahis@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 233 |  | Gx4 | 838161 | For maintenance,xpert check | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Maintenance and xpertcheck | All passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506220311317513593/5506220311317513593_signature_28.png | Monina Rica M. De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5506220311317513593/5506220311317513593_signature_31.png
| 2023-01-27 | 2023-01-26 08:52 | San Mateo rhu V | Modesta village san mateo | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23149 | PBSP | mmacahis@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 233 |  | Gx4 | 838161 | For maintenance,xpert check | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Maintenance and xpertcheck | All passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506220121317705108/5506220121317705108_signature_28.png | Monina Rica M. De Guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5506220121317705108/5506220121317705108_signature_31.png
| 2023-01-27 | 2023-01-27 14:43 | Don Emilio Valle Memorial Hospital | Ubay, Bohol | 7 | Bohol | Tb Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 23148 | PBSP | dedvmhlaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1909 | 2023-01-27 | Gx4 | 807603 | 4008 | Replaced module | Replace new module | Operational and 4 modules available |  | Module | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506111681024747862/5506111681024747862_signature_28.png | Sharmaine Pagolong | https://www.jotform.com/uploads/macarewecare/222927978372067/5506111681024747862/5506111681024747862_signature_31.png
| 2023-01-27 | 2023-01-27 14:43 | Don Emilio Valle Memorial Hospital | Ubay, Bohol | 7 | Bohol | Tb Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 23147 | PBSP | dedvmhlaboratory@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1909 | 2023-01-27 | Gx4 | 807603 | 4008 | Replaced module | Replace new module | Operational and 4 modules available |  | Module | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506111631023239096/5506111631023239096_signature_28.png | Sharmaine Pagolong | https://www.jotform.com/uploads/macarewecare/222927978372067/5506111631023239096/5506111631023239096_signature_31.png
| 2023-01-27 | 2023-01-27 14:04 | Santa Rosa RHU | Santa Rosa, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23146 | DOH | rhu1santarosa@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 157 ASSAY: Xpert MTB/RIF Ultra, COUNT: 84 | 2023-01-27 | Genexpert IV | 831599 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Unit is for Calibration this February | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506096233927054854/5506096233927054854_signature_28.png | Ana S. Caballong | https://www.jotform.com/uploads/macarewecare/222927978372067/5506096233927054854/5506096233927054854_signature_31.png
| 2023-01-27 | 2023-01-27 12:24 | Bansalan Rural Health Unit | Poblacion 2, Bansalan | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23145 | PBSP | ellen68nebrada@gmail.com | 29.9 | ASSAY: Xpert MTB/RIF, COUNT: 196 ASSAY: Xpert MTB/RIF Ultra, COUNT: 422 | 2023-01-27 | GX-IV | 838181 | For Calibration | Perform Calibration | Run XpertCheck Calibration | Module B1, B2 : Passed | High Temperature | Xpertcheck CE 5 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506028777168667219/5506028777168667219_signature_28.png | Ellen Nebrada Mayor RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5506028777168667219/5506028777168667219_signature_31.png
| 2023-01-27 | 2023-01-27 11:44 | Dauin Rural Health Unit | Dauin rhu | 7 | Negros Oriental | Laboratory | YES | MAINTENANCE Transfer New Location | CEPHEID | Jovy Noel |  |  |  |  | 23144 | PBSP | dkannahmae@gmail.com | 28 | ASSAY: , COUNT: 131 | 2023-01-27 | GX4 | 835059 | Transfer New Location | Transfer New Location | PM | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5506008476038220048/5506008476038220048_signature_28.png | Kannag Mae Dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5506008476038220048/5506008476038220048_signature_31.png
| 2023-01-27 | 2023-01-26 15:36 | San Fernando City Health Office (Dolores RHU) | Dolores, San Fernando, Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23143 | PBSP | celpineda1973@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 508 ASSAY: Xpert MTB/RIF Ultra, COUNT: 514 | 2023-01-26 | Genexpert IV | 804837 | Module 2 and 4 errors, UPS malfunction | Performed Xpertcheck on all modules Troubleshoot UPS | Performed Maintenance Gathered Reports Performed Calibration Xpertcheck A1 passed and A2, A3, A4 Service Required Performed Functionality Test | Unit is Operational | Ups for Replacement | Xpertcheck Kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505988924917901081/5505988924917901081_signature_28.png | Celia T. Pineda | https://www.jotform.com/uploads/macarewecare/222927978372067/5505988924917901081/5505988924917901081_signature_31.png
| 2023-01-27 | 2023-01-27 09:12 | Region 1 Medical Center | Arellano st., Dagupan City | 1 | Pangasinan | Molecular Lab | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23142 | PBSP | romarmtamayo@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF, COUNT: 619 ASSAY: Xpert MTB/RIF Ultra, COUNT: 886 ASSAY: Xpert SARS-CoV-2, COUNT: 4113 |  | GX16 | 738529 | Unit is Due for Calibration | Calibration | Calibration of the following Modules A1/883398 failed A2/883397 failed A3/883399 failed A4/874672 failed B1/670294 passed B2/883207 failed B3/883205 hardware fail B4/883450 hardware fail C1/883447 failed C2/883449 hardware failed C3/883448 failed C4/886691 hardware failed D1/886683 failed D2/617706 passed D3/883441 failed D4/883440 failed | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505926974913380933/5505926974913380933_signature_28.png | Romar Tamayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5505926974913380933/5505926974913380933_signature_31.png
| 2023-01-27 | 2023-01-27 09:15 | Eron Dumagieng | Pantal, Dagupan City | 1 | Pangasinan | eron dumagieng | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23141 | PBSP | romarmtamayo@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF, COUNT: 378 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1018 ASSAY: Xpert SARS-CoV-2, COUNT: 61 ASSAY: Xpert HIV-1 Viral Load, COUNT: 117 |  | GX4D | 804726 | Unit is due for Calibration | Calibration and Preventive Maintenance | Performed Calibration C1 / 643897 C2 / 640532 passed but become hardware failed after calibration C3 / 643898 C4/667002 Use additional 2 cartridge for the recall of C3 and C4 | Unit is operational | C1, C2, C3 and C4 needs service required/ module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505880274913870974/5505880274913870974_signature_28.png | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5505880274913870974/5505880274913870974_signature_31.png
| 2023-01-27 | 2023-01-27 09:15 | Eron Dumagieng | Pantal, Dagupan City | 1 | Pangasinan | eron dumagieng | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23140 | PBSP | romarmtamayo@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF, COUNT: 378 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1018 ASSAY: Xpert SARS-CoV-2, COUNT: 61 ASSAY: Xpert HIV-1 Viral Load, COUNT: 117 |  | GX4D | 804726 | Unit is due for Calibration | Calibration and Preventive Maintenance | Performed Calibration C1 / 643897 C2 / 640532 passed but become hardware failed after calibration C3 / 643898 C4/667002 Use additional 2 cartridge for the recall of C3 and C4 | Unit is operational | C1, C2, C3 and C4 needs service required/ module replacement |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505880164918266159/5505880164918266159_signature_28.png | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5505880164918266159/5505880164918266159_signature_31.png
| 2023-01-27 | 2023-01-27 08:01 | Cong. Lamberto L. Macias Memorial Hospital | Siaton | 7 | Negros Occidental | Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 23139 | DOH | cllmmhlaboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 177 | 2023-01-27 | GX4 | 838482 | Maintenance | Perform preventive maintenance | Pm | All Modules functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505873736825044163/5505873736825044163_signature_28.png | Marque Nicole M. Baldebrin | https://www.jotform.com/uploads/macarewecare/222927978372067/5505873736825044163/5505873736825044163_signature_31.png
| 2023-01-26 | 2023-01-26 19:38 | Mariano Marcos Memorial Hospital and Medical Center | San Julian, Batac City | 1 | Ilocos Norte | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23138 | PBSP | mmmhmc.molecularlab@gmail.com | 26.3 | ASSAY: Xpert HPV, COUNT: 157 ASSAY: Xpert MTB/RIF, COUNT: 74 ASSAY: Xpert MTB/RIF Ultra, COUNT: 198 ASSAY: Xpert SARS-CoV-2, COUNT: 1309 |  | Genexpert | 838168 | Unit is due for Module Replacement Unit is due for Calibration | Calibration and Module Replacement | Performed Module Replacement B4/612314 to B4/210008915 Performed Calibration B2/ 807925 Result: passed Performed Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505433424425888187/5505433424425888187_signature_28.png | Jaizel Aira Batara | https://www.jotform.com/uploads/macarewecare/222927978372067/5505433424425888187/5505433424425888187_signature_31.png
| 2023-01-26 | 2023-01-26 19:21 | Mariano Marcos Memorial and Medical Center | San Julian, Batac City | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23137 | DOH | mmmhmc.molecularlab@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF, COUNT: 196 ASSAY: Xpert HPV, COUNT: 531 ASSAY: Xpert MTB/RIF Ultra, COUNT: 771 ASSAY: Xpert SARS-CoV-2, COUNT: 3160 |  | GX4D | 831908 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Plunger Rod Cleaning Performed Filter Cleaning | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505422774427791719/5505422774427791719_signature_28.png | Jaizel Aira Batara | https://www.jotform.com/uploads/macarewecare/222927978372067/5505422774427791719/5505422774427791719_signature_31.png
| 2023-01-26 | 2023-01-26 14:55 | Toril A District Health Center | Rayas St. Toril Poblacion Davao City | 11 | Davao del Sur | TB DOTS Laboratory | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23136 | PBSP | leabuelis@yahoo.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 890 ASSAY: Xpert SARS-CoV-2, COUNT: 1069 | 2023-01-26 | GX-IV 4d | 838865 | For Calibration and Preventive Maintenance | Perfrom XPC Calibration And Preventive Maintenance | Perfrom XpertCheck Calibration, clean Surface Area, Cartridge Bay, Fan Filters, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Relocation of GeneXpert Position | Module B1 B2 B4 : Passed Module B3 : For Service Required Module B2 encounter multiple Temperature Errors | Relocation of GeneXpert, Assay Stat of Sars Cov from St. Elizabeth Hospital Inc | XPERTCHECK CE-5 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505322450912401905/5505322450912401905_signature_28.png | Lea Buelis, RMT, RN, ASCPI | https://www.jotform.com/uploads/macarewecare/222927978372067/5505322450912401905/5505322450912401905_signature_31.png
| 2023-01-26 | 2023-01-26 14:55 | Toril A District Health Center | Rayas St. Toril Poblacion Davao City | 11 | Davao del Sur | TB DOTS Laboratory | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23135 | PBSP | leabuelis@yahoo.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 890 ASSAY: Xpert SARS-CoV-2, COUNT: 1069 | 2023-01-26 | GX-IV 4d | 838865 | For Calibration and Preventive Maintenance | Perfrom XPC Calibration And Preventive Maintenance | Perfrom XpertCheck Calibration, clean Surface Area, Cartridge Bay, Fan Filters, Plunger Rod and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results Relocation of GeneXpert Position | Module B1 B2 B4 : Passed Module B3 : For Service Required Module B2 encounter multiple Temperature Errors | Relocation of GeneXpert, Assay Stat of Sars Cov from St. Elizabeth Hospital Inc | XPERTCHECK CE-5 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505322420918919567/5505322420918919567_signature_28.png | Lea Buelis, RMT, RN, ASCPI | https://www.jotform.com/uploads/macarewecare/222927978372067/5505322420918919567/5505322420918919567_signature_31.png
| 2023-01-26 | 2023-01-26 15:54 | Riverside Medical Center Inc | B.S Aquino Drive, Bacolod City | 6 | Negros Occidental | Laboratory | YES | MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23134 | COMMERCIAL | laboratory@rivermedcenter.net | 20 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 33 ASSAY: Xpert HPV, COUNT: 12 ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 19 ASSAY: Xpert HCV VL, COUNT: 40 |  | GX IV | 110005268 | Due date for preventive maintenance | Preventive Maintenance | Preventive Maintenance | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505291580319562604/5505291580319562604_signature_28.png | Dr. Pablo O.Torre Memorial Hosputal | https://www.jotform.com/uploads/macarewecare/222927978372067/5505291580319562604/5505291580319562604_signature_31.png
| 2023-01-26 | 2023-01-26 14:58 | Bicol Center for Health Development | Legazpi | 5 | Albay | Laboratory | NO | REPAIR | HAEMONETICS | John Monsour Doctor |  |  |  |  | 23133 | COMMERCIAL | versonolicia27@gmail.com | 23 |  | 2023-01-20 | MCS+ | 15H032G | Bios CRC test failed | Replaced defective CPU board. | Replaced CPU board (bios CRC test ok) Replaced urethane rollers Performed maintenance Performed calibration Checked centrifuge Check internal connections Verified calibration parameters | Bios CRC test ok after replacement of CPU board. Bowl could not be detected during calibration. | Recommend to pull out unir for further evaluation and problem isolation. | MCS+ CPU board assembly | SI No. 94654 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505272552481989043/5505272552481989043_signature_28.png | Verson Olicia | https://www.jotform.com/uploads/macarewecare/222927978372067/5505272552481989043/5505272552481989043_signature_31.png
| 2023-01-26 | 2023-01-26 14:00 | Romana Pangqn District Hospital | Porac Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23132 | PBSP | hayfamaealmario@gmail.com | 28.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 493 | 2023-01-11 | GXIV-D | 814378 | Module B3 for repoacement | Performed preventive maintenance Replqced defective module B3:691004 Replaced scanner Gathered reports Performed functionality test Old scanner SN:19255010556824 B4:691089 | Preventive maintenance and repair | Unit is operational | Replaced macare scanner | Genexpert module new sn: 751142 and new scanner sn:212770552309 | 38125 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505261827857653792/5505261827857653792_signature_28.png | Hayfa Mae Almario | https://www.jotform.com/uploads/macarewecare/222927978372067/5505261827857653792/5505261827857653792_signature_31.png
| 2023-01-26 | 2023-01-26 14:51 | Cordlife Philippines | Commonwealth Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | CALIBRATION | CYTIVA | Patrick Los Baños |  |  |  |  | 23131 | COMMERCIAL | katrina.castro@cordlife.com | 26 |  |  | Sepax 2 | 6220.4122 | None, for calibration | Performed unit calibration | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505256055564445118/5505256055564445118_signature_28.png | Julius Biteng | https://www.jotform.com/uploads/macarewecare/222927978372067/5505256055564445118/5505256055564445118_signature_31.png
| 2023-01-26 | 2023-01-26 14:51 | Cordlife Philippines | Commonwealth Ave., Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | YES | CALIBRATION | CYTIVA | Patrick Los Baños |  |  |  |  | 23130 | COMMERCIAL | katrina.castro@cordlife.com | 26 |  |  | Sepax 2 | 6220.4122 | None, for calibration | Performed unit calibration | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505256025568182166/5505256025568182166_signature_28.png | Julius Biteng | https://www.jotform.com/uploads/macarewecare/222927978372067/5505256025568182166/5505256025568182166_signature_31.png
| 2023-01-26 | 2023-01-26 14:39 | Cebu TB Reference Lab | B Rodriguez St. CEBU City | 7 | Cebu | Tb Laboratory | YES | INSTALLATION | CEPHEID | Mark Efren Ramas |  |  |  |  | 23129 | PBSP | tivalerio.pbsp@gmail.com | 22 | ASSAY: , COUNT: 1744, 4348,1717 | 2023-01-26 | Gx4 | 838263, 806201,838882 | Transfer only | Install the genxpert machine | Install Ups, Cpu, genxpert machine, printer, bar code reader, | Operational |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505248707228878888/5505248707228878888_signature_28.png | Titus Ian Valerii | https://www.jotform.com/uploads/macarewecare/222927978372067/5505248707228878888/5505248707228878888_signature_31.png
| 2023-01-26 | 2023-01-26 14:08 | Teresita Jalandoni Provincial Hospital | Silay, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | MAINTENANCE Upgrade version to 6.4 | CEPHEID | Raymond Laurito |  |  |  |  | 23128 | DOH | mrvillan.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 931 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2056 |  | GX IV | 838464 | Due date for Preventive Maintenance | Preventive Maintenance | Upgrade version to 6.4 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505238623511337255/5505238623511337255_signature_28.png | Ma. Arna R. Villan | https://www.jotform.com/uploads/macarewecare/222927978372067/5505238623511337255/5505238623511337255_signature_31.png
| 2023-01-26 | 2023-01-26 13:41 | Mexico RHU I | Parian, Mexico, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23127 | DOH | jasperdizon20@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 171 ASSAY: Xpert MTB/RIF Ultra, COUNT: 233 | 2023-01-26 | Genexpert IV | 831877 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operationql | For Calibration this March | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505213868792662991/5505213868792662991_signature_28.png | Jasper P. Dizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5505213868792662991/5505213868792662991_signature_31.png
| 2023-01-26 | 2023-01-26 12:18 | ST. PAUL’S HOSPITAL OF ILOILO | ILOILO CITY | 6 | Iloilo | BLOOD BANK | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Joseph Acebedo |  |  |  |  | 23126 | COMMERCIAL | sphi.bloodbank@yahoo.com | 25.3 |  |  | MCS+ | 98B020 | FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL | RECOMMENDATION: CONNECT TO UNINTERRUPTED POWER SUPPLY |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505164028272403445/5505164028272403445_signature_28.png | Den Marion Algarne | https://www.jotform.com/uploads/macarewecare/222927978372067/5505164028272403445/5505164028272403445_signature_31.png
| 2023-01-26 | 2023-01-26 11:26 | Paquibato District Health Center | Malabog Health Center, Davao City | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23125 | PBSP | abelongelfie@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF, COUNT: 233 ASSAY: Xpert MTB/RIF Ultra, COUNT: 16 | 2023-01-26 | GX-IV | 838538 | For Calibration | Perform XpertCheck Calibration | Perform XpertCheck Calibration | Module B1, B3 and B4 : Passed Module B2 : Service Required | N/a | Xpertcheck CE5 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505136700917088857/5505136700917088857_signature_28.png | Elfie Amor Abelong | https://www.jotform.com/uploads/macarewecare/222927978372067/5505136700917088857/5505136700917088857_signature_31.png
| 2023-01-26 | 2023-01-26 11:26 | Paquibato District Health Center | Malabog Health Center, Davao City | 11 | Davao del Sur | TB-DOTS | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23124 | PBSP | abelongelfie@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF, COUNT: 233 ASSAY: Xpert MTB/RIF Ultra, COUNT: 16 | 2023-01-26 | GX-IV | 838538 | For Calibration | Perform XpertCheck Calibration | Perform XpertCheck Calibration | Module B1, B3 and B4 : Passed Module B2 : Service Required | N/a | Xpertcheck CE5 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505136670913068918/5505136670913068918_signature_28.png | Elfie Amor Abelong | https://www.jotform.com/uploads/macarewecare/222927978372067/5505136670913068918/5505136670913068918_signature_31.png
| 2023-01-26 | 2023-01-26 11:03 | PHILIPPINE RED CROSS - ILOILO CHAPTER | ILOILO CITY | 6 | Iloilo | BLOOD BANK | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | Joseph Acebedo |  |  |  |  | 23123 | COMMERCIAL | renee_mari18@yahoo.com | 25 |  |  | MCS+ | 98G049 | -ERROR 466 WEIGHER PROBLEM -FOR CALIBRATION | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505116399278192355/5505116399278192355_signature_28.png | Girlie Marie Anas | https://www.jotform.com/uploads/macarewecare/222927978372067/5505116399278192355/5505116399278192355_signature_31.png
| 2023-01-26 | 2023-01-26 09:10 | Malilipot Rural Health Unit | Malilipot | 5 | Albay | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23122 | PBSP | faith.buban2124@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 526 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1727 |  | GX-IV-D | 808280 | Module B1 continuous booting. | Reseat internal connection. | Performed maintenance Reseat internal connection Generated reports | Module B1/655916 operational but under observation. Modules B2/625608 and B3/621226 for reporting. Module B4 operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107267902501647/5505107267902501647_signature_28.png | Fatima Buban, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107267902501647/5505107267902501647_signature_31.png
| 2023-01-26 | 2023-01-25 18:05 | Bicol Center for Health Development | Legazpi City | 5 | Albay | Laboratory | NO | REPAIR | HAEMONETICS | John Monsour Doctor |  |  |  |  | 23121 | COMMERCIAL | versonolicia27@gmail.com | 25 |  | 2023-01-20 | MCS+ | 15H032G | Bios CRC Test Failed | Replaced CPU board | Replaced CPU board Performed maintenance Calibrated the unit Troubleshoot bowl optics | Unit not operational. Bowl could not be detected. UPS not operational recommend to replace. | For further troubleshooting of bowl optics. | MCS+ CPU Board Assembly | SI No. 94654 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107267907324998/5505107267907324998_signature_28.png | Verson Olicia | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107267907324998/5505107267907324998_signature_31.png
| 2023-01-26 | 2023-01-25 14:02 | Tabaco City Health Unit | Tabaco | 5 | Albay | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23117 | PBSP | elyjoy17rmt@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 919 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2753 |  | GX-IV-D | 804257 | Module B1 and gateway board for replacement. | Replaced module B1 and gateway board | Performed maintenance Replaced module B1 Replaced gateway board. Generated reports | Module B1 operational. Module B2 and B4 operational but under observation. Module B3 not operational. | Old B1/640283 New B1/884721 |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107217905654276/5505107217905654276_signature_28.png | Ely Joy Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107217905654276/5505107217905654276_signature_31.png
| 2023-01-26 | 2023-01-25 10:19 | Sorsogon Provincial Hospital | Sorsogon City | 5 | Sorsogon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23118 | PBSP | wizard682000@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 75 ASSAY: Xpert MTB/RIF Ultra, COUNT: 820 |  | GX-IV-2-D | 835213 | Error 4008 on module B2 | Performed troubleshooting on module B2 | Swap module B2 to B1 Checked module fuse Check internal connections Performed maintenance Generated reports | Module B1 operational. Module B2 not operational. | Module B2/888907 reported already. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107217907728599/5505107217907728599_signature_28.png | Reynaldo B. Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107217907728599/5505107217907728599_signature_31.png
| 2023-01-26 | 2023-01-25 08:57 | Sorsogon Medical Mission Group Hospital and Health Services Cooperative | Sorsogon City | 5 | Sorsogon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23119 | PBSP | lotino.ana@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 590 |  | GX-IV-D | 707452 | Stuck cartridge on Module B2 | Removed stuck cartridge | Manual removal of stuck cartridge Performed maintenance Cleaned cartridge bay, Cleaned plunger rod and icore Generated reports. | Modules B1 and B2 operational. Modules B3 and B4 not operational. | B3/619380 and B4/644792 already reported. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107227906987485/5505107227906987485_signature_28.png | Ana Liza Lotino | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107227906987485/5505107227906987485_signature_31.png
| 2023-01-26 | 2023-01-24 15:08 | Bicol Diagnostic Reference Laboratory | Legazpi | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23120 | PBSP | rbalzaga.pbsp@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 6 ASSAY: Others, COUNT: 85 ASSAY: Xpert MTB/RIF Ultra, COUNT: 360 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 |  | GX-IV-D | 808636 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod and filter. | Modules B2 and B4 operational. Modules B1/617846 and B3/671780 not operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107227907278884/5505107227907278884_signature_28.png | Richard B. Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107227907278884/5505107227907278884_signature_31.png
| 2023-01-26 | 2023-01-24 16:03 | Bicol Diagnostic Reference Laboratory | Legazpi | 5 | Albay | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23116 | PBSP | rbalzaga.pbsp@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 805 ASSAY: Xpert MTB/RIF, COUNT: 520 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1954 ASSAY: Xpert HIV-1 Viral Load, COUNT: 163 |  | GX-IV-D | 807579 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod and filter. Generated reports | All modules operational. | Module D3/686316 has error 2014. Under observation. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107217906768550/5505107217906768550_signature_28.png | Richard B. Alzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5505107217906768550/5505107217906768550_signature_31.png
| 2023-01-26 | 2023-01-26 09:53 | Bacolod City Health Office | Bacolod City, Negros Occidental | 6 | Negros Occidental | Laboratory | YES | MAINTENANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23115 | DOH | cyrilcordova0515@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 824 |  | GX IV | 837612 | Due date for PM | Cleaning fan filter, clean module bay, PM | Cleaning fan filter, clean module bay, PM | Systems Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5505090144349331752/5505090144349331752_signature_28.png | Cyril Cordova | https://www.jotform.com/uploads/macarewecare/222927978372067/5505090144349331752/5505090144349331752_signature_31.png
| 2023-01-25 | 2023-01-25 09:07 | Rodriguez rhu | Rodriguez rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23114 | PBSP | deleonmykaela18@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 205 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2299 ASSAY: Xpert HIV-1 Viral Load, COUNT: 89 | 2023-01-23 | Gx4 | 838894 | For xpertcheck and Module replacement B2 862901 - 881532 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit performed module replacement | Module replacement,Maintenance | B4 module failed xpert check,b1,b2,b3passed |  | 900gx | G5184 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504498768677645605/5504498768677645605_signature_28.png | Mykaela Josepha De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5504498768677645605/5504498768677645605_signature_31.png
| 2023-01-25 | 2023-01-25 09:07 | Rodriguez rhu | Rodriguez rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23113 | PBSP | deleonmykaela18@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 205 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2299 ASSAY: Xpert HIV-1 Viral Load, COUNT: 89 | 2023-01-23 | Gx4 | 838894 | For xpertcheck and Module replacement B2 862901 - 881532 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit performed module replacement | Module replacement,Maintenance | B4 module failed xpert check,b1,b2,b3passed |  | 900gx | G5184 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504498718674105632/5504498718674105632_signature_28.png | Mykaela Josepha De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5504498718674105632/5504498718674105632_signature_31.png
| 2023-01-25 | 2023-01-25 09:18 | Rodriguez rhu | Montalban rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23112 | PBSP | deleonmykaela18@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 211 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2101 |  | Gx4 | 807580 | For xpert check and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpertcheck | Xpert check and maintenance | B1 and B4 failed,B2,B3 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504498268678144348/5504498268678144348_signature_28.png | Mykaela Josepha De leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5504498268678144348/5504498268678144348_signature_31.png
| 2023-01-25 | 2023-01-25 09:18 | Rodriguez rhu | Montalban rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23111 | PBSP | deleonmykaela18@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 211 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2101 |  | Gx4 | 807580 | For xpert check and maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpertcheck | Xpert check and maintenance | B1 and B4 failed,B2,B3 passed |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504498218672681011/5504498218672681011_signature_28.png | Mykaela Josepha De leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5504498218672681011/5504498218672681011_signature_31.png
| 2023-01-25 | 2023-01-25 13:42 | Antipolo city health office | Antipolo city | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23110 | PBSP | gwencb0720@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 1317 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7507 ASSAY: Xpert HIV-1 Viral Load, COUNT: 463 | 2023-01-23 | Gx4 | 808629 | For module replacement d4 668006(old) - d4 808034(new) | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed module replacement | Maintenance, module replacement | Operational |  | 900gx | G5107 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504497658672737677/5504497658672737677_signature_28.png | Gwendolyn Batingan | https://www.jotform.com/uploads/macarewecare/222927978372067/5504497658672737677/5504497658672737677_signature_31.png
| 2023-01-25 | 2023-01-25 13:42 | Antipolo city health office | Antipolo city | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23109 | PBSP | gwencb0720@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 1317 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7507 ASSAY: Xpert HIV-1 Viral Load, COUNT: 463 | 2023-01-23 | Gx4 | 808629 | For module replacement d4 668006(old) - d4 808034(new) | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed module replacement | Maintenance, module replacement | Operational |  | 900gx | G5107 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504497578676762011/5504497578676762011_signature_28.png | Gwendolyn Batingan | https://www.jotform.com/uploads/macarewecare/222927978372067/5504497578676762011/5504497578676762011_signature_31.png
| 2023-01-25 | 2023-01-25 17:39 | Sasa District Health Center | KM.9 Paradise Road, Sasa, Davao City | 11 | Davao del Sur | Laboratory | NO | CALIBRATION | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23108 | PBSP | mariarosario_mapayo@yahoo.com.sg | 25.7 | ASSAY: Xpert MTB/RIF, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 294 ASSAY: Xpert SARS-CoV-2, COUNT: 1048 | 2023-01-25 | GX-IV | 838889 | For Calibration | Perfrom XpertCheck Calibration | Calibrate the Module B1 s/n: 863324 B3: 863784 | Module B1 s/n: 863324 failed B3 s/n: 863784 retry and failed | Module B1 and B3 under observation | Xpertcheck CE-5 | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504491720651435628/5504491720651435628_signature_28.png | Maria Rosario Pino, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5504491720651435628/5504491720651435628_signature_31.png
| 2023-01-25 | 2023-01-24 16:14 | National children hospital | E.rodriguez ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 23107 | MACARE-HBDC | arellanojohneric09@gmail.com | 28 |  |  | Mcs + | 17H632SPG | For maintenance | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504490718677986259/5504490718677986259_signature_28.png | John Eric Arellano | https://www.jotform.com/uploads/macarewecare/222927978372067/5504490718677986259/5504490718677986259_signature_31.png
| 2023-01-25 | 2023-01-24 16:14 | National children hospital | E.rodriguez ave qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 23106 | MACARE-HBDC | arellanojohneric09@gmail.com | 28 |  |  | Mcs + | 17H632SPG | For maintenance | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504490468676462096/5504490468676462096_signature_28.png | John Eric Arellano | https://www.jotform.com/uploads/macarewecare/222927978372067/5504490468676462096/5504490468676462096_signature_31.png
| 2023-01-25 | 2023-01-25 15:43 | Philippine Red Cross - Ormoc Chapter | Ormoc City | 8 | Leyte | Laboratory | NO | REPAIR MAINTENANCE | HAEMONETICS | John Paul Labutap |  |  |  |  | 23105 | COMMERCIAL | hannah.declaro@redcross.org.ph | 24 |  | 2023-01-25 | MCS+ | 15H043G | For replacement of bloop pump and transfer pump rubber roller | #NAME? | #NAME? | Equipment is Operational |  | 4 new rubber roller | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504421806815378495/5504421806815378495_signature_28.png | hannah de claro | https://www.jotform.com/uploads/macarewecare/222927978372067/5504421806815378495/5504421806815378495_signature_31.png
| 2023-01-25 | 2023-01-25 15:43 | Philippine Red Cross - Ormoc Chapter | Ormoc City | 8 | Leyte | Laboratory | NO | REPAIR MAINTENANCE | HAEMONETICS | John Paul Labutap |  |  |  |  | 23104 | COMMERCIAL | hannah.declaro@redcross.org.ph | 24 |  | 2023-01-25 | MCS+ | 15H043G | For replacement of bloop pump and transfer pump rubber roller | #NAME? | #NAME? | Equipment is Operational |  | 4 new rubber roller | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504421666814448361/5504421666814448361_signature_28.png | hannah de claro | https://www.jotform.com/uploads/macarewecare/222927978372067/5504421666814448361/5504421666814448361_signature_31.png
| 2023-01-25 | 2023-01-25 15:00 | Mabalacat Rural Health 1 | Mabalacat RHU | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23103 | PBSP | perezriyadh@gmail.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 375 ASSAY: Xpert MTB/RIF Ultra, COUNT: 905 ASSAY: Xpert HIV-1 Viral Load, COUNT: 21 |  | GXIV-D | 807564 | Due for calibration Temperature error(2014,4009) modules B1 and B3 HW Failed for Module B2 Error 2008 for module B4 | Check-up of genexpert machine | Preventive maintenance and check-up | Unit is not operational modules B1-B4 for reporting |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504397678223185783/5504397678223185783_signature_28.png | Riyadh L. Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5504397678223185783/5504397678223185783_signature_31.png
| 2023-01-25 | 2023-01-25 14:49 | Nasipit Rural Health Unit RTDL | Brgy. Talisay, Nasipit | 13 | Agusan del Norte | RTDL | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 23102 | DOH | peterdolotina@gmail.com | 23.73 | ASSAY: Xpert MTB/RIF, COUNT: 249 ASSAY: Xpert MTB/RIF Ultra, COUNT: 995 |  | GX-4 | 837611 | None | None | Preventive Maintenance Data Collection | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504393049708455201/5504393049708455201_signature_28.png | Peter C. Dolotina | https://www.jotform.com/uploads/macarewecare/222927978372067/5504393049708455201/5504393049708455201_signature_31.png
| 2023-01-25 | 2023-01-25 14:16 | Cabanatuan CHO | Kapt. Pepe Subdivision, Cabanatuan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23101 | DOH | cholab2021@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1103 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1587 | 2023-01-25 | Genexpert IV - 6 colors | 831592 | N/a | N/a | Performed Maintenance ( Cleaned Icore, Cleaned Cartridge Bay, Cleaned Filters, Cleaned Plunger Rod) Performed Calibration Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck Result A1, A2, A3 passed the Xpertcheck | Xpertcheck Kit | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504380848795025518/5504380848795025518_signature_28.png | Maricris So | https://www.jotform.com/uploads/macarewecare/222927978372067/5504380848795025518/5504380848795025518_signature_31.png
| 2023-01-25 | 2023-01-25 13:10 | Malapatan Rural Health Unit | Malapatan, Sarangani | 12 | Sarangani | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23100 | DOH | maribelcaps@yahoo.com | 25.9 | ASSAY: Xpert MTB/RIF, COUNT: 127 ASSAY: Xpert MTB/RIF Ultra, COUNT: 116 |  | Gx-4 | 838419 | Module B2 and B3 h/w failed | Perform Genexpert machine Preventive maintenance | Perform Genexpert machine Preventive maintenance | Gx machine has 2 operational modules | Module B2 and B4, waiting for module to available |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504354433402353981/5504354433402353981_signature_28.png | Maribel M. Capilitan | https://www.jotform.com/uploads/macarewecare/222927978372067/5504354433402353981/5504354433402353981_signature_31.png
| 2023-01-25 | 2023-01-25 12:58 | Victorias City Health Office | Victorias City, Negros Occiddental | 6 | Negros Occidental | Laboratory | YES | REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 23099 | DOH | vincent_llamoleonor@yahoo.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 644 ASSAY: Xpert MTB/RIF Ultra, COUNT: 848 |  | GXIV | 838141 | Defective module 4 sn608861 | New module 4 replacement sn210041123 | Module replacement | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504324039222412300/5504324039222412300_signature_28.png | Vincent Leonor | https://www.jotform.com/uploads/macarewecare/222927978372067/5504324039222412300/5504324039222412300_signature_31.png
| 2023-01-25 | 2023-01-25 11:21 | Zamboanga City Medical Center | Dr evangelista street, zamboanga city, zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 23098 | PBSP | zcmctblab@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 304 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4167 | 2023-01-25 | GX16 | 839309 | None | Module Replacement | Replaced module 663088 with module 634434 | Module C4, D4, D1 and A4 are hardware fail. | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504313202865342050/5504313202865342050_signature_28.png | Aira Faith S. Runas | https://www.jotform.com/uploads/macarewecare/222927978372067/5504313202865342050/5504313202865342050_signature_31.png
| 2023-01-25 | 2023-01-25 09:52 | Mataas na Kahoy Rural Health Unit | Brgy II, Mataas na Kahoy | 4A | Batangas | Laboratory | NO | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23097 | PBSP | dmlescano1219@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 55 ASSAY: Xpert MTB/RIF Ultra, COUNT: 282 | 2023-01-23 | GXIV | 838325 | Unit for preventive maintenance. Unit for Xpertcheck. For module replacement of module B3. | Performed preventive maintenance. Performed module replacement on B3. Performed Xpertcheck. Generated reports. | Preventive maintenance, Xpertcheck and Module replacement. | Unit is operational. |  | GX module for B3(210068186) in replacement to defective module B3(860135). | G5112 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504283053963764095/5504283053963764095_signature_28.png | Debbie Lescano | https://www.jotform.com/uploads/macarewecare/222927978372067/5504283053963764095/5504283053963764095_signature_31.png
| 2023-01-25 | 2023-01-25 11:25 | Malvar Rural Health Unit | Malvar | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Shiedrick Daroya |  |  |  |  | 23096 | PBSP | lopezsheryl13@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 77 ASSAY: Xpert MTB/RIF Ultra, COUNT: 329 | 2023-01-09 | GXIV | 838191 | Unit for preventive maintenance. Unitmfor Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance amd Xpertcheck. | Unit is operational. All modules passed xpertcheck. |  | Xpertcheck | Monarch DR0037973 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504283053968651240/5504283053968651240_signature_28.png | Sheryl E. Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5504283053968651240/5504283053968651240_signature_31.png
| 2023-01-25 | 2023-01-25 09:41 | Cadiz City Health Office | Cadiz City, Negros Occidental | 6 | Negros Occidental | Laboratory | NO | Gathering data | CEPHEID | Raymond Laurito |  |  |  |  | 23095 | DOH | zhevebacoko@gmail.com | 20.9 | ASSAY: Xpert MTB/RIF, COUNT: 236 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1059 |  | GX IV | 838499 | Defective module 1 | Gathering data | Gathering data | Defective module 1 |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504270492112012703/5504270492112012703_signature_28.png | Zherlyn Eve Bacomo | https://www.jotform.com/uploads/macarewecare/222927978372067/5504270492112012703/5504270492112012703_signature_31.png
| 2023-01-25 | 2023-01-25 11:28 | Sarangani Health Care Facility | Alabel, Sarangani | 12 | Sarangani | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23093 | DOH | ronayadao@yahoo.com | 23.2 | ASSAY: Xpert MTB/RIF, COUNT: 365 ASSAY: Xpert MTB/RIF Ultra, COUNT: 634 |  | Gx-4 | 838542 | Module B3 - for Service Required Module B4 - for Service Required | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck at module B3 and B4 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Genexpert machine Calibration/Xpertcheck at module B3 and B4 for troubleshooting | Gx machine has 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504270429411429952/5504270429411429952_signature_28.png | Ronora G. Yadao | https://www.jotform.com/uploads/macarewecare/222927978372067/5504270429411429952/5504270429411429952_signature_31.png
| 2023-01-25 | 2023-01-25 10:47 | GUIMARAS PROVINCIAL HEALTH OFFICE | JORDAN, GUIMARAS | 6 | Guimaras | PHO LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23092 | PBSP | mylene210@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 261 ASSAY: Xpert MTB/RIF Ultra, COUNT: 651 |  | GX-IV | 807984 | #NAME? | -PERFORM MODULE B4 DEF SN: 664558 WITH NEW SN:649852 | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504254849479487813/5504254849479487813_signature_28.png | Mylene Publico | https://www.jotform.com/uploads/macarewecare/222927978372067/5504254849479487813/5504254849479487813_signature_31.png
| 2023-01-25 | 2023-01-25 10:40 | GUIMARAS PROVINCIAL HEALTH OFFICE | SAN MIGUEL,JORDAN GUIMARAS | 6 | Guimaras | PHO LABORATORY | NO | MAINTENANCE | CEPHEID | Joseph Acebedo |  |  |  |  | 23091 | PBSP | mylene210@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 261 ASSAY: Xpert MTB/RIF Ultra, COUNT: 651 |  | GX-IV | 838291 | PREVENTIVE MAINTANCE | #NAME? | -COLLECT IQR, SYSTEM LOG, AND ASSAY STAT -CONNECT NETWORK WITH GX MACHINE SN 807984 | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504254819472636243/5504254819472636243_signature_28.png | Mylene Publico | https://www.jotform.com/uploads/macarewecare/222927978372067/5504254819472636243/5504254819472636243_signature_31.png
| 2023-01-25 | 2023-01-25 09:55 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23090 | DOH | ammegzedd@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 587 ASSAY: Xpert MTB/RIF Ultra, COUNT: 401 | 2023-01-25 | GXIV | 838421 | Module B3 for expertcheck | #NAME? | #NAME? | #NAME? | Module B3 for service required | Xpertcheck Cartridge | LOT 05203 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5504214129157289720/5504214129157289720_signature_28.png | Gemma Malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5504214129157289720/5504214129157289720_signature_31.png
| 2023-01-24 | 2023-01-24 15:10 | San Luis RHU | San Luis, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23089 | DOH | mohan_santos@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 90 ASSAY: Xpert MTB/RIF Ultra, COUNT: 151 | 2023-01-24 | Genexpert IV | 831714 | N/a | N/a | Peformed Maintenance Gathered Reports Install HIV - Viral Load Performed Functionality Test | Unit is Operational | Unit is for Calibration this March | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503639253916426373/5503639253916426373_signature_28.png | Mohan S.Pangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5503639253916426373/5503639253916426373_signature_31.png
| 2023-01-24 | 2023-01-24 15:10 | San Luis RHU | San Luis, Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23088 | DOH | mohan_santos@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 90 ASSAY: Xpert MTB/RIF Ultra, COUNT: 151 | 2023-01-24 | Genexpert IV | 831714 | N/a | N/a | Peformed Maintenance Gathered Reports Install HIV - Viral Load Performed Functionality Test | Unit is Operational | Unit is for Calibration this March | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503639203913443888/5503639203913443888_signature_28.png | Mohan S.Pangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5503639203913443888/5503639203913443888_signature_31.png
| 2023-01-24 | 2023-01-24 13:19 | Hilongos Rural Health Unit | Hilongos Leyte | 8 | Leyte | Hilongos RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23087 | PBSP | leovillaflor@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 435 ASSAY: Xpert MTB/RIF Ultra, COUNT: 480 | 2023-01-24 | GXIV | 838143 | #NAME? | #NAME? | #NAME? | #NAME? |  | Xpert Check Lot 1000378669 | XPERTCHECK-CE-5 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503578063466070716/5503578063466070716_signature_28.png | Leonides Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5503578063466070716/5503578063466070716_signature_31.png
| 2023-01-24 | 2023-01-24 15:51 | Mati City Health Office | Madang, Poblacion, Mati City | 11 | Davao Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23086 | PBSP | jaslimjay@gmail.com | 21.5 | ASSAY: Xpert MTB/RIF, COUNT: 344 ASSAY: Xpert MTB/RIF Ultra, COUNT: 479 | 2023-01-24 | GX-IV - 2D | 835221 | Module B1 (S/N: 787801) subjected for replacement For Preventive Maintenance | Replaced the Defective Module B1 S/N: 787801 Perform Preventive Maintenance | Replaced the Defective Module B1 (S/N : 787801) with new Module S/N: 210068172 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is fully operational | N/a | 900-0491rs | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503563536281070183/5503563536281070183_signature_28.png | Jasper Jay Lim | https://www.jotform.com/uploads/macarewecare/222927978372067/5503563536281070183/5503563536281070183_signature_31.png
| 2023-01-24 | 2023-01-24 15:26 | DOH Treatment and Rehabilitation Center | Camp Bagong Diwa, Lower Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Eron Dumagieng |  |  |  |  | 23085 | PBSP | dohtrcbicutan.lab@yahoo.com | 26.2 | ASSAY: Xpert MTB/RIF, COUNT: 503 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1336 |  | GX4D | 807452 | Unit is due for troubleshooting | Troubleshooting ang diagnosting. | Troubleshooting of Module B3/ 801912, UPS and GX power Supply Performed Voltage Reading at UPS: While Plugged = 230 volts While Unplugged/ Backup output voltage = 215 volts Performed Voltage Reading at GX Main Power Board: L1 = 23 volts L2 = 23 volts Performed Continuity Testing at Module B3/ 801912 with Reference in B4/ 666170 B3 = infinity reading / open Ciruit (Busted Fuse) B4 = 0 ohms /short Circuit ( Good Fuse) Performed swapping of gateway lan of B3 and B4 to ensure same error will provide at other interface. Result: B3 at B4 hub provides same error 4008 | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503553191025502297/5503553191025502297_signature_28.png | April Chakee A. Javier | https://www.jotform.com/uploads/macarewecare/222927978372067/5503553191025502297/5503553191025502297_signature_31.png
| 2023-01-24 | 2023-01-24 14:57 | UP Manila - National Institutes of Health | Pedro Gil St, Manila | NCR | NCR 1st - Manila | IPS | NO | INSTALLATION | CARL ZEISS | Patrick Los Baños |  |  |  |  | 23084 | COMMERCIAL | gtdelatorre@up.edu.ph | 26 |  |  | Axioscope 5 | 3366001264 | None, for installation | None | #NAME? | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503531153718734978/5503531153718734978_signature_28.png | Gerwin Dela Torre | https://www.jotform.com/uploads/macarewecare/222927978372067/5503531153718734978/5503531153718734978_signature_31.png
| 2023-01-24 | 2023-01-24 14:13 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology Laboratory | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23082 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 125 ASSAY: Xpert MTB/RIF, COUNT: 1437 ASSAY: Xpert MTB/RIF Ultra, COUNT: 661 |  | GXIV-4D | 838285 | NA | NA | #NAME? | All four modules are operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503508515123478517/5503508515123478517_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5503508515123478517/5503508515123478517_signature_31.png
| 2023-01-24 | 2023-01-24 14:19 | Vicente Sotto Memorial Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23083 | DOH | vsmmcdeptpatho.microbiology@gmail.com | 27.3 | ASSAY: Xpert HPV, COUNT: 134 ASSAY: Xpert MTB/RIF, COUNT: 1197 ASSAY: Xpert MTB/RIF Ultra, COUNT: 621 |  | GXIV-4D | 820206 | None | None | #NAME? | All modules are functional. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503508515128628897/5503508515128628897_signature_28.png | Glydemaere Lecias | https://www.jotform.com/uploads/macarewecare/222927978372067/5503508515128628897/5503508515128628897_signature_31.png
| 2023-01-24 | 2023-01-24 13:57 | Talakag Rural Health Unit | Talakag | 10 | Bukidnon | TB DOTS | NO | CALIBRATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23081 | DOH | jhudielkathleen@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 156 |  | GeneXpert | 838566 | Failed modules | Recalibrating of failed modules | Calibration of modules B1 and B2 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503497654327249128/5503497654327249128_signature_28.png | Jhudiel Kathleen Agol | https://www.jotform.com/uploads/macarewecare/222927978372067/5503497654327249128/5503497654327249128_signature_31.png
| 2023-01-24 | 2023-01-24 12:50 | Bayambang District Hospital | Bayambang Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23080 | PBSP | bevzneri@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 679 | 2023-01-09 | GXIV-D | 812830 | B1 for module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Preventive maintenance and repair | Unit is operational | Unit is operational | Genexpert module | 37919 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503473702984398751/5503473702984398751_signature_28.png | Beverly Neri | https://www.jotform.com/uploads/macarewecare/222927978372067/5503473702984398751/5503473702984398751_signature_31.png
| 2023-01-24 | 2023-01-24 12:50 | Bayambang District Hospital | Bayambang Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23079 | PBSP | bevzneri@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 1 ASSAY: Xpert MTB/RIF Ultra, COUNT: 679 | 2023-01-09 | GXIV-D | 812830 | B1 for module replacement | Performed surface cleaning Performed preventive maintenance Performed module replacement Gathered reports Performed functionality test | Preventive maintenance and repair | Unit is operational | Unit is operational | Genexpert module | 37919 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503473682989430947/5503473682989430947_signature_28.png | Beverly Neri | https://www.jotform.com/uploads/macarewecare/222927978372067/5503473682989430947/5503473682989430947_signature_31.png
| 2023-01-24 | 2023-01-23 13:57 | Davao De Oro Provincial Hospital - Montevista | Bankeruhan, Montevista Davao De Oro | 11 | Davao de Oro | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23078 | PBSP | cecillejagurilles1122@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF, COUNT: 365 ASSAY: Xpert MTB/RIF Ultra, COUNT: 747 | 2023-01-23 | GX-IV | 808624 | Modules B1 (S/N: 618211) and B3 (S/N: 603402) for replacement For Preventive Maintenance | Module Replacement and Perform Preventive Maintenance | Replace the defective Module B1 (S/N:618211) with new module S/N :210068376 and Defective Module B3 (S/N:618211) with new Module S/N: 210068322 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is fully operational | N/A | 900-0491rs | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503472692219294889/5503472692219294889_signature_28.png | Cecille Jagurilles | https://www.jotform.com/uploads/macarewecare/222927978372067/5503472692219294889/5503472692219294889_signature_31.png
| 2023-01-24 | 2023-01-24 12:56 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Laboratory | YES | CALIBRATION TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 23076 | DOH | dulzna@gmail.com | 25.7 | ASSAY: Xpert MTB/RIF, COUNT: 402 ASSAY: Xpert MTB/RIF Ultra, COUNT: 308 |  | Gx-4 | 838570 | Module B4- for Service Required | Perform Genexpert machine Calibration/ Xpertcheck at module B4 for troubleshooting | Perform Genexpert machine Calibration/ Xpertcheck at module B4 for troubleshooting | Genexpert machine has 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503472415424751836/5503472415424751836_signature_28.png | Dulce Aurora N. Ariston, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5503472415424751836/5503472415424751836_signature_31.png
| 2023-01-24 | 2023-01-13 10:50 | Balamban Rural Health Unit | Balamban Cebu | 7 | Cebu | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 23075 | DOH | balambanrhu45@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 543 | 2023-01-13 | GX4 | 838231 | Preventive Maintenance | Perform preventive maintenance for all modules | Perform preventive maintenance | All modules are functional |  |  |  | Jovy D. Noel | https://www.jotform.com/uploads/macarewecare/222927978372067/5503412714912609343/5503412714912609343_signature_28.png | Mildred E. Climaco | https://www.jotform.com/uploads/macarewecare/222927978372067/5503412714912609343/5503412714912609343_signature_31.png
| 2023-01-24 | 2023-01-24 11:10 | Cebu TB Reference Lab | Cebu City | 7 | Cebu | Laboratory | YES | REPAIR | CEPHEID | Raymond Laurito |  |  |  |  | 23074 | PBSP | tivalerio.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1766 |  | Gx iv | 806201 | Defective module 1 sn601011 | Replace new module for module 1 sn670417 | New module 1 | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503392647461510781/5503392647461510781_signature_28.png | Titus Ian Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5503392647461510781/5503392647461510781_signature_31.png
| 2023-01-24 | 2023-01-24 10:39 | Rosario Super Health Center | Ortigas Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23073 | PBSP | eden.cawayan@yahoo.com | 26.3 | ASSAY: Xpert MTB/RIF, COUNT: 288 ASSAY: Xpert MTB/RIF Ultra, COUNT: 990 ASSAY: Xpert SARS-CoV-2, COUNT: 349 | 2023-01-23 | GX4D | 843924 | Unit is due for module replacement | Module Replacement Preventive Maintenance | Performed Module Replacement B4/ 893035 into B4/ 210068082 Performed Preventive Maintenance | Unit is operational |  | Genexpert module 6C / sn: 210068082 | G 5110 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503375765125920708/5503375765125920708_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5503375765125920708/5503375765125920708_signature_31.png
| 2023-01-24 | 2023-01-24 09:10 | Salvacion Oppus-Yniguez Memorial Provincial Hospital | Maasin City | 8 | Southern Leyte | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 23072 | PBSP | loreensalva@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 906 ASSAY: Xpert MTB/RIF Ultra, COUNT: 376 | 2023-01-24 | GXIV | 804262 | B3 for Module replacement | -Module replacement B3 (640302 into 688345) -Performed preventive maintenance -Generate reports | #NAME? | Operational |  | GeneXpert Module 6 Color (SN: 688345) | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5503336330866050155/5503336330866050155_signature_28.png | Loreen S. Agapay | https://www.jotform.com/uploads/macarewecare/222927978372067/5503336330866050155/5503336330866050155_signature_31.png
| 2023-01-23 | 2023-01-20 14:11 | East avenue medical center | East avenue q.c | NCR | NCR 2nd - Quezon City | Tb dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23071 | PBSP | eamcdots@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 154 ASSAY: Xpert MTB/RIF, COUNT: 1268 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2723 | 2023-01-20 | Gx 4 | 803087 | For maintenance,xpert check | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Pms,Xpertcheck | C4 661735 and C1 611387 failed xpertcheck,C2,C3 unable to retry lack of xpertcheck kits,C1 | C2 and c3 for retry xpertcheck lack of kits | Xpertcheck kits | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502823108678614596/5502823108678614596_signature_28.png | rose ann parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5502823108678614596/5502823108678614596_signature_31.png
| 2023-01-23 | 2023-01-20 14:11 | East avenue medical center | East avenue q.c | NCR | NCR 2nd - Quezon City | Tb dots | NO | CALIBRATION MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23070 | PBSP | eamcdots@gmail.com | 29 | ASSAY: Xpert HPV, COUNT: 154 ASSAY: Xpert MTB/RIF, COUNT: 1268 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2723 | 2023-01-20 | Gx 4 | 803087 | For maintenance,xpert check | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed xpert check | Pms,Xpertcheck | C4 661735 and C1 611387 failed xpertcheck,C2,C3 unable to retry lack of xpertcheck kits,C1 | C2 and c3 for retry xpertcheck lack of kits | Xpertcheck kits | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502823048676000264/5502823048676000264_signature_28.png | rose ann parnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5502823048676000264/5502823048676000264_signature_31.png
| 2023-01-23 | 2023-01-17 10:14 | National Tuberculosis Reference Laboratory | Filinvest | NCR | NCR 4th - Muntinlupa | Laboratory | NO | CALIBRATION | CEPHEID | [5 empty] | 23069 | PBSP | mspogosa.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1167 |  | GXIV-D | 807606 | Due for XpertCheck | Xpertcheck | Performed Maintenance Performed Xpertcheck Generate Reports | Unit is operational | 1Failed: B4/619472 |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5502722274119491418/5502722274119491418_signature_28.png | Maria Guileane S. Pogosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5502722274119491418/5502722274119491418_signature_31.png
| 2023-01-23 | 2023-01-18 14:57 | Cagayan Valley Medical Center | Tuguegarao City, Cagayan | 3 | Cagayan | Blood Bank | NO | MAINTENANCE TROUBLESHOOTING | HAEMONETICS | [5 empty] | 23068 | DOH | loleniamelad@ymail.com | 28 |  | 2023-01-18 | MCS+ | ISH0336 | N/a | Replaced Rollers on red Adjust the line sensor but not working Clean the machine | Performed Maintenance | Unit is Operational | Line sensor output value is low, 2pcs rollers for white is for replacement | Rollers 2pcs | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5502686552261284914/5502686552261284914_signature_28.png | Lolenia D. Melaf | https://www.jotform.com/uploads/macarewecare/222927978372067/5502686552261284914/5502686552261284914_signature_31.png
| 2023-01-23 | 2023-01-23 14:09 | Candaba RHU | Candaba, Pampangq | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Matthew Roxas |  |  |  |  | 23067 | DOH | rochelannlontoc@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 139 | 2023-01-23 | Genexpert IV | 831873 | Module A1 errors | Performed Xpertcheck on A1 module - service required | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck A1 - Service Required | Xpertcheck 1pc | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502638182267629027/5502638182267629027_signature_28.png | Rochel ann lontoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5502638182267629027/5502638182267629027_signature_31.png
| 2023-01-23 | 2023-01-13 15:30 | Santa Rosa RHU | Santa Rosa, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 23065 | DOH | rhu1santarosa@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 167 ASSAY: Xpert MTB/RIF Ultra, COUNT: 84 | 2023-01-13 | Genexpert IV | 831599 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5502633182268696355/5502633182268696355_signature_28.png | Ana Caballong | https://www.jotform.com/uploads/macarewecare/222927978372067/5502633182268696355/5502633182268696355_signature_31.png
| 2023-01-23 | 2023-01-16 15:26 | Sto. Domingo District Hospital | Sto. Domingo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 23066 | DOH | samatracyndrell@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 277 ASSAY: Xpert MTB/RIF Ultra, COUNT: 598 | 2023-01-16 | Genexpert IV | 831582 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5502633182269434967/5502633182269434967_signature_28.png | Cyndrell B. Samatra | https://www.jotform.com/uploads/macarewecare/222927978372067/5502633182269434967/5502633182269434967_signature_31.png
| 2023-01-23 | 2023-01-19 12:26 | Macalelon Rural Health Unit | Macalelon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 23063 | DOH | cheyster7388@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 399 |  | GXIV-D | 831636 | Creating Test Problem Barcode Scanner Problem | Troubleshoot Backup barcode scanner | Performed Maintenance Checked Functionality of Barcode Scanner Generate Reports | Barcode Scanner is defective. Modules are operational | Def: 18303010506478/ backup: 1514000502883 |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5502601991212720643/5502601991212720643_signature_28.png | Cheyster O. Agustin | https://www.jotform.com/uploads/macarewecare/222927978372067/5502601991212720643/5502601991212720643_signature_31.png
| 2023-01-23 | 2023-01-18 11:43 | Lian Rural Health Unit | Lian | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 23062 | PBSP | malc_rmt_09@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 170 | 2023-01-16 | GXIV-D - 2 | 835179 | Barcode Scanner Problem | Replacement | Replaced Barcode Scanner Performed functionality test Performed Maintenance Generate Reports | Unit is operational | 19075010504478 --> 21277010555478 | GX barcode scanner with stand | G 5047 | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5502601991215827106/5502601991215827106_signature_28.png | Mary Anne C. Sison | https://www.jotform.com/uploads/macarewecare/222927978372067/5502601991215827106/5502601991215827106_signature_31.png
| 2023-01-23 | 2023-01-20 09:21 | Damariñas City Health Office I | D. Mangubat, Dasmariñas | 4A | Cavite | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23064 | PBSP | cho1lab@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1633 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 |  | GXIV-D | 838893 | Unit is due for XpertCheck | XpertCheck | Performed Maintenance Performed XpertCheck Generate Reports Performed Self Test | Unit is operational. All modules passed xpertcheck. | Module 4 newly replaced |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5502601991216868092/5502601991216868092_signature_28.png | Elma R.Belostrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5502601991216868092/5502601991216868092_signature_31.png
| 2023-01-23 | 2023-01-23 10:41 | Urdaneta District Hospital | Urdaneta Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23061 | PBSP | warlyacosta@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 1076 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1360 | 2022-08-17 | GXIV-D | 805109 | Temperature error (1002) on modules B2 and B3 | Performed preventive maintenance Archived test results Performed calibration Gathered reports Performed functionality test | Preventive maintenance and calibration | Unit is operational except modulea B2 and B3 | 2 modules passed calibration, modules B2 and B3 for reporting | 2 Xpertcheck kits used | 34323 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502566182315981742/5502566182315981742_signature_28.png | Warlito Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5502566182315981742/5502566182315981742_signature_31.png
| 2023-01-23 | 2023-01-23 12:04 | Lambayong Rural Health Unit | Lambayong, Sultan kudarat | 12 | Sultan Kudarat | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23060 | PBSP | jeanalonzo1973@yahoo.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 115 ASSAY: Xpert MTB/RIF Ultra, COUNT: 314 |  | Gx-4 | 838159 | N/A | Genexpert machine Preventive maintenance Data collection for Calibration/Xpertcheck | Genexpert machine Preventive maintenance Data collection for Calibration/Xpertcheck | Gx machine has 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502564054336582244/5502564054336582244_signature_28.png | Jean K. Alonzo | https://www.jotform.com/uploads/macarewecare/222927978372067/5502564054336582244/5502564054336582244_signature_31.png
| 2023-01-23 | 2023-01-20 10:17 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23059 | PBSP | atumbaga.1079@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 992 ASSAY: Xpert SARS-CoV-2, COUNT: 5760 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GX XVI | 839292 | unit for calibration | Performed calibration | #NAME? | Unit is operational. All modules except B4 failed calibration due to calibration overdue / insufficient xpertcheck kit. For reporting. |  |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5502552071217948739/5502552071217948739_signature_28.png | Allan Tumbaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5502552071217948739/5502552071217948739_signature_31.png
| 2023-01-23 | 2023-01-20 10:17 | Sta. Ana Hospital | Sta. Ana, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23058 | PBSP | atumbaga.1079@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1105 ASSAY: Xpert MTB/RIF Ultra, COUNT: 992 ASSAY: Xpert SARS-CoV-2, COUNT: 5760 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GX XVI | 839292 | unit for calibration | Performed calibration | #NAME? | Unit is operational. All modules except B4 failed calibration due to calibration overdue / insufficient xpertcheck kit. For reporting. |  |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551481211275948/5502551481211275948_signature_28.png | Allan Tumbaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551481211275948/5502551481211275948_signature_31.png
| 2023-01-23 | 2023-01-20 14:24 | Mabini Rural Health Unit | Mabini Batangas | 4A | Batangas | Laboratory | NO | INSTALLATION CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23057 | PBSP | glee.maravilla@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 54 ASSAY: Xpert MTB/RIF Ultra, COUNT: 894 | 2023-01-16 | GXIV | 838867 | Unit for preventive maintenance. Unit for Xpertcheck. For installation of new workstation. | Performed installation of workstation. Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance, installation and Xpertcheck. | Unit is operational. Module B1(863439) and B2(863440) service required. |  | Lenovo CPU (CPH1121F28404) in replacement to defective Lenovo CPU (CPH0420F08168). | G 5046 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551221214792901/5502551221214792901_signature_28.png | Gliceria C. Maravilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551221214792901/5502551221214792901_signature_31.png
| 2023-01-23 | 2023-01-20 12:19 | Batangas Medical Center | Batangas City | 4A | Batangas | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 23056 | PBSP | batmc.tbculturelab@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 54 ASSAY: Xpert MTB/RIF, COUNT: 497 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1487 ASSAY: Xpert SARS-CoV-2, COUNT: 2379 ASSAY: Xpert HIV-1 Viral Load, COUNT: 554 | 2023-01-16 | GX16 | 844120 | Unit for preventive maintenance. For module replacement of A1 and A4. | Performed preventive maintenance. Performed module replacement on A1 and A4. | Preventive maintenance and module replacement. | Unit is operational. For reporting of module D4, H/W Failure. 15 working modules. |  | GX Module A1(644387) and A4(647244) in replacement for defective A1(892205) and A4(892244) | G5028 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551221215195372/5502551221215195372_signature_28.png | Patricia Louisse A. Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551221215195372/5502551221215195372_signature_31.png
| 2023-01-23 | 2023-01-19 14:02 | Calamba Medical Center | Calamba Laguna | 4A | Laguna | Microbiology | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23055 | COMMERCIAL | laboratory@cmc.ph | 29 | ASSAY: Xpert CT/NG, COUNT: 4 ASSAY: Xpert HPV, COUNT: 7 ASSAY: Xpert MTB/RIF, COUNT: 89 ASSAY: Xpert MTB/RIF Ultra, COUNT: 383 | 2023-01-19 | GXII | 812916 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B2 (686424) service required). |  | Xpertcheck c/o facility. | N/a | Shiedrick Daroy | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551221218123447/5502551221218123447_signature_28.png | Wilbert Siat | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551221218123447/5502551221218123447_signature_31.png
| 2023-01-23 | 2023-01-19 10:18 | Cabuyao City Health Office 1 | Cabuyao Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23054 | PBSP | shirleypalcantara18@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 154 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1160 | 2023-01-11 | GXIV | 812794 | Unit for preventive maintenance. Unit for Xpertcheck. For checking of Barcode Scanner. | Performed preventive maintenance. Performed Xpertcheck. Performed troubleshooting of barcode scanner. Generated reports. | Preventive maintenance, troubleshooting and Xpertcheck. | Unit is operational. Barcode scanner is malfunctioning and recommended for replacement. Modules B1, B3 and B4 H/W, unable to calibrate. Module B2 passed Xpertcheck. 1 module operational. |  | Xpertcheck | N/a | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551221212580126/5502551221212580126_signature_28.png | Shirley Alcantara | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551221212580126/5502551221212580126_signature_31.png
| 2023-01-23 | 2023-01-19 09:17 | Cabuyao City Health Office 2 | Cabuyao Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23049 | DOH | carmelaguanzon@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 297 | 2023-01-16 | GXIV | 821003 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. All modules passed Xpertcheck. |  | Xpertcheck | DR 9469 | Sheidrick | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171211114681/5502551171211114681_signature_28.png | Carmela G. Cantillon | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171211114681/5502551171211114681_signature_31.png
| 2023-01-23 | 2023-01-18 09:20 | Navotas Medical and Wellness Center | M. Naval St., Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23051 | PBSP | navotascitylab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 517 | 2023-01-18 | GXIV | 807592 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated report. | Preventive maintenance and Xpertcheck. | Unit is operational. Module B3 (665858), service required). |  | Xpertcheck. | N/a | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171215647627/5502551171215647627_signature_28.png | Ma. Arlene Teodoro | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171215647627/5502551171215647627_signature_31.png
| 2023-01-23 | 2023-01-18 09:15 | Navotas Medical and Wellness Center | M. Naval St., Navotas | NCR | NCR 3rd - Navotas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23050 | DOH | navotascitylab@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 517 | 2023-01-16 | GXIV | 831626 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Module A3 (693447), service required. |  | Xpertcheck | DR 9472 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171215966871/5502551171215966871_signature_28.png | Ma. Arlene Teodoro | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171215966871/5502551171215966871_signature_31.png
| 2023-01-23 | 2023-01-17 09:06 | Calaanan Health Center | Calaanan Caloocan City | NCR | NCR 3rd - Caloocan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23053 | DOH | rosaleslorina0917@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 161 ASSAY: Xpert MTB/RIF Ultra, COUNT: 623 | 2023-01-16 | GXIV | 831892 | Unit for PM. Unit for Xpertcheck. | Performed PM of GeneXpert. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Umit is operational. All modules passed Xpertcheck. |  | Xpertcheck Kit | DR 9471 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171216250240/5502551171216250240_signature_28.png | Lorina Rosales | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171216250240/5502551171216250240_signature_31.png
| 2023-01-23 | 2023-01-18 10:37 | Juan Posadas Health Center | Rodriguez St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23052 | PBSP | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 511 ASSAY: Xpert MTB/RIF Ultra, COUNT: 769 | 2023-01-18 | GXIV | 805060 | Unit for preventive maintenance. Unit for Xpertcheck. | Performed preventive maintenance. Performed Xpertcheck. Generated reports. | Preventive maintenance and Xpertcheck. | Unit is operational. Modules B3 (654188) and B4 (619323), service required. |  | Xpertcheck. | N/a | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171217322944/5502551171217322944_signature_28.png | Eliza junio | https://www.jotform.com/uploads/macarewecare/222927978372067/5502551171217322944/5502551171217322944_signature_31.png
| 2023-01-23 | 2023-01-12 15:36 | Amparo Health Center | Amparo, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 23048 | PBSP | jjpsjasperjoyce@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 239 ASSAY: Xpert MTB/RIF Ultra, COUNT: 800 |  | GXIV | 838873 | Unit for maintenance | Performed maintenance | #NAME? | All four modules are working. |  |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901213329111/5502550901213329111_signature_28.png | Jasper Joyce P. Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901213329111/5502550901213329111_signature_31.png
| 2023-01-23 | 2023-01-19 12:12 | Kamuning Health Center | Kamuning, Q.C. | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 23047 | DOH | vicvic.chua25@yahoo.com | 28 | ASSAY: Others, COUNT: 40) ASSAY: Xpert MTB/RIF, COUNT: 108 ASSAY: Xpert MTB/RIF Ultra, COUNT: 372 |  | GXIV | 831612 | A1 consecutive errors. A2, A3 temperature errors | Performed XpertCheck | #NAME? | Unit is operational. Module A1, A3 for reporting | Advised end-user to cover up windows to avoid GX sunlight exposure |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901214860972/5502550901214860972_signature_28.png | Jesusa Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901214860972/5502550901214860972_signature_31.png
| 2023-01-23 | 2023-01-18 13:06 | Manila Central University - Hospital | Caloocan City | NCR | NCR 3rd - Caloocan | Clinical Laboratory | NO | INSTALLATION | CEPHEID | [5 empty] | 23044 | COMMERCIAL | lab_mcuhospital@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 110018008 | UPS for installation | Installed UPS | - uninstalled temporary UPS - installed new UPS ( Brand: Eaton Model: NV 2000H SN: 83E500080) - performed back up mode test (Good) - Generated reports | Unit is operational. |  |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901215465278/5502550901215465278_signature_28.png | Gloria Caling | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901215465278/5502550901215465278_signature_31.png
| 2023-01-23 | 2023-01-18 14:11 | Potrero Health center | Malabon, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 23046 | PBSP | ailyn.santos19@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 289 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1357 ASSAY: Xpert HIV-1 Viral Load, COUNT: 18 |  | GXIV | 804715 | Unit for calibration | Performed calibration | #NAME? | Unit is operational. All modules passed calibration |  |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901217153170/5502550901217153170_signature_28.png | Ailyn Santos | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901217153170/5502550901217153170_signature_31.png
| 2023-01-23 | 2023-01-20 09:42 | Victor R. Potentiano Medical Center | Mandaluyong | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 23045 | CONSORTIUM | oro_janet@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 178 ASSAY: Xpert MTB/RIF Ultra, COUNT: 177 |  | GXIV | 835600 | Unit for maintenance | Performed maintenance | #NAME? | Unit is operational. No hardware problems found. Unit is due for annual calibration (last date of calibration 5-7-21). | Errors 5007 found on B3 will be forwarded to PAS |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901218570394/5502550901218570394_signature_28.png | Janet Oro | https://www.jotform.com/uploads/macarewecare/222927978372067/5502550901218570394/5502550901218570394_signature_31.png
| 2023-01-23 | 2023-01-23 10:01 | Cebu City Health office | Cebu city | 7 | Cebu | Laboratory | NO | TROUBLESHOOTING | CEPHEID | Raymond Laurito |  |  |  |  | 23043 | PBSP | laureanjo@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1700 |  | Gx iv | 808620 | stucked up carthridge in module 1 | Removing stucked up carthridge in module 1 | Systems operational | Good and operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502487410113227192/5502487410113227192_signature_28.png | Laurean Jo F. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5502487410113227192/5502487410113227192_signature_31.png
| 2023-01-23 | 2023-01-20 11:07 | Pasig City General Hospital | M. Eusebio St. | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23041 | PBSP | maniquizederlina@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 812 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1911 | 2023-01-16 | GX-IV-D | 805010 | Module B2 for replacement | Replaced module B2 | Performed maintenance Replaced module B2 Generated reports | All modules operational |  | Old B2/694296 New B2/637510 | G 5027 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521212290640/5502406521212290640_signature_28.png | Ederlina M. Sagales | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521212290640/5502406521212290640_signature_31.png
| 2023-01-23 | 2023-01-19 13:26 | Pangil Rural Health Unit | Pangil Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23040 | PBSP | tacile10@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 85 ASSAY: Xpert MTB/RIF Ultra, COUNT: 209 |  | GX-IV-D | 807661 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, Cleaned plunger, cleaned icore and filter Generated reports | All modules operational. |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521213698288/5502406521213698288_signature_28.png | Cecile R. Ssgritalo | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521213698288/5502406521213698288_signature_31.png
| 2023-01-23 | 2023-01-20 15:23 | Antipolo City Health Office | M. Santos St. | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23042 | PBSP | acho.laboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1344 ASSAY: Xpert MTB/RIF Ultra, COUNT: 7442 ASSAY: Xpert HIV-1 Viral Load, COUNT: 445 |  | GX-XVI-D | 842226 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore and plunger rod. Generated reports | All modules operational. | Module B4/886430 under observation. With error 4012. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521214379322/5502406521214379322_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521214379322/5502406521214379322_signature_31.png
| 2023-01-23 | 2023-01-19 11:40 | Kalayaan Rural Health Unit | Kalayaan | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | John Monsour Doctor |  |  |  |  | 23038 | DOH | gemmponce11@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 155 ASSAY: Xpert MTB/RIF Ultra, COUNT: 495 | 2022-09-29 | GX-IV-D | 821292 | For XpertCheck. Barcode scanner not working. | Run XpertCheck Checked barcose scanner | Performed maintenance Run XpertCheck Run reliability test Checked barcode scanner Generated reports | Modules A3 and 4 passed the XpertCheck and operational. Module A2 passed the XpertCheck but under observation. Module A1 not operational. Barcode scanner not operational. | Module A1/792273 already reported. | XpertCheck-CE-5 | DR#9215 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521218486264/5502406521218486264_signature_28.png | Gemma M.Ponce | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521218486264/5502406521218486264_signature_31.png
| 2023-01-23 | 2023-01-20 15:17 | Antipolo City Health Office | M. Santos St. | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 23039 | PBSP | acho.laboratory@gmail.com | 25 | ASSAY: Others, COUNT: N/A | 2023-01-16 | GX-IV-D | 808629 | Module 2 and 3 for replacement | Replaced defective module. | Performed maintenance Replaced defective module. Generated reports | Modules 1, 2, and 3 operational. Module 4 not operational. | Module 4 confirmed for replacement. | Old D2/668003, D3/668005 New D2/667219, D3/686854 | G 5033 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521218888862/5502406521218888862_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5502406521218888862/5502406521218888862_signature_31.png
| 2023-01-22 | 2023-01-22 22:05 | Chong Hua Hospital | Fuente Osmena, Cebu | 7 | Cebu | Bloodbank | YES | REPAIR | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 23037 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 24 |  | 2023-01-22 | MCS+ | 11J219 | Disposable ID Reader not Detected | Replaced Disposable ID with Cable | Replaced Disposable ID with cable | Operational | NA | Stock from Backup MCS+ | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5502058492023959818/5502058492023959818_signature_28.png | Alyanna casinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5502058492023959818/5502058492023959818_signature_31.png
| 2023-01-20 | 2023-01-20 15:31 | Gamu RHU | Gamu, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23036 | PBSP | gamugeneralclinicallaboratory@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 761 ASSAY: Xpert MTB/RIF Ultra, COUNT: 547 | 2023-01-19 | Genexpert IV | 833385 | UPS not Functioning | Troubleshoot the UPS | Performed Maintenance Performed Module Replacement Gathered Reports Performed Functionality Test | Unit is Operational | UPS for Replacement | Genexpert Module - 210068171 | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5500095186012257618/5500095186012257618_signature_28.png | Ederlyn S. Paracale | https://www.jotform.com/uploads/macarewecare/222927978372067/5500095186012257618/5500095186012257618_signature_31.png
| 2023-01-20 | 2023-01-20 08:47 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23035 | MACARE-HBDC | ruellamorena@gmail.com | 28 | ASSAY: Xpert SARS-CoV-2, COUNT: 445 | 2023-01-20 | Genexpert IV | 110008457 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5500090594526113703/5500090594526113703_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5500090594526113703/5500090594526113703_signature_31.png
| 2023-01-20 | 2023-01-20 09:19 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23034 | DOH | ruellamorena@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 680 ASSAY: Xpert MTB/RIF, COUNT: 1159 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2339 ASSAY: Xpert HIV-1 Viral Load, COUNT: 223 | 2023-01-20 | Genexpert IV | 831718 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5500090594529841652/5500090594529841652_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5500090594529841652/5500090594529841652_signature_31.png
| 2023-01-20 | 2023-01-20 15:22 | Malay Municipal Health Office | Malay, Aklan | 6 | Aklan | Clinical Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | Wynnona Reoma |  |  |  |  | 23033 | DOH | rhedal30@yahoo.com | 27.22 | ASSAY: Xpert MTB/RIF, COUNT: 401 ASSAY: Xpert MTB/RIF Ultra, COUNT: 240 | 2023-01-20 | GXIV-4D | 838457 | None | None | -Preventive Maintenance -Transfer GeneXpert to a well ventilated area (33C to 27C temp log) | All modules are operational. | -Previous GeneXpert area has a module temp ranging 32-34C | None | None |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5500090117194462181/5500090117194462181_signature_28.png | Redal D. Taytay | https://www.jotform.com/uploads/macarewecare/222927978372067/5500090117194462181/5500090117194462181_signature_31.png
| 2023-01-20 | 2023-01-20 14:11 | East Rembo Health Center | 22nd Street, East Rembo | NCR | NCR 4th - Makati | TB Laboratory | NO | MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23032 | PBSP | eastrembosatellitelaboratory@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF, COUNT: 1132 ASSAY: Xpert MTB/RIF Ultra, COUNT: 372 ASSAY: Xpert HIV-1 Viral Load, COUNT: 53 |  | GX4D | 807591 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Performed Plunger Rod Cleaning Performed Filter Cleaning | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5500046834385830761/5500046834385830761_signature_28.png | Blessed Joy M. Serrano | https://www.jotform.com/uploads/macarewecare/222927978372067/5500046834385830761/5500046834385830761_signature_31.png
| 2023-01-20 | 2023-01-20 13:55 | San Simon Rural Health Unit | San Jose, San simon Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23031 | PBSP | pingolseven@yahoo.com | 29.5 | ASSAY: Xpert MTB/RIF, COUNT: 92 ASSAY: Xpert MTB/RIF Ultra, COUNT: 156 | 2023-01-09 | GXIV-D | 838895 | Due for calibration | Performed preventive maintenance Archived test results Install Xpertcheck application Performed calibration Gathered reports Performed functionality test | Preventive maintenance and calibration | Unit is operational | All modules passee calibration | Calibration Kit 1 box | 38031 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5500046562315306859/5500046562315306859_signature_28.png | Virmar F. Pingol | https://www.jotform.com/uploads/macarewecare/222927978372067/5500046562315306859/5500046562315306859_signature_31.png
| 2023-01-20 | 2023-01-20 10:46 | National kidney and transplant institute | East avenue q.c | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | Noel Floresca |  |  |  |  | 23030 | MACARE-HBDC | molecularlabnkti@gmail.com | 29 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 754 ASSAY: Xpert CT/NG, COUNT: 1 ASSAY: Others, COUNT: 50 ASSAY: Xpert HBV VL, COUNT: 319 ASSAY: Xpert HPV, COUNT: 6 ASSAY: Xpert MTB/RIF, COUNT: 564 ASSAY: Xpert SARS-CoV-2, COUNT: 670 ASSAY: Xpert HCV VL, COUNT: 67 | 2023-01-20 | Gx4 | 838524 | B3 module 862548 emcounter continous 1002 and 4009 temp error | Performed re xpertcheck,Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance and troubleshooting | B1,B2,B4 operational B3 failed rexpertcheck |  | Xpertcheck kits | 0 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5500005283011300815/5500005283011300815_signature_28.png | Roina Estandarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5500005283011300815/5500005283011300815_signature_31.png
| 2023-01-20 | 2023-01-20 13:00 | Chong Hua Hospital | Fuente Osmena, Cebu City | 7 | Cebu | BLOODBANK | YES | REPAIR | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 23029 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 20 |  | 2023-01-20 | MCS+ | 11J219 | Disposable ID cannot be detected UPP 996E kit | Replaced Disposable ID reader | Install the Disposable ID bar code reader from the standby Machine | MCS+ can detect the Disposable ID |  | Disposable bar Code reader | Spare parts from the standby Mcs+ |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5500004734321104849/5500004734321104849_signature_28.png | Dave Gabini | https://www.jotform.com/uploads/macarewecare/222927978372067/5500004734321104849/5500004734321104849_signature_31.png
| 2023-01-20 | 2023-01-20 09:19 | Cagayan Valley Medical Center | Carig Sur, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 23028 | DOH | ruellamorena@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 680 ASSAY: Xpert MTB/RIF, COUNT: 1159 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2339 ASSAY: Xpert HIV-1 Viral Load, COUNT: 223 | 2023-01-20 | Genexpert IV | 831718 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499911884717705038/5499911884717705038_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5499911884717705038/5499911884717705038_signature_31.png
| 2023-01-20 | 2023-01-18 13:28 | DOH R02 Regional Blood Center - Cagayan Valley | Tuguegarao City | 2 | Cagayan | Bloodbank | NO | INSTALLATION MAINTENANCE | HAEMONETICS | Matthew Roxas |  |  |  |  | 23027 | DOH | acostaadelle@gmail.com | 27 |  | 2023-01-18 | MCS+ | 19G146SP6 | N/a | N/a | Assembled the stand Installed the MCS+ Run diagnostic Performed Maintenance | Unit is operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499911084713624817/5499911084713624817_signature_28.png | Adelina C. Acosta | https://www.jotform.com/uploads/macarewecare/222927978372067/5499911084713624817/5499911084713624817_signature_31.png
| 2023-01-20 | 2023-01-20 08:42 | Don Carlos Rural Health Unit | Don Carlos | 10 | Bukidnon | Laboratory | YES | Module Replacement | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23026 | PBSP | gibsz.l002@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 585 |  | GeneXpert | 807625 | H/W Fail on module C4/734479 | Module Replacement | Replaced module New Module S/N: C4/771729 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499891090915333049/5499891090915333049_signature_28.png | Loro, Gilbert E. | https://www.jotform.com/uploads/macarewecare/222927978372067/5499891090915333049/5499891090915333049_signature_31.png
| 2023-01-20 | 2023-01-20 09:20 | DOH Treatment and Rehabilitation Center | NCRPO, Lower Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23025 | PBSP | dohrrcbicutan.lab@yahoo.com | 26.3 | ASSAY: Xpert MTB/RIF, COUNT: 506 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1329 |  | GX4D | 807452 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Plunger Rod Cleaning Performed Filter Cleaning | Unit is Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499875566713156386/5499875566713156386_signature_28.png | Hellen Joy Rivera | https://www.jotform.com/uploads/macarewecare/222927978372067/5499875566713156386/5499875566713156386_signature_31.png
| 2023-01-20 | 2023-01-20 08:58 | Tagum City Health Office | Mabini St. Magupo, Tagum City | 11 | Davao del Norte | TB-DOTS Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23024 | DOH | azzi_belle06@yahoo.com | 27.6 | ASSAY: Xpert MTB/RIF, COUNT: 230 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1594 | 2023-01-20 | GX-IV | 838553 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Unit is fully operational | Air-conditioning unit not cold enough, for follow up | N/A | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499863672403733645/5499863672403733645_signature_28.png | Izzabelle Bartolome | https://www.jotform.com/uploads/macarewecare/222927978372067/5499863672403733645/5499863672403733645_signature_31.png
| 2023-01-20 | 2023-01-20 08:10 | Talisay District Hospital | Talisay City Cebu | 7 | Cebu | Tb Lab | YES | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 23023 | DOH | csmcmicro@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 693 | 2023-01-20 | GX4 | 838559 | Lost Communication | Arrange Lan cable | Arrange Lan cable and bar code scanner cable | All modules are operational |  | Na | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499830253669391446/5499830253669391446_signature_28.png | Joseph Lorenzo Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/5499830253669391446/5499830253669391446_signature_31.png
| 2023-01-20 | 2023-01-20 08:10 | Talisay District Hospital | Talisay City Cebu | 7 | Cebu | Tb Lab | YES | REPAIR TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 23022 | DOH | csmcmicro@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 693 | 2023-01-20 | GX4 | 838559 | Lost Communication | Arrange Lan cable | Arrange Lan cable and bar code scanner cable | All modules are operational |  | Na | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499829173663782069/5499829173663782069_signature_28.png | Joseph Lorenzo Mercado | https://www.jotform.com/uploads/macarewecare/222927978372067/5499829173663782069/5499829173663782069_signature_31.png
| 2023-01-19 | 2023-01-19 19:26 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23021 | PBSP | vincentmilestone20@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF, COUNT: 151 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2811 ASSAY: Xpert HIV-1 Viral Load, COUNT: 230 ASSAY: Xpert SARS-CoV-2, COUNT: 10422 | 2023-01-19 | GX-IV | 834440 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is fully operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499411531875713847/5499411531875713847_signature_28.png | Jenrh Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5499411531875713847/5499411531875713847_signature_31.png
| 2023-01-19 | 2023-01-19 19:26 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23020 | PBSP | vincentmilestone20@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF, COUNT: 151 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2811 ASSAY: Xpert HIV-1 Viral Load, COUNT: 230 ASSAY: Xpert SARS-CoV-2, COUNT: 10422 | 2023-01-19 | GX-IV | 834440 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is fully operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499411501875132233/5499411501875132233_signature_28.png | Jenrh Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5499411501875132233/5499411501875132233_signature_31.png
| 2023-01-19 | 2023-01-19 19:23 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23019 | PBSP | vincentmilestone20@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF, COUNT: 151 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2811 ASSAY: Xpert SARS-CoV-2, COUNT: 10422 ASSAY: Xpert HIV-1 Viral Load, COUNT: 230 | 2023-01-19 | GX-IV | 805063 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Unit is fully operational | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499410621876126536/5499410621876126536_signature_28.png | Jenry ken vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5499410621876126536/5499410621876126536_signature_31.png
| 2023-01-19 | 2023-01-19 19:23 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23018 | PBSP | vincentmilestone20@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF, COUNT: 151 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2811 ASSAY: Xpert SARS-CoV-2, COUNT: 10422 ASSAY: Xpert HIV-1 Viral Load, COUNT: 230 | 2023-01-19 | GX-IV | 805063 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Unit is fully operational | NA | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499410561873904712/5499410561873904712_signature_28.png | Jenry ken vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5499410561873904712/5499410561873904712_signature_31.png
| 2023-01-19 | 2023-01-19 19:19 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23017 | PBSP | vincentmilestone20@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF, COUNT: 151 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2811 ASSAY: Xpert SARS-CoV-2, COUNT: 10422 ASSAY: Xpert HIV-1 Viral Load, COUNT: 230 | 2023-01-19 | GX-IV 4D | 834626 | Module B3 for Module Replacement For Preventive Maintenance | Module Replacement and Perform Preventive Maintenance | Replace the service module S/N 674488 with new Module S/N: 656018 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate Pre and Post IQR, Assay Stat and System Log Archived Data test results | Unit is fullt operational | N/A | 900-0491rs | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499409681873456293/5499409681873456293_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5499409681873456293/5499409681873456293_signature_31.png
| 2023-01-19 | 2023-01-19 19:19 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23016 | PBSP | vincentmilestone20@gmail.com | 25.9 | ASSAY: Xpert MTB/RIF, COUNT: 151 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2811 ASSAY: Xpert SARS-CoV-2, COUNT: 10422 ASSAY: Xpert HIV-1 Viral Load, COUNT: 230 | 2023-01-19 | GX-IV 4D | 834626 | Module B3 for Module Replacement For Preventive Maintenance | Module Replacement and Perform Preventive Maintenance | Replace the service module S/N 674488 with new Module S/N: 656018 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate Pre and Post IQR, Assay Stat and System Log Archived Data test results | Unit is fullt operational | N/A | 900-0491rs | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499409611875190022/5499409611875190022_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5499409611875190022/5499409611875190022_signature_31.png
| 2023-01-19 | 2023-01-19 20:23 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23015 | PBSP | vincentmilestone20@gmail.com | 26.5 | ASSAY: , COUNT: N/A | 2023-01-19 | GX-XVI | 838756 | Returning Service Module SN 674488 | Return the Service module back to GX-XVI | Return the Service module back to GX-XVI | Unit not functional, powersupply unit is damaged | Waiting for Power Supply Unit parts | N/A | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499408671878069125/5499408671878069125_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5499408671878069125/5499408671878069125_signature_31.png
| 2023-01-19 | 2023-01-19 20:23 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 23014 | PBSP | vincentmilestone20@gmail.com | 26.5 | ASSAY: , COUNT: N/A | 2023-01-19 | GX-XVI | 838756 | Returning Service Module SN 674488 | Return the Service module back to GX-XVI | Return the Service module back to GX-XVI | Unit not functional, powersupply unit is damaged | Waiting for Power Supply Unit parts | N/A | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499408551874858917/5499408551874858917_signature_28.png | Jenry Ken Vincent Mibato | https://www.jotform.com/uploads/macarewecare/222927978372067/5499408551874858917/5499408551874858917_signature_31.png
| 2023-01-19 | 2023-01-19 14:37 | Esperanza health center | Teresa st sta mesa manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23013 | DOH | jong.agdamag@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF, COUNT: 382 ASSAY: Xpert MTB/RIF Ultra, COUNT: 771 |  | Gx4 | 831891 | For maintenance and stuck cartridge | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,fixed stuck cartridge | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499391088671505611/5499391088671505611_signature_28.png | Victor Agdamag jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5499391088671505611/5499391088671505611_signature_31.png
| 2023-01-19 | 2023-01-19 14:37 | Esperanza health center | Teresa st sta mesa manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23012 | DOH | jong.agdamag@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF, COUNT: 382 ASSAY: Xpert MTB/RIF Ultra, COUNT: 771 |  | Gx4 | 831891 | For maintenance and stuck cartridge | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,fixed stuck cartridge | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499390838672584011/5499390838672584011_signature_28.png | Victor Agdamag jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5499390838672584011/5499390838672584011_signature_31.png
| 2023-01-19 | 2023-01-19 14:37 | Esperanza health center | Teresa st sta mesa manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23011 | DOH | jong.agdamag@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF, COUNT: 382 ASSAY: Xpert MTB/RIF Ultra, COUNT: 771 |  | Gx4 | 831891 | For maintenance and stuck cartridge | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,fixed stuck cartridge | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499390888673845400/5499390888673845400_signature_28.png | Victor Agdamag jr | https://www.jotform.com/uploads/macarewecare/222927978372067/5499390888673845400/5499390888673845400_signature_31.png
| 2023-01-19 | 2023-01-19 16:23 | Proj.7 social hygiene | Bansalangin st. Qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | Noel Floresca |  |  |  |  | 23010 | MACARE-HBDC | tonettechacon@gmail.com.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 59 ASSAY: Xpert MTB/RIF Ultra, COUNT: 68 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1144 |  | Gx4 10c | 845446 | For maintenance | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Maintenance | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499389838679293028/5499389838679293028_signature_28.png | Anthony Reyes | https://www.jotform.com/uploads/macarewecare/222927978372067/5499389838679293028/5499389838679293028_signature_31.png
| 2023-01-19 | 2023-01-19 15:37 | Manuel J. Santos Hospital | Montilla Blvd. Butuan City | 13 | Agusan del Norte | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Faizar Avenido |  |  |  |  | 23009 | COMMERCIAL | molecularlab@mjsantos.com.ph | 24.375 | ASSAY: Xpert SARS-CoV-2, COUNT: 4090 |  | GX-4 | 831292 | Screen Monitor Glitching Jumpy Mouse Pointer Movement | Hard drive checking Troubleshooting | Troubleshooting Preventive Maintenance Data Collection | Defective monitor All modules in genexpert are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499243339517181599/5499243339517181599_signature_28.png | Jubily Dyanne A. Yap | https://www.jotform.com/uploads/macarewecare/222927978372067/5499243339517181599/5499243339517181599_signature_31.png
| 2023-01-19 | 2023-01-18 13:38 | Amai Pakpak Medical Center | Marawi City | BARMM | Lanao del Sur | Laboratory | NO | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 23008 | DOH | apmcmollab@gmail.com | 27 | ASSAY: Xpert SARS-CoV-2, COUNT: 8918 |  | GeneXpert | 834444 | H/W Fail Module D1/804652, D3/765468, D4/729707 | Module replacement | Module Replacement New Modules: D1/671961 D3/645283 D4/693683 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499235239681549811/5499235239681549811_signature_28.png | Rayan Cosari | https://www.jotform.com/uploads/macarewecare/222927978372067/5499235239681549811/5499235239681549811_signature_31.png
| 2023-01-19 | 2023-01-19 14:38 | Love YourSelf - Whitehouse | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | Jovy Noel |  |  |  |  | 23007 | COMMERCIAL | rmission@loveyourself.ph | 26 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 259 |  | Gx4 | 814150 | Preventive Maintenance | Clean all Modules and change filter Save File IQR, System Log Assay Stat,GXX file | Pm | All modules are functional |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499198244773351336/5499198244773351336_signature_28.png | Ralph Pearvin G. Mission, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5499198244773351336/5499198244773351336_signature_31.png
| 2023-01-19 | 2023-01-19 07:23 | University of Santo tomas hospital | España ,manila | NCR | NCR 1st - Manila | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | Noel Floresca |  |  |  |  | 23006 | COMMERCIAL | usth_bloodbank@yahoo.com | 27 |  |  | Mcs + | 02F069 | For Pm and Calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | PM and Calibration | Operational |  |  |  | Noel flores a | https://www.jotform.com/uploads/macarewecare/222927978372067/5499195013019106289/5499195013019106289_signature_28.png | Mylene Balb0a | https://www.jotform.com/uploads/macarewecare/222927978372067/5499195013019106289/5499195013019106289_signature_31.png
| 2023-01-19 | 2023-01-19 14:24 | Malibay Health Center | Malibay | NCR | NCR 4th - Pasay | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 23005 | PBSP | ag_abad@yahoo.com | 27.3 | ASSAY: Xpert MTB/RIF, COUNT: 153 ASSAY: Xpert MTB/RIF Ultra, COUNT: 593 ASSAY: Xpert HIV-1 Viral Load, COUNT: 138 |  | GX4D | 838886 | Unit is due for Xpertcheck | Calibration | XpertCheck Performed Maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499189441026780748/5499189441026780748_signature_28.png | Aljandro Abad | https://www.jotform.com/uploads/macarewecare/222927978372067/5499189441026780748/5499189441026780748_signature_31.png
| 2023-01-19 | 2023-01-19 12:19 | Mlang Rural Health Unit | Mlang, North Cotabato | 12 | Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 23004 | PBSP | kathlynlacabe7@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 429 |  | Gx-4 | 815434 | Module B3 H/W failed multiple temperature error | Perform Genexpert machine Preventive maintenance Perform Genexpert calibration/Xpertcheck | Perform Genexpert machine Preventive maintenance Perform Genexpert calibration/Xpertcheck | Gx machine has 3 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499189108778203224/5499189108778203224_signature_28.png | Kathlyn Mae Y. Lacabe | https://www.jotform.com/uploads/macarewecare/222927978372067/5499189108778203224/5499189108778203224_signature_31.png
| 2023-01-19 | 2023-01-19 14:06 | Cebu TB reference lab | Cebu City | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23003 | PBSP | asmantiquilla.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1766 |  | GX IV | 808620 | No power in display monitor | Taking up video of monitor and picture of serial number of monitor. | Taking up video of monitor and picture of serial number of monitor. Sending data to jotform by the End User | No power of display monitor |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499178946343678623/5499178946343678623_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5499178946343678623/5499178946343678623_signature_31.png
| 2023-01-19 | 2023-01-19 13:56 | Cebu TB Reference Lab | Cebu City | 7 | Cebu | Laboratory | NO | ON-SITE ASSISTANCE | CEPHEID | Raymond Laurito |  |  |  |  | 23002 | PBSP | asmantiquilla.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1766 |  | Gx iv | 808620 | No power in display monitor | Taking up video of monitor and Serial number picture send to jotform | Troubleshooting of monitor | No power in display monitor |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499174746346550320/5499174746346550320_signature_28.png | ALLAN S. MANTIQUILLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5499174746346550320/5499174746346550320_signature_31.png
| 2023-01-19 | 2023-01-19 12:08 | Central Luzon Doctor's Hospital | Hospital Drive Tarlac City | 3 | Tarlac | Molecular | NO | REPAIR Check-up of CPU not booting up | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23001 | COMMERCIAL | rohmhardameg@gmail.com | 17.1 | ASSAY: Xpert MTB/RIF, COUNT: 28 ASSAY: Xpert SARS-CoV-2, COUNT: 464 |  | GXII-D | 830841 | CPU not booting up | Troubleshoot probolem (did not boot up) Checked connection then try to boot up( booted-up) Performed functionality test | Troubkeshoot and repair | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499111742313377448/5499111742313377448_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/5499111742313377448/5499111742313377448_signature_31.png
| 2023-01-19 | 2023-01-19 10:23 | Tarlac Provincial Hospital | San Vicente | 3 | Tarlac | MolecularbTB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 22999 | PBSP | molecularlab.tph@gmail.com | 22.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 437 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 |  | GXIV-D | 805108 | Due for Calibration C2-Temperature error 1002/HW Failed C3-Error 2097,2032 | Performed preventive maintenance Installation of Xpertcheck Performed Calibration Gathered reports Performed functionality test | Preventive maintenance and calibration | Unit is operational except module C2 and C3 | Module C2 and C3 for reporting, modules passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499105272317752997/5499105272317752997_signature_28.png | Camille Ann Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5499105272317752997/5499105272317752997_signature_31.png
| 2023-01-19 | 2023-01-19 11:10 | Tarlac Provincial Hospital | San Vicente Tarlac City | 3 | Tarlac | MOLECULAR-PCR | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 23000 | PBSP | molecularlab.tph@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 1505 |  | GXIV-D | 838142 | Due for calibration | Performed prevenyive maintenance Installation of xpertcheck application Performed calibration Gathered reports Perform functionality test | Preventive maintenance and calibration | Unit is operational | All modules passed callibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499105282318987273/5499105282318987273_signature_28.png | Camille Ann Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5499105282318987273/5499105282318987273_signature_31.png
| 2023-01-19 | 2023-01-19 10:23 | Tarlac Provincial Hospital | San Vicente | 3 | Tarlac | MolecularbTB DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 22998 | PBSP | molecularlab.tph@gmail.com | 22.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 437 ASSAY: Xpert HIV-1 Viral Load, COUNT: 78 |  | GXIV-D | 805108 | Due for Calibration C2-Temperature error 1002/HW Failed C3-Error 2097,2032 | Performed preventive maintenance Installation of Xpertcheck Performed Calibration Gathered reports Performed functionality test | Preventive maintenance and calibration | Unit is operational except module C2 and C3 | Module C2 and C3 for reporting, modules passed calibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499087944688292072/5499087944688292072_signature_28.png | Camille Ann Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5499087944688292072/5499087944688292072_signature_31.png
| 2023-01-19 | 2023-01-19 11:10 | Tarlac Provincial Hospital | San Vicente Tarlac City | 3 | Tarlac | MOLECULAR-PCR | NO | CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 22997 | PBSP | molecularlab.tph@gmail.com | 24 | ASSAY: Xpert SARS-CoV-2, COUNT: 1505 |  | GXIV-D | 838142 | Due for calibration | Performed prevenyive maintenance Installation of xpertcheck application Performed calibration Gathered reports Perform functionality test | Preventive maintenance and calibration | Unit is operational | All modules passed callibration |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499087074682546208/5499087074682546208_signature_28.png | Camille Ann Cariaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5499087074682546208/5499087074682546208_signature_31.png
| 2023-01-19 | 2023-01-19 11:12 | Nabas Rural Health Unit | Nabas, Aklan | 6 | Aklan | TB Dots | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 22996 | PBSP | lambito.godelyn61@gmail.com | 24 | ASSAY: Others, COUNT: 0 |  | GXIV-4D | 815387 | #NAME? | None | #NAME? | All modules are functional. GeneXpert has no CPU. | CPU has been pulled out by RHU for repair to Kalibo. End user has not reported the incident to Macare. Waiting for repair of CPU. Wasn’t able to collect data for lack of CPU. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499074668308668921/5499074668308668921_signature_28.png | Godelyn Lambito | https://www.jotform.com/uploads/macarewecare/222927978372067/5499074668308668921/5499074668308668921_signature_31.png
| 2023-01-19 | 2023-01-19 10:50 | Dr. Elvira Lagrosa Health Center | 164 Dolores Hidalgo St. Brgy. 66 | NCR | NCR 4th - Pasay | Main Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Eron Dumagieng |  |  |  |  | 22995 | PBSP | rhemejetbotaquia@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF, COUNT: 94 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1138 |  | GX4D | 812833 | Unit due for Calibration/ Xpert Check | Calibration | XpertCheck of the following Module: B1/ 685785 Failed B2/ 639114 Passed B3/768349 Passed | Unit is operational | B1/685785 service required, module is for replacement. |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499065495027610764/5499065495027610764_signature_28.png | Rheme Botaquia | https://www.jotform.com/uploads/macarewecare/222927978372067/5499065495027610764/5499065495027610764_signature_31.png
| 2023-01-19 | 2023-01-19 10:45 | MANDAUE CITY HEALTH OFFICE | MANDAUE CITY, CEBU | 7 | Cebu | LABORATORY | NO | CALIBRATION | CEPHEID | Joseph Acebedo |  |  |  |  | 22994 | DOH | elma.bragat@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1406 ASSAY: Xpert MTB/RIF Ultra, COUNT: 864 ASSAY: Xpert HIV-1 Viral Load, COUNT: 98 |  | GX-IV | 838508 | #NAME? | -PERFORM MODULE B3 SN 862394 CALIBRATION : RETRY THEN FAILED | #NAME? | -MODULE B1, B4 - USED; MODULE B2,B3 - DISABLED |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499057428814016044/5499057428814016044_signature_28.png | ELMA B. GONZALES | https://www.jotform.com/uploads/macarewecare/222927978372067/5499057428814016044/5499057428814016044_signature_31.png
| 2023-01-19 | 2023-01-19 10:34 | Pikit Rural Health Unit | Pikit, north Cotabato | 12 | Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 22992 | PBSP | litosepelagior@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 223 | 2023-01-17 | Gx -2 | 835181 | N/A | Perform Genexpert machine Preventive maintenance Perform Genexpert calibration/Xpertcheck | Perform Genexpert machine Preventive maintenance Perform Genexpert calibration/Xpertcheck | Gx machine has 2 operational modules |  | 1000378669 | XpertCheck -CE-5 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499054513669674836/5499054513669674836_signature_28.png | Lito Sepelagio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5499054513669674836/5499054513669674836_signature_31.png
| 2023-01-19 | 2023-01-19 09:33 | Kidapawan City Health Office | Kidapawan City, North Cotabato | 12 | Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | Jesson Tupas |  |  |  |  | 22991 | DOH | sg93rmt@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF, COUNT: 790 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1497 |  | Gx-4 | 838545 | Module B3 H/W failed due to multiple temperature error Module B4 H/W failed due to multiple temperature error | Perform Genexpert machine Preventive maintenance Perform Genexpert calibration/Xpertcheck to module B3 and B4 for troubleshooting | Perform Genexpert machine Preventive maintenance Perform Genexpert calibration/Xpertcheck to module B3 and B4 for troubleshooting | Gx machine has 2 operational modules | Module B3 passed and module B4 for Service Required |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499034732117061778/5499034732117061778_signature_28.png | Tessa Mae E. Saguindan | https://www.jotform.com/uploads/macarewecare/222927978372067/5499034732117061778/5499034732117061778_signature_31.png
| 2023-01-19 | 2023-01-19 09:56 | Kumalarang Rural Health Unit | Poblacion kumalarang Zamboanga Del Sur | 9 | Zamboanga del Sur | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | Gian Carlo Aizon |  |  |  |  | 22990 | PBSP | roseimelda1228@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1005 | 2023-01-19 | GX 4 | 815278 | Module 2 temperature error | Replaced module 2 (695642) with Module (655869) | Preventive Maintenance , Module Replacement | All 4 modules are available | None | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5499030835509169757/5499030835509169757_signature_28.png | Rose imelda P. Uy | https://www.jotform.com/uploads/macarewecare/222927978372067/5499030835509169757/5499030835509169757_signature_31.png
| 2023-01-19 | 2023-01-19 08:05 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 22989 | DOH | burauenlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 128 |  | GXIV | 838551 | For PM | #NAME? | Preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498983623119317729/5498983623119317729_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5498983623119317729/5498983623119317729_signature_31.png
| 2023-01-19 | 2023-01-19 08:05 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 22988 | DOH | burauenlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 140 ASSAY: Xpert MTB/RIF Ultra, COUNT: 128 |  | GXIV | 838551 | For PM | #NAME? | Preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498983533115360271/5498983533115360271_signature_28.png | Mary Ann Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5498983533115360271/5498983533115360271_signature_31.png
| 2023-01-18 | 2023-01-18 21:52 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 22987 | DOH | ammegzedd@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 604 ASSAY: Xpert MTB/RIF Ultra, COUNT: 353 | 2023-01-16 | GXIV | 838421 | For PM | #NAME? | Preventive maintenance | Module B2 & B3 are Disabled (Waiting for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498594626712669543/5498594626712669543_signature_28.png | Gemma Malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5498594626712669543/5498594626712669543_signature_31.png
| 2023-01-18 | 2023-01-18 21:52 | Tacloban City Health Office | Tacloban City | 8 | Leyte | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 22986 | DOH | ammegzedd@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 604 ASSAY: Xpert MTB/RIF Ultra, COUNT: 353 | 2023-01-16 | GXIV | 838421 | For PM | #NAME? | Preventive maintenance | Module B2 & B3 are Disabled (Waiting for module replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498594296715527269/5498594296715527269_signature_28.png | Gemma Malate | https://www.jotform.com/uploads/macarewecare/222927978372067/5498594296715527269/5498594296715527269_signature_31.png
| 2023-01-18 | 2023-01-18 16:28 | DOH R02 | Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 22984 | DOH | cvchdcovidlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 997 ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 ASSAY: Xpert SARS-CoV-2, COUNT: 5968 | 2023-01-18 | Genexpert XVI | 842894 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module C2, D1, D3, D4 for reporting | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498406135424665885/5498406135424665885_signature_28.png | Queenie Cantil | https://www.jotform.com/uploads/macarewecare/222927978372067/5498406135424665885/5498406135424665885_signature_31.png
| 2023-01-18 | 2023-01-18 16:28 | DOH R02 | Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 22985 | DOH | cvchdcovidlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 997 ASSAY: Xpert MTB/RIF Ultra, COUNT: 874 ASSAY: Xpert SARS-CoV-2, COUNT: 5968 | 2023-01-18 | Genexpert XVI | 842894 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module C2, D1, D3, D4 for reporting | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498406155424981501/5498406155424981501_signature_28.png | Queenie Cantil | https://www.jotform.com/uploads/macarewecare/222927978372067/5498406155424981501/5498406155424981501_signature_31.png
| 2023-01-18 | 2023-01-18 16:33 | DOH R02 | Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 22983 | PBSP | cvchdcovidlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 394 ASSAY: Xpert MTB/RIF Ultra, COUNT: 331 ASSAY: Xpert SARS-CoV-2, COUNT: 2306 | 2023-01-18 | Genexpert IV | 834434 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module B4 for reporting | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498405575423717265/5498405575423717265_signature_28.png | Queenie Cantil | https://www.jotform.com/uploads/macarewecare/222927978372067/5498405575423717265/5498405575423717265_signature_31.png
| 2023-01-18 | 2023-01-18 16:33 | DOH R02 | Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 22982 | PBSP | cvchdcovidlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 394 ASSAY: Xpert MTB/RIF Ultra, COUNT: 331 ASSAY: Xpert SARS-CoV-2, COUNT: 2306 | 2023-01-18 | Genexpert IV | 834434 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Module B4 for reporting | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498405515427521333/5498405515427521333_signature_28.png | Queenie Cantil | https://www.jotform.com/uploads/macarewecare/222927978372067/5498405515427521333/5498405515427521333_signature_31.png
| 2023-01-18 | 2023-01-18 16:36 | DOH R02 | Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | MAINTENANCE | CEPHEID | Matthew Roxas |  |  |  |  | 22981 | PBSP | cvchdcovidlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 394 ASSAY: Xpert MTB/RIF Ultra, COUNT: 331 ASSAY: Xpert SARS-CoV-2, COUNT: 2306 | 2023-01-18 | Genexpert IV | 808628 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498404935421809276/5498404935421809276_signature_28.png | Queenie Cantil | https://www.jotform.com/uploads/macarewecare/222927978372067/5498404935421809276/5498404935421809276_signature_31.png
| 2023-01-18 | 2023-01-18 15:34 | Dr. Rafael S. Tumbokon Memorial Hospital | Kalibo, Aklan | 6 | Aklan | Laboratory | YES | CALIBRATION TROUBLESHOOTING | CEPHEID | Wynnona Reoma |  |  |  |  | 22980 | PBSP | gonzagaevarmt@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1267 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2965 ASSAY: Xpert HIV-1 Viral Load, COUNT: 40 | 2023-01-18 | GXIV-4D | 804680 | Module D1 - Ultrasonic Horn Error; Module D2 - for rerunning of Xpertcheck | #NAME? | #NAME? | D2 - disabled; D1 - under observation | Gateway board for replacement. Module D1 under observation if for replacement or not. Two GeneXpert machines connected in one CPU. | Xpert Check Lot No. 05502 Exp. 2024 02 11 | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498378309677500811/5498378309677500811_signature_28.png | Eva Raquel Gonzaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5498378309677500811/5498378309677500811_signature_31.png
| 2023-01-18 | 2023-01-18 15:17 | Libertad Main Health Center | Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | CALIBRATION MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 22979 | DOH | suzzettekho34@gmail.com | 28.75 | ASSAY: Xpert MTB/RIF, COUNT: 338 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1012 |  | GX-4 | 838554 | Module B2 encounters error 1002 | Xpertcheck Calibration | Calibration of defective module B2 Monthly Preventive Maintenance Data Collection | 3 modules working and 1 module service required for replacement |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498359284918878208/5498359284918878208_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/5498359284918878208/5498359284918878208_signature_31.png
| 2023-01-18 | 2023-01-18 15:17 | CEBU DOCTORS’ UNIVERSITY HOSPITAL | Cebu City | 7 | Cebu | MOLECULAR LAB | NO | TRANSFER | CEPHEID | Joseph Acebedo |  |  |  |  | 22978 | COMMERCIAL | bjohnduke@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 11 |  | GX-IV | 110008524 | TRANSFER | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498357823404466645/5498357823404466645_signature_28.png | John Duke S. Bacarisas | https://www.jotform.com/uploads/macarewecare/222927978372067/5498357823404466645/5498357823404466645_signature_31.png
| 2023-01-18 | 2023-01-18 14:59 | Reproductive Health and Wellness Center - Jacinto Health Center | Jacinto St. Poblacion Davao City | 11 | Davao del Sur | HIV Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 22977 | DOH | rhwclab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 475 | 2023-01-18 | GX-IV 4D | 838262 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is Fully Operational | N/A | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498349623424427426/5498349623424427426_signature_28.png | Cherry Anne R. Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5498349623424427426/5498349623424427426_signature_31.png
| 2023-01-18 | 2023-01-18 14:59 | Reproductive Health and Wellness Center - Jacinto Health Center | Jacinto St. Poblacion Davao City | 11 | Davao del Sur | HIV Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 22976 | DOH | rhwclab@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 475 | 2023-01-18 | GX-IV 4D | 838262 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is Fully Operational | N/A | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498349583427491879/5498349583427491879_signature_28.png | Cherry Anne R. Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5498349583427491879/5498349583427491879_signature_31.png
| 2023-01-18 | 2023-01-18 14:52 | Reproductive Health and Wellness Center - Jacinto Health Center | Jacinto St. Poblacion, Davao City | 11 | Davao del Sur | HIV Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 22975 | MACARE-HBDC | rhwclab@yahoo.com | 24.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 23 ASSAY: Xpert HIV-1 Viral Load, COUNT: 433 | 2023-01-18 | GX-IV 4D | 839236 | For Preventive Maintenance Multiple Errors / Hardware Failure on Module B1 S/N: 869012 | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Multiple Temperature Error at Module B1 | Errors sent to Cepheid | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498348903425812238/5498348903425812238_signature_28.png | Cherry Anne R. Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5498348903425812238/5498348903425812238_signature_31.png
| 2023-01-18 | 2023-01-18 14:52 | Reproductive Health and Wellness Center - Jacinto Health Center | Jacinto St. Poblacion, Davao City | 11 | Davao del Sur | HIV Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 22974 | MACARE-HBDC | rhwclab@yahoo.com | 24.9 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 23 ASSAY: Xpert HIV-1 Viral Load, COUNT: 433 | 2023-01-18 | GX-IV 4D | 839236 | For Preventive Maintenance Multiple Errors / Hardware Failure on Module B1 S/N: 869012 | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Multiple Temperature Error at Module B1 | Errors sent to Cepheid | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498348873424431733/5498348873424431733_signature_28.png | Cherry Anne R. Asis | https://www.jotform.com/uploads/macarewecare/222927978372067/5498348873424431733/5498348873424431733_signature_31.png
| 2023-01-18 | 2023-01-18 14:57 | Santolan Super Health Center | Evangelista St. | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR CALIBRATION | CEPHEID | John Monsour Doctor |  |  |  |  | 22973 | PBSP | samandlet17@gmail.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2664 | 2023-01-16 | GX-IV-D | 838898 | Module B4 for replacement, for xpertcheck | Replaced defective module, performed XpertCheck | Performed maintenance Replaced defective module Run XpertCheck Generated reports | All modules operational |  | Old B4/863777 New B4/738573 | G 5042 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498347839215858339/5498347839215858339_signature_28.png | Lolita Formalejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5498347839215858339/5498347839215858339_signature_31.png
| 2023-01-18 | 2023-01-18 14:47 | Iligan Society of Internists | Iligan City | 10 | Lanao del Norte | Laboratory | YES | INSTALLATION | CEPHEID | Kim Ryan Arrojado |  |  |  |  | 22972 | PBSP | sdramayan@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2633 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 |  | GeneXpert | 707842 | Damaged GeneXpert power supply | GeneXpert Power supply replacement | Replacement of power supply Preventive Maintenance Network 2 GeneXpert equipment SN: 805090 SN: 707842 | Operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498342771027258271/5498342771027258271_signature_28.png | Sandra D. Ramayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5498342771027258271/5498342771027258271_signature_31.png
| 2023-01-18 | 2023-01-18 14:42 | Kabacan Rural Health Unit | Kabacan, North Cotabato | 12 | Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 22971 | PBSP | yayo_kmsi@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 355 ASSAY: Xpert MTB/RIF Ultra, COUNT: 613 |  | Gx-2 | 835182 | N/A | Genexpert machine Preventive maintenance | Genexpert machine Preventive maintenance | Gx machine has 2 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498336098916378010/5498336098916378010_signature_28.png | Joan A. Sangines | https://www.jotform.com/uploads/macarewecare/222927978372067/5498336098916378010/5498336098916378010_signature_31.png
| 2023-01-18 | 2023-01-18 13:19 | Paniqui Rural Health Unit | Paniqui Tarlac | 3 | Tarlac | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 22970 | PBSP | forylab76@yahoo.com | 29.4 | ASSAY: Xpert MTB/RIF, COUNT: 300 ASSAY: Xpert MTB/RIF Ultra, COUNT: 414 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2023-01-17 | GXIV-D | 812792 | Module replacement and calibration | Performed preventive maintenance Performed module replacement Performed calibration Gathered reports Performed functionality test | Repair and calibration | Unit is operational | All modules passed calibration | Genexpert module | 38195 | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5498323545321864283/5498323545321864283_signature_28.png | Fortune Abutin | https://www.jotform.com/uploads/macarewecare/222927978372067/5498323545321864283/5498323545321864283_signature_31.png
| 2023-01-18 | 2023-01-18 13:19 | Paniqui Rural Health Unit | Paniqui Tarlac | 3 | Tarlac | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 22969 | PBSP | forylab76@yahoo.com | 29.4 | ASSAY: Xpert MTB/RIF, COUNT: 300 ASSAY: Xpert MTB/RIF Ultra, COUNT: 414 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2023-01-17 | GXIV-D | 812792 | Module replacement and calibration | Performed preventive maintenance Performed module replacement Performed calibration Gathered reports Performed functionality test | Repair and calibration | Unit is operational | All modules passed calibration | Genexpert module | 38195 | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5498323525328096545/5498323525328096545_signature_28.png | Fortune Abutin | https://www.jotform.com/uploads/macarewecare/222927978372067/5498323525328096545/5498323525328096545_signature_31.png
| 2023-01-18 | 2023-01-18 13:46 | Lapulapu City Health Office | Lapulapu City Health office | 7 | Cebu | Tb laboratory | YES | REPAIR | CEPHEID | Mark Efren Ramas |  |  |  |  | 22968 | PBSP | tanya_pax@yahoo.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 1322 | 2023-01-18 | Gx4 | 808637 | Stucked cartridge due to Error 2037 | Remove cartridege | Cartridge removal | Only to 2 modules available |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498304769024848002/5498304769024848002_signature_28.png | Maria Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5498304769024848002/5498304769024848002_signature_31.png
| 2023-01-18 | 2023-01-18 11:37 | Dr. Amado Diaz Provincial Foundation Hospital | Midsayap, North Cotabato | 12 | Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 22967 | PBSP | charlenemae_amatac@yahoo.com | 29.5 | ASSAY: Xpert MTB/RIF, COUNT: 477 ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 |  | Gx-4 | 807623 | Module B2 under observation due to multiple temperature error last September 2022, | Genexpert machine Preventive maintenance | Genexpert machine Preventive maintenance | Gx machine has 4 operational modules |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498226724661159419/5498226724661159419_signature_28.png | Charlene Mae A. Amatac | https://www.jotform.com/uploads/macarewecare/222927978372067/5498226724661159419/5498226724661159419_signature_31.png
| 2023-01-18 | 2023-01-18 11:27 | Amang Rodriguez Memorial Medical Center | Sumlong Hwy Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR | CEPHEID | John Monsour Doctor |  |  |  |  | 22966 | PBSP | pathologymole.armmc@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 3681 ASSAY: Others, COUNT: 1 | 2023-01-16 | GX-XVI-D | 844113 | Module D4 For replacement | Replaced defective module | Performed maintenance Replaced defective module Generated reports | All modules operational |  | Old: D4/886142 New: D4/210036542 | G 5025 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498221397611642863/5498221397611642863_signature_28.png | Racquet Rhys C. Verzosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5498221397611642863/5498221397611642863_signature_31.png
| 2023-01-18 | 2023-01-18 11:16 | Amang Rodriguez Memorial Medical Center | Sumulong highway Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | John Monsour Doctor |  |  |  |  | 22965 | PBSP | maricel_c27@tahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 650 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1578 |  | GX-IV-D | 807982 | Module B3 error 1001/1002/4012 | Performed maintenance | Performed maintenance Checked module B3 Generated reports | Modules B1,B2 and B4 operational | Module B3/608694 for reporting |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498213442329242924/5498213442329242924_signature_28.png | Maricel S.Xakderin | https://www.jotform.com/uploads/macarewecare/222927978372067/5498213442329242924/5498213442329242924_signature_31.png
| 2023-01-18 | 2023-01-18 10:12 | Midsayap Rural Health Unit | Midsayap, North Cotabato | 12 | Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | Jesson Tupas |  |  |  |  | 22964 | PBSP | rhumidsayap@gmail.com | 26.3 | ASSAY: Xpert MTB/RIF, COUNT: 220 ASSAY: Xpert MTB/RIF Ultra, COUNT: 428 | 2023-01-18 | Gx- 2 | 835183 | Module B1 and B2 for Service Required | Genexpert machine Calibration/ExpertCheck Genexpert machine Preventive maintenance | Genexpert machine Calibration/ExpertCheck Genexpert machine Preventive maintenance | Gx machine has 2 operational modules |  | 1000378669 | ExpertCheck -CE-5 |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498189584668867116/5498189584668867116_signature_28.png | Rhea A Drilon | https://www.jotform.com/uploads/macarewecare/222927978372067/5498189584668867116/5498189584668867116_signature_31.png
| 2023-01-18 | 2023-01-18 10:12 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 22963 | DOH | sarah_yangzon@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 325 ASSAY: Xpert HIV-1 Viral Load, COUNT: 278 |  | GXIV | 838406 | For PM | #NAME? | Preventive maintenance | Module B3 Disabled (waiting for replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498176831013383375/5498176831013383375_signature_28.png | Sarah V Yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5498176831013383375/5498176831013383375_signature_31.png
| 2023-01-18 | 2023-01-18 10:12 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | TB-DOTS | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 22962 | DOH | sarah_yangzon@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 325 ASSAY: Xpert HIV-1 Viral Load, COUNT: 278 |  | GXIV | 838406 | For PM | #NAME? | Preventive maintenance | Module B3 Disabled (waiting for replacement) |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5498176721016483376/5498176721016483376_signature_28.png | Sarah V Yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5498176721016483376/5498176721016483376_signature_31.png
| 2023-01-17 | 2023-01-17 14:00 | Buhangin District Health Center | NHA Buhangin, Davao City | 11 | Davao del Sur | TB-DOTS Laboratory | NO | MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 22961 | DOH | cborre19@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF, COUNT: 104 ASSAY: Xpert MTB/RIF Ultra, COUNT: 587 | 2023-01-17 | GX-IV 4D | 838232 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is fully operational | N/A | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497454013426608516/5497454013426608516_signature_28.png | Catherine B. Deocampo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5497454013426608516/5497454013426608516_signature_31.png
| 2023-01-17 | 2023-01-17 14:05 | Lapulapu City Health office | Lapulapu City | 7 | Cebu | Tb laboratory | YES | TROUBLESHOOTING | CEPHEID | Mark Efren Ramas |  |  |  |  | 22960 | PBSP | cuyoscarlgerard@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 1325 | 2023-01-17 | Gx4 | 808637 | 1.Valve error on Module 2 Error 2006 2. Hang Up problem on Module 4 during Mtb ultra running | Check valve position and restart machine | Get I QR and Sys log Generate Gxx, And Assay Statistics | Module 3 is excluded during HWF during running process | Only 3 modules are operational | NA | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497453881653819852/5497453881653819852_signature_28.png | Lucille Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5497453881653819852/5497453881653819852_signature_31.png
| 2023-01-17 | 2023-01-17 13:58 | Lapulapu City Health office | Poblacion Lapulapu city | 7 | Cebu | Tb laboratory | YES | MAINTENANCE | CEPHEID | Mark Efren Ramas |  |  |  |  | 22959 | DOH | cuyosgerard@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 1093 | 2023-01-17 | GX4 | 837661 | No errors | Check and monitor unit and UPS | Conduct check up and monitoring of Ups and genxpert machine | Operational |  | Na | Na |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497446891657919818/5497446891657919818_signature_28.png | Lucelle Arong | https://www.jotform.com/uploads/macarewecare/222927978372067/5497446891657919818/5497446891657919818_signature_31.png
| 2023-01-17 | 2023-01-17 12:45 | Calbiga Rural Health Unit | Calbiga Samar | 8 | Samar | TB-DOT | NO | MAINTENANCE | CEPHEID | John Paul Labutap |  |  |  |  | 22958 | DOH | lornsabaigar@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 183 ASSAY: Xpert MTB/RIF Ultra, COUNT: 234 |  | GXIV | 838409 | for PM | #NAME? | Preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497402570719347692/5497402570719347692_signature_28.png | Lorna C.Abaigar | https://www.jotform.com/uploads/macarewecare/222927978372067/5497402570719347692/5497402570719347692_signature_31.png
| 2023-01-17 | 2023-01-17 12:20 | RHU Magsaysay | Brgy. Poblacion, Magsaysay | 10 | Misamis Oriental | TB-DOTS | NO | MAINTENANCE | CEPHEID | Faizar Avenido |  |  |  |  | 22957 | DOH | jkakilgan12@gmail.com | 25.5 | ASSAY: Xpert MTB/RIF, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 94 |  | GX-4 | 838557 | None | Nine | Preventive Maintenance Data Collection Site Visit | All modules are working |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497400629606026284/5497400629606026284_signature_28.png | Jesusa N. Kakilgan | https://www.jotform.com/uploads/macarewecare/222927978372067/5497400629606026284/5497400629606026284_signature_31.png
| 2023-01-17 | 2023-01-17 12:13 | Zamboanga City Health Office | Pettit Baracks Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 22956 | PBSP | florestresitaj25@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 | 2023-01-17 | GX 4 | 838128 | Xpertcheck Calibration | Undergone Xpertcheck Calibration | Conducted Xpertcheck Calibration | All modules are available | Module 1,2 and 3 passed. Module 4 service required | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497389025326594255/5497389025326594255_signature_28.png | Teresita J. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5497389025326594255/5497389025326594255_signature_31.png
| 2023-01-17 | 2023-01-17 11:39 | Zamboanga City Health Office | Pettit Baracks Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 22955 | PBSP | teresitafloresj25@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 934 | 2023-01-17 | GX 4 | 818902 | System unit unable to boot up. | Changed CMOS battery, changed boot up priority. | Changed CMOS battery and Preventive Maintenance | All 4 modules are available | Module 3 under observation | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497387585321744066/5497387585321744066_signature_28.png | Teresita J. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5497387585321744066/5497387585321744066_signature_31.png
| 2023-01-17 | 2023-01-17 12:13 | Zamboanga City Health Office | Pettit Baracks Zamboanga City | 9 | Zamboanga del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | Gian Carlo Aizon |  |  |  |  | 22954 | PBSP | florestresitaj25@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 501 | 2023-01-17 | GX 4 | 838128 | Xpertcheck Calibration | Undergone Xpertcheck Calibration | Conducted Xpertcheck Calibration | All modules are available | Module 1,2 and 3 passed. Module 4 service required | N/A | N/A |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497385025328945091/5497385025328945091_signature_28.png | Teresita J. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5497385025328945091/5497385025328945091_signature_31.png
| 2023-01-17 | 2023-01-17 12:04 | Urdaneta Sacred Heart Hospital | Urdaneta Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | Mark Joseph Venzon |  |  |  |  | 22953 | COMMERCIAL | annabelle_llaneta@yahoo.com | 27.9 | ASSAY: Xpert MTB/RIF, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 145 |  | GXII-D | 820303 | Due for maintenance | Performed Surface cleaning Cleaning of plunger rod,filter,cartridge bay and icore Archived test results Gathered reports Performed functionality test | Preventive maintenance | Unit is operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497380840423000740/5497380840423000740_signature_28.png | Annabelle Catbagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5497380840423000740/5497380840423000740_signature_31.png
| 2023-01-17 | 2023-01-17 11:08 | Tomas Claudio Health Center - District A | Corner E. Quirino, Camus Street, Davao City | 11 | Davao del Sur | TB DOTS - Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | Roi Kenneth Fagutao |  |  |  |  | 22952 | DOH | arleen_p_alconera@yahoo.com | 28.9 | ASSAY: Xpert MTB/RIF, COUNT: 266 ASSAY: , COUNT: 211 | 2023-01-17 | GX-IV 4D | 838492 | Loose power chord of GeneXpert For Preventive Maintenance | Check the wiring, power chord and Perform Preventive Maintenance | Swapped the power chord of the desktop system unit with the GeneXpert Machine power cord Replug the power chords for checking for faulty/loose connection Clean Surface Area, Cartridge Bay, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is under observation | If problem still exist, will replaced the new powerchord | N/A | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497347513425990177/5497347513425990177_signature_28.png | Arleen Alconera | https://www.jotform.com/uploads/macarewecare/222927978372067/5497347513425990177/5497347513425990177_signature_31.png
| 2023-01-17 | 2023-01-17 09:19 | Chong Hua Hospital | Fuente Cebu City | 7 | Cebu | Bloodbank | YES | REPAIR | HAEMONETICS | Mark Efren Ramas |  |  |  |  | 22951 | COMMERCIAL | Chhc_bloodbank@chonghua.com.ph | 21 |  | 2023-01-17 | MCS+ | 11J219 | DUAL MANIFOLD PUMP CANNOT BE DETECTED Line sensor cannot be detected | Reset machine and swap protocol card of UPP | Troubleshoot problem of UPP protocol Reset machine Clean up dual pump bar code reader | Operational and can be used |  | NA | NA |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5497279021656252100/5497279021656252100_signature_28.png | Mayleen F. Jakosalem | https://www.jotform.com/uploads/macarewecare/222927978372067/5497279021656252100/5497279021656252100_signature_31.png
| 2023-01-16 | 2023-01-16 14:42 | Cebu TB reference Lab | Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | Raymond Laurito |  |  |  |  | 22950 | DOH | tivalerio.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1783 |  | GX IV | 838263 | Re-installation | Re-installation | Re-installation | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5496610503521882795/5496610503521882795_signature_28.png | Titus Ian Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5496610503521882795/5496610503521882795_signature_31.png
| 2023-01-16 | 2023-01-16 14:39 | Cebu TB reference lab | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION | CEPHEID | Raymond Laurito |  |  |  |  | 22949 | PBSP | tivalerio.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1772 |  | Gx IV | 838882 | Re-installation | Re-installation | Re-installation | Systems operational |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5496605503522904735/5496605503522904735_signature_28.png | Titus Ian Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5496605503522904735/5496605503522904735_signature_31.png
| 2023-01-16 | 2023-01-16 14:31 | Cebu TB Reference Lab | Cebu City | 7 | Cebu | Laboratory | NO | INSTALLATION | CEPHEID | Raymond Laurito |  |  |  |  | 22948 | PBSP | tivalerio.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1772 |  | GX IV | 806201 | Re-installation | Re-installation | Re-installation | Module E2 hardware fail |  |  |  |  | https://www.jotform.com/uploads/macarewecare/222927978372067/5496603009915726298/5496603009915726298_signature_28.png | Titus Ian Valerio | https://www.jotform.com/uploads/macarewecare/222927978372067/5496603009915726298/5496603009915726298_signature_31.png
| 2023-01-16 | 2023-01-16 13:25 | Koronadal City Health Office | Koronadal City | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22947 | PBSP | calvokenneth12@gmail.com | 23.1 | ASSAY: Xpert MTB/RIF, COUNT: 851 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1052 | 2023-01-16 | Gx-4 | 707529 | Module A2 H/W failed, waiting for module to available | Genexpert machine Preventive maintenance Replace module A3 SN: 717592 with new Module SN: 210019125 | Genexpert machine Preventive maintenance Replace module A3 SN: 717592 with new Module SN: 210019125 | Gx machine has 3 operational modules |  | 900-0491RS | Genexpert module 6-color | Jesson Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5496563668318651240/5496563668318651240_signature_28.png | Kenneth Paolo T. Calvo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5496563668318651240/5496563668318651240_signature_31.png
| 2023-01-16 | 2023-01-16 11:37 | Santa Cruz Rural Health Unit | Poblacion, Sta. Cruz, Davao del Sur | 11 | Davao del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22946 | DOH | rhustacruzdavsur@gmail.com | 26.9 | ASSAY: Xpert MTB/RIF, COUNT: 172 ASSAY: Xpert MTB/RIF Ultra, COUNT: 499 | 2023-01-16 | GX-IV 4D | 838505 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters and Icore Lenses Generate IQR, Assay Stat and System Log Archived Data test results | Unit is fully operational | N/A | N/A | N/A | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5496557044915668253/5496557044915668253_signature_28.png | DOROTHY A. DALAGAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5496557044915668253/5496557044915668253_signature_31.png
| 2023-01-16 | 2023-01-13 14:15 | Asian Hospital and Medical Center | Filinvest | NCR | NCR 4th - Muntinlupa | Molecular Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22942 | MACARE-HBDC | labteamlead@asianhospital.com | 24 | ASSAY: Others, COUNT: 33 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 ASSAY: Xpert MTB/RIF, COUNT: 106 ASSAY: Xpert MTB/RIF Ultra, COUNT: 947 |  | GXIV-L | 803869 | Cannot Detect GX Modules | Troubleshoot | Checked Connections and GX Setup Used Different Cables Performed Maintenance Generate Reports | Unit is operational. UPS is Defective | UPS S/N: 1435400116PS27M |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5496373951213326629/5496373951213326629_signature_28.png | Christine Fallaria | https://www.jotform.com/uploads/macarewecare/222927978372067/5496373951213326629/5496373951213326629_signature_31.png
| 2023-01-16 | 2023-01-12 11:13 | Molino III CHU | Molino 3, Bacoor | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22944 | DOH | joanrnd@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 460 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 |  | GXIV-D | 821221 | Monitor has lcd problem | Replacement | Replaced Monitor Performes Maintenance Generate reports | Unit is operational | CNK7470PJ --> QSDN8HA087659 |  |  | Brenguiler I Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5496373951214181690/5496373951214181690_signature_28.png | Jonalyn Aranda | https://www.jotform.com/uploads/macarewecare/222927978372067/5496373951214181690/5496373951214181690_signature_31.png
| 2023-01-16 | 2023-01-13 10:11 | Jose Reyes Memorial Medical Center | Rizal ave, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22943 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 871 ASSAY: Xpert SARS-CoV-2, COUNT: 2988 |  | GX16-D | 110013856 | Temperature error on modules | Replaced Module | Performed Module replacement Generate reports | 5 modules for reporting and 11 modules are operational | 210069291--> 210077080 |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5496373951216786906/5496373951216786906_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5496373951216786906/5496373951216786906_signature_31.png
| 2023-01-16 | 2023-01-13 11:27 | Kalina Health Hub | Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22939 | CONSORTIUM | jasmin.catral@innovations.org | 25 | ASSAY: Xpert MTB/RIF, COUNT: 332 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1088 |  | GX-IV-D | 110008418 | For Maintenance | Performed maintenance | Performed maintenance cleaned cartridge bay, plunger rod, icore and filter Generated reports | All modules operational |  |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5496372511213552337/5496372511213552337_signature_28.png | Jasmin Catral RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5496372511213552337/5496372511213552337_signature_31.png
| 2023-01-16 | 2023-01-12 15:02 | Rosario Super Health Center | Rosario | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22940 | WHO | eden.cawayan@yahoo.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 927 ASSAY: Xpert SARS-CoV-2, COUNT: 526 |  | GX-IV-D | 843924 | Stuck cartridge on module 1 | Removed stuck cartridge | Performed maintenance Removed stuck cartridge Generated reports | Modules B1,B2, and B3 operational. Module B4 H/W failed reported already. |  |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5496372511215265195/5496372511215265195_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5496372511215265195/5496372511215265195_signature_31.png
| 2023-01-16 | 2023-01-13 13:43 | Pasig Treatment Hub | Champ Bldg. | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22941 | MACARE-HBDC | jhozza.josermt@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 50 ASSAY: Xpert HIV-1 Viral Load, COUNT: 815 |  | GX-IV-L | 821648 | Module A1/73328 | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod. Generated reports. | Module A1 not operational. Modules A2, A3 and A4 operational. | Module A1 disabled. |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5496372511216339022/5496372511216339022_signature_28.png | Jhozza Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5496372511216339022/5496372511216339022_signature_31.png
| 2023-01-16 | 2023-01-13 11:24 | Kalinga Health Hub | Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22938 | PBSP | jasmin.catral@innovations.org | 25 | ASSAY: Others, COUNT: N/A |  | GX-IV-L | 805211 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter. Generated reports | All modules operational. |  |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5496372511218469282/5496372511218469282_signature_28.png | Jasmin Catral RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5496372511218469282/5496372511218469282_signature_31.png
| 2023-01-13 | 2023-01-13 16:49 | Saint Lukes Extension Center | J. Bocobo Street, Ermita | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22937 | MACARE-HBDC | laboratory@slec.ph | 21.4 | ASSAY: Xpert MTB/RIF, COUNT: 1652 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2952 ASSAY: Xpert CT/NG, COUNT: 4073 | 2023-01-11 | Genexpert Infinity | 809643 | Unit is due for Module Replacement Unit is due for Preventive Maintenance | Module Replacement | Preventive Maintenance Module Replacement From F1/ 880991 into 702116 From F3/ 642037 into F3 / 699635 From H1/723917 / inti H1/ 608345 | Unit is Operational. A4 and H4 newly reported |  | GeneXpert Module 6 SN 702116, 608345, 699635 | G 5022 | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5494099072592304078/5494099072592304078_signature_28.png | Maria Teresita Baclig | https://www.jotform.com/uploads/macarewecare/222927978372067/5494099072592304078/5494099072592304078_signature_31.png
| 2023-01-13 | 2023-01-13 13:44 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | [5 empty] | 22936 | PBSP | laureanjo@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 975 ASSAY: Xpert SARS-CoV-2, COUNT: 36 |  | GX IV | 808620 | Transfer from CTRL to Cebu CHO | Installation | Installation | All modules functional |  |  |  | Raymond G. Laurito | https://www.jotform.com/uploads/macarewecare/222927978372067/5493986437129731092/5493986437129731092_signature_28.png | Laurean Jo F. Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5493986437129731092/5493986437129731092_signature_31.png
| 2023-01-13 | 2023-01-13 13:09 | Sta Rosa Community District Hospital | Sta Rosa Poblacion Olango Island | 7 | Cebu | TB Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22935 | DOH | ariestumz@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 97 | 2023-01-13 | GX4 | 838534 | No errors | Conduct Preventive Maintenance Clean Icore, Clean Filter Arcihive Test Results Back up database Generate IQR, Sys Log, and Assay Stat | Conduct Preventive mainrenance | All modules are operational | NA | Na | Na | Mark Efren Ramas | https://www.jotform.com/uploads/macarewecare/222927978372067/5493981645486679251/5493981645486679251_signature_28.png | Ariston Tumulak | https://www.jotform.com/uploads/macarewecare/222927978372067/5493981645486679251/5493981645486679251_signature_31.png
| 2023-01-13 | 2023-01-13 13:37 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | [5 empty] | 22934 | DOH | laureanjo@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 8 ASSAY: Xpert SARS-CoV-2, COUNT: 40 |  | Gx IV | 838527 | Transfer from CTRL to Cebu CHO | Installation | Installation | Module B4 H/W fail |  |  |  | Raymond Laurito | https://www.jotform.com/uploads/macarewecare/222927978372067/5493978677128867167/5493978677128867167_signature_28.png | Laurean Jo Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5493978677128867167/5493978677128867167_signature_31.png
| 2023-01-13 | 2023-01-13 13:16 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22933 | DOH | dulzna@gmail.com | 25.3 | ASSAY: Xpert MTB/RIF, COUNT: 425 ASSAY: Xpert MTB/RIF Ultra, COUNT: 284 |  | Gx-4 | 838570 | Module B4 waiting for module to available | Genexpert machine Preventive maintenance | Genexpert machine Preventive maintenance | Gx machine has 3 operational modules |  |  |  | Jesson Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5493965614646381135/5493965614646381135_signature_28.png | Dulce Aurora N. Ariston, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5493965614646381135/5493965614646381135_signature_31.png
| 2023-01-13 | 2023-01-13 13:10 | West Dadiangas Rural Health Unit | Brgy. West Dadiangas, General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22932 | PBSP | dulzna@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF, COUNT: 207 ASSAY: Xpert MTB/RIF Ultra, COUNT: 265 |  | Gx-4 | 815428 | N/A | Genexpert machine Preventive maintenance | Genexpert machine Preventive maintenance | Gx machine has 4 operational modules |  |  |  | Jesson Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5493961764647580140/5493961764647580140_signature_28.png | Dulce Aurora N. Ariston, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5493961764647580140/5493961764647580140_signature_31.png
| 2023-01-13 | 2023-01-13 12:41 | Binan City Health Office 1 | Binan Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22931 | PBSP | evalexaa29@gmail.com | 25 | ASSAY: Others, COUNT: Backup laptop | 2023-01-11 | GXIV | 812780 | Unit for PM. For installation of backup laptop. For module replacement of module 1. For checking of UPS. | Performed installation of backup laptop. Performed module replacement on module 1. Performed PM. Performed checking of UPS. | PM, module replacement, installation and troubleshooting. | Unit is operational. 4 operational modules. UPS's power button is malfunctioning, and is recommended for replacement. |  | GX module D1 (886541) in replacement for defective D1 (615860). Backup laptop LenovoThinkpad (CPH1019F03266) in replacement for defective CPU Lenovo thinkcentre (CPH0421F20801). | G 5023 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5493951761422687469/5493951761422687469_signature_28.png | Alexandra Mae Evangelista | https://www.jotform.com/uploads/macarewecare/222927978372067/5493951761422687469/5493951761422687469_signature_31.png
| 2023-01-13 | 2023-01-13 11:55 | San Fernando RHU 2 | Sindalan San Fernando Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22930 | PBSP | punoricahrd@gmail.com | 26.7 | ASSAY: Xpert MTB/RIF, COUNT: 233 ASSAY: Xpert MTB/RIF Ultra, COUNT: 614 |  | GXIV-D | 838879 | Due for maintenance and calibration | Performed preventive maintenance Archived test results Installation and setup of Xpertcheck Performed calibration Gathered reports Performed functionality test | Preventive maintenance and calibration | Unit is operational | All modules passed calibration |  |  | Mark Joseph C. Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5493949876019651824/5493949876019651824_signature_28.png | Ruchard M. Puno | https://www.jotform.com/uploads/macarewecare/222927978372067/5493949876019651824/5493949876019651824_signature_31.png
| 2023-01-13 | 2023-01-13 10:52 | Medical Center Manila | Kalaw, Ermita | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22929 | CONSORTIUM | mdsegui.manilamed@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 10 |  | GX4D | 835540 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed suface Cleaning Performed Plunger Rod Cleaning Performed Filter Cleaning | Unit is operational Module 2 and 3 are for reporting |  |  |  | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5493881994029126069/5493881994029126069_signature_28.png | Michelle Segui | https://www.jotform.com/uploads/macarewecare/222927978372067/5493881994029126069/5493881994029126069_signature_31.png
| 2023-01-13 | 2023-01-13 10:11 | San Pedro Rural Health Unit 2 | Brgy Langgam San Pedro Laguna | 4A | Laguna | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | [5 empty] | 22928 | DOH | splhealth2@gmail.com | 27 | ASSAY: Xpert HPV, COUNT: 853 ASSAY: Xpert MTB/RIF, COUNT: 259 ASSAY: Xpert MTB/RIF Ultra, COUNT: 722 ASSAY: Xpert HIV-1 Viral Load, COUNT: 70 | 2023-01-09 | GXIV | 821189 | Unit for PM. For installation of new monitor. | Performed PM of Genexpert. Installed new monitor. | PM and installation. | Unit is operational. 2 operational modules. |  | Aoc monitor (sn QSDN8HA087213) in replacement for defective HP Monitor (sn CNK831082X). | DR 9448 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5493859455309621658/5493859455309621658_signature_28.png | john carlo insorio | https://www.jotform.com/uploads/macarewecare/222927978372067/5493859455309621658/5493859455309621658_signature_31.png
| 2023-01-13 | 2023-01-12 15:02 | Davao Oriental Provincial Medical Center | National Highway, Brgy. Dahican, Mati City | 11 | Davao Oriental | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22927 | PBSP | canenciaemar@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 173 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1333 | 2023-01-12 | GX-IV 4D | 804693 | Module A3/ SN: 688471 for replacement | Module Replacement Preventive Maintenance | Replaced the Defective Module A3/ SN: 688471 with new Module SN: 725010 | Unit is fully operational | N/A | 900-0491rs | N/a | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5493801716764740600/5493801716764740600_signature_28.png | Mary Rose Salasinas | https://www.jotform.com/uploads/macarewecare/222927978372067/5493801716764740600/5493801716764740600_signature_31.png
| 2023-01-12 | 2023-01-12 20:13 | Taguig Pateros District Hospital | West Service Road, Western Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22926 | DOH | labdepttpdh@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 859 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1525 ASSAY: Xpert HIV-1 Viral Load, COUNT: 677 |  | GX4D | 831638 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Filter Cleaning Performed Plunger Rod Cleaning | Unit is operationanal |  |  |  | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5493368865892871897/5493368865892871897_signature_28.png | Angelie Alviar | https://www.jotform.com/uploads/macarewecare/222927978372067/5493368865892871897/5493368865892871897_signature_31.png
| 2023-01-12 | 2023-01-12 20:13 | Taguig Pateros District Hospital | West Service Road, Western Bicutan | NCR | NCR 4th - Taguig | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22925 | DOH | labdepttpdh@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 859 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1525 ASSAY: Xpert HIV-1 Viral Load, COUNT: 677 |  | GX4D | 831638 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Filter Cleaning Performed Plunger Rod Cleaning | Unit is operationanal |  |  |  | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5493367733522216325/5493367733522216325_signature_28.png | Angelie Alviar | https://www.jotform.com/uploads/macarewecare/222927978372067/5493367733522216325/5493367733522216325_signature_31.png
| 2023-01-12 | 2023-01-12 15:27 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | YES | INSTALLATION | CEPHEID | [5 empty] | 22924 | PBSP | laureanjo@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1796 |  | GX IV | 804840 | Installation | Installation | Installation | Systems operational |  |  |  | Raymond Laurito | https://www.jotform.com/uploads/macarewecare/222927978372067/5493179981165342135/5493179981165342135_signature_28.png | Laurean Jo Cabase | https://www.jotform.com/uploads/macarewecare/222927978372067/5493179981165342135/5493179981165342135_signature_31.png
| 2023-01-12 | 2023-01-11 14:06 | General Tinio RHU | General Tinio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22923 | DOH | jadelyn_pajarillaga@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 174 ASSAY: Xpert MTB/RIF Ultra, COUNT: 288 | 2023-01-11 | Genexpert IV | 831695 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Unit is for calibration this February | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5493178760123608534/5493178760123608534_signature_28.png | Jadelyn B. Pajarillaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5493178760123608534/5493178760123608534_signature_31.png
| 2023-01-12 | 2023-01-12 14:20 | Cabanatuan CHO | Kapitan Pepe, Cabanatuan ,Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22922 | DOH | cholab2021@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1481 ASSAY: Xpert MTB/RIF, COUNT: 1136 | 2023-01-12 | Genexpert IV | 831592 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit Is Operational | Unit is due for Calibration | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5493178760128946514/5493178760128946514_signature_28.png | Maricris SO | https://www.jotform.com/uploads/macarewecare/222927978372067/5493178760128946514/5493178760128946514_signature_31.png
| 2023-01-12 | 2023-01-12 14:15 | Macabebe Rural Health Unit | Macabebe Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22921 | DOH | salvamari13@yahoo.com | 20.3 | ASSAY: Xpert MTB/RIF, COUNT: 284 ASSAY: Xpert MTB/RIF Ultra, COUNT: 554 | 2022-12-20 | GXIV-D | 831893 | Error 4012 on module A1 | Performed preventive maintenance Archived test results Performed calibration on module A1 Gathered reports Performed functionality test | Preventive maintenance and calibration | Unit is operational except modules A2 and A3, sahaky monitor display, alignment error (Auto-adjustment error) | Passed Xpertcheck;Monitor for replacement | 2 Xpertcheck kit | 9410 | Mark Joseph C. Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5493159542515289676/5493159542515289676_signature_28.png | Salva Sese | https://www.jotform.com/uploads/macarewecare/222927978372067/5493159542515289676/5493159542515289676_signature_31.png
| 2023-01-12 | 2023-01-12 09:57 | PHILIPPINE RED CROSS-ZAMBOANGA CITY CHAPTER | ZAMBOANGA CITY | 9 | Zamboanga del Sur | BLOOD BANK | NO | REPAIR CALIBRATION MAINTENANCE | HAEMONETICS | [5 empty] | 22920 | COMMERCIAL | zamboangacity@redcross.org.ph | 28 |  |  | MCS+ | 01M256 | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  | JOSEPH ACEBEDO / GIAN AIZON | https://www.jotform.com/uploads/macarewecare/222927978372067/5492987697025869841/5492987697025869841_signature_28.png | Joan A. Luna | https://www.jotform.com/uploads/macarewecare/222927978372067/5492987697025869841/5492987697025869841_signature_31.png
| 2023-01-11 | 2023-01-10 15:00 | Domingo santiago health center | D.santiago samapaloc mainila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22918 | DOH | tdhiragizawa@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 372 ASSAY: Xpert MTB/RIF Ultra, COUNT: 842 | 2023-01-10 | Gx4 | 831527 | A2 753754 continous 2014 error,A4 unable to perform xpertcheck bec of continous booting | Performed rexpert check, gathered reports, | Rexpert check | A3,A4 disable | Passed rexpertcheck A2 | Xpertcheck kits | 0 | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5492459386428071094/5492459386428071094_signature_28.png | Flory fe Paas | https://www.jotform.com/uploads/macarewecare/222927978372067/5492459386428071094/5492459386428071094_signature_31.png
| 2023-01-11 | 2023-01-10 16:41 | UERMMMC | Aurora blvd qc | NCR | NCR 2nd - Quezon City | Bloodbank | NO | TROUBLESHOOTING | HAEMONETICS | [5 empty] | 22919 | COMMERCIAL | jpppancho11@gmail.com | 0 |  |  | mcs+ | 13C270 | No display | Check unit, check cable, performed simulation run, performed functionality test | Troubleshooting | Operational |  |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5492459386428100774/5492459386428100774_signature_28.png | Jasper Pancho | https://www.jotform.com/uploads/macarewecare/222927978372067/5492459386428100774/5492459386428100774_signature_31.png
| 2023-01-11 | 2023-01-11 14:06 | General Tinio RHU | General Tinio, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22917 | DOH | jadelyn_pajarillaga@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 174 ASSAY: Xpert MTB/RIF Ultra, COUNT: 288 | 2023-01-11 | Genexpert IV | 831695 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Unit is for calibration this February | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5492295420122509141/5492295420122509141_signature_28.png | Jadelyn B. Pajarillaga | https://www.jotform.com/uploads/macarewecare/222927978372067/5492295420122509141/5492295420122509141_signature_31.png
| 2023-01-11 | 2023-01-11 14:36 | Braulio E. Dujali Rural Health Unit | Poblacion, Braulio E. Dujali. Davao del Norte | 11 | Davao del Norte | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22916 | PBSP | sblacorda@yahoo.com | 26.7 | ASSAY: Xpert MTB/RIF, COUNT: 172 ASSAY: Xpert MTB/RIF Ultra, COUNT: 652 | 2023-01-11 | GX-IV 2D | 835216 | Module B1/ SN: 788491 for module replacement | Module Replacement Preventive Maintenance | Replacement of Defective Module B1/ SN 788491 with new Module SN:672770 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lens Generate IQR, Assay Stat and System Log Archived Test Results | Unit is fully operational | N/A | 900-0491rs | N/a | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5492284888911803351/5492284888911803351_signature_28.png | Sharon Rose B. Lacorda RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5492284888911803351/5492284888911803351_signature_31.png
| 2023-01-11 | 2023-01-11 13:25 | Cebu City Medical Center | Cebu City, Cebu | 7 | Cebu | Microbiology Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22915 | DOH | ccmclaboratory01@gmail.com | 24.8 | ASSAY: Xpert MTB/RIF, COUNT: 780 ASSAY: Xpert MTB/RIF Ultra, COUNT: 699 |  | GXIV-4D | 838561 | Due for PM | #NAME? | N/A | Module A2 & A4 - waiting for module allocation. Module A1 & A3 - operational |  |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5492243491618772383/5492243491618772383_signature_28.png | Fatimah Maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5492243491618772383/5492243491618772383_signature_31.png
| 2023-01-11 | 2023-01-11 13:22 | Tampakan Rural Health Unit | Poblacion Tampakan | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22914 | DOH | jmedtech_lab@yahoo.com | 27.2 | ASSAY: Xpert MTB/RIF, COUNT: 123 ASSAY: Xpert MTB/RIF Ultra, COUNT: 385 |  | Gx-4 | 838544 | N/A | Genexpert machine Preventive maintenance | Genexpert machine Preventive maintenance | Gx machine has 4 operational modules |  |  |  | Jesson Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5492242054211889032/5492242054211889032_signature_28.png | Jovi G Gelio | https://www.jotform.com/uploads/macarewecare/222927978372067/5492242054211889032/5492242054211889032_signature_31.png
| 2023-01-11 | 2023-01-11 12:27 | Pilar Rural Health Unit | Pilar Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22913 | DOH | aimna_24@yahoo.com | 26.4 | ASSAY: Xpert MTB/RIF, COUNT: 158 ASSAY: Xpert MTB/RIF Ultra, COUNT: 77 |  | GXIV-D | 831613 | Due for maintenance | Performed surface cleaning Archived test results Performed Preventive maintenance Gatherd reports Performed functionality test | Preventive maintenance | Unit is operational except module A3 |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5492214358811638185/5492214358811638185_signature_28.png | Marciana Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5492214358811638185/5492214358811638185_signature_31.png
| 2023-01-11 | 2023-01-11 12:27 | Pilar Rural Health Unit | Pilar Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22912 | DOH | aimna_24@yahoo.com | 26.4 | ASSAY: Xpert MTB/RIF, COUNT: 158 ASSAY: Xpert MTB/RIF Ultra, COUNT: 77 |  | GXIV-D | 831613 | Due for maintenance | Performed surface cleaning Archived test results Performed Preventive maintenance Gatherd reports Performed functionality test | Preventive maintenance | Unit is operational except module A3 |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5492214338814809362/5492214338814809362_signature_28.png | Marciana Mendiola | https://www.jotform.com/uploads/macarewecare/222927978372067/5492214338814809362/5492214338814809362_signature_31.png
| 2023-01-11 | 2023-01-11 10:43 | Minglanilla Rural Health Unit | Ward 1 Poblacion, Minglanilla Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22911 | DOH | brendagonzagaparedes@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 317 | 2023-01-11 | GXIV | 838512 | No Errors For Preventive Maintenance | FOR Preventive maintenance | 1. Conduct Preventive Maintenance 2. Clean Icore, Clean Plunger Rod 3. Remove and filter and replaced new one 4..Generate IQR, Sys Log, and Gxx files 5.. Back up database 6. Ups and avr 7. Check Cpu and monitor display 8..check bar code scanner | All modules are operational |  | No parts used | No parts used | Mark Efren Ramas | https://www.jotform.com/uploads/macarewecare/222927978372067/5492145751013095852/5492145751013095852_signature_28.png | Brenda G. Paredes | https://www.jotform.com/uploads/macarewecare/222927978372067/5492145751013095852/5492145751013095852_signature_31.png
| 2023-01-11 | 2023-01-11 10:20 | Mingllanilla Rural Health Unit | Ward 1, Mingllanilla Cebu | 7 | Cebu | Tb laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22910 | DOH | brendagonzagaparedes@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 317 | 2023-01-11 | GXIV | 838512 | NO Errors | PREVENTIVE MAINTENANCE | 1. Conduct Preventive Maintenance 2. Clean Icore, Clean Cartridge Bay, Clean Plunger 3. Clean and change Air Filter 4. Check Ups and AVR 5. Generate IQR, Sys Log, Assay statistics 6. Archive Test results and back Up database | All Modules are operational |  | No parts | No parts | Mark Efren Ramas | https://www.jotform.com/uploads/macarewecare/222927978372067/5492140341014367559/5492140341014367559_signature_28.png | Brenda Paredes, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5492140341014367559/5492140341014367559_signature_31.png
| 2023-01-11 | 2023-01-11 10:35 | Cebu TB reference Lab | Cebu city | 7 | Cebu | Laboratory | YES | Pull out | CEPHEID | [5 empty] | 22909 | DOH | asmantiquilla.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1743 |  | Gx IV | 838527 | Pull out | Pull out | Pull out | Systems operational |  |  |  | Raymond Laurito | https://www.jotform.com/uploads/macarewecare/222927978372067/5492140212112300090/5492140212112300090_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5492140212112300090/5492140212112300090_signature_31.png
| 2023-01-11 | 2023-01-11 10:30 | Cebu TB reference lab | Cebu city | 7 | Cebu | Laboratory | YES | Pull out | CEPHEID | [5 empty] | 22908 | PBSP | asmantiquilla.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1743 |  | Gx iv | 804840 | Pull out | Pull out | Pull out | Systems operational |  |  |  | Raymond Laurito | https://www.jotform.com/uploads/macarewecare/222927978372067/5492136792119747274/5492136792119747274_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5492136792119747274/5492136792119747274_signature_31.png
| 2023-01-11 | 2023-01-10 13:14 | West Rembo Health Center | West Rembo | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22906 | DOH | mcabalosjrpbsp@gmail.com | 21 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 594 |  | GX-IV-D | 831712 | Modules A2 and A3 for Re-XpertCheck | Performed XpertCheck | Performed maintenance Cleaned cartridge bay, icore, plunger rod Run XpertCheck on module Generated reports | Module A1 operational. Module A2/757446 passed and operational, under observation. Module A4/757448 failed, for reporting. Module A3/757447 already reported. |  |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5492088644113806708/5492088644113806708_signature_28.png | Modesto Abalos | https://www.jotform.com/uploads/macarewecare/222927978372067/5492088644113806708/5492088644113806708_signature_31.png
| 2023-01-11 | 2023-01-10 10:49 | MR Flores Health Center | 321 P. Herrera | NCR | NCR 4th - Pateros | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22907 | PBSP | paterosclinicallaboratory@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 121 ASSAY: Xpert MTB/RIF Ultra, COUNT: 371 ASSAY: Xpert HIV-1 Viral Load, COUNT: 2 |  | GX-IV-2-D | 834803 | Stuck cartridge | Removed stuck cartridge | Removed stuck cartridge Performed maintenance Cleaned cartridge bay, icore, plunger rod and filter. Generated reports | Unit is operational. |  |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5492088644116042797/5492088644116042797_signature_28.png | JOHN-KARL JIMENEZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5492088644116042797/5492088644116042797_signature_31.png
| 2023-01-11 | 2023-01-10 12:24 | Seamen's Hospital | Intramuros, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22904 | MACARE-HBDC | shmolecularlab@gmail.com | 22 | ASSAY: Xpert SARS-CoV-2, COUNT: 101 |  | GXIV-D | 838836 | Unit is due for Maintenance | Maintenance | Performed Maintenance Generate reports Performed selft test and functionality | Unit is operational |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5492078044114462181/5492078044114462181_signature_28.png | Queenie Lea Serquina | https://www.jotform.com/uploads/macarewecare/222927978372067/5492078044114462181/5492078044114462181_signature_31.png
| 2023-01-11 | 2023-01-10 09:04 | San Lazaro Hospital | Quiricada, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22902 | PBSP | eisel21@yahoo.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 810 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1431 ASSAY: Xpert HIV-1 Viral Load, COUNT: 654 |  | GXIV-D | 804766 | Unit is due for Maintenance | Maintenance | Performed Maintenance Generate Reports Performed Functionality Test | Unit is operational. |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5492078044117514341/5492078044117514341_signature_28.png | Leslie Grace Bundalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5492078044117514341/5492078044117514341_signature_31.png
| 2023-01-11 | 2023-01-10 09:08 | San Lazaro Hospital | Quiricada, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22903 | PBSP | eisel21@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 810 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1431 ASSAY: Xpert HIV-1 Viral Load, COUNT: 654 |  | GXIV-D | 707843 | Unit is due for maintenance | Maintenance | Performed Maintenance Generate Reports Performed Functionality Test | Unit is operational |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5492078044119353751/5492078044119353751_signature_28.png | Leslie Grace Bundalian | https://www.jotform.com/uploads/macarewecare/222927978372067/5492078044119353751/5492078044119353751_signature_31.png
| 2023-01-11 | 2023-01-10 12:21 | Seamen's Hospital | Intramuros, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22905 | MACARE-HBDC | shmolecularlab@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 2028 |  | GX16-D | 838777 | Unit is due for Maintenance | Maintenance | Performed Maintenance Generate Reports Performed Self-Test and Functionality | Unit is Operational |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5492078054119842433/5492078054119842433_signature_28.png | Queenie Lea Serquina | https://www.jotform.com/uploads/macarewecare/222927978372067/5492078054119842433/5492078054119842433_signature_31.png
| 2023-01-10 | 2023-01-10 12:41 | Samal RHU | Samal Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22901 | DOH | vickytholits2002@gmail.com | 27.8 | ASSAY: Xpert MTB/RIF, COUNT: 303 ASSAY: Xpert MTB/RIF Ultra, COUNT: 249 |  | GXIV-D | 831729 | Due for maintenance | Performed surface cleaning Archived test results Performed preventive maintenance Gathered reports Performed functionality test | Preventive maintenance | Unit is operqtional except A1 |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5491523108818652950/5491523108818652950_signature_28.png | Maria Victoria Banzuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5491523108818652950/5491523108818652950_signature_31.png
| 2023-01-10 | 2023-01-10 11:39 | Manila Central University Hospital | MCU Hospital, Morning Breeze Subdivision, Caloocan City | NCR | NCR 3rd - Caloocan | Clinical Laboratory | NO | INSTALLATION | CEPHEID | [5 empty] | 22900 | CONSORTIUM | lab_mcuhospital@yahoo.com | 25 | ASSAY: Others, COUNT: NA |  | GXIV | 110018008 | Unit for installation | - provided 3 1/2 in. of space, table against wall (aircon) - asked facility to provide at least 1 KVA UPS (650VA currently installed) to proceed with training with actual running of test - advised biomed that there is no power coming from outlet (fixed) | - checked GX's environment * Power source = 220 V * Room temperature = 25°C * Heat sources = NA * Table = Sturdy - installed unit - Performed self test on all modules - Generated reports | Unit is operational. All 4 modules are functional. 1 KVA UPS recommended for operation. | - table dimensions (16×57×36 in.) |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5491506930229080246/5491506930229080246_signature_28.png | Gloria Caling | https://www.jotform.com/uploads/macarewecare/222927978372067/5491506930229080246/5491506930229080246_signature_31.png
| 2023-01-10 | 2023-01-10 16:09 | Graman Medical and Maternity Hospital | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22899 | COMMERCIAL | gramanmolecularlab@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 17 |  | Gx Iv | 110008530 | Errors 2008, 2037 | Performed maintenance | #NAME? | Unit is operational. No hardware problems found. Errors possibly due to cartridge. | For reporting |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5491506920223260303/5491506920223260303_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5491506920223260303/5491506920223260303_signature_31.png
| 2023-01-10 | 2023-01-10 15:22 | Talavera General Hospital | Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22898 | MACARE-HBDC | talaverashope@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 100 ASSAY: Xpert HIV-1 Viral Load, COUNT: 49 | 2023-01-10 | Genexpert IV | 110013773 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5491494913913124503/5491494913913124503_signature_28.png | Fitz Gerald Manuzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5491494913913124503/5491494913913124503_signature_31.png
| 2023-01-10 | 2023-01-10 15:22 | Talavera General Hospital | Talavera, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22897 | MACARE-HBDC | talaverashope@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 100 ASSAY: Xpert HIV-1 Viral Load, COUNT: 49 | 2023-01-10 | Genexpert IV | 110013773 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5491494863918747355/5491494863918747355_signature_28.png | Fitz Gerald Manuzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5491494863918747355/5491494863918747355_signature_31.png
| 2023-01-10 | 2023-01-10 14:18 | Davao TB Reference Laboratory | DOH Compound, J.P Laurel Ave. Bajada, Davao City | 11 | Davao del Sur | TB Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22896 | WHO | dtrlmark.pbsp@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF, COUNT: 206 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1086 ASSAY: Others, COUNT: 604 | 2023-01-10 | GX - IV 4D 10-Colors | 848248 | For Preventive Maintenance | Perfrom Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat and System Log | Unit is fully operational | Other Assay Stat : MTB XDR | N/a | N/A | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5491477445226126057/5491477445226126057_signature_28.png | Joseph Mark Anque | https://www.jotform.com/uploads/macarewecare/222927978372067/5491477445226126057/5491477445226126057_signature_31.png
| 2023-01-10 | 2023-01-10 14:11 | Davao TB Reference Laboratory | DOH Compound, J.P Laurel Ave. Bajada. Davao City | 11 | Davao del Sur | TB Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22895 | PBSP | dtrlmark.pbsp@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 229 | 2023-01-10 | GX - IV 4L R1 | 707513 | Module B1 (S/N: 636452) encounter Hardware Failure | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Icore Lenses Generate IQR, Assay Stat, and System Log Archive Data Test Results | Module B1 (S/N: 636452) encounter multiple Temperature Errors, subjected for Module Replacement | N/A | N/a | N/a | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5491476425227880260/5491476425227880260_signature_28.png | Joseph Mark Anque | https://www.jotform.com/uploads/macarewecare/222927978372067/5491476425227880260/5491476425227880260_signature_31.png
| 2023-01-10 | 2023-01-10 15:46 | ZAMBOANGA CITY MEDICAL CENTER | ZAMBOANGA CITY | 9 | Zamboanga del Sur | TB LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22893 | PBSP | jbcua.pbsp@gmail.com | 23.3 | ASSAY: Xpert HPV, COUNT: 265 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4447 ASSAY: Xpert SARS-CoV-2, COUNT: 23356 ASSAY: Xpert HIV-1 Viral Load, COUNT: 219 |  | GX-XVI | 839309 | #NAME? | -PERFORM MODULE C1 REPLACEMEN DEF SN:870952 NEW SN:914832 | #NAME? | #NAME? | #NAME? |  |  | JOSEPH ACEBEDO / GIAN AIZON | https://www.jotform.com/uploads/macarewecare/222927978372067/5491475381915378923/5491475381915378923_signature_28.png | James peniel b cua | https://www.jotform.com/uploads/macarewecare/222927978372067/5491475381915378923/5491475381915378923_signature_31.png
| 2023-01-10 | 2023-01-10 15:22 | Cebu Tb reference lab | Cebu city | 7 | Cebu | Laboratory | YES | Pull out machine transfer to cebu cho | CEPHEID | [5 empty] | 22892 | PBSP | asmantiquilla.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1780 |  | Gx iv | 808620 | Pull out machine | Pull out machine | Pull out machine | Good |  |  |  | Raymond laurito | https://www.jotform.com/uploads/macarewecare/222927978372067/5491449494561473396/5491449494561473396_signature_28.png | Allan S. Mantiquilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5491449494561473396/5491449494561473396_signature_31.png
| 2023-01-10 | 2023-01-10 15:08 | Samboan Rural Health Unit | Poblacion Samboan, Cebu | 7 | Cebu | Tb Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22891 | DOH | sheneannesenagan@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 305 | 2023-01-10 | GX IV | 838513 | No Error | Preventive Maintenance | 1. Conduct Preventive Maintenance 2. Clean Icore area, clean cartridge bay, 3. Clean Plunger, clean Filter 4. Generate IQR, Sys Log, Gxx files, 5. Back Up database, get Assay statistics 4. | All modules are operational |  | NA | No parts replaced | Mark Efren Ramas | https://www.jotform.com/uploads/macarewecare/222927978372067/5491444091324146465/5491444091324146465_signature_28.png | Shene Anne Senagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5491444091324146465/5491444091324146465_signature_31.png
| 2023-01-10 | 2023-01-10 14:56 | Juan Posadas Health Center | Rodriguez corner Nepa, Balut Tondo, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22890 | DOH | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 705 |  | GX4D | 805060 | Unit is due for troubleshooting Preventive Maintenance Lag while getting archive file Mouse is also lag and qr code scanner also lag | Troubleshooting Preventive Maintenance | Performed Preventive Maintenance Performed Troubleshooting Reset RAM/plug unplug ram while shutdown | Unit is operational | Unit is required CPU replacement |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5491437797317575366/5491437797317575366_signature_28.png | ELIZA JUNIO | https://www.jotform.com/uploads/macarewecare/222927978372067/5491437797317575366/5491437797317575366_signature_31.png
| 2023-01-10 | 2023-01-10 14:56 | Juan Posadas Health Center | Rodriguez corner Nepa, Balut Tondo, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22889 | DOH | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 705 |  | GX4D | 805060 | Unit is due for troubleshooting Preventive Maintenance Lag while getting archive file Mouse is also lag and qr code scanner also lag | Troubleshooting Preventive Maintenance | Performed Preventive Maintenance Performed Troubleshooting Reset RAM/plug unplug ram while shutdown | Unit is operational | Unit is required CPU replacement |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5491437697315947583/5491437697315947583_signature_28.png | ELIZA JUNIO | https://www.jotform.com/uploads/macarewecare/222927978372067/5491437697315947583/5491437697315947583_signature_31.png
| 2023-01-10 | 2023-01-10 14:56 | Juan Posadas Health Center | Rodriguez corner Nepa, Balut Tondo, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22888 | DOH | elizajunio05@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 705 |  | GX4D | 805060 | Unit is due for troubleshooting Preventive Maintenance Lag while getting archive file Mouse is also lag and qr code scanner also lag | Troubleshooting Preventive Maintenance | Performed Preventive Maintenance Performed Troubleshooting Reset RAM/plug unplug ram while shutdown | Unit is operational | Unit is required CPU replacement |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5491436127316682194/5491436127316682194_signature_28.png | ELIZA JUNIO | https://www.jotform.com/uploads/macarewecare/222927978372067/5491436127316682194/5491436127316682194_signature_31.png
| 2023-01-10 | 2023-01-10 14:06 | Calumpang RHU | Brgy. Calumpang, General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22887 | DOH | jan19.cam@gmail.com | 25.6 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 747 ASSAY: Xpert MTB/RIF, COUNT: 730 |  | Gx-4 | 838504 | N/A | Genexpert machine Preventive maintenance | Genexpert machine Preventive maintenance | In good condition |  |  |  | Jesson Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5491404304122723904/5491404304122723904_signature_28.png | Camiña, Janice ,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5491404304122723904/5491404304122723904_signature_31.png
| 2023-01-10 | 2023-01-10 08:50 | Philippine Red Cross Laguna | Port Area Manila | NCR | NCR 1st - Manila | Blood bank | NO | MAINTENANCE | HAEMONETICS | [5 empty] | 22886 | COMMERCIAL | lalaine.ilagan@redcross.org.ph | 27 |  |  | MCS+ | 12D063 | Unit for preventive maintenance. | Performed preventive maintenance of Haemonetics MCS+. Cleaned rotor pumps and unit. Verified calibration parameters. Performed simulation running. | Preventive maintenance. | Unit is operational. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5491379393823868418/5491379393823868418_signature_28.png | Lalaine ilagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5491379393823868418/5491379393823868418_signature_31.png
| 2023-01-10 | 2023-01-10 11:16 | Tondo Medical Center | H. Lopez Tondo, Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22885 | MACARE-HBDC | microscan@tmc.doh.gov.ph | 24.6 | ASSAY: Xpert HPV, COUNT: 301 ASSAY: Xpert MTB/RIF, COUNT: 352 ASSAY: Xpert SARS-CoV-2, COUNT: 4045 |  | GX4D | 831872 | Unit is due for Preventive Maintenance Unit is due for Xpertcheck | Preventive Maintenance Calibration | Performed Maintenance Performed Calibration A2/612830 Service Required A3/633967 Service Required | Unit is operational |  |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5491300596812168725/5491300596812168725_signature_28.png | Analyn N.Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5491300596812168725/5491300596812168725_signature_31.png
| 2023-01-10 | 2023-01-10 11:03 | Tondo Medical Center | H. Lopez Highway Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22884 | MACARE-HBDC | microscan@tmc.doh.giv.ph | 24.6 | ASSAY: Xpert HPV, COUNT: 441 ASSAY: Xpert MTB/RIF, COUNT: 487 ASSAY: Xpert MTB/RIF Ultra, COUNT: 757 ASSAY: Xpert SARS-CoV-2, COUNT: 10759 |  | GX16 | 740011 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Plunger Rod Cleaning Performed Filter Cleaning | Unit is operational |  |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5491294886816544756/5491294886816544756_signature_28.png | Analyn N.Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5491294886816544756/5491294886816544756_signature_31.png
| 2023-01-10 | 2023-01-10 11:03 | Tondo Medical Center | H. Lopez Highway Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22883 | MACARE-HBDC | microscan@tmc.doh.giv.ph | 24.6 | ASSAY: Xpert HPV, COUNT: 441 ASSAY: Xpert MTB/RIF, COUNT: 487 ASSAY: Xpert MTB/RIF Ultra, COUNT: 757 ASSAY: Xpert SARS-CoV-2, COUNT: 10759 |  | GX16 | 740011 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Plunger Rod Cleaning Performed Filter Cleaning | Unit is operational |  |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5491294686815097682/5491294686815097682_signature_28.png | Analyn N.Facto | https://www.jotform.com/uploads/macarewecare/222927978372067/5491294686815097682/5491294686815097682_signature_31.png
| 2023-01-10 | 2023-01-09 12:05 | Cuyapo RHU | Cuyapo, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22882 | DOH | shane.verlaine@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 598 ASSAY: Xpert MTB/RIF Ultra, COUNT: 499 | 2023-01-09 | Genexpert IV | 831672 | N/a | N/a | Performed Maintenance Gathered Reports Replaced Monitor Perform Functionality Test | Unit is Operational | N/a | AOC Monitor - QSDN8HA087200 | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5491241773912109385/5491241773912109385_signature_28.png | Verlaine Ascencion Nervi Verueco Lico | https://www.jotform.com/uploads/macarewecare/222927978372067/5491241773912109385/5491241773912109385_signature_31.png
| 2023-01-10 | 2023-01-10 09:31 | San Francisco Rural Health Unit | Brgy. Anao-aon, San Francisco | 13 | Surigao del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22881 | PBSP | yakin920@yahoo.com | 27.1 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 133 |  | GX-2 | 835175 | None | None | Replacing Module B1 with SN: 788751 to SN: 906758 | All modules working |  |  |  | Engr. Faizar S. Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5491238573024460184/5491238573024460184_signature_28.png | Nikki Melody Relliquete | https://www.jotform.com/uploads/macarewecare/222927978372067/5491238573024460184/5491238573024460184_signature_31.png
| 2023-01-10 | 2023-01-09 15:05 | Medilinx | Amorsolo st makati city | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22880 | COMMERCIAL | ma.michelle.guarnes@medilinx.com.ph | 27 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 8 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 6 ASSAY: Xpert MTB/RIF Ultra, COUNT: 19 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 2 |  | Gx4 (10) | 110001647 | For troubleshooting barcode scanner | Configure barcode scanner,gathered reports,test run scanner,performed functionality test, pull out back up bar code scanner # 1403500505966 | Troubleshooting | Operational | Return barcode scanner # 21118010556725 |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5491193374112106526/5491193374112106526_signature_28.png | Ma. Michelle A. Guarnes | https://www.jotform.com/uploads/macarewecare/222927978372067/5491193374112106526/5491193374112106526_signature_31.png
| 2023-01-09 | 2023-01-06 11:31 | Tanauan City Health Office | Tanauan Batangas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22879 | DOH | ramilomher27@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 379 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1189 | 2022-12-21 | GXIV | 821299 | Unit for Preventive Maintenance. For Xpertcheck of modules A1, A2 and A4. | Performed Preventive maintenance of GeneXpert. Performed troubleshooting on module A1. Performed Xpertcheck on Modules A2 and A4. | Preventive maintenance, Xpertcheck and troubleshooting. | Unit is operational. Module A1(731303) is h/w fail and can't be Xpertchecked even after resetting the module. Modules A2(793145) and A4(760160) are Service Required. 1 operational module. |  | Xpertcheck 2pcs | DR 9416 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5490761263825214394/5490761263825214394_signature_28.png | EDITA D. RAMILO | https://www.jotform.com/uploads/macarewecare/222927978372067/5490761263825214394/5490761263825214394_signature_31.png
| 2023-01-09 | 2023-01-06 09:59 | Talisay Rural Health Unit | Talisay Batangas | 4A | Batangas | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22878 | DOH | medcertrhutalisay@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 58 ASSAY: Xpert MTB/RIF Ultra, COUNT: 341 | 2022-12-21 | GXIV | 820201 | Unit for preventive maintenance. For Xpertcheck of module A2. | Performed preventive maintenance of GeneXpert. Performed XpertCheck on module A2. | Preventive maintenance and Xpertcheck. | Unit is operational. Module A2 (721636) service required on Xpertcheck. 3 modules operational. |  | Xpertcheck | DR 9416 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5490761263827358753/5490761263827358753_signature_28.png | Adettlou Rodriguez | https://www.jotform.com/uploads/macarewecare/222927978372067/5490761263827358753/5490761263827358753_signature_31.png
| 2023-01-09 | 2023-01-05 13:13 | BOCAUE RURAL HEALTH UNIT I | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22876 | DOH | zaychi20@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 346 ASSAY: Xpert MTB/RIF Ultra, COUNT: 560 |  | GXIV | 831691 | UMIT FOR XPERTCHECK | PERFORMED CALIBRATION | #NAME? | UNIT IS OPERATIONAL. ALL 4 MODULES ARE FUNCTIONAL. |  |  |  | RONNEL BELIZAR | https://www.jotform.com/uploads/macarewecare/222927978372067/5490619518831870752/5490619518831870752_signature_28.png | Princess Castillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5490619518831870752/5490619518831870752_signature_31.png
| 2023-01-09 | 2023-01-06 15:02 | Eastern Pangasinan District Hospital | Eastern Pangasinan | 3 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22877 | PBSP | epdhlaboratory@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 70 |  | GXIV | 815394 | CPU for installation. | Installed CPU | #NAME? | unit is operational. | Module D4 for follow up. |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5490619518832356265/5490619518832356265_signature_28.png | FergyI Aries Escano | https://www.jotform.com/uploads/macarewecare/222927978372067/5490619518832356265/5490619518832356265_signature_31.png
| 2023-01-09 | 2023-01-09 14:15 | Sasa District Health Center | Sasa, Davao City | 11 | Davao del Sur | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22874 | PBSP | mariarosario_mapayo@yahoo.com.sg | 27.6 | ASSAY: Xpert MTB/RIF, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 275 ASSAY: Xpert SARS-CoV-2, COUNT: 2039 | 2023-01-09 | GX-IV 4D | 838889 | Modules B2 (S/N: 863783) and B4 (S/N: 863785) are for replacement. For Preventive Maintenance | Perform Module Replacement and Prevent Maintenence | Replaced the Defective Module B2 (S/N: 863783) with new Module S/N: 789427 Replaced the defective Module B4 (S/N:863785) with new Module S/N: 635986 Clean Surface Area, Cartridge Bay, Plunger Rod, Fan Filters, Icore Lenses, Generate IQR, System Log, Assay Stat Archived Test Results | Unit is fully Operational | Assay stat Sars CoV count is from SEHI | 900-0491rs | N/a | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5490586187126264049/5490586187126264049_signature_28.png | Maria Rosario Mapayo Pino | https://www.jotform.com/uploads/macarewecare/222927978372067/5490586187126264049/5490586187126264049_signature_31.png
| 2023-01-09 | 2023-01-09 14:01 | Guimba Community Hospital | Guimba, Nueva Ecija | 3 | Nueva Ecija | Laboratoey | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22873 | PBSP | marquezandrelyn18@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 327 ASSAY: Xpert MTB/RIF Ultra, COUNT: 656 | 2023-01-09 | Genexpert IV | 807560 | N/a | N/a | Performed Maintenance Performed Calibration Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck B3 and B4 Failed and B1 and B2 passed | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5490548183328162046/5490548183328162046_signature_28.png | Andrelyn N. Marquez | https://www.jotform.com/uploads/macarewecare/222927978372067/5490548183328162046/5490548183328162046_signature_31.png
| 2023-01-09 | 2023-01-09 13:25 | Norala District Hospital | Poblacion Norala | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22872 | DOH | NoralaDHlaboratory@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF, COUNT: 414 ASSAY: Xpert MTB/RIF Ultra, COUNT: 566 |  | Gx -4 | 838543 | N/A | Genexpert machine Preventive maintenance | Genexpert machine Preventive maintenance | In good condition |  |  |  | Jesson Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5490516124589820585/5490516124589820585_signature_28.png | Reziel Deanon Sabido | https://www.jotform.com/uploads/macarewecare/222927978372067/5490516124589820585/5490516124589820585_signature_31.png
| 2023-01-09 | 2023-01-09 12:08 | Hermosa Rural Health Unit | Hermosa Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22871 | DOH | kryztynangela@gmail.com | 26.4 | ASSAY: Xpert MTB/RIF, COUNT: 239 ASSAY: Xpert MTB/RIF Ultra, COUNT: 191 |  | GXIV-D | 831645 | Due for maintenance | Performed surface cleaning Cleaning of filter,plunger rod,iqor and cartridge bay Archived test results Gathered reports Performed functionality test | Preventive maintenance | Unit is operational |  |  |  | Mark Joseph C. Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5490471228815095087/5490471228815095087_signature_28.png | Mary Ann Patambang | https://www.jotform.com/uploads/macarewecare/222927978372067/5490471228815095087/5490471228815095087_signature_31.png
| 2023-01-09 | 2022-11-17 17:07 | Bicol South Luzon Subnational Reference Laboratory/BRDL | BRTTH Compound, Legazpi City | 5 | Albay | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22870 | PBSP | lrnosares.pbsp@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2022-11-17 | Genexpert 4D | 808636 | Unit is due for CPU replacement | CPU Replacement | Performed CPU Replacement Performed Preventive Maintenance | Unit is operational |  | Lenovo ThinkCenter CPU/ CPH1121F27252 | PBSP | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5490453311214639661/5490453311214639661_signature_28.png | LHOUIELYN NOSARES | https://www.jotform.com/uploads/macarewecare/222927978372067/5490453311214639661/5490453311214639661_signature_31.png
| 2023-01-09 | 2022-11-17 17:07 | Bicol South Luzon Subnational Reference Laboratory/BRDL | BRTTH Compound, Legazpi City | 5 | Albay | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22869 | PBSP | lrnosares.pbsp@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2022-11-17 | Genexpert 4D | 808636 | Unit is due for CPU replacement | CPU Replacement | Performed CPU Replacement Performed Preventive Maintenance | Unit is operational |  | Lenovo ThinkCenter CPU/ CPH1121F27252 | PBSP | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5490453201212835733/5490453201212835733_signature_28.png | LHOUIELYN NOSARES | https://www.jotform.com/uploads/macarewecare/222927978372067/5490453201212835733/5490453201212835733_signature_31.png
| 2023-01-09 | 2023-01-06 11:50 | Naic Rural Health Unit | Capt C Nazareno St, Naic | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22864 | PBSP | jengcasamar2007@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 866 |  | GXIV-D | 812006 | Stuck Cartridge in Module 2 | Removed Cartridge Manually | Performed maintenance. Removed Stuck Cartridge and checked horn cable. Run Reliability Test Generate reports. | 2 modules are defective, 2 modules are opertional |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416914644954/5490449416914644954_signature_28.png | Jennifer Casamar | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416914644954/5490449416914644954_signature_31.png
| 2023-01-09 | 2023-01-05 12:27 | Rosario Rural Health Unit | Rosario municipal compound | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | Brenguiler Bagbaguen |  |  |  |  | 22866 | DOH | arlenepenaarmijo@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 220 ASSAY: Xpert MTB/RIF Ultra, COUNT: 823 |  | GXIV-D | 838537 | Lost comm in module 2 Temoerature error in module 4 | Open gx machine Xpertcheck | Reset gateway board Performed xpertcheck in module 4 Performed maintenance Generate reports | 2 modules defective, 2 module operational |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416915172633/5490449416915172633_signature_28.png | Arlene P. Armijo | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416915172633/5490449416915172633_signature_31.png
| 2023-01-09 | 2023-01-04 11:30 | Icasiano Health Center | Pedro gil | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22865 | DOH | lorrainefabros@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 525 ASSAY: Xpert MTB/RIF Ultra, COUNT: 961 |  | GXIV-D | 831684 | Cpu and Monitor Problem | Troubleshoot Maintenance | Troubleshoot CPU and monitor Performed maintenance Take pictures of problem | Not operational. CPU and monitor for Replacement. |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416917670495/5490449416917670495_signature_28.png | Lorraine Lucila L.Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416917670495/5490449416917670495_signature_31.png
| 2023-01-09 | 2023-01-04 09:39 | San Diego Health Center | San Diego, Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22867 | MACARE-HBDC | johndarrellturato@gmail.com | 28 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 699 |  | GXIV-D 10-C | 848262 | Module 3 with heater problem | Maintenance and XpertCheck | Performed Maintenance Performed Xpertcheck in module 3 Generate reports | 1 defective, 3 operational | B3/210002569 - Failed |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416918896084/5490449416918896084_signature_28.png | John Darrell Turato | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416918896084/5490449416918896084_signature_31.png
| 2023-01-09 | 2023-01-05 14:23 | General Trias City Health Office | General Trias | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22868 | PBSP | mackchico00@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1068 |  | GXIV-D | 804691 | Stuck Cartridge in Module 1 | Self-Test in module | Performed Maintenance Performed Self-Test in Module 1 Generate Reports | All modules are operational. Gx unit from Antipolo CHO |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449426915895809/5490449426915895809_signature_28.png | Mack Alfred L. Chico | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449426915895809/5490449426915895809_signature_31.png
| 2023-01-09 | 2023-01-04 11:22 | Pedro Gil Health Center | Pedro gil | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22863 | PBSP | lorrainefabros@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV-D | 812829 | CPU problem | Replaced CPU Performed Maintenance | Performed maintenance Replaced with New CPU Generate reports | Unit is operational |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416919998994/5490449416919998994_signature_28.png | Lorraine Lucila L.Fabros | https://www.jotform.com/uploads/macarewecare/222927978372067/5490449416919998994/5490449416919998994_signature_31.png
| 2023-01-09 | 2023-01-06 13:38 | Nagpayong Super Health Center | Nagpayong, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22862 | DOH | michelindg23@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 544 | 2022-09-29 | GX-IV-D | 831905 | For XpertCheck Module B2 error 2014. Module B3 error 1001/1002 | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. B3/786244 and B4/758212 failed in XpertCheck. Module B2/758210 under observation. | Modules B2 B3 and B4 not operational. UPS not powering on. | XpertCheck-CE-5 | DR#9217 | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433271215618451/5490433271215618451_signature_28.png | Michelle de guzman | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433271215618451/5490433271215618451_signature_31.png
| 2023-01-09 | 2023-01-06 10:33 | Pasig Treatment Hub | Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22861 | MACARE-HBDC | jhozza.josermt@gmaol.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 50 ASSAY: Xpert HIV-1 Viral Load, COUNT: 803 |  | GX-IV-L | 821648 | Module A1/733228 | Run XpertCheck on module A1 | Performed maintenance Run XpertCheck on module A1 Generated reports | Unit is operational. | Module A1/733228 passed the XpertCheck. Under observation. |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433271216145360/5490433271216145360_signature_28.png | Jhozza Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433271216145360/5490433271216145360_signature_31.png
| 2023-01-09 | 2023-01-05 09:27 | Cainta City Health Office | Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22856 | DOH | rowenatec@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 478 ASSAY: Xpert HIV-1 Viral Load, COUNT: 31 |  | Gx-IV-D | 831623 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay,icore,plunger rod Generated reports | Unit is operational. | A3/ 754520 HW Failed |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261213096553/5490433261213096553_signature_28.png | Rowena T. Laxamana | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261213096553/5490433261213096553_signature_31.png
| 2023-01-09 | 2023-01-05 13:45 | Marikina City Health Office | Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22859 | DOH | amymitra22@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1142 ASSAY: Xpert HIV-1 Viral Load, COUNT: 85 |  | GX-IV-D | 831588 | Module A3/657408 has error 1001/1002 | Run XpertCheck on module A3 | Performed maintenance Run XpertCheck on module A3 Cleaned cartridge bay, icore, plunger rod Generated reports. Run reliability test on Module A3. | Unit is operational. Module A2 and A4 not operational already reported. UPS not powering on. | Module A3 passed Xpertcheck. Under observation. Failed in self test. |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261213626801/5490433261213626801_signature_28.png | Amelia Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261213626801/5490433261213626801_signature_31.png
| 2023-01-09 | 2023-01-05 09:40 | Cainta Health Office | Cainta,Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22860 | PBSP | rowenatec@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 446 ASSAY: Xpert HIV-1 Viral Load, COUNT: 25 |  | GX-IV-D | 838182 | For Maintenance | Performed Maintenance | Performed Maintenance cleaned cartridge bay, icore,plunger rod Generated reports | Unit is operational. | A2/714059 continuous booting |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261213720839/5490433261213720839_signature_28.png | Rowen T. LAxamana | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261213720839/5490433261213720839_signature_31.png
| 2023-01-09 | 2023-01-05 13:51 | Marikina City Health Office | Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22858 | PBSP | amymitra22@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1237 ASSAY: Xpert HIV-1 Viral Load, COUNT: 126 |  | GX-IV-D | 804254 | For Maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, icore, plunger rod Generated reports | Unit is operational. | Module D2 not operational. Already reported. Module |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261217006531/5490433261217006531_signature_28.png | Amelia Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261217006531/5490433261217006531_signature_31.png
| 2023-01-09 | 2023-01-05 09:34 | Cainta City Health Office | Cainta, Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22857 | PBSP | rowenatec@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 688 |  | GX-IV-D | 804687 | for Maintenance | Performed Maintenance | performed maintenance cleaned cartridge bay, icore, plunger rod Generated reports | All modules operational. |  |  |  | Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261219940532/5490433261219940532_signature_28.png | Rowena T. Laxamana | https://www.jotform.com/uploads/macarewecare/222927978372067/5490433261219940532/5490433261219940532_signature_31.png
| 2023-01-09 | 2023-01-09 10:14 | PHILIPPINE RED CROSS - ORMOC CHAPTER | ORMOC CITY | 8 | Leyte | BLOOD BANK | NO | CALIBRATION MAINTENANCE | HAEMONETICS | [5 empty] | 22855 | COMMERCIAL | limroseline25@gmail.com | 23 |  |  | MCS+ | 15H043G | #NAME? | #NAME? | #NAME? | #NAME? | #NAME? |  |  | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5490404878816025333/5490404878816025333_signature_28.png | ROSELINE L. PURA | https://www.jotform.com/uploads/macarewecare/222927978372067/5490404878816025333/5490404878816025333_signature_31.png
| 2023-01-07 | 2023-01-06 11:15 | Cuyegkeng health center | Layug st pasay city | NCR | NCR 4th - Pasay | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22853 | PBSP | LBCANICULA@gmail.com | 32 | ASSAY: Xpert MTB/RIF, COUNT: 128 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1555 | 2022-12-19 | Gx4 | 805106 | For module replacement, b1 609217 - 807850 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit,performed module replacement | Module replacement,maintenance | Operational |  | 900gx | G4978 | Noel floresca/ Eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5488657776427347267/5488657776427347267_signature_28.png | Lorna Canicula | https://www.jotform.com/uploads/macarewecare/222927978372067/5488657776427347267/5488657776427347267_signature_31.png
| 2023-01-07 | 2023-01-06 14:53 | Makati medical center | Amorsolo st makati city | NCR | NCR 4th - Makati | Bloodbank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | [5 empty] | 22854 | COMMERCIAL | biomedsupervisor@makatimed.net.ph | 0 |  |  | Mcs + | 97H330 | Due for maintenance and calibration | cleaned unit, checked valve,rollers,line sensor,display panel,dpm,spm,cuff,performed calibration,performed functionality test | PM and calibration | Operational |  |  |  | Noel floresca/Eron dumaguing | https://www.jotform.com/uploads/macarewecare/222927978372067/5488657776427813164/5488657776427813164_signature_28.png | SHIERWEN CONCRENIO L | https://www.jotform.com/uploads/macarewecare/222927978372067/5488657776427813164/5488657776427813164_signature_31.png
| 2023-01-06 | 2023-01-06 18:36 | Medilinx | Makati city | NCR | NCR 4th - Makati | Molecular | NO | Pullout | CEPHEID | [5 empty] | 22852 | COMMERCIAL | John.Manzano@sodexo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 0 |  | Gx 4 | 805365 | For pull out gx4 805365,1kva ups at1000170815st0163, hp laptop elite book cnu4120vp | Pullout | Pullout | Pullout |  |  |  | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5488109703424870251/5488109703424870251_signature_28.png | Medilinx | https://www.jotform.com/uploads/macarewecare/222927978372067/5488109703424870251/5488109703424870251_signature_31.png
| 2023-01-06 | 2023-01-06 17:36 | Makati medical center | Amorsolo st makati city | NCR | NCR 4th - Makati | Molecular | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22851 | COMMERCIAL | John.Manzano@sodexo.com | 28 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 2 ASSAY: Xpert C. difficile BT, COUNT: 9 ASSAY: Xpert CT/NG, COUNT: 2 ASSAY: Others, COUNT: 7 ASSAY: Xpert HBV VL, COUNT: 10 ASSAY: Xpert HPV, COUNT: 3 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6 ASSAY: Others, COUNT: 2 ASSAY: Xpert HCV VL, COUNT: 2 ASSAY: Xpert HIV-1 Viral Load, COUNT: 11 |  | Gx4 | 110001647 | Bar code scanner encounter double entry 21118010556725 | Check cable,performed troubleshooting,set default | Check up | For observation barcode scanner |  |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5488080843426607583/5488080843426607583_signature_28.png | John Paul S. Manzano | https://www.jotform.com/uploads/macarewecare/222927978372067/5488080843426607583/5488080843426607583_signature_31.png
| 2023-01-06 | 2023-01-06 13:13 | Abucay RHU | Abucay Bataan | 3 | Bataan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22850 | DOH | parreraledelina22176@gmal.com | 29.5 | ASSAY: Xpert MTB/RIF, COUNT: 212 ASSAY: Xpert MTB/RIF Ultra, COUNT: 157 | 2022-12-20 | GXIV-D | 831878 | Error 2014 on moduke B1, Re-Xpertcheck for module B2 | Performed preventive maintenance Archived test result Peroformed Xpertcheck for modules B1 and B2 Performed functionality test | Preventive maintenance and calibration | Unit is operational except A4; under observation. | Passed Xpertcheck | 2 Xpertcheck kits | 9410 | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5488010719242398052/5488010719242398052_signature_28.png | Ledelina M. Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5488010719242398052/5488010719242398052_signature_31.png
| 2023-01-06 | 2023-01-06 10:50 | Bongabon District Hospital | Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22849 | DOH | teodysantuajr@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 166 ASSAY: Xpert MTB/RIF Ultra, COUNT: 137 | 2023-01-05 | Genexpert IV | 831562 | Module 4 temperature error | Performed xpertcheck on B4 module | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck B4 - Failed | Xpertcheck 2 pc | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5487996450021462012/5487996450021462012_signature_28.png | Teodorico Santua | https://www.jotform.com/uploads/macarewecare/222927978372067/5487996450021462012/5487996450021462012_signature_31.png
| 2023-01-06 | 2023-01-06 14:17 | Bansalan Rural Health Unit | Poblacion, Bansalan | 11 | Davao del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22848 | PBSP | ellen68nebrada@gmail.com | 27.4 | ASSAY: Xpert MTB/RIF, COUNT: 209 ASSAY: Xpert MTB/RIF Ultra, COUNT: 366 | 2023-01-06 | GX-IV 4D | 838181 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Cartridge Bay, replaced Fan Filters, Plunger Rod, Icore Lenses Generate IQR, Assay Stat, and System Log Archived Test Results | Module B3 and B4 encounter multiple Temperature Error, | Errors reported through JotForm | N/A | N/A | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5487958776712601147/5487958776712601147_signature_28.png | ELLEN N. MAYOR | https://www.jotform.com/uploads/macarewecare/222927978372067/5487958776712601147/5487958776712601147_signature_31.png
| 2023-01-06 | 2023-01-06 11:35 | Tablon Health Center | Tablon, Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | NO | GeneXpert Equipment Site Transfer | CEPHEID | [5 empty] | 22847 | DOH | chocdodiagnosticlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2274 |  | GeneXpert | 838435 | N/A | N/A | GeneXpert Equipment Transfer | Operational |  |  |  | Kim Ryan Arrojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5487869500967140689/5487869500967140689_signature_28.png | Charisa Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5487869500967140689/5487869500967140689_signature_31.png
| 2023-01-06 | 2023-01-06 11:40 | Cagayan de Oro City Health Office | Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | INSTALLATION | CEPHEID | [5 empty] | 22846 | DOH | chocdodiagnosticlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2274 |  | GeneXpert | 838441 | N/A | N/A | Install additional Genexpert equipment | 2 GeneXpert equipment installed are operational |  |  |  | Kim Ryan Arrojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5487869050967367567/5487869050967367567_signature_28.png | Charisa Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5487869050967367567/5487869050967367567_signature_31.png
| 2023-01-06 | 2023-01-06 11:20 | Bataan General Hospital | Balanga Bataan | 3 | Bataan | PCR Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22845 | MACARE-HBDC | honeyeichel@gmail.com | 27.35 | ASSAY: Xpert SARS-CoV-2, COUNT: 17894 | 2022-12-20 | GXIV-D | 829390 | B2 | Performed preventive maintenance Archived test results Performed xpertcheck Gathered reports Performed functionality test | Preventive maintenance and calibration | Unit is operational except for modukes B2,B3 and D2 | Failed Xpertcheck on module B2: | Xpertcheck | 9410 | Mark Joseph C. Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5487868234416548596/5487868234416548596_signature_28.png | Honey Eichel M. Caringal | https://www.jotform.com/uploads/macarewecare/222927978372067/5487868234416548596/5487868234416548596_signature_31.png
| 2023-01-06 | 2023-01-06 11:48 | Cebu City Health Office | Cebu City | 7 | Cebu | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22844 | COMMERCIAL | jonnifercodoyceniza@gmail.com | 25 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 1065 |  | GX IV | 829419 | None | None | PM | Good and operational |  |  |  | Raymond Laurito | https://www.jotform.com/uploads/macarewecare/222927978372067/5487866577179223604/5487866577179223604_signature_28.png | Jonnifer c. Ceniza | https://www.jotform.com/uploads/macarewecare/222927978372067/5487866577179223604/5487866577179223604_signature_31.png
| 2023-01-06 | 2023-01-06 11:48 | Cagayan de Oro City Health Office | Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | NO | INSTALLATION GeneXpert Equipment Network | CEPHEID | [5 empty] | 22843 | DOH | chocdodiagnosticlab@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2274 |  | GeneXpert | 838435 | N/A | N/A | Additional installation of GeneXpert Equipment using a network switch | Both Genexpert equipment are operational |  |  |  | Kim Ryan Arrojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5487865580967006519/5487865580967006519_signature_28.png | Charisa Galacio | https://www.jotform.com/uploads/macarewecare/222927978372067/5487865580967006519/5487865580967006519_signature_31.png
| 2023-01-06 | 2023-01-06 10:46 | Lianga Rural Health Unit | Purok 2, Poblacion, Lianga | 13 | Surigao del Sur | RTDL | NO | REPAIR | CEPHEID | [5 empty] | 22842 | PBSP | gyvxs@yahoo.com | 24.9 | ASSAY: Xpert MTB/RIF, COUNT: 228 ASSAY: Xpert MTB/RIF Ultra, COUNT: 200 |  | GX-2 | 835174 | None | None | Replacing Module B1 w/ SN: 788658 to SN: 669053 | All modules working |  |  |  | Engr. Faizar S. Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5487833535919977284/5487833535919977284_signature_28.png | Gyver I. Sistona | https://www.jotform.com/uploads/macarewecare/222927978372067/5487833535919977284/5487833535919977284_signature_31.png
| 2023-01-06 | 2023-01-06 10:32 | Lake Sebu RHU | Lake Sebu | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22841 | DOH | jennypingkian23@yahoo.com | 25.7 | ASSAY: Xpert MTB/RIF, COUNT: 339 ASSAY: Xpert MTB/RIF Ultra, COUNT: 206 |  | Gx-4 | 838569 | N/A | Genexpert machine Preventive maintenance | Genexpert machine Preventive maintenance | Gx machine In Good condition |  |  |  | Jesson Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5487826823849783775/5487826823849783775_signature_28.png | Jenny A. Pingki-an | https://www.jotform.com/uploads/macarewecare/222927978372067/5487826823849783775/5487826823849783775_signature_31.png
| 2023-01-06 | 2023-01-06 09:45 | Roxas Rural Health Unit | Langatian pres. Manuel A. Roxas, Zamboanga del norte | 9 | Agusan del Norte | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | [5 empty] | 22840 | PBSP | Maythel.tres@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 683 | 2023-01-06 | GX 4 | 834449 | None | Module replacement, preventive maintenance | Module replacement, preventive maintenance | All 4 modules are working | None | N/A | N/A | Gian Carlo Aizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5487796476319513722/5487796476319513722_signature_28.png | Maythel c. Tomimbang | https://www.jotform.com/uploads/macarewecare/222927978372067/5487796476319513722/5487796476319513722_signature_31.png
| 2023-01-06 | 2023-01-06 09:56 | Davao del Norte Hospital - Igacos | Island Garden City of Samal, Davao del Norte | 11 | Davao del Norte | Laboratory | YES | REPAIR CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22839 | PBSP | allensoriano92@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 155 ASSAY: Xpert MTB/RIF Ultra, COUNT: 995 |  | GXIV-4D | 815383 | -Module B1/602699 - Xpertcheck service required -2 modules retry for calibration | -Replace defective module SN: 602699 with new module SN: 913008 -Run Xpertcheck on module B3 & B4 | Preventive maintenance | Module B1 & B2 are functional. Module B3 & B4 service required/for replacement. | GeneXpert area accumulates dust quickly. |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5487796297729480323/5487796297729480323_signature_28.png | Allen Glenn Soriano | https://www.jotform.com/uploads/macarewecare/222927978372067/5487796297729480323/5487796297729480323_signature_31.png
| 2023-01-06 | 2023-01-05 15:20 | Kolambugan Rural Health Unit | Poblacion, Kolambugan, Lanao del Norte | 10 | Misamis Oriental | TB DOTS | YES | INSTALLATION Module Replacement | CEPHEID | [5 empty] | 22838 | PBSP | yojecinaj@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 517 |  | GeneXpert | 815272 | H/W Fail Module A2 SN : 695606 LCD Monitor Display Issue | Module Replacement Monitor Replacement | Module Replacement GeneXpert 6-Color Module SN: 657677 Monitor Replacement Monitor 21.5" Flat Panel | Operational |  |  |  | Kim Ryan Arrojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5487781562026795383/5487781562026795383_signature_28.png | Janice Joy R. Siso | https://www.jotform.com/uploads/macarewecare/222927978372067/5487781562026795383/5487781562026795383_signature_31.png
| 2023-01-05 | 2023-01-05 16:08 | National kidney and transplant institute | East ave. Diliman q.c | NCR | NCR 2nd - Quezon City | Covid lab | NO | REPAIR | CEPHEID | [5 empty] | 22836 | MACARE-HBDC | molecularlabnkti@gmail.com | 24 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 144 ASSAY: Others, COUNT: 16 ASSAY: Others, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 9 ASSAY: Xpert MTB/RIF, COUNT: 1002 ASSAY: Xpert SARS-CoV-2, COUNT: 15782 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Others, COUNT: 9 | 2023-01-05 | Gxxvi | 110005587 | For module replacement a4 210028984 - 882420 D3 210028995 -800477 | Performed module replacement Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit | Module replacement,maintenance | Operational |  | 900gx | G5014 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5487224980017251749/5487224980017251749_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5487224980017251749/5487224980017251749_signature_31.png
| 2023-01-05 | 2023-01-05 13:14 | Bagong pag Asa health center | Quezon | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22837 | DOH | erlindalorilla026@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 60 ASSAY: Xpert MTB/RIF, COUNT: 250 ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 |  | Gx4 | 831597 | Loss communication during process module A2 | Cleaned unit,filter,cartridge bay,icore,plunger rod,gathered reports,test run unit Swap A2 module location to A1 position,replaced cable | Troubleshooting | For observation A2 module |  |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5487224980018582249/5487224980018582249_signature_28.png | Ellendale Lorilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5487224980018582249/5487224980018582249_signature_31.png
| 2023-01-05 | 2023-01-05 14:34 | Central Luzon Doctor's Hospital | Hospital drive Tarlac City | 3 | Tarlac | Molecular Laboratory | NO | INSTALLATION REPAIR | CEPHEID | [5 empty] | 22835 | COMMERCIAL | rohmhardameg@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 30 ASSAY: Xpert SARS-CoV-2, COUNT: 535 |  | GXIV-D | 830841 | CPU replacement | Performed CPU replacement Backup data Performed functionqlity test | Repair and installation | Unit is operational except B1 | Backup old files to follow |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5487135547153820721/5487135547153820721_signature_28.png | Rohmhar Dameg | https://www.jotform.com/uploads/macarewecare/222927978372067/5487135547153820721/5487135547153820721_signature_31.png
| 2023-01-05 | 2023-01-05 14:34 | Philippine Orthopedic Center | Banawe st. Quezon City | NCR | NCR 2nd - Quezon City | TB Dots Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22834 | PBSP | cmblanceta@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2023-01-05 | GX4D | 803088 | Unit is due for backup laptop installation | Laptop Installation | Backup Laptop Replacement | Unit is operational |  | Laptop Lenovo Thinkpad | G-5013 | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5487100018916604204/5487100018916604204_signature_28.png | Stella Maree C. Bernardo,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5487100018916604204/5487100018916604204_signature_31.png
| 2023-01-05 | 2023-01-05 13:58 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22833 | DOH | wcpatron11x4@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 513 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1399 | 2023-01-03 | GXIV | 831534 | Unit for Preventive Maintenance. For Xpertcheck of Modules 1, 2 and 3. | Performed preventive maintenance. Performed Xpertcheck on Modules 1, 2 and 3. | Preventive Maintenance and Xpertcheck. | Unit is operational. Module A2 (753694) service required. For replacement of keyboard, some keys are not functioning. 3 working modules. |  | Xpertcheck 3pcs | DR 9416 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5487096039818280371/5487096039818280371_signature_28.png | Wilma Patron | https://www.jotform.com/uploads/macarewecare/222927978372067/5487096039818280371/5487096039818280371_signature_31.png
| 2023-01-05 | 2023-01-05 13:52 | San Francisco Rural Health Unit | Purok 3, Barangay 4, San Francisco | 13 | Agusan del Sur | RTDL | NO | REPAIR | CEPHEID | [5 empty] | 22832 | PBSP | periabrasangelica@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF, COUNT: 466 ASSAY: Xpert MTB/RIF Ultra, COUNT: 229 |  | GX-2 | 835078 | None | None | Replacing Module B1 w/ SN: 682799 to 674763 | All modules working |  |  |  | Engr. Faizar S. Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5487081794865061343/5487081794865061343_signature_28.png | Angelica Periabras | https://www.jotform.com/uploads/macarewecare/222927978372067/5487081794865061343/5487081794865061343_signature_31.png
| 2023-01-05 | 2023-01-05 08:47 | Cotabato City Health Office | RH10, COTABATO CITY | BARMM | Maguindanao | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22831 | DOH | onehope834@gmail.com | 20.9 | ASSAY: Xpert MTB/RIF, COUNT: 932 ASSAY: Xpert MTB/RIF Ultra, COUNT: 680 | 2023-01-05 | GX IV- 4D | 838528 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Cartridge Bay, surface Area, Plunger Rod, Fan Filters, and Icore Lenses Generate IQR, Assay Stat, and System Log Archived Data test results | Unit is fully operational | N/A | N/A | N/a | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5486913409973643053/5486913409973643053_signature_28.png | EMELYN AGUSTIN CADIZ, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5486913409973643053/5486913409973643053_signature_31.png
| 2023-01-04 | 2023-01-04 07:50 | Hi-Precision Diagnostics | N.S. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22830 | CONSORTIUM | pcr@hi-precision.com.ph | 27 | ASSAY: Xpert HPV, COUNT: 1 ASSAY: Xpert HPV, COUNT: 42 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3165 |  | GXIV | 835602 | Unit for Preventive Maintenance. For troubleshooting of Module A2. | Performed Preventive Maintenance. Cleaned docking bay, plunger, icore & filter. Performed troubleshooting of module A2. Performed selftest & functionality test. | Preventive Maintenance and troubleshooting. | Unit is operational. Enabled module 2 for running. 4 modules operational. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5486437659819095693/5486437659819095693_signature_28.png | Rheizza Daphne Salvilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5486437659819095693/5486437659819095693_signature_31.png
| 2023-01-04 | 2023-01-05 09:49 | Tondo Health Center | Gagalangin Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22828 | DOH | ma.lorinamanera@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 449 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1078 | 2023-01-03 | GXIV | 831620 | Unit for Preventive Maintenance. Unit for Xpertcheck. | Performed Preventive Maintenance of GeneXpert. Performed Xpertcheck. | Preventive Maintenance and Xpertcheck. | Unit is operational. Module A1 (638087), A2 (678005) and A3 (702275) are Service Required. |  | Xpertcheck 4pcs | DR 9416 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5486437689816165285/5486437689816165285_signature_28.png | Lontoc, Ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5486437689816165285/5486437689816165285_signature_31.png
| 2023-01-04 | 2023-01-04 14:11 | Parola Health Center | Parola Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22829 | DOH | friaslynette@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: N/a | 2023-01-03 | GXIV | 831871 | Unit for PM. For installation of backup CPU. For checking of monitor. | Performed PM of GeneXpert. Performed installation of backup CPU. Checked monitor. | Preventive Maintenance and installation of CPU. | Unit is operational. Monitor has dead pixel on bottom part horizontally and is subject for replacement. |  | HP Z240 SFF Workstation (sn 2UA8471VR2) as backup in replacement for defective CPU (sn 2UA8471VZN). | N/a | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5486437699817139537/5486437699817139537_signature_28.png | Lynette Olarte | https://www.jotform.com/uploads/macarewecare/222927978372067/5486437699817139537/5486437699817139537_signature_31.png
| 2023-01-04 | 2023-01-04 07:58 | Hi Precision Diagnostics | N. Amoranto St. Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22827 | CONSORTIUM | pcr@hi-precision.com.ph | 27 | ASSAY: Xpert HPV, COUNT: 176 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4002 |  | GXIV | 835601 | Unit for Preventive Maintenance. | Performed Preventive Maintenance. Cleaned docking bay, plunger, icore and filter. | Preventive Maintenance | Unitmis operational. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5486437659819568277/5486437659819568277_signature_28.png | Rheizza Daphne Salvilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5486437659819568277/5486437659819568277_signature_31.png
| 2023-01-04 | 2023-01-04 15:23 | Cotabato Sanitarium | Sultan kudarat | BARMM | Maguindanao | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22826 | PBSP | normiaando.tahir@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 463 ASSAY: Xpert MTB/RIF Ultra, COUNT: 215 | 2023-01-04 | Gx -4 | 815277 | Modules B1 and B2 for Service Required | Replace module B1 SN:695616 with new Module SN: 801637 Replace module B2 SN: 695619 with new Module SN: 657296 Genexpert machine Preventive maintenance | Replace module B1 SN:695616 with new Module SN: 801637 Replace module B2 SN: 695619 with new Module SN: 657296 Genexpert machine Preventive maintenance | Gx machine has 4 operational modules |  | 900-0491RS | N/A | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5486284266904509174/5486284266904509174_signature_28.png | Normia A. Tahir | https://www.jotform.com/uploads/macarewecare/222927978372067/5486284266904509174/5486284266904509174_signature_31.png
| 2023-01-04 | 2023-01-04 13:23 | Tarlac City Health Center 3 | San Miguel Tarlac City | 3 | Tarlac | Laboratory | NO | INSTALLATION REPAIR | CEPHEID | [5 empty] | 22825 | DOH | daganlea572@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 247 ASSAY: Xpert MTB/RIF Ultra, COUNT: 141 | 2022-12-20 | GXIV-D | 831697 | Defective monitor | Performed preventive maintenance Replaced defective monitor Installed HIV Viral Load assay Gathered reports Performed functionality test | Imstallation and repair | Unit is operational | Monitor serial number QSDN8HA087212 | Monitor | 9409 | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5486251642613396285/5486251642613396285_signature_28.png | Lea Dagan | https://www.jotform.com/uploads/macarewecare/222927978372067/5486251642613396285/5486251642613396285_signature_31.png
| 2023-01-04 | 2023-01-04 14:57 | PERPETUAL SUCCOUR HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | HAEMONETICS | [5 empty] | 22824 | COMMERCIAL | rodaflores1038@yahoo.com.ph | 30 |  |  | MCS+ | 11H307 | #NAME? | #NAME? | #NAME? | #NAME? |  |  |  | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5486248378813006614/5486248378813006614_signature_28.png | DEBBIE JOYCE A. BAGA,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5486248378813006614/5486248378813006614_signature_31.png
| 2023-01-04 | 2023-01-04 14:26 | The Medical City - Clark | Clark Global City Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22823 | COMMERCIAL | tmcclarkbiomed@themedicalcity.com | 25 | ASSAY: Xpert HPV, COUNT: 8 ASSAY: Xpert MTB/RIF, COUNT: 33 ASSAY: Xpert MTB/RIF Ultra, COUNT: 130 ASSAY: Xpert CT/NG, COUNT: 4 | 2023-01-04 | Genexpert IV | 839235 | N/a | N/a | Performed Maintenance (Cleaned Icore, Cleaned Plunger Rod, Cleaned Filter, Cleaned Cartridge Bay) Gathered Reports Insatalled HCV assay Performed Functionality Test | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5486232064913988496/5486232064913988496_signature_28.png | Vince Jilmer C. Nambio | https://www.jotform.com/uploads/macarewecare/222927978372067/5486232064913988496/5486232064913988496_signature_31.png
| 2023-01-04 | 2023-01-04 13:38 | Bislig District Hospital | National Higway, Poblacion, Bislig City | 13 | Surigao del Sur | RTDL | NO | REPAIR | CEPHEID | [5 empty] | 22822 | PBSP | gnovero@yahoo.com | 23.69 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 360 |  | GX-4 | 805280 | None | None | Replacing Module B2 w/ SN: 743259 to SN: 765175 Replacing Module B3 w/ SN: 647942 to SN: 723764 | All modules are now operational |  |  |  | Faizar S. Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5486220444026072208/5486220444026072208_signature_28.png | Georgina Novero | https://www.jotform.com/uploads/macarewecare/222927978372067/5486220444026072208/5486220444026072208_signature_31.png
| 2023-01-03 | 2022-12-22 13:45 | Cebu City Medical Center | Cebu City | 7 | Cebu | Microbiology | NO | MAINTENANCE | CEPHEID | [5 empty] | 22821 | DOH | ccmclaboratory01@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF, COUNT: 786 ASSAY: Xpert MTB/RIF Ultra, COUNT: 641 |  | GXIV-4D | 838561 | Due for PM | Preventive Maintenance | NA | A1 & A3 are functional. Waiting for module allocation for A2 & A4 |  |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5485455748815145840/5485455748815145840_signature_28.png | Fatimah maraviles | https://www.jotform.com/uploads/macarewecare/222927978372067/5485455748815145840/5485455748815145840_signature_31.png
| 2023-01-03 | 2023-01-03 09:14 | ZAMBOANGA DEL NORTE MEDICAL CENTER | DIPOLOG CITY | 9 | Zamboanga del Norte | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22820 | PBSP | ferranbryder@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 240 |  | GX-IV | 815285 | -FOR B1 MODULE REPLACEMENT -MODULE B3 STUCK CARTRIDGE -MODULE B4 ERROR 1001 | -REPLACE DEFECTIVE MODULE B1 SN: 695735 WITH NEW SN:635857 -STUCK CARTRIDGE ISSUE RESOLVED | #NAME? | #NAME? | #NAME? |  |  | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5485179598818889747/5485179598818889747_signature_28.png | THELMA Y. MORGIA WABE; RMT; RN | https://www.jotform.com/uploads/macarewecare/222927978372067/5485179598818889747/5485179598818889747_signature_31.png
| 2023-01-03 | 2023-01-03 08:51 | DIPOLOG CITY HEALTH OFFICE | DIPOLOG CITY | 9 | Zamboanga del Norte | TB DOTS | NO | REPAIR | CEPHEID | [5 empty] | 22819 | PBSP | jz.ebron@yahoo.com.ph | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 628 ASSAY: Xpert HIV-1 Viral Load, COUNT: 41 |  | GX-II | 834826 | FOR MODULE B1 REPLACEMENT | -PERFORM MODULE B1 REPLACEMENT DEFECTIVE SN: 787167 NEW SN: 611328 | #NAME? | #NAME? |  |  |  | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5485175408811719255/5485175408811719255_signature_28.png | JAMES ZEBEDEE C. EBRON | https://www.jotform.com/uploads/macarewecare/222927978372067/5485175408811719255/5485175408811719255_signature_31.png
| 2022-12-29 | 2022-12-29 09:08 | Pamplona rhu | Pamplona camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22818 | PBSP | francislozano226@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 381 ASSAY: Xpert MTB/RIF Ultra, COUNT: 626 | 2022-12-29 | Gx4 | 812774 | For module replacment a3 758215 - 637299 a4 731314 - 711531 | Performed module replacement ,cleaned unit,filter,icore,cartridge bay,plunger,gathered reports,test run unit | Module replacement, maintenance | Operational |  | 900gx | 0 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5480857833322521619/5480857833322521619_signature_28.png | Francis S. Lozano | https://www.jotform.com/uploads/macarewecare/222927978372067/5480857833322521619/5480857833322521619_signature_31.png
| 2022-12-29 | 2022-12-29 08:37 | IBA Rural Health Unit | Iba Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22817 | PBSP | garadoshanna@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF, COUNT: 745 ASSAY: Xpert MTB/RIF Ultra, COUNT: 942 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 | 2022-12-12 | GXIV-D | 805070 | Module replacement for defective module B1 | Performed Preventive maintenance Generate Pre IQR Performed module replacement Generate post IQR and gathered reports Performed functionality test | Preventive maintenance and repair | Unit us operational | Defective module SN:713337 | Genexpert module New SN:654694 | 37723 | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5480838573225980329/5480838573225980329_signature_28.png | Hanna C. Garados | https://www.jotform.com/uploads/macarewecare/222927978372067/5480838573225980329/5480838573225980329_signature_31.png
| 2022-12-28 | 2022-12-28 16:11 | Bongabon District Hospital | Bongabon, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22816 | DOH | faye_kharizze@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 170 ASSAY: Xpert MTB/RIF Ultra, COUNT: 129 | 2022-12-28 | Genexpert IV | 841562 | Module 4 | Performed Xpertcheck on A4 module | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | Xpertcheck A4 - retry | Xpertcheck 1pc | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5480326046319697976/5480326046319697976_signature_28.png | KHARIZZE FAJARDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5480326046319697976/5480326046319697976_signature_31.png
| 2022-12-28 | 2022-12-28 16:35 | Bicol medical center | Naga city | 5 | Camarines Sur | Bacteriology | NO | REPAIR | CEPHEID | [5 empty] | 22815 | MACARE-HBDC | maryelainetolidanes@gmail.com | 22 | ASSAY: Xpert HPV, COUNT: 1771 ASSAY: Xpert MTB/RIF Ultra, COUNT: 34 ASSAY: Xpert SARS-CoV-2, COUNT: 192 | 2022-12-27 | Gx4 | 110008458 | For module replacement b2 220001508 - 886130 | Performed maintenance,cleaned filter,plunger,cartridge bay,gathered reports,perforemed module replacement | Module replacement,maintenance | Operational |  | 900gx | G5003 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5480282999587543686/5480282999587543686_signature_28.png | MARY ELAINE TOLIDANES | https://www.jotform.com/uploads/macarewecare/222927978372067/5480282999587543686/5480282999587543686_signature_31.png
| 2022-12-28 | 2022-12-28 14:15 | Mauban Municipal Health Office Molecular Laboratory | Signal Street, Mauban, Quezon | 4A | Quezon | Molecular Laboratory | NO | INSTALLATION | CEPHEID | [5 empty] | 22814 | MACARE-HBDC | maubanmhomolecular.laboratory@gmail.com | 23.9 | ASSAY: Others, COUNT: 0 | 2022-12-27 | GeneXpert GXIV-D-10C | 110015847 | None | None | Installed GeneXpert unit Configured accounts Setup the software Conducted basic operators training Perfirmed functionality test | Unit is operational with 4 modules |  | GeneXpert GXIV-D-10C Set with UPS and Printer | SI 94545 and DR 9430 | Carlos Leniel C. Origenes | https://www.jotform.com/uploads/macarewecare/222927978372067/5480179841224643501/5480179841224643501_signature_28.png | Lady Angel R. Astoveza | https://www.jotform.com/uploads/macarewecare/222927978372067/5480179841224643501/5480179841224643501_signature_31.png
| 2022-12-28 | 2022-12-28 14:11 | Lambayong Rural Health Unit | Poblacion, Lambayong Sultan Kudarat | 12 | Sultan Kudarat | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22813 | PBSP | jeanalonzo73@gmail.com | 29.7 | ASSAY: Xpert MTB/RIF, COUNT: 160 ASSAY: Xpert MTB/RIF Ultra, COUNT: 278 | 2022-11-28 | Gx IV - 4d | 838159 | For Preventive Maintenance | Performed Preventive Maintenance | Clean Surface Area, Cartridge Bay, Fan Filters, Plunger Rod, and Icore Lens Generate IQR, Assay Stat, System Log Archived Test Results | Unit is fully Operational | N/A | N/A | N/a | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5480176076226953734/5480176076226953734_signature_28.png | Jean K. Alonzo RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5480176076226953734/5480176076226953734_signature_31.png
| 2022-12-28 | 2022-12-28 14:01 | Tinambac rhu | Tinambac camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22812 | PBSP | maelagarde27@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 411 ASSAY: Xpert MTB/RIF Ultra, COUNT: 804 | 2022-12-28 | Gx4 | 808285 | For module replacement b3 638250 - 699314 | Performed maintenance cleaned filter,cartridge bay,plunger,icore,gathered reports,performed module replacement | Module replacement,maintenance | Operational |  | 900 gx | 0 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5480170410021025956/5480170410021025956_signature_28.png | Mae L. Aguinillo | https://www.jotform.com/uploads/macarewecare/222927978372067/5480170410021025956/5480170410021025956_signature_31.png
| 2022-12-28 | 2022-12-28 12:23 | Mauban Rural Health Unit | Mauban, Quezon | 4A | Quezon | TB DOTS | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22811 | PBSP | rhumauban.laboratory@gmail.com | 26.5 | ASSAY: Xpert HPV, COUNT: 159 ASSAY: Xpert MTB/RIF, COUNT: 373 ASSAY: Xpert BCR-ABL Ultra, COUNT: 950 | 2022-12-28 | GXIV | 814375 | Modules B1 and B2 for replacement | Replaced modules B1 and B2. | Performed module repkacement. | All 4 modules operatioanl. |  | New modules SN: 724421, 719993 | N/A | Harold B. TOCO | https://www.jotform.com/uploads/macarewecare/222927978372067/5480113790421666336/5480113790421666336_signature_28.png | Rhandy Buraga | https://www.jotform.com/uploads/macarewecare/222927978372067/5480113790421666336/5480113790421666336_signature_31.png
| 2022-12-28 | 2022-12-28 11:33 | Masinloc RHU | Masinloc Zambales | 3 | Zambales | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22810 | DOH | mariapaz.ferrer08@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF, COUNT: 614 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1147 |  | GXIV-D | 831608 | Due For Maintenance | #NAME? | Preventive maintenance | Unit is operational except module A3 HW Failed already reported |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5480080699257869983/5480080699257869983_signature_28.png | Maria Paz Ferrer | https://www.jotform.com/uploads/macarewecare/222927978372067/5480080699257869983/5480080699257869983_signature_31.png
| 2022-12-28 | 2022-12-28 11:28 | Presentacion Rhu | Presentacion camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22809 | PBSP | tabo.juliet@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 196 ASSAY: Xpert MTB/RIF Ultra, COUNT: 408 | 2022-12-28 | Gx2 | 834808 | For module replacement b1 787133 -643219 b2 787135 -693691 | Performed module replacement, cleaned filter plunger,cartridge bay | Module replacement | Operational |  | 900 gx module | 0 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5480078780023570480/5480078780023570480_signature_28.png | Julieta G Tabo | https://www.jotform.com/uploads/macarewecare/222927978372067/5480078780023570480/5480078780023570480_signature_31.png
| 2022-12-28 | 2022-12-28 08:23 | Allen District Hospital | Allen Northern Samar | 8 | Northern Samar | Laboratory | YES | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22808 | PBSP | aracelienzon@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 941 | 2022-12-28 | GXIV - 4D | 815276 | CPU error booting operating system | #NAME? | #NAME? | Unit is in good condition |  | CR2032 | Energizer Battery 3 volt | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5479978094637230292/5479978094637230292_signature_28.png | Araceli Enzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5479978094637230292/5479978094637230292_signature_31.png
| 2022-12-27 | 2022-12-20 10:31 | Lung center of the philippines | Quezon ave,quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | REPAIR | CEPHEID | [5 empty] | 22807 | PBSP | madzrmt@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1653 ASSAY: Xpert SARS-CoV-2, COUNT: 1744 | 2022-12-19 | GXXVI | 801080 | For module replacement D4 696340 - 870126 | Performed module replacement,cleaned plunger,cartridge bay,icore,filter,gathered reports | Module replacement | C2 disabled, 15 module operational |  | Module | G4977 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5479364872948612311/5479364872948612311_signature_28.png | Madeline Gonzalez | https://www.jotform.com/uploads/macarewecare/222927978372067/5479364872948612311/5479364872948612311_signature_31.png
| 2022-12-27 | 2022-12-20 15:36 | Philippine heart center | East ave diliman q.c | NCR | NCR 2nd - Quezon City | Molecar | NO | REPAIR | CEPHEID | [5 empty] | 22806 | PBSP | phcmolecularlab@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 264 ASSAY: Others, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 2789 | 2022-12-19 | GX4 | 834443 | For module replacement b1 643530 - 603916 | Performed module replacement,maintenance,gathered reports | Module replacement | B2 reported 3 module operational |  | 900gx | G4976 | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5479364702949273427/5479364702949273427_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5479364702949273427/5479364702949273427_signature_31.png
| 2022-12-27 | 2022-12-21 15:15 | Western pangasinan district hospital | Alaminos pangasinan | 1 | Pangasinan | Lab | NO | REPAIR | CEPHEID | [5 empty] | 22805 | PBSP | wpdhlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 855 ASSAY: Xpert MTB/RIF Ultra, COUNT: 682 | 2022-12-21 | GX4 | 807598 | For module replacement a2 662108 - 884568 a4 662109 - 687202 | Performed module replacement,maintenance cleaned filter,plunger,icore,cartridge bay | Module replacement,maintenance | Operational |  | 900gx | 0 | Noel floresca | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5479364282941094847/5479364282941094847_signature_28.png | Alysa Kere. C. Apalis | https://www.jotform.com/uploads/https:/www.jotform.com/uploads/macarewecare/222927978372067/5479364282941094847/5479364282941094847_signature_31.png
| 2022-12-27 | 2022-12-27 15:38 | Basud municipal health office | Basud, camarines norte | 5 | Camarines Norte | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22804 | PBSP | gbechano@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 432 ASSAY: Xpert MTB/RIF Ultra, COUNT: 879 | 2022-12-27 | Gx4 | 807600 | For module replacement A2 662500 - 680631 | Performed module replacement,performed maintenance,cleaned filter,cartridge bay,icore,plunger,gathered reports,test run unit | Module replacement | 3 module operational, module 4 disabled encountered error 4004 |  | 900 gx | 0 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5479363322948117031/5479363322948117031_signature_28.png | GRACE B. ECHANO | https://www.jotform.com/uploads/macarewecare/222927978372067/5479363322948117031/5479363322948117031_signature_31.png
| 2022-12-27 | 2022-12-27 15:27 | Palayan CHO | Palayan City, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22803 | DOH | riazarizzamae@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 352 ASSAY: Xpert MTB/RIF Ultra, COUNT: 566 | 2022-12-27 | Genexpert IV | 831637 | N/a | N/a | Performed Maintenance Gathered Report Disabled Module 4 Performed Functionality Test | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5479356195818256677/5479356195818256677_signature_28.png | Rizza Mae A. Riaza | https://www.jotform.com/uploads/macarewecare/222927978372067/5479356195818256677/5479356195818256677_signature_31.png
| 2022-12-27 | 2022-12-27 14:38 | DOH REGIONAL OFFICE 2 | Maharlika Highway, Tuguegarao City, Cagayan | 2 | Cagayan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22802 | PBSP | cvchdcovidlab@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 432 ASSAY: Xpert MTB/RIF Ultra, COUNT: 293 ASSAY: Xpert SARS-CoV-2, COUNT: 2742 |  | GX4D | 834434 | Unit is due for module replacement | Module Replacement | Performed Module Replacement B1/776202 into 805485 B3/779575 into 616313 | Unit is operational |  |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5479328503641583685/5479328503641583685_signature_28.png | Queenie Cantil | https://www.jotform.com/uploads/macarewecare/222927978372067/5479328503641583685/5479328503641583685_signature_31.png
| 2022-12-27 | 2022-12-23 12:36 | Matanao Rural Health Unit | Poblacion, Matanao, Davao del Sur | 11 | Davao del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22801 | DOH | gmdomingso13@gmai.com | 28.7 | ASSAY: Xpert MTB/RIF, COUNT: 141 ASSAY: Xpert MTB/RIF Ultra, COUNT: 240 | 2022-12-23 | Gx-IV 4D | 838456 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Fan Filters, Plunger Rod, Icore Lenses Generate Reports; IQR, System Log, and Assay Stat Archived Test Results | Unit is fully operational | High temp | N/A | N/A | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5479319382211641001/5479319382211641001_signature_28.png | Dominador S. Castro III | https://www.jotform.com/uploads/macarewecare/222927978372067/5479319382211641001/5479319382211641001_signature_31.png
| 2022-12-27 | 2022-12-27 13:20 | CEBU DOCTORS UNIVERSITY HOSPITAL | CEBU CITY | 7 | Cebu | BLOOD BANK | NO | REPAIR MAINTENANCE | HAEMONETICS | [5 empty] | 22800 | COMMERCIAL | loveolive798@gmail.com | 20 |  | 2022-12-23 | MCS+ | 17H639SPG | LIQUID/SPILL DETECTOR ERROR ISSUE | #NAME? | -CHECK CENTRIFUGE SPILL DETECTOR FUNCTIONALITY AND DIAGNOSTICS, PASSED -COLLECT LAB HUMIDITY DATA USING DATA LOGGER, 65.8% | #NAME? |  | N/A | N/A | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5479298728812514990/5479298728812514990_signature_28.png | ERICKA JANE CORRO | https://www.jotform.com/uploads/macarewecare/222927978372067/5479298728812514990/5479298728812514990_signature_31.png
| 2022-12-27 | 2022-12-27 13:45 | President Ramon Magsaysay Memorial Hospital | Iba Zambales | 3 | Zambales | Molecular Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22799 | PBSP | prmmhmoleclab@gmail.com | 29.75 | ASSAY: Xpert SARS-CoV-2, COUNT: 4910 | 2022-12-27 | GXIV-D | 834622 | Module replacement for moduke B1 | Performed Preventive maintenance Generate Pre-IQR Replaced defective module B4 Generate post-IQR and generate report file Performed functionality test | Preventive maintenance and repair | Unit us operational | Defective module SN:784809 | Genexpert module SN:694489 | N/a | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5479297430711503562/5479297430711503562_signature_28.png | Angelina Cuaton | https://www.jotform.com/uploads/macarewecare/222927978372067/5479297430711503562/5479297430711503562_signature_31.png
| 2022-12-27 | 2022-12-27 13:10 | Eron Dumagieng | Maharlika Highway, Tuguegarao City, Cagayan | 2 | Cagayan | Molecular Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22798 | MACARE-HBDC | cvmc.mollab0803@gmail.com | 25.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 359 | 2022-12-27 | GX4D | 110008457 | Unit is due for module replacement | Module Replacement | Performed Module Replacement B2/ 220001502 into 687144 Performed Maintenance | Unit is operational |  | Genexpert Module 6C 687144 | G-5002 | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5479278623641908437/5479278623641908437_signature_28.png | Ruel Lamorena | https://www.jotform.com/uploads/macarewecare/222927978372067/5479278623641908437/5479278623641908437_signature_31.png
| 2022-12-27 | 2022-12-27 11:30 | Calauag municipal hospital | Calauag quezon | 4A | Quezon | Laboratory | NO | CALIBRATION | CEPHEID | [5 empty] | 22797 | DOH | jamila17carmel@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 170 ASSAY: Xpert MTB/RIF Ultra, COUNT: 453 | 2022-12-27 | Gx4 | 831910 | A3 module continous 4017 error | Check unit, performed re xpert check,gathered reports | Re xpert check | 3 module operational | A3 failed re xpertcheck | Xpertcheck check | 0 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5479219497317167762/5479219497317167762_signature_28.png | Jamila Carmel Pedrezuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5479219497317167762/5479219497317167762_signature_31.png
| 2022-12-27 | 2022-12-27 10:57 | Subic Rural Health Unit | Subic Zambales | 3 | Zambales | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22796 | PBSP | salonga.jrs97@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 1457 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1144 ASSAY: Xpert HIV-1 Viral Load, COUNT: 35 |  | GXIV-D | 812788 | Module replacement for defective module A4 | Performed Preventive maintenance Generate pre IQR Replaced defective module A4 Gathered post IQR and generate reports Performed functionality test | Preventive maintenance and repair | Unit is operational |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5479209700713042110/5479209700713042110_signature_28.png | Jessa Rowe Salonga | https://www.jotform.com/uploads/macarewecare/222927978372067/5479209700713042110/5479209700713042110_signature_31.png
| 2022-12-27 | 2022-12-27 08:49 | Batangas Medical Center | Batangas City | 4A | Batangas | Covid Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22795 | PBSP | batmc.tbculturelab@gmail.com | 25 | ASSAY: Others, COUNT: New CPU, no data | 2022-12-19 | GXIV | 834627 | Unit for PM. For installation of new CPU. For checking of UPS. | Performed PM of GeneXpert. Performed installation of CPU. Installed assays and generated reports. Performed troubleshooting of UPS. Checked monitor. | PM, installation and troubleshootiny | Unit is operational. Module B2 (784301), H/W Fail. 3 modules operational. UPS (Powercom SN: 40224771704) is defective. Monitor (HP Prodisplay P19A SN: CNK91915SX) has deadpixels and for replacement. | Pulled out defective CPU (SN: CPH1119F03592) and backup Laptop | Lenovo Configured Desktop (SN: CPH1221F29351) in replacement for defecfive CPU (SN: CPH1119F03592) | G4982 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5479151135961504907/5479151135961504907_signature_28.png | Patricia Louisse A. Baltazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5479151135961504907/5479151135961504907_signature_31.png
| 2022-12-27 | 2022-12-27 09:04 | Kalinga Provincial Hospital | Bolaoan, Tabuk City, Kalinga | CAR | Kalinga | TB Dots Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22794 | DOH | tuguinaymicahjanelle@gmail.com | 22.7 | ASSAY: Xpert MTB/RIF, COUNT: 342 ASSAY: Xpert MTB/RIF Ultra, COUNT: 713 |  | GX4D | 838466 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Filter Cleaning Performed Plunger Rod Cleaning | Unit is operational |  |  |  | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5479126043425802417/5479126043425802417_signature_28.png | Micah Janelle B. Tuguinay | https://www.jotform.com/uploads/macarewecare/222927978372067/5479126043425802417/5479126043425802417_signature_31.png
| 2022-12-27 | 2022-12-27 08:56 | Kalinga Provincial Hospital | Bolaoan, Tabuk City, Kalinga | CAR | Kalinga | TB Dots Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22793 | PBSP | whinzie18@gmail.com | 24.3 | ASSAY: Xpert MTB/RIF, COUNT: 399 ASSAY: Xpert MTB/RIF Ultra, COUNT: 667 |  | GX4D | 807622 | Unit is due for module replacement | Module Replacement | Performed Module Replacement A1/678260 into A1/794446 Performed surfaces cleaning Performed filter cleaning Performed Plunger Rod Cleaning | Unit is Operational |  |  |  | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5479121863426231395/5479121863426231395_signature_28.png | Wenona P.Baculi | https://www.jotform.com/uploads/macarewecare/222927978372067/5479121863426231395/5479121863426231395_signature_31.png
| 2022-12-26 | 2022-12-26 13:49 | Rizal Medical Center | Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22792 | MACARE-HBDC | molecularlab@rmc.doh.gov.ph | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 ASSAY: Xpert SARS-CoV-2, COUNT: 16101 ASSAY: Others, COUNT: 3 |  | GX XVI-D | 838627 | For maintenance | Performed maintenance | Performed maintenance Cleaned: plunger rod, icore, cartridge bays Generated reports | Unit is operational. | Module B4/860923 disabled. For XpertCheck |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5478444513519037555/5478444513519037555_signature_28.png | Jamaine A. Lomibao | https://www.jotform.com/uploads/macarewecare/222927978372067/5478444513519037555/5478444513519037555_signature_31.png
| 2022-12-26 | 2022-12-26 13:53 | Rizal Medical Center | Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22791 | PBSP | molecularlab@rmc.doh.gov.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 324 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2358 ASSAY: Xpert SARS-CoV-2, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 19 |  | GX-IV-D | 814374 | For maintenance | Performed maintenance | Performed maintenance Cleaned: plunger rod, icore, cartridge bays Generated reports | Unit is operational. | Modules under observation due to error 2008. |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5478444093515339792/5478444093515339792_signature_28.png | Jamaine A. Lomibao | https://www.jotform.com/uploads/macarewecare/222927978372067/5478444093515339792/5478444093515339792_signature_31.png
| 2022-12-26 | 2022-12-26 10:57 | Department of Health - Treatment and Rehabilitation Center Bicutan | Bicutan, Taguig | NCR | NCR 4th - Taguig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22790 | PBSP | dohtrcbicutan.lab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 519 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1265 | 2022-12-19 | GeneXpert GX-IV-D | 807452 | Module B1 for replacement | Replaced module B1 | Performed maintenance Replaced defective module Generated reports | Unit is operational. | Module B3/801912 not operational. | Old: B1/661843 New: B1/644488 | G 4975 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5478331454016780765/5478331454016780765_signature_28.png | April Chakee A. Javier | https://www.jotform.com/uploads/macarewecare/222927978372067/5478331454016780765/5478331454016780765_signature_31.png
| 2022-12-23 | 2022-12-23 17:31 | Philippine Heart Center | East Ave., Quezon City | NCR | NCR 2nd - Quezon City | Bloof Bank | NO | REPAIR | HAEMONETICS | [5 empty] | 22789 | COMMERCIAL | baluyotjeric@gmail.comcom | 28 |  | 2022-12-21 | MCS+ | 01M201 | Defective CPU | CPU replacement | Replaced CPU board. Performed calibration..performed functionality test and simulation run. | Unit is operational. | None | CPU Board | DR # 9420 | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5475981954221794757/5475981954221794757_signature_28.png | Jeric Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5475981954221794757/5475981954221794757_signature_31.png
| 2022-12-23 | 2022-12-23 13:57 | Bataan General Hospital | Balanga Bataan | 3 | Bataan | Bacteriology Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22788 | PBSP | melodyjoy.flores16@yahoo.com | 27.65 | ASSAY: Xpert MTB/RIF, COUNT: 1279 ASSAY: Xpert MTB/RIF Ultra, COUNT: 997 ASSAY: Xpert HIV-1 Viral Load, COUNT: 70 ASSAY: Xpert HCV VL, COUNT: 62 | 2022-12-20 | GXIV-D | 805124 | Module replacement for A1 and A2 | Performed Preventive maintenance Gathered Pre IQR Replaced defective modules A1 and A2 Gathered Post IQR Generate reports Performed functionality test | Preventive maintenance and repair | Unit is operational |  | Genexpert module Defective A1:698583;New Module:724871 A2:603170;New Module:210021632 | G4987 | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5475851998944082099/5475851998944082099_signature_28.png | Melody Joy C. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5475851998944082099/5475851998944082099_signature_31.png
| 2022-12-23 | 2022-12-23 14:05 | Makati medical center | Salcedo vill makati | NCR | NCR 4th - Makati | Stem cell | YES | MAINTENANCE | OTHERS |  |  |  |  | Miltenyi biotech | 22787 | COMMERCIAL | biomedsupervisor@makatimed.net.ph | 0 |  |  | Clinimacs | 797 | For maintenance | Performed maintenance,check unit, rollers,sensor,display,programs parameter,performed functionality test,visual check | Maintenance | Operational |  |  |  | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5475851569715917162/5475851569715917162_signature_28.png | Mark lenard latay | https://www.jotform.com/uploads/macarewecare/222927978372067/5475851569715917162/5475851569715917162_signature_31.png
| 2022-12-23 | 2022-12-23 11:47 | Dingalan Community Hospital | Dingalan, Aurora | 3 | Aurora | Laboratroy | NO | MAINTENANCE | CEPHEID | [5 empty] | 22786 | DOH | bnjlvrz9@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 214 ASSAY: Xpert MTB/RIF Ultra, COUNT: 244 | 2022-12-23 | Genexpert IV | 831621 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit Is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5475767466111717223/5475767466111717223_signature_28.png | Benjo Alvarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5475767466111717223/5475767466111717223_signature_31.png
| 2022-12-23 | 2022-12-23 11:22 | Camarin Health Center | Camarin, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22785 | PBSP | cindy.carpo06@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 51 ASSAY: Xpert MTB/RIF Ultra, COUNT: 481 | 2022-12-19 | GXIV | 815395 | Modules B1 (686345), B3 (635192) for replacement | Replaced B1, B3 | #NAME? | Unit is operational. 3 modules are working. | Module B2 for reporting/follow-up. | GeneXpert Module SN: 656113, 711296 | G 4985 | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5475766580816280591/5475766580816280591_signature_28.png | Cindy D. Carpo | https://www.jotform.com/uploads/macarewecare/222927978372067/5475766580816280591/5475766580816280591_signature_31.png
| 2022-12-23 | 2022-12-23 11:06 | Tiaong Rural Health Unit | Do V. Robles St. Poblacion 3, Tiaong, Quezon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22784 | PBSP | ilabrhutiaong@gmail.com | 22.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 652 ASSAY: Xpert MTB/RIF, COUNT: 166 | 2022-12-23 | GXIV | 838899 | Module B1 for module replacement. For maintenance | Performed maintenance and cleaning..Replaced module B1 | Replaced module B1. Performed finctionality test.. | All 4 modules operational. | None | New module sn: 788784 | None | Harold B. Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5475746180228899427/5475746180228899427_signature_28.png | Maria Teresa Tolentino | https://www.jotform.com/uploads/macarewecare/222927978372067/5475746180228899427/5475746180228899427_signature_31.png
| 2022-12-23 | 2022-12-23 10:25 | Bagong Silang Ph. 1 Health Center | Bagong Silang, Caloocan | NCR | NCR 3rd - Caloocan | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | [5 empty] | 22783 | PBSP | alcaydeirene07@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 274 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1128 |  | GXIV | 838856 | Unit for installation | Installed unit | #NAME? | Unit is operational. All modules are working. | Unit was temporarily from Camarin Health Center. |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5475717990819110127/5475717990819110127_signature_28.png | Irene D. Alcayde | https://www.jotform.com/uploads/macarewecare/222927978372067/5475717990819110127/5475717990819110127_signature_31.png
| 2022-12-23 | 2022-12-23 09:58 | Carigara District Hospital | Carigara Leyte | 8 | Leyte | TB-DOTS | YES | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22782 | PBSP | donitarose.supolmo@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 753 | 2022-12-23 | GXIV-4D | 807619 | For module replacement (A1,A2) | PM and Module replacement | Module replacement A1 def. SN 662507 - new SN: 775309 A2 def. SN 662506 - new SN: 868467 - Preventive Maintenance -Generate reports | Module A3 and A4 received multiple temperature error | A3 and A4 waiting for module replacement | 900-0491RS | GeneXpert 6 color module SN: 775309 , 868467 | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5475709357137505639/5475709357137505639_signature_28.png | Donita Rose Supolmo | https://www.jotform.com/uploads/macarewecare/222927978372067/5475709357137505639/5475709357137505639_signature_31.png
| 2022-12-23 | 2022-12-21 11:53 | Imus City Health Office I | Velarde subd. | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22780 | PBSP | armndlasq@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 658 ASSAY: Xpert HIV-1 Viral Load, COUNT: 136 | 2022-12-19 | GXIV-D | 814168 | Due for module replacement | Replaced module | Performed Maintenance Replaced Defective module | All modues are operational |  | 900-0491RS | G 4971 | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689946732671283/5475689946732671283_signature_28.png | Armand G. Lasquete | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689946732671283/5475689946732671283_signature_31.png
| 2022-12-23 | 2022-12-22 10:45 | Tanza Rural Health Unit | Tanza municipal compound | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22781 | PBSP | gutierrezrhian@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 285 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1438 | 2022-12-19 | GXIV-D | 812835 | Due for module replacement | Module replacement | Performed Maintenance Module replacement on modules 1 and 4 Generate reporta | 3 modules are operational |  | Gx module 6-C | G 4970 | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689946739514153/5475689946739514153_signature_28.png | MARY ANN C. GUTIERREZ | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689946739514153/5475689946739514153_signature_31.png
| 2022-12-23 | 2022-12-23 09:25 | General Emilio Aguinaldo Memorial Hospital | Trece-Indang Road, Trece Martires | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22777 | PBSP | djaemiel@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 465 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1495 ASSAY: Xpert HIV-1 Viral Load, COUNT: 37 | 2022-12-19 | GXIV-L | 802860 | Due for Module Replacement | Replaced module | Performed Maintenance Module Replacement on module 4 Generate reports | 3 modules are operational, 1 module defective |  | Genexpert module 6-C | G 4967 | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689936731813731/5475689936731813731_signature_28.png | JAEMIEL JUVIRENS A. DIMARANAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689936731813731/5475689936731813731_signature_31.png
| 2022-12-23 | 2022-12-20 15:17 | Putatan Health Center | Putatan | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22778 | PBSP | apmolera@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 21 ASSAY: Xpert MTB/RIF, COUNT: 281 ASSAY: Xpert MTB/RIF Ultra, COUNT: 910 |  | GXIV-D | 834690 | For module replacement | Replaced defective module | Performed Maintenance Replaced module 1 Generate reports | All modules are operational |  |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689936738236427/5475689936738236427_signature_28.png | Alelli P. Molera | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689936738236427/5475689936738236427_signature_31.png
| 2022-12-23 | 2022-12-21 13:37 | Dasmariñas City Health Office I | D mangubat | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22779 | PBSP | cho1lab@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 224 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1458 ASSAY: Xpert HIV-1 Viral Load, COUNT: 24 | 2022-12-19 | GXIV-D | 838893 | Due for module replacement | Replaced module 4 | Performed maintenance Replaced defective module Generate reports | All modules are operational |  | 900-0491RS | G 4969 | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689946732469738/5475689946732469738_signature_28.png | Elma R.Belostrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5475689946732469738/5475689946732469738_signature_31.png
| 2022-12-22 | 2022-12-22 15:12 | Sorsogon Provincial Hospital | Sorsogon City | 5 | Sorsogon | TB DOTS | NO | REPAIR CALIBRATION | CEPHEID | [5 empty] | 22776 | PBSP | wizard682000@yahoo.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 77 ASSAY: Xpert MTB/RIF Ultra, COUNT: 726 | 2022-12-12 | GXIV | 835213 | Module B2 for replacement | Replaced Module B2 | Performed preventive maintenance, replaced module B2 (789033 > 888907), performed XpertCheck | Unit is operational, XpertCheck passed |  | GeneXpert Module 6-color S/N:888907 | Monarch DR 37717 | Patrick Los Baños | https://www.jotform.com/uploads/macarewecare/222927978372067/5475028520924095466/5475028520924095466_signature_28.png | Reynaldo B Gonzalgo | https://www.jotform.com/uploads/macarewecare/222927978372067/5475028520924095466/5475028520924095466_signature_31.png
| 2022-12-22 | 2022-12-22 14:02 | Mariano Marcos Memorial Hospital and Medical Center | Philippine National Road, MMHMC compound, Batac City, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22775 | DOH | mmmhmc.molecularlab@gmail.com | 24.3 | ASSAY: Xpert HPV, COUNT: 506 ASSAY: Xpert MTB/RIF, COUNT: 257 ASSAY: Xpert MTB/RIF Ultra, COUNT: 700 ASSAY: Xpert SARS-CoV-2, COUNT: 3026 |  | GX4D | 831908 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Plunger Rod Cleaning Performed Filter Cleaning | Unit is operational |  |  |  | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5474990488817038661/5474990488817038661_signature_28.png | J-Bie Alvarado | https://www.jotform.com/uploads/macarewecare/222927978372067/5474990488817038661/5474990488817038661_signature_31.png
| 2022-12-22 | 2022-12-22 13:53 | Mariano Marcos Memorial Hospital and Medical Center | Philippine National Rd. MMHMC Compound , Batac City, Ilocos Norte | 1 | Ilocos Norte | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22774 | PBSP | mmmhmc.molecularlab@gmail.com | 23.4 | ASSAY: Xpert HPV, COUNT: 132 ASSAY: Xpert MTB/RIF, COUNT: 92 ASSAY: Xpert MTB/RIF Ultra, COUNT: 154 ASSAY: Xpert SARS-CoV-2, COUNT: 1426 |  | GX4D | 838168 | Unit is due for module Replacement | Module Replacement | Performed Module Replacement B1/807091 into 603891 B3/807093 into 870925 Performed Preventive Maintenance | Unit is good condition |  |  |  | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5474989498816978866/5474989498816978866_signature_28.png | J-Bie Alvarado | https://www.jotform.com/uploads/macarewecare/222927978372067/5474989498816978866/5474989498816978866_signature_31.png
| 2022-12-22 | 2022-12-22 14:07 | Lubao Rural Health Unit 3 | Lubao Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22773 | PBSP | nathaliecoleencovar@gmail.com | 28.2 | ASSAY: Xpert MTB/RIF, COUNT: 318 ASSAY: Xpert MTB/RIF Ultra, COUNT: 425 |  | GXIV-D | 812834 | Module replacement for A3 | #NAME? | Preventive maintenance and repair | Unit is operational |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5474989314117490371/5474989314117490371_signature_28.png | Nathalie Covar | https://www.jotform.com/uploads/macarewecare/222927978372067/5474989314117490371/5474989314117490371_signature_31.png
| 2022-12-22 | 2022-12-22 13:28 | Bukidnon Provincial Medical Center | Malaybalay, Bukidnon | 10 | Bukidnon | Laboratory | YES | Module Replacement | CEPHEID | [5 empty] | 22772 | PBSP | maureenlopez65@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1450 ASSAY: Xpert HIV-1 Viral Load, COUNT: 7 |  | GeneXpert | 804708 | H/W Fail module B4/603321 | Module Replacement | Replaced Module New Module SN: 805519 | Operational |  |  |  | Kim Ryan Arrojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5474984184647287132/5474984184647287132_signature_28.png | Maureen Lopez | https://www.jotform.com/uploads/macarewecare/222927978372067/5474984184647287132/5474984184647287132_signature_31.png
| 2022-12-22 | 2022-12-22 12:46 | To Love Yourself | Shaw Blvd., Mandaluyong City | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | [5 empty] | 22771 | COMMERCIAL | laboratory@loveyourself.ph | 25 |  |  | Primo Star | 3134003748 | Unit for calibration. | Calibrated unit. | #NAME? | Unit is operational | Calibration certificate to follow |  |  | Ronnel Belizar, John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5474947250814357274/5474947250814357274_signature_28.png | Mariel Chari Canoxa | https://www.jotform.com/uploads/macarewecare/222927978372067/5474947250814357274/5474947250814357274_signature_31.png
| 2022-12-22 | 2022-12-22 12:46 | To Love Yourself | Shaw Blvd., Mandaluyong City | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | CALIBRATION MAINTENANCE | CARL ZEISS | [5 empty] | 22770 | COMMERCIAL | laboratory@loveyourself.ph | 25 |  |  | Primo Star | 3134003748 | Unit for calibration. | Calibrated unit. | #NAME? | Unit is operational | Calibration certificate to follow |  |  | Ronnel Belizar, John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5474947220819198774/5474947220819198774_signature_28.png | Mariel Chari Canoxa | https://www.jotform.com/uploads/macarewecare/222927978372067/5474947220819198774/5474947220819198774_signature_31.png
| 2022-12-22 | 2022-12-21 12:28 | Bureau of Jail and Penology Infirmary Calamba Laguna | Turbina Calamba Laguna | 4A | Laguna | Infirmary Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22769 | PBSP | rivainfirmary@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1131 | 2022-12-19 | GXIV | 814156 | Unit for PM. For Module Replacement of Modules C1 and C4. | Performed PM of Genexpert. Performed module replacement on C1 and C4. | PM and Module Replacement. | Unit is operational. 4 working modules. |  | Gx module C1 (600850) and C4 (702817) in replacement for defective module C1 (793458) and C4 (698578). | G4980 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5474942211913532862/5474942211913532862_signature_28.png | JO1 Rovicson R Monterubio | https://www.jotform.com/uploads/macarewecare/222927978372067/5474942211913532862/5474942211913532862_signature_31.png
| 2022-12-22 | 2022-12-22 11:42 | Leyte Provincial Hospital | Pawing, Palo Leyte | 8 | Leyte | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22768 | PBSP | maryllelegio96@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 474 | 2022-12-22 | GeneXpert GXIV-4D | 815267 | #NAME? | #NAME? | Module replacement - Def. module SN: 662971 - new module SN: 627304 -Preventive Maintenance -Generate reports | unit is in good condition |  | 900-0491RS | GeneXpert Module 6 Color SN: 627304 | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5474904085024369319/5474904085024369319_signature_28.png | Maryll faye O. Elegio | https://www.jotform.com/uploads/macarewecare/222927978372067/5474904085024369319/5474904085024369319_signature_31.png
| 2022-12-22 | 2022-12-20 14:35 | Eastern Samar Provincial Hospital | Borongan City | 8 | Eastern Samar | TB DOTS | YES | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22767 | PBSP | elvalynglom@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 549 ASSAY: Xpert MTB/RIF Ultra, COUNT: 315 | 2022-12-20 | GXIV - 4D | 804256 | For module replacement (B1,B2,B3,B4) | #NAME? | -Module Replacement B1 def. SN 639236 - new SN 750949 B2 def. SN 639233 - new SN 668765 B3 def. SN 639232 - new SN 669662 B4 def. SN 685270 - new SN 730823 - Preventive Maintenance -Generate Reports | Unit is now in good condition |  | GeneXpert 6C module | 900-0491RS | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5474893005029821788/5474893005029821788_signature_28.png | ELVALYN GRACIA G. LOMUNTAD | https://www.jotform.com/uploads/macarewecare/222927978372067/5474893005029821788/5474893005029821788_signature_31.png
| 2022-12-22 | 2022-12-22 11:09 | Koronadal City Health Office | Koronadal City | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22766 | PBSP | calvokenneth12@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF, COUNT: 1012 ASSAY: Xpert MTB/RIF Ultra, COUNT: 903 | 2022-12-22 | Gx-4 | 707529 | Module A1 and A3 H/W failed | Replace module A1 SN:754528 with new module SN:866833 Genexpert machine Preventive Maintenance | Replace module A1 SN:754528 with new module SN:866833 Genexpert machine Preventive Maintenance | Gx machine has 3 Operational modules | Module A3 waiting for module to available |  |  | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5474883440124154533/5474883440124154533_signature_28.png | Kenneth Paolo T. Calvo, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5474883440124154533/5474883440124154533_signature_31.png
| 2022-12-22 | 2022-12-22 10:48 | National Tuberculosis Reference Laboratory | Filinvest, Muntinlupa City | NCR | NCR 4th - Muntinlupa | TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22765 | PBSP | ntrlphil_cs@yahoo.com.ph | 25 | ASSAY: Xpert MTB/RIF, COUNT: 14 ASSAY: Xpert MTB/RIF Ultra, COUNT: 11 ASSAY: Others, COUNT: 2971 | 2022-12-22 | GXIV 10C | 848244 | Module B4 | Replaced Module B4 210038638 with 210075576 | Performed Maintenance | Unit is Operational with 4 Modules | Good module only borrowed from ZCMC | 210075576 10C Module | PAL AWB 079-08465936 | Carlos Origenes | https://www.jotform.com/uploads/macarewecare/222927978372067/5474871485989864081/5474871485989864081_signature_28.png | JENKIN A. MENDOZA | https://www.jotform.com/uploads/macarewecare/222927978372067/5474871485989864081/5474871485989864081_signature_31.png
| 2022-12-22 | 2022-12-22 10:05 | Valenzuela Medical Center | Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22764 | PBSP | kuliscool4@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 791 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1318 ASSAY: Xpert HIV-1 Viral Load, COUNT: 8 | 2022-12-19 | GXIV | 805045 | Module B1 for replacement | Replaced module B1 | - generated pre reports - replaced module B1 (620354) - generated post reports - checked cartridge bay, plunger rod, filter - performed functionality test | All 4 modules operational | NA | GeneXpert Module SN: 645161 | G 4984 | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5474842062229287664/5474842062229287664_signature_28.png | Francis Paul D. Salmos | https://www.jotform.com/uploads/macarewecare/222927978372067/5474842062229287664/5474842062229287664_signature_31.png
| 2022-12-22 | 2022-12-22 09:09 | Don Carlos Rural Health Unit | Don Carlos, Bukidnon | 10 | Bukidnon | Laboratory | YES | Module Replacement | CEPHEID | [5 empty] | 22763 | PBSP | gibsz.l002@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 556 |  | GeneXpert | 807625 | Error 1002 on Module C3/722694 | Module replacement | Replaced Module | Module C1,C2 and C3 are operational Module C4 H/W Fail waiting for module replacement |  |  |  | Kim Ryan Arrojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5474823086912110164/5474823086912110164_signature_28.png | Gilbert Loro | https://www.jotform.com/uploads/macarewecare/222927978372067/5474823086912110164/5474823086912110164_signature_31.png
| 2022-12-22 | 2022-12-22 09:08 | Rosario Super Health Center | Rosario, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22762 | WHO | eden.cawayan@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 306 ASSAY: Xpert MTB/RIF Ultra, COUNT: 878 ASSAY: Xpert SARS-CoV-2, COUNT: 551 | 2022-12-19 | GeneXpert GX-IV-D | 843924 | Modules B1, B2, B3 for replacement | Replaced modules B1,B2,B3 | Performed maintenance Replaced defective modules Generated reports | Unit is operational. | Module B4 not operational. Already reported. | Old modules: B1/893044,B2/893050,B3/893099 New modules: B1/638232,B2/674387,B3/699915 | G 4973 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5474809706018800291/5474809706018800291_signature_28.png | Eden R. Cawayan, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5474809706018800291/5474809706018800291_signature_31.png
| 2022-12-22 | 2022-12-22 08:55 | Cauayan City Health Office | Cauayan City, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22761 | PBSP | cho1lab14@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 643 | 2022-12-22 | Genexpert IV | 814373 | Sometimes Lost Communication error | Troubleshoot the cables | Performed Maintenance Gathered Reports Performed Module Replacement B1 - old - 691571 - new - 696408 Performed Functionality Test | Unit is Operational | N/a | Genexpert Module - 696408 | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5474800465321163312/5474800465321163312_signature_28.png | Arlyn C. Daggao | https://www.jotform.com/uploads/macarewecare/222927978372067/5474800465321163312/5474800465321163312_signature_31.png
| 2022-12-21 | 2022-12-21 12:01 | Graman Medical And Maternity Hospital | Malolos Bulacan | 3 | Bulacan | Molecular Laboratory | NO | INSTALLATION | CEPHEID | [5 empty] | 22760 | COMMERCIAL | robertsayo72@gmail.com | 22 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1 |  | GXIV | 110008530 | None | None | Performed unit installation. Performed funstionality test. Run actual sample with acceptable result. | Unit is operational. | None |  |  | Harold Toco | https://www.jotform.com/uploads/macarewecare/222927978372067/5474213763517818828/5474213763517818828_signature_28.png | Robert Sayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5474213763517818828/5474213763517818828_signature_31.png
| 2022-12-21 | 2022-12-21 14:18 | University of the Philippines - Diliman | UP Diliman | NCR | NCR 2nd - Quezon City | Marine Science Institute | YES | INSTALLATION | CARL ZEISS | [5 empty] | 22759 | COMMERCIAL | iccatolico@up.edu.ph | 25 |  | 2022-12-21 | Primostar 3 | 3185001467 | For Installation | Installed the unit | Delivery and installation of unit Functionality test | Unit is operational. |  |  | SI# 94528 | John Monsour Doctor/Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5474144674225407584/5474144674225407584_signature_28.png | Luigi C Catolico | https://www.jotform.com/uploads/macarewecare/222927978372067/5474144674225407584/5474144674225407584_signature_31.png
| 2022-12-21 | 2022-12-21 13:32 | Mangatarem Rural Health Unit | Mangatarem Pangasinan | 1 | Pangasinan | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22758 | DOH | magallionjason07@gmail.com | 25.4 | ASSAY: Xpert MTB/RIF, COUNT: 215 ASSAY: Xpert MTB/RIF Ultra, COUNT: 221 |  | GXIV-D | 838284 | HW Failed Module B3 | -Performed surface cleaning -Performed cleaning of cartridge bay,icore,plunger rod and filter -Archived test result -Update software to 1.5a -Performed Xpertcheck -Generate reports -Performed functionality test | Preventive maintenance and Calibrtion | Unit is operational except Module B3 | B3 for reporting |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5474115414331682562/5474115414331682562_signature_28.png | Jayson M. Magallon | https://www.jotform.com/uploads/macarewecare/222927978372067/5474115414331682562/5474115414331682562_signature_31.png
| 2022-12-21 | 2022-12-21 13:34 | La Union Medical Center | Philippine National Road, Agoo La Union | 1 | La Union | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22757 | PBSP | jubilocatherine@gmail.com | 25.2 | ASSAY: Xpert MTB/RIF, COUNT: 882 ASSAY: Xpert MTB/RIF Ultra, COUNT: 900 |  | GX4D | 707400 | Unit Is Due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Plunger Rod Cleanin Performed Filter Cleaning | Unit is operational |  |  |  | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5474105154818881888/5474105154818881888_signature_28.png | Catherine Joy Jubilo | https://www.jotform.com/uploads/macarewecare/222927978372067/5474105154818881888/5474105154818881888_signature_31.png
| 2022-12-21 | 2022-12-21 12:37 | Amang Rodriguez Memorial Medical Center | Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22756 | PBSP | pathologymole.armmc@gmail.com | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 3807 ASSAY: Others, COUNT: 1 | 2022-12-19 | GeneXpert GX-XVI-D | 844113 | Modules B2,B4,D2 for replacement | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. |  | Old modules: B2/891862,B4/884017,D2/892186 Nee modules: B2/628603, B4/804084,D2/647024 | G 4974 | John Monsour Doctor, Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5474068985921011362/5474068985921011362_signature_28.png | Ma Lanica Manda | https://www.jotform.com/uploads/macarewecare/222927978372067/5474068985921011362/5474068985921011362_signature_31.png
| 2022-12-21 | 2022-12-21 12:12 | Agno rhu | Agno pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22755 | PBSP | rena_humilde@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 77 ASSAY: Xpert MTB/RIF Ultra, COUNT: 87 | 2022-12-21 | GX4 | 834837 | For module replacement b2 784784 - 210057605 | Performed module replacement,cleaned unit,filter,icore,plunger rod,cartridge bay,gathered reports | Module replacement, maintenance | Operational b2, b1 reported by end user thru jotform |  | 900 gx | 0 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5474055779127111296/5474055779127111296_signature_28.png | Maria Cris Rena H. Arenas | https://www.jotform.com/uploads/macarewecare/222927978372067/5474055779127111296/5474055779127111296_signature_31.png
| 2022-12-21 | 2022-12-20 15:36 | Philippine heart center | East ave diliman q.c | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | [5 empty] | 22754 | MACARE-HBDC | phcmolecularlab@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 13 ASSAY: Xpert MTB/RIF Ultra, COUNT: 696 ASSAY: Others, COUNT: 14 ASSAY: Xpert SARS-CoV-2, COUNT: 17432 | 2022-12-20 | GXXVI | 838655 | module replacement b1 210068311 - 647383 | Performed maintenance, performed module replacement,gathered reports | Module replacement | 12 module operational c1 for reporting c3,cr,d2 reported |  | 900gx | G4986 | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5474051439127966374/5474051439127966374_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5474051439127966374/5474051439127966374_signature_31.png
| 2022-12-21 | 2022-12-21 11:33 | Ilocos Training and Regional Medical Center | Philippine National Road, San Fernando City, La Union | 1 | La Union | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22753 | PBSP | mfmadayag31@gmail.com | 24.1 | ASSAY: Xpert HPV, COUNT: 126 ASSAY: Xpert MTB/RIF Ultra, COUNT: 270 ASSAY: Xpert SARS-CoV-2, COUNT: 27 ASSAY: Xpert HIV-1 Viral Load, COUNT: 44 |  | GX4L | 707524 | Unit is due for module replacement | Module Replacement | Module Replacement B3/717580 into 716353 B4/688184 into 798629 | Unit is Operational |  |  |  | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5474039234811470542/5474039234811470542_signature_28.png | Michelle F. Madayag | https://www.jotform.com/uploads/macarewecare/222927978372067/5474039234811470542/5474039234811470542_signature_31.png
| 2022-12-21 | 2022-12-21 11:19 | Gilbert O. Teodoro Memorial Hospital | Malacampa Camiling Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22752 | PBSP | almabellapdancel3@gmail.com | 28.6 | ASSAY: Xpert MTB/RIF, COUNT: 282 ASSAY: Xpert MTB/RIF Ultra, COUNT: 259 |  | GXIV-D | 812838 | Due for maintenance | #NAME? | Preventive maintenance | Unit is operational |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5474025044332648499/5474025044332648499_signature_28.png | Alma Dancel | https://www.jotform.com/uploads/macarewecare/222927978372067/5474025044332648499/5474025044332648499_signature_31.png
| 2022-12-21 | 2022-12-21 10:55 | Ilagan City Health Office | Ilagan City, Isabela | 2 | Isabela | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22751 | PBSP | cho1ilagan.laboratory@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 902 ASSAY: Xpert MTB/RIF Ultra, COUNT: 671 | 2022-12-21 | GENEXPERT IV | 808635 | Module 3 Hardware Failure UPS not working | Troubleshoot the UPS | Performed Maintenance Gathered Reports Performed Module Replacement B1 - old - 671265 - new - 802431 Performed Functionality Test | Unit Is Operational | UPS needs to be replaced | Genexpert Module - 802431 | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5474018734911921041/5474018734911921041_signature_28.png | April Jenneth Chu | https://www.jotform.com/uploads/macarewecare/222927978372067/5474018734911921041/5474018734911921041_signature_31.png
| 2022-12-21 | 2022-12-21 09:14 | The Medical City South Luzon | Sta. Rosa Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22750 | COMMERCIAL | jddaquis@themedicalcity.com | 25 | ASSAY: Xpert HPV, COUNT: 5 ASSAY: Xpert MTB/RIF, COUNT: 54 ASSAY: Xpert MTB/RIF Ultra, COUNT: 92 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 |  | GXIV | 839233 | Unit for Preventive Maintenance. | Performed Preventive Maintenance. Cleaned docking bay, plunger, icore and filter. Generated reports. | Preventive Maintenance. | Unit is operational. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5474003828696229096/5474003828696229096_signature_28.png | JOHN MICHAEL DAQUIS | https://www.jotform.com/uploads/macarewecare/222927978372067/5474003828696229096/5474003828696229096_signature_31.png
| 2022-12-21 | 2022-12-21 09:14 | Bayugan City Health Office | Poblacion, Bayugan City | 13 | Agusan del Sur | RTDL | NO | REPAIR | CEPHEID | [5 empty] | 22749 | PBSP | diolalourdes@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 2 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1079 |  | GX-4 | 815391 | None | None | Replacing B2 with SN : 720096 to 901636 Replacing B4 with SN: 624234 to 707736 | 2 modules not functional and 2 modules been replaced |  |  |  | Engr. Faizar Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5473974294421112904/5473974294421112904_signature_28.png | Lourdes Diola | https://www.jotform.com/uploads/macarewecare/222927978372067/5473974294421112904/5473974294421112904_signature_31.png
| 2022-12-21 | 2022-12-21 09:32 | Kidapawan City Health Office | Kidapawan City | 12 | Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22748 | DOH | malufranz5@gmail.com | 23.7 | ASSAY: Xpert MTB/RIF, COUNT: 904 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1328 | 2022-12-21 | Gx - 4 | 838545 | Module B3 and B4 H/W failed | Genexpert machine Preventive Maintenance | Genexpert machine Preventive Maintenance | Gx machine has 2 operational modules |  |  |  | Jesson Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5473962532537031796/5473962532537031796_signature_28.png | Ma. Luisa M. Francisco, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5473962532537031796/5473962532537031796_signature_31.png
| 2022-12-21 | 2022-12-20 12:44 | Cotabato Provincial Hospital | Brgy. Amas, Kidapawan City | 12 | Cotabato | Laboratory | YES | REPAIR CALIBRATION | CEPHEID | [5 empty] | 22747 | PBSP | jydm2073@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 196 | 2022-12-20 | Gx-4 | 807615 | Module B2 and B4 H/W FAILED | Replace module B2 SN: 784365 with new module SN: 868129 Replace module B4 SN: 662444 with new module SN: 690460 Genexpert machine Preventive Maintenance | Replace module B2 SN: 784365 with new module SN: 868129 Replace module B4 SN: 662444 with new module SN: 690460 Genexpert machine Preventive Maintenance | Genexpert machine in Good condition | Gx machine has 4 operational modules |  |  | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5473955727751113183/5473955727751113183_signature_28.png | Joey L. Marte, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5473955727751113183/5473955727751113183_signature_31.png
| 2022-12-21 | 2022-12-21 08:47 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22746 | PBSP | marichu_latoy@yahoo.com.ph | 22 | ASSAY: Xpert MTB/RIF, COUNT: NA ASSAY: Xpert MTB/RIF Ultra, COUNT: NA |  | GXIV | 805112 | Gathering of additional data for troubleshooting | Taken a picture of gateway board with serial numbers | #NAME? | Unit is operational. | Cords of the unit are currently installed on their other unit (SN: 807466) |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5473931966911176624/5473931966911176624_signature_28.png | MARIA YVETTE C MENDOZA | https://www.jotform.com/uploads/macarewecare/222927978372067/5473931966911176624/5473931966911176624_signature_31.png
| 2022-12-20 | 2022-12-20 11:22 | PERPETUAL SUCCOUR HOSPITAL | GORORDO AVE., CEBU CITY | 7 | Cebu | BLOOD BANK | YES | REPAIR | HAEMONETICS | [5 empty] | 22744 | COMMERCIAL | rodaflores1038@yahoo.com.ph | 24 |  | 2022-12-15 | MCS+ | 11H307 | FOR LINE SENSOR WITH COVER ASSEMBLY REPLACEMENT | #NAME? | #NAME? | #NAME? |  | LINE SENSOR WITH COVER ASSEMBLY | SI NO 94500 | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5473359888814814592/5473359888814814592_signature_28.png | RODA FLORES, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5473359888814814592/5473359888814814592_signature_31.png
| 2022-12-20 | 2022-12-20 16:30 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Microbiology | NO | MAINTENANCE | CEPHEID | [5 empty] | 22743 | PBSP | marichu_latoy@yahoo.com.ph | 21 | ASSAY: Xpert MTB/RIF, COUNT: 602 ASSAY: Xpert MTB/RIF Ultra, COUNT: 609 ASSAY: Xpert HCV VL, COUNT: 1 ASSAY: Xpert HIV-1 Viral Load, COUNT: 139 |  | GXIV | 807466 | Unit for maintenance | NA | #NAME? | Unit is operational. 3 modules are working. Module D2 for follow up. | D2 is disabled. Unit is a rovering unit (from Tarlac) |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5473346408958565143/5473346408958565143_signature_28.png | Marichu Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5473346408958565143/5473346408958565143_signature_31.png
| 2022-12-20 | 2022-12-20 15:07 | Caraga Regional Hospital | Rizal St. Brgy. Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | CALIBRATION | CEPHEID | [5 empty] | 22742 | PBSP | abpabellan.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1017 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2495 |  | GX-4 | 807602 | All modules are service required | None | Xpertcheck Calibration of back-up unit formerly ( Zamboanga City Medical Center) | All modules are service required. Replacement Needed. |  |  |  | Engr. Faizar Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5473302559367679918/5473302559367679918_signature_28.png | Andrew Pabellan | https://www.jotform.com/uploads/macarewecare/222927978372067/5473302559367679918/5473302559367679918_signature_31.png
| 2022-12-20 | 2022-12-20 14:53 | Caraga Regional Hospital | Rizal St. Brgy. Washington, Surigao City | 13 | Surigao del Norte | RTDL | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22741 | PBSP | abpabellan.pbsp@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1017 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2495 ASSAY: Xpert HIV-1 Viral Load, COUNT: 29 |  | GX - 4 | 815489 | D1, D2 and D4 temperature related errors | Reported to Cepheid. 3 modules prior for replacements. | Replacing Module D1 with SN: 710720 to SN: 677347 Replacing Module D2 with SN: 608331 to SN: 798340 Replacing Module D4 with SN: 861748 to SN: 702788 | All modules working |  |  |  | Engr. Faizar S. Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5473292269363058552/5473292269363058552_signature_28.png | Andrew Pabellan | https://www.jotform.com/uploads/macarewecare/222927978372067/5473292269363058552/5473292269363058552_signature_31.png
| 2022-12-20 | 2022-12-20 12:25 | Tanay Rural Health Unit | Tanay, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22740 | PBSP | maryjunnette@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 377 ASSAY: Xpert MTB/RIF Ultra, COUNT: 571 | 2022-12-19 | GeneXpert GX-IV-D | 807593 | Module A2 for replacement | Replaced module A2 | Performed maintenance Replaced module A2 Generated reports Replaced bios battery | All modules operational. | UPS for replacement. | Old module: A2/677922 New module: A2/862098 | G 4972 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5473247012519592480/5473247012519592480_signature_28.png | Mary Junnette Fulgueras | https://www.jotform.com/uploads/macarewecare/222927978372067/5473247012519592480/5473247012519592480_signature_31.png
| 2022-12-20 | 2022-12-20 13:29 | Central Luzon Doctor's Hospital | Hospital Drive Tarlac City | 3 | Tarlac | Molecular Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22739 | COMMERCIAL | moleclab@cldh.com.ph | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 3 ASSAY: Xpert SARS-CoV-2, COUNT: 75 |  | GXIV-D | 830841 | Module A1 HW FAILED | Performed Preventive Maintenance Checked system log Performed functionality test | Preventive maintenance Troubleshooting | Unit is operational except Module A1:210070792 | For reporting A1 for replacement. |  |  | MARK JOSEPH VENZON | https://www.jotform.com/uploads/macarewecare/222927978372067/5473243804337387170/5473243804337387170_signature_28.png | Justine Pader | https://www.jotform.com/uploads/macarewecare/222927978372067/5473243804337387170/5473243804337387170_signature_31.png
| 2022-12-20 | 2022-12-20 13:14 | Rogaciano M. Mercado Memorial General Hospital | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22738 | PBSP | rhia.aragon@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 754 ASSAY: Xpert MTB/RIF Ultra, COUNT: 988 | 2022-12-19 | GXIV | 805252 | Module A3 (696903) for replacement | Replaced module A3 | #NAME? | All 4 modules are operational. |  | GX Module SN: 896226 | G 4983 | Ronnep Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5473241808959755553/5473241808959755553_signature_28.png | Rhia Aragon | https://www.jotform.com/uploads/macarewecare/222927978372067/5473241808959755553/5473241808959755553_signature_31.png
| 2022-12-20 | 2022-12-20 12:36 | Belmonte Health Center | Smapaloc Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22737 | PBSP | jpmarcos.29@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 363 ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 ASSAY: Xpert HIV-1 Viral Load, COUNT: 13 | 2022-12-19 | GXIV | 814376 | Unit for PM. For module replacement of A2. | Performed PM of Genexpert. Performed module replacement on A2. | PM and module replacement. | Unit is operational. |  | GX module for A2 (671861) in replacement for defective A2 (691293). | G4979 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5473215997212289352/5473215997212289352_signature_28.png | John Patrick Marcos | https://www.jotform.com/uploads/macarewecare/222927978372067/5473215997212289352/5473215997212289352_signature_31.png
| 2022-12-20 | 2022-12-20 12:55 | NTRL | Filinvest | NCR | NCR 4th - Muntinlupa | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22736 | PBSP | mspogosa.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 839 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1087 |  | GXIV-D | 707512 | 2032 errors | Additional troubleshooting by cepheid | Performed maintenance Checked gateway board Generate reports | 2 modules are defective |  |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5473215749807852532/5473215749807852532_signature_28.png | Maria Guileane Pogosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5473215749807852532/5473215749807852532_signature_31.png
| 2022-12-20 | 2022-12-20 12:13 | Baguio Health Service Office | T. Alonso St., New Lucban, Baguio city | CAR | Benguet | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22735 | DOH | sherryannemarrero@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 92 ASSAY: Xpert MTB/RIF Ultra, COUNT: 646 ASSAY: Xpert HPV, COUNT: 3 |  | GXIV | 838281 | Unit is due for reexpert checy Unit is due for Preventive Maintenance | Expert Check and Preventive Maintenance | Performed Surface Cleaning Performed plunger Rod Cleaning Peroformed Filter Cleaning | Unit is operational |  |  |  | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5473194982015564255/5473194982015564255_signature_28.png | Sherry Anne Marrero | https://www.jotform.com/uploads/macarewecare/222927978372067/5473194982015564255/5473194982015564255_signature_31.png
| 2022-12-20 | 2022-12-20 12:04 | Northern Mindanao Medical Center | Cagayan de Oro City | 10 | Misamis Oriental | Laboratory | YES | Module Replacement | CEPHEID | [5 empty] | 22734 | PBSP | kimlestherbautista@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 6743 ASSAY: Xpert HIV-1 Viral Load, COUNT: 421 ASSAY: Xpert HPV, COUNT: 186 ASSAY: Xpert SARS-CoV-2, COUNT: 12629 | 2022-12-19 | GeneXpert | 844118 | H/W Fail Module C1 SN: 805858 | Module Replacement | Replaced Module C1/805858 | Operational | Module C1 ready for use. | GeneXpert 6-Color Module | N/A | Kim Ryan Arrojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5473187793427652512/5473187793427652512_signature_28.png | Bautista, Kim | https://www.jotform.com/uploads/macarewecare/222927978372067/5473187793427652512/5473187793427652512_signature_31.png
| 2022-12-20 | 2022-12-20 10:58 | Dimasalang Health Center | Dimasalang Sta. Cruz, Manil | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22733 | PBSP | kikokendz_060830@yahoo.com | 25 | ASSAY: Xpert HPV, COUNT: 26 ASSAY: Xpert MTB/RIF, COUNT: 395 ASSAY: Xpert MTB/RIF Ultra, COUNT: 492 | 2022-12-19 | GXIV | 707072 | Unit for PM. For module replacement of Modules b1 and b3. | Performed PM. Performed module replacement on b1 and b3. Generated reports. | PM and Module Replacement | Unit is operational. |  | GX Module B1 (702411) and B3 (804103) in replacement for defective B1 (796662) and B3 (687758). | G4981 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5473149417216297977/5473149417216297977_signature_28.png | Maria joy cleofas m. Abenales | https://www.jotform.com/uploads/macarewecare/222927978372067/5473149417216297977/5473149417216297977_signature_31.png
| 2022-12-20 | 2022-12-20 09:18 | Baguio General Hospital and Medical Center | Gov. Pack rd. Bgh Compound, Baguio City | CAR | Benguet | TB LABORATORY | NO | MAINTENANCE | CEPHEID | [5 empty] | 22732 | PBSP | keithtagudar@gmail.com | 24.3 | ASSAY: Xpert HPV, COUNT: 414 ASSAY: Xpert MTB/RIF, COUNT: 807 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1236 |  | GXIV | 834685 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed plunger Rod Cleaning Performed filter Cleaning | Unit is operational |  |  |  | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5473094813318544383/5473094813318544383_signature_28.png | Keith William Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5473094813318544383/5473094813318544383_signature_31.png
| 2022-12-20 | 2022-12-20 09:09 | Baguio General Hospital and Medical Center | Gov. Pack rd., Bgh Compund, Baguio City | CAR | Benguet | TB LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22731 | PBSP | keithtagudar@gmail.com | 23.4 | ASSAY: Xpert HPV, COUNT: 208 ASSAY: Xpert MTB/RIF, COUNT: 427 ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 | 2022-12-20 | GXIV | 805212 | Unit is due for Module Replacement | Module Replacement | Replace Modules: B3/758560 into 875228 B4/682707 into 713404 | Unit is operational |  | Genexpert module 6C 875228 713404 | PBSP | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5473093753312969309/5473093753312969309_signature_28.png | Keith William Tagudar | https://www.jotform.com/uploads/macarewecare/222927978372067/5473093753312969309/5473093753312969309_signature_31.png
| 2022-12-19 | 2022-12-19 10:43 | Tunasan Health Center | Tunasan | NCR | NCR 4th - Muntinlupa | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22730 | PBSP | gigietan0911@xn--gmai-jqa.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 283 ASSAY: Xpert MTB/RIF Ultra, COUNT: 892 ASSAY: Xpert HIV-1 Viral Load, COUNT: 111 |  | GXIV-D | 807613 | GeneXpert connection problem | Troubleshoot connections and reset gateway board | Performed maintenance Check connections Used different cables Reset Gateway Board Generate Reports | Unit is operational but with IP configuration. Possible replacement of cpu or gateway board. | For observation of CPU and GeneXpert |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5472684454704205354/5472684454704205354_signature_28.png | Geraldine Tan | https://www.jotform.com/uploads/macarewecare/222927978372067/5472684454704205354/5472684454704205354_signature_31.png
| 2022-12-19 | 2022-12-19 16:31 | Philippine heart center | East ave diliman quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | Ups replacement | HAEMONETICS | [5 empty] | 22729 | MACARE-HBDC | sheild419@gmail.com | 26 |  | 2022-10-12 | Mcs + | 01M201 | Defective ups unable to backup | Replaced 1kva ups | Replacement of ups | Unoperational defective cpu board |  | Ups 1 kva | 9281 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5472644969469609882/5472644969469609882_signature_28.png | Sheila Decena | https://www.jotform.com/uploads/macarewecare/222927978372067/5472644969469609882/5472644969469609882_signature_31.png
| 2022-12-19 | 2022-12-19 16:35 | Philippine heart center | East ave diliman quezon city | NCR | NCR 2nd - Quezon City | Bloodbank | NO | Ups replacement | HAEMONETICS | [5 empty] | 22728 | MACARE-HBDC | sheild419@gmail.com | 26 |  | 2022-10-12 | Mcs+ | 02D168 | Ups unable to back up | Replace defective ups | Ups replacement | Operational |  | Ups 1kva | 9280 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5472644559466830819/5472644559466830819_signature_28.png | Sheila Decena | https://www.jotform.com/uploads/macarewecare/222927978372067/5472644559466830819/5472644559466830819_signature_31.png
| 2022-12-19 | 2022-12-19 11:22 | San Isidro RHU | Poblacion, San Isidro, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22727 | PBSP | urian_marlyn@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 352 ASSAY: Xpert MTB/RIF, COUNT: 46 | 2022-12-19 | Genexpert IV | 805107 | UPS automatic shutdown | Troubleshoot the UPS | Performed Maintenance Gathered Reports Performed Module Replacement Installed Printer Drivers Performed Functionality Test | Unit ia operational | UPS needs to be REPLACED | 900-0491RS | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5472298388517774229/5472298388517774229_signature_28.png | Marlyn M. Urian | https://www.jotform.com/uploads/macarewecare/222927978372067/5472298388517774229/5472298388517774229_signature_31.png
| 2022-12-19 | 2022-12-09 14:33 | Pascual Laboratories | McArthur Highway, Balagtas | 3 | Bulacan | QC | YES | CALIBRATION MAINTENANCE | CARL ZEISS | [5 empty] | 22726 | COMMERCIAL | mbjimenez@pascuallab.com | 25 |  |  | Axioscope A1 | 3321001650 | For PM | Performed maintenance and calibration | Performed maintenance and calibration Cleaned the microscope body Cleaned lenses Cleaned objectives Checked the light source Check illumination | Unit is operational. For bulb replacement. | Low light output. |  |  | Jogn Monsour Doctor/ Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5472239484113106312/5472239484113106312_signature_28.png | Maureen Jimenez | https://www.jotform.com/uploads/macarewecare/222927978372067/5472239484113106312/5472239484113106312_signature_31.png
| 2022-12-13 | 2022-12-13 09:15 | St Luke's Extension Clinic | Bocobo St, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22725 | MACARE-HBDC | laboratory@slec.ph | 25 | ASSAY: Xpert CT/NG, COUNT: 4003 ASSAY: Xpert MTB/RIF, COUNT: 2035 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2475 |  | GeneXpert Infinity 48s | 809643 | LIS Connectivity | Transferred connection to different Ethernet port | Performed troubleshooting with in-house IT | Unit is operational, LIS connectivity under observation |  |  |  | Patrick Los Baños | https://www.jotform.com/uploads/macarewecare/222927978372067/5467037652419849455/5467037652419849455_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5467037652419849455/5467037652419849455_signature_31.png
| 2022-12-09 | 2022-12-09 14:56 | Cuyegkeng health center | Pasay city | NCR | NCR 4th - Pasay | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22724 | PBSP | LBCANICULA@gmail.com | 32 | ASSAY: Xpert MTB/RIF, COUNT: 236 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1461 |  | Gx 4 | 805106 | For maintenance | Performed maintenance,cleaned plunger,cartridge bay,icore,filter,gathered reports,test run unit | Maintenance | B1 failed xpertcheck,3 modules operational |  |  |  | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5463899736427413494/5463899736427413494_signature_28.png | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5463899736427413494/5463899736427413494_signature_31.png
| 2022-12-09 | 2022-12-09 11:59 | Doña marta health center | 37 Revilla st pasay city | NCR | NCR 4th - Pasay | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22723 | DOH | cerissalynsabado@yahoo.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 461 |  | Gx4 | 831700 | For maintenance | Performed maintenance,cleaned filter,plunger,icore,cartridge bay,gathered report,test run unit | Maintenance | Operational |  |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5463899166424223713/5463899166424223713_signature_28.png | Cerissa Lyn Sabado | https://www.jotform.com/uploads/macarewecare/222927978372067/5463899166424223713/5463899166424223713_signature_31.png
| 2022-12-09 | 2022-12-09 15:41 | Santa Rosa RHU | Santa Rosa Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22722 | DOH | rhu1santarosa@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-12-09 | Genexpert IV | 831599 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5463813854919875214/5463813854919875214_signature_28.png | Ana S Caballong | https://www.jotform.com/uploads/macarewecare/222927978372067/5463813854919875214/5463813854919875214_signature_31.png
| 2022-12-09 | 2022-12-09 14:53 | Panabo City Health Office | Panabo City, Davao del Norte | 11 | Davao del Norte | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22721 | DOH | nadineeust67@gmail.com | 21.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1005 | 2022-12-09 | Gx- IV 4D | 838295 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Cartridge Bay, surface area, Plunger Rod, fan filters, Icore Generate IQR, Assay Stat, System Log Archived Data Results | 3Module are fully operational Module B2 / SN:808297 encounter multiple Temperature Error GeneXpert Monitor Display is Malfunctioned | Module B2/ SN:808297 under observation ; Monitor was temporarily replaced by the End User | N/A | N/A | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5463794536301159805/5463794536301159805_signature_28.png | Nadine Gemma S. Eustaquio, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5463794536301159805/5463794536301159805_signature_31.png
| 2022-12-09 | 2022-12-09 14:42 | Philippine General Hospital | Rizal Ave. Ermita | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22720 | DOH | rcharlamagnenabong@gmail.com | 25.8 | ASSAY: Xpert HPV, COUNT: 342 ASSAY: Xpert SARS-CoV-2, COUNT: 25947 ASSAY: Others, COUNT: 10 |  | GX16 | 844975 | Unit is due for troubleshootin. no saving of results in database. | Backup and Purge Test | Transfer file from pc to usb with supervision of end user. Performed Troubleshooting | Unit is operational |  |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5463781133831737372/5463781133831737372_signature_28.png | Rhon Charlamagne Nabong | https://www.jotform.com/uploads/macarewecare/222927978372067/5463781133831737372/5463781133831737372_signature_31.png
| 2022-12-09 | 2022-12-09 13:45 | Tarlac City Health Center 3 | San Miguel Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE Monitor Check-up | CEPHEID | [5 empty] | 22719 | DOH | eunicadomingo@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 339 ASSAY: Xpert MTB/RIF Ultra, COUNT: 130 |  | GXIV-D | 831697 | Sudden shutdown of monitor and auto adjustment | #NAME? | Preventive maintenance and check-up | Unit is operational except faulty monitor |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5463742972897975427/5463742972897975427_signature_28.png | Frances Eunica Domingo | https://www.jotform.com/uploads/macarewecare/222927978372067/5463742972897975427/5463742972897975427_signature_31.png
| 2022-12-09 | 2022-12-09 12:08 | Soccsksargen General Hospital | Surallah | 12 | South Cotabato | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22718 | PBSP | johnellamarisbautista@gmail.com | 22.1 | ASSAY: Xpert MTB/RIF, COUNT: 309 ASSAY: Xpert MTB/RIF Ultra, COUNT: 652 | 2022-12-09 | Gx-4 | 815432 | Module B1 H /w failed | Replace module B1 SN: 696324 with new module SN:870231 Genexpert machine Preventive Maintenance | Replace module B1 SN: 696324 with new module SN:870231 Genexpert machine Preventive Maintenance | Genexpert machine in Good condition |  |  |  | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5463685245222802596/5463685245222802596_signature_28.png | Johnella Maris Bautista | https://www.jotform.com/uploads/macarewecare/222927978372067/5463685245222802596/5463685245222802596_signature_31.png
| 2022-12-09 | 2022-12-09 10:41 | San Pedro Rural Health Unit 2 | Langgam San Pedro Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22717 | DOH | leniemahalin@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 764 ASSAY: Xpert MTB/RIF, COUNT: 356 ASSAY: Xpert MTB/RIF Ultra, COUNT: 651 ASSAY: Xpert HIV-1 Viral Load, COUNT: 11 |  | GXIV | 821189 | Unit for PM. For Xpertcheck of Module A3 due to error. For checking of monitor | Performed PM of GeneXpert. Performed Xpertcheck on Module A3. Checked monitor. Generated reports. | PM an XpertCheck. | Unit is operational. Module A3 passed Xpertcheck. 3 working modules, Module A4 defective and was already reported previously. Monitor has lining on the bottom part and for replacement. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5463653926019210442/5463653926019210442_signature_28.png | Ma. Celenita Mahalin | https://www.jotform.com/uploads/macarewecare/222927978372067/5463653926019210442/5463653926019210442_signature_31.png
| 2022-12-09 | 2022-12-09 10:44 | Navotas Medical and Wellness Center | Naval, San Jose | NCR | NCR 3rd - Navotas | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22716 | PBSP | alejandrinoknowell@gmail.com | 26.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 633 |  | GX4D | 807592 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed filter Cleaning Performed Plunger Rod Cleaning | Unit is operational. |  |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5463632753622341946/5463632753622341946_signature_28.png | Noel Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5463632753622341946/5463632753622341946_signature_31.png
| 2022-12-09 | 2022-12-09 10:35 | Navotas Medical and Wellness Center | Naval St. San Jose | NCR | NCR 3rd - Navotas | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22715 | DOH | alejandrinoknowell@gmail.com | 26.25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 633 |  | GX4D | 831626 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed plunger Rod Cleaning Performed filter Cleaning | Unit is operational |  |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5463630733629619087/5463630733629619087_signature_28.png | Noel Alejandrino | https://www.jotform.com/uploads/macarewecare/222927978372067/5463630733629619087/5463630733629619087_signature_31.png
| 2022-12-09 | 2022-12-09 10:22 | LAPU-LAPU CITY HEALTH OFFICE | LAPU LAPU CITY CEBU | 7 | Cebu | LABORATORY | NO | MAINTENANCE | CEPHEID | [5 empty] | 22714 | DOH | tanya_pax@yahoo.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 1083 ASSAY: Xpert MTB/RIF Ultra, COUNT: 758 |  | GX-IV | 837661 | #NAME? | N/A | #NAME? | #NAME? |  |  |  | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5463620608814099060/5463620608814099060_signature_28.png | MARIA LUCILLE P. ARONG | https://www.jotform.com/uploads/macarewecare/222927978372067/5463620608814099060/5463620608814099060_signature_31.png
| 2022-12-09 | 2022-12-09 10:06 | LAPU-LAPU CITY HEALTH OFFICE | LAPU LAPU CITY CEBU | 7 | Cebu | LABORATORY | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22713 | PBSP | tanya_pax@yahoo.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 1337 ASSAY: Xpert MTB/RIF Ultra, COUNT: 498 |  | GX-IV | 808637 | #NAME? | -CHECK MODULE B3 ERRORS, FOUND ERROR 4009 AND 1002 ERRORS OF TEST FROM VIEW RESULTS | #NAME? | #NAME? | #NAME? |  |  | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5463617468812002256/5463617468812002256_signature_28.png | MARIA LUCILLE P. ARONG | https://www.jotform.com/uploads/macarewecare/222927978372067/5463617468812002256/5463617468812002256_signature_31.png
| 2022-12-07 | 2022-12-07 13:53 | Davao Oriental Provincial Hospital - Cateel | Cateel, Davao Oriental | 11 | Davao Oriental | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22712 | PBSP | saiespirituoso@gmail.com | 27.6 | ASSAY: Xpert MTB/RIF, COUNT: 146 ASSAY: Xpert MTB/RIF Ultra, COUNT: 404 | 2022-12-07 | Gx IV - 4D | 807432 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Surface Area, Cartridge Bay, Plunger Rod, Icore, Fan Filters Generate IQR, System Log, Assay Stat | Machine is fully operational | N/A | N/A | N/A | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5462350637016540176/5462350637016540176_signature_28.png | SUZAINE ESPIRITUOSO | https://www.jotform.com/uploads/macarewecare/222927978372067/5462350637016540176/5462350637016540176_signature_31.png
| 2022-12-07 | 2022-12-07 09:46 | Caraga Rural Health Unit | Poblacion, Caraga, Davao Oriental | 11 | Davao Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22711 | PBSP | rhizabenaning06@gmail.com | 26.9 | ASSAY: Xpert MTB/RIF, COUNT: 333 | 2022-12-07 | Gx-IV 2D | 835217 | Module B1/ SN: 788653 for replacement | Module Replacement Perform Preventive Maintenance | Replaced Module B1/ SN:788653 with new Module SN: 663639 | Unit is fully Operational | N/A | 900-0491rs | PL#9546 | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5462350287018526360/5462350287018526360_signature_28.png | Rhiza Mae M. Benaning | https://www.jotform.com/uploads/macarewecare/222927978372067/5462350287018526360/5462350287018526360_signature_31.png
| 2022-12-07 | 2022-12-07 16:25 | Cabanatuan CHO | Cabanatuan Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22710 | DOH | armi_adecer@yahoo.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-12-07 | Genexpert IV | 831592 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality test | Unit is operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5462109492022721647/5462109492022721647_signature_28.png | Arminda A. Adecer | https://www.jotform.com/uploads/macarewecare/222927978372067/5462109492022721647/5462109492022721647_signature_31.png
| 2022-12-07 | 2022-12-07 13:50 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22709 | DOH | abigailgsula@gmail.com | 22 | ASSAY: Xpert MTB/RIF, COUNT: 365 ASSAY: Xpert MTB/RIF Ultra, COUNT: 718 |  | GXIV-D | 831912 | Temperature error module 1 (1002) | Performed Preventive maintenance Performed xpertcheck Performed functionality test | Preventive maintenance and Calibrtion | Unit is operational except A1 and A2 | Failed xpertcheck result for module A1 (HW failed) |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5462017111216841844/5462017111216841844_signature_28.png | Abigail Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5462017111216841844/5462017111216841844_signature_31.png
| 2022-12-07 | 2022-12-06 14:39 | Mati City Health Office | Madang, Mati City, Davao Oriental | 11 | Davao Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22708 | PBSP | cholaboratory@yahoo.com | 27.1 | ASSAY: Xpert MTB/RIF, COUNT: 430 ASSAY: Xpert MTB/RIF Ultra, COUNT: 361 | 2022-12-06 | Gx IV - 2D | 835221 | Module B1: SN 787801 Encounter An Error | Troubleshooting Module B1 | Reset Gateway Board Connections Clean Icore Lens Perform Xpert Check on Module B1 Only | Module B1/SN:787801 for Service Required | Only 1 Module is Working, Module B1 is disabled | XPERTCHECK-CE-5 | N/A | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5461994798683720689/5461994798683720689_signature_28.png | Liezl Anne Alago | https://www.jotform.com/uploads/macarewecare/222927978372067/5461994798683720689/5461994798683720689_signature_31.png
| 2022-12-07 | 2022-12-07 12:53 | Dasmariñas City Health Office II | Area E | 4A | Cavite | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22707 | DOH | hazelmilan0777@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 476 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2059 ASSAY: Xpert HIV-1 Viral Load, COUNT: 46 |  | GX4-D | 838535 | Temperature errors in module 4 | Maintenance machine | Performed Maintenance Performed Xpertcheck on module 4 | 1 module operational, 3 modules for replacement | Xpertcheck: B4/654888 - Failed |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5461987792424533675/5461987792424533675_signature_28.png | Hazel Milan | https://www.jotform.com/uploads/macarewecare/222927978372067/5461987792424533675/5461987792424533675_signature_31.png
| 2022-12-07 | 2022-12-07 12:02 | Tatalon hc | Victory ave. Qc | NCR | NCR 2nd - Quezon City | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22706 | DOH | tatalonlab@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 267 ASSAY: Xpert MTB/RIF Ultra, COUNT: 235 |  | Gx4 | 831610 | A3 continous temp error | Checked unit,cleaned filter,plunger,cartridge bay,performed rexpert check | Pm,rexpert check | A4 disabled reported,A3 754333 failed rexpertcheck |  |  |  | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5461962522725619479/5461962522725619479_signature_28.png | Miguel Joseph C. Aquino | https://www.jotform.com/uploads/macarewecare/222927978372067/5461962522725619479/5461962522725619479_signature_31.png
| 2022-12-07 | 2022-12-07 12:06 | St. Elizabeth Hospital Inc | General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22705 | COMMERCIAL | laboratory@sehi.ph.com | 22.6 | ASSAY: Xpert SARS-CoV-2, COUNT: 7769 ASSAY: Xpert MTB/RIF Ultra, COUNT: 62 | 2022-12-07 | Gx-4 ( 10C) | 110003713 | Module B4 H/W failed due to multiple temperature error | Genexpert machine Preventive Maintenance | Genexpert machine Preventive Maintenance | Gx machine has 3 Operational modules | Module B4 subject for module replacement due to H/W FAILED |  |  | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5461955302916895846/5461955302916895846_signature_28.png | Joana May Simplicio | https://www.jotform.com/uploads/macarewecare/222927978372067/5461955302916895846/5461955302916895846_signature_31.png
| 2022-12-07 | 2022-12-07 12:01 | St. Elizabeth Hospital inc. | General Santos City | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22704 | COMMERCIAL | laboratory@sehi.ph.com | 26.3 | ASSAY: Xpert SARS-CoV-2, COUNT: 7593 ASSAY: Xpert MTB/RIF Ultra, COUNT: 38 | 2022-12-07 | Gx-4 | 831313 | Module C3 multiple temperature error | Genexpert machine Preventive Maintenance | Genexpert machine Preventive Maintenance | Gx machine has 3 Operational modules | Module C3 subject for module replacement due to multiple temperature errors |  |  | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5461952022913425245/5461952022913425245_signature_28.png | Joana May Simplicio | https://www.jotform.com/uploads/macarewecare/222927978372067/5461952022913425245/5461952022913425245_signature_31.png
| 2022-12-06 | 2022-12-06 13:13 | East Rembo Health Center | East Rembo, Makati | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22703 | PBSP | rtmendros@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1179 ASSAY: Xpert MTB/RIF Ultra, COUNT: 145 |  | GX-IV-D | 807591 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, icore and filter Generated reports | Unit is operational. |  |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5461348302819998413/5461348302819998413_signature_28.png | Roxan T. Mendros | https://www.jotform.com/uploads/macarewecare/222927978372067/5461348302819998413/5461348302819998413_signature_31.png
| 2022-12-06 | 2022-12-06 15:03 | East avenue medical center | East ave diliman q.c | NCR | NCR 2nd - Quezon City | Molecular lab | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22702 | MACARE-HBDC | eamcmolecularlab@gmail.com | 24 | ASSAY: Xpert HPV, COUNT: 129 ASSAY: Xpert SARS-CoV-2, COUNT: 558 |  | Gx4 10c | 110014085 | 1002 error b4 210069598 | Check unit,fan,filter,performed selftest error still occut | Checked up | 3 module operational b4 disabled | B4 for reporting |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5461330939883979477/5461330939883979477_signature_28.png | Camille L. Fernando | https://www.jotform.com/uploads/macarewecare/222927978372067/5461330939883979477/5461330939883979477_signature_31.png
| 2022-12-06 | 2022-12-06 13:23 | Philippine heart center | East ave diliman q.c | NCR | NCR 2nd - Quezon City | Blood bank | NO | TROUBLESHOOTING | HAEMONETICS | [5 empty] | 22701 | MACARE-HBDC | baluyotjeric@gmail.com | 0 |  |  | Mcs+ | 01M201 | Crc bios | Performed cpu reset but error still occur, swap cpu board to other unit, error still occur in unit no 2, try install good cpu board no error display | Check up troubleshooting | Unoperational | For replacement of cpu board |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5461330659884502778/5461330659884502778_signature_28.png | Jeric C. Baluyot | https://www.jotform.com/uploads/macarewecare/222927978372067/5461330659884502778/5461330659884502778_signature_31.png
| 2022-12-06 | 2022-12-06 15:20 | Munoz RHU | Munoz Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22700 | DOH | mimdsantiago@gmail.com | 20 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-12-06 | Genexpert IV | 831673 | Module 2 and 4 Temperature Errors | Performed Xpertcheck on B2 and B4 | Performed Maintenance Gathered Reports Performed Xpertcheck on B2 and B4 module Performed Functionality Test | Unit is operational | B2 and B4 passed | Xpertcheck 2 pcs | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5461267148205158183/5461267148205158183_signature_28.png | Geneve Ann Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5461267148205158183/5461267148205158183_signature_31.png
| 2022-12-06 | 2022-12-06 14:05 | Bagumbayan Rural Health Unit | Bagumbayan | 12 | South Cotabato | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22699 | DOH | marizalenoid@yahoo.com | 27.3 | ASSAY: Xpert MTB/RIF, COUNT: 567 ASSAY: Xpert MTB/RIF Ultra, COUNT: 348 | 2022-12-06 | Gx-4 | 838261 | N/A | Perform Genexpert Preventive Maintenance | Perform Genexpert Preventive Maintenance | Perform Genexpert Preventive Maintenance |  |  |  | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5461165571325447012/5461165571325447012_signature_28.png | MA. RIZA R. DIONELA | https://www.jotform.com/uploads/macarewecare/222927978372067/5461165571325447012/5461165571325447012_signature_31.png
| 2022-12-06 | 2022-12-06 13:32 | Bustos Rural Health Unit | Bustos, Bulacan | 3 | Bulacan | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22698 | DOH | ronald_galvez2002@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 239 ASSAY: Xpert MTB/RIF Ultra, COUNT: 308 |  | GX IV | 831874 | Hardware failure on module B4 | Performed xpertcheck on B4 | #NAME? | Unit is oprational. All modules except B4 are functional. | B4 for reporting. |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5461150327471198460/5461150327471198460_signature_28.png | Ronald Galvez | https://www.jotform.com/uploads/macarewecare/222927978372067/5461150327471198460/5461150327471198460_signature_31.png
| 2022-12-06 | 2022-12-06 13:26 | Almanza Uno Health Center | Almanza Uno | NCR | NCR 4th - Las Piñas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22697 | DOH | lpc.socialhygiene@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF, COUNT: 370 ASSAY: Xpert MTB/RIF Ultra, COUNT: 918 ASSAY: Xpert HIV-1 Viral Load, COUNT: 107 |  | GX4-D | 831918 | Temperature Errors | Module replacement | Replaced Modules 1 and 3 Performed Maintenance Performed self test Generate reports | 2 modules are newly replaced, 2 modules for replacement | B1:758515 to 664554 , B3: 783832 to 210052509 |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5461143719725337813/5461143719725337813_signature_28.png | Milcer Pasao Lladoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5461143719725337813/5461143719725337813_signature_31.png
| 2022-12-06 | 2022-12-06 12:38 | Concepcion District Hospital | Concepcion Tarlac | 3 | Tarlac | Labortory | NO | Check up | CEPHEID | [5 empty] | 22696 | PBSP | cdh.laboratorydept@gmail.com | 18 | ASSAY: Xpert MTB/RIF, COUNT: 415 ASSAY: Xpert MTB/RIF Ultra, COUNT: 639 ASSAY: Xpert HIV-1 Viral Load, COUNT: 12 |  | GXIV-D | 807562 | Temperature error on module A1(1002) Error 2037 on module A4 | #NAME? | Preventive maintenance and check-up | Unit is operational except module A1 and A3, A4 under observation |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5461118082211676378/5461118082211676378_signature_28.png | Josette Bondoc | https://www.jotform.com/uploads/macarewecare/222927978372067/5461118082211676378/5461118082211676378_signature_31.png
| 2022-12-06 | 2022-12-06 11:39 | West Rembo Health Center | West Rembo, Makati | NCR | NCR 4th - Makati | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22695 | PBSP | danise_er_19@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 601 ASSAY: Xpert MTB/RIF Ultra, COUNT: 561 |  | GX-IV-D | 831712 | Modules A2, A3, A4 temp error | Checked modules | Performed maintenance Checked errors Generated reports | Unit is operational. | Modules A3 and A4 not operational. |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5461100249225607049/5461100249225607049_signature_28.png | Danis Maebel V. Calimag | https://www.jotform.com/uploads/macarewecare/222927978372067/5461100249225607049/5461100249225607049_signature_31.png
| 2022-12-06 | 2022-12-05 14:39 | Tarlac Cith Health Center X | Tibag Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE CPU Check up | CEPHEID | [5 empty] | 22694 | DOH | rodelie220@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 448 ASSAY: Xpert MTB/RIF Ultra, COUNT: 275 |  | GXIV-D | 831587 | N/A | #NAME? | Preventive maintenance and check-up | Unit is operational |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5461047132216858644/5461047132216858644_signature_28.png | Rodelie Verayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5461047132216858644/5461047132216858644_signature_31.png
| 2022-12-06 | 2022-12-06 09:44 | Jose Reyes Memorial Medical Center | Rizal Ave | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22691 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 774 ASSAY: Xpert SARS-CoV-2, COUNT: 2646 |  | GX16-D | 110013856 | Due for maintenance | Performed Maintenance | Performed Maintenance Performed Self Test Generate Reports | All modules are operational | C2 for observation |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5461016288722278400/5461016288722278400_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5461016288722278400/5461016288722278400_signature_31.png
| 2022-12-06 | 2022-12-06 10:00 | Jose Reyes Memorial Medical Center | Rizal Ave. | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22692 | DOH | covidlabjrrmmc@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 787 ASSAY: Xpert SARS-CoV-2, COUNT: 2671 |  | GX4-D | 831530 | Due for Maintenance | Performed Maintenance | Performed Maintenance Performed Self Test Generate Reports | All modules are operational |  |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5461016288722372604/5461016288722372604_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5461016288722372604/5461016288722372604_signature_31.png
| 2022-12-06 | 2022-12-06 09:57 | Jose Reyes Memorial Medical Center | Rizal Ave | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22693 | MACARE-HBDC | covidlabjrrmmc@gmail.com | 26 | ASSAY: Xpert HPV, COUNT: 167 ASSAY: Xpert SARS-CoV-2, COUNT: 136 |  | GX4-D-10C | 110014099 | Due for maintenance | Performed maintenance | Performed Maintenance Performed Self Test Generate reports | All modules are operational |  |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5461016288727584129/5461016288727584129_signature_28.png | Nelson Mark Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5461016288727584129/5461016288727584129_signature_31.png
| 2022-12-06 | 2022-12-06 09:13 | St.lukes q c | E.rodriguez | NCR | NCR 2nd - Quezon City | Microbiology | NO | Check up | CEPHEID | [5 empty] | 22690 | COMMERCIAL | biomed.qc@stlukes.com.ph | 26 | ASSAY: Xpert C. difficile BT, COUNT: 3 ASSAY: Xpert CT/NG, COUNT: 203 ASSAY: Others, COUNT: 4 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2559 |  | Gx 4 | 839234 | Error 5007 | Cleaned icore,filter,plunger,cartridge | Checked up | Operational ready to use |  |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5461013963223055495/5461013963223055495_signature_28.png | Maria Eliza Apo | https://www.jotform.com/uploads/macarewecare/222927978372067/5461013963223055495/5461013963223055495_signature_31.png
| 2022-12-06 | 2022-12-06 09:27 | San Lazaro Hospital | Quirecada, Sta. Cruz | NCR | NCR 1st - Manila | TB Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22689 | PBSP | acong.pbsp@gmail.com | 26.2 | ASSAY: Xpert MTB/RIF, COUNT: 989 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1249 ASSAY: Xpert HIV-1 Viral Load, COUNT: 402 |  | Gx4D | 707843 | Unit is due for Preventive Maintenance | Preventive Maintenance | Performed Surface Cleaning Performed Filter Cleaning Performed Plunger Rod Cleaning | Unit is operational |  |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5460995817713950447/5460995817713950447_signature_28.png | Ashley Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5460995817713950447/5460995817713950447_signature_31.png
| 2022-12-06 | 2022-12-06 09:16 | San Lazaro Hospital | Quirecada St., Sta. Cruz, | NCR | NCR 1st - Manila | Microbiology TB Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22688 | PBSP | acong.pbsp@gmail.com | 26.7 | ASSAY: Xpert HIV-1 Viral Load, COUNT: 989 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1249 |  | GX4D | 804766 | Unit is due for xpertcheck | Xpert Check and Preventive Maintenance | PerformedXpert Check Performed Surface Cleaning Performed Filter Cleaning Performed Plunger Rod Cleaning | Unit is Operational Xpert check results: 778497/B1 passed 682955/B2 failed 685262/B3 failed 606103/B4 failed |  |  |  | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5460992577712567433/5460992577712567433_signature_28.png | Ashley Ong | https://www.jotform.com/uploads/macarewecare/222927978372067/5460992577712567433/5460992577712567433_signature_31.png
| 2022-12-06 | 2022-12-01 16:58 | Tingloy Rural Health Unit | Brgy 14, Poblacion, Tingloy | 4A | Batangas | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22687 | PBSP | hazelgarcebinay@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 196 ASSAY: Xpert MTB/RIF Ultra, COUNT: 100 | 2022-11-25 | GX4D | 804685 | Unit is due for Module replacement | Module Replacement | Performed Module Replacement B2/614518 into B2/775567 Performed Preventive Maintenance Gather Reports | Unit is Operational |  | Genexepert Module 6C SN: 775567 | PBSP | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5460973637718546984/5460973637718546984_signature_28.png | Hazel Grace B. Sawali | https://www.jotform.com/uploads/macarewecare/222927978372067/5460973637718546984/5460973637718546984_signature_31.png
| 2022-12-05 | 2022-12-05 08:08 | National Tuberculosis Reference Laboratory | Filinvest | NCR | NCR 4th - Muntinlupa | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22686 | PBSP | mspogosa.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 914 ASSAY: Xpert MTB/RIF Ultra, COUNT: 936 |  | GXIV- D | 707512 | Due for Xpertcheck | Performed XpertCheck | Performed Maintenance Performed Xpertcheck Generate reports | 1 module for replacement. 3 modules operational | 695979/A1 - Failed |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5460398811217344008/5460398811217344008_signature_28.png | Maria Guileane Pogosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5460398811217344008/5460398811217344008_signature_31.png
| 2022-12-05 | 2022-12-05 13:58 | South Cotabato Provincial Hospital | Koronadal City | 12 | South Cotabato | Laboratory | YES | REPAIR | CEPHEID | [5 empty] | 22685 | PBSP | anamariermt22@gmail.com | 27.2 | ASSAY: Xpert MTB/RIF, COUNT: 398 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1591 ASSAY: Xpert HIV-1 Viral Load, COUNT: 426 | 2022-12-05 | Gx-4 | 815380 | Cpu problem / Cpu Bios Problem | Change bios Battery Fix Bios Set-up Data Collection - IQR, S. LOG, Assay statistics | Change bios Battery Fix Bios Set-up Data Collection - IQR, S. LOG, Assay statistics | Gx machine has in Good Condition | N/A |  |  | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5460296112313619825/5460296112313619825_signature_28.png | Ana Marie Grace C. Egasan | https://www.jotform.com/uploads/macarewecare/222927978372067/5460296112313619825/5460296112313619825_signature_31.png
| 2022-12-05 | 2022-12-05 13:58 | Sto. Domingo District Hospital | Sto. Domingo District Hospital | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE ON-SITE ASSISTANCE | CEPHEID | [5 empty] | 22684 | DOH | ojie_obejas@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-12-05 | Genexpert IV | 831582 | N/a | N/a | Performed Maintenance Gathered Reports Installed HIV Viral Load Assay Backup and archive test Performed Functionality Test | Unit is operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5460295750718943714/5460295750718943714_signature_28.png | Adora P. Obejas | https://www.jotform.com/uploads/macarewecare/222927978372067/5460295750718943714/5460295750718943714_signature_31.png
| 2022-12-05 | 2022-12-05 12:03 | South Cotabato Provincial Hospital | Koronadal City | 12 | South Cotabato | Molecular Laboratory | YES | CALIBRATION MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22683 | COMMERCIAL | rocaroland@yahoo.com | 22.4 | ASSAY: Xpert SARS-CoV-2, COUNT: 10695 | 2022-12-05 | Gx-16 | 839314 | Module A3, C2, D1, and D4 multiple Temperature Errors | Perform Genexpert machine Preventive Maintenance Perform Xpertcheck/Calibration to Module A3, C2, D1, and D4 for Troubleshooting | Perform Genexpert machine Preventive Maintenance Perform Xpertcheck/Calibration to Module A3, C2, D1, and D4 for Troubleshooting | Gx machine has 13 operational modules | Module A3, C2, D1 and D4 subject for module replacement |  |  | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5460226515564928973/5460226515564928973_signature_28.png | Roland Roca | https://www.jotform.com/uploads/macarewecare/222927978372067/5460226515564928973/5460226515564928973_signature_31.png
| 2022-12-02 | 2022-12-02 16:20 | Munoz RHU | Munoz Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22682 | DOH | mimdsantiago@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-12-02 | Genexpert IV | 831673 | Module 2 and 4 Temperature error | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | 2nd module 3 consecutive errors | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5457803153331127241/5457803153331127241_signature_28.png | Geneve Ann Santiago | https://www.jotform.com/uploads/macarewecare/222927978372067/5457803153331127241/5457803153331127241_signature_31.png
| 2022-12-02 | 2022-12-02 14:42 | Coron Rural Health Unit | Coron, Palawan | 4B | Palawan | TB Dots | YES | MAINTENANCE | CEPHEID | [5 empty] | 22681 | PBSP | lycking@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 359 ASSAY: Xpert MTB/RIF Ultra, COUNT: 334 |  | GXIV-4D | 807612 | Due for maintenance | Preventive Maintenance | #NAME? | Module B3- service required; Module B4- H/W failed |  |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5457727678814184575/5457727678814184575_signature_28.png | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5457727678814184575/5457727678814184575_signature_31.png
| 2022-12-02 | 2022-12-02 14:39 | Culion Sanitarium and General Hospital | Culion, Palawan | 4B | Palawan | Molecular Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22680 | PBSP | culsanlab@gmail.com | 24.2 | ASSAY: Xpert MTB/RIF, COUNT: 75 ASSAY: Xpert MTB/RIF Ultra, COUNT: 50 ASSAY: Xpert SARS-CoV-2, COUNT: 2367 |  | GXIV-2D | 835101 | Module A2 - temperature error | None | #NAME? | All two modules failed XpertCheck- for replacement |  |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5457725508813253841/5457725508813253841_signature_28.png | Rose Eve Calinawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5457725508813253841/5457725508813253841_signature_31.png
| 2022-12-02 | 2022-12-02 10:59 | Culion Sanitarium and General Hospital | Culion, Palawan | 4B | Palawan | Molecular Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22679 | DOH | culsanlab@gmail.com | 24.2 | ASSAY: Xpert MTB/RIF, COUNT: 218 ASSAY: Xpert MTB/RIF Ultra, COUNT: 123 ASSAY: Xpert SARS-CoV-2, COUNT: 5564 | 2022-11-21 | GXIV-4D | 847725 | Module B2 SN:913442 defective | Replace defective module B2 SN: 913442 with new module SN: 210068598 | Preventive Maintenance | All four modules are functional |  | 900-0508, 1 GeneXpert Module-6C SN. 210068598 | G4948 | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5457723178813396875/5457723178813396875_signature_28.png | Rose Eve Calinawan | https://www.jotform.com/uploads/macarewecare/222927978372067/5457723178813396875/5457723178813396875_signature_31.png
| 2022-12-02 | 2022-12-02 14:18 | Nazareth Health Center | Nazareth, Cagayan de Oro City | 10 | Misamis Oriental | TB DOTS | YES | REPAIR | CEPHEID | [5 empty] | 22678 | DOH | mymy.cruz0824@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 686 |  | GeneXpert | 838269 | Stuck Cartridge | Adjust valve drive to proper lining and force out stucked cartridge | Adjust valve drive | Good |  |  |  | Kim Ryan Arrojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5457714883317724058/5457714883317724058_signature_28.png | Michelle Therese Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5457714883317724058/5457714883317724058_signature_31.png
| 2022-12-02 | 2022-12-02 14:08 | San Luis Rural Health Unit | San Luis Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22677 | DOH | mohan_santos@yahoo.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 123 |  | GxIV-D | 831714 | For maintenance | #NAME? | Preventive maintenance | Unit is operational |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5457708455816026043/5457708455816026043_signature_28.png | Mohan Pangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5457708455816026043/5457708455816026043_signature_31.png
| 2022-12-02 | 2022-12-02 13:18 | Lumban Rural Health Unit | Lumban, Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22676 | PBSP | rizalynesacdalan@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 447 ASSAY: Xpert MTB/RIF Ultra, COUNT: 627 |  | GX-IV-D | 838148 | Module B3 Error 2014/4009 | Checked module B3 | Performed maintenance Checked module B3 Generated reports | Module B3 not operational. |  |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5457688570759948681/5457688570759948681_signature_28.png | Rizalyne sacdalan | https://www.jotform.com/uploads/macarewecare/222927978372067/5457688570759948681/5457688570759948681_signature_31.png
| 2022-12-02 | 2022-12-02 11:18 | Kalayaan Rural Health Unit | Kalayaan, Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22675 | DOH | gemmponce11@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 189 ASSAY: Xpert MTB/RIF Ultra, COUNT: 457 |  | GX-IV-D | 821292 | For XpertCheck, barcode scanner not working, printer not powering on | Checked barcode scanner, checked printer | Performed maintenance Tried to reconfigure barcode scanner Checked printer and cables. Generated reports. | Modules A1/792273 and A2/783209 not operational. | Barcode scanner and printer not operational. Unable to perform XpertCheck due to barcode scanner problem. |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5457683750758330394/5457683750758330394_signature_28.png | Gemma M.Ponce | https://www.jotform.com/uploads/macarewecare/222927978372067/5457683750758330394/5457683750758330394_signature_31.png
| 2022-12-02 | 2022-12-02 10:29 | Bureau of Quarantine | Manila, Metro Manila | NCR | NCR 1st - Manila | COVID Laborstory | NO | CALIBRATION | CEPHEID | [5 empty] | 22674 | COMMERCIAL | dklo.lab@quarantine.doh.gov.ph | 28 | ASSAY: Xpert SARS-CoV-2, COUNT: 1137 |  | GX XVI | 844418 | Unit for calibration. | Performed XpertCheck | #NAME? | Unit is operational. 15 modules are functional, D2 is disabled. | All modules except D2 passed calibration. D2 for reporting. |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5457611382518492642/5457611382518492642_signature_28.png | Dulce Karmela Lo | https://www.jotform.com/uploads/macarewecare/222927978372067/5457611382518492642/5457611382518492642_signature_31.png
| 2022-12-02 | 2022-12-02 11:02 | Vicente Sotto Memorial Medical Center | B. Rodriguez St. , Cebu City | 7 | Cebu | Bloodbank | YES | CALIBRATION MAINTENANCE | HAEMONETICS | [5 empty] | 22673 | COMMERCIAL | ertorreonrmt@gmail.com | 24 |  |  | MCS+ | 02A090 | DUE FOR CALIBRATION AND PREVENTIVE MAINTENANCE | CONDUCT CALIBRATION AND PREVENTIVE MAINTENANCE | CALIBRATION AND PREVENTIVE MAINTENANCE | OPERATIONAL |  |  |  | MARK EFREN RAMAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5457595658817519345/5457595658817519345_signature_28.png | Mark Efren Delos Santos RAMAS | https://www.jotform.com/uploads/macarewecare/222927978372067/5457595658817519345/5457595658817519345_signature_31.png
| 2022-12-02 | 2022-12-02 10:47 | Southern Palawan Provincial Hospital | Brooke’s Point, Palawan | 4B | Palawan | Laboratory | YES | REPAIR | CEPHEID | [5 empty] | 22672 | PBSP | frederickmadjad@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 1119 ASSAY: Xpert MTB/RIF Ultra, COUNT: 587 |  | GXIV- 4D | 804271 | Module B1 - XpertCheck Service Required | Replace defective module SN: 662744 with new module SN: 891079 | Preventive Maintenance | All four modules are functional |  |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5457591598811261025/5457591598811261025_signature_28.png | Anates Abon | https://www.jotform.com/uploads/macarewecare/222927978372067/5457591598811261025/5457591598811261025_signature_31.png
| 2022-12-02 | 2022-12-02 09:21 | Tondo Foreshore Super Health Center | Tondo, Manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22671 | MACARE-HBDC | sdaroya.macare@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 36 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1314 ASSAY: Others, COUNT: 89 |  | GX IV | 110001650 | Module A2 (210010360) for replacement | Replaced module A2 by (210077583) | #NAME? | Unit is operational. All 4 modules functional. | NA |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5457537112514543516/5457537112514543516_signature_28.png | Moral Joseph | https://www.jotform.com/uploads/macarewecare/222927978372067/5457537112514543516/5457537112514543516_signature_31.png
| 2022-12-02 | 2022-12-01 11:29 | Plaridel Rural Health Unit IV | Plaridel, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22670 | DOH | yellei0908@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 257 ASSAY: Xpert MTB/RIF Ultra, COUNT: 399 |  | GX IV | 831711 | UNIT FOR MAINTENANCE | CLEANED FAN FILTER | #NAME? | UNIT IS OPERATIONAL. ALL 4 MODULED ARE IN GOOD CONDITION. | NA |  |  | RONNEL BELIZAR | https://www.jotform.com/uploads/macarewecare/222927978372067/5457533022519406923/5457533022519406923_signature_28.png | Ahlelei Hernandez | https://www.jotform.com/uploads/macarewecare/222927978372067/5457533022519406923/5457533022519406923_signature_31.png
| 2022-12-01 | 2022-12-01 15:14 | Kalinga Health Hub | Marikina | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE Setup of GeneXpert system | CEPHEID | [5 empty] | 22669 | PBSP | jasmin.catral@innovationsch.org | 25 | ASSAY: Others, COUNT: N/A |  | GX-IV-D | 805211 | For maintenance | Performed maintenance | Setup the GeneXpert system Performed maintenance Generated reports | Unit is operational. Ups not powering on. | No barcode scanner. |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5456909845431307627/5456909845431307627_signature_28.png | Jasmin Catral | https://www.jotform.com/uploads/macarewecare/222927978372067/5456909845431307627/5456909845431307627_signature_31.png
| 2022-12-01 | 2022-12-01 15:46 | Philippine heart center | East ave diliman q.c | NCR | NCR 2nd - Quezon City | Molecular | NO | CALIBRATION | CEPHEID | [5 empty] | 22668 | MACARE-HBDC | molecularlab@phc.gov.ph | 26 | ASSAY: Xpert MTB/RIF, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 595 ASSAY: Others, COUNT: 14 ASSAY: Xpert SARS-CoV-2, COUNT: 17844 |  | GXXVI | 838655 | For rexpertcheck c4 and d2 | Re xpertcheck | Performed xpert check cleaned plunger ,icore,cartridge bay | B1,c3,c4,d2 disabled, 12 module operational | Failed xpertcheck c4 and d2 |  |  | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5456907791516109502/5456907791516109502_signature_28.png | Ison san jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5456907791516109502/5456907791516109502_signature_31.png
| 2022-12-01 | 2022-12-01 14:24 | Cuyapo RHU | Cuyapo, Nueva Ecija | 3 | Nueva Ecija | Laboratoy | NO | MAINTENANCE | CEPHEID | [5 empty] | 22667 | DOH | shane.verlaine@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-12-01 | Genexpert IV | 831672 | Monitor automatic turn off | Check the wires, condition of the monitor | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | Monitor for observation | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5456856040817135663/5456856040817135663_signature_28.png | Verlaine Ascencion Nervi Verueco-Lico | https://www.jotform.com/uploads/macarewecare/222927978372067/5456856040817135663/5456856040817135663_signature_31.png
| 2022-12-01 | 2022-12-01 14:21 | DATU HALUN SAKILAN MEMORIAL HOSPITAL | LAMION BONGAO, TAWI-TAWI | BARMM | Tawi-Tawi | LABORATORY | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22666 | PBSP | rappaiballaho@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 176 ASSAY: Xpert MTB/RIF Ultra, COUNT: 290 |  | GX-IV | 808643 | -MODULE B2 FOR MODULE REPLACEMENT -MODULE B4 ERROR 4017 | -REPLACE MODULE B2 DEF S/N: 808643 WITH NEW S/N 210044995 -PERFORM PREVENTIVE MAINTENANCE AND CLEANING | #NAME? | #NAME? | #NAME? |  |  | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5456853808816266744/5456853808816266744_signature_28.png | RAPPAI BALLAHO | https://www.jotform.com/uploads/macarewecare/222927978372067/5456853808816266744/5456853808816266744_signature_31.png
| 2022-12-01 | 2022-12-01 13:07 | CHONG HUA HOSPITAL - MANDAUE | MANDAUE CITY, CEBU | 7 | Cebu | BLOOD BANK | NO | CALIBRATION MAINTENANCE | HAEMONETICS | [5 empty] | 22665 | COMMERCIAL | bloodbankmandaue@gmail.com | 22 |  |  | MCS+ | 11J217 | #NAME? | #NAME? | #NAME? | EQUIPMENT IS OPERATIONAL |  |  |  | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5456806788812014073/5456806788812014073_signature_28.png | AIRA GABRIELLE MABINI | https://www.jotform.com/uploads/macarewecare/222927978372067/5456806788812014073/5456806788812014073_signature_31.png
| 2022-12-01 | 2022-12-01 12:47 | Hagonoy Rural Health Unit | Poblacion Hagonoy, Davao Del Sur | 11 | Davao del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22664 | WHO | rhuhagonoy@gmail.com | 28.5 | ASSAY: Xpert MTB/RIF, COUNT: 395 | 2022-12-01 | Gx - IV 4D | 847256 | For Preventive Maintenance | Performed Preventive Maintenence | Clean Surface Area, Cartridge Bay, Plunger Rod, Icore, and fan filters Generate IQR, Assay Stat and System Log | Unit is fully operational | N/A | N/A | N/a | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5456801389718577546/5456801389718577546_signature_28.png | ARNEL CRESAR G. NEROSA | https://www.jotform.com/uploads/macarewecare/222927978372067/5456801389718577546/5456801389718577546_signature_31.png
| 2022-12-01 | 2022-12-01 12:22 | Macabebe Rural Health Unit | Macabebe Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22663 | DOH | salvamari13@yahoo.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 356 ASSAY: Xpert MTB/RIF Ultra, COUNT: 463 |  | GXIV-D | 831893 | Temperature error on modules A2 and A3 | #NAME? | Preventive maintenance | Unit is operational except for modules A2 and A3 |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5456784104018716795/5456784104018716795_signature_28.png | Salva Sese | https://www.jotform.com/uploads/macarewecare/222927978372067/5456784104018716795/5456784104018716795_signature_31.png
| 2022-12-01 | 2022-12-01 12:19 | Oriental Mindoro Provincial Hospital | Calapan City Oriental Mindoro | 4B | Oriental Mindoro | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22662 | PBSP | mariatheresayaco@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 158 ASSAY: Xpert SARS-CoV-2, COUNT: 736 ASSAY: Xpert HIV-1 Viral Load, COUNT: 6 | 2022-12-01 | GXIV | 834446 | Unit for PM. For module replacemeng of Module C4. | Performed PM of GeneXpert. Performed Module Replacement of GeneXpert. | PM and Module Replacement | Unit is operational. 4 working modules. |  | GX Module (210068698) in replacement for defective module (604614 | N/A | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5456783905508349058/5456783905508349058_signature_28.png | Maria Theresa M. Yaco, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5456783905508349058/5456783905508349058_signature_31.png
| 2022-12-01 | 2022-12-01 12:23 | Oriental Mindoro Provincial Hospital | Calapan Ciry Oriental Mindoro | 4B | Oriental Mindoro | Covid Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22661 | PBSP | mariatheresayaco@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 813 ASSAY: Xpert SARS-CoV-2, COUNT: 6557 | 2022-12-01 | GX16 | 842229 | Unit for PM. For module replacement of B1, B3 and C4. | Performed PM of GeneXpert. Performed module replacement. | PM and Module Replacement | Unit is operational. 16 working modules. |  | GX Modules (210068601, 210068595 and 210068343) in replacement for defective modules (884307, 863398 and 681335). | N/A | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5456783115509143345/5456783115509143345_signature_28.png | Maria Theresa M. Yaco,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5456783115509143345/5456783115509143345_signature_31.png
| 2022-11-29 | 2022-11-29 16:26 | Sto. Domingo District Hospital | Sto. Domingo Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22660 | DOH | lhaydeesel@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-11-29 | Genexpert IV | 831582 | N/a | N/a | Performed Maintenance Cleaned Icore Cleaned Cartridge Bay Cleaned Plunger Rod Cleaned Filters Gathered Reports Perfomed Functionality Test | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5455216314529919520/5455216314529919520_signature_28.png | Haydee G.Alejo | https://www.jotform.com/uploads/macarewecare/222927978372067/5455216314529919520/5455216314529919520_signature_31.png
| 2022-11-29 | 2022-11-29 12:41 | Tarlac City Health Center II | Matatalaib Tarlac City | 3 | Tarlac | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22659 | DOH | clarizannebeltran1996@yahoo.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 308 ASSAY: Xpert MTB/RIF Ultra, COUNT: 343 |  | GXIV-D | 831917 | N/A | #NAME? | Preventive maintenance | Unit is operational |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5455063050018371036/5455063050018371036_signature_28.png | Clariza Beltran | https://www.jotform.com/uploads/macarewecare/222927978372067/5455063050018371036/5455063050018371036_signature_31.png
| 2022-11-28 | 2022-11-28 16:17 | National kidney and transplant institute | East ave,diliman quezon city | NCR | NCR 2nd - Quezon City | Molecular | NO | REPAIR | CEPHEID | [5 empty] | 22658 | MACARE-HBDC | molecularlabnkti@gmail.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 120 ASSAY: Others, COUNT: 16 ASSAY: Others, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert MTB/RIF, COUNT: 763 ASSAY: Xpert SARS-CoV-2, COUNT: 13600 ASSAY: Xpert HCV VL, COUNT: 4 ASSAY: Others, COUNT: 9 | 2022-11-28 | GXXVI | 110005587 | For module replacement D2 210028991 - 633953 B4 210028981 - 636241 | Module replacement,cleaned icore,plunger rod, cartridge bay | Module replacement | A4 and d3 reported |  | Genexpert 6 color module | G4955 | Noel floresca/Eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5454406966329104762/5454406966329104762_signature_28.png | Ma. Odessa baretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5454406966329104762/5454406966329104762_signature_31.png
| 2022-11-28 | 2022-11-28 15:49 | Talavera General Hospital | Talavera Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22657 | DOH | pjgtehlab@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 4 | 2022-11-28 | Genexpert IV | 110013773 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is operational | Unit is Operational | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5454316664718125700/5454316664718125700_signature_28.png | John Carlo P. Cuanan | https://www.jotform.com/uploads/macarewecare/222927978372067/5454316664718125700/5454316664718125700_signature_31.png
| 2022-11-28 | 2022-11-28 13:45 | Angono Rural Health Unit | Angono Rizal | 4A | Rizal | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22656 | DOH | aimielyvonnedla@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 210 ASSAY: Xpert MTB/RIF Ultra, COUNT: 529 |  | GX-IV-D | 819864 | For maintenance | Performed maintenance | Performed maintenance Cleaned cartridge bay, plunger rod, filter Generated reports | 3 modules operational. Module A2/686363 H/W Failed already reported. |  |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5454276033318588267/5454276033318588267_signature_28.png | Aimiel Yvonne De Las Alas | https://www.jotform.com/uploads/macarewecare/222927978372067/5454276033318588267/5454276033318588267_signature_31.png
| 2022-11-28 | 2022-11-28 14:18 | Tondo Health Center | Gagalangin Manila | NCR | NCR 1st - Manila | Laboratory | NO | INSTALLATION CALIBRATION | CEPHEID | [5 empty] | 22655 | DOH | ma.lorinamanera@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 491 ASSAY: Xpert MTB/RIF Ultra, COUNT: 982 | 2022-11-28 | GXIV | 831620 | For installation of monitor. For Xpertcheck of module A3 due to errors. | Performed installation of monitor. Performed Xpertcheck on module A3. | Installation of monitor. Xpertcheck | Unit is operational. 2 working modules. Module A2, already reported. Module A3 (702275), service required. |  | HP ProDisplay P19A (sn CNK9430DYM) in replacement for defective monitor (sn CNK83224QC). | N/A. Delivered by PBSP | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5454259727809333431/5454259727809333431_signature_28.png | Patricio G. Mariano, Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5454259727809333431/5454259727809333431_signature_31.png
| 2022-11-28 | 2022-11-28 13:51 | Velasquez Health Center | Varona St. Tondo Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22654 | PBSP | velasquezhealthcenter2054@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 460 ASSAY: Xpert MTB/RIF Ultra, COUNT: 420 |  | GXIV | 812561 | No power. For troubleshooting. | Performed troubleshooting. Checked voltages on power supply. | Troubleshooting. | Unit is not operational. For replacement of power supply. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5454253727801617430/5454253727801617430_signature_28.png | Bona Requiero | https://www.jotform.com/uploads/macarewecare/222927978372067/5454253727801617430/5454253727801617430_signature_31.png
| 2022-11-28 | 2022-11-25 14:05 | Jala-Jala Rural Health Unit | Jala-Jala, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22653 | DOH | abegailjoeycapongga@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 164 ASSAY: Xpert MTB/RIF Ultra, COUNT: 62 |  | GX-IV-D | 820184 | For XpertCheck | Performed XpertCheck | Performed maintenance Run XpertCheck Generated reports | All modules operational. | All modules passed. |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5454242293312789355/5454242293312789355_signature_28.png | Abegail joey | https://www.jotform.com/uploads/macarewecare/222927978372067/5454242293312789355/5454242293312789355_signature_31.png
| 2022-11-28 | 2022-11-28 13:32 | Guagua Rural Health Unit | Guagua Pampanga | 3 | Pampanga | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22652 | DOH | abigailgsula@gmail.com | 18 | ASSAY: Xpert MTB/RIF, COUNT: 381 ASSAY: Xpert MTB/RIF Ultra, COUNT: 692 |  | GXIV-D | 831912 | Multiple errors Modules (A1 and A3) A2 already reported | #NAME? | Preventive maintenance and Calibrtion | Unit is operational except for modules A1:687131 and A3:752683 | A1 and A3 Xpertcheck passed |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5454240328121131040/5454240328121131040_signature_28.png | Abigail G. Sula | https://www.jotform.com/uploads/macarewecare/222927978372067/5454240328121131040/5454240328121131040_signature_31.png
| 2022-11-28 | 2022-11-28 13:15 | Padada Rural Health Unit | Padada, Davao Del Sur | 11 | Davao del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22651 | DOH | torresrmt2016@gmail.com | 27.5 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 154 | 2022-11-28 | GX-IV 4D | 838511 | For Preventive Maintenence | Perfrom Preventive Maintenance | Clean Cartridge Bay, Fan Filters, Icore, Plunger Rod, Generate Report; IQR, Assay Stat, System Log | Unit is fully Operational | N/A | N/A | N/A | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5454225339215834278/5454225339215834278_signature_28.png | RHU PADADA /DWIGHT TORRES | https://www.jotform.com/uploads/macarewecare/222927978372067/5454225339215834278/5454225339215834278_signature_31.png
| 2022-11-28 | 2022-11-28 12:41 | Malapatan Rural Health Unit | Poblacion Malapatan | 12 | Sarangani | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22650 | DOH | jasminrayne21@gmail.com | 29.1 | ASSAY: Xpert MTB/RIF, COUNT: 291 ASSAY: Xpert MTB/RIF Ultra, COUNT: 71 | 2022-11-28 | GX-4 | 838419 | Module B3 Multiple Temperature Errors | Perform Xpertcheck for troubleshooting at module B3 Perform Genexpert machine Preventive Maintenance | Perform Xpertcheck for troubleshooting at module B3 Perform Genexpert machine Preventive Maintenance | Gx machine has 2 operational modules | Module B2 and B3 Subject for module Replacement |  |  | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5454205300711988810/5454205300711988810_signature_28.png | Jasmin Rayne Africa | https://www.jotform.com/uploads/macarewecare/222927978372067/5454205300711988810/5454205300711988810_signature_31.png
| 2022-11-28 | 2022-11-28 12:38 | Pulilan Rural Health Unit | Pulilan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22649 | DOH | celsptcruz19@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 250 ASSAY: Xpert MTB/RIF Ultra, COUNT: 216 |  | GX IV | 831915 | Unit for maintenance | NA | #NAME? | Unit is operational. All 4 modules are working. | NA |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5454198187226681174/5454198187226681174_signature_28.png | Maria Cecilia T. Cruz | https://www.jotform.com/uploads/macarewecare/222927978372067/5454198187226681174/5454198187226681174_signature_31.png
| 2022-11-28 | 2022-11-28 11:49 | De La Salle Health and Sciences Institute | D. Mangubat Drive, Dasmariñas City | 4A | Cavite | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22648 | PBSP | acpabruada@dlshsi.edu.ph | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 793 |  | GXIV-L | 707514 | For Laptop Replacement | Replaced Laptop | Performed maintenance Replaced Laptop Generate Reports | All modules are operational | CPH1019F03472-->CPH0222F32306 |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5454175265808129562/5454175265808129562_signature_28.png | ANGIE C. PABRUADA | https://www.jotform.com/uploads/macarewecare/222927978372067/5454175265808129562/5454175265808129562_signature_31.png
| 2022-11-28 | 2022-11-25 09:40 | Rosario Reyes Health Center | San Andres | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22647 | DOH | moninabatil16@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 930 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1443 |  | GXIV-D | 831843 | Modules have 5007 and temp error | Performed xpertcheck on module 4 | Performed Maintenance Xpertcheck on module 4 | 2 modules for observation, 2 modules for replacement | Xpertcheck: 757669/A4 - Failed |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5454175175806178907/5454175175806178907_signature_28.png | Monina Batil | https://www.jotform.com/uploads/macarewecare/222927978372067/5454175175806178907/5454175175806178907_signature_31.png
| 2022-11-28 | 2022-11-25 16:31 | Santa Rosa RHU | Santa Rosa Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22646 | DOH | diamatrowena11@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-11-25 | Genexpert IV | 831599 | N/a | N/a | Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5454044914817519461/5454044914817519461_signature_28.png | Rowena Diamat | https://www.jotform.com/uploads/macarewecare/222927978372067/5454044914817519461/5454044914817519461_signature_31.png
| 2022-11-25 | 2022-11-25 15:22 | National kidney and transplant institute | East ave diliman quezon city | NCR | NCR 2nd - Quezon City | Molecular | NO | MAINTENANCE | CEPHEID | [5 empty] | 22645 | MACARE-HBDC | molecularlabnkti@gmail.com | 23 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 120 ASSAY: Others, COUNT: 16 ASSAY: Others, COUNT: 1 ASSAY: Xpert HBV VL, COUNT: 8 ASSAY: Xpert MTB/RIF, COUNT: 745 ASSAY: Xpert SARS-CoV-2, COUNT: 13439 |  | GXXVI | 110005587 | A4 210028984 intermittent temp error D3 210028995 intermittent temp error | Performed maintenance,cleaned filter,plunger,cartridge bay,gathered reports | Maintenance | A4,B4,D2,D3 disabled, 12 module operational |  |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5451751311916260536/5451751311916260536_signature_28.png | Ma.Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5451751311916260536/5451751311916260536_signature_31.png
| 2022-11-25 | 2022-11-25 16:03 | NATIONAL KIDNEY AND TRANSPLANT INSTITUTE | East Ave., Diliman, Quezon City | NCR | NCR 2nd - Quezon City | eron dumagieng | NO | MAINTENANCE | CEPHEID | [5 empty] | 22644 | MACARE-HBDC | molecularlabnkti@gmail.com | 25.3 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 600 ASSAY: Others, COUNT: 57 ASSAY: Xpert HBV VL, COUNT: 197 ASSAY: Others, COUNT: 1 ASSAY: Xpert MTB/RIF, COUNT: 682 ASSAY: Xpert SARS-CoV-2, COUNT: 670 ASSAY: Xpert HCV VL, COUNT: 61 |  | GX4D | 838524 | Unit is Due for preventive maintenance | Preventive Maintenance | Performed Surface cleaning Performed Plunger rod cleaning Performed Filter Cleaning Gather Reports | Unit is operational |  |  |  | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5451746341514941070/5451746341514941070_signature_28.png | Ma. Odessa Barretto | https://www.jotform.com/uploads/macarewecare/222927978372067/5451746341514941070/5451746341514941070_signature_31.png
| 2022-11-25 | 2022-11-25 13:56 | Bocaue Rural Health Unit I | Bocaue, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22643 | DOH | msa_28@yahoo.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 310 |  | GX IV | 831691 | Unit for maintenance. | NA | #NAME? | Unit is operational. All 4 modules operational. | Temp. Errors on module A4. A4 in under observation. |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5451651975121664112/5451651975121664112_signature_28.png | Maryson alisuag | https://www.jotform.com/uploads/macarewecare/222927978372067/5451651975121664112/5451651975121664112_signature_31.png
| 2022-11-25 | 2022-11-25 11:37 | Antipolo City Health Office | Antipolo, Rizal | 4A | Rizal | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22642 | PBSP | acho.laboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1611 ASSAY: Xpert MTB/RIF Ultra, COUNT: 6315 ASSAY: Xpert HIV-1 Viral Load, COUNT: 282 | 2022-11-24 | GX-IV-D | 804691 | All modules for replacement | Replaced defective modules | Performed maintenance Replaced defective modules Generated reports | All modules operational. |  | Old: D1/634210 D2/710229 D3/604831 D4/756956 New: D1/746455 D2/804835 D3/714780 D4/874662 | G 4952 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5451642098223741593/5451642098223741593_signature_28.png | Christine Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5451642098223741593/5451642098223741593_signature_31.png
| 2022-11-25 | 2022-11-25 11:54 | Schistosomiasis Research and Training Center | Palo Leyte | 8 | Leyte | Laboratory | YES | INSTALLATION MAINTENANCE | CEPHEID | [5 empty] | 22641 | PBSP | mbesplago.pbsp@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1834 | 2022-02-17 | GXIV-4D | 707521 | Laptop/PC Defective (Old version OS) | -Laptop Replacement (Defective) SN: 36N2SM1 (New) SN: CPH0222F31968 | -Install new Laptop -Install assay (MTB/RIF and MTB/RIF ULTRA) -Archive all test results and retrieve to new laptop -Back up archive files | Good |  | Latitude 5420 (Dell Laptop) | SN: CP0222F31968 | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5451626645579026973/5451626645579026973_signature_28.png | Maria Chrissa E. Villaflor | https://www.jotform.com/uploads/macarewecare/222927978372067/5451626645579026973/5451626645579026973_signature_31.png
| 2022-11-25 | 2022-11-25 12:19 | Bulacan Medical Center | Malolos, Bulacan | 3 | Bulacan | Microbiology | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22640 | PBSP | bmcmicro_lab@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 1109 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1255 ASSAY: Xpert HCV VL, COUNT: 14 ASSAY: Xpert HIV-1 Viral Load, COUNT: 230 |  | GX IV | 805112 | A1 temperature error A2 ultrasonic horn error A4 temperature error | Generated reports for reporting | #NAME? | Unit is operational. Only 1 module available. | A1, A2, A4 have hardware problem. To be reported/followed-up immediately. |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5451598435129898291/5451598435129898291_signature_28.png | Trixie Dela Fuente | https://www.jotform.com/uploads/macarewecare/222927978372067/5451598435129898291/5451598435129898291_signature_31.png
| 2022-11-25 | 2022-11-25 10:45 | Davao Regional Medical Center | Apokon, Tagum City, Davao del Norte | 11 | Davao del Norte | OPD OB | YES | MAINTENANCE | MEDGYN | [5 empty] | 22639 | COMMERCIAL | oconc303@gmail.com | 27 |  | 2022-11-25 | MGC-200 Cryotherapy System | 11263 | For Preventive Maintenance | Perform Preventive Maintenance | Perform Preventive Maintenance Visual inspection and checking of the Cryo Tip Rings | Unit was not able to test its functionality due to lost/misplaced CO2 cylinder Tank | Lost Cylinder CO2 Tank | NA | NA | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5451552946342434051/5451552946342434051_signature_28.png | Drmc | https://www.jotform.com/uploads/macarewecare/222927978372067/5451552946342434051/5451552946342434051_signature_31.png
| 2022-11-25 | 2022-11-25 10:45 | Davao Regional Medical Center | Apokon, Tagum City, Davao del Norte | 11 | Davao del Norte | OPD OB | YES | MAINTENANCE | MEDGYN | [5 empty] | 22638 | COMMERCIAL | oconc303@gmail.com | 27 |  | 2022-11-25 | MGC-200 Cryotherapy System | 11263 | For Preventive Maintenance | Perform Preventive Maintenance | Perform Preventive Maintenance Visual inspection and checking of the Cryo Tip Rings | Unit was not able to test its functionality due to lost/misplaced CO2 cylinder Tank | Lost Cylinder CO2 Tank | NA | NA | Roi Kenneth R Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5451552906348099093/5451552906348099093_signature_28.png | Drmc | https://www.jotform.com/uploads/macarewecare/222927978372067/5451552906348099093/5451552906348099093_signature_31.png
| 2022-11-25 | 2022-11-25 08:44 | Davao Regional Medical Center | Apokon, Tagum City, Davao del Norte | 11 | Davao del Norte | OPD OB | YES | MAINTENANCE | MEDGYN | [5 empty] | 22637 | COMMERCIAL | oconc303@gmail.com | 26 |  | 2022-11-25 | Digital Colposcope - AL 106 | 111AEQ0004 | For preventing maintenance | Perform Preventive Maintenance | Perfrom Preventive Maintenance | Unit cannot be test its functionality since laptop inclusion and tv power adapter was misplaced | Laptop Inclusion and Tv Power Adapter was lost/misplaced | N/a | N/A | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5451535096347649936/5451535096347649936_signature_28.png | DRMC | https://www.jotform.com/uploads/macarewecare/222927978372067/5451535096347649936/5451535096347649936_signature_31.png
| 2022-11-25 | 2022-11-25 09:41 | Philippine Orthopedic Center | Banawe Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22636 | PBSP | cmblanceta@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 824 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1194 | 2022-11-24 | GXIV | 803088 | Unit for PM. For module replacement of Module 4. For installation of new gatewat board. | Performed PM of GeneXpert. Performed Module Replacement on Module 4. Performed installation of gateway board. | PM, Module replacement and installation of gateway board | Unit is operational. |  | GX Module (626989), replaced defective module (657586). Gateway board kit (6104S20460031) | G 4953 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5451523549914752305/5451523549914752305_signature_28.png | Stella Maree C. Bernardo | https://www.jotform.com/uploads/macarewecare/222927978372067/5451523549914752305/5451523549914752305_signature_31.png
| 2022-11-25 | 2022-11-24 15:14 | Tropical disease foundation | Amorsolo st. Makati city | NCR | NCR 4th - Makati | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22635 | COMMERCIAL | rrjflores@tdf.org.ph | 24 | ASSAY: Xpert MTB/RIF, COUNT: 854 ASSAY: Xpert MTB/RIF Ultra, COUNT: 559 | 2022-11-24 | Gx4 | 812560 | For xpert check | Cleaned unit,filter,icore,cartridge bay,plunger,performed xpert check | Pm, xpertcheck | All module passed |  | Xpertcheck ce-5 | 94356 | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5451512829881583940/5451512829881583940_signature_28.png | Rupert Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5451512829881583940/5451512829881583940_signature_31.png
| 2022-11-25 | 2022-11-24 11:20 | Tropical disease foundation | Amorsolo st makati city | NCR | NCR 4th - Makati | Laboratory | YES | REPAIR | CEPHEID | [5 empty] | 22634 | COMMERCIAL | rrjflores@tdf.org.ph | 28 | ASSAY: Xpert MTB/RIF, COUNT: 476 ASSAY: Xpert MTB/RIF Ultra, COUNT: 228 | 2022-11-24 | GX4 r1 | 705175 | For module replacement A2 607603 - 697695 | Peformed module replacement,cleaned unit,plunger,icore,cartridge bay | Repair,pms | Operational |  | 900 0491rs | 94356 | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5451512319882267660/5451512319882267660_signature_28.png | Rupert Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5451512319882267660/5451512319882267660_signature_31.png
| 2022-11-25 | 2022-11-25 09:37 | Cordlife Philippines | Commonwealth Ave, Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | Protocol update | CYTIVA | [5 empty] | 22633 | COMMERCIAL | julius.biteng@cordlife.com | 25 |  |  | Sepax 2 | 6220.4122 | None, UCB and UCBHES protocols for update | Updated UCB and UCBHES protocols | Updated UCB from v318 to v338 Updated UCBHES from v314 to v316 | Unit is operational |  |  |  | Patrick Los Baños | https://www.jotform.com/uploads/macarewecare/222927978372067/5451497791835883746/5451497791835883746_signature_28.png | Julius Marion Biteng | https://www.jotform.com/uploads/macarewecare/222927978372067/5451497791835883746/5451497791835883746_signature_31.png
| 2022-11-25 | 2022-11-23 10:50 | De La Salle Health And Sciences Institute | D. Mangubat, Dasmariñas | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22632 | CONSORTIUM | acpabruada@dlshsi.edu.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 760 |  | GXIV-D | 812380 | Gx Modules cannot be detected | Reset gateway board | Performed Maintenance Reset gateway board Check connections Generate reports | 3 modules are operational, 1 module to be reported |  |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5451478660412595959/5451478660412595959_signature_28.png | Angelita C. Pabruada | https://www.jotform.com/uploads/macarewecare/222927978372067/5451478660412595959/5451478660412595959_signature_31.png
| 2022-11-25 | 2022-11-22 12:10 | Almanza Uno Health Center | Almanza Uno, Las Piñas | NCR | NCR 4th - Las Piñas | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22631 | DOH | milcerpl@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 407 ASSAY: Xpert MTB/RIF Ultra, COUNT: 918 ASSAY: Xpert HIV-1 Viral Load, COUNT: 107 |  | GXIV-D | 831918 | Modules 2 and 4 have temperature errors | Xpertcheck on reported modules | Performed Maintenance XpertCheck modules 2 and 4 Generate Reports | 4 modules for replacement. | Xpertcheck: B2/635070 , B4/758518 -Failed |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5451478660416800013/5451478660416800013_signature_28.png | MILCER P. LLADOC | https://www.jotform.com/uploads/macarewecare/222927978372067/5451478660416800013/5451478660416800013_signature_31.png
| 2022-11-24 | 2022-11-24 17:10 | THE LOVE YOURSELF | 715, Shaw Boulevard | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22630 | MACARE-HBDC | laboratory@loveyourself.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 616 ASSAY: Xpert HIV-1 Viral Load, COUNT: 552 ASSAY: Xpert CT/NG, COUNT: 631 | 2022-11-14 | GX4L | 807112 | Unit is due for Module Replacement | Module Replacement | Performed Module Replacement Performed Preventive Maintenance Gather Reports | Unit is operational |  | GX modules 882201 901520 910995 | 9344 | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5450906798518182612/5450906798518182612_signature_28.png | CALAMA-AN, MAVIENE | https://www.jotform.com/uploads/macarewecare/222927978372067/5450906798518182612/5450906798518182612_signature_31.png
| 2022-11-24 | 2022-11-24 13:34 | Gapan RHU | San Vicente, Gapan Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE Cpu Installation | CEPHEID | [5 empty] | 22629 | PBSP | emmaliwaggutierrez@gmail.com | 32 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-11-24 | Genexpert IV | 814182 | N/a | N/a | Installed Genexpert Computer Performed Maintenance Gathered Reports Performed Functionality Test | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5450781131219369196/5450781131219369196_signature_28.png | Emma L Gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5450781131219369196/5450781131219369196_signature_31.png
| 2022-11-24 | 2022-11-24 13:28 | Orani Rural Health Unit | Orani Bataan | 3 | Bataan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22628 | DOH | belinda_guinto@yahoo.com | 19 | ASSAY: Xpert MTB/RIF, COUNT: 220 ASSAY: Xpert MTB/RIF Ultra, COUNT: 141 |  | GXIV-D | 831668 | Error 2008 on module 1 | #NAME? | Preventive maintenance | Unit is operational |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5450776532817930608/5450776532817930608_signature_28.png | Belinda Guinto | https://www.jotform.com/uploads/macarewecare/222927978372067/5450776532817930608/5450776532817930608_signature_31.png
| 2022-11-24 | 2022-11-24 12:11 | Davao Regional Medical Center | Apokon, Tagum City, Davao del Norte | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22627 | PBSP | covidlab@drmc.doh.gov.ph | 25.7 | ASSAY: Xpert MTB/RIF Ultra, COUNT: N/A | 2022-11-24 | GX - XVI | 838756 | Module A3 for replacement | Module replacement | Replaced the defective module s/n 866561 with module s/n 674488 | Machine not operational, waiting for powersupply unit replacement parts | Module s/n 674488 is used as service module to GX-IV s/n 834626 | 900-0491rs | N/a | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5450734367846547215/5450734367846547215_signature_28.png | Carlo Milan Famoso | https://www.jotform.com/uploads/macarewecare/222927978372067/5450734367846547215/5450734367846547215_signature_31.png
| 2022-11-24 | 2022-11-24 10:10 | Pililla Rural Health Unit | Pililla, Rizal | 4A | Rizal | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22626 | DOH | tbdotspililla@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 236 ASSAY: Xpert MTB/RIF Ultra, COUNT: 231 | 2022-09-29 | GX-IV-D | 821293 | For XpertCheck | Performed XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. All modules passed. |  | XpertCheck kit | DR# 9219 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5450680351787452974/5450680351787452974_signature_28.png | Karren Abegail R. Olaya | https://www.jotform.com/uploads/macarewecare/222927978372067/5450680351787452974/5450680351787452974_signature_31.png
| 2022-11-24 | 2022-11-24 09:33 | Masambong Health Center | Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION | CEPHEID | [5 empty] | 22625 | PBSP | cocristy987@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 25 | 2022-11-24 | GXIV | 814160 | For installation of CPU. | Performed installation of CPU | Installation of CPU. | Unit is operational. Pulled out back-up | Pulled out back-up unit (CPH0420F08151) | GX-D-M70 ¦ GeneXpert Desktop (Lenovo ThinkCentre M70)1unit CPH1121F28199 | G 4951 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5450637039918808464/5450637039918808464_signature_28.png | Cristy cortezano | https://www.jotform.com/uploads/macarewecare/222927978372067/5450637039918808464/5450637039918808464_signature_31.png
| 2022-11-23 | 2022-11-09 14:36 | Lung center of the philippines | Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22624 | MACARE-HBDC | lcpmicrolab@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 124 ASSAY: Xpert MTB/RIF Ultra, COUNT: 451 ASSAY: Xpert SARS-CoV-2, COUNT: 1034 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 | 2022-11-09 | GX 4 | 838837 | B4 h/failed for rexpert check | Performed re xpert check | Cleaned unit,filter plunger,cartridge bay,icore | 3 modules operation, b4 780860 failed xpert check disabled h/w failed |  | Xpertcheck kits | N/a | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5437861826598500563/5437861826598500563_signature_28.png | Lawrence Raquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5437861826598500563/5437861826598500563_signature_31.png
| 2022-11-23 | 2022-11-23 15:55 | Philippine heart center | East ave.diliman,quezon city | NCR | NCR 2nd - Quezon City | Molecular | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22623 | MACARE-HBDC | molecularlab@phc.gov.ph | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 29 ASSAY: Xpert MTB/RIF Ultra, COUNT: 576 ASSAY: Others, COUNT: 14 ASSAY: Xpert SARS-CoV-2, COUNT: 18026 |  | GXXVI | 838655 | B1 210068311 H/W failed,C4 628HW/failed D2 653968 TEMP ERROR 1002 | Gathered report, restart gx system,checked unit | Checked up | C3,C4,D2,B1 disabled, 12 module operational | B1 and C3 reported already |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5450092839884381395/5450092839884381395_signature_28.png | Ison San Jose | https://www.jotform.com/uploads/macarewecare/222927978372067/5450092839884381395/5450092839884381395_signature_31.png
| 2022-11-23 | 2022-11-23 13:27 | Batis Health Center | Batis, San Juan | NCR | NCR 2nd - San Juan | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22622 | DOH | almacenaio16@gmai.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 846 ASSAY: Xpert MTB/RIF Ultra, COUNT: 706 ASSAY: Xpert HIV-1 Viral Load, COUNT: 62 |  | GX-IV-D | 831598 | Module A1/754112 has errors 1001/1002/2014 | Run XpertCheck on module A1 | Performed maintenance Generated reports Run XpertCheck on module A1 | Unit is operational. Module A1/754112 failed in XpertCheck. | Module A1/754112 and A3/677100 not operational |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5449949874113823426/5449949874113823426_signature_28.png | Rosario Almacen | https://www.jotform.com/uploads/macarewecare/222927978372067/5449949874113823426/5449949874113823426_signature_31.png
| 2022-11-23 | 2022-11-23 13:52 | Dingalan Community Hospital | Caragsacan, Dingalan, Aurora | 3 | Aurora | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22621 | DOH | bnjlvrz9@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-11-23 | Genexpert IV | 831621 | N/a | N/a | Performed Maintenance Gathered Reports Performed Xpertcheck on all modules Performed Functionality test | Unit is Operational | N/a | Xpertcheck Kit | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5449929446116014364/5449929446116014364_signature_28.png | Benjo A. Alvarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5449929446116014364/5449929446116014364_signature_31.png
| 2022-11-23 | 2022-11-23 12:11 | Marulas 3S Health Center | Marulas, Valenzuela | NCR | NCR 3rd - Valenzuela | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22620 | DOH | melissamartin0828@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 608 ASSAY: Xpert MTB/RIF Ultra, COUNT: 620 |  | GX IV | 831696 | Unit for maintenance | NA | #NAME? | Unit is operational. All 4 modules working. | No room air-condition. Advised end user to use external fan while waiting for air-condition repair. |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5449862139615082142/5449862139615082142_signature_28.png | Melissa M. Galit | https://www.jotform.com/uploads/macarewecare/222927978372067/5449862139615082142/5449862139615082142_signature_31.png
| 2022-11-23 | 2022-11-23 12:03 | Tupi Rural Health Unit | Tupi | 12 | South Cotabato | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22619 | PBSP | pedzella@gmail.com | 23.4 | ASSAY: Xpert MTB/RIF, COUNT: 200 ASSAY: Xpert MTB/RIF Ultra, COUNT: 350 | 2022-11-23 | Gx-4 (2 modules) | 835184 | N/A | Genexpert machine Calibration Genexpert machine Preventive Maintenance | Genexpert machine Calibration Genexpert machine Preventive Maintenance | All modules are Operational | Module B2 passed the annual calibration |  | Xpertcheck CE-5 | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5449861815659116398/5449861815659116398_signature_28.png | Ella Esmailene S. Pedregosa | https://www.jotform.com/uploads/macarewecare/222927978372067/5449861815659116398/5449861815659116398_signature_31.png
| 2022-11-23 | 2022-11-23 11:28 | Abucay Rural Health Unit | Abucay Bataan | 3 | Bataan | Laboratory | YES | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22618 | DOH | parreraledelina22176@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 232 ASSAY: Xpert MTB/RIF Ultra, COUNT: 820 |  | GXIV-D | 831878 | Temperature error Module 2 Re-expertcheck for module 4 | #NAME? | Preventive Maintenance and Calibration | Unit is operational except A4 | A2 passed Xpertcheck A4 failed xpertcheck with serial number 757994 |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5449852141915665644/5449852141915665644_signature_28.png | Ledelina Parrera | https://www.jotform.com/uploads/macarewecare/222927978372067/5449852141915665644/5449852141915665644_signature_31.png
| 2022-11-22 | 2022-11-22 13:00 | Tondo Health Center | Gagalangin Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22617 | DOH | ma.lorinamanera@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 2 ASSAY: Xpert MTB/RIF, COUNT: 499 ASSAY: Xpert MTB/RIF Ultra, COUNT: 958 |  | GXIV | 831620 | Unit for PM. | Performed Preventive Maintenance. Generated reports for Module A3, temperature error. | Preventive Maintenance. | Unit is operational. 2 modules operational. Defective A2 (678005), for reporting of A3 (702275) with temp error. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5449037157808338426/5449037157808338426_signature_28.png | Lontoc, Ma. Lorina M. | https://www.jotform.com/uploads/macarewecare/222927978372067/5449037157808338426/5449037157808338426_signature_31.png
| 2022-11-22 | 2022-11-22 12:36 | Maco Rural Health Unit | Binuangan, Maco, Davao De Oro | 11 | Davao de Oro | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22616 | DOH | kaiprado26@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 344 ASSAY: Xpert MTB/RIF Ultra, COUNT: 249 | 2022-11-22 | Gx-IV 4D | 838500 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Cartridge Bay, replace fan filters, Clean Plunger Rod, Clean Icore Generate Reports -IQR, System Log, Assay Stat | Unit is fully operational | None | N/A | N/a | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5449020364373161699/5449020364373161699_signature_28.png | Edlien Kai Prado | https://www.jotform.com/uploads/macarewecare/222927978372067/5449020364373161699/5449020364373161699_signature_31.png
| 2022-11-22 | 2022-11-22 11:10 | Atang dela Rama Health Center | Gagalangin Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22615 | DOH | monskie_22@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 515 ASSAY: Xpert MTB/RIF Ultra, COUNT: 776 |  | GXIV | 831644 | Unit for PM. | Performed Preventive Maintenance. | Preventive Maintenance. | Unit is operational. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5448999598581453197/5448999598581453197_signature_28.png | Monaliza Taruc | https://www.jotform.com/uploads/macarewecare/222927978372067/5448999598581453197/5448999598581453197_signature_31.png
| 2022-11-22 | 2022-11-22 10:29 | Rosales Rural Health Unit | Rosales Pangasinan | 1 | Pangasinan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22614 | DOH | allan.vinluan2013@gmail.com | 20 | ASSAY: Xpert MTB/RIF, COUNT: 508 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 |  | GXIV-D | 838472 | For Maintenance | #NAME? | Preventive Maintenance | Unit is opertional |  |  |  | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5448991718615225399/5448991718615225399_signature_28.png | Allan B. Vinluan | https://www.jotform.com/uploads/macarewecare/222927978372067/5448991718615225399/5448991718615225399_signature_31.png
| 2022-11-21 | 2022-11-21 13:47 | Davao Chest Center | Villa Abrille St, Davao City | 11 | Davao del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22613 | PBSP | dcclabdept@gmail.com | 24.7 | ASSAY: Xpert MTB/RIF, COUNT: 481 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1407 | 2022-11-21 | Gx IV - 4D | 804835 | For Preventive Maintenance | Perform Preventive Maintenance | Clean Cartridge Bay, Plunger Rod, Fan Filter, Icore Generate Installation Qualification Report, Assay Statistics Report, System Log | Unit is fully operational | N/A | N/A | N/A | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5448289739757227418/5448289739757227418_signature_28.png | Maria Lisa M. Paras, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5448289739757227418/5448289739757227418_signature_31.png
| 2022-11-21 | 2022-11-21 14:28 | Tarlac City Helth Center X | Tibag Tarlac City | 3 | Tarlac | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22612 | DOH | jayteecornelius@gmail.com | 18 | ASSAY: Xpert MTB/RIF, COUNT: 471 ASSAY: Xpert MTB/RIF Ultra, COUNT: 246 | 2022-11-21 | GXIV-D | 831587 | #NAME? | #NAME? | Troubleshooting | Unit is operational but under observation | Kernel power ( Blue screen of death) | N/A | N/A | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5448236246118545742/5448236246118545742_signature_28.png | Jaytee Cornelius Bachiller | https://www.jotform.com/uploads/macarewecare/222927978372067/5448236246118545742/5448236246118545742_signature_31.png
| 2022-11-21 | 2022-11-21 14:19 | Burauen District Hospital | Burauen Leyte | 8 | Leyte | TB-DOTS | YES | MAINTENANCE | CEPHEID | [5 empty] | 22611 | DOH | buraeunlab@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: |  | GXIV-4D | 838551 | N/A | Change setup position for lower temperature results | Preventive Maintenance | Good |  |  |  | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5448218325845912667/5448218325845912667_signature_28.png | Mary Ann C. Sudario | https://www.jotform.com/uploads/macarewecare/222927978372067/5448218325845912667/5448218325845912667_signature_31.png
| 2022-11-21 | 2022-11-21 13:56 | Iligan Society of Internists | Iligan City | 10 | Lanao del Norte | Laboratory | YES | Monitor Replacement | CEPHEID | [5 empty] | 22610 | PBSP | sdramayan@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2773 ASSAY: Xpert HIV-1 Viral Load, COUNT: 43 | 2022-09-19 | GeneXpert | 805090 | Defective Monitor | Replaced Monitor | Monitor Replacement Preventive maintenance Xpertcheck Calibration | Good | All modules passed the calibration | GeneXpert Monitor Display 21.5" ThinkVision S22e-20 Flat Panel Monitor Serial No. V5CYY987 | G 4715 | Kim Ryan Arrojado | https://www.jotform.com/uploads/macarewecare/222927978372067/5448200761411114162/5448200761411114162_signature_28.png | Sandra D. Ramayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5448200761411114162/5448200761411114162_signature_31.png
| 2022-11-21 | 2022-11-18 13:44 | Imus City Health Office II | Plaridel Subd., Imus | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22608 | DOH | amielynnmangulabnan1217@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 327 ASSAY: Xpert MTB/RIF Ultra, COUNT: 747 |  | GXIV-D | 821284 | Module 3 with Temperature Errors | Xpertcheck in module 3 | Performed Maintenance Xpertcheck on Reported Module Generate Reports | 2 modules are operational, 1 module for replacement and 1 module for observation | Xperrcheck : A3/730981-Passed |  |  | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5448182851215749837/5448182851215749837_signature_28.png | Amielynn Mangulabnan | https://www.jotform.com/uploads/macarewecare/222927978372067/5448182851215749837/5448182851215749837_signature_31.png
| 2022-11-21 | 2022-11-15 09:33 | Fabella Memorial Hospital | Tayuman | NCR | NCR 1st - Manila | Molecular | NO | MAINTENANCE | CEPHEID | [5 empty] | 22607 | MACARE-HBDC | fabella.mollabgene@gmail.com | 26 | ASSAY: Xpert SARS-CoV-2, COUNT: 1500 |  | GXIV-L | 834953 | Due for Maintenance | Maintenance GX | Performed Maintenance. Generate Reports. | 3 modules are operational, 1 module is defective |  |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5448182851211394220/5448182851211394220_signature_28.png | Christopher Panti | https://www.jotform.com/uploads/macarewecare/222927978372067/5448182851211394220/5448182851211394220_signature_31.png
| 2022-11-21 | 2022-11-17 11:36 | Kawit Rural Health Unit | Kawit | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22609 | DOH | Alambraaljhon0@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 248 ASSAY: Xpert MTB/RIF Ultra, COUNT: 626 |  | GXIV-D | 820705 | Module 4 with temp errors | Xpertcheck reported module | Performed Maintenance Performed Xpertcheck on reported module | 3 modules operational, A4 for replacement | Xpertcheck: A4/723150 - Failed |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5448182851216252195/5448182851216252195_signature_28.png | Aljhon G. Alambra | https://www.jotform.com/uploads/macarewecare/222927978372067/5448182851216252195/5448182851216252195_signature_31.png
| 2022-11-19 | 2022-11-18 13:22 | Libmana rhu | Libmanan camarines sur | 5 | Camarines Sur | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22606 | PBSP | mdelizo.megaop3@gmail.com | 30.5 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 534 | 2022-11-18 | Gx4 | 805260 | Gateway defective,for module replacement b3 684618 - 693766 | Performed maintenance,cleaned plunger,filter,cartridge bay,icore,back up gxx file | Pm, module replacement | Operational |  | Module 900 gx and gateway | 0 | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5446740269887629437/5446740269887629437_signature_28.png | Mary Rose C. Delizo | https://www.jotform.com/uploads/macarewecare/222927978372067/5446740269887629437/5446740269887629437_signature_31.png
| 2022-11-18 | 2022-11-18 16:53 | St Lukes Extension Clinic | Bocobo St, Manila | NCR | NCR 1st - Manila | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22605 | COMMERCIAL | laboratory@slec.ph | 24 | ASSAY: Others, COUNT: |  | Genexpert Infinity 48s | 809643 | Module F3 and H1 experiencing hardware failure | Performed troubleshooting | Xpertcheck troubleshooting | Unit is operational, modules F1, F3, H1 are in hardware failure |  |  |  | Patrick Los Baños | https://www.jotform.com/uploads/macarewecare/222927978372067/5445712906227408501/5445712906227408501_signature_28.png | Ma. Sarah R. Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5445712906227408501/5445712906227408501_signature_31.png
| 2022-11-18 | 2022-11-18 15:14 | Libertad Main Health Center | 5th Street, Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22604 | DOH | suzzettekho34@gmail.com | 27.9 | ASSAY: Xpert MTB/RIF, COUNT: 448 ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 | 2022-11-14 | GX - 4 | 838554 | Module B2 with SN: 862237 was found E1002 | troubleshooting | Replacing B1 with SN: 877068 to SN: 669367 Preventive Maintenance Troubleshooting | 3 modules working, 1 modules for observation | None | GeneXpert Module 6- colors | G 4934 | Engr. Faizar Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5445696056318787753/5445696056318787753_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/5445696056318787753/5445696056318787753_signature_31.png
| 2022-11-18 | 2022-11-18 15:14 | Libertad Main Health Center | 5th Street, Brgy. Libertad, Butuan City | 13 | Agusan del Norte | RTDL | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22603 | DOH | suzzettekho34@gmail.com | 27.9 | ASSAY: Xpert MTB/RIF, COUNT: 448 ASSAY: Xpert MTB/RIF Ultra, COUNT: 831 | 2022-11-14 | GX - 4 | 838554 | Module B2 with SN: 862237 was found E1002 | troubleshooting | Replacing B1 with SN: 877068 to SN: 669367 Preventive Maintenance Troubleshooting | 3 modules working, 1 modules for replacement | None | GeneXpert Module 6- colors | G 4934 | Engr. Faizar Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5445657435208547019/5445657435208547019_signature_28.png | Suzzette Kho | https://www.jotform.com/uploads/macarewecare/222927978372067/5445657435208547019/5445657435208547019_signature_31.png
| 2022-11-18 | 2022-11-18 15:19 | Santa Rosa RHU | Santa Rosa Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22602 | DOH | fcm_vizconde@yahoo.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-11-18 | Genexpert aiV | 831599 | N/a | N/a | Performed Maintenance Gathered Reports Perform Self Test Perform Functionality Test | Unit is operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5445655736715362509/5445655736715362509_signature_28.png | Flor Carmina Vizconde | https://www.jotform.com/uploads/macarewecare/222927978372067/5445655736715362509/5445655736715362509_signature_31.png
| 2022-11-18 | 2022-11-18 14:19 | Eastern Pangasinan District Hospital | Pangasinan | 1 | Pangasinan | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22601 | PBSP | epdhlaboratory@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: NA | 2022-11-14 | GX IV | 815394 | #NAME? | #NAME? | #NAME? | Unit is operational. D1, D2, D3 are functional. | Waiting for module D4's replacement module. | GeneXpert Module SN: 678145, 716521 Lenovo Laptop SN: CPH1019F03266 | G 4937 | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5445618665116251372/5445618665116251372_signature_28.png | Fergyl Aries Escaño | https://www.jotform.com/uploads/macarewecare/222927978372067/5445618665116251372/5445618665116251372_signature_31.png
| 2022-11-18 | 2022-11-18 14:01 | The Medical City | Ortigas Ave. Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22600 | COMMERCIAL | moleculardx@themedicalcity.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 7 ASSAY: Others, COUNT: 16 ASSAY: Others, COUNT: 3 ASSAY: Xpert HBV VL, COUNT: 85 ASSAY: Xpert HPV, COUNT: 170 ASSAY: Xpert MTB/RIF, COUNT: 144 ASSAY: Xpert MTB/RIF Ultra, COUNT: 2064 ASSAY: Xpert HCV VL, COUNT: 29 ASSAY: Xpert HIV-1 Viral Load, COUNT: 259 | 2022-11-14 | GX-IV-L | 802630 | Module A2 and A4 for replacement | Replaced module A2 and A4 | Performed maintenance Replaced defective module Generated reports | All modules operational. |  | Old: A2/698743 A4/627869 New: A2/210070909 A4/210071314 | DR# 9345 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5445610744314037575/5445610744314037575_signature_28.png | Stephanie Obida | https://www.jotform.com/uploads/macarewecare/222927978372067/5445610744314037575/5445610744314037575_signature_31.png
| 2022-11-18 | 2022-11-18 14:01 | The Medical City | Ortigas Ave., Pasig City | NCR | NCR 2nd - Pasig | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22599 | COMMERCIAL | moleculardx@themedicalcity.com | 25 | ASSAY: Xpert BCR-ABL Ultra, COUNT: 3 ASSAY: Xpert CT/NG, COUNT: 8 ASSAY: Xpert HBV VL, COUNT: 15 ASSAY: Xpert HPV, COUNT: 57 ASSAY: Xpert MTB/RIF, COUNT: 53 ASSAY: Xpert MTB/RIF Ultra, COUNT: 189 ASSAY: Xpert HCV VL, COUNT: 12 ASSAY: Xpert HIV-1 Viral Load, COUNT: 85 | 2022-11-14 | GXIV | 802117 | Unit for PM and for Module Replacement of A1 and A3. | Performed PM. Performed Module Replacement on A1 and A3. | Preventive Maintenance and Module Replacement. | Unit is Operational. |  | GX Module. DEF A1 (624081), NEW A1 (210071653). DEF A3 (637349), NEW A3 (210071792) | DR#9345 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5445610025719163292/5445610025719163292_signature_28.png | Stephanie Obida | https://www.jotform.com/uploads/macarewecare/222927978372067/5445610025719163292/5445610025719163292_signature_31.png
| 2022-11-18 | 2022-11-17 15:31 | Philippine Red Cross - Sta. Cruz | Sta. Cruz Laguna | 4A | Laguna | Blood bank | NO | TROUBLESHOOTING | HAEMONETICS | [5 empty] | 22598 | MACARE-HBDC | lagunablood@redcross.org.ph | 25 |  |  | MCS- | 02J046 | Centrifuge not working | Troubleshooting of centrifuge | Remove centrifuge for troubleshooting Checked for debris Checked internal connections | Unit is not operational. For centrifuge replacement |  |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5445604740027530097/5445604740027530097_signature_28.png | Hiltor Jonh M. Abustan | https://www.jotform.com/uploads/macarewecare/222927978372067/5445604740027530097/5445604740027530097_signature_31.png
| 2022-11-18 | 2022-11-18 13:14 | Paniqui Rural Health Unit | Paniqui Tarlac | 3 | Tarlac | Laboratory | NO | MAINTENANCE TROUBLESHOOTING Check-up | CEPHEID | [5 empty] | 22597 | PBSP | forylab76@yahoo.com | 18 | ASSAY: Xpert MTB/RIF, COUNT: 380 ASSAY: Xpert MTB/RIF Ultra, COUNT: 242 | 2022-11-19 | GXIV-D | 812792 | Module 1 Hardware failure | Performed PM Perforned Self test Performed gx reliability test Generate reports Performed functionality test | Troubleshooting Check-up | Unit is operational except module 1 | Deffective module 1 for reporting | N/A | N/A | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5445582598328379310/5445582598328379310_signature_28.png | Fortune Abutin | https://www.jotform.com/uploads/macarewecare/222927978372067/5445582598328379310/5445582598328379310_signature_31.png
| 2022-11-18 | 2022-11-18 12:08 | Tabaco CHU I | Tobaco Municipality Hall, Tobaco City, Albay | 5 | Albay | Laboratory | NO | Software Upgrade | CEPHEID | [5 empty] | 22596 | PBSP | elyjoy17rmt@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 69 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 |  | Genexpert 4D | 804257 | Unit is due for Software Upgrade | Intalling updated software of GXDX | Install Genexpert Software Version 6.4 Import Reliability Test gxx file Install Xpertcheck Version 1.5a | Unit is operational |  |  |  | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5445542437116381153/5445542437116381153_signature_28.png | Ely Joy C. Hilario | https://www.jotform.com/uploads/macarewecare/222927978372067/5445542437116381153/5445542437116381153_signature_31.png
| 2022-11-18 | 2022-11-18 11:26 | Talomo North Health Center | SIR, Matina, Davao City | 11 | Davao del Sur | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22595 | PBSP | alfredojrsantos@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 280 ASSAY: Xpert MTB/RIF Ultra, COUNT: 970 | 2022-11-18 | Gx IV 4D | 815385 | GeneXpert Desktop has error upon logging in windows; Windows Prompt Error "User Profile Service failed the log on" "User Profile Cannot be Loaded" | Diagnose the windows start up and check the user administration settings | Enable and access windows in safe mode Back up data Try to reset the user admin password but still the error exist Create a new profile | GeneXpert desktop can be access with new user profile | Old user profile cannot be accessed | Na | Na | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5445526689424950024/5445526689424950024_signature_28.png | Alfredo H Santos, Jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5445526689424950024/5445526689424950024_signature_31.png
| 2022-11-18 | 2022-11-18 08:56 | Gandara Rural Health Unit | Gandara Samar | 8 | Samar | RTDL Gandara | YES | MAINTENANCE | CEPHEID | [5 empty] | 22594 | DOH | sarah_yangzon@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 356 ASSAY: Xpert MTB/RIF Ultra, COUNT: 193 |  | GXIV-4D | 838406 | N/A | N/A | #NAME? | Module B3 defective (waiting for module replacement) |  |  |  | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5445441851944790279/5445441851944790279_signature_28.png | Sarah Yangzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5445441851944790279/5445441851944790279_signature_31.png
| 2022-11-18 | 2022-11-17 12:25 | San Jose Rural Health Unit | San Jose Northern Samar | 8 | Northern Samar | TB - DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22593 | PBSP | amancitahoney324@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 106 | 2024-02-11 | GXIV - 2D | 834809 | N/A | N/A | -XpertCheck Calibration -Install Xpertcheck version 1.5a -Import new assay (MTB/RIF ultra) -Preventive Maintenance -Generate IQR and System log | Good |  | Xpert Check | Lot 1000378669 | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5445441851945266245/5445441851945266245_signature_28.png | Amancita Honey | https://www.jotform.com/uploads/macarewecare/222927978372067/5445441851945266245/5445441851945266245_signature_31.png
| 2022-11-17 | 2022-11-17 17:07 | Bicol South Luzon Subnational Reference Laboratory/BRDL | BRTTH Compound, Legazpi City | 5 | Albay | Laboratory | NO | INSTALLATION Installation of new CPU | CEPHEID | [5 empty] | 22592 | PBSP | lrnosares.pbsp@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2022-11-17 | Genexpert 4D | 808636 | Unit is due for CPU replacement | CPU Replacement | Performed CPU Replacement Performed Preventive Maintenance | Unit is operational |  | Lenovo ThinkCenter CPU/ CPH1121F27252 | PBSP | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5444859563313376115/5444859563313376115_signature_28.png | LHOUIELYN NOSARES | https://www.jotform.com/uploads/macarewecare/222927978372067/5444859563313376115/5444859563313376115_signature_31.png
| 2022-11-17 | 2022-11-17 17:07 | Bicol South Luzon Subnational Reference Laboratory/BRDL | BRTTH Compound, Legazpi City | 5 | Albay | Laboratory | NO | Installation of new CPU unit | CEPHEID | [5 empty] | 22591 | PBSP | lrnosares.pbsp@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 18 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2022-11-17 | Genexpert 4D | 808636 | Unit is due for CPU replacement | CPU Replacement | Performed CPU Replacement Performed Preventive Maintenance | Unit is operational |  | Lenovo ThinkCenter CPU/ CPH1121F27252 | PBSP | ERON DUMAGIENG | https://www.jotform.com/uploads/macarewecare/222927978372067/5444858523315743271/5444858523315743271_signature_28.png | LHOUIELYN NOSARES | https://www.jotform.com/uploads/macarewecare/222927978372067/5444858523315743271/5444858523315743271_signature_31.png
| 2022-11-17 | 2022-11-17 11:19 | Sta. Cruz Rural Health Unit | Sta. Cruz, Laguna | 4A | Laguna | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22590 | DOH | evetrams1234@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 85 ASSAY: Xpert MTB/RIF Ultra, COUNT: 616 | 2022-09-29 | GX-IV-D | 820229 | For XpertCheck | Run XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. Module D1/722064 and D4/722098 under observation. | All modules passed. | XpertCheck kit | DR #9220 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5444794686416159750/5444794686416159750_signature_28.png | Evelyn Ramos | https://www.jotform.com/uploads/macarewecare/222927978372067/5444794686416159750/5444794686416159750_signature_31.png
| 2022-11-17 | 2022-11-17 14:39 | Cotabato Regional and Medical Center | Cotabato City | BARMM | Maguindanao | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22589 | PBSP | tindocvelche7@yahoo.com | 22.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 7 ASSAY: Xpert MTB/RIF, COUNT: 7 ASSAY: Xpert HIV-1 Viral Load, COUNT: 79 | 2022-11-17 | GX-4 | 807463 | Module C2 for Service Required | Genexpert machine Preventive Maintenance Replace module C2 SN:677827 with new module SN:210068722 | Genexpert machine Preventive Maintenance Replace module C2 SN:677827 with new module SN:210068722 | GX machine has 3 Operational modules | Module C3 for Service required |  | 9000-0508 | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5444769816512544254/5444769816512544254_signature_28.png | VELCHE TINDOC | https://www.jotform.com/uploads/macarewecare/222927978372067/5444769816512544254/5444769816512544254_signature_31.png
| 2022-11-17 | 2022-11-17 14:28 | Cabanatuan CHO | Kapitan Pepe Cabanatuan City Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22588 | DOH | cholab2021@gmail.com | 31 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 4 | 2022-11-17 | Genexpert IV | 831592 | N/a | Gathered Reports Check the unit's condition | Performed Maintenance Cleaned Filter Cleaned Icore Cleaned Plunger Rod Cleaned Cartridge Bay Performed Functionality Test | Unit is operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5444762809326377590/5444762809326377590_signature_28.png | Maricris C. So | https://www.jotform.com/uploads/macarewecare/222927978372067/5444762809326377590/5444762809326377590_signature_31.png
| 2022-11-17 | 2022-11-17 13:43 | Tomas Claudio Health Center - District A | Camus St. Corner E. Quirino, Davao City | 11 | Davao del Sur | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22587 | DOH | arleen_p_alconera@yahoo.com | 31.4 | ASSAY: Xpert MTB/RIF, COUNT: 280 ASSAY: Xpert MTB/RIF Ultra, COUNT: 89 | 2022-11-17 | Gx IV- 4D | 838492 | For Preventive Maintenance | Perfrom Preventive Maintenance | Clean Cartridge Bay, Fan filter, Icore, Plunger Rod Generate Reports - IQR, Assay stat, System Log | Unit is fully operational | High Room Temperature | N/a | N/a | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5444739885018604722/5444739885018604722_signature_28.png | Arleen P. Alconera, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5444739885018604722/5444739885018604722_signature_31.png
| 2022-11-17 | 2022-11-15 16:00 | Siay Rural Health Unit | Poblacion, Siay Zamboanga Sibugay | 9 | Zamboanga Sibugay | Laboratory | NO | Module Replacement | CEPHEID | [5 empty] | 22586 | PBSP | juvyouano@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 887 | 2022-10-18 | GX IV | 815304 | Module B3 temperature error | Replaced module 695770 with 759896 | Module Replacement | All four 4 modules are available | Monitor LCD Broken | 900-0491RS | G 4876 | Gian Carlo Aizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5444734508385741905/5444734508385741905_signature_28.png | Juvy E. Ouano | https://www.jotform.com/uploads/macarewecare/222927978372067/5444734508385741905/5444734508385741905_signature_31.png
| 2022-11-17 | 2022-11-17 13:37 | Nabua rhu | Nabua, camarines sur | 5 | Camarines Sur | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22585 | PBSP | fevalenzuela1965@gmail.com | 28.3 | ASSAY: Xpert MTB/RIF, COUNT: 313 ASSAY: Xpert MTB/RIF Ultra, COUNT: 488 |  | Gx4 | 834620 | CPU unable to boot and power | Check cpu, reset memory card,check cable connection | Check up | Operational |  |  |  | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5444732457616243231/5444732457616243231_signature_28.png | Fe L. Valenzuela | https://www.jotform.com/uploads/macarewecare/222927978372067/5444732457616243231/5444732457616243231_signature_31.png
| 2022-11-17 | 2022-11-17 13:06 | Bureau of Quarantine | Intramuros Manila | NCR | NCR 1st - Manila | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22584 | COMMERCIAL | dklo.lab@quarantine.doh.gov.ph | 25 | ASSAY: Xpert SARS-CoV-2, COUNT: 1143 |  | GX16 | 844418 | Unit for PM. | Performed Preventive Maintenance. | Preventive Maintenance | Unit is operational. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5444712987804887359/5444712987804887359_signature_28.png | Dulce Karmela Lo | https://www.jotform.com/uploads/macarewecare/222927978372067/5444712987804887359/5444712987804887359_signature_31.png
| 2022-11-17 | 2022-11-17 12:50 | Dr. Emigdio C. Cruz Memorial Hospital | Arayat Pampanga | 3 | Pampanga | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | [5 empty] | 22582 | PBSP | leilani101683@yahoo.com | 18 | ASSAY: Xpert MTB/RIF, COUNT: 344 ASSAY: Xpert MTB/RIF Ultra, COUNT: 21 ASSAY: Xpert HIV-1 Viral Load, COUNT: 3 | 2022-11-10 | GXIV-D | 812790 | Deffective CPU | #NAME? | #NAME? | Unit is operational | N/ | CPU | G4940 | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5444708164025806751/5444708164025806751_signature_28.png | Leilani Marie P. Apostol | https://www.jotform.com/uploads/macarewecare/222927978372067/5444708164025806751/5444708164025806751_signature_31.png
| 2022-11-17 | 2022-11-17 10:29 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | [5 empty] | 22581 | MACARE-HBDC | roselyn.alua@redcross.org.ph | 25 |  |  | MCS+ | 12D063 | Unit for PM and Calibration. | Performed PM and Calibration. | Preventive Maintenance and Calibration. | Unit is operational |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5444635390619788686/5444635390619788686_signature_28.png | Roselyn Jra. C. Alua | https://www.jotform.com/uploads/macarewecare/222927978372067/5444635390619788686/5444635390619788686_signature_31.png
| 2022-11-17 | 2022-11-17 10:32 | Philippine Red Cross | Intramuros Manila | NCR | NCR 1st - Manila | Blood Bank | NO | CALIBRATION MAINTENANCE | HAEMONETICS | [5 empty] | 22580 | MACARE-HBDC | roselyn.alua@redcross.org.ph | 25 |  |  | MCS+ | 02A087 | Unit for PM and Calibration | Performed PM and Calibration. | Preventive Maintenance and Calibration. | Unit is operational. |  |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5444634950619109817/5444634950619109817_signature_28.png | Roselyn Jra. C. Alua | https://www.jotform.com/uploads/macarewecare/222927978372067/5444634950619109817/5444634950619109817_signature_31.png
| 2022-11-17 | 2022-11-17 10:29 | Bicol medical center | Bmc road naga city | 5 | Camarines Sur | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22579 | PBSP | bmc.bacteriology@gmail.com | 21.5 | ASSAY: Xpert HPV, COUNT: 1043 ASSAY: Xpert MTB/RIF, COUNT: 2453 ASSAY: Xpert MTB/RIF Ultra, COUNT: 3775 ASSAY: Xpert SARS-CoV-2, COUNT: 3260 | 2022-11-17 | GXVI | 842225 | For module replacement b1 886287-210068158 | Performed module replacement,cleaned filter ,cartridge bay,plunger,icore,gathered reports,test run unit | Module replacement, PMS | Operational |  | 900gx genexpert 6 color module | O | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5444631044231689250/5444631044231689250_signature_28.png | Mary joy Nisola | https://www.jotform.com/uploads/macarewecare/222927978372067/5444631044231689250/5444631044231689250_signature_31.png
| 2022-11-16 | 2022-11-16 18:19 | San Jacinto Rural Health Unit | Espinosa st. District 3, Poblacion, San Jacinto | 5 | Masbate | Laboratory | NO | INSTALLATION CALIBRATION | CEPHEID | [5 empty] | 22578 | PBSP | janet.bragais@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2022-11-16 | Genexpert 4D | 812793 | Unit is due for CPU Replacement Unit is due for module Replacement Unit is due for xpertcheck | CPU Replacement Module Replacement | Performed CPU Replacement Performed Module Replacement B1- 685729 old B2- 686291 old Performed XpertCheck | Unit is operational |  | Lenovo Desktop CPU Genexpert Module/ B1-694090 new Genexpert Module/ B2-712979 new | PBSP | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5444042412327072395/5444042412327072395_signature_28.png | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5444042412327072395/5444042412327072395_signature_31.png
| 2022-11-16 | 2022-11-16 16:27 | Sasa District Health Center | Sasa, Davao City | 11 | Davao del Sur | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22577 | PBSP | mariarosario_mapayo@yahoo.com.sg | 20.1 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 239 ASSAY: Xpert SARS-CoV-2, COUNT: 3926 | 2022-11-16 | Gx IV - 4D | 838889 | Module B4 encounter an error | Perform/ used new cartridge to test Module B4 | Relocate the GeneXpert Set Up Test and observe Module B4 | Machine has only two functional Module: Modules B2 and B4 are disabled | Assay SARS COV from St. Elizabeth Hospital Inc. : Modules Recommended for calibration | N/A | N/A | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5443994826221505611/5443994826221505611_signature_28.png | Maria Rosario M. Pino | https://www.jotform.com/uploads/macarewecare/222927978372067/5443994826221505611/5443994826221505611_signature_31.png
| 2022-11-16 | 2022-11-16 15:58 | Palayan CHO | Palayan City Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22576 | DOH | jazer.palasigue30@gmail.com | 30 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2022-11-16 | Genexpert IV | 831637 | Module 4 error 1002 | Performed xpertcheck in module 4 | Performed maintenance Gathered reports Performed xpertcheck in module 4 Performed functionality test | Unit is operational | Xpertcheck B4 - passed | Xpertcheck 1pc | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5443959656817917006/5443959656817917006_signature_28.png | Jazer M. Palasigue | https://www.jotform.com/uploads/macarewecare/222927978372067/5443959656817917006/5443959656817917006_signature_31.png
| 2022-11-16 | 2022-11-16 13:47 | PORAC RURAL HEALTH UNIT | Genarl Luna St. Poboacion Porac Pampanga | 3 | Pampanga | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22575 | DOH | rhulabporac@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 581 ASSAY: Xpert MTB/RIF Ultra, COUNT: 293 | 2022-11-16 | GXIV-D | 831911 | -Error 2008 (Module C2) -C3 temperature error -C4 error 4017 | #NAME? | Maintenance and troubleshooting | Unit is operational | C3-784539 and C4-588984 already reported. Waiting for module replacement. | N/A | N/A | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5443882081837407482/5443882081837407482_signature_28.png | Marife Lansangan | https://www.jotform.com/uploads/macarewecare/222927978372067/5443882081837407482/5443882081837407482_signature_31.png
| 2022-11-16 | 2022-11-16 11:28 | Laoang Rural Health Unit | Laoang northern samar | 8 | Northern Samar | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22574 | DOH | roseadarayan26@gmail.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 479 ASSAY: Xpert MTB/RIF Ultra, COUNT: 198 |  | GXIV-4D | 838202 | N/A | N/A | Preventive Maintenance | Good |  |  |  | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5443837159617169839/5443837159617169839_signature_28.png | Rosario Adarayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5443837159617169839/5443837159617169839_signature_31.png
| 2022-11-16 | 2022-11-16 09:49 | Dr. Jose Rizal Memorial Hospital | Dapitan City, Zamboanga del Norte | 9 | Zamboanga del Norte | TBDots | YES | REPAIR | CEPHEID | [5 empty] | 22573 | PBSP | lovinoantonieta@gmail.com | 24 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1377 |  | GXIV-4D | 804698 | Module C1 (607887) defective | -Replace defective module C1 with new module (691298) | Replace fan filter | All four modules are functional. | Module C2 has frequent temp error - for observation. CPU has Trojan malware (usb shortcut virus), recommend to avoid plugging in unauthorized flash drives |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5443728218811562543/5443728218811562543_signature_28.png | Antonieta Lovino | https://www.jotform.com/uploads/macarewecare/222927978372067/5443728218811562543/5443728218811562543_signature_31.png
| 2022-11-16 | 2022-11-16 09:43 | Zamboanga del Norte Medical Center | Dipolog City, Zamboanga del Norte | 9 | Zamboanga del Norte | TBDots | YES | CALIBRATION | CEPHEID | [5 empty] | 22572 | PBSP | ferranbryder@gmail.com | 23.2 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 191 |  | GXIV-4D | 815285 | #NAME? | Run XpertCheck on all four modules | Preventive Maintenance | Xpert Check Result: -Module B1 service required -Module B2, B3, B4 passed |  |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5443724968814289336/5443724968814289336_signature_28.png | Ferran Tomoy | https://www.jotform.com/uploads/macarewecare/222927978372067/5443724968814289336/5443724968814289336_signature_31.png
| 2022-11-16 | 2022-11-16 09:33 | Dipolog City Health Office | Dipolog City, Zamboanga del Norte | 9 | Zamboanga del Norte | TBDots | YES | CALIBRATION | CEPHEID | [5 empty] | 22571 | PBSP | jz.ebron@yahoo.com.ph | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 581 ASSAY: Xpert HIV-1 Viral Load, COUNT: 1 |  | GXIV-4D | 834826 | Due for XpertCheck | Run Xpert Check on all modules | Preventive Maintenance | Module B1 - service required, Module B2 - pass |  |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5443721928815123661/5443721928815123661_signature_28.png | James Zebedee Ebron | https://www.jotform.com/uploads/macarewecare/222927978372067/5443721928815123661/5443721928815123661_signature_31.png
| 2022-11-16 | 2022-11-16 09:21 | Hi-Precision Diagnostics | N. Amoranto St. QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22570 | CONSORTIUM | pcr@hi-precision.com.ph | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 2626 ASSAY: Xpert HPV, COUNT: 52 ASSAY: Others, COUNT: 1 |  | GXIV | 835602 | Unit for PM. | Performed PM. | Preventive Maintenance | Unit is operational. 3 modules operational, Module A2 (791652) is disabled. | For request of Xpertcheck. Due on May 2022. |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5443719740116432327/5443719740116432327_signature_28.png | Joselito T. Chua | https://www.jotform.com/uploads/macarewecare/222927978372067/5443719740116432327/5443719740116432327_signature_31.png
| 2022-11-16 | 2022-11-16 09:26 | Hi-Precision Diagnostics | N. Amoranto St. QC | NCR | NCR 2nd - Quezon City | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22569 | CONSORTIUM | pcr@hi-precision.com.ph | 25 | ASSAY: Xpert HPV, COUNT: 240 ASSAY: Xpert MTB/RIF Ultra, COUNT: 4067 |  | GXIV | 835601 | Unit for PM. | Performed PM. | Preventive Maintenance. | Unit is operational. 2 modules operational, Module B2 (791454) & B3 (791649) is disabled. | For request of XpertCheck. Due on May 2022. |  |  | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5443718510112801156/5443718510112801156_signature_28.png | Joselito T. Chus | https://www.jotform.com/uploads/macarewecare/222927978372067/5443718510112801156/5443718510112801156_signature_31.png
| 2022-11-16 | 2022-11-16 09:26 | Sindangan Rural Health Unit | Sindangan, Zamboanga del Norte | 9 | Zamboanga del Norte | TBDots | YES | REPAIR | CEPHEID | [5 empty] | 22568 | PBSP | lryano0618@gmail.com | 27 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1680 |  | GXIV-4D | 807596 | Module B1 - SN 644791 | Replace defective module B1 with new module SN: 861036 | Preventive Maintenance | All modules are functional |  |  |  | Wynnona Reoma | https://www.jotform.com/uploads/macarewecare/222927978372067/5443715018816748990/5443715018816748990_signature_28.png | Leni Riza Yano | https://www.jotform.com/uploads/macarewecare/222927978372067/5443715018816748990/5443715018816748990_signature_31.png
| 2022-11-16 | 2022-11-16 08:57 | San Roque RHU | San Roque Northern Samar | 8 | Northern Samar | Laboratory | YES | MAINTENANCE | CEPHEID | [5 empty] | 22567 | DOH | lynnjarito@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 315 ASSAY: Xpert MTB/RIF Ultra, COUNT: 219 |  | GXIV - 4D | 838498 | Low ink cartridge printer (brand: Canon Pixma iP2770 SN: KMMK16785 | N/A | Preventive Maintenance | Good |  |  |  | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5443698089551225205/5443698089551225205_signature_28.png | Lynn N. Jarito | https://www.jotform.com/uploads/macarewecare/222927978372067/5443698089551225205/5443698089551225205_signature_31.png
| 2022-11-15 | 2022-11-15 15:45 | Bagong pag asa health center | Bagong pag asa q.c | NCR | NCR 2nd - Quezon City | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22566 | DOH | erlindalorilla026@gmail.com | 25.6 | ASSAY: Xpert HPV, COUNT: 10 ASSAY: Xpert MTB/RIF, COUNT: 250 ASSAY: Xpert MTB/RIF Ultra, COUNT: 694 | 2022-11-15 | Gx4 | 831597 | For module replacement temp error | Module replacement a2 754336>698524 a3 754337>632618 a4 754338>697973 | Module replacement and maintenance | A1 disable reported already A2,A3,A4 operational |  | 900gx module 3pcs | 9341 | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5443252689889350144/5443252689889350144_signature_28.png | Erlinda S. Lorilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5443252689889350144/5443252689889350144_signature_31.png
| 2022-11-15 | 2022-11-15 14:59 | Rosario Super Health Center | Rosario, Pasig | NCR | NCR 2nd - Pasig | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22565 | WHO | cawayanjodo@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 306 ASSAY: Xpert MTB/RIF Ultra, COUNT: 878 ASSAY: Xpert SARS-CoV-2, COUNT: 551 |  | GX-IV-D | 843924 | Module B4 H/W failed | Check modules | Performed maintenance Cleaned plunger rod, icore and cartridge bays Generated reports | All modules not operational. |  |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5443084708115776539/5443084708115776539_signature_28.png | Eden Cawayan | https://www.jotform.com/uploads/macarewecare/222927978372067/5443084708115776539/5443084708115776539_signature_31.png
| 2022-11-15 | 2022-11-15 14:56 | CANDIJAY COMMUNITY HOSPITAL | CANDIJAY BOHOL | 7 | Bohol | LABORATORY | NO | REPAIR | CEPHEID | [5 empty] | 22564 | DOH | angelinaredulla@yahoo.com.ph | 30 | ASSAY: Xpert MTB/RIF, COUNT: 687 | 2022-11-10 | GX-IV | 838396 | AVR UPS POWER ISSUE | -CHECK AVR FUNCTIONALITY, PASSED ; UPS - PASSED -CHECK UPS AND AVR OUTPUT VOLTAGE: WITHIN RANGE | -CHECK OUTPUT VOLTAGE AT LAB 320-350V READING -HOSPITAL ELECTRICIAN PERFORM ADJUSTMENT, VOLTAGE READING AT 250V | ALL EQUIPMENT ARE FUNCTIONAL, LABORATORY LINE VOLTAGE AT 320-350V READING | #NAME? | N/A | N/A | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5443065878817713346/5443065878817713346_signature_28.png | ANGELINA O. REDULLA | https://www.jotform.com/uploads/macarewecare/222927978372067/5443065878817713346/5443065878817713346_signature_31.png
| 2022-11-15 | 2022-11-15 14:55 | Davao Regional Medical Center | Apokon, Tagum City | 11 | Davao del Norte | Molecular Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22563 | PBSP | mls021493@gmail.com | 25.8 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 985 | 2022-06-17 | GX-IV 4D | 834626 | Module B3 for Replacement due to H/W Failure | Module Replacement | Replaced defective module with s/n 893589 with service module s/n 674488 | Unit is fully operational | Service module is from gx16 | 900-GxlX | 4604 | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5443052552116440672/5443052552116440672_signature_28.png | Rowena Bancairen | https://www.jotform.com/uploads/macarewecare/222927978372067/5443052552116440672/5443052552116440672_signature_31.png
| 2022-11-15 | 2022-11-15 14:54 | Davao Regional and Medical Center | Tagum City | 11 | Davao del Norte | Laboratory | YES | REPAIR | CEPHEID | [5 empty] | 22562 | PBSP | mls021493@gmail.com | 25.8 | ASSAY: Xpert SARS-CoV-2, COUNT: 4000 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1000 | 2022-11-15 | GX-16 | 838756 | Power supply Replacement | Replace Gx-16 Power Supply | Replace Gx-16 Power Supply | None Operational Due Power Supply Problem | Unsuccessfully installed Due to Incompatibility of power supply | N/A | N/A | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5443050963183189167/5443050963183189167_signature_28.png | Rowena Bancairen | https://www.jotform.com/uploads/macarewecare/222927978372067/5443050963183189167/5443050963183189167_signature_31.png
| 2022-11-15 | 2022-11-15 14:35 | Marilao Rural Health Unit I | Marilao, Bulacan | 3 | Bulacan | Laboratory | NO | Check up | CEPHEID | [5 empty] | 22561 | DOH | lab.rhu1marilao@gmail.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 894 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1039 ASSAY: Xpert HIV-1 Viral Load, COUNT: 52 |  | GX IV | 831897 | Module A1 for check up. | NA | #NAME? | Unit is operational. All 4 modules seems functional. | Module A1 is temp error free for a month. Module is under observation. |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5443045741718515541/5443045741718515541_signature_28.png | JONATHAN JENILO | https://www.jotform.com/uploads/macarewecare/222927978372067/5443045741718515541/5443045741718515541_signature_31.png
| 2022-11-15 | 2022-11-15 14:52 | DON EMILIO DEL VALLE MEMORIAL HOSPITAL | UBAY,BOHOL | 7 | Bohol | LABORATORY | NO | ON-SITE ASSISTANCE | CEPHEID | [5 empty] | 22560 | PBSP | example@example.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 2123 ASSAY: Xpert MTB/RIF Ultra, COUNT: 165 | 2022-11-09 | GX-IV | 807603 | MODULE B3 ERROR 4008, MODULE B4 ERROR 4012 | RESET GX CONNECTIONS | PERFORM GX CLEANING, COLLECT IQR, SYSTEM LOG, AND ARCHIVE | B1,B2 - Passed ; B3,B4 - H/W FAILED | N/A | N/A | N/A | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5443045538819934129/5443045538819934129_signature_28.png | JAN ARKID A. DIACOR,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5443045538819934129/5443045538819934129_signature_31.png
| 2022-11-15 | 2022-11-15 14:24 | Lianga Rural Health Unit | Purok 2, Poblacion, Lianga | 13 | Surigao del Sur | RTDL/TB-DOTS | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22559 | PBSP | gyvxs@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 249 ASSAY: Xpert MTB/RIF Ultra, COUNT: 118 |  | GX-4 with 2 modules | 835174 | Module B1 with SN: 788658 was found Service Required | None | XpertCheck Calibration and Preventive Maintenance | 2 modules working but 1 module was prior for replacement | N/A |  |  | Engr. Faizar S. Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5443031123323880062/5443031123323880062_signature_28.png | Gyver I. Sistona | https://www.jotform.com/uploads/macarewecare/222927978372067/5443031123323880062/5443031123323880062_signature_31.png
| 2022-11-15 | 2022-11-15 14:01 | Dingalan Community Hospital | Caragsacan Dingalan Aurora | 3 | Aurora | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22558 | DOH | bnjlvrz9@gmail.com | 33 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2022-11-15 | Genexpert gx4 | 831621 | Genexpert is due for Calibration | Gathered Reports | Performed Maintenance Cleaned Filters Cleaned Cartridge bay Cleaned Icore Cleaned Plunger Rod Performed functionality Test | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5443014656118923964/5443014656118923964_signature_28.png | Benjo Alvarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5443014656118923964/5443014656118923964_signature_31.png
| 2022-11-15 | 2022-11-15 13:40 | Tarlac City Health Center X | ICS Tibag Tarlac City | 3 | Tarlac | Laboratory | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22557 | DOH | rodelie220@gmail.com | 18 | ASSAY: Xpert MTB/RIF, COUNT: 486 ASSAY: Xpert MTB/RIF Ultra, COUNT: 240 | 2022-11-15 | GXIV-D | 831587 | #NAME? | -Checked power connections -Troubleshoot ups (working) -Troubleshoot CPU (working) -Performed GxReliability test (error 2005) -Performed functionality test | Troubleshooting | Unit is operational | Under observation | N/A | N/A | Mark Joseph Venzon | https://www.jotform.com/uploads/macarewecare/222927978372067/5443014413668178278/5443014413668178278_signature_28.png | Rodelie Verayo | https://www.jotform.com/uploads/macarewecare/222927978372067/5443014413668178278/5443014413668178278_signature_31.png
| 2022-11-15 | 2022-11-15 13:44 | Masambong Health Center | Masambong Quezon City | NCR | NCR 2nd - Quezon City | Laboratory | NO | INSTALLATION MAINTENANCE | CEPHEID | [5 empty] | 22556 | PBSP | michellebonrostro@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 | 2022-11-15 | GXIV | 814160 | Unit for PM and for installation of new CPU | PM and Installation of CPU | PM and Installation of CPU | Unit is operational. |  | Lenovo ThinkCentre CPU (SN:CPH0420F08151) in replacement for defective CPU HP Z240 (SN: 2UA7021MTS) | N/A | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5443005690116752408/5443005690116752408_signature_28.png | Michelle Bonrostro | https://www.jotform.com/uploads/macarewecare/222927978372067/5443005690116752408/5443005690116752408_signature_31.png
| 2022-11-15 | 2022-11-15 13:09 | Meycauayan City Health Center VII | Meycauayan, Bulacan | 3 | Bulacan | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22555 | DOH | ronamarieiris_25@yahoo.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 326 ASSAY: Xpert MTB/RIF Ultra, COUNT: 625 ASSAY: Xpert HIV-1 Viral Load, COUNT: 22 |  | GX IV | 831904 | Unit for maintenance. | NA | #NAME? | Unit is operational. 3 modules are functional. | Waiting for module for A3 replacement. |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5442984421718513595/5442984421718513595_signature_28.png | Rona Marie Iris R. De Leon | https://www.jotform.com/uploads/macarewecare/222927978372067/5442984421718513595/5442984421718513595_signature_31.png
| 2022-11-15 | 2022-11-15 10:40 | Antipolo City Health Office | Antipolo, Rizal | 4A | Rizal | Laboratory | NO | INSTALLATION | CEPHEID | [5 empty] | 22554 | PBSP | acho.laboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1683 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5870 ASSAY: Xpert HIV-1 Viral Load, COUNT: 249 |  | GX-XVI | 842226 | For installation | Installed the GeneXpert unit. | Installed the GX-XVI Functionality test Run reliability test Generated reports | Unit is operational. | Machine only. |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5442894822816622314/5442894822816622314_signature_28.png | Christine S. Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5442894822816622314/5442894822816622314_signature_31.png
| 2022-11-14 | 2022-11-14 14:52 | San Isidro | Poblacion,San Isidro, Nueva Ecija | 3 | Nueva Ecija | Laboratory | NO | MAINTENANCE | CEPHEID | [5 empty] | 22553 | PBSP | rmumali44@gmail.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2022-11-14 | Genexpert gx4 | 805107 | Module 1 and 2 errors | Gathered Reports to Report | Performed Maintenance Cleaned Filters Cleaned Cartridge bay Cleaned Plunger Rod Cleaned Icore | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5442182369738327377/5442182369738327377_signature_28.png | Rowena Umali | https://www.jotform.com/uploads/macarewecare/222927978372067/5442182369738327377/5442182369738327377_signature_31.png
| 2022-11-14 | 2022-11-11 13:12 | GENERAL EMILIO AGUINALDO MEMORIAL HOSPITAL | Trece Martires | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22549 | PBSP | djaemiel@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 658 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1312 |  | GXIV-L | 802860 | GX Software Problem Module 3 lost comm error | Reset Connections Purged Tests | Maintenance Reset Gateway Board Connections | 2 modules are operational. 2 modules for replacement | B3/632179 and B4/631101 ultrasonic horn cables problem |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152824116014538/5442152824116014538_signature_28.png | JAEMIEL JUVIRENS A. DIMARANAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152824116014538/5442152824116014538_signature_31.png
| 2022-11-14 | 2022-11-09 13:37 | Alabat District Hospital | Alabat Island | 4A | Quezon | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22551 | DOH | letpiquero@yahoo.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 388 |  | GXIV-D | 831614 | Due for Xpertcheck | Performed XpertCheck | XpertCheck Maintenance | 3 modules operational, 1 Module Failed | C2/754442 - Failed |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152824118484772/5442152824118484772_signature_28.png | Carmelita S. Piquero | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152824118484772/5442152824118484772_signature_31.png
| 2022-11-14 | 2022-11-11 11:11 | Dasmariñas City Health Office II | Area E, Dasmariñas | 4A | Cavite | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22550 | DOH | betoy.nerissa.iris10@gmail.com | 26 | ASSAY: Xpert MTB/RIF, COUNT: 568 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1916 ASSAY: Xpert HIV-1 Viral Load, COUNT: 30 |  | GXIV-D | 838535 | Module 2 with Temp Errors | Xpertcheck on Module 2 | Maintenance XpertCheck on B2 | 2 modules are operational. 2 modules for replacement | Xpertcheck: B2/862309 - Failed |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152824112658044/5442152824112658044_signature_28.png | Roberto c. Borilla | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152824112658044/5442152824112658044_signature_31.png
| 2022-11-14 | 2022-11-10 09:38 | Quezon Medical Center | Lucena | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22552 | PBSP | joannpalermo0613@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 158 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1618 ASSAY: Xpert HIV-1 Viral Load, COUNT: 41 |  | GXIV-D | 804692 | Temp errors on module 4 | Replaced module 4 | Maintenance Module Replacement | 4 modules are operational |  |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152824119362184/5442152824119362184_signature_28.png | JoAnn Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152824119362184/5442152824119362184_signature_31.png
| 2022-11-14 | 2022-11-11 14:41 | Marikina CHO | Marikina City | NCR | NCR 2nd - Marikina | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22548 | DOH | amymitra22@gmail.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 667 ASSAY: Xpert MTB/RIF Ultra, COUNT: 997 ASSAY: Xpert HIV-1 Viral Load, COUNT: 61 |  | GX-IV-D | 831588 | Module A3 error 1001/1002/4017 | Run XpertCheck on module A3 | Performed maintenance Run XpertCheck on module A3 Generated reports | Unit is operational. | Module A3 passed. Under observation. |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152744119820139/5442152744119820139_signature_28.png | Amelia Mitra | https://www.jotform.com/uploads/macarewecare/222927978372067/5442152744119820139/5442152744119820139_signature_31.png
| 2022-11-14 | 2022-11-14 13:49 | DON EMILIO DEL VALLE MEMORIAL HOSPITAL | UBAY,BOHOL | 7 | Bohol | LABORATORY | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | [5 empty] | 22547 | PBSP | example@example.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 2123 ASSAY: Xpert MTB/RIF Ultra, COUNT: 165 | 2022-11-09 | GX-IV | 807603 | MODULE B3 ERROR 4008, MODULE B4 ERROR 4012 | COLLECT IQR,SYSTEM LOG, AND ARCHIVE | PERFORM CLEANING, RESET GX CONNECTIONS | B1,B2-OK; B3,B4- H/W FAILED |  | N/A | N/A | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5442144684818509182/5442144684818509182_signature_28.png | JAN ARKID A. DIACOR,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5442144684818509182/5442144684818509182_signature_31.png
| 2022-11-14 | 2022-11-14 10:36 | DON EMILIO DEL VALLE MEMORIAL HOSPITAL | UBAY,BOHOL | 7 | Bohol | LABORATORY | NO | TROUBLESHOOTING ON-SITE ASSISTANCE | CEPHEID | [5 empty] | 22546 | PBSP | example@example.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 2123 ASSAY: Xpert MTB/RIF Ultra, COUNT: 165 | 2022-11-09 | GX-IV | 807603 | MODULE B3 ERROR 4008, MODULE B4 ERROR 4012 | COLLECT IQR,SYSTEM LOG, AND ARCHIVE | PERFORM CLEANING, RESET GX CONNECTIONS | B1,B2-OK; B3,B4- H/W FAILED |  | N/A | N/A | JOSEPH ACEBEDO | https://www.jotform.com/uploads/macarewecare/222927978372067/5442140214817612087/5442140214817612087_signature_28.png | JAN ARKID A. DIACOR,RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5442140214817612087/5442140214817612087_signature_31.png
| 2022-11-14 | 2022-11-10 09:38 | Quezon Medical Center | Lucena | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22545 | PBSP | joannpalermo0613@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 158 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1618 ASSAY: Xpert HIV-1 Viral Load, COUNT: 41 |  | GXIV-D | 804692 | Temp errors on module 4 | Replaced module 4 | Maintenance Module Replacement | 4 modules are operational |  |  |  | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5442131211217730822/5442131211217730822_signature_28.png | JoAnn Yamballa | https://www.jotform.com/uploads/macarewecare/222927978372067/5442131211217730822/5442131211217730822_signature_31.png
| 2022-11-12 | 2022-11-10 18:30 | Felipe Abrigo Memorial Hospital | Guiuan, Eastern Samar | 8 | Eastern Samar | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22539 | PBSP | dagantejas@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 460 ASSAY: Xpert MTB/RIF Ultra, COUNT: 226 |  | GXIV - 4D | 8152711 | UPS not working | Check Voltage source (215 V ) | -Troubleshooting UPS ( SN: 40223321704) -Preventive Maintenance -Generate IQR and System log | UPS battery not charging |  |  |  | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5440727675611857939/5440727675611857939_signature_28.png | JASMIN JOY NAVIDAD | https://www.jotform.com/uploads/macarewecare/222927978372067/5440727675611857939/5440727675611857939_signature_31.png
| 2022-11-12 | 2022-11-10 18:30 | Felipe Abrigo Memorial Hospital | Guiuan, Eastern Samar | 8 | Eastern Samar | Laboratory | YES | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22538 | PBSP | dagantejas@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 460 ASSAY: Xpert MTB/RIF Ultra, COUNT: 226 |  | GXIV - 4D | 8152711 | UPS not working | Check Voltage source (215 V ) | -Troubleshooting UPS ( SN: 40223321704) -Preventive Maintenance -Generate IQR and System log | UPS battery not charging |  |  |  | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5440584735612876619/5440584735612876619_signature_28.png | JASMIN JOY NAVIDAD | https://www.jotform.com/uploads/macarewecare/222927978372067/5440584735612876619/5440584735612876619_signature_31.png
| 2022-11-11 | 2022-11-11 16:30 | ST. LUKE'S EXTENSION CLINIC | BOCOBO ST. ERMITA, MANILA | NCR | NCR 1st - Manila | LABORATORY | NO | REPAIR | CEPHEID | [5 empty] | 22537 | COMMERCIAL | laboratory@slec.ph | 21.6 | ASSAY: Xpert CT/NG, COUNT: 3811 ASSAY: Xpert MTB/RIF, COUNT: 2384 | 2022-11-11 | INFINITY | 809643 | MODULES FOR REPLACEMENT | PERFORMED MODULE REPLACEMENT | PERFORMED PM GENERATED REPORTS ARCHIVED RESULTS PERFORMED FUNCTIONALITY TEST | UNIT IS OPERATIONAL. HW FAILED F3 642037 F1 880991 OUT OF BOX |  | 900-GX OLD - 635774, 710642, 688434, 681658, 684184, 687517 NEW - 804179, 647149, 880991, 693672, 896445, 877673 | 9326 | CIENNE MAE VALERIO | https://www.jotform.com/uploads/macarewecare/222927978372067/5439655176023751313/5439655176023751313_signature_28.png | Ma. Sarah Salvador | https://www.jotform.com/uploads/macarewecare/222927978372067/5439655176023751313/5439655176023751313_signature_31.png
| 2022-11-11 | 2022-11-11 13:53 | Cainta CHO | Cainta, Rizal | 4A | Rizal | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22536 | PBSP | jmdoctor@macare-medicals.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 355 ASSAY: Xpert MTB/RIF Ultra, COUNT: 566 | 2022-11-07 | GX-IV-D | 804687 | Module B2 for replacement | Replaced module B2 | Performed maintenance Replaced defective module Generated reports | Unit is operational | Old: B2/685113 New: B2/888288 | GeneXpert module | G 4922 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5439555691811712266/5439555691811712266_signature_28.png | Rowena Laxamana | https://www.jotform.com/uploads/macarewecare/222927978372067/5439555691811712266/5439555691811712266_signature_31.png
| 2022-11-11 | 2022-11-11 13:48 | Southern Isabela Medical Center | Rosario Santiago City Isabela | 2 | Isabela | Engineering | NO | REPAIR MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22535 | PBSP | tonz23_mad@yahoo.com | 29 | ASSAY: Xpert MTB/RIF Ultra, COUNT: | 2022-11-11 | Genexpert gx4 | 805104 | Module 4 stock on booting | Performed Maintenance Gathered Reports | Performed Troubleshooting Performed Expertcheck Plugged in and out internal connections Performed Functionality test Enabled module 4 | Unit is Operational | N/a | N/a | N/a | Matthew Roxas | https://www.jotform.com/uploads/macarewecare/222927978372067/5439553367948669586/5439553367948669586_signature_28.png | Mary Antoinette Pastor | https://www.jotform.com/uploads/macarewecare/222927978372067/5439553367948669586/5439553367948669586_signature_31.png
| 2022-11-11 | 2022-11-11 13:10 | Rogaciano M. Mercado Memorial Hospital | Sta. Maria, Bulacan | 3 | Bulacan | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22534 | PBSP | pascualmaribeth72@gmail.co.com | 29 | ASSAY: Xpert MTB/RIF, COUNT: 908 ASSAY: Xpert MTB/RIF Ultra, COUNT: 795 | 2022-11-07 | GX IV | 805252 | Monitor for replacement (LG Flatron E1910P-BN, SN: 312NDDM03410 | Replaced monitor by Lenovo SN: V5DMY041 | #NAME? | Unit is operational. | NA | Lenovo Monitor | G 4925 | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5439549405172463702/5439549405172463702_signature_28.png | Maribeth Bernardo | https://www.jotform.com/uploads/macarewecare/222927978372067/5439549405172463702/5439549405172463702_signature_31.png
| 2022-11-11 | 2022-11-11 13:16 | Esperanza health center | Teresa st sta mesa | NCR | NCR 1st - Manila | Lab | NO | MAINTENANCE | CEPHEID | [5 empty] | 22533 | DOH | esperanzahealthcenter2000@gmail.com | 28 | ASSAY: Xpert HPV, COUNT: 23 ASSAY: Xpert MTB/RIF, COUNT: 446 ASSAY: Xpert MTB/RIF Ultra, COUNT: 631 |  | Gx4 | 831891 | N/a | Cleaned unit,filter,cartridge bay,plunger,icore | Maintenance | Operational |  |  |  | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5439533146592283699/5439533146592283699_signature_28.png | Victor agdamag jr. | https://www.jotform.com/uploads/macarewecare/222927978372067/5439533146592283699/5439533146592283699_signature_31.png
| 2022-11-11 | 2022-11-11 12:31 | Batangas Medical Center | Kumintang Ibaba, Batangas City | 4A | Batangas | Laboratory | NO | INSTALLATION Installation of backup laptop | CEPHEID | [5 empty] | 22532 | PBSP | batmcxpertxpresscovidlab@gmail.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 ASSAY: Xpert HPV, COUNT: 0 ASSAY: Xpert SARS-CoV-2, COUNT: 0 | 2022-11-10 | GX4D | 834627 | Unit is due for laptop backup installation | Intalled backup laptop | Performed Backup Laptop Installation | Unit is good and operational |  | Lenovo thinkpad Laptop | G 4929 | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5439506153712239543/5439506153712239543_signature_28.png | eron dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5439506153712239543/5439506153712239543_signature_31.png
| 2022-11-11 | 2022-11-10 12:21 | Zamboanga City Health Office | Rizal Street, Zamboanga City, Zamboanga del sur | 9 | Zamboanga del Sur | Laboratory | NO | Module Replacement | CEPHEID | [5 empty] | 22531 | PBSP | teresitaflores99@yahoo.com | 23.3 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 670 | 2022-11-10 | GX IV | 807599 | None | Module Replacement | Module Replacement, Preventive Maintenance | All Four (4) Modules are available | Replaced module 662297, 662296 with 817600, 892103. | N/A | N/A | Gian Carlo Aizon | https://www.jotform.com/uploads/macarewecare/222927978372067/5439475643874831855/5439475643874831855_signature_28.png | Teresita J. Flores | https://www.jotform.com/uploads/macarewecare/222927978372067/5439475643874831855/5439475643874831855_signature_31.png
| 2022-11-11 | 2022-11-11 10:44 | Davao Oriental Provincial Medical Center | Mati City, Davao Oriental | 11 | Davao Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22530 | PBSP | ofeliamendiola3@gmail.com | 29.9 | ASSAY: Xpert MTB/RIF, COUNT: 397 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1140 |  | GX IV - 4D | 804693 | Module A1 and A3 encounter an error 2120 | Check for lossy connection, Perfrom Preventive Maintenance | Reset Gateway board Connection Clean Fan filter, Cartridge Bay, iCore Generate IQR, Assay Stat, System Log | A3 subjected for Replacement 3 modules are operational |  |  |  | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5439447435228419951/5439447435228419951_signature_28.png | Ofelia Mendiola, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5439447435228419951/5439447435228419951_signature_31.png
| 2022-11-11 | 2022-11-11 10:13 | PARANG RURAL HEALTH UNIT | Parang | BARMM | Maguindanao | Laboratory | YES | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22529 | PBSP | Nadjimbagua20@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 210 ASSAY: Xpert MTB/RIF Ultra, COUNT: 222 | 2022-11-11 | Gx-4 | 838135 | Module B2 H/W FAILED | Genexpert machine Preventive Maintenance Replace module B2 SN:806694 with new module SN:866870 | Genexpert machine Preventive Maintenance Replace module B2 SN:806694 with new module SN:866870 | In good condition - All modules are Functional | N/A | N/A | 900-0491RS | Jesson P. Tupas | https://www.jotform.com/uploads/macarewecare/222927978372067/5439424296028596010/5439424296028596010_signature_28.png | Nadjim A. Bagua | https://www.jotform.com/uploads/macarewecare/222927978372067/5439424296028596010/5439424296028596010_signature_31.png
| 2022-11-11 | 2022-11-11 10:12 | Mabalacat RHU I | Mabalacat, Pampanga | 3 | Pampanga | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22528 | PBSP | perezriyadh@gmail.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 190 ASSAY: Xpert MTB/RIF Ultra, COUNT: 329 ASSAY: Xpert HIV-1 Viral Load, COUNT: 4 | 2022-11-07 | GX IV | 835104 | Module B2 (787982) for replacement. | Replaced module B2 by 880771 | #NAME? | Uni is operational. | NA | GeneXpert Module | G 4919 | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5439424154689239715/5439424154689239715_signature_28.png | Riyadh L. Perez | https://www.jotform.com/uploads/macarewecare/222927978372067/5439424154689239715/5439424154689239715_signature_31.png
| 2022-11-11 | 2022-11-11 09:36 | Antipolo CHO | Antipolo, Rizal | 4A | Rizal | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22527 | PBSP | acho.laboratory@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1704 ASSAY: Xpert MTB/RIF Ultra, COUNT: 5819 ASSAY: Xpert HIV-1 Viral Load, COUNT: 229 | 2022-11-07 | GX-IV-D | 812006 | Module B3 and B4 for replacement | Replaced defective modules | Perform maintenance Replaced defective modules Generated reports | Unit is operational. | Old modules:B3/623350,B4/792345 New modules: B3/654640,B4/651850 | GeneXpert modules. | G 4921 | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5439403523219487339/5439403523219487339_signature_28.png | Christine S. Suarez | https://www.jotform.com/uploads/macarewecare/222927978372067/5439403523219487339/5439403523219487339_signature_31.png
| 2022-11-11 | 2022-11-10 11:09 | Sariaya RHU | Sariaya, Quezon | 4A | Quezon | Laboratory | NO | CALIBRATION | CEPHEID | [5 empty] | 22526 | DOH | marielinnetabril@yahoo.com | 24 | ASSAY: Xpert MTB/RIF, COUNT: 365 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1276 |  | GX-IV-D | 831890 | For XpertCheck | Performed XpertCheck | Performed maintenance Run XpertCheck Generated reports | Unit is operational. | Module A4 service required. Modules A1,A2,A3 passed. |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5439397333128491085/5439397333128491085_signature_28.png | Marie linnet abril | https://www.jotform.com/uploads/macarewecare/222927978372067/5439397333128491085/5439397333128491085_signature_31.png
| 2022-11-11 | 2022-11-11 08:57 | Domingo santiago health center | Domingo santiago st.sampaloc manila | NCR | NCR 1st - Manila | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22525 | DOH | fei_paas@yahoo.com | 32 | ASSAY: Xpert MTB/RIF, COUNT: 454 ASSAY: Xpert MTB/RIF Ultra, COUNT: 750 | 2022-11-07 | Gx4 | 831527 | Defective monitor | Delivered and replaced monitor | Repair | 2 module operational a3 673872module failed xpertcheck a4 module booting issue unable to performed rexpertcheck |  | Monitor | D.r 9322 | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5439383876591486595/5439383876591486595_signature_28.png | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5439383876591486595/5439383876591486595_signature_31.png
| 2022-11-10 | 2022-11-10 16:30 | Mati City Health Office | Mati City, Davao Oriental | 11 | Davao Oriental | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22524 | PBSP | cholaboratory@yahoo.com | 20.7 | ASSAY: Xpert MTB/RIF, COUNT: 475 ASSAY: Xpert MTB/RIF Ultra, COUNT: 306 | 2022-11-10 | GX-IV 2D | 835221 | Module B2 subjected for replacement | Module Replacement | Replaced defective module S/N : 789059 with New Module S/N: 687398 Perform Preventive Maintenance Clean Cartridge Bay, Icore, Fan Filters, Plunger Rod Generate IQR, System Log, Assay Stat | Unit is fully operational |  | 900-0491rs | 0 | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5438845975225495495/5438845975225495495_signature_28.png | Liezl Ann Alago, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5438845975225495495/5438845975225495495_signature_31.png
| 2022-11-10 | 2022-11-10 14:26 | Asian hospital and medical center | Alabang muntinlupa | NCR | NCR 4th - Muntinlupa | Bloodbank | NO | MAINTENANCE | HAEMONETICS | [5 empty] | 22523 | COMMERCIAL | bmanlangit@asianhospital.com | 26 |  |  | Mcs + | 01F116 | Spillage | Checked bowl optic,performed maintenance,performed simulation running,checked line sensor,weigher,valve,replaced rollers | Checked up and Maintenance | Operational |  |  |  | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5438755556594955936/5438755556594955936_signature_28.png | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5438755556594955936/5438755556594955936_signature_31.png
| 2022-11-10 | 2022-11-10 15:06 | Mandalulong CHO | Mandaluyong City Hall | NCR | NCR 2nd - Mandaluyong | Laboratory | NO | REPAIR Module Replacement | CEPHEID | [5 empty] | 22522 | PBSP | recie.mendoza@yahoo.com | 21 | ASSAY: Xpert MTB/RIF, COUNT: 79 ASSAY: Xpert MTB/RIF Ultra, COUNT: 0 ASSAY: Xpert HIV-1 Viral Load, COUNT: 0 | 2022-11-10 | GX4D | 805062 | Unit is due for module replacement/A3 | Module Replacement | Performed Module Replacement | Unit is operational |  | Genexpert Module 6C | G 4920 | Eron Dumagieng | https://www.jotform.com/uploads/macarewecare/222927978372067/5438737894659510444/5438737894659510444_signature_28.png | Eresita Mendoza | https://www.jotform.com/uploads/macarewecare/222927978372067/5438737894659510444/5438737894659510444_signature_31.png
| 2022-11-10 | 2022-11-10 14:20 | Bislig District Hospital | National Highway, Poblacion, Bislig City | 13 | Surigao del Sur | RTDL | NO | INSTALLATION CALIBRATION | CEPHEID | [5 empty] | 22521 | PBSP | gvnovero@yahoo.com | 23 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 912 | 2022-10-21 | GX - 4 | 805280 | Module B2 and B3 found service required from calibration | Calibration | Replacing backup unit Laptop w/ new cpu lenovo thinkcentre with SN: CPH1121F27721 | 3 modules working, 1 module h/w failed. |  | 800-0560GXDX | 36242 | Engr. Faizar S. Avenido | https://www.jotform.com/uploads/macarewecare/222927978372067/5438710389028714911/5438710389028714911_signature_28.png | Georgina V. Novero | https://www.jotform.com/uploads/macarewecare/222927978372067/5438710389028714911/5438710389028714911_signature_31.png
| 2022-11-10 | 2022-11-09 09:38 | SALVACION OPPUS-YNIGUEZ MEMORIAL PROVINCIAL HOSPITAL | Maasain City | 8 | Southern Leyte | Laboratory | YES | CALIBRATION MAINTENANCE Module Replacement | CEPHEID | [5 empty] | 22520 | PBSP | pangilinanyulaida@yahoo.com | 27 | ASSAY: Xpert MTB/RIF, COUNT: 1008 ASSAY: Xpert MTB/RIF Ultra, COUNT: 112 | 2022-08-23 | GXIV - 4D | 804262 | Module B2 defective Module B3 H/W Failed | #NAME? | - Module Replacement Def. module SN: 630185 New module SN : 757944 - Run Xpertcheck on Module B3 | Module B3 for service required |  | 900- 0491RS | GeneXpert 6 color SN : 757944 | JOHN PAUL LABUTAP | https://www.jotform.com/uploads/macarewecare/222927978372067/5438581077913170405/5438581077913170405_signature_28.png | YULAIDA M. PANGILINAN | https://www.jotform.com/uploads/macarewecare/222927978372067/5438581077913170405/5438581077913170405_signature_31.png
| 2022-11-10 | 2022-11-10 10:44 | Carigara District Hospital | Carigara Leyte | 8 | Leyte | TB DOTS | YES | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22519 | PBSP | maervyong22@gmail.com | 30 | ASSAY: Xpert MTB/RIF, COUNT: 766 | 2022-11-10 | GXIV - 4D | 807619 | Module (A3,A4) received multiple temperature error | Check ambient temperature | #NAME? | Module (A3,A4) temperature error | #NAME? | Xpertcheck cartridges | Lot 05203 | John Paul Labutap | https://www.jotform.com/uploads/macarewecare/222927978372067/5438581077919256405/5438581077919256405_signature_28.png | Mae Ervy Ong Ariza | https://www.jotform.com/uploads/macarewecare/222927978372067/5438581077919256405/5438581077919256405_signature_31.png
| 2022-11-10 | 2022-11-10 10:11 | M.icasiano health center | Pedro gil st. Paco | NCR | NCR 1st - Manila | Lab | NO | TROUBLESHOOTING | CEPHEID | [5 empty] | 22518 | DOH | lorrainefabros@yahoo.com | 27.8 | ASSAY: Xpert MTB/RIF, COUNT: 618 ASSAY: Xpert MTB/RIF Ultra, COUNT: 952 |  | Gx4 | 831684 | Monitor issue | Checked monitor,cable and power cord | Checked up | Replaced temporary monitor | For reporting monitor |  |  | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5438561146592500599/5438561146592500599_signature_28.png | Noel Floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5438561146592500599/5438561146592500599_signature_31.png
| 2022-11-10 | 2022-11-10 09:42 | San Sebastian District Hospital | Sablayan, Occidental Mindoro | 4B | Occidental Mindoro | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | [5 empty] | 22517 | PBSP | merj3101996@yahoo.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 1000 ASSAY: Xpert MTB/RIF Ultra, COUNT: 693 | 2022-10-18 | GXIV | 812791 | Unit for PM and for Module Replacement of B1. | Performed PM and Module Replacement | PM and Module Replacement | Unit is operational. | For reporting of module B4. 3 modules operational. | GX Module | G 4863 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5438541842123816275/5438541842123816275_signature_28.png | Merjorie Ann Salazar | https://www.jotform.com/uploads/macarewecare/222927978372067/5438541842123816275/5438541842123816275_signature_31.png
| 2022-11-09 | 2022-11-09 19:03 | ITRMC | San Fernando, La Union | 1 | Ilocos Sur | Laboratory | NO | INSTALLATION Laptop Installation | CEPHEID | [5 empty] | 22516 | PBSP | itrmc.tblaboratory@gmail.com | 23 | ASSAY: Xpert HPV, COUNT: 6 ASSAY: Others, COUNT: 44 ASSAY: Xpert MTB/RIF, COUNT: 621 ASSAY: Xpert MTB/RIF Ultra, COUNT: 385 |  | GX IV | 707524 | Laptop for installation | Installed laptop (Model: Dell, SN: CPH1221F30262 | #NAME? | Unit is operational | NA |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5438034152154286593/5438034152154286593_signature_28.png | Marie Jennifer Culo | https://www.jotform.com/uploads/macarewecare/222927978372067/5438034152154286593/5438034152154286593_signature_31.png
| 2022-11-09 | 2022-11-09 16:17 | Sta. Cruz MHO | Sta. Cruz, Ilocos Sur | 1 | Ilocos Sur | Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22515 | PBSP | gmracuya.pbsp@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 369 ASSAY: Xpert MTB/RIF Ultra, COUNT: 173 |  | GX IV | 808622 | Module B2(660571) for replacement. | Replaced module B2 | - generated pre reports - replaced module B2 by 675724 - generated new reports - performed functionality test | Unit is operational. | NA |  |  | Ronnel Belizar | https://www.jotform.com/uploads/macarewecare/222927978372067/5437915834252762358/5437915834252762358_signature_28.png | Gladys Mae Racuya | https://www.jotform.com/uploads/macarewecare/222927978372067/5437915834252762358/5437915834252762358_signature_31.png
| 2022-11-09 | 2022-11-09 13:55 | Sasa District Health Center | Sasa, Davao City | 11 | Davao del Sur | Laboratory | NO | CALIBRATION MAINTENANCE | CEPHEID | [5 empty] | 22514 | PBSP | decebendavid@gmail.com | 29.6 | ASSAY: Xpert MTB/RIF, COUNT: 10 ASSAY: Xpert MTB/RIF Ultra, COUNT: 229 ASSAY: Xpert SARS-CoV-2, COUNT: 3926 | 2022-11-09 | Gx-IV 4D | 838889 | Module B2 and B4 encounter an Error | Perform Preventive Maintenance | Clean Cartridge Bay, Fan Filter, Plunger Rod Perform xpertcheck on module B4 only | Unit has 3 Functional Modules | Sars Cov2 Assay Stat: From St. Elizabeth Hospital Inc | 1pc xpertcheck cartridge | 0 | Roi Kenneth R. Fagutao | https://www.jotform.com/uploads/macarewecare/222927978372067/5437902616441361878/5437902616441361878_signature_28.png | Deceben V. David, RMT | https://www.jotform.com/uploads/macarewecare/222927978372067/5437902616441361878/5437902616441361878_signature_31.png
| 2022-11-09 | 2022-11-09 14:36 | Lung center of the philippines | Quezon city | NCR | NCR 2nd - Quezon City | Microbiology | NO | Re xpert check | CEPHEID | [5 empty] | 22513 | MACARE-HBDC | lcpmicrolab@gmail.com | 25.6 | ASSAY: Xpert MTB/RIF, COUNT: 124 ASSAY: Xpert MTB/RIF Ultra, COUNT: 451 ASSAY: Xpert SARS-CoV-2, COUNT: 1034 ASSAY: Xpert HIV-1 Viral Load, COUNT: 14 | 2022-11-09 | GX 4 | 838837 | B4 h/failed for rexpert check | Performed re xpert check | Cleaned unit,filter plunger,cartridge bay,icore | 3 modules operation, b4 780860 failed xpert check disabled h/w failed |  | Xpertcheck kits | N/a | Noel floresca | https://www.jotform.com/uploads/macarewecare/222927978372067/5437861826598500563/5437861826598500563_signature_28.png | Lawrence Raquindanum | https://www.jotform.com/uploads/macarewecare/222927978372067/5437861826598500563/5437861826598500563_signature_31.png
| 2022-11-09 | 2022-11-09 12:04 | Tanza Rural Health Unit | Tanza Cavite | 4A | Cavite | Laboratory | NO | MAINTENANCE Module replacement | CEPHEID | [5 empty] | 22512 | PBSP | gutierrezrhian@gmail.com | 25 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 1210 | 2022-11-07 | GXIV | 812835 | Unit for PM and for module replacement of B3 | Performed PM and Module Replacement | PM and Module Replacement | Unit is operational | Module 1 anf 4 waiting for replacement, already reported. 2 modules operational. Def B3 685749, New B3 799469 | GX Module | G 4939 | Shiedrick C. Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5437770793488340810/5437770793488340810_signature_28.png | Mary Anne C. Gutierrez | https://www.jotform.com/uploads/macarewecare/222927978372067/5437770793488340810/5437770793488340810_signature_31.png
| 2022-11-09 | 2022-11-09 10:13 | MMHMC | Batac, Ilocos Norte | 1 | Ilocos Norte | Micro | NO | REPAIR CALIBRATION | CEPHEID | [5 empty] | 22511 | DOH | dennisoncatindig@gmail.com | 25 | ASSAY: Xpert HPV, COUNT: 416 ASSAY: Xpert MTB/RIF, COUNT: 326 ASSAY: Xpert MTB/RIF Ultra, COUNT: 579 ASSAY: Xpert SARS-CoV-2, COUNT: 2815 | 2022-11-07 | GX IV | 831908 | Unit for xpertcheck | Replaced module 4 | Checked and cleaned cartidge bay, plunger lod, icore slot, filter Performed xpertcheck (Module 4 failed XC) Generated reports Replaced module 4 | Unit is operational. | Module 4 failed xpertcheck | GeneXpert Module | 9320 | Ronnel | https://www.jotform.com/uploads/macarewecare/222927978372067/5437730514258073029/5437730514258073029_signature_28.png | DENNISON B. CATINDIG | https://www.jotform.com/uploads/macarewecare/222927978372067/5437730514258073029/5437730514258073029_signature_31.png
| 2022-11-09 | 2022-11-09 09:45 | MMHMC | Batac, Ilocos Norte | 1 | Ilocos Norte | Micro | NO | MAINTENANCE | CEPHEID | [5 empty] | 22510 | PBSP | dennisoncatindig@gmail.com | 25 | ASSAY: Others, COUNT: 108 ASSAY: Xpert MTB/RIF, COUNT: 126 ASSAY: Xpert MTB/RIF Ultra, COUNT: 135 | 2022-11-09 | GX IV | 838168 | Unit for maintenance | Will create a follow up report regarding modules | Checked and cleaned cartridge bay, plunger rod, icore slot, filter Generated reports Performed functionality test | Unit is operational. Awaiting for modules. | B4 Hardware Failure | NA | NA | Ronnel | https://www.jotform.com/uploads/macarewecare/222927978372067/5437682824256321467/5437682824256321467_signature_28.png | DENNISON B. CATINDIG | https://www.jotform.com/uploads/macarewecare/222927978372067/5437682824256321467/5437682824256321467_signature_31.png
| 2022-11-09 | 2022-11-09 08:23 | Lopez RHU | Lopez, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE TROUBLESHOOTING | CEPHEID | [5 empty] | 22509 | DOH | janette_villasanta@yahoo.com | 23 | ASSAY: Xpert MTB/RIF, COUNT: 334 ASSAY: Xpert MTB/RIF Ultra, COUNT: 584 |  | GX-IV-D | 831889 | Temp error on module C3 | Performed XpertCheck on reported module | XpertCheck on module C3 Performed maintenance Generated reports | Two modules operational. | Module C3 failed in XpertCheck. |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5437657827913333542/5437657827913333542_signature_28.png | Janette Villasanta | https://www.jotform.com/uploads/macarewecare/222927978372067/5437657827913333542/5437657827913333542_signature_31.png
| 2022-11-09 | 2022-11-08 15:02 | Calauag Municipal Hospital | Calauag | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22508 | DOH | kimmanlogon953@gmail.com | 28 | ASSAY: Xpert MTB/RIF, COUNT: 293 ASSAY: Xpert MTB/RIF Ultra, COUNT: 386 | 2022-10-04 | GXIV-D | 831910 | For module replacement | Replaced Failed Module | Maintenance Module Replacement | All modules are operational |  | 900-0491R | DR#9249 | Brenguiler Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5437614815515459699/5437614815515459699_signature_28.png | Kimberly Rose J. grimaldo | https://www.jotform.com/uploads/macarewecare/222927978372067/5437614815515459699/5437614815515459699_signature_31.png
| 2022-11-08 | 2022-11-08 15:59 | Southern Philippines Medical Center | Bajada, Davao City | 11 | Davao del Sur | TB-HIV Laboratory | NO | REPAIR | CEPHEID | [5 empty] | 22507 | PBSP | spmcrhivda@gmail.com | 28.9 | ASSAY: Xpert MTB/RIF, COUNT: 264 ASSAY: Xpert MTB/RIF Ultra, COUNT: 1309 ASSAY: Xpert HIV-1 Viral Load, COUNT: 291 | 2022-11-08 | GENEXPERT | 808298 | Module B4 subjected for Module replacement | Replace Defective Module with new module allocation | Replace defective module with serial no. 692011 with new module serieal no. 804865 | All Modules are fully operational |  | 900-0491RS | 0 | ROI KENNETH R. FAGUTAO | https://www.jotform.com/uploads/macarewecare/222927978372067/5437044303125008442/5437044303125008442_signature_28.png | Flordeliza B. Go, RMT, MPH | https://www.jotform.com/uploads/macarewecare/222927978372067/5437044303125008442/5437044303125008442_signature_31.png
| 2022-11-08 | 2022-11-08 12:24 | Bondoc Peninsula District Hospital | Catanauan, Quezon | 4A | Quezon | Laboratory | NO | MAINTENANCE Pull out Defective laptop | CEPHEID | [5 empty] | 22506 | PBSP | trinity_carisse136@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 17 |  | GX-IV-L | 805339 | For laptop replacement | Replaced laptop | Performed maintenance Generated reports Pull out Defective laptop Checked modules B1,B3,B4 | Unit is operational. Module B1 and B3 H/W failed. Module B4 under observation |  |  |  | John Monsour Doctor | https://www.jotform.com/uploads/macarewecare/222927978372067/5436912069906939669/5436912069906939669_signature_28.png | Trinity Carisse Puyal | https://www.jotform.com/uploads/macarewecare/222927978372067/5436912069906939669/5436912069906939669_signature_31.png
| 2022-11-08 | 2022-11-08 10:44 | BJMP Infirmary 4a | Turbina, Calamba Laguna | 4A | Laguna | Laboratory | NO | MAINTENANCE Module Replacement | CEPHEID | [5 empty] | 22505 | PBSP | jhonzelborja@gmail.com | 25 | ASSAY: Xpert MTB/RIF, COUNT: 98 ASSAY: Xpert MTB/RIF Ultra, COUNT: 965 | 2022-11-07 | GXIV | 814156 | For module replacement of Module 3 | Performed Module Replacement and PM | Module Replacement and PM | Unit is Operational | Module 1 and 4 pending for replacement. | GeneXpert module | G 4924 | Shiedrick Daroya | https://www.jotform.com/uploads/macarewecare/222927978372067/5436850908202640889/5436850908202640889_signature_28.png | Justine Jhonzel Borja | https://www.jotform.com/uploads/macarewecare/222927978372067/5436850908202640889/5436850908202640889_signature_31.png
| 2022-11-08 | 2022-11-08 10:36 | Macalelon Rural Health Unit | Macalelon | 4A | Quezon | Laboratory | NO | REPAIR MAINTENANCE | CEPHEID | [5 empty] | 22504 | DOH | cheyster7388@gmail.com | 28 | ASSAY: Xpert MTB/RIF Ultra, COUNT: 308 ASSAY: Xpert MTB/RIF, COUNT: 100 | 2022-10-04 | GXIV-D | 831636 | For Module Replacement | Replaced Failed Module | Module Replacement Preventive Maintenance | All Modules are operational |  | Module 6-C | DR 9245 | Brenguiler I. Bagbaguen | https://www.jotform.com/uploads/macarewecare/222927978372067/5436846202351078547/5436846202351078547_signature_28.png | Cheyster O. Agustin | https://www.jotform.com/uploads/macarewecare/222927978372067/5436846202351078547/5436846202351078547_signature_31.png